## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities
Z-0490-2010 | 53845 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | LACTOSORB SYSTEM 1.5MM ADJUSTABLE SELF-DRILLING TAP FOR POWER DRIVER, REF: 915-2185, NON-STERILE PRODUCT, |  |  | September 15, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling LactoSorb Self Drilling Adjustable Tap Model Numbers 915-2075 Lot # 064350 and 915-2185 Lot # 885580. The two lots identified have a new style stop and an old style tap; they are not compatible, as the depth etchings do not line up correctly. Depending on the tap (1.5mm VS 2.0mm) you may drill too shallow or too deep. | Component change control | Biomet Microfixation notified consignees of the recalled product beginning September 2009 via telephone and e-mail correspondence. Users were asked to return the affected product to the firm.For further information, contact Biomet Microfixation at 1-904-741-9468. |  | 7 | 7 pieces | USA - FL, MO, and KS; Foreign: Brazil and China. | Terminated | January 23, 2010 | 130 | 22
Z-0589-2010 | 53843 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Maquet Servo Ventilator SV900C/D/E (SV900) |  |  | November 17, 2009 | 2009 | MAQUET Inc. | Some Servo Ventilators 300/300A and Servo Ventilator 900C/D/E should not be used with a system that may generate negative pressure below -100 cm H2O (Closed system suctioning) due to the risk of damaging the ventilator's pressure transducers. | Device Design | Maquet sent Device Correction letters commencing 11/17/09 and ending 11/23/09 by Federal Express. |  | 29364 | 29364 (8438 of those in US) | Worldwide distribution | Terminated | June 04, 2012 | 930 | 50201
Z-0590-2010 | 52207 | No | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Hocoma Lokomat-System with Pediatric Module |  |  | May 27, 2009 | 2009 | Hocoma AG | Changes to the stored "sensitivity" values may occur and leads to higher limits in the force and trajectory deviation | Software design | Hocoma issued a letter dated May 27, 2009 to Lokomat users informing them of the affected device. Instructions were provided to avoid further problems with the product and users were asked to complete and return the provided Declaration to Field Safety Notice.For further information, contact Hocoma at info@hocoma.com. |  | 7 | 7 | Nationwide Distribution -- including IL, MA, MI, NY and OH. | Terminated | February 22, 2010 | 271 | 7
Z-0504-2010 | 53386 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Accolade System Surgical Technique; |  |  | August 31, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Certain surgical protocols do not provide specific instructions for initiating power to the calcar planer prior to engagement with bone. | Error in labeling | "Urgent Product Correction - Surgical Protocol" letters and Product Accountability Forms were sent to consignees via Fed Ex on August 31, 2009 and September 1, 2009 with return receipt. Notification letters described the issue, protocols affected, potential hazard and harm and risk mitigation. Direct questions about the recall by calling 1-201-972-2100 or Customer Service 1 -866-OR-ASSIST (1 -866-672-7747). |  | 5670 | 5,670 | Worldwide Distribution. | Terminated | July 17, 2013 | 1416 | 14396
Z-0592-2010 | 53957 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Signa Excite 1.5T, Signa HD 1.5T, Signa HDx 1.5T, Signa HDxT 1.5T, Signa HDe, GE DISCOVERY MR450 and the Signa 3.0T MR750 MRI systems when used in conjunction with Confirma's CADstream (These systems contain the PURE software option.) |  |  | October 06, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware that the use of Confirma, Inc.'s CADstream in conjunction with the GE Healthcare's PURE (Phased Array Uniformity Enhancement) software for dynamic MRI imaging (Signa Excite 1.5T, Signa HD 1.5T, Signa HDx 1.5T, Signa HDxT 1.5T, Signa HDe, GE DISCOVERY MR450 and the Signa 3.0T MR750 MRI systems ) may require a modification to the study preference settings to p | Software design | GE Healthcare issues an "Important Product Information" letter dated October 23, 2009. The letter was addressed to Hospital Administrators, Managers of Radiology and Radiologists. The letter listed the Issue, Affected Product Details, Instructions, Product correction and Contact Information.For further information, contact GE Healthcare at 1-262-548-2731. |  | 5460 | 5460 | Worldwide Distribution -- All States in the US including DC and Puerto Rico. OUS: ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BELGIUM, BOLIVIA, BOSNIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, ISLAMIC IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA REPUBLIC, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB, LITHUANIA, LUXEMBOURG, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB, UNITED KINGDOM, UZBEKISTAN, VENEZUELA and VIETNAM. | Terminated | August 03, 2012 | 1032 | 5460
Z-0594-2010 | 53983 | Yes | OAD |  |  | Cardiovascular | PMA | Powerlink System with IntuiTrak Delivery System (Endovascular Graft): Model Number 34-34-80LE, LOTS W09-2967 AND W09-3000. |  |  | November 25, 2009 | 2009 | Endologix Inc | This action is being taken because the product has the potential for the polyimide tubing to detach from the inner core due to the bond failure. If this occurs, treatment may be delayed or interrupted. The firm's initial health hazard analysis indicated the event may cause temporary or medically reversible adverse health consequences or an outcome where the probability of serious adverse health | Process design | The recall communication was initiated on 11/25/09 by phone call and followed by recall letter via Fed-Ex Overnight. The recall letter informed the consignees of the reason for recall and products affected. Users were instructed to share this information with physicians that performed the affected procedures at their facility. Endologix informed their consignees that their personnel would retrieve the affected devices and return it to Endologix at no cost and would replace the affected units. Consignees with questions or comments about the notice are directed to (800) 983-2284. |  | 28 | 14 UNITS DISTRIBUTED: 3 LOT W09-2967 AND 1 LOT W09-3000 STILL AT CONSIGNEES. 10 UNITS USED. | Recalled devices were distributed to medical centers and hospitals in OR, PA, FL, TX, MI and LA. | Terminated | May 07, 2010 | 163 | 34
Z-0591-2010 | 53949 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Bausch &amp; Lomb SoFlex SE Foldable Intraocular Lens. |  |  | November 19, 2009 | 2009 | Bausch &amp; Lomb Inc | Bausch &amp; Lomb Incorporated, Rochester, NY is recalling their Silicone L161SE Intraocular Lens Lot 4916928 (23.5D) after receiving three customer complaints for what appears to be opacification of the intraocular lens. The opacification occurred within one week of implantation. In each reported case the surgeon explanted the lens. | Other | Bausch &amp; Lomb, Inc. notified consignees via phone beginning November 19, 2009 and by follow-up letter dated November 23, 2009. Titled "Urgent - Medical Device Recall", the letter instructed users to discontinue implantation of the affected product and return any unused lenses along with the enclosed acknowledgement form to the firm.For further information, contact Bausch &amp; Lomb at 1-585-338-6612. |  | 85 | 85 lenses | Lenses were distributed to one distributor and 35 retailers. Worldwide Distribution -- Distributed to Canada, Taiwan and throughout the United States. | Terminated | July 22, 2010 | 245 | 85
Z-0595-2010 | 53982 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker MIS Bur, 2.0mm Neuro Diamond, sterile, Stryker Instruments, Kalamazoo, MI; REF 8420-107-120. |  |  | July 06, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | The bur may fracture, resulting in fragments within the surgical site, injury to the patient or surgical team, or both. | Device Design | Consignees were sent a letter dated 7/10/09 instructing them to locate and return product on hand for replacement. Surgeons who use the product were notified via a separate, untitled letter, dated July 13, 2009. A speed warning will be added to the product labels. Questions are directed to the firm at: 1-800-800-4236 ext. 3808. |  | 5438 | 5438 of all products | Nationwide distribution: Australia, Canada, England, France, Hong Kong, Japan, Netherlands, South Africa, Sweden, Switzerland and Taiwan. | Terminated | October 25, 2010 | 476 | 5438
Z-0605-2010 | 53781 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AMO Tecnis 1-Piece Intraocular Lens (Model ZCB00) |  |  | November 11, 2009 | 2009 | Abbott Medical Optics Inc (AMO) | Abbott Medical Optics (AMO) has received complaints of some Tecnis 1-Piece IOLs (model ZCB00) have adhered to the lens optic after insertion into the eye. This can lead to a need for additional manipulation to separate the haptics from the optic during surgery which could lead to an injury. | Process design | Beginning on Friday, November 6, 2009, AMO began distribution of the Recall Notification Letter via Federal Express to the 874 AMO Tecnis 1-piece IOL customer accounts in the US. Included with the Recall Notification letter is a facsimile delivery confirmation that each customer was instructed to fax back to AMO to confirm receipt of the Recall Notification Letter. --- Beginning on Monday, November 9, 2009, AMO began distribution of the Recall Notification letter to the AMO Tecnis 1-piece IOL customer/distributors accounts outside the USA."THIS RECALL NOTIFICATION DOES NOT INVOLVE TECNIS 1-PIECE IOL THAT HAVE BEEN IMPLANTED. ONCE THE HAPTICS HAVE BEEN SUCCESSFULLY DEPLOYED, THESE LENSES FUNCTION PROPERLY."Questions should be directed to an AMO Customer Service Representative at 1-877-AMO-4LIFE (1-877-266-4543). |  | 246000 | 246000 Lenses | AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY, Hawaii, and Puerto Rico.INTERNATIONAL: Austria, Azerbaijan, Belgium, Canada, Switzerland, Cyprus, Czech Republic, Germany, Estonia, Spain, Finland, France, United Kingdom, Hungary, Ireland, Israel, Italy, Lebanon, Lichtenstein, Libya, Martinique, Netherlands, Poland, Palestine, Portugal, Saudi Arabic, Sweden, Slovenia, Turkey, and South Africa. | Terminated | July 25, 2012 | 987 | 246000
Z-0607-2010 | 54031 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Medtronic Sofamor Danek Extended Length Trocar, Size : L=50MM, Item No.: 7578014, Mat'l: Stainless Steel, non-sterile, Rx only, Trocars used with the CD HORIZONÂż LEGACY" Percutaneous Peek Spinal System |  |  | December 07, 2009 | 2009 | Medtronic Sofamor Danek USA Inc | The Trocar may not fully engage with the Rod Inserter or may prematurely disengage. | Device Design | On 12/7/2009, the Sales Representatives were notified by conference call and email regarding the recall. The recall notices were mailed directly to Risk Managers and Sales Representatives, with Recall Questionnaires and Confirmation forms included respectively, on 12/09/2009 with a delivery date of 12/10/09. |  | 266 | 266 | Nationwide distribution, and the Netherlands | Terminated | January 14, 2011 | 403 | 465
Z-0616-2010 | 53788 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Omega IV and Omega V tables on GE Advantx-E imaging system, used for general purpose diagnostic angiographic fluoroscopy and radiographic studies. |  |  | December 22, 2009 | 2009 | GE Medical Systems, LLC | Potential table rotation of Omega IV and Omega V tables affecting Advantx-E and Innova 2000, 2100IQ, 3100, 3100IQ, 4100, 4100IQ, 2121IQ and 3131IQ cardiovascular X-ray imaging systems.Due to potential weakness of some table rotational brakes, an unintended table rotation may occur while the table is locked. This issue could lead to a patient fall if the unintended motion occurs during patient | Device Design | Consignees were sent on 12/22/09 a GE Healthcare " Urgent Medical Device Correction" letter dated 12/1/09. The letter was addressed to Hospital Administrators / Risk Mangers, Mangers of Radiology / Cardiology, Radiologists / Cardiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. |  | 1502 | 1502 | Worldwide distribution: USA, SRI LANKA, ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, BELARUS, BELGIUM, BOLIVIA, BOSNIA &amp; Herzegovina, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, GERMANY, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and YUGOSLAVIA. | Terminated | November 13, 2015 | 2152 | 3635
Z-0611-2010 | 54114 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet 2.9 Micromax L-15 suture anchor, flexible long shaft, #2 blue/white maxbraid, PLLA/PGA L-15, sterile, Biomet Sports Medicine, Warsaw, IN; REF 905456. |  |  | December 11, 2009 | 2009 | Biomet, Inc. | The anchor label states it contains a flexible shaft, but this lot actually contians a rigid shaft. | Labeling mix-ups | Biomet Sports Medicine notified the single consignee by letter dated December 11, 2009. The firm requested that users locate and return the affected product and return a business reply card.For further information, contact Biomet at 1-800-348-9500 or 1-574-372-3983. |  | 6 | 6 | Nationwide Distribution -- Alabama. | Terminated | October 18, 2010 | 311 | 6
Z-0623-2010 | 53953 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Cordis Aquatrack Hydrophilic Nitinol Guidewire, for use in angiographic procedures, REF C3526-SSA, Guidewire Diameter .035" Tip Length Flexible 3 cm, Sterile. |  |  | August 25, 2009 | 2009 | Cordis Corporation | Some of the guidewires in these two lots have a straight tip (incorrect) rather than an angled tip (correct configuration for that catalog number). | Process control | Recall for Lot 80000521 was initiated on August 25, 2009 with Cordis Medical Device Recall letters being mailed for delivery on August 25, 2009. Response form was also provided to obtain product disposition.Recall for Lot 80000584 was initiated on October 19, 2009 with Cordis Medical Device Recall letters being mailed for delivery on October 20, 2009. Response form was also provided to obtain product disposition. |  | 145 | 145 units | Nationwide distribution. | Terminated | August 11, 2010 | 351 | 145
Z-0631-2010 | 53980 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | GE ApexPro Oximeter systems (ApexPro CH or ApexPro FH with ApexPro Telemetry version 3.9 software) |  |  | September 25, 2009 | 2009 | General Electric Medical Systems Information Technology | Potential misuse of the ApexPro Telemetry System and incorrect message and alarm information in the ApexPro Operator's Manual may impact patient safety. When ApexPro Telemetry systems are used with SpO2 oximeters:1. ApexPro FH Transceiver can operate with the ApexPro FH Xpod connected, but without ECG lead wires attached to the transceiver. This is unintended use of the device that is not do | Labeling design | Consignees were sent on 10/27/09 a GE Healthcare "Urgent Medical Device Correction" letter dated September 25, 2009. The letter was addressed to Risk Manger, Chief of Nursing, director of Biomedical Engineering. The letter described the Safety Issues, Affected Product Details, Product Correction, Safety Instructions and Contact Information. "ApexPro Telemetry System Operator's Manual Supplement for SpO2 was enclosed with the letter. |  | 8305 | 8305 | Worldwide distribution: USA, AUSTRALIA, AUSTRIA, BELGIUM, BERMUDA, CANADA, CHILE, CHINA, COSTA RICA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUAM, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, MALAYSIA, MALTA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PHILIPPINES, POLAND, QATAR, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED KINGDOM, and VENEZUELA. | Terminated | May 24, 2012 | 972 | 8305
Z-0622-2010 | 53981 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Foundation Tibia Baseplate, Catalog number 324-01-108. The tibial is modular in design, such that plastic inserts of various thicknesses and intramedullary stem are attached to the metal baseplate intraoperatively. |  |  | December 03, 2009 | 2009 | Encore Medical, Lp | Product does not meet specifications for a size 8 Right Foundation Tibia Baseplate. | Mixed-up of materials/components | Firm notified consignees by an Urgent Field Safety Notice dated 12/3/2009. The letter identified the product, stated the problem, and the risk associated with the problem. Customers are to immediately return the affected product and notify all of their customers if product was further distributed. Customers are to complete and return the enclosed response form. Questions should be directed to David Harris, Manager of Customer Service at 512-834-6330 or Valerie Moulton, International Customer Service at 512-834-6275. |  | 17 | 17 | Worldwide Distribution -- USA, including states of GA, OK, FL, CA, and SD, and countries of Germany and Turkey. | Terminated | May 05, 2010 | 153 | 17
Z-0609-2010 | 53994 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific COGNIS CRT-D, Cardiac Resynchronization Therapy Defibrillator, models N118, N119. the following models are not available in the US: models N106, N107, P106, P107, P108. Sterilized using ethylene oxide. Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. |  |  | December 01, 2009 | 2009 | Boston Scientific CRM Corp | BSC CRM provided important patient management information regarding Boston Scientific COGNIS CRT-Ds and TELIGEN ICDs. Engineering simulations and field reports indicate that in some cases, implanted devices may encounter sufficient mechanical stress to weaken the bond between the header and case when positioned subpectorally. A weakened header bond may alter lead impedance and introduce noise | Device Design | Boston Scientific CRM issued an "Urgent Medical Device Correction" letter dated December 01, 2009. The letter described the issue and potential device behavior. The letter provided recommendations to avoid issues relating to the affected device until improved devices become available.For further information, contact your local Boston Scientific CRM representative, United States Technical Services at 1-800-227-3422, or European Technical Services at +32-2-416-7222. |  | 39873 | 39,873 | Worldwide Distribution -- United States including Guam, Puerto Rico, and U.S. - Virgin Islands. OUS to include: Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Belgium, Bermuda, Brazil, Canada, Cayman Islands, Chile, China, Cocos (Keeling) Islands, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Finland, Gabon, Germany, Greece, Hong Kong, Hungary, India, Iran, Ireland, Israel, Jordan, Kuwait, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Trinidad &amp; Tobago, Tunisia, United Arab Emirates, United Kingdom and Venezuela. | Terminated | November 06, 2012 | 1071 | 100084
Z-0637-2010 | 53890 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gyrus ACMI PKS Seal Open Forceps, model 915010PK, 9-3/4" Angled Jaw. Sterile EO. The device is intended to be used with the G400 electrosurgical generator. Gyrus ACMI, Inc., 136 Turnpike Road, Southborough, MA 01772-2104 USA. |  |  | August 18, 2009 | 2009 | Gyrus ACMI Corporation | Gyrus ACMI received complaints of a problem with 3103PK, 3104PK, 3105PK, 915000PK, 915005PK, 915010PK - Gyrus ACMI PKS Seal Open Forceps. The metal shim between the jaws may detach during the procedure and fall into the surgical field. The detached shim may be undetected and be left behind in the patient. This circumstance could subsequently cause complications. Further use of this product sh | Environmental control | An Urgent Medical Device Recall letter, dated 08/15/09 was sent to customers notifying them of the issue. The letter identified affected product, asked that they cease further use, remove affected devices from stockrooms, contact Gyrus ACMI customer service for a return goods authorization and return product to Gyrus ACMI at no charge. A reply form is also asked to be returned to Gyrus ACMI via fax.Questions and requests for a Return Goods Authorization are directed to Gyrus ACMI customer service at 1-888-524-7266. |  | 93665 | 93,665 total (63,418 US, 30,247 OUS) | US and International distribution. OUS to Include: Australia, Belgium, Bolivia, Bulgaria, Canada, Chile, China, Columbia, Croatia, Czech, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece , Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Panama, Puerto Rico, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, United Kingdom. | Terminated | March 06, 2012 | 931 | 93665
Z-0642-2010 | 54020 | No | LYJ |  |  | Neurology | PMA | VNS Therapy Demipulse Generator, Model 103 and VNS Therapy Demipulse Duo Generator, Model 104. |  |  | November 20, 2009 | 2009 | Cyberonics, Inc | Under certain conditions, product's battery life can be reduced. | Other | Firm sent "An Important Safety Notice" letter to all physicians explaining events, risks presented and recommended actions. Letter was mailed November 20, 2009. Physicians should complete and return the effectiveness card as soon as possible to acknowledge that you have read and understood the Safety Alert Letter. If further information is needed, customers in the US can contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com. |  | 8667 | 8,667 units | Worldwide Distribution -- USA, Argentina, Austria, Belgium, Canada, Columbia, Costa Rica, Croatia, Czech Republic, Denmark, Finland, France, Greece, Hungary, Ireland, Iceland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Netherlands, Norway, Panama, Portugal, Republic of Germany, Saudi Arabia, Sweden, Slovakia, Spain, Switzerland, Taiwan, UAE, United Kingdom, and Venezuela. | Terminated | July 24, 2012 | 977 | 8667
Z-0630-2010 | 54000 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sorin / Clearglide Precision Bipolar, REF PBD01, Sterile R, Manufactured for: Sorin Group USA, Inc., 14401 W.65th Way, Arvada, CO 80004. |  |  | November 25, 2009 | 2009 | Sorin Group USA, Inc. | Vessel clamping device may break during use. | Component change control | Consignees were notified by "Urgent Field Safety Notice" dated November 24, 2009 with instructions to examine the device before, during and after use. Customers were instructed to not use the device, if it is damaged. Customers were also instructed to contact the Sorin Group Customer Service at 1-800-650-2623 to arrange shipment of replacement product and requested to complete and return the response form. Direct questions to the Sorin Group Customer Service at 1-800-650-2623. |  | 756 | 756 units | Worldwide Distribution -- United States, Australia, Canada and Italy. | Terminated | October 04, 2010 | 313 | 5817
Z-0643-2010 | 53755 | No | KZI | Enriched culture medium. | Microbiology | Microbiology | 510(K) Exempt | BBL (TM) Brain Heart Infusion, Catalog number 221778, package of 100 tubes, 0.5 mL tubes. Product is labeled in part *** Becton Dickinson and Company, 7 Loveton Circle, Sparks, Maryland 21152 USA*** 800-638-8663 |  |  | February 19, 2009 | 2009 | Becton Dickinson &amp; Co. | In vitro diagnostic test reagent may contain residual material that inhibits the growth of gram positive/gram negative organisms. | Equipment maintenance | The recalling firm notified direct accounts by an Urgent Product Recall letter on 02/19/09. The notification identified the affected product and stated the problem. The letter also advised that users discontinue use of the product and to discard any remaining packages. The attached response form should be completed and returned for replacement of product. If further assistance is needed regarding replacements, customers can contact BD Customer Service Department at 1-800-675-0908. |  | 144 | 144 boxes | Nationwide Distribution -- Including states of CA, CO, FL, GA, IL, MI, MN, MT, NY, OH, PA, TN, TX, UT, and VA. | Terminated | February 12, 2010 | 358 | 144
Z-0646-2010 | 53945 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Vanguard Lock-On Femoral Impactor, Biomet Orthopedics. Warsaw. IN; REF 32-486206. |  |  | December 01, 2009 | 2009 | Biomet, Inc. | The locking hook tabs may fracture and fragments remain in the patient post-operatively. | Device Design | Urgent Medical Device Recall Notices were sent to distributors on 12/01/2009 with a notice for them to provide to medical facilities where distributor-owned inventory is stored. Distributors were instructed to immediately locate and remove the instruments from circulation; follow the instructions on the "FAX Back Response Form" and fax a copy back prior to returning the product; and to return the product to the firm. Questions should be directed to 800-348-9500. |  | 92 | 92 | Worldwide Distribution -- USA, Canada, China, Costa Rica, Denmark, Germany, Netherlands, New Zealand, and United Kingdom. | Terminated | October 18, 2010 | 321 | 92
Z-0652-2010 | 53960 | No | BZA | Airway connector. | Anesthesiology | Anesthesiology | 510(K) Exempt | Hudson RCI One Way Valve, 22mm E.D. to 22mm O.D., REF 1665, Rx Only, Teleflex Medical, Research Triangle park, NC 27709 |  |  | December 02, 2009 | 2009 | Telefelx Medical | Products 1664 and 1665 are each one-way valves that offer low flow resistance, each in one, opposite direction. Each of these valves demonstrate the appropriate direction with a red arrow.The reason for this recall is incorrect packaging; product code 1664 was incorrectly packaged as product code 1665.If the red arrows are not followed, the valve may be placed in the circuit in the wrong | Packaging | Consignees were notified by letter via next day delivery. Each consignee was sent a letter requesting acknowledgement of the recall and a request to return all of the affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. A second mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the second mailing or as an alternative, they may be contacted by phone, e-mail, or fax. A record of these contacts will be maintained. Affected products received by Teleflex will be destroyed or reworked and replacement product or credit will be provided to the consignees. Records of destruction or rework will be maintained.Questions are directed to the firm's sales representatives or to Teleflex Customer Service at 1-866-246-6990 |  | 2450 | 2450 eaches | CA, IL, NE, NJ, TX, VA, WI and Canada. | Terminated | August 11, 2010 | 252 | 2450
Z-0644-2010 | 54042 | No | JTZ | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | remel Mueller Hinton Agar w/4% NaCl w/6 mcg/ml Oxacillin, Ref R01626, 10/box. The firm name on the label is Remel, Lenexa, KS. |  |  | December 08, 2009 | 2009 | Remel, Inc | The product may fail to adequately grow methicillin resistant Staphylococcus aureus ATCC 43300 and Staphylococcus aureus ATCC 33591. | Nonconforming Material/Component | The recalling firm issued two different recall letters dated December 8, 2009 via regular mail. One letter was addressed to Microbiology Laboratory Supervisor and the other to Distributors. Both letters explained the reason for recall, requested the customer discard all remaining units in their possession, and complete the enclosed Product Inventory Checklist to indicate the amount of product that was discarded. The letter issued to distributors also requests they notify their customers to the laboratory level and the Product Inventory Checklist has a block to be marked indicating they have notified their customers. Direct questions to the Technical Services Department at 1-800-255-6730 (United States) or 1-913-888-0939 (International). |  | 267 | 267 units | Nationwide. | Terminated | July 12, 2011 | 581 | 267
Z-0641-2010 | 53968 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Orthofix Spinal Implants; Pillar SA PEEK Spacer System Instrument Case. |  |  | November 19, 2009 | 2009 | Orthofix, Inc | The Pillar SA Implant Insertion Instrument 49-1100 may experience resistance in turning the activation knob while releasing an implant after intra-operative positioning. | Nonconforming Material/Component | A notification letter dated November 19, 2009 was sent to customers via Federal Express. The letter described the recalled product (PILLAR SA Implant Insertion Instrument, 49-1100) and problem and actions for customers to take. The affected product, PILLAR SA lmplant lnsertion lnstrument, 49-1100 is in the PILLAR SA set (set part number 49-0021) which was shipped to customers. An enclosed attachment provide instructions on product removal and return. Customers were instructed to complete the enclosed Tracking and Verification Form and fax to Orthofix Spinal Implants at 1-973-406-2912. A replacement instrument will be provided as available. In the interim, the Implant Inserter (Part Number 49-1000) included in the PILLAR SA UPE set (set number 49-0020) that was provided earlier this year may be used as an alternative to the affected product. For questions regarding the removal and return of the recalled product to Orthofix Spinal Implants, please contact your Account Services Representative.Direct other questions to the Regulatory Affairs Department at 1-973-406-2812. Direct product questions to Marketing at 1-862-221-1335 or Product Development at 1-973-406-2824. |  | 29 | 29 instruments | United States (IA, NV, CO, AZ, NY, MD, IL, VA, MN, MI, NC, CT, LA, AR, GA, CA, MO, WA, OK and AL). | Terminated | February 02, 2010 | 75 | 29
Z-0651-2010 | 53940 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical Medfusion model 3010a syringe Infusion pump. REF 3010A. Rx Only. Smiths Medical MD, Inc., St. Paul, Minnesota 55112 (USA) |  |  | November 16, 2009 | 2009 | Smiths Medical ASD, Inc. | Smiths Medical has become aware of an increased trend in reports of Motor Not Running (MNR) and Motor Rate Error (MRE) alarm message events with certain MedfusionÂż Syringe Infusion Pumps, Models 3010A and 3500 (Pump). MNR/ MRE are high priority alarm messages which can occur for a variety of reasons as part of the usual self-test functionality of the Pump. The Pump is designed so that when its so | Device Design | An Urgent Medical Device Recall letter, dated November 16, 2009, was sent to customers beginning November 16, 2009. The letter described the issue, identified affected devices, and asked that affected devices be returned to Smiths Medical to be inspected, repaired, if necessary, and returned back to medical facility. A Confirmation Form is also asked to be completed and returned to Smiths Medical.To return Pumps or for questions regarding this Urgent Recall Notice, please contact Smiths Medical Customer Service Department at 1~866-270-7060.Please report any issues with these Pumps please contact Smiths Medical Technical Customer Service Department at 1866-216-8806. |  | 43 | 43 | AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY,OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV,including Guam, Puerto Rico. OUS to include: Canada, Chile, Czech Republic,, England, France, Germany, Hong Kong, Ireland, Israel, Lebanon, Malaysia, Poland, Saudi Arabia, Scotland, Singapore, Sweden, Switzerland, United Arab Emirates. | Terminated | December 18, 2011 | 762 | 5436
Z-0647-2010 | 53779 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Model 91367 SL 2200 Compact Monitor |  |  | November 06, 2009 | 2009 | Spacelabs Healthcare, Incorporated | Potential for monitor to shut down and not being able to power on again. | Component design/selection | On 11/6/09 the firm sent Urgent - Device Correction notifications to customers by certified mail informing them of the monitor failure and that a field service engineer will contact the customer to fix the affected device at no cost. On 11/13/2009 the firm's Global Tech Support sent notification via email to all foreign subsidiaries and distributors informing them of the monitor failure and that a field service engineer will contact the customer to fix the affected device at no cost. Questions are to be directed to Spacelabs Medical's Technical Support at 1-800-522-7025, and 1-425-657-7200 x5089 for those customers outside of the US. |  | 422 | 422 units | Worldwide Distribution | Terminated | July 21, 2011 | 622 | 3413
Z-0656-2010 | 54036 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Tornier Arthro Tunneler. For transossous fixation of tendons to bone in rotator cuff repair. |  |  | December 04, 2009 | 2009 | TAG Medical Corp. | Device failed to deploy the distal loop | Process change control | Medicsense, the US Agent notified the sole distributor Tornier Inc., Edina, MN of the recall and requested notification to users requesting to discontinue and return product. Letter to customers dated December 4, 2009. |  | 60 | 60 UNITS | Nationwide | Terminated | February 29, 2016 | 2278 | 60
Z-0655-2010 | 53199 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Complete SE Biliary Stent System, SC680L (6 mm x 120 mm), lot number 0000983806, UPC code 006133994186669, use by date 03/17/2011, manufactured by Medtronic CardioVascular, Santa Rosa, CA. |  |  | July 17, 2009 | 2009 | Medtronic CardioVascular | Mislabeling-- A lot of 5 mm x 12 mm stents were labeled as 6 mm x 80 mm stents and vice versa. | Labeling mix-ups | Medtronic, Inc. issued an "Urgent - Voluntary Medical Device Recall" communication dated July 17, 2009 asking consignees to segregate labeled lot numbers until a representative from the firm arrived to retrieve the products.For further information, contact your local Medtronic representative or call Medtronic, Inc. at 1-707-525-0111. |  | 20 | 20 units | Product was distributed to 24 consignees in TX, MO, PA, MN, NY, IL, CA, IN, AZ, TN, OH, MI, and CT. | Terminated | June 21, 2010 | 339 | 20
Z-0004-2010 | 53491 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC 9900 Elite ESP |  |  | August 12, 2009 | 2009 | Ge Healthcare | A software defect - may result in a false indication and warning on the OEC 9900 Workstation and C-arm display that the X-Ray tube Anode is over-heated. This will also cause the displayed air kerma rate and cumulative air kerma to be inaccurate. | Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter will be sent to customers identifying the issues, interim solution, and permanent solution. Questions or concerns are to be directed to the GEHC OEC service team at 1-800-874-7378. |  | 1411 | 1411 units total | Nationwide and Canadian Distribution | Terminated | April 23, 2012 | 985 | 1411
Z-0577-2010 | 54019 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Ellex Motorised Safety Filter, Zeiss; Model number 6339585. |  |  | February 11, 2009 | 2009 | Ellex Medical Pty Ltd | Potential for two screws important in the design of the safety filter to come loose. | Radiation Control for Health and Safety Act | Ellex issued a bulletin stating the problem and identifying the affected products. Customers are to remove the Motorised Safety Filter from clinical use if the green laser light is visible while viewing through the microscope at any time during use of the Solitaire laser. The Bulletin also states that customers should contact an Ellex approved Service Engineer to investigate and perform the Field Corrective Action listed. The manufacturer will remedy the defect or bring the product into compliance free of charge. Ellex will inspect and rework all affected safety filters sold prior to February 2009. |  | 200 | 200 units total | Worldwide Distribution. | Terminated | September 07, 2012 | 1304 | 200
Z-0645-2010 | 53164 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare MAC 800 RESTING ECG ANALYSIS SYSTEM. |  |  | December 10, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has become aware of an issue with MAC800 laser printouts. When more than one ECG report in the file manager is batch printed in laser print, the potential exists for an incorrect patient identification or name to be displayed on the print out. Misidentification of patient data potentially contributing to delayed or incorrect treatment is then possible. | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated December 10, 2009. The letter was addressed to Healthcare Administrator/Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.GE Healthcare will provide a software upgrade to resolve the issue. A GE Healthcare representative will contact customers to schedule implementation of the correction when it is available for installation at the customer facility.Direct questions to the GE Healthcare Technologies Support Center at 1-800-558-7044, enter Option 2 and ask for Technical Support or your local GE Field Service Representative. |  | 567 | 567 | Worldwide Distribution -- United States (CA, CO, CT, FL, IL,MD, NY, NC, OH, TN, TX, WA and WI), UNITED KINGDOM, UNITED ARAB EMIRATES, THAILAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVAKIA, SINGAPORE, ROMANIA, QATAR, PORTUGAL NETHERLANDS, LITHUANIA, KOREA, JAPAN, ITALY, ISRAEL, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, CZECH REPUBLIC, CROATIA, CHINA, BAHRAIN, AUSTRIA and AUSTRALIA. | Terminated | August 06, 2013 | 1335 | 567
Z-0336-2010 | 53704 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Imager" II Angiographic Catheters, Sterilized with ethylene oxide gas, Manufactured for Boston Scientific Corporation, One Boston scientific Place, Natick, MA 01760-1537, USA, Made in Ireland, Unit 7,8 &amp; 9 Annacotty Business Park, Annocotty, Co. Limerick, Ireland: |  |  | October 22, 2009 | 2009 | Boston Scientific Corporation | The sterile barrier in the packaging (pouch) that contains the catheter may be breached. The breach of sterility could lead to a contamination of the device with subsequent transfer of infectious agents to the patient. | Packaging process control | Two Boston Scientific "Urgent Voluntary Medical Device Recall" letters were sent to consignees on October 29, 2009. The letters described the problem and the affected product. Consignees were instructed to cease use and distribution of the product and to complete and return the Account Reply Verification Tracking Form. The first letter with the heading, "Boston Scientific IMAGER II Angiographic Catheter (Peripheral Intervention)", was addressed to the Risk Manager/Field Action Contact. The second letter with the heading, "Imager II Urology Torque Catheters" was addressed to the Risk/Materials Manager. Direct questions about the recall to your local Sales Representative or Boston Scientific Corporation by calling at 1-763-494-1700. |  | 28321690 | 8845 boxes (5/box) US ; 3202 boxes (5/box) OUS | Worldwide Distribution -- United States (AL, AZ, AR, CA, CO,CT, DE, DC, FL, GA, HI, IL IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO,MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SC, TN, TX, UT, VT, VA, WA, WV and WI), United Arab Emir., Dutch Antilles, Austria, Australia, Belgium, Brazil, Canada, Switzerland, Chile, Colombia, Cyprus, Czech Republic, Germany, Denmark, Egypt, Spain, Finland, France, Great Britain, Greece, Hong Kong, Hungary, Israel, Italy, South Korea, Kuwait, Lebanon, Lithuania, Mexico, Malaysia, Netherlands, Norway, New Zealand, Poland, Puerto Rico, Portugal, Qatar, Russian Fed., Saudi Arabia, Sweden, Singapore, Slovakia, Syria, Thailand, Tunisia, Turkey, Trinidad, Tobago, Taiwan, and South Africa. | Terminated | December 12, 2011 | 781 | 28323022
Z-0587-2010 | 54153 | No | NVK | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | Dental | 510(k) | SoftLase Pro Dental Lasers: Zap SoftLase, SoftLase Pro, OrthoLase, and HygieneLase models, Model numbers ZA-SLP, ZA-ORT, ZA-HYG, G22002SLP, ZA-SOFTLASE2WSII, ZA-SOFTLASER2W, Manufactured by Zap Lasers LLC, Pleasant Hill, CA. |  |  | August 08, 2008 | 2008 | Zap Lasers LLC | The SoftLase Lasers are in need of compliance upgrades due to lack of a remote interlock connector, an emission delay, user guide labels and locations, and calibration procedures and schedule. | Radiation Control for Health and Safety Act | Zap Lasers sent a letter to affectedcustomers stating that their SoftLase diode laser is eligible for a compliance maintenance upgrade. The upgrades will be free of charge and will enhance the safety of the laser. Customers are encouraged to call the firm's customer service line at 888-876-4547 to schedule their RMA number and shipping instructions. The letter also lists the upgrades that are needed. |  | 1375 | 1375 units | Worldwide Distribution to 1397 consignees | Terminated | May 25, 2010 | 655 | 1375
Z-0495-2010 | 53883 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Infinia with Hawkeye 4, model numbers: H3000WC and H3000WD. GE Healthcare, Waukesha, Wisconsin 53118. |  |  | August 01, 2009 | 2009 | GE Medical Systems, LLC | Failure to provide the complete User Manualinformation. | Radiation Control for Health and Safety Act | A Field Modification Instruction will be released in February 2010. GE Healthcare will correct the Operator Manuals in forward production and the installed base units. |  | 343 | 343 total | Nationwide Distribution | Terminated | February 04, 2012 | 917 | 343
Z-0613-2010 | 53941 | No | GKG | Hematocrit measuring device. | Hematology | Hematology | 510(K) Exempt | Haematokrit 24 Centrifuge with 2050 Rotor |  |  | November 21, 2009 | 2009 | Hettich Instruments Management I | Centrifuge Rotor Failure: may crack, break, and explode outside the centrifuge. | Device Design | Hettich Instruments notified Distributors by email on 11/21/09, to the Request For Urgent Action notifications, dated 11/23/2009. There is a description of how to identify a recalled rotor as well as a description of the replacement metal rotor that is to be used in place of the list of centrifuges affected by this recall. Customers are to cease using the setup described and contact their distributor or the firm. The firm will replace the plastic housing with a metal housing. Questions should be directed to 978-232-3957. |  | 46 | 46 units total | Worldwide Distribution -- USA, including states of CT and IN, and country of Canada. | Terminated | October 25, 2011 | 703 | 46
Z-0638-2010 | 54203 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Exel Huber Needles (Right angle and Straight Models with Catalog Numbers: 26901CE, 26902CE, 26903CE, 26904CE, 26905CE, 26906CE, 26907CE, 26908CE, 26909CE, 26910CE, 26911CE, 26921CE, 26922CE, 26923CE, 26924CE, 26925CE, 26926CE, 26927CE, 26928CE, and 26929CE). |  |  | December 23, 2009 | 2009 | Exelint International Co. | Non-coring needles that, when inserted into an implanted port, were shown to core the silicone port's septum. Coring of port's septum may lead to: decreased effectiveness of port, infection, necrosis, swelling, and embolization, which may result in pulmonary embolism, limb ischemia, stroke, local extravasation of medication such as chemotherapy or other caustic material; may lead to local tissue | Device Design | Distributors have been advised to cease distribution of any recalled needles and infusion sets; to quarantine any remaining; to document amount on Customer Response Form; to return form to Exel (firm); to arrange for return of product to Exel for credit. Distributors also are advised to share recall notice with all affected customers. Contact phone: 727-827-1922x0 (M-F 9a-5p EST). Contact email: info@exelint.com |  | 849000 | 849,000 | US (VA TX NV KY OH NH MS NY OH NM DE WA AZ IL PA OR CO FL GA LA UT MI WI NV CA NJ MO IA) and Canada. | Terminated | June 06, 2011 | 530 | 2010360
Z-0662-2010 | 53802 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | AC Power Cords, 110V Manufactured by Electri-Cord Manufacturing Company in Westfield, PA. Power Cord #4100079 used with Kinair IV Model 22600, Kinair IV 60601 Model 216050, Triadyne Proventa Model 406000, Therapulse ATP/II Models 227800 and 227080.S, Kinair Medsurg Model 201001W, and Kinair Medsurg Pulse Model 40900. |  |  | November 04, 2009 | 2009 | KCI USA, Inc. | Power cords may crack and fail inside plug with potential for fire hazard. | Nonconforming Material/Component | Kinetics Concepts, Inc. (KCI) issued an "Urgent Medical Device Correction" letter dated November 4, 2009. Consignees were informed of the affected product and instructed to take necessary action to avoid any further issues associated with the product. For further information, contact KCI's technical support team at 1-800-275-4524 option #3. |  | 4637 | 4,637 units | Nationwide and Canada | Terminated | January 21, 2011 | 443 | 79267
Z-0716-2010 | 53955 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Outlook Pump Nitroglycerin Set w/Universal Spike &amp; Spin Lock Connector, 135 in |  |  | November 24, 2009 | 2009 | B. Braun Medical, Inc. | System error 9 alarms. A System Error 9 alarm indicates a potential pressure loss in the pump cassette delivery chamber (chamber closest to the patient). This alarm is generated during the pumps initiation sequence and will not allow the set to be used. This requires a new set to be obtained, thus creating a potential for delay in therapy. | Nonconforming Material/Component | The recalling firm issued a Product Removal Notice via US Postal Service Certified Mail with registered return receipt or FedEx to all customers and distributors in receipt of the affected product. The notices were sent 11/30/09 and 12/1/09. Direct customers also received letters addressed to Material Management that informed them of the recall and provided instruction for disposition of affected product in their inventory. Distributors received the direct customer recall letter along with a cover letter outlining their responsibilities to execute the recall and notify their customers.Questions are directed to the company's Customer Support Departments: US customers please call (800) 227-2862. Canadian customers, please call (800) 624-2920.International customers, please call (610) 691-5400. |  | 57480 | 57480 units | The products were shipped to distributors and medical facilities nationwide. The product was also shipped to Afghanistan, Canada and LU. | Terminated | February 14, 2011 | 447 | 5435249
Z-0653-2010 | 53554 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Edwards Lifesciences Aquarius System Model GEF09800. |  |  | March 16, 2009 | 2009 | Edwards Lifesciences, Llc | There have been reports of clinically significant fluid imbalance when a certain level of fluid imbalance is detected the Aquarius will trigger an alarm, however, users are able to override this alarm and continue therapy. It is possible to remove too much fluid from or replace too much fluid to the patient. In extreme cases, this may result in serious injuries or death. | Employee error | Edwards is continuing a field corrective action related to the possibility of fluid imbalance when using the Aquarius. In order to mitigate some of the risk associated with this problem, Edwards sent notifications to consignees advising of mitigation steps. Consignees were provided: a description of the affected product, problem, and potential hazard; advice on action to be taken; and revised device labeling. Additionally, consignees were notified of a planned software upgrade to prevent users from bypassing the fluid balance alarm more than five times in a 20-minute period. Although Edwards does not believe there is a need to stop using the Aquarius system at this time, clinicians should weigh the risks/benefits to the patients before continuing the use of the device Edwards first provided some of this information to consignees in a Field Safety Notice on March 16, 2009. On January 11, 2010, Edwards provided an Urgent Product Recall letter to consignees which updated and expanded the earlier communication. The FDA classified this as a Class I recall. |  | 135 | 135 units | Nationwide Distribution. | Terminated | January 18, 2012 | 1038 | 135
Z-0602-2010 | 53692 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | DPX-NT series, GE Lunar DEXA bone densitometers operating with enCORE software versions 12.0, 12.1, 12.2,13.0, &amp; 13.1. |  |  | February 01, 2009 | 2009 | Ge Healthcare | Failure to properly pre-indicate the technique factors to be used during a patient scan. | Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter will be sent to customers. The letter will discuss the defect description and related hazards, the affected product details, safety instructions, and product correction. Customers are to contact their local GE Healthcare service representative with questions. |  | 1764 | 1,764 total | Nationwide Distribution | Terminated | December 31, 2011 | 1063 | 1764
Z-0748-2010 | 51104 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Jeanes Hospital Custom Cath Procedure Tray, K09T-08945A, Sterile EO, Merit Medical Systems, Inc., 1600 West Merit Parkway, South Jordan, Utah 84095. |  |  | January 29, 2009 | 2009 | Merit Medical Systems, Inc. | A manufacturing defect may prevent adequate sterilization of the entire device. | Process control | Applicable Merit sales representatives and distributors were notified by phone and letter beginning January 29, 2009 and instructed to immediately contact all affected consignees. Consignees were to immediately quarantine any remaining product and arrange for return to the firm.For further information, contact Merit Medical Systems, Inc. at 1-800-356-3748 extensions 4929 or 4998. |  | 157 | 157 units | Worldwide Distribution -- United States ( PA and TX) and Japan. | Terminated | June 16, 2011 | 868 | 910
Z-0749-2010 | 53607 | Yes | NIM |  |  | Cardiovascular | PMA | RX ACCULINK Carotid Stent System, 10x40 mm, Product Number: 1011344-40, Lot Number: 9081851. |  |  | October 01, 2009 | 2009 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | The recall was initiated after Abbott Vascular discovered, through internal testing, that the affected lot may not meet our quality specifications for catheter shaft tensile strength. | Process control | Abbott Vascular's Sales Representatives initiated customer contact with the identified consignees beginning October 1, 2009 . They provided the consignees with a recall letter with attached fax-back effectiveness check form. The recall letter instructed the consignees to work with their local account representative to review their inventory, complete the attached recall effectiveness check form and return identified products to Abbott Vascular.For further information, contact Abbott Vascular Customer Service at 1-800-227-9902. |  | 22 | 22 units | Nationwide Distribution. | Terminated | September 27, 2011 | 726 | 22
Z-0732-2010 | 54149 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Stryker CBCII ConstaVac Blood Conservagtion System w/ Quick Disconnect and 1/8in. (3.2mm) PVC Drain, sterile, Stryker Instruments, Kalamazoo, MI REF 225-28-914. |  |  | November 12, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact. | Packaging process control | Stryker issued a "Medical Device Recall Notification" via email to foreign consignees dated November 12, 2009. Stryker further issued recall letters via overnight mail to domestic consginees beginning November 20, 2009. These notifications informed consignees of the affected devices instructing them to destroy all affected product and return a completed Business Reply Form to the firm.For further information, contact Stryker Instruments at 1-800-800-4236 extension 3808. |  | 79617 | 79,617 | Worldwide Distribution. | Terminated | October 25, 2010 | 347 | 79617
Z-0757-2010 | 54138 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Total Stabilizer Femoral Component; Size # 2 |  |  | December 10, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Triathlon TS femoral components may have been damaged during manufacturing, possibly preventing the assembly of a stem extension. | Nonconforming Material/Component | An "Urgent Product Recall" Letter dated December 10, 2009 was sent by Federal Express to all Stryker Branches, Chief of Orthopaedics, Hospital Risk Management and Surgeons. The letter described the affected product, issue and potential hazards. Consignees were instructed to complete Product Recall Acknowledgement Form and return to Stryker Orthopaedics via fax at 1-201-831-6069. All questions regarding the recall should be directed to Stryker Orthopaedics at 1-201-972-2100. |  | 5 | 5 | Worldwide Distribution -- United States (DE, NC, OH. MI, TX, CA, MA, NJ, TN, KS, GA, FL, AL, IN, MN, MD, NE, AR, OK, CO, AZ, OR, and MA), Australia, United Kingdom and Canada. | Terminated | August 28, 2012 | 992 | 90
Z-0755-2010 | 53749 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | MPS Acacia Extension Set, 15cm Small Bore, Spin Nut w/BD Q-Syte, REF 385151 |  |  | October 19, 2009 | 2009 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter, possibly resulting in an air embolism. | Process control | Consignees were notified by B-D Urgent Medical Device Recall letter on 10/28/2009 and asked to return all unused affected lots. For information, call the company at 800-453-4538, option 2, extension 2341. Another letter was sent to all affected customers on 02/08/2010 notifying them of the extension of the recall to include Nexiva products. A response card was included to be returned whether or not product was on hand. B-D issued a press release on Febriuary 8, 2010. It can be found on www.bd.com at http://www.bd.com/contentmanager/b_article.asp?Item_ID=24486&ContentType;_ID=1&BusinessCode;=20001&d;=BD+Worldwide&s;=&dTitle;=&dc;=&dcTitle;= |  | 45350 | 45,350 units | Nationwide distribution. | Terminated | October 19, 2010 | 365 | 7895560
Z-0779-2010 | 54141 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet bone mulch screw, 10.5mm x 20 mm, TI-6AL-4V alloy, sterile, Biomet Sports Medicine, Warsaw, IN; REF 907320. |  |  | December 22, 2009 | 2009 | Biomet, Inc. | The screw may not have the hex dimension inside of the head of the screw, which will prevent the device from being implantable. This may cause a delay in the procedure while another screw is obtained. | Employee error | Hospitals, distributors, and other customers were notified by letter dated 12/22/09 and instructed to locate and return these lots of product. Questions related to this notice should be directed to the firm at 800-348-9500 or 574-372-3983, Monday through Friday, 8AM to 5PM. |  | 169 | 169 | Arkansas, California, Delaware, Florida, Kansas, Kentucky, Michigan, Missouri, New York, Pennsylvania, South Dakota, Texas, Washington, Wisconsin, Chile, Germany, Italy, Japan and Norway. | Terminated | October 22, 2010 | 304 | 169
Z-0764-2010 | 54172 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Percutaneous Catheter Introducer Set, VNUS Medical Technologies Cath Shth Intro 7F x 7cm, catalog number IS-7F07, Disposable, Single Use, Do Not Re-sterilize, Manufactured in USA for VNUS, 5799 Fontanus Way, San Jose CA 95138 USA; General: 408-360-7200. |  |  | January 06, 2009 | 2009 | Argon Medical Devices, Inc | One lot of Percutaneous Catheter Introducer Set contained incorrect size needle. | Other | Firm notified sole consignee by an Urgent Medical Device Recall certified letter dated 1/6/2010. The letter identified the product being recalled and asked the consignee to immediately quarantine any product remaining at their facility. The recall should be conducted to the user level. Affected products at the consignee's facility and with their customers, should be returned using the instructions on the response form. The response forms are to completed and returned by fax or e-mail as indicated. Questions or the need for additional information concerning this recall should be directed to Sarah Larson at 903-677-9351 or Sarah.Larson@argonmedical.com. |  | 500 | 500 units | Nationwide Distribution -- State of CA. | Terminated | July 14, 2010 | 554 | 500
Z-0661-2010 | 54193 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Thomas Medical Products Safesheath CSG (Coronary Sinus Guide) Hemostatic Tear away Introducer System with Infusion Side Port, for the introduction of various types of pacing or defibrillator leads and catheters. The product is shipped 5 per carton. Product is distributed by Pressure Products. Catalog numbers: CSG/MSP-00-09, CSG/90-09, CSG-KR-90-09, CSG/WORLEY-1-09, CSG-WORLEY-Bcor-1-09, CSG/WORLEY/L/Bcor-1-09, CSG/WORLEY-2-09, CSG/WORLEY/Bcor-2-09, CSGMSP0009M, CSGWORLEY109M, CSGWORL19M, CSGWORBC19M, CSGWORLBC19M, CSGWORBC29M, CSG/Worley/BCor-1-07, CSG/Worley/BCor-2-07, CSG/WORLEY/BCOR-M-09, and CSG/Worley/R-1-09. |  |  | December 17, 2009 | 2009 | Thomas Medical Products Inc | Radiopaque sheath tip may fracture | Component design/selection | GE Healthcare Thomas Medical Products firm issued an Urgent Medical Device Correction email on 12/23/09 to the distributor to discontinue distribution and return any affected products. The distributor was instructed to conduct a sub recall regarding the issue. Press issued 1/29/2010. |  | 60982 | 60,982 units | Nationwide. | Terminated | November 01, 2010 | 319 | 60982
Z-0668-2010 | 54178 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Discovery CT 750 HD model numbers: 5232083, 5232083-2, 5232083-3, and 5232083-4. |  |  | October 01, 2009 | 2009 | Ge Healthcare | Failure to properly document the CTDI in the Technical Reference or User Manual. | Radiation Control for Health and Safety Act | GE is sending an Important Electronic Product Radiation Warning letter to customers. The letter will describe the defect, identify the affected product details, provide instructions, state the product correction, and the necessary contact information. |  | 111 | 111 units total | Nationwide Distribution -- Including states of CT, SD, VA, WA, MA, PA, GA, IL, CA, MD, NY, NC, SC, AZ, TX, MT, FL, OH, OK, MI, CO, ME, WV, TN, KS, and WI. | Terminated | December 10, 2011 | 800 | 111
Z-0782-2010 | 53947 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | TempPadÂż Cold Therapy Pad. |  |  | November 14, 2009 | 2009 | Cincinnati Sub-Zero Products Inc | Product is labeled as "Sterile". It was observed that incomplete seals were created in the packaging materials during production. When the seal is compromised, the sterility of the product inside the package cannot be assured. | Other | On November 14, 2009, Cincinnati Sub-zero Products, Inc. (CSZ) sent a recall notification letter and response letter to customers. A second notice dated December 10, 2009 was also issued to customers. Customers were instructed in the notifications to assume the recalled product is not sterile and do not use. CSZ requested all affected product be returned to their firm for replacement. Customers are to complete and return the response letter acknowledging receipt of the recall letter. Direct questions to CSZ by calling Customer Service at 1-800- 989-7373. Customers may also call 1-800-989-7373, extension 3234 or 1-51 3-71 9-3234 for recall instructions. Representatives will be available to take your call from 8:30 a.m. until 4:30 p.m. eastern standard time, Monday through Friday. |  | 6006344 | 8,872 cartons of ten + 677 each (89,397 total) | Worldwide Distribution -- United States (AL, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MT, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI &amp; WV), Canada, France, Germany, Ireland, Israel, New Zealand, Singapore, and The Netherlands. | Terminated | January 25, 2012 | 802 | 6007439
Z-0608-2010 | 54226 | No | REG | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Model ST-400 Solafeet Foot portable tanners; Solafeet, Inc. |  |  | September 08, 2009 | 2009 | Solafeet, Inc. | Product failed to have a label that contained a recommended exposure schedule; the operator's manual provided inadequate instructions for use; and label implication may result in overexposure. | Radiation Control for Health and Safety Act | Solafeet sent letters to customers informing them that the labeling and instructions included with the product were not in accordance with the FDA requirements. Attached with the letter was a replacement label and a replacement set of instructions. Customers are to attach the new label on top of the old label and use the new instructions, noting the change in exposure limits. Questions should be directed to 941-359-2608 or info@solafeet.com. |  | 72 | 72 units | Nationwide Distribution | Terminated | August 21, 2012 | 1078 | 72
Z-0789-2010 | 54274 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Xia Precision System; Jam Shidi 10 Gauge 9 inch; Sterile; Manufactured for: Stryker Spine SAS, Cestas, France; Distributed in the US by Howmedica Osteonics Corp, 325 Corporate Drive, Mahwah, NJ. |  |  | January 18, 2010 | 2010 | Stryker Spine | Stryker Spine received reports of packaging failures during shipping for the Jam Shidi instruments, which may compromise the sterility of the product. | Package design/selection | Stryker Spine issued an "Urgent Product Recall" notification dated January 15, 2010 via Federal Express to all consignees. The letter informed users of the issue and potential hazards of the affected product.For further information, contact your local Stryker Spine Sales Representative or call 201-760-8206. |  | 2727 | 2727 total - all sizes. | Nationwide distribution. | Terminated | September 06, 2012 | 962 | 2727
Z-0794-2010 | 50099 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Navigation Reamer T-handle; Non Sterile; Howmedica Osteonics, 325 Corporate Drive, Mahwah, NJ 07430. |  |  | August 08, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics learned that the T-handles have the potential to not engage and not attach to the reamer properly. | Device Design | Stryker Orthopaedics issued notification dated August 8, 2005 via Federal Express. Consignees were informed of the affected product and necessary steps to take for product replacement.For further information, contact Stryker at (201) 831-5825. |  | 22 | 22 | Worldwide Distribution -- United States, Germany, The Netherlands, Spain, Italy, Switzerland, Poland, Romania, South Africa, England, Columbia, Venezuela, Japan and Australia. | Terminated | February 18, 2010 | 1655 | 295
Z-0795-2010 | 50101 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Scorpio PS Femoral Waffle (No post) and LFIT; Use only with Scorpio PS Tibial Bearing inserts and Scorpio Patellas, Stryker Orthopaedics. |  |  | July 30, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that a box packaged and labeled as a Scorpio PS Femoral Waffle (No Post) with LFIT, may actually contain a Scorpio Femoral Waffle (With Post) and LFIT. | Labeling mix-ups | Stryker Orthopaedics issued an "Important Product Removal" letter via Federal Express dated July 30, 2007 to all direct accounts. Users were asked to identify and return all affected product to the firm.For further information, contact Stryker Orthopaedics at (201) 831-5825. |  | 7 | 7 | Worldwide Distribution -- United States and France. | Terminated | February 19, 2010 | 935 | 7
Z-0796-2010 | 54198 | No | HTJ | Depth gauge for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Depth Gauge; Catalog Number: 703707; |  |  | December 22, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The Depth Measuring Gauges have 5 scale marking instead of the required 4 scale markings. | Process control | Urgent Product Recall letters, dated December 22, 2009, were sent to all Stryker Branches/Agencies. Letters were also sent to hospitals and surgeons that use the VariaAx Elbow System were also notified. All letters were sent by Federal Express. The letters identified the affected product, stated the issue, and the hazards involved. It also asked customers to examine their inventory and hospital locations to identify affected products. All affected products are to be retrieved and returned to their branch or agency warehouse for reconciliation. All products should be reconciled on the attached Product Accountability Form and a copy should be faxed to 201-831-6069 within 5 days of receipt of the notice. Follow the instructions for shipping the affected product back to the firm. Questions should be directed to 201-972-2100. |  | 49 | 49 | Nationwide Distribution | Terminated | March 16, 2010 | 84 | 49
Z-0780-2010 | 53989 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MRI Eclipse 1.5T System Model #781230. |  |  | December 07, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | Polymeric RF connector block on the tabletop overheated and eventually ignited during scanning due to combination of arcing of electrical components associated with the RF trap assembly located below the connector and flame rating of the connector block material. | Other | An "Urgent-Field Safety Notice" dated December 4, 2009 was issued via certified mail to customers. The notification letter described the affected product, problem, hazard involved and action to be taken by customer. The customer will ensure to maintain frequent contact with patients during scanning and that the patient comfort kit is installed for all patient studies. The Customer will be visited by a Philips Service Representative to inspect their MRI System and install the newly remanufactured RF Connector Block. Additionally, step-by-step trap/cable installation instructions including pictures are being drafted to ensure correct installation of cabling. If you need any further information or support concerning this issue, please contact your local Philips representative or 1-800-722-9377. |  | 290 | 290 units | Worldwide distribution- United states: (AL, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, LA, MA, MD, ME, MN, MO, NE, NH, NJ, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA &amp; WI), ARGENTINA, BRAZIL, CANADA, CHINA, CYPRUS, FRANCE, GERMANY, GREECE, ITALY, JAPAN, KOREA, NEW ZEALAND, PAKISTAN, POLAND, RUSSIA, SWITZERLAND, TAIWAN &amp; UNITED KINGDOM. | Terminated | September 19, 2012 | 1017 | 290
Z-0798-2010 | 54233 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | 2-Channel Laryngeal Surface Electrode (Model LSE500DCS and LSE500DCL) |  |  | January 11, 2010 | 2010 | Neurovision Medical Products Inc | This recall has been initiated due to a conductive glue used to attach the lead wire to electrode, which has been found in a small number of electrodes, to cause a short circuit in one of the channels. This is due to migration of the adhesive from one pad to the other creating an electrical short. Use of one of these defective electrodes could lead to a failure to monitor. Our investigation has c | Other | URGENT RECALL: Medical Device Letter dated 1/7/10.The letter identified the affected product and provided a list of necessary actions for customers to take.The customers should immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, please identify their customers and notify them at once of this product recall. Their notification to their customers may be enhanced by including a copy of this recall notification letter.This recall should be carried out to the user level. Your assistance is appreciated and necessary to prevent any failure to monitor.Please complete and return the enclosed response form as soon as possible. If you have any questions, call Ryan Rea or Kevin Rea, at 866-815-6999. |  | 4143 | 4143 | Nationwide: CA CO FL GA IL LA MI ND NY OH PA TX UT WA WI) | Terminated | April 21, 2011 | 465 | 4143
Z-0811-2010 | 54005 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Biphasic LIFEPAK 12 defibrillator / monitor. |  |  | March 31, 2009 | 2009 | Physio Control, Inc. | Affected devices contain a Biphasic PCBA that may have a solder defect on the H Bridge component that could prevent the device from providing a defibrillation shock. There could be a delay in therapy or prevention of defibrillation therapy if the defibrillator were needed in a cardiac arrest situation. Delay or prevention of defibrillation therapy may contribute to the death of a viable patient. | Nonconforming Material/Component | Consignees were sent a notification letter dated March 2009 that explained firm representatives would follow up with each customer within 30 days to arrange an update to affected devices.Affected Biphasic PCBAs will be replaced. Customers advised to keep the devices in use and to continue to perform daily User Test. |  | 1718 | 1718 | Distributed domestically and internationally to hospitals, clinics, etc.Foreign distribution to the following countries:Australia, Brazil, Canada, Colombia, Costa Rica, Germany, Hong Kong, Mexico, Netherlands, and Puerto Rico. | Terminated | August 08, 2013 | 1591 | 1718
Z-0808-2010 | 54195 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LifeShield Latex-Free Primary SYMBIQ Set, Piggyback with Backcheck Valve, 3 Prepierced Y-Sites, 106 inch, Non-DEHP, 15 drops/mL; an Rx sterile single use fluid path way for administration of I.V. fluids using the Symbiq infusion pump; 48 sets per case; Hospira, Inc., Lake Forest, IL 60045 USA, Made in Costa Rica; list number 16096-28. |  |  | January 08, 2010 | 2010 | Hospira Inc. | There is a potential for fluid to leak from the backcheck valve which could result in inaccurate delivery and occupational exposure. | Process control | Hospira Medical issued an "Urgent Device Recall" notification dated January 8, 2010 informing consignees of the affected product. The accounts were instructed to examine their inventory and quarantine affected product immediately and arrange for its return. Users were also requested to complete the enclosed reply form and return it to the firm. For further information, contact Hospira Medical Communications at 1-800-615-0187. |  | 47904 | 47,904 sets | Minnesota and Utah | Terminated | July 20, 2011 | 558 | 49440
Z-0810-2010 | 54074 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | LOCATOR Restorative Abutment, ILOA002, LOT 877423, Certain LOCATOR ABUTMENT, 4.1 mm(D) X 2mm(H), STERILE Made in USA, Sterile. |  |  | November 23, 2009 | 2009 | Biomet 3i, LLC | Package labeled as sterile may have an incomplete or partial seal thereby possibly compromising sterility. | Process control | Each US customer will be contacted by telephone using a prepared phone script. Following telephone contact, Biomet 3i will perform a follow-up of each call with a faxed letter. The Urgent: Medical Device Recall letter, dated 11/23/2009, identified the affected product and asked customers to return unused product for replacement. The letter also explained the clinical implications and asked customers to respond with the attached response fax and then return the abutment(s) to BIOMET 3i. Questions or concerns can be directed to Kelly Taylor at 561-776-6906. |  | 97 | 97 | Worldwide Distribution -- USA, Middle East, Italy, Europe, and Greece. | Terminated | December 15, 2010 | 387 | 97
Z-0806-2010 | 54485 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Endotek Alimaxx-B Uncovered Biliary Stent for Endoscopic Delivery, REF 70129-210, 10 x 80, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095. |  |  | January 28, 2010 | 2010 | Merit Medical Systems, Inc. | Degradation of the outer portion of the delivery catheter may prevent proper deployment of the biliary stent. | Component change control | Merit Medical Systems, Inc initiated customer notification by phone and letter beginning January 28, 2010. Consignees were instructed to immediately contact their affected customers, advise them of the recall activity and to identify and quarantine any unused inventory. A Recall Notification Form is to be completed and signed by the sales rep and by the site representative. Affected product is to returned to the firm. For further information, contact Merit Medical Systems, Inc. at 1-801-208-4228. |  | 51 | 51 units | Nationwide Distribution -- AL, CA, CO, FL, IL, KY, NY, MA, MD, ME, MN, MO, PA, TX, UT, VA, WI, VA, TX and WI. | Terminated | April 27, 2010 | 89 | 165
Z-0813-2010 | 52754 | No | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | Abbott i-STAT ACT Celite Cartridges. |  |  | August 18, 2009 | 2009 | Abbott Point of Care Inc. | Labeling discrepancy: The manual instruction state that the line should be flushed with 5 mL of saline and the first 5 mL of blood or six dead space volumes should be discarded. The Cartridge Test Information sheets that it should be 2 mL of blood - which is incorrect. | Labeling False and Misleading | Abbott Point of Care issued an "Important Product Information" letter beginning August 18, 2009. Consignees were instructed on how to reconcile all affected product.For further information, contact Abbott Point of Care Technical Support at 1-800-366-8020, Option 1. |  | 696675 | 696,675 Cartridges | Worldwide distribution. | Terminated | May 18, 2010 | 273 | 1807450
Z-0819-2010 | 53025 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CellSearch Circulating Tumor Cell Kit (Epithelial) IVD |  |  | August 12, 2009 | 2009 | Veridex, LLC | Veridex has received 4 complaints regarding reagent bottles having been incorrectly placed or duplicate bottles places within the reagent tray of CellSearch Circulating Tumor Cell Kit. Delayed detection of disease relapse or progression may occur due to false low CTC results. | Other | Urgent Product Correction Notification letters were sent to all customers on August 12, 2009 by Federal Express. Customers were advised to inspect the kits to assure the proper reagents are in the proper position. If not, do not use the kits and contact the company at 877-837-4339 (chose option 1). |  | 2021 | 1641 in US, 380 internationally. | Product was distributed worldwide, US (one government account: National Cancer Institute, Frederick, MD). Also 8 international accounts in Belgium, France, Germany, Italy, Spain, United Kingdom, Netherlands, Norway and Israel. | Terminated | July 18, 2012 | 1071 | 2021
Z-0817-2010 | 53553 | No | JWX | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Oxoid dry Spot Staphytect Plus, Catalog #DR0100M, 120 tests/kit. Oxoid Ltd., Wade road, Basingstoke, Hants, UK. |  |  | July 20, 2009 | 2009 | Remel, Inc | The products may give false positive results with some non-Staphylococcus aureus strains | Process control | Thermo Fischer Scientific, Remel Inc. issued letters dated July 20, 2009 via regular mail explaining the reason for recall, risks involved, and requested they discard all remaining product. A Product Inventory Checklist was enclosed to report the amount of product their customer discarded. A second letter dated October 19, 2009 was issued via certified mail.For further information, contact Thermo Fischer Scientific, Remel Inc. at 1-800-255-6730, Extension 3304185. |  | 27 | 27 kits | Nationwide Distribution -- MN, CT, FL, GA, CA, NY, MI, MO, PA, and IL. | Terminated | April 25, 2012 | 1010 | 77
Z-0812-2010 | 53335 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT 1 Analyzer, Abbott Point of Care Inc., Abbott Park, IL 60064 USA, Product of USA. |  |  | September 09, 2009 | 2009 | Abbott Point of Care Inc. | The storage temperature on the label on the side of the i-Stat Portable Clinical Analyzer shipping box does not match product specification. | Labeling mix-ups | An "Urgent Recall Notice" was issued to customers during September 2009. The letter described the affected product, issue and recommended actions. Customers were asked to verify storage conditions on the box are correct and if not, discard the outside box. If the customer requires a replacement box, the customer must contact Abbott Point of Care Technical Support. Direct questions and requests for additional information to the Abbott Point of Care Technical Support by calling 1-800-366-8020, Option 1 or contact your Abbott Point of Care representative. |  | 3918 | 3,918 analyzers | Worldwide Distribution (including the United States). | Terminated | December 08, 2014 | 1916 | 3918
Z-0827-2010 | 54184 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | RambaChrom KPC M/N 11500. RambaCHROM KPC is for detection of gram negative bacteria with a reduced susceptibility to most of the carbapenem agents. |  |  | December 17, 2009 | 2009 | Gibson Laboratories Inc | Discrepancies between the product package label and technical insert are conflicting and may be misleading to user. The technical insert included the statement for In Vitro Diagnostic Use. The product insert should have indicated for Research Use Only instead of for In Vitro Diagnostic Use. | Employee error | On 12/10/2009 the firm sent letters to their customers notifying them of the problem and instructing them to replace insert with a new one. They were also instructed to complete and return a "Recall Response Form." |  | 3440 | 3,440 units | Distribution of the product was limited to AL, FL, IL, KY, MA, MI, MO, NY, OH, PA, &amp; VA and one customer in France. | Terminated | September 21, 2012 | 1009 | 18770
Z-0823-2010 | 54271 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker XIA 3 Polyaxial Screw, 8.5mm x 65 mm. Catalog Number: 482318565. |  |  | December 30, 2009 | 2009 | Stryker Spine | The Xia 3 8.5mm Polyaxial screws, lot A91109 were anodized turquoise instead of Fuchsia. The turquoise color is reserved for the 9.5 mm Xia 3 screws. | Process control | Stryker Spine branches that received the recalled product were notified of the recall by an Urgent Product Recall letter, dated January 15, 2010, sent via Federal Express. The letter identified the affected product and the issue involved. It also discussed the potential hazards, risk mitigation, and the return process. Customers are to examine their inventory and hospital locations to identify the product. Customers are also to reconcile the product by utilizing the Customer Response Form and distribution sheet. In addition, they are to retrieve and return the affected product. Questions should be directed to Tiffani Rogers at 201-760-8206. |  | 21 | 21 | Nationwide Distribution -- IL, MO, DA, MA and CO. | Terminated | September 06, 2012 | 981 | 21
Z-0826-2010 | 54197 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Encore Medical Foundation Knee Revision Baseplate, Part Number 321-01-106, Size 6, Left |  |  | December 11, 2009 | 2009 | Encore Medical, Lp | Partial lot of product was not manufactured to specification. Central locking screw was missing threads. There is a potential for dissociation because the screw is not able to be used as a secondary fixation method. | Process control | Firm notified consignees by letter on 12/11/2009. The letter informs physicians that patients with implanted devices should counseled about the issue and should report any changes to their treating physician. |  | 17 | 17 | TX, FL, GA, CA, AZ, Japan | Terminated | September 24, 2010 | 287 | 17
Z-0825-2010 | 54238 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes In Situ Bender/Cutter Kit-Wide-Sterile. Catalog Number 530.521S. |  |  | January 06, 2010 | 2010 | Synthes USA (HQ), Inc. | There is the potential for the bender/cutter attachment to continue heating after release of the powerbutton. There is potential for patient and/or staff injury if continuous heating is unrecognized. | Other | An "Urgent: Medical Device Recall" letter dated January 8, 2010 was issued to customers via USPS Certified mail. The letter described the issue and affected product. Customers were notified to cease use of the recalled product, examine inventory and follow instructions for return to the firm as indicated in the recall letter. For questions, please call 1-800-620-7025 x 5452 or 1-610-719-5452 or contact your Synthes CMF Sales Consultant. |  | 187 | 187 units | Nationwide Distribution. | Terminated | October 06, 2010 | 273 | 187
Z-0822-2010 | 53702 | No | KIR | Cultured animal and human cells. | Hematology | Pathology | 510(K) Exempt | Rabbit Kidney Culture, brand name: 'FreshCells'. The product is labeled as 'Rabbit Kidney' also known as 'Rbk' , 2-square cm area of a cell monolayer covered with 2-mL culture medium. The product is packed in glass culture tubes 16x125mm, with screw caps. The product is distributed in EPS shipping racks which can hold up to 50 tubes each. |  |  | January 21, 2009 | 2009 | Diagnostic Hybrids Inc | During QC post release monitoring, the firm determined that the monolayer culture of rabbit kidney cells previously distributed by the firm, appeared to contain another type of cells (monkey kidney cells). | Process control | Diagnostic Hybrids issued a "Product Notification" letter dated January 23, 2009 in addition to verbal communication of the affected product beginning January 21, 2009. The firm requested that all affected product be discarded /destroyed and not use for diagnostic testing. Consignees were asked to complete and return a Fax-Back form.For further information, contact Diagnostic Hybrids Technical Services at 1-866-344-3477 or 1-740-589-3300. |  | 1127 | 1127 units | Worldwide Distribution -- United States (CA, FL, GA, MA, MO, NY, NJ, TN, TX and VA), Ontario Canada and Manitoba Canada. | Terminated | September 21, 2010 | 608 | 1127
Z-0846-2010 | 54249 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker T2 Ankle Arthrodesis Nail, right; 012 x 300 mm; Catalog number: 18191230S; |  |  | December 29, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that there is the potential for damage to the sterile packaging of the product during transport. | Packaging | Urgent Product Recall notification letters were sent to branches on 12/29, 2009, with a corrected version that included a list of catalog numbers and code numbers on January 11, 2009. Letters were also sent by Fed Ex on January 11, 2009 to Hospital Risk Management, Chief of Orthopaedics and surgeons who may have used the device. The letters identified the affected product, described the issue, and the hazards involved. The letter also asked customers to examine their inventory and hospital locations to identify the affected product. They are to retrieve all affected product and return it to their branch or agency warehouse for reconciliation. Customers are to reconcile all products utilizing the Product Recall Accountability Form. Questions should be directed to 201-972-2100. |  | 702 | 702 total, all sizes. | Nationwide Distribution. | Terminated | August 06, 2012 | 951 | 702
Z-0853-2010 | 53785 | No | DXH | Telephone electrocardiograph transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Cybernet Medical MedStar Telemedicine System, Manufactured by Cybernet Medical, Ann Arbor, MI; Model 533-380. |  |  | January 23, 2008 | 2008 | Cybernet Systems Corp. | A software problem renders these devices incapable of capturing lung function, oxygen saturation and pulse data collected from multiple tests in accordance with the instructions for use. | Component design/selection | The firm sent out a recall letter dated 10/27/09. Consignees were asked to return the devices in prepaid shipping containers. The recalling firm would do a software upgrade and return the device to the consignee. The firm expects it to take 3-6 months to upgrade all devices.Contact Cybernet Medically immediately, at 1-800-CYBERNET, with the serial number and code versions of the devices in question and Cybernet will start the RMA process for returning the devices for the updates, free of cost. |  | 24 | 24 | Nationwide distribution: California, Florida, Georgia, Illinois, Louisiana, Michigan, Mississippi, New York, Oklahoma and West Virginia. | Terminated | October 22, 2010 | 1003 | 88
Z-0850-2010 | 48978 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | COULTER LH SlideMaker, LH Barcode Specification, LH SM Intended Use, LH Slide and Sample Flow, Part Number 6605633, P/N 4277248, 4277299, and 277299. |  |  | May 29, 2008 | 2008 | Beckman Coulter Inc | Beckman Coulter has become aware of a potential issue in which the Sample ID information embedded in the barcode affixed to a slide made by the Coulter GEN*S and LH SlideMaker does not match the text. | Pending | An Urgent: Product Corrective Action letter, dated June 9, 2008, was sent to customers. The letter identified the affected products, stated the background, discussed the issue, and stated the actions to be taken. Customers are to: 1. Only use upper case alpha characters in the Sample Identifier2. If you must use lower case alpha characters in the Sample Identifier, configure the SlideMaker Slide labels without a bar-code. Questions regarding the Product Corrective Action should be directed to Beckman Coulter Customer Service at 1-800-526-7694 or your local Beckman Coulter Rep. |  | 843 | 843 | Worldwide Distribution -- USA and Canada. | Terminated | November 29, 2011 | 1279 | 847
Z-0858-2010 | 53026 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Flex Reagent Cartridge Cardiac Tropinin-I, Catalog Number: RF421C and RF521. |  |  | March 02, 2007 | 2007 | Siemens Healthcare Diagnostics, Inc. | Tests exhibit falsely elevated results. | Material/Component Contamination | The recalling firm issued a Field Correction letter dated 3/2/07 informing user of the problem and to contact the firm at 800-241-0420 for replacements. Customers are to follow the precautions provided in the letter. Questions should be directed to Dade Behring Technical Assistance Center at 800-441-9250. The notification should be forwarded to anyone to whom the product may have been further distributed to. |  | 9767 | 9767 cartons | Worldwide Distribution -- USA, Canada, Germany, Egypt, New Zealand, Saudi Arabia, Singapore, and BS. | Terminated | February 25, 2010 | 1091 | 9767
Z-0856-2010 | 53679 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxI 800 and 600 Access Immunoassay Systems, UniCel DxC 880i, 860i, 660i, and 680i SYNCHRON Access Clinical Systems with Dxl System Software versions 4.2 or 4.2.1. |  |  | October 28, 2009 | 2009 | Beckman Coulter Inc. | This recall was initiated because of software issue was identified in System Software 4.2 and 4.2.1, which may result in the incorrect handling of a wash arm dispense plate motion failure. Instead of generating a red warning event and stopping sample processing, the system will post a yellow caution event and continue to operate with the dispense plate in an elevated position until the instrument | Software design | Beckman Coulter issued an "Urgent: Product Corrective Action" letter dated October 28, 2009. Customers were notified of the affected product and how the issue will be resolved.For further information, contact Beckman Coulter Inc. Technical Support at 1-800-854-3633. |  | 377 | 377 Instruments | Worldwide Distribution -- United States and Canada. | Terminated | May 07, 2012 | 922 | 377
Z-0859-2010 | 53810 | No | LBR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid HDL (PEG) Cholesterol Reagent Set. Pointe Scientific HDL Cholesterol Reagent Set, Pointe Scientific, Inc, Canton, MI; Catalog # H7511-60 and Catalog # H7511-SAM. HDL Cholesterol Reagent Set, Pointe Scientific, Inc, Canton, MI; . (Clinical chemistry). |  |  | November 10, 2009 | 2009 | Pointe Scientific, Inc. | The firm initiated the recall due to the presence of visible contamination of the product. | Process control | Consignees were notified by an Urgent: Medical Device Recall letter (sent by email, fax or through the mails) dated 11/10/09. The letter identified the affected product and instructed customers to immediately examine their inventory. If the product was further distributed their customers should be identified and notified of the recall. Customers are instructed to dispose of any product in stock. Questions should be directed to the Technical Support Department at 800-757-5313 or 734-487-8300. |  | 379 | 379 sets | Worldwide Distribution -- USA, Bolivia, Ghana, Jamaica, Kenya, Korea, Malaysia, Mexico, Nicaragua, Nigeria, Philippines, Poland, Sri Lanka and Trinidad. | Terminated | February 11, 2011 | 458 | 625
Z-0857-2010 | 54364 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Cardinal Health EnVe Ventilator Cat # 19250-001, Manufactured by Cardinal Health, Minneapolis, MN 55447-1341. |  |  | January 08, 2010 | 2010 | CareFusion 203, Inc. | CareFusion is recalling the EnVe Ventilator because the EnVe Ventilator does not fully meet applicable standards for immunity to electro-Magnetic Interference (EMI) when tested in the laboratory environment. CareFusion has identified certain device malfunctions that may affect device performance and patient safety in the intended use environment. | Component design/selection | CareFusion issued an "Urgent Medical Device Recall notification" letter dated January 07, 2010. The letter was addressed to the Director of Respiratory Care describing the problem and the product. Consignees were told that a Service Representative will contact them to arrange for the prompt return of the affected product for correction.For further information, contact CareFusion Technical Support at 1-800-231-2446, Option 1. |  | 9 | 9 | MN | Terminated | December 17, 2011 | 708 | 9
Z-0870-2010 | 53800 | No | JXB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | ERIC (Electronic RapID Compendium) Version 1.0.77, Catalog Number R8323600, used in conjunction with: remel RapID Yeast Plus System is a qualitative micromethod employing conventional and chromogenic substrates for the identification of medically important yeast, yeast-like, and related organisms isolated from human clinical specimens. A complete listing of the organisms addressed by the RapID Yeast Plus System is given in the RapID Yeast Plus Differential Chart. |  |  | October 28, 2009 | 2009 | Remel, Inc | There is an error in Version 1.0.77 of the electronic code compendium (ERIC) software. The error is limited to the database corresponding to the R8311007 RapID Yeast Plus System. The database erroneously includes Candida dubliniensis as well as displays an incorrect temperature range as an accessory test for codes with an overlap between Candida albicans and Candida dubliniensis. When using the | Software design | The firm issued an "IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" dated October 29, 2009. After completion of the firm's investigation and risk assessment, Remel decided to notify customers of this situation and quarantine any remaining units on hand. The notification provided the consignees with instructions for obtaining and downloading the software patch. A response form was included with the recall letter and the consignee was asked to complete the form and return it to Remel.The consignees should contact the Technical Services Department at 800-255-6730 (USA) or 913-888-0939 (International), if they have inquiries concerning this advisory. |  | 143 | 143 units of software | Worldwide Distribution: USA, Europe and Canada | Terminated | March 03, 2010 | 126 | 143
Z-0864-2010 | 53323 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista System CHEM 1 CAL, Catalog Number KC110, Lot Number 9DM001 |  |  | August 31, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics, Inc. has confirmed high accuracy shifts with Dimension Vista URCA Flex reagent cartridge (various lots) when calibrating with some vials from CHEM 1 calibrator lot 9DM001. | Other | An Urgent Field Safety Notice dated August 31, 2009, was issued to all affected customers. Customers were instructed to discontinue the use of Dimension Vista System CHEM 1 Cal lot 9DM001 and discard any remaining product. If customers have any question they can contact Technical Solutions at 1-800-241-0420. |  | 1474 | 1474 Cartons | Nationwide and Foreign: Australia, Canada, Denmark, and Malaysia. | Terminated | December 13, 2010 | 469 | 1474
Z-0861-2010 | 53341 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Huber Plus Non-Coring Needle Safety Infusion Set, 19g, 1 inch, Needless Y-site, REF 011901NY, Sterile EO, Assembled In Mexico, Bard Access Systems, Inc., SLC, UT 84116. |  |  | January 14, 2010 | 2010 | Bard Access Systems | After activation of the safety mechanism, the tip of the needle protrudes from the safety sheath creating the potential for needlestick injury. | Component change control | Bard Access Systems, Inc. issued and "Urgent: Huber Plus Non-Coring Needle Safety Infusion Set Recall" notification beginning January 2010. Users were instructed to quarantine all affected product and to notify any of their customers that had received the product. Quarantined product is to be returned to the firm. For further information, contact Bard Access Systems, Inc. at 1-800-290-1689. |  | 4975 | 4,975 units | Worldwide Distribution -- United States (VA, KY, MA, NJ and TN) and Belgium. | Terminated | May 14, 2010 | 120 | 5750
Z-0865-2010 | 53483 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | i-STAT HL7 Interface |  |  | September 29, 2009 | 2009 | Abbott Point of Care Inc. | Abbott Point of Care has determined that it is possible under certain circumstances that some test records could have their status update messages mis-matched with another record, resulting in the incorrect patient name, order number of comment being included in the CDS test record only. | Software design | Urgent Recall Notification letters were sent by Fed Ex to all US customers on September 29, 2009 and worldwide on October 1, 2009. Questions should be directed to Abbott Point of Care Technical Support at 1-800-366-8010, Option 1. |  | 36 | 36 | Nationwide and Foreign: Germany, the United Kingdom, Switzerland and the Netherlands. | Terminated | July 01, 2010 | 275 | 36
Z-0868-2010 | 54237 | No | KRI | Cardiopulmonary bypass accessory equipment. | Cardiovascular | Cardiovascular | 510(K) Exempt | IDEAL MAXPAC TOTAL SYSTEM (USA) , Sterile and Nonpyrogenic, ETO or Gamma Sterilized, Medtronic Perfusion Systems, Model TS9453R29, TS9453R30, TS1470R1, TS1442R1, TS9800R32, TS8299R4. |  |  | December 14, 2009 | 2009 | Medtronic Inc | Medtronic has identified a small number of Autotransfusion One Source Kits and MAXPAC/Total System lots with incorrect Use Before Dates (UBD) on the outer kit package label. The UBD on the outer label of the kit/pack does not accurately reflect the UBD of the components within the kit/Pack. For affected product, the outer label lists a date which may be later than the USB of the components inclu | Incorrect or no expiration date | Consignees were sent a Medtronic "Urgent Medical Device Recall Notice" dated December 14, 2009. The letter was addressed to "Valued Customer." The letter described the product and the problem and requested them to quarantine the product and to complete and sign the attached certificate, "Medtronic Recall Certificate (FCA 1002)." For additional information, contact the firm at 763-514-9605. |  | 24 | 24 | CA, CO, JAPAN, CHINA. | Terminated | April 20, 2011 | 492 | 36
Z-0867-2010 | 53649 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Variant II Turbo Hemoglobin A1c Program, model number 270 2417, manufactured by Bio-Rad Laboratories, Inc, Hercules, CA. |  |  | October 12, 2009 | 2009 | Bio-Rad Laboratories Inc | Product update kit CD rom will not upload. As a result, the device cannot be run. | Other | The firm initiated this action on 10/12/2009. Consignees were contacted by phone with a fax follow up. Subsidiaries, as well as consignees that received the product, were provided with a Medical Device Correction notice, dated October 13, 2009, and response form. Each subsidiary will contact customers. The letter asks customers to destroy the affected lot included with their reorder packs and replace them with the lot number that's enclosed. Customers are to complete the Customer Medical Device Correction Response Form and fax it to Bio-Rad CSD Regulatory Affairs Department at the number provided. Questions should be directed to customers' regional Bio-Rad office. |  | 21 | 21 kits | Worldwide Distribution -- USA, Hong Kong, and New Zealand. | Terminated | December 17, 2010 | 431 | 21
Z-0866-2010 | 54355 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | EXXCEL Soft Standard Wall EPTFE Vascular Graft; |  |  | January 20, 2010 | 2010 | Maquet Cardiovascular, LLC | Labeling mix up: Shelf Carton label for Exxcel Soft Thin Wall Vascular Grafts, lot # 12633379 were mislabeled as Exxcel Soft Standard Wall Grafts. Bleeding could occur if wrong product is used. | Labeling mix-ups | Recall letters were sent to all US customers on January 20, 2010 by certified mail. Letters for foreign consignees were to be sent as soon as translation was completed. Customer are directed to call the company at 973-709-7515 for more information. |  | 18 | 18 units | Product was distributed to hospitals in TX, TN, CT, AR, AZ as well as to Austria and France. | Terminated | July 06, 2010 | 167 | 18
Z-0895-2010 | 54487 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Electri-Cord Manufacturing Corporation (312 E. Main St., Westfield, PA 16950) AC power cords used with either Bair PawsÂż Temperature Management Units: Models 850 and 875 |  |  | December 17, 2009 | 2009 | Arizant Inc | Arizant Healthcare Inc. is initiating a nationwide voluntary recall of 100/110-volt power cords (Arizant part number 502221 Model 90024) attached to its temperature management units and pressure infusion power pack (Bair PawsÂż Temperature Management Units: Models 850 and 875Bair HuggerÂż Temperature Management Units: Models 500/OR, 505 (Human &amp; Vet.), 750 (Human &amp; Vet) and 775RangerÂż Blood/Flui | Nonconforming Material/Component | An "Urgent Medical Device Power Cord Recall" letter dated December 17, 2009 was sent to customers. The letter was addressed to Recall Coordinator/Biomed Dept. The letter described the affected product, problem, hazard involved, and action to be taken by customer. The customers are being ask to inspect their temperature management units (as described in FDA's recommendation dated October 19, 2009) and contact the firm with the results of their inspection and provide a contact name via website www.arizant.com/powercords or phone (877) 947-1487. Replacement cords and instructions will be provided to customers based on priority listed in the letter and availability of replacement cords. If the customer does not have the means to make the changes, alternate means will be taken (i.e., an Arizant employee or outside service organizations may be brought in to do the replacements).If you have any questions, please contact the Aziant at (877) 47-1487. |  | 84537 | Total distributed for all devices: 84,537 US, 7,380 OUS. | Worldwide Distribution: USA and ARGENTINA, THE BAHAMAS, CANADA, COLOMBIA, COSTA RICA, CUBA, DOMINICAN REPUBLIC, ECUADOR, GERMANY, GUAM, JAPAN, LEBANON, MEXICO, NICARAGUA, PANAMA, PUERTO RICO, QATAR, SAUDI ARABIA, SOUTH KOREA, SURINAME , TAIWAN, TRINIDAD AND TOBAGO, QATAR. | Terminated | April 19, 2012 | 854 | 84537
Z-0871-2010 | 53768 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PTS Panels Cholesterol test strips for self-testing or professional use with CardioChek analyzers, Manufactured by Polymer Technology Systems, Inc., Indianapolis, IN.Catalog # 1711. |  |  | April 02, 2009 | 2009 | Polymer Technology Systems, Inc. | The memory chip was coded with incorrect product information. | Process control | Polymer Technology Systems, Inc. issued an "Urgent - Field Correction" letter dated April 2, 2009. Consignees were advised of the problem and asked to identify how many kits were in inventory so that the firm could replace product.For further information, contact Polymer Technology Systems, Inc. Customer Service at 1-877-870-5610. |  | 486 | 486 vials | Nationwide Distribution | Terminated | October 20, 2010 | 566 | 486
Z-0800-2010 | 54285 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice PRO Automated Peritoneal Dialysis Systems; Catalog Numbers; 5C8310, 5C8310R, R5C8320, R5C8320R, T5C8300, and T5C8300R. |  |  | January 08, 2010 | 2010 | Baxter Healthcare Renal Div | Potential overfill of the peritoneal cavity, also referred to as Increased Intraperitoneal Volume (IIPV). IIPV can result in serious injury or death. | Labeling design | Urgent Product Recall letters dated January 8, 2010 were mailed to all Peritoneal Dialysis Clinicians on January 8, 2010 and letters dated January 12, 2010 to all Home Care Coordinators and Home Dialysis Patients on January 12, 2010, informing them of steps to take to help reduce or eliminate overfill, also referred to as Increased Intraperitoneal Volume, associated with HomeChoice/Homechoice Pro cyclers. Baxter Healthcare Corporation explained that they had received complaints of IIPV. The accounts were requested to complete and return the enclosed response form, indicating that they received and understood the notification, and if they still had a HomeChoice System or not. Any questions from the Home Care Coordinators are to be directed to the Baxter Global Technical Services at 1-800-553-6898. |  | 32678 | 32,678 units | Worldwide Distribution -- United States, Austria, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Iraq, Ireland, Israel, Italy, South Korea, Kuwait, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Portugal, Qatar, Russia, South Africa, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Thailand, Taiwan, Tunisia, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | June 21, 2017 | 2721 | 253508
Z-0908-2010 | 49449 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | Rubella IgG EIA test kit, containing 1 Rubella IgG Microplate, 1 Negative Control, 1 High Positive Control, 1 Low Positive Control, 1 Substrate, 1 Stop Reagent. Labeled: Distributed by: Bio-Rad Laboratories-Redmond, WA 98052 |  |  | August 20, 2008 | 2008 | Bio-Rad Laboratories | Elevated optical density (OD) values with the kit Negative Control. | Employee error | An "Urgent Product Correction" letter dated August 26, 2008 was sent via overnight carrier to all customers. The letter described the affected product, problem, and action to be taken by customer. The customers are to review their current inventory of the product for any remaining kits of the specified lot number. The customer is to complete the form and fax to the Technical Support department if replacement kits were needed. The remaining kits will be replaced once the form is completed and faxed.If you have any questions, please contact Technical Support at 1-800-224-6723, options 2 then 4. |  | 82 | 82 kits | Nationwide distribution: NE, NY, FL, GA, MA, OR, WI, MN, MI, MS, NC, CA, MD, MO, IA, DC | Terminated | April 08, 2010 | 596 | 82
Z-0921-2010 | 48648 | No | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | IMMAGE Immunochemistry Systems C-Reactive Protein (CRP) Reagent, Part Number 447280. |  |  | January 10, 2008 | 2008 | Beckman Coulter Inc | IMMAGE CRP Reagent lots M705441, M709279, &amp; M709319, may demonstrate a prozone or hook effect on samples with extremely high CRP analyte concentration. | Other | Customers were sent Urgent: Product Corrective Action letters dates, 1/22/2008. The letter identified the issue, discussed the impact, and listed the following actions/resolution.* To eliminate the hook effect when using these lots, run CRP samples beginning with the 1:216 dilution. ~ A limited correlation of patient samples run in the 1:36 dilution compared to the same samples run in the 1:216 dilution verifies acceptable performance (there may be a shift of up to 0.8 mg/dL for samples between 1.0 and 8.0 mg/dL).* Controls run in the 1:216 dilution may recover differently than those run in the default 1:36 dilution. A limited number of control samples were also run in the 1:216 dilution.* The IMMAGE 800 system can be set to default to a non-standard dilution for each chemistry. Change CRP to a Non-Standard Dilution of 1:216 in the Setup section of the software. Detailed instructions for configuring a Non-Standard dilution as default are located in chapter 5 of the IMMAGE 800 Immunochemistry System Operations Manual.* The IMMAGE system does not allow resetting the default dilution of a chemistry. Instructions for running CRP samples beginning in a Non-Standard Dilution are enclosed.* Complete and return the enclosed response form within 10 days so we may maintain our records.Customers are to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If you have forwarded any of the affected product listed above to another laboratory, please provide a copy of this letter to them. Questions regarding this letter should be directed to 800-854-3633 or customers can contact their local Beckman Coulter Representative |  | 878 | 878 kits | Nationwide Distribution -- Including states of AL, CA, CT, DC, FL, FL ,GA, HI, IA, IL ,IN, KY ,MA, MD, ME, MI, MO, NE, NJ, NY, OH, OR, PA, RI, TX, VA, WI, and WV. | Terminated | July 02, 2012 | 1635 | 878
Z-0910-2010 | 50150 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon X3 UHMWPE Tibial Inserts and Patellar Components, multiple sizes, Stryker Orthopaedics, Mahwah, NJ. |  |  | February 17, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware of a potential anomaly related to the packaging process of referenced lots of X# polyethylene used on Scorpio, Triathlon and Trident inserts. | Process design | Stryker Orthopaedics issued an "Important Market Withdrawal" notification dated February 17, 2007 via federal express. Consignees were informed of the affected product and asked to retrieve, reconcile and return product via Return Material Authorization to the firm.For further information, contact Stryker Orthopaedics Quality Assurance at 1-201-831-5825. |  | 594 | 594 units total - all products. | Worldwide Distribution -- United States and Canada. | Terminated | March 03, 2010 | 1475 | 594
Z-0914-2010 | 50703 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Peristaltic Pump component of: UniCel DxI 800 Access Immunoassay Systems, UniCel DxI 600 Access Immunoassay Systems, UniCel DxC 880i SYNCHRON Access Clinical Systems. |  |  | October 09, 2008 | 2008 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports of a new premature failure of the upper aspirate peristaltic pump in the systems listed above. Premature failure of the pump may negatively affect precision. The failure is limited to those UniCel DXI 800, DXI 600, and DxC 880i systems which have a pump identified with one of the new pump manufacturing numbers (Mfg. Nos.) listed ab | Other | Beckman Coulter contacted consignees via "Urgent Product Corrective Action" letter dated October 16, 2008 to inform them of the recall and to arrange a date to service the affected product as soon as possible. For further information, contact Beckman Coulter: (a)In the U.S.A. or Canada, contact Technical Support at 800-854-3633 or online at www.beckman.com/customersupport. (b) Outside the U.S.A. and Canada, contact your technical support representative. |  | 107 | 107 (US), 7 (OUS) | Worldwide Distribution -- United States and Canada. | Terminated | May 04, 2012 | 1303 | 107
Z-0885-2010 | 54228 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Mistique Infusion Catheter, MIC5-135-20-038, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095 |  |  | January 06, 2010 | 2010 | Merit Medical Systems, Inc. | Infusion catheters may become brittle and break during use. | Storage | All affected sales reps were notified by phone and a Mistique Infusion Catheter - Product Recall letter on 01/06/2010 and instructed to contact their accounts, inform them of the recall situation, and complete the Recall Notification form with a site representative. Accounts were to immediately cease use of any affected units and to arrange for return to Merit. Recall actions are to be completed by Jan 22, 2010. Questions should be directed to Greg Turner at 801-316-4998 or Liz Pratt at 801-208-4828. |  | 18 | 18 uinits | Worldwide Distribution -- USA, including states of AL, CA, IN, LA, MO, MS, NC, NY, OH, PA, SC, UT, and VA and countries of India, Japan, Poland, San Marino, and UK. | Terminated | September 13, 2010 | 250 | 1389
Z-0907-2010 | 54258 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion One Whole Body X-Ray CT Scanner |  |  | January 07, 2010 | 2010 | Toshiba American Medical Systems Inc | When W-Volume is performed in Prospective CTA or Target CTA studies with ECG gating, unnecessary intervals may be set before starting the volume scan. As a result, scanning cannot be performed at the appropriate timing of constant enhancement and clinically useful images may not be obtained. | Device Design | A corrective measure will be installed to prevent measures to avoid this phenomenon until corrective actions have been taken.Customers are advised: (1) When performing scanning in Prospective CTA or Target CTA studies with ECG gating, do not use W-Volume scanning but use Helical Scanning. (2) Please share this information with all users and reviewing radiologists as well as clinical engineering or Biomedical group at facility (3) Please complete and return the attached form and fax it to the toll-free number at the top or by email to raffairs@tams.com. Contact: Paul Biggins |  | 28 | 28 units | Nationwide: AL AR AZ CA CO FL GA IA MA MD ND NJ NV NY OH PA PR andVA | Terminated | September 12, 2012 | 979 | 28
Z-0899-2010 | 53872 | No | JIY | Iron (non-heme) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific Iron/TIBC Reagent Set, for the quantitative determination of serum iron and iron binding capacity, Pointe Scientific, Inc, Canton, MI, Catalog #s: I7504-60, I7504-400, I7504-SAM, 10-I7504-60. |  |  | November 16, 2009 | 2009 | Pointe Scientific, Inc. | Turbidity formation could occur, resulting in invalid test results being reported out. | Device Design | Consignees were notified by recall letter dated 11/16/09 and instructed to discard any remaining product. Please call 1-800-757-5313 if you have questions. |  | 417 | 417 sets | Nationwide, Australia, Austria, Bolivia, Canada, Cyprus, Egypt, Ghana, Greece, Guyana, India, Indonesia, Israel, Jordan, Kenya, Korea, Lebanon, Macedonia, Malaysia, Mexico, Nicaragua, Nigeria, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Trinidad, Turkey, Ukraine and Venezuela. | Terminated | June 06, 2011 | 567 | 1480
Z-0915-2010 | 54359 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NextGen Knee Gender Solutions CR-Flex Femoral Component, Precoat , Size D, Left. 00-5750-014-01 Zimmer, Inc. Warsaw, IN. |  |  | January 27, 2010 | 2010 | Zimmer Inc. | The package should contain the left knee component but actually contains the right knee component. Delay in completing surgery may result, and a new or different implant may need to be identified. | Labeling mix-ups | The firm sent an immediate notification by e-mail to all first consignees on 1/12/2010. They sent a formal recall notification on 1/27/2010. Customers should contact their Zimmer Sales representative with additional questions. |  | 14 | 14 units | US and Korea. | Terminated | October 17, 2010 | 263 | 14
Z-0913-2010 | 54112 | No | KNK | Uric acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific Liquid Uric Acid (R1) Reagent Set, Pointe Scientific, Inc., Canton, MI; Catalog # HU482-R1. |  |  | December 01, 2009 | 2009 | Pointe Scientific, Inc. | A process deviation occurred during the production of the R1 component which may affect its performance. | Device Design | Pointe Scientific, Inc. issued an "Urgent: Medical Device Recall" notification dated December 1, 2009. Consignees were instructed to examine their inventory for the affected product and discard the affected product immediately.For further information, contact Pointe Scientific, Inc. Technical Support at 1-800-757-5313 or 1-734-487-8300. |  | 1 | 1 kit | Worldwide Distribution -- United States (California, Florida, Kentucky, Michigan, Ohio, Oklahoma and Virginia), Chile and Malaysia. | Terminated | October 08, 2010 | 311 | 13
Z-0922-2010 | 54124 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | DiaSorin Rubella IgG ELISA Kit, Cat. # 7545, 2 x 6 Wells, DiaSorin Inc., 1951 northwestern Avenue, Stillwater, MN 55082-0285, Lot 123127A, Exp. 2010//07/28. |  |  | December 16, 2009 | 2009 | Diasorin Inc. | DiaSorin Rubella IgG kits of lot 123127A were packed with one coated microtiter plate instead of two microtiter plates as indicated in the product labeling. The risk is the possibility of delayed patient test results. | Employee error | Consignees were sent a DiaSorin Customer Notification letter dated December 16, 2009. The Letter was addressed to Customer. The letter described the product and the problem and indicated a replacement product will be provided. Requested consignees to complete the acknowledgement receipt form. For questions or concerns please contact DiaSorin Inc., Technical Services at 1-800-328-1482 or 651-439-9710. |  | 10 | 10 kits | US distribution: CA, CT, MA, MD, IL, OR | Terminated | March 29, 2012 | 834 | 10
Z-0916-2010 | 53431 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Herculite Ultra (Part Number 34333), dental composite resorative material. Manufactured by Kerr Corporation, Orange, CA. |  |  | February 23, 2009 | 2009 | Sybron Dental Specialties | Sybron Dental Specialties is initiating a field correction on the Herculite Ultra, a dental composite restorative material, due to a mistake in the directions for use located in the Herculite technique card. | Labeling mix-ups | Kerr Corporation issued a letter dated February 23, 2009. Users were informed of the affected product and instructed on proper return.For further information, contact Kerr Customer Care at 1-800-537-7123. |  | 442 | 442 | Nationwide distribution | Terminated | March 19, 2012 | 1120 | 442
Z-0919-2010 | 54360 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | BCI 3180 Pulse Oximeter uses power cord WW3005 (Electri-Cord Manufacturing Co., 312 E. Main St., Westfield, PA 16950). |  |  | January 15, 2010 | 2010 | Smiths Medical PM, Inc. | Smiths Medical is conducting a voluntary recall of a limited number of medical device AC power cords, manufactured by Electri-Cord Manufacturing Co., for use with BCI 3180 Pulse Oximeter(Human), 9200 Advisor Vital Signs Monitor(Human or Veterinary), 8400 Capnocheck II Capnometer(Human or Veterinary) and V6400 Invasive Pressure Monitor (Veterinary).The affected Electri-Cord AC power cords are eq | Nonconforming Material/Component | Consignees were sent a Smiths Medical PM, Inc letter dated January 12, 2010 starting on 1/15/10. The letter was addressed to Risk/Safety Managers, Distributors, Clinicians, Critical Care Physicians, Pediatricians, Pediatric Intensivists, Neonatologists, and Nursing, Clinical/Biomedical Engineering, and Anesthesia Professionals. The letter described the problem and products involved. It provided "Advice on Action to be Taken by the User" and requested the return of the "Urgent Recall Notice Confirmation Form." Customers are advised to monitor affected equipment. To return Power Cords, for questions regarding this Urgent Recall Notice, or to report any issues with these Power Cords, please contact the Smiths Medical Technical Service Department at 1-800-558-2345. |  | 1358 | 1358 cords for all devices (1032 US, 326 OUS) (871 Human, 487 Vet), (Cord # 8404 5 Vet &amp; 325 Human), (Cord # WW3005 482 Vet &amp; 546 Human) | US distribution: AL, AZ, AR, CA, CO, CT, DE, FL, GA, HI, IL, IN, IA, KS, KY, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, WI, WY.Worldwide distribution: Trinidad and Tobago, Mexico, Taiwan, Ecuador, Norway, Colombia, Brazil, France, Japan, Canada. | Terminated | March 14, 2012 | 789 | 1358
Z-0917-2010 | 53480 | No | KIF | Root canal filling resin. | Dental | Dental | 510(k) | Sealapex Express (Part Number: 33639), root canal filling material. Manufactured by SybronEndo, Glendora, CA. |  |  | February 23, 2009 | 2009 | Sybron Dental Specialties | Sybron Dental Specialties is initiating a recall on the Sealapex Xpress due to the presence of crystals in the catalyst. This crystallization may affect proper canal filling by preventing Gutta Percha Points from reaching the bottom of the root canal which may lead to an inadequate seal. | Process control | SybronEndo initiated an "Urgent Medical Device Recall" notification dated February 19 , 2009. Consignees were instructed to complete a "return form" and return any affected product in their inventory to the firm. Any product returned will be replaced at no charge.For further information, contact SybronEndo Customer Care at 1-800-346-3636. |  | 1309 | 1309 units | Worldwide Distribution -- United States, Australia, Canada, United Arab Emirates, Belgium, Bulgaria, Germany, Spain, France, UK, Greece, Italy, Kuwait, Lebanon, Portugal, Russia, Sweden, Hong Kong, India, Malaysia and Taiwan. | Terminated | March 21, 2012 | 1122 | 1309
Z-0968-2010 | 52901 | No | NGE | Luteinizing hormone test system. | Clinical Chemistry | Toxicology | 510(K) Exempt | OvuChek Saliva Ovulation Disposable Test Strips; 5 disposable test strips per pouch, packaged in a box with a specimen cup and instructions for use, 12 boxes per case; a rapid colorimetric indicator for the qualitative changes (ferning) of salt content in saliva to aid in the prediction of ovulation; Distributed by Quest Products, Libertyville, IL |  |  | November 07, 2008 | 2008 | Quest Products Inc. | The ovulation test strips were shipped without 510(k) premarket notification clearance from the FDA. | Other | Quest Products notified their major customer by telephone and e-mail on 11/7/08, requesting the return of all inventory of the OvuChek Saliva Ovulation Test Strips due to a lack of 510(k) clearance for marketing of the device. |  | 3930 | 3,930 cases | Nationwide | Terminated | March 03, 2010 | 481 | 3930
Z-0969-2010 | 54152 | No | MKZ | Chlamydia serological reagents. | Microbiology | Microbiology | 510(k) | BD Probetec" ET CT/GC Endocervical Specimen Collection and Dry Transport System-Pink Cap, Catalog number 440476, packed in shelfpack unit of 100 swabs. |  |  | September 18, 2009 | 2009 | Becton Dickinson &amp; Co. | Product was shipped beyond its expiration date. | Employee error | The recalling firm notified end users by letter flagged as "Urgent Product Recall" on 09/18/09. The notification advised that the referenced lot number had been correctly labeled with expiration date but had been shipped after the labeled date. End users were instructed to discontinue use discard product for replacement and return enclosed response form. If further assistance is needed regarding replacements, please contact SD Customer ServiceDepartment at 1-800-675-0908. For all other inquiries, please contact SD Technical Services Department at 1-800-638-8663. |  | 242 | 242 packs | The product was sold for distribution to clinical laboratories nationwide and to the firm's international affiliate in Australia. | Terminated | March 24, 2010 | 187 | 242
Z-0954-2010 | 54269 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | JOURNEY " UNI TIBIAL INSERT RIGHT MEDIAL/ LEFT LATERAL, A/P 52 MM, M/L 30 MM, REF 71422265, QTY: (1), STERILE EO, SMITH &amp; NEPHEW, INC. MEMPHIS, TN 38116 |  |  | January 23, 2009 | 2009 | Smith &amp; Nephew Inc | The firm has received complaints of the base plates breaking in situ. If a breakage occurs, the broken baseplate must be revised and/or changed. If left unrevised a broken base could lead to instability and premature wear of the construct. | Device Design | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall overnight mail on 12/23/2009. All affected Smith &amp; Nephew International Distributors were notified via e-mail and telephone on 01/23/2009. They were instructed to immediately quarantine the product for return. For questions, please contact the firm, Leah Easley at 901-399-6137. |  | 1344 | 1344 units | US and Argentina, Switzerland, Australia, Canada, Belgium, Italy, Spain, Thailand, Austria; PROC, Finland, Germany, Malaysia, Portugal, Netherlands, Singapore, and England | Terminated | March 25, 2011 | 791 | 36181
Z-0923-2010 | 54381 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Orthopedic Dovetail Clamp, REF 9730289, Medtronic Navigation, Louisville, CO 80027. |  |  | November 03, 2009 | 2009 | Medtronic Navigation, Inc | Orthopedic clamp may slip if used with non-indicated system, resulting in possible inaccurate navigation information being displayed. | Error in labeling | All affected users were notified by a Product Correction Notification letter on 11/03/2009 providing a description of the issue and informing them of the updated instructions for use for the Orthopedic Dovetail Clamp. Customers are to ensure that all users read the letter, place a copy of the letter with the IFUs for the related software applications listed in the letter, and fill out the attached response form indicating the receipt of the information. Further information is available at 720-890-3309. |  | 244 | 244 units | Worldwide Distribution -- USA, including VA, IA, MN and OK and countries of Australia, Austria, Belarus, Belgium, Brazil, Canada, China, Croatia, Czech Republic, Denmark, France, Germany, Hong Kong, Iceland, India, Israel, Italy, Japan, Korea, Lebanon, Malta, Mexico, Netherlands, Poland, Russian Federation, Saudi Arabia, South Africa, Switzerland, Turkey, United Arab Emirates, and UK. | Terminated | August 13, 2010 | 283 | 244
Z-0978-2010 | 54068 | No | JTO | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BBL(tm) Lactophenol Cotton Blue Stain Droppers, Catalog number 261188, shelf package of 50 droppers, incorrectly labeled in part ***Methylene Blue Loeffler***Becton Dickinson and Company***Packaged in Mexico*** |  |  | September 28, 2009 | 2009 | Becton Dickinson &amp; Co. | An in-vitro diagnostic reagent used by clinical laboratories was mislabeled. | Labeling mix-ups | An "URGENT PRODUCT RECALL" notice dated September 2009 was issued via UPS letter to customers. The notification letter described the product, problem and the action to be taken by customer. The recalling firm notified distributors and end users to discontinue use of the identified product and discard for replacement. The customers were requested to return an enclosed response form. The customers were requested to provide, within one week, an excel file or a list with name, address and phone number of end-users who were shipped the product from their distribution branches so that the BD may inform them of the issue.For further assistance regarding product replacement contact our Customer Service Department at 1-800-675-0908. For all other inquiries call our Technical Services Department at 1-800-638-8663. |  | 100 | 100 shelf packs | Worldwide distribution: USA and Puerto Rico, Korea, and Singapore. | Terminated | March 24, 2010 | 177 | 100
Z-0974-2010 | 53560 | No | JCF | Lymphocyte separation medium. | Hematology | Hematology | 510(K) Exempt | Sigma-Aldrich Histopaque-1077, Ref. 10771, packaged in 500-ml. btls. The firm name on the label is Sigma-Aldrich, Inc., St. Louis, MO. |  |  | September 04, 2009 | 2009 | Sigma-Aldrich Mfg LLC | Some of the product was contaminated with Pseudomonas mendocina | Other | An "IMPORTANT PRODUCT NOTIFICATION" dated September 4, 2009 was sent to customers via regular mail. The notification letter described product, problem and action to be taken. The customers were requested to immediately discontinue use of the product and discard all remaining material. Subrecall was requested if the product was further distributed. The customer is to complete and fax the notification form to (314) 286-7829 to assure proper replacement for material they have not used and which is to be discarded.Should you have any questions, please contact Signma-Aldrich Technical Service at (800) 325-5832. |  | 6808 | 6,808/500-ml. bottles | Worldwide distribution: USA and Great Britain, France, China, Germany, Australia, Canada, Mexico, Korea, Singapore, India, Japan, Colombia, Hong Kong, Peru, Taiwan, and Venezuela. | Terminated | April 10, 2012 | 949 | 6808
Z-0775-2010 | 53827 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | GE Responder 2019198 automated external defibrillator. This semi-automatic model requires the user to press a shock button in order for the device to deliver defibrillation energy (when appropriate). |  |  | November 13, 2009 | 2009 | Cardiac Science Corporation | Potential for devices not to deliver therapy | Process control | Cardiac Science issued an "Urgent - Voluntary Medical Device Correction AED Resistor Issue" letter dated November 13, 2009. Consignees were informed of the affected product and instructed on preventive measures.For further information, contact Cardiac Science at 1-877-901-1788 in the US and +44-161-926-0011 for outside the US. |  | 4518 | 4518 | Worldwide Distribution -- UNITED STATES, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BELARUS, BELGIUM, BERMUDA, BOLIVIA, BRAZIL, BULGARIA, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, CZECHOSLOVAKIA, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, GRECE, GUAM, GUATEMALA, HOLLAND, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, LEBANON, LIBYA, LITHUANIA, MALAYSIA, MALTA, MEXICO, MONACO, MORROCCO, NAMIBIA, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SCOTLAND, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD &amp; TOBAGO, TURKEY, UAE, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM, VIRGIN ISLANDS AND YEMEN. | Terminated | October 24, 2014 | 1806 | 288468
Z-0973-2010 | 53750 | No | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | "K-ASSAY CRP (3) Calibrator E". |  |  | May 01, 2009 | 2009 | Kamiya Biomedical Company, LLC | The long-term stability of level 2 (1.0 mg/dL) may not be within the specifications due to a decrease in CRP concentration. | Other | On May 1, 2009 the recalling firm, Kamiya Biomedical Company contacted the 3 consignees by telephone and notified them of the issue. A follow-up e-mail notification on May 4, 2009 was sent to consignees requesting to return the affected product to Kamiya Biomedical Company for a free replacement kit. After the e-mail notifications, an Urgent Medical Device Recall Notice dated May 13, 2009 was also sent to consignees describing the issue and actions to take for returning opened and unopened affected product to the recalling firm for replacement. Direct questions to the Kamiya Biomedical Company by calling 1- 206-575-8068. |  | 15 | 3 kits each containing 5 x 2 ml vials | United States (NY, NJ, and NC). | Terminated | January 21, 2011 | 630 | 15
Z-0824-2010 | 54167 | No | DSQ |  |  | Cardiovascular | PMA | Abiomed AB5000 Circulatory Support System, Catalog Number: 0015-0000. Manufactured by Abiomed, Inc. Danvers, MA. |  |  | December 16, 2009 | 2009 | Abiomed, Inc. | Console shut down with no audible alarm | Process control | Abiomed initiated the recall by telephone contact and followed with a Recall Letter dated January 27, 2010. Consignees were informed that a Field Service Engineer will be in contact and schedule a time to visit and correct the affected device.For further information, contact Abiomed Quality Assurance at 1-978-656-1543 or Abiomed Field Service at 1-800-554-8666. |  | 496 | 496 units | Worldwide Distribution -- United States, Australia, Brazil, China, Hong Kong, Japan, Mexico and Germany. | Terminated | July 02, 2014 | 1659 | 496
Z-0976-2010 | 53956 | No | JCF | Lymphocyte separation medium. | Hematology | Hematology | 510(K) Exempt | Sigma-Aldrich Histopaque-1077, Ref. 10771, packaged in 100- and 500-ml. btls. and a 6x100-ml. bottle package. The firm name on the label is Sigma-Aldrich, Inc., St. Louis, MO. |  |  | December 08, 2009 | 2009 | Sigma-Aldrich Mfg LLC | Instructions for Use (IFU) was revised to recommend sterile filtering of the product under certain circumstances. | Other | The recalling firm issued an "Urgent Field Correction" notice dated December 7, 2009 via regular mail on December 8, 2009 which explains the reason for the field correction and instructs the user to review the enclosed, revised Instructions for Use (IFU) prior to further use of the product. The customer is requested to complete the acknowledgement information and return it to the recalling firm via FAX or e-mail. Direct questions to Sigma-Aldrich by calling 1-800-505-6275, extension 1232. |  | 63281175 | 11,015/100-ml. btls; 5,745/6/100-ml btl packages; and 14,872/500-ml. btls. | Worldwide Distribution -- United States, Canada, Mexico, United Arab Emirate, Australia, Bolivia, Brazil, Chile, China, Colombia, Costa Rica, Germany, Dominican Republic, Ecuador, France, Great Britain, Hong Kong, Israel, India, Italy, Jamaica, Japan, Korea, Panama, Peru, Paraguay, Saudi Arabia, Singapore, Thailand, West Indies, Taiwan, Venezuela, and South Africa. | Terminated | April 10, 2012 | 854 | 90393495
Z-0807-2010 | 54365 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | StatSpin Express 4 Horizontal Centrifuge |  |  | January 27, 2010 | 2010 | StatSpin, Inc., dba Iris Sample Processing | Centrifuge shield micro-switch failed due to damage, and the unit opened, ejecting pieces. | Device Design | Iris Sample Processing notified distributors by an Urgent Device Recall letter advising of the recall and the defect. Users are to immediately quarantine the affected products and choose between the two options listed in regards to correcting the affected centrifuge. Customers may return the unit or an upgrade kit with instructions will be sent to upgrade the unit. Questions should be directed to Gail Magrath at 800-782-8774 ext. 6123. |  | 1510 | 1510 units | Worldwide Distribution -- USA, CANADA, AUSTRALIA, JORDAN, UK, BELGIUM, FRANCE, CZECH, RUSSIA, KUWAIT', NORWAY, GREECE, TURKEY, DENMARK, FINLAND, SEOUL, KOREA, and HONG KONG. | Open, Classified |  |  | 1510
Z-0797-2010 | 54356 | No | ETD | Tympanostomy tube. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | MICRONÂż Âż BOBBIN VENT TUBE, 1.27 MM I.D., TITANIUM, REF 145281-ENT, QTY 1, STERILE EO, GYRUS ACMI INC., 136 TURNPIKE ROAD, SOUTHBOROUGH, MA 01772 USA. |  |  | January 08, 2010 | 2010 | Gyrus ACMI Corporation | Units of the product in lot MH136952 may have been shipped without being sterilized. | Pending | The firm initiated their recall on 01/04/2010 by telephone or visit to their consignees. A follow-up Urgent: Medical Device Recall correspondence, dated 1/14/2010, was sent on 01/27/2010, return receipt requested, to both the consignees and the implanting surgeons. The follow-up letter identified the affected product and described the issue. Customers are to immediately cease any further use of any affected product and quarantine it. Also, they are to obtain a Return Goods Authorization and return the affected product to the firm. The reply form should be completed and also returned. Questions should be directed to customer service at 1-800-773-4301. |  | 114 | 19 boxes of 6 units | Nationwide Distribution -- Including states of FL, GA, OH, TX, RI, IL, MI, and KS. | Terminated | July 27, 2010 | 200 | 114
Z-0863-2010 | 54344 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter UniCel DxC Synchron Clinical System ISE Flow Cell. |  |  | January 04, 2010 | 2010 | Beckman Coulter Inc. | Excessive Buildup of protein, bacteria and sample tube additives in the ISE flow cell may cause erroneous NA (sodium) results. | Device Design | An Urgent: Product Corrective Action letter dated January 4, 2010, was mailed to customers advising of erroneous sodium results derived from build up of protein, bacteria and additives in the ISE flow cell. The letter discussed the impact and actions to be taken along with some additional recommendations. Customers are advised to follow the new maintenance instructions included with the letter. The information in the letter should be shared with laboratory staff and the notification should be retained as part of the Quality System documentation. A copy of the letter should be provided to any other laboratory that affected product was forwarded to. The enclosed response form should be completed and returned. Questions should be directed to the Customer Support Center at 1-800-854-3633. |  | 2630 | 2630 | Worldwide Distribution -- USA and Canada. | Terminated | April 17, 2012 | 834 | 2630
Z-0815-2010 | 54525 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Transseptal Sheath Introducer Kits (under trade names: HeartSpan, Channel FX, Torflex, Braided Guiding Introducer Kit). The product is shipped 5 per carton. |  |  | February 01, 2010 | 2010 | Thomas Medical Products Inc | Radiopaque tip may fracture | Component design/selection | Thomas Medical Products, A GE Healthcare Company, issued an "Urgent Medical Device Correction" letter dated February 1, 2010. Consignees were informed of the affected product then asked to cease distribution, return any product to firm and contact all sub-accounts.For further information, contact Thomas Medical Products Customer Service at 1-866-446-3003. |  | 13973 | 13,973 units | Worldwide Distribution -- United States (CA, MA and NJ), Canada, Germany, and the United Kingdom. | Terminated | January 05, 2011 | 338 | 13973
Z-0788-2010 | 53954 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Multi-Lumen Central Venous Catheterization Kit with Blue Flex-Tip(R), ARROWg+ard Blue PLUS (R) Catheter, Sharps Safety Features and Maximal Barrier Precautions. (contains Q-Syte) |  |  | November 24, 2009 | 2009 | Arrow International Inc | Q-Syte Component contained in kits was recalled by manufacturer due to the potential for embolism if used with a central venous catheter. | Nonconforming Material/Component | The recalling firm issued an Urgent Medical Device Recall letter dated 11/24/09 to their customers informing them of the problem the need to return any kits in inventory and notify their customers of the recall. For additional information customers can contact the Customer Service Department at 800-523-8446. |  | 180 | 180 kits | The products were shipped to distributors in GA, MD, and VA. | Terminated | January 18, 2012 | 785 | 608
Z-1001-2010 | 50691 | No | LJB | Rubeola (measles) virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Measles IgM ELISA 96 Well Kit, Catalog Number: MS019M |  |  | August 27, 2008 | 2008 | Calbiotech Inc | The recall has been initiated due to the "For Export Only" labeling found on some Calbiotech, Inc. ELISA kits. Although there is no potential health hazard related to this recall, Calbiotech is fully committed to compliance with all State and Federal regulations. | Labeling False and Misleading | Immediately examine your inventory and quarantine product subject to recall. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.Please return any product you currently have in stock which is subject to this recall and noted on the attached list. Any product you return will be replaced free of charge. Please use our FedEx account: 2284-6203-9 and send any packages via "FedEx Saver" service.This recall should be carried out to the wholesale level in order to comply with Federal regulations.Please complete and return the enclosed response form as soon as possible. If you have any questions, please contact David Barka at 619-660-6162. |  | 3 | 3 Kits | FL, and VA | Terminated | March 05, 2010 | 555 | 651
Z-1015-2010 | 51983 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Reference Electrode Membranes for use with blood gas analyzer ABL 800 FLEX, when used on ABL 827 and ABL 837 with Creatinine modules (8x7). The device is packaged four (4) units per box. |  |  | March 26, 2009 | 2009 | Radiometer America Inc | Analysis of customer data, (for ABL 8X7 customers) has shown that the reference membranes are out of analytical specifications when replaced as specified in the device labeling. These errors in analytical specifications could lead to inadequate or clinical interventions (therapy) or misleading interpretation of patient status in clinical settings. | Other | The recalling firm contacted their customers and notified them of this recall with a letter on 3/26/2009. The letter informs the customer that Radiometer has determined that the Reference Membrane D711 for use in the ABL8x7 Flex Creatinine Analyzer has a shorter in-use lifetime than the previously recommended 1-month interval listed in the Operators Manual. It states that consequently, the firm has changed the recommended replacement intervals as they correspond to the number of samples per day analyzed by the customer and a revised insert is being supplied to the customer which should be placed in the Operators Manual to reflect this new information. The recall notification is also accompanied by an attached response FAX FORM which the customer is asked to complete and to fax back to Radiometer so as to acknowledge receipt of the recall notification letter and the revised insert.Please direct questions to the company TechLine at 1-800Âż736Âż0600 opt 2. |  | 50 | 50 membranes in U. S. distribution | The recalled product was distributed customers located in the following states within the U.S.: CT, FL, GA, IL, MA, MD, MN, MO, NC, NM, OH, PA, VA, WA, WI, and WV. Product was also sold in the following countries: Algeria, Australia, Austria, Bahrain, Brazil, Czech Rep., Denmark, England, Estonia, France, Germany, Greece, Hungary, India, Italy, Japan, Netherlands, Norway, Poland, Russia, Spain, Sweden, Switzerland &amp; Turkey. | Terminated | May 09, 2012 | 1140 | 50
Z-1005-2010 | 52101 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Whatman Body Fluid Collection Paper 903, Ref: 10534790, for IVD use. Whatman Inc., Florham Park, NJ. |  |  | May 19, 2009 | 2009 | Whatman, Inc. | '903' newborn Neonate screening cards have duplicate numbers assigned on cards and may result in cross-over of patient test results | Process control | Whatman issued an "Urgent Medical Device Correction " letter on May 19, 2009. The consignee was instructed to remove cards with duplicate numbers. For further information, contact Whatman Inc. Quality Assurance at 1-207-459-7557. |  | 11300 | 11,300 Cards | Nationwide Distribution -- CA. | Terminated | September 28, 2015 | 2323 | 11300
Z-1014-2010 | 52851 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pointe Scientific Direct Bilirubin Reagent Set, Pointe Scientific Inc., Canton, MI., Catalog # HB936-1250. |  |  | July 29, 2009 | 2009 | Pointe Scientific, Inc. | R2 reagent component may fail because it may contain R1 reagent. The reagent was rendered inoperative and would not pass the required quality control (calibration and quality control testing) before beginning running the bilirubin assay. | Error in labeling | Consignee was notified by letter dated 7/29/2009 and instructed to identify and notify them at once of this product recall. Product is to be returned or destroyed. Customers with questions should contact the firm directly. |  | 2 | 2 | California. | Terminated | December 06, 2010 | 495 | 2
Z-1012-2010 | 53246 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS Panels CHOL+HDL+GLU Panel test strips for professional use with CardioChek P-A analyzers, Manufactured by Polymer Technology Systems, Inc., Indianapolis, IN. |  |  | August 04, 2009 | 2009 | Polymer Technology Systems, Inc. | Possible under recovery of cholesterol, HDL and glucose. A loss of recoverable cholesterol and HDL cholesterol is a manageable to moderate health risk. The clinical applications analysis shows that lot 1891 shows such a loss of activity. A loss of recoverable Glucose measurement is a moderate to high risk, based upon the degree of under-recovery, and the frequency of occurrence. | Other | Consignees were notified by letter dated 7/28/09 to discontinue use of, and to destroy, the reagent and to contact the firm for replacement product. For questions, please contact the company Customer Service toll-free at 877-870-5610 or + 1-317-870-5610 (outside the US). |  | 962 | 962 | Nationwide. | Terminated | October 20, 2010 | 442 | 962
Z1006-2010 | 53708 | No | NPR | Factor V Leiden DNA mutation detection systems. | Hematology | Hematology | 510(k) | FACTOR II (PROTHROMBIN) G20210A Kit, LightCycler 2.0 Instrument. Material Number 03610195 001 |  |  | March 12, 2009 | 2009 | Roche Molecular Systems, Inc. | Within the Precautions and Warning Section, Polymorphism heading, of the Product's Package Insert, it stated that mutations at positions 20207, 20209, 20218 and 20221 exist and these mutations are spanned by the mutation probe. These rare mutations will lead to an unknown result after performing genotyping. | Labeling design | Product Advisory Notices were sent via e-mail on 3/12/09 to Roche Local Safety Officers and Roche county General Manager. Local Safety Officers and General Managers were to communicate to the end users with the Important Product Notice. In addition, Reagent Bulletin 09-052, dated 3/27/09, was issued to Roche U.S. Affiliate of U.S. customers. |  | 9624 | 9624 kits | Nationwide Distribution | Terminated | June 09, 2010 | 454 | 9624
Z-1013-2010 | 52719 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio-Rad VARIANT II Hemoglobin A1c Program Recorder Pack with Calibrator/Diluent Set, Model Number 270-2101NU Manufactured by Bio-Rad Laboratories, Hercules, CA |  |  | June 15, 2009 | 2009 | Bio-Rad Laboratories Inc | Calibration Failure-- due to high intercept values, the device fails calibration and users are unable to run samples. | Other | Recall initiated June 15, 2009, with notification initially conducted by phone with Fax Notification follow up domestically. Each subsidiary will be given an urgent notice, subsidiary response form, customer notification, customer acknowledgement form. Each subsidiary is to contact customers and translate notification and response form as needed. Contact the recalling firm for additional information. |  | 613 | 613 units | Product was distributed to 127 customers throughout the US and internationally to Argentina, Australia, Canada, Chile, Czech Republic, France, Germany, Hong Kong, Hungary, India, Italy, New Zealand, Poland, Singapore, Sweden, Taiwan, and the United Kingdom | Terminated | February 02, 2011 | 597 | 613
Z-0980-2010 | 54039 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista LOCI Reaction Vessel |  |  | November 24, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Some vessels may leak. If leakage occurs, patient results will be flagged and are not reportable. | Nonconforming Material/Component | The recalling firm issued an Urgent Field Safety Notice dated 11/2009 to all affected customers. Customers have been told to discard any bags of the affected lot. If the affected lot is the only lot the customer has, customers can examine a vessel from the bag. If the vessel mold cavity identity is B8, customers should discard the bag. Thus customers can use the lot if the cavity mold is anything other than B8. In addition, customers have been told to empty the LOCI vessel holder on all systems and to discard vessels in case any affected vessels have been loaded onto the system.For technical questions or concerns, please contact the Siemens Technical Solutions Center at 800-441-9250 for further assistance. |  | 408 | 408 units | The device was shipped to medical facilities in CA, CO, FL, GA, IL, LA, MD, MI, MS, NE, NC, NJ, NY, OH, SC, TN, TX, VA, WA, and WI. The product was also shipped to Belgium, France, Germany, Italy, Netherlands, Portugal, Spain, and Switzerland. | Terminated | September 08, 2011 | 653 | 408
Z-1008-2010 | 53961 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BACTEC(TM) Plus + Aerobic/F Medium, catalog 442192, packed in tubes, 50 tubes/case and labeled in part ***Becton Dickinson and Company, Sparks, MD 21152*** |  |  | November 11, 2009 | 2009 | Becton Dickinson &amp; Co. | The draw volume on test tubes of blood culture medium is marked with graduations that may be confusing for healthcare practitioners using the tubes for appropriate collection of patient samples for microbiological laboratory testing. | Component change control | An "URGENT CUSTOMER NOTIFICATION" was sent via UPS to customers. The notification letter described the product, problem and action to be taken. The letter alerted customers that the firm had recently manufactured identified lots of BACTEC(tm) Blood Culture Media with a new label containing a red line which should not be used as an accurate indication for blood draw volume. The firm also advised that the manufacturer is working to build stock with a revised label, but all of the identified lots distributed had passed testing and release criteria and will continue to perform as expected when used as instructions indicate. The customers are being urged to pay close attention to their BACTEC blood collection procedure until they receive bottles with the revised labels. Drawing too much blood may contribute to false negative or false positive results. The customer is being ask to communicate this information to all specimen collection sites within their facility.If further assistance is needed, please contact BD Technical Services Department at 1-800-638-8663. |  | 2940473 | 2,940,473 bottles | Worldwide Distribution: USA and Canada, Japan, Australia, Singapore, Korea, Brazil, Taiwan, China, Mexico, Germany, New Zealand, Hong Kong, Uruguay | Terminated | May 13, 2010 | 183 | 5025698
Z-1020-2010 | 53541 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | RapidPointÂż 405 System Analyzers |  |  | October 30, 2009 | 2009 | Siemens Healthcare Diagnostics Inc | Software: Reference sensor failure may result in inaccurate pH reuslts | Component design/selection | An "Urgent Field Safety Notice" dated September 2009 was sent via overnight mail to customers and included software Version 3.5.2 with instructions for install that will detect and prevent the reporting of potentially errant test results. The customers are requested to please install the included software update immediately and return the attached Fax back form. Siemens will provide the mandatory update to all customers with the Rapidpoint 400 /405 test system.If you have any questions, please contact your local Technical Support at 508-359-3877. |  | 4095 | 4095 units | Worldwide distribution: USA and Australia, Austria, Bahamas, Barbados, Belgium, Bosnia Herzeg, Brazil, Canada, Canary Islands, Chile, Croatia, Czech Republic, Egypt, France, Germany, Holy See (Vatican City State)Hong Kong, Ireland, Italy, Japan, Macedonia, Malaysia, MÂżxico, New Zealand Not assigned P.R. China, Poland, Portugal, Puerto Rico, Refurb.Systems U.S.ARepublic Korea, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, U.A.E., United Kingdom | Terminated | September 04, 2015 | 2135 | 5093
Z-1017-2010 | 52958 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista System BUN Flex Reagent Cartridge, Catalog number K1021, in vitro diagnostic. |  |  | June 16, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Incorrect result: patient sample BUN results reported may be higher or lower than actual level. | Nonconforming Material/Component | Siemens issued an Urgent Field Safety Notice dated 6/16/09 to all affected customers instructing them to discontinue use of the Dimension Vista BUN Flex reagent cartridge lot 09061AA. |  | 1254 | 1254 cartons | Worldwide distribution: USA, Australia, Canada, Germany, Japan, and Korea. | Terminated | November 30, 2010 | 532 | 1254
Z-0972-2010 | 53546 | No | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | Roche/Hitachi GENT2 ONLINE TDM Gentamicin for Roche cobas c 501 analyzer systems; Catalog no. 04490843190. |  |  | October 07, 2009 | 2009 | Roche Diagnostics Operations, Inc. | A high recovery of the assay may be generated, which would lead to an underdosage of gentamicin being administered to the patient. | Other | Consignees were notified by recall letter dated October 7, 2009. The letter described the affected products, issue and instructed consignees to immediately discontinue use of the products. Please contact Roche Diagnostics Technical Support at 1-800-428-2336 if you have questions aboutthe information contained in the recall letter. |  | 13187 | 13,187 of all products. | Nationwide Distribution. | Terminated | December 10, 2010 | 429 | 13187
Z-1024-2010 | 53650 | No | JSM | Transport culture medium. | Microbiology | Microbiology | 510(k) | BD Vacutainer Urine Complete Kit with UA Tube and S&C; Preservative Tube for Midstream specimens; |  |  | October 12, 2009 | 2009 | Becton Dickinson &amp; Company | A small amount of discoloration was noted on Towelettes included in BD Vacutainer Urine collection kits. The entire kitcomponents retain sterile interiors and no viable growth was detected on the towelettes. | Other | Recall notification letters were sent on 10/12/09 with return receipt requested. Customers were advised to contact your SO Sales Consultant for additional information, via e-mail at Vacutainer_TechServices@SD.com or via telephone at 1-800-6310174. |  | 588900 | 588900 | Nationwide distribution to hospitals and medical device distributors. | Terminated | March 10, 2010 | 149 | 1217900
Z-1018-2010 | 54131 | No | LCX | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Mainline Confirms hCG Urine Pregnancy Test Catalog # 6008C. |  |  | October 07, 2009 | 2009 | Mainline Technology, Inc. | Lack of assurance of safety and efficacy: unapproved for marketing in the U.S., non-compliance with cGMP regulations and no stability data to support labeled expirations dates. | PMA | Consignees were notified via recall letters dated October 7, 2009 to cease using and to return the recalled products to Mainline Technology. Direct questions about the recall to Mainline Technology by calling 1-800-541-4467 or 1-734-930-2043. |  |  | Information not provided. | Worldwide Distribution -- United States, Belgium and Switzerland. | Terminated | February 04, 2011 | 485
Z-1021-2010 | 54089 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | SmartInfuser PainPump, part numbers P49230, P49220, P49224-D, P49224N, P49230-D and P49234. |  |  | January 08, 2009 | 2009 | Hsmg, Inc. D/b/a Smartinfuser Usa | Product leaks associated with a new pump reservoir, and incident related to regulator fast priming feature. | Process control | SmartInfuser USA sent letters to consignees on 1/8/2009 and 1/22/2009. |  | 6393 | 6,393 units | Nationwide. | Terminated | April 25, 2011 | 837 | 6837
Z-1016-2010 | 51296 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Xmatrx FISH (fluorescent in situ hybridization) Automated Staining System; Model Number/Catalog Number: AS4010B; |  |  | January 23, 2009 | 2009 | Biogenex Laboratories | Under certain conditions, product system failure has the potential to cause operator injury due to splash of Hemo-De solution. | Device Design | On 1/23/09, the firm issued an Advisory Notice to all its direct consignees, informing them of the affected product and providing instructions on the recall.Please contact BioGenex Technical Support at 1-800-421-4149 for any assistance you may need. |  | 12 | 12 units | The firm distributed the product to one (1) distributor as a consignee of our product. The firm indicated that no customer has received the product directly from BioGenex, due to the fact that the product is exclusive distributed by Abbott Molecular. | Terminated | December 02, 2010 | 678 | 12
Z-1040-2010 | 54232 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Teleflex Medical, Hudson RCI, Concha Therm Neptune, Heated Humidifier. Product numbers: 425-00, 425-10 and 525-30. |  |  | January 12, 2010 | 2010 | Telefelx Medical | When turned on, the Neptune heater may cause an artifact on the patient's heart monitor. | Device Design | Teleflex consignees were notified via Urgent Medical Device Recall letter via next day delivery on/about 01/12/2010. They were asked to return all of the affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. A second mailing to non-responders on the consignee list will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the second mailing or as an alternative, they may be contacted by phone, e-mail or fax. A record of these contacts will be maintained. Affected products received by Teleflex Medical will be reworked with software upgrade and placed back into distribution. Consignees will either receive a software upgrade on site by a third party or replacement product with the software upgrade. Records will be maintained accordingly. |  | 15134 | 15,134 units | U.S. Canada, India, Japan, Puerto Rico, Germany, France, Italy and Ireland. | Terminated | April 05, 2011 | 448 | 15134
Z-0979-2010 | 54194 | No | JSC | Microbiological assay culture medium. | Microbiology | Microbiology | 510(K) Exempt | Oxoid CM0225 Brain Heart Infusion IVD. The firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, England. |  |  | December 21, 2009 | 2009 | Remel, Inc | Product failure - Failure to grow quality control strains of Streptococcus pneumoniae. | Other | The recalling firm issued separate recall letters dated December 21, 2009 addressed to Microbiology Laboratory Supervisors or Distributors explaining the reason for recall and requesting all remaining product be discarded. A Product Inventory Checklist was enclosed to report the amount of product that was discarded by the customer. The Distributor letter requests the consignee contact their customers to notify them of the recall. Direct inquiries concerning this recall to the Technical Services Department at 1-800-255-6730 (United States) or 1-913-888-0939 (International). |  | 105 | 105 bottles | United States (FL, NY, MO, NY, DE, and PA). | Terminated | July 12, 2011 | 568 | 105
Z-1031-2010 | 53893 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Exacta Mix TPN (total parenteral nutrition) Bag, Model 739, 1000mL, Baxa Corporation |  |  | November 12, 2009 | 2009 | Baxa Corporation | TPN bags may leak fluid due to inadequate sealing. | Process control | Consignees were notified by Baxa Urgent Exactamix Bag Recall letter on 11/12/2009 and then again on 11/17/2009 when more lots were found to be affected. Instructions were to stop use and return all affected product to Baxa. More information is available at 800-678-2292. |  | 105500 | 2110/50 bag cases | Worldwide distribution, including USA, Argentina, Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, Hong Kong, Japan, Malaysia, Mexico, New Zealand, Panama, Peru, Singapore, South Korea, Taiwan, and Uruguay. | Terminated | February 24, 2011 | 469 | 293300
Z-1036-2010 | 54094 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pointe Scientific Creatinine Reagent R2, Pointe Scientific, Inc., Canton, MI; Catalog 3-C7539-L, 7-C7539-R2-30, 7-C7539-R2-125, 7-C7539-R2-250, 8-C7539-R2-30, 8-C7539-R2-125 and 8-7539-R2-250. |  |  | December 01, 2009 | 2009 | Pointe Scientific, Inc. | An additional 10% picric acid was added to the R2 component during production. | Process change control | Consignees were notified by An Urgent: Medical Device Recall letter, dated 12/1/09, that identified the affected product. Customers were asked to immediately examine their inventory for the affected product and to notify any customers that the product may have been further distributed to by giving them a copy of the recall notice. Customers are to dispose of the affected product immediately. Questions should be directed to the Technical Support Department at (800)757-5313 or (734)487-8300. There is also a Recall Response form attached with this notice. |  | 625 | 625 vials | Worldwide Distribution -- USA, Bolivia, Greece, Guyana, India, Indonesia, Israel, Kenya, Korea, Malaysia, Mexico, Pakistan, Philippines, Russia, Serbia, Sri Lanka, Tanzania, Trinidad and Ukraine. | Terminated | October 15, 2010 | 318 | 1517
Z-1053-2010 | 54080 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Latex-Free Symbiq Pump Set, convertible pin, 104 inch piggyback with backcheck valve, 2 prepierced injection sites, 0.2 micron filter and OPTION-LOK, non-DEHP, 15 drops/mL; a sterile Rx single use device; Hospira Inc, Lake Forest, IL USA. Made in Costa Rica; product 160410428, list number 16041-28. |  |  | November 24, 2009 | 2009 | Hospira Inc. | Company has confirmed reports of backflow of fluid past the backcheck valve on Symbiq administration sets. | Process control | Hospira, Inc. issued an "Important Administration Set Information" letter dated November 24, 2009 to their direct accounts, informing them of the affected product. Consignees were instructed to return the defective product to the firm.For further information, contact Hospira Medical Communications at 1-800-615-0187. |  | 21312 | 21,312 sets | Worldwide Distribution -- United States, Australia, Canada, Costa Rica, India, Mexico, Netherlands, Philippines, Singapore and United Arab Emirates. | Terminated | January 29, 2012 | 796 | 3558848
Z-1027-2010 | 54187 | eosin y | HYB | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | SS-035C, Hematology Reagent C Eosin Stain, 500 mL, IVD, |  |  | January 06, 2010 | 2010 | Wescor, Inc | Stains used for diagnostic use may be contaminated with bacteria. | Component design/selection | Customers were notified by letter on Jan. 6, 2010 and asked to destroy any remaining product. A response form was included to note whether or not customer had any remaining stock for replacement purposes. Additional information is available at 800-453-2725 or regaff@wescor.com. |  | 24466 | 24,466 bottles | Nationwide, Foreign distribution to Algeria, Australia, Austria, Canada, Cyprus, Denmark, France, Germany, Hungary, Ireland, Italy, Japan, Kuwait, Malta, Netherlands, Oman, Panama, Poland, Qatar, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, United Arab Emirates, UK, and Venezuela. | Terminated | April 20, 2010 | 104 | 26962
Z-1039-2010 | 53686 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista V-LYTE Fluids: Diluent catalog number K835 |  |  | October 22, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Results can elevate after sensor replacement. Sodium urineresults can be falsely elevated 15 mmol/L within 24 hours after installation of a new sensor. | Other | The recalling firm issued an Urgent Field Safety Notice dated October 2009 to all affected customers. Customers were given instruction which minimize urine specimen use and report urine specimen results when QC is acceptable. Sensors are to be replaced when urine QC is out. Replacement sensors are available by contacting the firm at 800-241-0420. Please contact the Siemens Technical Solutions Center at 800-441-9250 for further assistance. |  | 150 | 150 units | The product was shipped to medical facilities nationwide. The product was also shipped to Australia, Canada, Belgium, France, Germany, Italy, Japan, Malaysia, New Zealand, Netherlands, Portugal, Spain, South Korea, and Switzerland. | Terminated | December 16, 2010 | 420 | 150
Z-1043-2010 | 53784 | No | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcium (CPC) R2 Reagent Set. Pointe Scientific Liquid Calcium (CPC) R2 Reagent Set, Pointe Scientific, Inc, Canton, MI; Catalog #'s HC402-R2. Calcium CPC Hitachi Reagent Set for repackaging, Catalog # 3-HC902-R2. |  |  | November 10, 2009 | 2009 | Pointe Scientific, Inc. | Non-compliance with the Good Manufacturing Practice regulations, as the manufacturing process for this product had not been validated. | Process control | Consignees were notified by "Urgent: Medical Device Recall" letter dated November 11, 2009 and instructed to conduct a recall to the retail level, that users should dispose of the product immediately, and that a business reply card should be returned to account for products. If you have any questions, please call Pointe Scientific, Inc. Technical Support Department at 1-800-757-5313 or 1-734-487-8300. |  | 13 | 4 sets and 9 L | Worldwide Distribution (United States, Chile, India, Republic of Yemen and Trinidad). | Terminated | December 10, 2010 | 395 | 84
Z-1033-2010 | 53675 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | IMMULITE 2000 and IMMULITE 2500 |  |  | October 21, 2009 | 2009 | Siemens Healthcare Diagnostics Inc. | The recall was initiated because Siemens Healthcare Diagnostics has identified there is the potential for patient results to be biased low (-0.3 Sample to Cutoff [S/CO] ratio units) due to incompatibility of the Toxoplasma IgM kits with this particular lot of IgG/IgM Sample Diluent. All available lots of the IMMULITE 2000/2500 Toxoplasma IgM kits (L2KTM/L5KTM) are incompatible with Diluent lot 150 | Nonconforming Material/Component | The customer notification was initiated on 10/22/2009 with the firm forwarding an Urgent Field Safety notice with attached Field Correction Effectiveness Check form to the affected consignees informing them of the issue and the products affected. The consignees were informed that Siemens is initiating the Field Correction (FC) to prevent IMMULITE 2000/2500 Toxoplasma IgM from being run with Diluent lot 150.The consignees were instructed in the interim, until replacement kit lots of RUB IgG and CMV IgG are received, your laboratory can continue using Rubella IgG, CMV IgG, or Toxoplasma IgM as long as Diluent Lot 150 is not loaded on the analyzer when running Toxoplasma IgM. Diluent Lot150 can be used with RUB IgG and CMV IgG.Once the replacements lots are received, any unopened kit lots listed in the above table MUST BE RETURNED to Siemens per instructions that will accompany a separate Fed Ex Call Tag that will explain the proper return of the affected product.The consignees were also instructed to complete the attached Field Correction Effectiveness Check form and fax to the firm's technical solutions center. Additional questions should be directed to the firm at 973-927-2828. |  | 557 | 557 units | Nationwide | Terminated | April 22, 2010 | 183 | 567
Z-1064-2010 | 53942 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer NexGen Complete Knee Solution Minimally Invasive Solutions Procedures Trabecular Metal Technology Tibial Tray, fixed bearing, size 8, sterile, Zimmer, Warsaw, Indiana; REF 00-5954-057-02. |  |  | November 16, 2009 | 2009 | Zimmer Inc. | The titanium portion of the implant may separate from the trabecular metal material. | Other | On December 1, 2009, an e-mail was sent to consignees for lot 60866613 to quarantine and return the product. On December 2, 2009 an e-mail was sent to U.S. distributors, and on December 3, 2009 and e-mail was sent to international distributors, instructing them to suspend usage and return all lots of size 8 trays. On December 9, 2009, a recall letter was sent to U.S. user accounts who received the products directly from Zimmer, U.S. distributors and Zimmer's U.S. sales force informing them of the recall. Direct questions about this recall to Zimmer, Inc. by calling 1-800-613-6131. |  | 846 | 846 | Worldwide Distribution -- United States, Australia, Austria, Belgium, Canada, Denmark, France, Germany, Italy, Saudi Arabia, Spain, Sweden, Switzerland and the United Kingdom. | Terminated | March 23, 2011 | 492 | 846
Z-0816-2010 | 54526 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Torflex Transseptal Guiding Sheath, Baylis Medical Company Inc., 5959 Trans-Canada Highway, Montreal, QC, Canada H4T1A1, (514) 488-9802; INCLUDES: One (1) Radiopaque Sheath, One (1) Radiopaque Dilator; And One (1) B* x 135 cm Guidewire. |  |  | February 08, 2010 | 2010 | Baylis Medical Corp * | Radiopaque soft tip may fracture. | Component design/selection | Baylis Medical Issued an "Urgent Medical Device Correction" letter via overnight courier beginning February 8, 2010. Consignees were instructed to segregate and return all affected product to the firm.For further information, contact Baylis Medical Regulatory Scientific Affairs at 1-905-602-4875 Extension 252. |  | 1085 | 1,085 units (508 in US, 577 Foreign) | Worldwide Distribution -- United States, Canada, Chile, China, Colombia, Denmark, Egypt, France, Germany, Hong Kong, Italy, Mexico, Netherlands, New Zealand, Portugal, Qatar, Russia, Singapore, South Africa, Sweden, Switzerland, United Arab Emirates and the United Kingdom. | Terminated | September 13, 2010 | 217 | 1085
Z-1044-2010 | 53896 | No | BSN | Anesthesia conduction filter. | Anesthesiology | Anesthesiology | 510(k) | Millipore Millex Syringe Filter Unit, 33 mm Durapore PVDF 0.22um, Sterile, 50 Filters per box. Catalog Number: SLGVM33RS. |  |  | November 25, 2009 | 2009 | Millipore Corporation | Membrane integrity defects may allow the passage of viable or non-viable contaminants to pass through the filter. | Process control | Millipore Corporation issued an "Urgent: Medical Device Recall" letter dated November 24, 2009 requesting users to examine inventory and immediately discontinue use and return all affected product to the firm.For further information, contact Millipore Technical Services at 1-800-645-5476. |  | 26950 | 539 Boxes - 50 per box | Worldwide Distribution -- United States, Japan and Hungary. | Terminated | May 23, 2011 | 544 | 26950
Z-1067-2010 | 53841 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | CHROMagar VRE Base; |  |  | November 09, 2009 | 2009 | DRG International Inc | CHROMagar Products are labeled incorrectly. Label currently states that the product is for IVD use, when it should state that the product is intended For Research Use, Not for use in diagnostic procedures. | Other | Two recall letters were sent via certified mail to customers. One letter was dated November 9, 2009 referencing the CHROMagar VRE product and the other letter dated November 12, 2009 referencing the CHROMagar 0157 product. The letters describe the affected product, problem, and action to be taken by the customer. The customers are to please STOP the use or distribution of the listed products and IMMEDIATELY examine your stock and return any un-expired CHROMagar product slated for recall for replacement the DRG International, Inc. The customers are being requested to please fill out and return the disposition sheet even if you have none of these items on hand since the information is important.If you have any questions, please contact customer service at (908) 233-0758. |  | 325 | 325 (5 x 5000ml), 65 5 L packs | Nationwide distribution | Terminated | May 14, 2012 | 917 | 2561
Z-1081-2010 | 54253 | No | GXQ | Dura substitute. | Neurology | Neurology | 510(k) | Suturable DuraGen Dural Regeneration Matrix; |  |  | January 13, 2010 | 2010 | Integra LifeSciences Corp | One lot of Suturable DuraGen Dural Regeneration Matrix (# 1091575) did not pass stability testing at certain time points for Suture Retention Strength and Shrink Temperature specifications. | Other | An "URGENT PRODUCT RECALL NOTIFICATION" dated January 13, 2010, was sent via Federal Express to customers. The notification described the product, problem, hazard involved and action to be taken by customers. In the event that the customers have in their possession any products from the specific lot number 1091576, the customers are being ask to please immediately quarantine all stock on hand and contact their local Integra Sales representative for further assistance and for arrangements of product return. The customers are being ask to return the enclosed "Recall Acknowledge and Return Form" indicating whether or not they possess any of the affected product identified. Integra will replace any product affected by this Recall that is returned.Please feel free to contact me at 609-936-5407 should you have any additional questions. |  | 94 | 94 | Nationwide distribution:Ca, NY, TX, CO, VA, SC, IA, MO, IL, LA, OR, OH, WI, FL, DC, DE, MD, MN. | Terminated | March 18, 2010 | 64 | 94
Z-1069-2010 | 54076 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CLAY ADAMSÂż DynacÂż III Centrifuge, 120V, Model Number 420104 |  |  | August 17, 2009 | 2009 | Becton Dickinson &amp; Co. | Medical laboratory centrifuge component may have been assembled incorrectly and cause disabling of the safety feature which prevents customers from opening the centrifuge while it is spinning. | Nonconforming Material/Component | An "Urgent Product Recall" letter dated August 2009 was sent to the customers. The letter described the affected product, problem and the action to be taken by customers. The customers were advised to discontinue use of the centrifuge for repair or replacement with refurbished unit , provide an Excel file or a list with name, address and telephone number of end-users who were shipped this product from their distribution branches so that BD may inform them of the issue,and return the enclosed response form.For further assistance regarding product replacement and returns contact our Technical Services Department at 1-800-638-8663. |  | 411 | 411 units | Worldwide distribution: USA and Mexico, Phillipines, Columbia, Chile, Ecuador, Venezuela, Puerto Rico, Canada and Singapore. | Terminated | April 22, 2010 | 248 | 411
Z-1100-2010 | 53892 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | PET-511BTM Multi-plane transesophageal transducer (TEE) used with Toshiba Diagnostic Ultrasound System |  |  | May 11, 2009 | 2009 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems Inc is removing the PET-511BTM Multi-plane Transesophageal Transducer from the market after receiving several reports of injury regarding the transducer. Bending or buckling of the flexible portion of the PET-511 BTM multi-plane transesophageal transducer (TEE) may result in injury including perforation. | Device Design | Toshiba America Medical Systems sent a Field Safety Notice to all customers on 5 MAY 2009 indicating its intent to remove the units from the market. The letter requested immediate cessation of use and return to the manufacturer site. It announced plans to replace unit with a new multi-plane transducer at no addle charge.The notice also included a description of the problem that occurred, the proper insertion technique into the esophageal passage, how to avoid and detect bending or buckling of the transducer, and what to do if the transducer does buckle. A revised safety manual and poster reflecting these precautions for the PET-51 IBTM was sent to each customer using this transducer. Shipments of the transducer from December 2008 also contained the Field Safety Notification materials. For additional information, contact the firm at (714) 669 2458. |  | 70 | 70 | Nationwide | Terminated | March 22, 2012 | 1046 | 70
Z-1102-2010 | 54230 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Howmedica Osteonics Corp. Restoration Modular Revision Hip System, Plasma Distal Stem; Titanium, Alloy and Hydroxyapatite. Catalog number: 6276-5-014. |  |  | December 30, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The Restoration Modular Hip Stem may have two different package labels. The 14 mm diameter stem may be labeled with both a size 14 mm diameter and a 26 mm diameter. | Process control | Stryker Branches were notified by e-mail on December 30, 2009. This was followed by a formal recall letter by Fed Ex sent on 1/11/10 to branches as well as the hospital and surgeon that used the device. |  | 1 | 1 | 1 device distributed to a hospital in GA. | Terminated | September 14, 2012 | 989 | 1
Z-1071-2010 | 54113 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | DeRoyal TOTAL KNEE TRACECART, REF 53-1818.13, Rx ONLY, NON-STERILE. Custom surgical kit containing Stryker CBCII Blood Conservation Kit. |  |  | December 09, 2009 | 2009 | DeRoyal Industries Inc | Custom surgical kits contain Stryker CBCII Blood Conservation Kit which may lack sterility | Nonconforming Material/Component | DeRoyal initiated its recall by visit on 12/09/2009 and followed by letter dated 12/18/2009 |  | 194 | 194 units | LA &amp; MS | Terminated | May 20, 2010 | 162 | 313
Z-1065-2010 | 54117 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom TotalCare Bed System, Hill-Rom, Batesville, IN; Model P1900 and P1840. |  |  | January 01, 2010 | 2010 | Hill-Rom, Inc. | The controls for nurse call, bed movement and entertainment functions on the intermediate rail may fail to function. | Nonconforming Material/Component | The firm sent a letter to each consignee dated January 4, 2010 advising them that a firm representative will visit the facility to make corrections. Hospitals were requested to take all of the affected beds out of service, if possible, until they are corrected, or to inspect and remove those with a problem out of service until they are corrected if all cannot be removed from service. If you have any questions concerning the recall, please contact Hill-Rom Technical Support at 1-800-445-3720 (U.S.). |  | 185 | 185 | Worldwide Distribution -- United States, Australia, Canada, France, Germany, Mexico, United Arab Emirates and United Kingdom. | Terminated | February 17, 2012 | 777 | 185
Z-1103-2010 | 54358 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | NIPRO, GLUCOPRO INSULIN SYRINGES, SINGLE USE INSULIN SYRINGE with needle, 1CC 30g 5/16", U-100 INSULIN, LATEX FREE. Distributed by: NIPRO MEDICAL CORPORATION 3150 N.W. 107th Ave., Miami, Florida 33172. |  |  | January 13, 2010 | 2010 | Nipro Medical Corporation | Nipro Medical Corporation, Miami FL is initiating a nationwide recall of all GlucoPro Insulin syringes (This does not include the GlucoPro syringe specific for use with the amigo Insulin pump). These syringes may have needles that detach from the syringe. If the needle becomes detached from the syringe during use, it can become stuck in the insulin vial, push back into the syringe, or remain in | Nonconforming Material/Component | Each consignee will receive a Recall Notification Letter and response form by either fax or e-mail.Product should be returned to point of sale.Product not distributed should be collected and returned to Nipro Medical Warehouse in Memphis TN and Carolina Puerto Rico.Any questions or concerns can be addressed to the company at 787.769.2522. |  | 2268600 | 2,268,600 pieces | US (including Puerto Rico) Foreign Distribution: Canada, Barbados, Honduras, Jamaica, Panama, Trinidad and Tobago, and Costa Rica. | Terminated | September 15, 2010 | 245 | 2268600
Z-1083-2010 | 54512 | No | CGZ | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OPTIÂżÂż LION Cassettes, E-Plus, Model BP7507 (25 per box), OSME TECH critical care. |  |  | April 20, 2009 | 2009 | OPTI Medical Systems, Inc | Outer box labeling shows an incorrect storage temperature range. | Process control | OPTI Medical Systems, Inc. issued a "Medical Device Correction" letter dated April 20, 2009. Consignees were instructed on how to corrected the affected product and were asked to return a completed Correction Response Card.For further information, contact OPTI Medical Systems, Inc. at +01 770 510 4444 (outside US) or 1-800-490-6784 (US). |  | 3175 | 127 boxes of 25 cassettes each for a total of 3175 cassettes. | Worldwide Distribution -- United States (CA, KY, MA and OK), Djouti (Africa), Egypt, Guatemala, India, Iraq, Pakistan, United Arab Emirates and Vietnam. | Terminated | April 17, 2012 | 1093 | 3175
Z-1082-2010 | 54482 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Electri-Cord AC Power Cords, IEC 60320 Coupler, Product Reorder Number 21-2145-01 only for use with Smiths Medical CADDÂżÂż-Solis ambulatory infusion pumps, Model 2100. These cords were sold as an accessory for these pumps under Manufacturing Number A23.164.098.0A, Electri-Cord part Number PC2000403. |  |  | January 27, 2010 | 2010 | Smiths Medical ASD, Inc. | Potential for the power cord's prongs to crack and fail at/or inside the plug. Other medical device manufacturers have reported incidents of sparking, charring, and fires from the affected power cords used with their devices.Smiths Medical has received no reports of incidents regarding the use of the affected Electri-Cord powercords with their devices. | Nonconforming Material/Component | Consignees were sent a Smiths Medical, Inc. "Urgent Medical Device Recall" letter dated January 27, 2010. The letter was addressed to Risk/Safety Managers, Clinicians, Critical Care Physicians, Pediatricians, Pediatric Intensivists, Nenatologists, and Nursing, Pharmacy, Clinical/Biomedical Engineering, and Anesthesia Professionals. The letter described the problem, products involved and provided "Advice on Action to be Taken by the User". Consignees were requested to complete and return the "Urgent Recall Notice Confirmation Form". To return Power Cords, for questions regarding the recall notice or to report any issues with the affected product, please contact the Smiths Medical Technical Service Department at 1-800-558-2345. |  | 62 | 62 | United States (KY and OH) and Canada. | Terminated | December 10, 2011 | 682 | 62
Z-1070-2010 | 54251 | No | JOX | Automated heparin analyzer. | Hematology | Hematology | 510(k) | Medtronic HMS PLUS, Hemostasis Management System, Model 30514. |  |  | December 14, 2009 | 2009 | Medtronic Inc | Recently the US Food and Drug Administration (FDA) issued an alert, communicating a change in the United States Pharmacopeia (USP) monograph for heparin, introducing a new USP reference standard that will in effect, change potency, and harmonize the USP unit dose with the WHO International Standard (IU) unit dose. Users of the Medtronic Hemostasis Management System (HMS Plus) must be aware o | Other | Consignees were sent a Medtronic "Urgent Medical Device Notice' dated December 14, 2009. The letter was addressed to "Valued Customer". The letter described the product and the problem. Advised customers to configure the HMS Plus appropriately.The Field Notification was sent to all current HMS Plus users. The current HMS Plus users were identified by the sales of HMS Plus disposable product -Heparin Assay Cartridges from the period of August 1, 2008 to date. This time was considered appropriate, being more than twice the longest expiration date of any of the HMS disposable. |  | 580 | 580 | All states in USA (Including DC) except RI and VT. | Terminated | December 24, 2011 | 740 | 580
Z-1101-2010 | 54369 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | AGA Medical Corporation, Amplatzer TorqVue Delivery System 45-degree. REF: 9-1TV10F45/80. Designed to facilitate the introduction of transvenous devices to chambers and coronary vasculature of the heart and for introducing therapeutic devices to a location within the peripheral vasculature. |  |  | December 18, 2009 | 2009 | AGA Medical Corporation | In this lot of TorqVue Delivery Systems (PN: 9-ITV10F45/80, LN: 0906151781), the dilator cannot be fully secured to the sheath as directed by the instructions for use. | Process design | An AGA Medical Product Recall Notice, dated December 18, 2009 was sent to customers on 12/18/2009. The letter identified the affected product, described the issue and asked for devices to be returned. Customers are asked to complete a form attached to the product recall notice and return it to AGA via fax or email. |  | 57 | 57 | Worldwide distribution: USA, Colombia, Greece, Italy, Japan, Mexico, Slovakia, and Taiwan. | Terminated | March 29, 2012 | 832 | 57
Z-1097-2010 | 54250 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Smart Perfusion Pack, REF: 088501900, Sterile EO, Sorin Group USA, Inc., 14401 W. 65th Way, Arvada, CO 80004 |  |  | December 30, 2009 | 2009 | Sorin Group USA, Inc. | Tubing set packaging seal between lid and tray may not be adequately sealed, compromosing sterility. | Packaging process control | Customers were notified by letter on December 30, 2009 and asked to examine all inventory for seal voids and arrange for return if any were found defective. Examination instructions were included. The two distributors were given the same instructions and were told to share all the information with any of their customers. |  | 9 | 9 units | AZ, CA, FL, MO, NJ, SC, SD, TN, TX, UT, VA and WI. | Terminated | March 15, 2010 | 75 | 280
Z-1111-2010 | 54267 | No | JWX | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Wampole Staph Latex Test, 6.0 ml -120 Test Kit |  |  | January 15, 2010 | 2010 | NERL Diagnostics Corporation | Unreactive Staph latex reagent | Process control | NERL Diagnostics notified distributors on January 15, 2010, by letter to advise of the recall and request they examine inventory, discontinue use of affected lot numbers and discard any remaining inventory. In addition, if product was further distributed, notify the customers of the product recall. If you have any questions, please telephone 1-401-824-2046. |  | 146 | 146 kits | NJ | Terminated | October 05, 2011 | 628 | 271
Z-1133-2010 | 53674 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Martel Printer |  |  | November 06, 2009 | 2009 | Abbott Point of Care Inc. | Some Martel Printers used with the i-STAT 1 system have the potential of being damaged due to the use or rechargeable battery packs that do not contain a fuse. | Component design/selection | An "Urgent Recall Notice" dated October 2009 was sent to customers. The notice described the product, problem and action to be taken by the customers. The customers are requested to remove the printers from use, if they have a Martel printer with the specified serial numbers, that do not have a battery door for access to the rechargeable battery pack. The customers are ask to turn in to Abbott Point of Care or a designated repair facility, all affected Martel printers for battery pack replacement. Abbott will return the Martel printers to your facility at no charge. The customers are ask to please follow the instructions on the enclosed customer reply cared. If in doubt regarding identification of the affected Martel printers, contact Abbott Point of Care Technical Support. If you have any questions regarding this information or if removing the affected Martel printers from use will negatively impact patient care, please contact Abbott Point of Care Technical Support at 800-366-8020, option 1 or your Abbott Point of Care representative. |  | 6432 | 6432 printers (1208 have already been returned and repaired or scrapped) | Worldwide distribution | Terminated | September 14, 2011 | 677 | 6432
Z-1132-2010 | 53759 | No | MRR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PTS Panels LDL Cholesterol test strips for professional use with CardioChek P-A test systems, Manufactured by Polymer Technology Systems, Inc., Indianapolis, IN; Catalog No. 1753. |  |  | May 29, 2008 | 2008 | Polymer Technology Systems, Inc. | The expiration date was extended by 20 weeks without justification. | Process change control | An "URGENT-FIELD CORRECTION" letter dated May 29, 2008, was sent to customers. The letter described the product, problem and action to be taken by the customers. The customers are to please immediately identify the number of remaining LDL kits they currently have remaining in inventory of the specified lots and indicate the amount on the included Customer Acknowledgement and Confirmation letter and return via fax to 217-870-5608. The firm will ship a new lot of the LDL Cholesterol strips with new expiration dates to the customers. The customers are to please save the updated information and make the appropriate changes to their records. If you have any questions, please contact Customer Service at 317-870-5610 or toll-free at 877-870-5610. |  |  | not available. | Worldwide distribution | Terminated | October 20, 2010 | 874
Z-1137-2010 | 54231 | No | MOV | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Modified Novy Cornual Cannulation Set with clear tip, G17558, REF J-NCS-504070, Cook OB/GYN. |  |  | January 08, 2010 | 2010 | Cook Women'S Health | The distal tip may separate from the catheter while the catheter is in the patient. | Process control | Cook Women's Health sent Urgent Recall Notice letters dated 1/8/2010. Consignees were instructed to return product to the manufacturer for credit. |  | 721 | 721 both products | USA, Ireland, China, Canada, India, Australia, Hong Kong. | Terminated | December 10, 2010 | 336 | 3596
Z-1135-2010 | 54282 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips Switched Internal Defibrillator Paddles (used with Philips HeartStart XL and HeartStart MRx monitor/defibrillators) used during open heart surgery. |  |  | January 19, 2010 | 2010 | Philips Healthcare Inc. | Paddle shock switch may fail to actuate and delay or prevent delivery of defibrillation therapy | Nonconforming Material/Component | Philips issued Urgent Medical Device Recall notification to customers by letter on 1/25/10 via UPS advising of this defect and providing instructions on how to determine if they have a defective paddle. The customers will be assisted by field service engineers to remove and replace these defective units (those within the date range of September 1, 2008 to December 30, 2008). Customers are asked to follow the Actions to be taken by customer/user section of the Urgent Medical Device Recall Notification.Removed units will be replaced free of charge and will be destroyed by Philips. Distribution of the Urgent Medical Device Recall Notification will be managed by Philips representatives in each affected geography.Customers are asked to follow the Actions to be taken by customer/user section of the Urgent Medical Device Recall Notification. |  | 1592 | 1592 units | Worldwide distribution: USA, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Dominican Republic, Egypt, Estonia, Finland, France, Georgia, Germany, HONG KONG, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Republic of Korea, Lebanon, Lithuania, Malaysia, Mauritius, Netherlands, NICARAGUA, Norway, Peru, Philippines, POLAND, Portugal, Puerto Rico, Russian Federation, Saudi Arabia, Serbia &amp; Montenegro, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad/Tobago, Turkey, UKRAINE, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, andVietnam. | Open, Classified |  |  | 1592
Z-1136-2010 | 53773 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PTS Panels Lipids Panel test strips for self-testing or professional use with CardioChek P-A analyzers, Manufactured by Polymer Technology Systems, Inc. Catalog # 1710. |  |  | June 05, 2009 | 2009 | Polymer Technology Systems, Inc. | Exhibits under-recovery of all analytes. | Nonconforming Material/Component | Consignees were notified by letter dated 6/5/09, advised of the problem, and that the firm would replace product. |  |  | Not available. | Nationwide and international. | Terminated | October 20, 2010 | 502
Z-1139-2010 | 54281 | No | JSM | Transport culture medium. | Microbiology | Microbiology | 510(k) | remel A.C.T. II Sterile Pack Tube, Collection and Transport System, Ref 12402, 10/pk, each Tyvek pouch contains 1 tube of medium and 1 rayon-tipped sterile swab. The firm name on the label is remel, Lenexa, KS. |  |  | January 18, 2010 | 2010 | Remel, Inc | Pouch may be inadequately heat sealed compromising sterility | Packaging | An IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" dated January 18, 2010, was sent to customers via regular mail. The notice described the product, problem, and action to be taken by the customers. The customers are requested to: review their inventory for the indicated product, discard all remaining units of the product lot, complete the "Product Inventory Checklist" enclosed to report the amount of product they wanted replaced or credited, and return by February 1, 2010 regardless if they have affected product and/or have verbally contacted the firm.If you have any inquiries concerning this advisory, please contact our Technical Services Department at 800-255-6730 (USA) or 913-888-0930 (International). |  | 154 | 154/10-unit packages | Nationwide distribution: MI, MA, WV, CO, ME, FL, NJ, MD, WA, AZ, VA, CA, KS, IL, MO, AL, TN | Terminated | April 25, 2012 | 828 | 154
Z-1131-2010 | 54191 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Optovue RTVue Optical Coherence Tomography (OCT), Model-RT100 with software versions 3.0.x.x and higher, with Normative Database, manufactured by Optovue, Fremont, CA |  |  | January 08, 2010 | 2010 | Optovue Inc. | Marketed without proper 510(k)-- No 510(k) clearance for Optovue RTVue Model-RT100 with Software Versions 3.0.x.x and higher which includes the Normative Database. | Other | The firm is preparing a software "Roll Back" version of software to disable the NDB and other features, ensuring that this version will work properly. A customer notification letter is provided as well as a correction verification form, installation instructions and user manual.If there are any questions, please call 1-866-344-8948. |  | 1800 | 1800 units | Nationwide and Worldwide | Terminated | March 24, 2011 | 440 | 1800
Z-1140-2010 | 54623 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Centurion Healthcare Products Pump Fill Kit DT10175. CONTENTS: 1 CSR WRAP, 2 GLOVES, 1 FENESTRATED DRAPE, 1 EXTENSION LINE Approx. Priming Vol.: 0.3mL, 1 SYRINGE, 1 Exel Huber NEEDLE, 1 ChloraPrep(R), 1 TRAY |  |  | February 01, 2010 | 2010 | Centurion Medical Products | A supplier notified this firm of a recall of their Huber needles because the needles that are labeled as non-coring could core 60-72% of the time. Centurion is doing a sub recall of their kits that contain the referenced needles. | Other | An "URGENT MEDICAL DEVICE WITHDRAWAL" notice dated January 26, 2010, was sent to customers via certified mail. The notice described the product, problem and action to be taken by the customers. The customers were instructed: DO NOT USE THESE AFFECTED PRODUCTS. PLEASE RETURN ALL AFFECTED PRODUCTS IN ACCORDANCE WITH THE INSTRUCTIONS PROVIDED. The customers were also instructed to report any adverse reactions experienced with the use of the product and/or quality problems to MedWatch Program by phone at 1-800-FDA-1088, by fax at 1-800-FDA-1078 by mail at MedWatch, HF-2, FDA, 5600 Fishers Lane, Rockville, MD 20852-9787 or on the MedWatch website at www.fda.gov/medwatch. The customers were also instructed to check the lot numbers of products in inventory to determine if they have products affected by this withdrawal and complete and return the enclosed Customer Response Form by fax or email.Direct all inquiries to: EXEL International Customer Service 727-827-1922 ext. 0 Monday through Friday 9am-5pm EST or E-mail info@exelint.com |  | 880 | 880 kits | USA only | Terminated | November 08, 2010 | 280 | 880
Z-1108-2010 | 54622 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ACM Bio-Prep with Merocel pkg/6, Rx, Sterile, REF 0306-708-000S4, Stryker Instruments, Kalamazoo, MI. |  |  | January 25, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | The firm did not have packaging validation to support the labeled shelf life. There is potential for a breach in sterility specific to the sterile packaged femoral sponge which may result in a non-sterile product. | Other | A "Medical Device Recall Notification" dated January 21, 2010 was issued to customers. The notification described the affected product, issue and actions to take. Consignees were asked to forward the notification to all their affected locations. Customers were instructed to immediately check inventory and quarantine any recalled product found. Consignees were instructed to destroy all product and report how much product was destroyed. For questions regarding this recall, please contact Stryker Instruments by calling 1-800-800-4236, ext. 4354 |  | 9274 | 9,274 all products | Worldwide Distribution -- United States, Canada, Mexico, Australia, Brazil, Switzerland, Japan, Netherlands, England, UK, Sweden, Hong Kong, Korea, Malaysia, Singapore, Argentina, Spain, and Italy. | Terminated | October 25, 2010 | 273 | 9274
Z-1154-2010 | 48976 | No | JEC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stableloc (TM) Wrist External Fixation System, Quantity 1. Packaged labeled, "Stableloc Wrist Fixator System... Desc:stabeloc (TM) External Fixator Kit... Ref FX-4001-S... Manufactured or Distributed by: Acumed (R) 5885 NW Cornelius Pass Road, Hillsboro, OR 97124...STERILE" |  |  | July 09, 2008 | 2008 | Acumed LLC | Parts are non-functional. The user is unable to fully pass the fixator pins through the body of the Stableloc assembly as required. The kit may also be missing the set screw. | Process control | The firm sent "Urgent Field Safety Notice: Medical Device Recall" letters to customers 7/9-11/08. In addition, the firm sent the letter via e-mail to some of the customers. The notification advised the customer of the recall, identified the product, explained the reason, and asked the customer to locate and segregate any affected product. The customer was asked to notify their customers who they had distributed the affected product to and asked them to complete and return an enclosed response form. The firm provided a model recall letter for its agents/distributors to issue to users. Upon receipt of the response form, the firm will contact customers to arrange for replacement of any affected product. Questions should be directed to the "Distributor phone number" or Richard Snyder, Quality Manager at 1-888-627-9957. |  | 80 | 80 kits | Worldwide Distribution -- USA, including states of VA, NC, FL, ME, WA, GA, TX and PA and countries of Spain, Japan and Greece. | Terminated | May 21, 2010 | 681 | 80
Z-1141-2010 | 50102 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Howmedica Osteonics Corp., Scorpio NRG Femoral Size 7 Right; Knee prosthesis (Stryker France). Product Code: 814407R. |  |  | August 20, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | A box labeled as a Scorpio NRG Femoral Size 7 Right, may actually contain a Scorpio PS Femoral Size 5 Right. | Process design | Stryker Important Product Removal Letters were sent via Federal Express on 8/20/2007. |  | 4 | 4 units | Product went to 2 US Stryker branches. | Terminated | March 16, 2010 | 939 | 4
Z-1109-2010 | 52618 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PrecisePLAN 2.15 Treatment Planning System. |  |  | September 30, 2008 | 2008 | Elekta, Inc. | The new "DICOM CT with Body Outline" method to create a patient case in PrecisePLAN 2.15 may produce incorrectresults for some customers, given several input conditions encountered in the field. Discontinue use of this method ofcreating a patient case until FCO 772 00 540 064, (entitled "Update for Patient creation method using DICOM CT with Body Outline"), is received from Elekta to update th | Other | Important Notice A297, "PrecisePLAN release 2.15 with DICOM CT with Body Outline" method of patient creation (dated May 8, 2007) was distributed to all customers. An additional Field Change Order 540 073, dated September 30, 2008 was distributed to customers. Important Notice A316 "Cancellation of Important Notice for PrecisePLAN" was issued. The Important Notice instructed affected users to dispose of Important Notice A297 and replace it with Important Notice A316.If you have any questions regarding the notification, please contact your local Elekta representative directly. |  | 280 | 280 units | United States -- (AL, AK, AZ, AR, CA, FL, GA, HI, IL, IN, IA, KY, LA, MD, MA, MI, MS, MO, NE, NV, NJ, NM, NY, , NC, OH, OK,OR, PA, SC, TN, TX, VA, WA and WI) &amp; France. | Terminated | December 28, 2010 | 819 | 280
Z-1146-2010 | 53597 | No | MVC |  |  | Pathology | PMA | VENTANA PATHWAY anti-HER-2/neu (4B5) RABBIT MONOCLONAL PRIMARY ANTIBODY |  |  | September 17, 2009 | 2009 | Ventana Medical Systems Inc | Incorrect result: Light staining inconsistencies with lot number 680957 of the Ventana PATHWAY HER-2/neu (4B5) Rabbit Monoclonal Primary Antibody. | Nonconforming Material/Component | The recall was initiated on 09/17/2009 with Ventana Medical Systems forwarding Urgent Voluntary Medical Device Recall letters to all the affected customers advising them to discontinue the use of lot number 680957 and have their laboratory retest with product from a different lot of PATHWAY HER2 (4B5) any breast cases scored as 0, 1+, or 2+ that were not reflexed to an approved in situ hybridization (ISH) method.Customers were also instructed to contact Ventana Customer Support at 1-800-227-2155 and select 1 for Technical Support to speak to a Customer Support agent who will work with then to replace or credit any product from this specific lot number. |  | 433 | 433 units | Worldwide distribution: USA, Austria, Denmark, Germany, Hungary, Netherlands, Sweden, United Kingdom, Brazil, Canada,India, Korea, Morocco &amp; Thailand | Terminated | April 22, 2010 | 217 | 433
Z-1148-2010 | 54190 | formulations, paraffin, all | KEO | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Paraffin Pretreatment Reagent Kit III; each kit contains the following: |  |  | January 22, 2010 | 2010 | Abbott Molecular | The Paraffin Pretreatment Reagent kits do not have any hazardous or MSDS information included in their labeling. | Package design/selection | Abbott Molecular sent Field Correction Letter/Urgent Field Safety Notice letters dated 1/21/10 to the Paraffin Pretreatment Regent Kit customers on 1/22/10, informing them that the MSDA symbology for hazardous and corrosive material is missing from the Paraffin Pretreatment Regent Kit II and Paraffin Pretreatment Regent Kit III labels. The information is missing from the vial label, kit label and package insert. The letter listed the risk and safety statements that should be included in the package insert, and advised the customers should use the kits in accordance with the listed risk and safety statements. The accounts were instructed to review this information with laboratory personnel and retain the letter for future reference, making copies of the letter and placing them in each kit in stock. Any questions were directed to their local Abbott Molecular Representative. Abbott Molecular plans to telephone each customer to ensure that they have received the letter and have disseminated the information to their laboratory personnel. |  | 17 | 17 kits | Worldwide distribution: USA, Argentina, Austria, Belarus, Belgium, Canada, Colombia, France, Germany, Greece, Hong Kong, Hungary, India, Italy, Japan, Kuwait, Lebanon, Lithuania, the Netherlands, Panama, Poland, Singapore, Slovenia, South Africa, Sweden, Switzerland, Taiwan, Turkey, United Kingdom and Vietnam. | Terminated | January 29, 2012 | 737 | 322
Z-1155-2010 | 53848 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Piccolo Comprehensive Metabolic Panel Reagent Disc, 10-pack box, for use with the Piccolo Blood Chemistry Analyzer or Piccoloxpress Chemistry Analyzer, Catalog No. 400-0028, manufactured by Abaxis Inc., Union City, CA |  |  | November 16, 2009 | 2009 | Abaxis Inc | Wrong barcode applied to a lot of items, which may result in incorrect calibration factors and results. | Labeling mix-ups | An "URGENT: PRODUCT RECALL NOTICE" dated November 16, 2009, was sent to customers. The notification described the product, issue/impact, and the action/resolution to be taken by the customers. The customers are to discontinue use of the specified lot, contact Abaxis for immediate replacement and complete and return the enclosed Response form within 5 working days. They may fax the form to Abaxis Technical Support at 1-510-315-3140. The customer are requested to ensure all laboratory staff are informed about this notice and retain this notification as a part of their laboratory Quality System documentation. The firm is currently contacting all direct shipment customers.If you have any questions regarding this Product Recall Notice, please contact Abaxis Medical Technical Support at 800-822-2947, extension 6505, or at medtechsupport@abaxis.com. |  | 1010 | 1010 boxes | Worldwide distribution: USA and Canada, EU and Asia | Terminated | December 14, 2010 | 393 | 1010
Z-1150-2010 | 54490 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Bond-1 Primer/Adhesive. |  |  | August 24, 2009 | 2009 | Sybron Dental Specialties | Some of the material has been found to be gelled and therefore becomes unusable. | Nonconforming Material/Component | An "Urgent Medical Device Recall" letter and "Recall Return Form" were sent by first class mail on August 24, 2009 to consignees. The letter acknowledged the "gelling" issue of identified lots shipped between July 2009 &amp; August 2009. Customers were requested to return any of the remaining affected lots in their inventory. Contact Customer Service by calling 1-800-551-0283, (option 1) directly to handle the arrangements of a quick return and replacement. |  | 1400 | 1,400 | Worldwide Distribution -- United States, Canada, Netherlands, Sweden, Taiwan and Thailand. | Terminated | February 24, 2012 | 914 | 1400
Z-1157-2010 | 54537 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Auto Suture ProTack Fixation 5 mm |  |  | February 08, 2010 | 2010 | Covidien LP | Fixation device may fail to fire and staple. | Process control | Covidien notified accounts by "Urgent Medical Device Recall" letters dated 2/8/10. The letter identified the affected product and the reason for recall. Customers are to examine their inventory, and remove and return recalled lots. Customers are to follow the instructions provided in the letter to return the affected product. Questions should be directed to a Covidien Representative. On June 23, 2010, the firm issued a letter to expand the recall to include 5 additional lots. |  | 5638 | 5,638 units | Nationwide Distribution | Terminated | February 28, 2013 | 1116 | 5638
Z-1153-2010 | 53477 | Yes | NIM |  |  | Cardiovascular | PMA | Xact 10 mm x 20 mm x 136 cm (Part Number 82099-01) |  |  | March 19, 2009 | 2009 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Abbott Vascular is initiating a recall on the XACT Carotid Stent System because the product was distributed prior to approval of a PMA supplement for a manufacturing line move between sites. There are no product quality issues associated with this action. | PMA | Notification was sent out to all consignees about the product recall. Representatives from Abbott Vascular personally visited consignees and removed product from site. |  | 49 | 49 | Nationwide: GA, NY, IN, LA, MA and PA | Terminated | March 30, 2012 | 1107 | 49
Z-1142-2010 | 54363 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Neutralect Diathermy Cable, |  |  | January 22, 2010 | 2010 | ConvaTec | Mislabeling: Boxes labeled with the order code 3508M may contain a cable for the order code 3405M, and boxes labeled with order code 3405M may contain a cable for order code 3508M. | Other | Unomedical Urgent Medical Device Correction letters were sent to distributors on January 22, 2010 by certified mail. Letters state that the only way to determine if the incorrect product is contained in the carton is to open the box. Distributors are asked to inspect all Diathermy Cables in their possession with order code 3508M or 3405M to ensure that the correct cable is inside. The number of mislabeled product is to be recorded on the Questionnaire for Distributors, and then mislabeled product is to be destroyed by cutting the product in half. The questionnaires include an area for names of customers stating that the distributor will contact them, or requesting that Unomedical contact them directly. Copies of the correction notice and questionnaire for end users are to be forwarded to all customers, asking them to complete the questionnaire and return it to the distributors. Once received by the distributors, end users can discard the cables by cutting them in half. All completed forms should be returned to Unomedical and local reps are to be contacted to arrange for replacement of mislabeled stock and/or credits. |  | 6400 | 5130 cables worldwide, 1270 in US. | Worldwide distribution. | Terminated | June 10, 2010 | 139 | 6400
Z-1144-2010 | 54507 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Ascent Posterior Occipital Cervical Thoracic System, Single Looped Cable with Two Crimps, One Each. Distributed by Blackstone Medical, Springfield, MA. |  |  | February 02, 2010 | 2010 | Blackstone Medical, Inc. | The Ascent Posterior Occipital Cervical Thoracic System Single Looped Cable with two crimps may be made of stainless steel instead of Titanium. | Labeling mix-ups | Orthofix Spinal Implants contacted customers by phone beginning December 22, 2009 and issued a Notification of Product Removal dated February 2, 2010. Consignees were instructed to Identify, Remove and Return all affected product to the firm.For further information, contact Orthofix Spinal Implants, Regulatory Affairs Department at 1-973-406-2812. |  | 25 | 25 | Worldwide Distribution -- United States (Texas and California), Australia, Colombia and Greece. | Terminated | April 09, 2010 | 66 | 25
Z-1149-2010 | 54374 | No | LIA | Shigella spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD Shigella poly Group D Antiserum S. sonnei; types I,II, packaged in 3 mL vials, Catalog #24103, and labeled in part. Becton, Dickinson and Company, 7 Loveton Circle, Sparks, MD 21152 USA. |  |  | January 25, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic test reagent may exhibit decreased reactivity. | Nonconforming Material/Component | BD Diagnostic Systems issued an "Urgent Product Recall" notification dated January 2010. Consignees were advised to discontinue use of the affected product and discard for replacement or credit.For further information, contact BD Diagnostic Systems, Customer Service at 1-800-675-0908. |  | 1096 | 1096 units | Nationwide Distribution | Terminated | May 21, 2010 | 116 | 1096
Z-1143-2010 | 54278 | No | GZL | Depth electrode. | Neurology | Neurology | 510(k) | FHC microTargeting Guideline 4000 ver 1.4 (GL4K) Software, |  |  | December 28, 2009 | 2009 | FHC, Inc. | The raster tab for the GL4K version 1.4 software may assign incorrect depths for waveform recordings and if used to identify the implantation target for the DBS lead, the implant may be placed in the wrong location. | Software change control | FHC notified customers by telephone and by an Urgent: Medical Device Correction email, dated December 2009, beginning December 28th. Customers have been instructed to keep the instruction with the GL4K until FHC can provide a revision to the software; and, to use only "distance from zero" or "distance traveled" modes (not "distance from target" mode) during procedures using the raster tab. Customers are to notify their FHC representative to confirm receipt of the notification. A software patch has been verified and validated and will be installed on all affected units. The new version of the corrected software is Guideline 4000 version 1.41. Questions should be directed to Jason Moyer, product manager or Keri Seitz, recall coordinator at 207-666-8190. |  | 10 | 10 | Worldwide Distribution -- USA, including states of GA, MA, NC, NJ, OH, RI, TN, TX, and WV and Bucharest, Romania. | Terminated | March 18, 2010 | 80 | 10
Z-1145-2010 | 52889 | No | DHC | Systemic lupus erythematosus immunological test system. | Immunology | Immunology | 510(k) | BIO-RAD 425-2020 Autoimmune EIA Anti-Cardiolipin IgM, 96 Tests...2010 02 16...Lot CL-863. Positive Control included in lot is identified as "REF 425-2025" and "LOT CPM-591". in vitro diagnostic. |  |  | July 14, 2009 | 2009 | Bio-Rad Laboratories Inc | The firm received customer complaints of microbial contamination in Positive Controls in its Anti-Cardiolipin IgM Test Kit Lot CL-863. | Nonconforming Material/Component | Bio-Rad initiated its field correction on July 14, 2009. The firm notified domestic consignees by phone followed by fax, and each subsidiary will be given an urgent notice, subsidiary response form, customer medical device correction notification, acknowledgement form and a listing of consignees. The subsidiaries must contact customers as instructed. |  | 213 | 213 kits | Worldwide distribution: USA, Australia, Canada, France, Italy, Korea, Mexico, Panama, Singapore, and Trinidad. | Terminated | March 18, 2011 | 612 | 213
Z-1156-2010 | 53463 | No | LPB |  |  | Cardiovascular | PMA | CircuCool Fluid Circulation Pump used in the Chilli Cooled Ablation System, Model Number M00480050, manufactured by Barnant ThermoFisher Scientific, Barrington, IL. Distributed by Boston Scientific Corporation, San Jose, CA. |  |  | September 22, 2009 | 2009 | Boston Scientific Corp | Product failure-- Product may cease operation as a result of potential speed control board failure when the device is shut off, and may not turn back on. A component on the circuit board fails unexpectedly because of a power surge, causing a diode to burn out. | Nonconforming Material/Component | The firm has issued notification letters via Federal express with reply verification tracking forms. A follow up letter and telephone survey as needed will be sent to follow up with non-responders. The firm is asking its consignees to return any recalled product.If there are any questions, please call 1.408.935.4952. |  | 75 | 75 units | US and to Spain, France, Germany, Italy, Russia, Poland, the Netherlands, Israel and Belgium | Terminated | December 27, 2010 | 461 | 75
Z-1151-2010 | 54372 | No | JEY | Bone plate. | Dental | Dental | 510(k) | BIOMET MICROFIXATION, "1.5MM" SYSTEM BONE SCREW LP, SELF DRILL, X-DRIVE, SGL.LEAD AUGER POINT, 1.5 X 3.5MM, TITANIUM, NON STERILE PRODUCT. Model 01-7293 and 99-7293-01. BIOMET MICROFIXATION, 1520 TRADEPORT DRIVE, JACKSONVILLE, FL 32218 USA. |  |  | October 06, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling Model 01-7293 an Model 99-7293-01, 1.5X3.5mm Self Drilling, Low Profile, Cross Drive screw, 5/PK was found to contain a 91-6104 (1.5X4mm High Torque, Self Drilling, Cross Drive screw. | Process control | Biomet Microfixation Issued an "Urgent Medical Device Recall Notice". Consignees were notified of the recall instructed to indentify and return all affected product to the firm. An Inventory Reconciliation sheet should be completed and faxed to 1-904-741-9425. For further information, contact Biomet Microfixation, Global Complaint Coordinator at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 831 | 831 screws | Worldwide Distribution -- United States (WI), Spain, Italy, Japan, Brazil and Korea. | Terminated | April 07, 2010 | 183 | 831
Z-1158-2010 | 54554 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion, Article #715000. |  |  | July 02, 2007 | 2007 | Elekta, Inc. | Need to modify the closing speed of the shielding doors in the event of an emergency exit | Other | A field action (FCO TN-07-0007) has been performed to modify the shielding door speed on all sites in clinical use. A second Field Change Order 200 069 "Replacement of Door Servo Board" dated 9/16/08 was released. |  | 19 | 19 | Worldwide Distribution -- USA, including states of AR, AZ, CA, FL, IL, MI, MN, MS, NJ, OH, OR, PA, UT, VA and countries of France and United Kingdom. | Terminated | March 25, 2010 | 997 | 19
Z-1163-2010 | 54634 | [6 empty] | Olympus Trocar Spikes, Model numbers: A5821, A5948, A5949, WA58341L, and WA58343L |  |  | February 08, 2010 | 2010 | Olympus America Inc. | weak weld seams | Nonconforming Material/Component | The recalling firm issued a Medical Device Recall letter dated February 8, 2010, informing their customers of the problem and the need to return the product. |  | 45 | 45 | Nationwide: AL, AR, CA, FL, MI, MN, MO, NY, PA, OH, TX, and WY. Foreign: Canada | Terminated | May 25, 2010 | 106 | 45
Z-1159-2010 | 54559 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta SL Series/Precise Digital Accelerator. |  |  | October 01, 2005 | 2005 | Elekta, Inc. | Customer reported larger than expected variations in wedge factory with gantry rotation. | Device Design | Important Notice A280 "Wedge Matching Procedure" (dated 11/18/05) was sent to all effected customers. The Important Notice described the issue, warnings and cautions and corrective instructions/Maintenance to take to eliminate occurrence. Direct questions to your local Elekta Limited representative. |  | 315 | 315 units | United States, Canada and Mexico. | Terminated | March 25, 2010 | 1636 | 315
Z-1160-2010 | 54619 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ATEC 0909-12 Breast Biopsy and Excision System , 9 gauge, 9 centimeter handpiece (part number 0909-12), Rx, Sterile, Suros Surgical Systems, Inc. 6100 Technology Center DR. Indianapolis, IN 46278. |  |  | February 01, 2010 | 2010 | Hologic, Inc. | The product is labeled as ATEC0909-12 but it actually contains 0909-20. | Employee error | Formal recall notification letters and response forms were sent to customers via FEDEX on February 1, 2010. If there are any questions, please call 317.344.7630. |  | 190 | 190 | US and Puerto Rico. | Terminated | October 08, 2010 | 249 | 190
Z-1162-2010 | 52299 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | SilverHawkÂż Peripheral Cutter Driver Plaque Excision System |  |  | January 21, 2010 | 2010 | Ev3, Inc. | ev3 Inc. is conducting a voluntary recall of specific lots of the SilverHawkÂż Cutter Drivers (Model Number FG 02550) because of damage to the packaging. | Packaging change control | Consignees were sent an ev3 "Medical Device Recall" letter dated January 21, 2010. The letter was addressed to Cath Lab Manger/Risk Manager. The letter described the Issue Summary, Affected Product, and Required Action. The letter instructed customers to locate and remove from use the specified Cutter Driver(s). Customers are to complete and return the "Device Recall Field Action Reconciliation Form." An ev3 rep will assist in arranging for the return of all unused product and a replacement will be provided. Questions regarding the recall should be directed to ev3 Customer Service at 1-800-716-6700. |  | 8330 | 8,330 | Worldwide Distribution -- USA, with the exception of AK and VT, and countries of AUSTRIA, BELGIUM, BRAZIL, CANADA, CHINA, CZECH REPUBLIC, FRANCE, GERMANY, INDIA, IRELAND, ISRAEL, ITALY, JORDAN, LEBANON, POLAND, SAUDI ARABIA, SOUTH KOREA, SPAIN, SWEDEN, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | April 26, 2012 | 826 | 8330
Z-1164-2010 | 52825 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion 32 CT System; Model number TSX-101 A/E |  |  | January 08, 2009 | 2009 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems Inc (TAMS) has initiated a field correction on the Aquilion 32 CT Scanners due to the potential for the bolus tracking software (SureStart) to cause delay of the helical scan, up to 10 seconds, when large volumes of data are being reconstructed on the secondary console. | Component design/selection | Customer notification was sent and a software upgrade will be issued to all customers affected by the field correction.If you have questions, please call (800) 521-1968. |  | 65 | 65 | Nationwide | Terminated | June 01, 2012 | 1240 | 692
Z-1171-2010 | 52832 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Aplio 50; SSA-700A; software version 5.5r002 |  |  | November 12, 2007 | 2007 | Toshiba American Medical Systems Inc | Toshiba America Medical System Inc initiated a field corrective action on Aplio 50; SSA-700A; software version 5.5r002 and later, Aplio 80; SSA-700A; software version 5.5r002 and later, and Xario; SSA-660A; software version 1.0 and later, because the product does not meet the Safety Standard requirements. | Component design/selection | Toshiba America Medical Systems (TAMS) has issued a Field Modification Instruction notification to all of its customers about the affected devices. The correct software will be administered through TAMS' representatives. Reimbursement for the modification will be given upon receipt of the reimbursement claim invoice. |  | 8 | 8 units | Nationwide Distribution | Terminated | April 18, 2012 | 1619 | 200
Z-1174-2010 | 54703 | No | MEQ |  |  | Gastroenterology/Urology | PMA | UROLOGIXÂż CTC Advance" Standard Microwave Catheter (3.0-5.0 cm), CTC Advance" Long Microwave Catheter (4.5+ cm), CTC Advance" Short Microwave Catheter (2.5 -3.5 cm), RTU Plus Balloon, and RTU Plus Handle, Sterile EO, Manufactured by UROLOGIXÂż, INC, 14405 Twenty-First Avenue North, Minneapolis, MN 55447. |  |  | February 18, 2010 | 2010 | Urologix, Inc. | UrolgixÂż is recalling certain lots of CTC Advance" catheters and RTU Plus devices which are integral components of the TargisÂż System. The reason for this Recall is a labeling error only; the issue involves a discrepancy in the"Use Before" date printed on the product label. In the affected lots, the "Use Before" date , commonly known as the expiration date, is printed as 2012-12 when it should r | Incorrect or no expiration date | Consignees were sent a UROLOGIX " Important Recall Information" letter dated February 18, 2010. The letter was addressed to the client's name. The letter described the problem and product involved. It provided Instructions for "Returning and Replacing Affected Product" and requested the return of the "Customer Acknowledgement Form".If you have questions, please call (763)-475-1400. |  | 154 | 154 | AL, AZ, CA, FL, IN, MI, MN, MS, NY, OK PA, TN, TX, and WI. | Terminated | March 08, 2012 | 749 | 154
Z-1177-2010 | 52816 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems Inc (TAMS) Digital Radiography System Radrex-i DRAD-3000E; Electrostatic x-ray imaging system |  |  | May 22, 2009 | 2009 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems Inc (TAMS) has initiated a field correction on the DRAD-3000E; Radrex-I, Electrostatic x-ray imaging system due to a software bug that results in a loss of image. | Component design/selection | An Urgent - Medical Device Correction letter was sent to customers. The letter described the issue, the corrective action, and the avoidance. Toshiba America Medical Systems informed customers that software was being prepared to address the issue and planned to send representatives to each customer affected by the field correction to provide a software upgrade to the DRAD-3000E; Radrex-I, Electrostatic x-ray imaging system. Customers were to complete the attached customer reply form. Questions were to be directed to Aaron Ybarra, X-ray Product Manager at 800-421-1968, ext. 2492. |  | 20 | 20 | Nationwide Distribution -- Including states of PA, VA, OH, TN, LA, IL, CA, MT, AZ, and TX. | Terminated | March 21, 2012 | 1034 | 20
Z-1178-2010 | 53474 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Peripheral Dilatation Catheter, Viatrac 14 Plus 5.0mm x 15mm x 80cm, Part Number 1008194-15. Manufactured by Abbott Vascular, Temecula, CA. |  |  | May 06, 2009 | 2009 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Abbott Vascular is initiating a recall on the Viatrac 14 Plus 5.0x15 mm Peripheral Dilation Catheter due to a mistake in labeling/packaging. | Labeling mix-ups | Abbott Vascular issued an "Urgent Device Recall" notification dated May 7, 2009. Consignees were informed of the affected product and instructed to work with local sales representatives in identifying any unused devices and arrange for return to the firm.For further information, contact Abbott Vascular Customer Service at 1-800-227-9902. |  | 26 | 26 | Nationwide Distribution -- IL, FL, CA, WA, AL, OH, NY, MA, TX and IN. | Terminated | March 30, 2012 | 1059 | 26
Z-1175-2010 | 54134 | gown, examination | FME | Examination gown. | General Hospital | General Hospital | 510(K) Exempt | Thermoflect Product Line including items for adults and pediatrics. |  |  | December 23, 2009 | 2009 | Encompass Group LLC | The products are not compatible for use in the Magnetic Resonance Imaging (MRI) environment. | Other | A press release distributed through specialized news media, and trade magazines, on/about 12/26/2009. The press release was also posted on the thermoflect.com website. Letters were sent via first class mail to customers, hospitals and doctors in radiology on/about 12/23/2009. In addition to the written notification to the user institutions, Encompass provided adhesive labels sent along with the letter notification. The notice instructed that the labels should be affixed to remaining inventories of the products(s), specifically to the individual package, which holds 5 blankets and to the outermost packaging or carton. The label reads Not for use in MRI, and a symbol showing a circle with a diagonal line through the letters MR. Thermoflect product literature has been quarantine at this time. Instructions were provided showing where the labels where the labels were to be affixed. Encompass also provided a poster indicating the products are for use in the MR environment which is to be displayed in the MR area. All new product literature will indicate that the product is not for ruse in the MR environment. Thermoflect sales representatives will visit with each user institution to confirm their receipt of the recall notice and to be available to answer any question concerning the recall. The sales representatives all of whom have been briefed on the recall strategy will be available to provide training for Radiation Safety Officers and radiology staff, if requested. Following the recall notice, Encompass will conduct Effectiveness Checks to assure that all purchasers/users of the Thermoflect product have been notified and have taken appropriate action to notify their radiology centers and staff. The Encompass recall team will maintain records of the contacts for review by the FDA. A Status Report on the effectiveness of the recall effort will be filed with the Atlanta District Office on a monthly basis until the recall is terminated. |  | 700000 | 700,000 units (of various product lines) | AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY and Switzerland. | Terminated | October 27, 2010 | 308 | 700000
Z-118-2010 | 54538 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | AOS Trochanteric Nail, 13mm x 20cm x 130Âż, Product Part Number: 1037-200. |  |  | January 14, 2010 | 2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail, where one piece was found to be incorrectly bent. If this issue were to recur, surgery time could be extended, but only slightly as nail exchanges take place commonly in a surgery of this type. | Process control | The recall was initiated with the firm forwarding a recall notice with attached customer response form on January 14, 2010 to the affected consignees. The recall notice informed the consignees of the reason for recall and gave them three steps to follow to inspect the affected Short Trochanteric Nail products. Consignees were provided with product return instruction for any incorrectly bent products found during their inspection. Customers with questions are instructed to contact the firm at (310) 533-9966. |  | 65 | 65 | Nationwide and Japan | Terminated | April 26, 2012 | 833 | 1981
Z-1176-2010 | 54633 | No | NFF | Medical image storage device. | Radiology | Ophthalmic | 510(K) Exempt | Image Consultant (IC) Software |  |  | February 05, 2010 | 2010 | R.O. Gulden &amp; Co., Inc. | Software not validated; inaccurate results may occur. | Other | On 2/5/10, the recalling firm issued letters via certified mail to its consignees informing them of the problem and to cease use of the software for diagnosis. |  | 20 | 20 | The product was shipped to eye doctors/clinics in AL, FL, LA, MA, NC, OK, OR, PA, TX, WI, and WA. | Terminated | June 14, 2010 | 129 | 20
Z-1198-2010 | 50777 | No | HIF | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Pnemosure Insufflators, model numbers 620-040-600 and 620-040-610, manufactured by Stryker Endoscopy, San Jose, CA. |  |  | December 29, 2008 | 2008 | Stryker Endoscopy | Unexpected powering down/failure to boot up: A problem with a transistor in the heater/temperature sensor component in the device may cause the device to power off during surgery or to fail to boot up at all. The risk includes delay and interruption of ongoing surgical procedures. | Component design/selection | Recall initiated 12/29/2009. Sales representatives were notified via phone and e-mail, while product advisory notices were sent to consignees via FedEx. Consignees are provided with a response form, and are instructed to run a software correction. |  | 884 | 884 devices | Product was distributed throughout the US and to Argentina, Brazil, Canada, Chile and Latin America, as well as to 16 government facilities. | Terminated | March 24, 2011 | 815 | 884
Z-1197-2010 | 54773 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Exacta-Mix 2400 Valve Set, 24 - Port Valve Assembly, REF: 724, STERILE, 10 per box. Baxa Corporation, Englewood CO. |  |  | February 10, 2010 | 2010 | Baxa Corporation | Molding defect in the valve body causes leaking. | Nonconforming Material/Component | Baxa Corporation notified customers of the affected product by letter beginning February 10, 2010. Users were instructed to return all affected product to the firm.For further information, contact Baxa Corporation at 1-303-617-2242. |  | 104329000 | 3230 cases; 32300 units (10 units/case) | Nationwide Distribution -- Including AL, AR, AZ, CA, CO, DC, MN, IL, NY, DE, FL, PA, IN, TN, MS, MO, TX, MI, MA, MT, MO, SC, VA, WV, NJ, OH, CT, MO, NV, LA, WA, WI, GA, OK, NE, NH, NC, MD, KS, UT, IA, KY, HI, OR, SC, ME and SD. | Terminated | April 09, 2012 | 789 | 104329000
Z-0847-2010 | 54641 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DigitalDiagnost. General radiographic examinations and applications wherever a solid state X-ray imaging device can be used. |  |  | January 21, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Device failed to display appropriate labels indicating Radiation Symbol along with the required wording. | Radiation Control for Health and Safety Act | Urgent - Field Safety Notice letters, dated January 20 &amp; 21, 2010, were sent to customers. The letter identified the affected products, gave a description of the problem &amp; the hazard involved, stated how to identify affected products, actions to be taken by the customer or user, and actions planned by the firm. Customers are to make sure that the relevant safety precautions are observed and check their operators console (EEC) if the labeling is present as shown in the drawing. If the labeling is missing, a Philips rep should be contacted. For further information or support concerning the issue, a Philips rep should be contacted at 800-722-9377. |  | 117 | 117 units | Devices were distributed through out the US in the following states:AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TX, UT, VA, WA, WI, and WY. | Terminated | January 25, 2011 | 369 | 144
Z-1196-2010 | 53816 | Yes | NSA | Automated external defibrillator system. | Cardiovascular | Cardiovascular | 510(k) | Heartstart HS1 Defibrillator Model M5068A |  |  | November 02, 2009 | 2009 | Philips Medical Systems | Potential for failure of a capacitor during use which would prevent the AED from delivering effective defibrillation therapy when indicated. | Nonconforming Material/Component | The firm notified their US distributors of the recall and instructed them to contact customers and provide information how to return the product to Philips for a free replacement. Firm contacted distributors via phone and email on 11/2/09 and by letter on 11/9/09. A web site providing information related to the recall was activated on 11/3/09. |  | 34 | 34 units distributed domestically | Worldwide Distribution. | Terminated | May 23, 2013 | 1298 | 2588
Z-1202-2010 | 54663 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Tessier Osseous Microtome TOM(R) Bone Mill, 01-15402, Stryker Portage, MI. |  |  | February 18, 2010 | 2010 | Stryker Craniomaxillofacial Division | The firm has received 328 customer complaints for this device. Most of the complaints were involving non-functioning of the ratcheting handle or complete device, or the presence of metal shavings or discoloration. | Equipment maintenance | An "URGENT MEDICAL DEVICE CORRECTION" letter dated February 18, 2010, was sent to the customers. The letter describes the product, problem, scope of affected products, potential patient risks and actions to be taken by the customers. The customers should review the letter with all users of the TOM(r) Bone Mill and follow the instructions on the proper inspection and maintenance of the device enclosed with the letter. The customers should fax the accountability form within 5 days to help Stryker ensure their receipt of the information. If the accountability form is not received, addition communication will be sent.If you have any urgent questions or concerns, please contact us at (269) 323-4258. Please ref: PFA #2009-128. Our normal business hours are Monday-Friday 8am-5pm (GMT+1:00)On 7/8/2010 the firm sent an Urgent Medical Device Recall letter to its customers. The letter stated thatt upon consultation with regulatory agencies outside of the United States, the firm had intiated a devicd removal. The letter instructed the customers to return any devices that they had to the recalling firm. |  | 596 | 596 units | Worldwide distribution: USA and Belarus, Bulgaria, Denmark, Finland, Hungary, Jordan, Norway, Russia, Saudi Arabia and the UAE. | Terminated | November 24, 2010 | 279 | 596
Z-1200-2010 | 54845 | No | DTB | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | PMA | Medtronic, Sutureless, unipolar, myocardial, screw-in pacing lead, 5071-15, 5071-25, 5071-35 &amp; 5071-53, Medtronic, Inc., Minneapolis, MN 55432, Manufactured in: Villalba, Puerto Rico. |  |  | February 25, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic has identified a packaging issue for a subset of Epicardial Leads Model 5071. They determined that specific package seals could be compromised. Medtronic is not aware of any patient injury due to this packaging issue and their evaluation indicated the risk of patient injury is remote. However, a compromised package seal could potentially affect product sterility. | Packaging | Consignees were sent a Medtronic "Important Medical Device Recall" Letter dated February 2010 on 2/25/10. The letter described the problem and the product. Medtronic recommended to immediately remove the affected leads from inventory. Customers with questions, please contact your Medtronic Representative or Medtronic Technical Services at 1-800-505-4636. |  | 4863 | 4,863 (US); 6,328 (OUS) | All States in continental USA including DC except Alaska.OUS: Algeria, Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Belgium, Bolivia, Bosnia And Herzegovina, Brazil, Canada, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, El Salvador, Finland, France, Germany, Greece, Guam, Guatemala, Honduras, Hong Kong, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lithuania, Luxembourg, Malaysia, Malta, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Vatican City, Venezuela, and Yemen. | Terminated | January 07, 2012 | 681 | 4863
Z-1199-2010 | 54453 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Avanta Fluid Management Injection System, Avanta Multi-Patient Disposable Set (MPAT) High Pressure Check Valve, Catalog number AVA 500 MPAT, Material number 3018231. |  |  | January 06, 2010 | 2010 | Medrad Inc | The product may be defective resulting in a reduction of the saline delivery rate and inadequate air purging. | Nonconforming Material/Component | The recalling firm issued an Urgent Medical Device Recall letter dated 1/6/10 to all customers that received the affected lots. The customers were instructed to determine whether or not they had the affected product lots. Customers are to discontinue use and contact MEDRAD Customer Support or their MIP Rep to make arrangements for return of the recalled product. Any affected product returned will be replaced at no charge to the customer. Customers are also to complete and sign the attached form and fax it back to the firm. Questions concerning this recall are to be directed to Deb Easler, Product Manager, at 412-767-2400. |  | 2625 | 2625 pieces | Worldwide Distribution -- USA, including stated of AL, AZ, CA, DC, FL, GA, IL, KS, MD, MO, NM, NY, OH,OK, PA, TN, TX, and WI. and countries of Belgium, Germany, and Italy. | Terminated | April 08, 2010 | 92 | 2625
Z-1205-2010 | 53321 | No | CDT | Triglyceride test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista System StreamLAB Analytical Workcell. Dimension Vista System Part Numbers 1000035896, 1000035899, 1000036563, and 1000036567. StreamLAB Analytical Workcell Part Number 746002.901, for in vitro diagnostic use. |  |  | September 01, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Incorrect result: An intermittent low result may be seen with the last test run from an aliquot on a Dimension Vista System when connected to a StreamLAB Analytical Workcell. The sample delivery sequence may deliver insufficient surplus volume to the low sensitivity aliquot (second aliquot). The surplus sample volume is occasionally insufficient to and can cause short sampling on the last test. | Software design | Siemens healthcare Diagnostics issued an Urgent Field Safety Notice letter to each customer regarding the issue dated September 2009. The letter informed the user of the problem and volume change instructions. Customers can call Technical Solutions Center at 800-441-9250 |  | 81 | 81 units | Worldwide distribution: USA, Australia, Canada, France, Germany, Malaysia, Netherlands, Spain, and Switzerland. | Terminated | September 15, 2010 | 379 | 81
Z-1203-2010 | 54481 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Vygon Nutrisafe 2 syringe 5 ml, For Oral use only with Nutrisafe 2 |  |  | January 29, 2010 | 2010 | Churchill Medical Systems, Inc. | Mislabeled feeding syringes : 6 ml syringe mislabeled as 5 ml. | Labeling Change Control | Churchill Medical issued an Urgent: Product Advisory Customer Notification on January 29, 2010 to customers and distributors via overnight delivery. The letter identified the affected product along with the corresponding lot numbers and explained the problem. Accounts are requested to complete the enclosed form and fax it to 215-672-6740, attention Marie Benson. The affected products should be quarantined and customer service should be contacted to arrange return of product. Questions should be directed to Marie Benson at 800-473-5414. |  | 29700 | 29,700 eaches | Nationwide Distribution -- CA, NM, NJ, NY, PA, VA, and TX. | Terminated | September 23, 2010 | 237 | 129600
Z-1206-2010 | 53501 | No | EXX | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | VIPA Veronikis Intra-Fiberoptic Pelvic Assistant, Set Catalog No 350-370. |  |  | April 09, 2009 | 2009 | Marina Medical Instruments Inc. | Marina Medical, Sunrise, FL initiated recall of 46 different 350 series devices and 47 different 270 series products because the material used to manufacture the products was not proven to be biocompatible, resulting in risk of temporary skin or tissue irritation. | Nonconforming Material/Component | A "PRODUCT RECALL NOTICE" letter dated April 8, 2009, was sent to the customers. The letter describes the product, problem and the action to be taken by the customers. The customers are asked to identify the subject part in their instrument inventory and quarantine it. The Quality Control team at Marina Medical will contact the customer with instructions to return the part. When contacted, the customer should report any incidents of adverse reactions that have been observed at their institution. If you have any questions, please call 1-800-697-1119 ore (954) 924-4418. |  | 47 | 47 | Nationwide distribution. | Terminated | September 03, 2010 | 512 | 96
Z-1210-2010 | 54669 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific, Ultra-thin" Diamond" Balloon Dilatation Catheters, M001164040 (DT/3-4/5/40, 16-404), M001164360 (DT/5-2/5/40, 16-436),M001164620 (DT/6-4/5/40, 16-462), Made in Ireland, Ballybrit Business Park, Galway, Ireland |  |  | February 18, 2010 | 2010 | Boston Scientific Corporation | Boston Scientific Corporation is conducting a Medical Device Recall Removal of specific lots/batches of the Blue Max and Ultrathin Diamond Balloon Dilatation Catheters. Boston Scientific has determined that, the sterile barrier in the catheter packaging (i.e. outer pouch side seal) of the affected products may be compromised. An undetected breach in sterility of the outer pouch may lead to contami | Packaging | Two "URGENT MEDICAL DEVICE RECALL REMOVAL -IMMEDIATE ACTION REQUIRED" letters dated February 17, 2010 and February 18, 2010 were sent to the customers. The letters described the Action Description, Affected Product and Recall Instructions. The customers should immediately discontinue use of and remove all affected recalled product from your inventory, segregate the product in a secure location for return to Boston Scientific, complete and return the enclosed Account Reply Verification Tracking Form (even if you do not have any product to return) -fax to: Customer Service Call Center (866) 213-1806, and package/ship the recalled product per instructions. Your local sales Representative can answer any questions that you may have regarding this Recall Removal.If you have any question, please contact (763) 494-7971. |  | 99 | 99 | Worldwide distribution: USA-AL, CA, FL, IL, MI, MS, NJ, NY, NC, PA, RI, SC, TN, VA, WI, and CANADA, JAPAN, GREAT BRITAIN, NETHERLANDS. | Terminated | December 10, 2011 | 660 | 172
Z-1209-2010 | 55008 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, Release 4.51.00. |  |  | February 22, 2010 | 2010 | Computerized Medical Systems Inc | Inconsistent MLC positions, dose, and monitor units | Software design | Elekta contacted customers by phone and mail beginning February 22, 2010, explaining the affected product and providing software update information.For further information, contact Elekta at 1-408-830-8023. |  | 25 | Planning system is loaded 25 in locations | Worldwide Distribution -- United States (PA, AZ, CA, MN, OH, WI, NC, FL, TX, NY and WI), Australia, The Netherlands and India. | Terminated | August 24, 2011 | 548 | 25
Z-1212-2010 | 53430 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Zimmer PTW4321 20 Ncm Dental Torque Wrench, QTY 1 Pro-Dex Inc, Irvine, CA 92614 |  |  | January 19, 2009 | 2009 | Pro-Dex Inc dba Dyna-Dent | Pro-Dex has initiated a recall on their Dental Torque Wrench. The recall was initiated due to an inaccurate laser etch identification of the torque value. | Error in labeling | A "Voluntary Firm Initiated Recall" letter dated January 22, 2009, was sent to the customer. The letter describes the product, problem and action to be taken by the customer. The customer should immediately cease distribution or use of the noted lot; evaluate inventory and distribution records to ensure that this lot is removed from any potential point of use, and return the product to Returns Department at Pro-Dex, Inc . and include the RMA number on the shipping box. If you have any questions or concerns regarding this matter, please call (949) 769-3200. |  | 70 | 70 | Nationwide distribution: CA | Terminated | July 06, 2011 | 898 | 70
Z-1213-2010 | 54616 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Axiom Iconos R200 X-Ray |  |  | January 29, 2010 | 2010 | Siemens Medical Solutions USA, Inc | table can tilt/move when it is not intended | Software design | The recalling firm has issued a letter to affected customers and will send technicians to apply the fix to each defective device. |  | 169 | 169 units | Nationwide Distribution. | Terminated | August 11, 2010 | 194 | 169
Z-1214-2010 | 54477 | No | LKK |  |  | General Hospital | PMA | Medtronic N'Vision 8870 Application Card for use in the Medtronic model 8840 Physician Programmer. Manufactured for: Medtronic, Inc., Minneapolis, MN 55432-5604 USA. |  |  | March 23, 2009 | 2009 | Medtronic Neuromodulation | Routine monitoring of Medtronic's vigilance program identified user related programming errors. Medtronic has received reports of both overdose and underdose that were the result of programming use of SynchroMed implantable infusion pumps. These errors were most commonly associated with: --Priming Bolus Programming; and --Bridge Bolus Programming. Pump programming errors that result in overdos | Other | Medtronic is exchanging 8870 versions AAG, AAJ and BBG and replacing it with the updated versions AAH01 and BBH01. The updated versions contain user interface improvements to reduce the chance of user related programming errors when programming SynchroMed II and SynchroMed EL Infusion Pump system.A Medtronic "Medical Device Correction" letter was left behind for consignees. The letter was addressed to "Healthcare Provider" and was dated February 2009. The letter provided information regarding changes to the programming software used in SynchorMed Infusion System. For additional information, please contact your Medtronic Representative or Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 8499 | 8499 US | All States in the US including DC.OUS: AustriaBelgiumCroatiaCyprusCzech RepublicDenmarkEstoniaFaroe IslandsFinlandFranceFrench PolynesiaGermanyGreeceHungaryIndiaIrelandIsraelItalyKuwaitLebanonMaltaMartiniqueNetherlandsNew CaledoniaNorwayPolandPortugalReunionRussian FederationSan MarinoSaudi ArabiaSouth AfricaSpainSwedenSwitzerlandTurkeyUnited KingdomVatican City StateCanadaMexicoBrasil | Terminated | July 10, 2012 | 1205 | 8499
Z-1079-2010 | 54181 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Acacia Catalog # 385165 "Y" IV Extension Set 20cm Macro Bore, Spin Nut w/BD Q-Syte |  |  | December 18, 2009 | 2009 | Medical Product Specialist Inc dba MPS Acacia | Some lots of IV Extension Sets with BD Q-Syte Luer Access Split Septum have a manufacturing deviation which could cause the product to not function properly, possibly resulting in an air embolism. | Process design | A letter dated December 18, 2009 was sent to Acacia customers with the following instructions. 1. Immediately review your inventory of Acacia IV Extension Sets with BD Q-Syte for REF numbers and lot numbers listed and remove them from all inventory printed on the unit packaging , shelf box and shelf shipper containers.2. Complete the Distributor Recall response Card (Attachment 1) and return by fax immediately to the number on the Distributor Recall Response Card for product in your facility at the time of receipt of this letter. This form must be completed and f axed to Acacia, Inc even if there is no product in inventory.3. Acacia, Inc is currently working to ensure that replacement products are readily available. Once you have identified how much product you will be returning, please call your Customer Representative at 1-800-453-4538, opt. 2, ext. 2534, to arrange for product replacement.Questions are to be directed to the Customer Representative at 1-800-453-4538, opt. 2, ext. 2534, between 8:00AM and 5:00PM MST, Monday-Friday. |  | 69726 | 69,726 | US Only | Terminated | January 05, 2011 | 383 | 136441
Z-1124-2010 | 54723 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 10 Drop IV Sets, with 3 Pre-Pierced Y-Injection Sites and 3-Gang Pressure Activated Manifold Valve. (Approx. 10 Drops/ML); with Two Y-Injection Sites and Rotating Luer Lock (Approx. 10 Drop/ML); with Positive Pressure Pump with 4-Way "Hi-Flo" Stopcock and 4 Pre-Pierced "Y" Injection Sites (approx. 10 Drop/mL). |  |  | February 01, 2010 | 2010 | Arrow International Inc | Sterility may be compromised. | Component design/selection | The recalling firm issued an "Urgent Medical Device Recall" letter dated February 19, 2010 to their consignees to inform them of the issue. Customers were instructed to check their inventory for the affected products, cease use and distribution, and quarantine all affected product immediately. Consignees were further instructed the following: 1) Complete the enclosed Recall Acknowledgement &amp; Stock Status Form and immediately fax the completed form to Arrow International at 1-800-343-2935 and 2) Return all affected product to Arrow International, Inc. 312 Commerce Place, Asheboro, NC 27203 Attn: Asheboro Quality Assurance Manager. If additional information or clarification regarding the recall is needed, please contact Arrow International, Customer Service Department at 1-800-523-8446. Details about the recall can be accessed on the firm's website at http://www.teleflexmedical.com/spotlight/ivTubingRecall/index.html. |  | 70959969 | 70,959,969 of total units | Worldwide Distribution -- United States, Canada, Columbia, South Africa, Turkey, Kazakhstan, Venezuela, Germany, Cuba, Panama, El Salvador, Costa Rica, Singapore, and United Kingdom. | Terminated | November 08, 2012 | 1011 | 70959969
Z-1073-2010 | 53880 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medical Action Industries, IV Start Bundle Kit, using MPS Acacia with BD Q-Syte (BD #385151), Part number: 69946B. |  |  | November 06, 2009 | 2009 | Medical Action Industries Inc | Manufacturing defect may allow air entry into a central venous catheter, possibly resulting in an air embolism. | Nonconforming Material/Component | On 11/6/2009 a letter was drafted and sent by email to three (3) Medical Action Industries customers. They were instructed to return the recalled products in inventory to MAI. They further advised, if the product was distributed to notify their customer, inform them of the recall and instruct them to return the affected lots to MAI. |  | 12000 | 12,000 kits | Nationwide distribution | Terminated | June 01, 2011 | 572 | 18800
Z-1215-2010 | 54677 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ADVANCE TIBIAL WEDGE AUGMENT, REF KTAG-W315, SIZE 3/2+, Thickness 15 DEG, SURFACE NON-POROUS, 1 EACH, Wright Medical Technology, 4677 Airline Road, Arlington, TN 38002. |  |  | February 19, 2010 | 2010 | Wright Medical Technology Inc | The product contained screws with the incorrect length (incorrectly contained 15 mm length screws instead of the correct 5 mm length screws). | Employee error | The firm initiated their recall of the product by letter on February 19, 2010. The notification letters were addressed to distributors, hospital administrators, and surgeons with a fax back response form. The letter described the affected product and reason for the recall. Customers were instructed to examine their inventory and return the recalled product to Wright Medical Technology, Inc. Customers may contact Customer Service by calling 1-800-238-7117 for return instructions and replacement inventory. Direct questions regarding this matter to Wright Medical Technology, Inc. by calling 1-800-874-5630. |  | 47 | 47 units | Worldwide Distribution -- United States, Canada, Netherlands, Germany, Japan, S. Korea, UK, Czech Republic, Turkey and Brazil. | Terminated | July 08, 2011 | 504 | 47
Z-1217-2010 | 54655 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Live Better 1cc, 31 gauge x 5/16 inch U-100 Insulin Syringe short needle. 100 syringes, Manufactured by UltiCare. |  |  | February 05, 2010 | 2010 | Ulti Med, Inc. | UltiMed, Inc. is recalling one lot of 1cc, 31 gauge x 5/16 inch U-100 Insulin Syringe Live Better because it is mislabeled and actually contains 1/2 cc, 30 gauge x 1/2 inch U-100 insulin Syringe. | Employee error | An " IMPORTANT - DEVICE PRODUCT RECALL"letter dated February 3, 2010, was sent to customers. The letter describes the product, problem and action to be taken by customers. The customers should check their inventory for the affected lot, complete the recall tracking form and return by mail or fax and contact UltiMed for return instructions.If you have any questions concerning this recall, please call UltiMed, Inc. at 877-854-3434 and ask to speak to Jill. |  | 2500 | 25 units (each dispense and dispose sharps unit contains 100 syringes) | Nationwide distribution: CT and NJ | Terminated | May 07, 2011 | 456 | 2500
Z-1216-2010 | 55074 | No | FNM | Alternating pressure air flotation mattress. | General Hospital | General Hospital | 510(K) Exempt | AccuMax Quantum Complete (Model CU3/CU4), an alternating pressure with turn assist &amp; on demand low air loss system; Encompass Therapuetic Support Systems, Glendale, CA. |  |  | March 05, 2010 | 2010 | Encompass Therapeutic Support Systems dba BG North America | Potential out gasing of the battery and short term ignition of the gas within the control unit enclosure. This failure mode includes a risk for fire. | Nonconforming Material/Component | An "URGENT RECALL NOTIFICATION" letter dated March 5, 2010, was sent to the customers. The letter describes the product, problem and action to be taken by the customers. The customers should immediately discontinue use of the control unit, turn off the unit and disconnect from the wall socket and quarantine the unit away from any patient's room. The customers should immediately examine your inventory and quarantine the products listed; identify and notify any facility where they may have further distributed or transferred the listed products; and complete and return the REPLY/FAX-BACK FORM as soon as possible to Fax: (626) 796-7856 or mail to EncompassTSS, 100 E. Corson Street, Ste 310, Pasadena, CA 91103.Please contact EncompassTSS Customer Service at 800-822-8288 if you have any questions regarding this recall notice. |  | 349 | 349 | Nationwide distribution: CA, TX, IN, OR, and MO. | Terminated | March 21, 2012 | 747 | 349
Z-1218-2010 | 54742 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Harmony Lux Advantage LA Surgical Lighting and Visualization System, also known as the Harmony LA 500, B129382391. |  |  | February 23, 2010 | 2010 | Steris Corporation | Premature Bulb Failure | Device Design | STERIS issued an "Urgent Field Correction Notice" dated February 23, 2010 via Fed Ex. Consignees were informed of the affected product and were instructed on proper handeling of the device until a STERIS Service Technician contacts the consignee to schedule a service visit.For further information, contact STERIS Field Service Dispatch at 1-800-8828. |  | 10433 | 10,433 units | Worldwide Distribution -- United States, Canada, Algeria, Australia, Bahrain, Bangladesh, Brazil, Chile, China, Costa Rica, Ecuador, El Salvador, France, Germany, Guam, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Malaysia, Mexico, New Zealand, Nigeria, Pakistan, Philippines, Poland, Qatar, Russia, Saudi Arabia, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Arab Emirates, United Kingdom, Italy, Yemen and United Arab Emirates. | Terminated | June 04, 2012 | 832 | 29460
Z-1179-2010 | 54719 | No | KLC | Dental chair and accessories. | Dental | Dental | 510(k) | Pelton &amp; Crane Dental Chairs, Chairman 5000 with magnetic headrest, Pelton &amp; Crane, Charlotte, NC 28273. |  |  | February 24, 2010 | 2010 | Pelton &amp; Crane Company | Magnetic headrests may affect patients with implantable pacemakers and defibrillators. | Device Design | Pelton &amp; Crane Issued an "Urgent Field Safety Alert" notification dated February 24, 2010. Consignees were instructed of the affected product and informed of proper actions to take. For further information, contact Pelton &amp; Crane at 1-704-588-2126. |  | 498 | Total: 498 chairs and 619 replacement upholstery spare parts | Worldwide Distribution -- United States, Australia, the Netherlands, Canada and Germany. | Terminated | October 28, 2010 | 246 | 498
Z-0975-2010 | 53761 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PTS Panels CHOL+HDL Panel test strips for professional use with CardioChek P-A analyzer, Manufactured by Polymer Technology Systems, Inc., Indianapolis, IN.Catalog # 1795. |  |  | February 20, 2009 | 2009 | Polymer Technology Systems, Inc. | Extended expiration date; the product expired on February 12, 2009 but was labeled with a April 4, 2009 expiration date. | Process control | Consignees were notified by letter dated Febuary 20, 2009, advised of the problem, and that the firm would replace product not consumed by the true expiration date. Direct questions to Polymer Technology Systems, Inc. |  | 3748 | 3,748 vials | Worldwide Distribution. | Terminated | October 20, 2010 | 607 | 3748
Z-1222-2010 | 54656 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Publix Sterile Preserved Saline Solution 12oz, Product SKU#041415-003767. |  |  | February 03, 2010 | 2010 | K C Pharmaceuticals Inc | The recall was initiated because the affected lots were manufactured between two failed media fills; therefore the lots produced cannot be assured as being free of contamination and may not be sterile. | Other | An "Urgent: Medical Device Recall Notification" letter dated February 17, 2010 was issued to customers. Consignees were instructed to immediately hold and discontinue selling the existing, affected stock.Consignees were also instructed to check inventories at all distribution locations for affected products shipped between June 8, 2OO9 and August 7, 2009 and perform a SUBRECALL to the retail level and report back to KC Pharmaceuticals. If the retailers indicate they have any of the affected lots of product, please have them contact KC Pharmaceuticals at 1-888-527-4276 for a shipping authorization number.Consignees were asked to complete the enclosed Recall Effectiveness Check Form, even if they did not have the products in question and return the form to KC Pharmaceuticals via mail within five (5) business days. The recalled lot and completed packing slips were asked to be shipped to:KC Pharmaceuticals-SALP Recall,3201 Producer Way,Pomona, CA 91768.Direct questions to KC Pharmaceuticals by calling 1-888-527-4276, extension 27. |  | 2004 | 2,004 units | Nationwide Distribution. | Terminated | January 05, 2011 | 336 | 46428
Z-1228-2010 | 51926 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 10 degree Tibial Augment Wedge, a component of Ortho Balanced Knee Modular Tibial System. |  |  | March 19, 2009 | 2009 | Ortho Development Corporation | A component of knee prosthesis was assembled incorrectly. | Device Design | E-mail was sent to all agents and distributors on March 19, 2009 with illustrations and instructions on how to detect the upside down assembly and how to correct the problem if detected during surgery. Direct questions to the Ortho Development Corporation by calling 1-801-619-3444. |  | 422 | 422 units | United States (AZ, FL, OR, TX and UT) and Japan. | Terminated | June 30, 2010 | 468 | 790
Z-1161-2010 | 53214 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System IS2000 Endoscopic Instrument Control System, rev. A51_P7, manufactured by Intuitive Surgical, Sunnyvale, CA. |  |  | July 15, 2009 | 2009 | Intuitive Surgical, Inc. | Gripper or Scissor jaws may close inadvertently, and will not open on command and various other reported modes of failure. Control by surgeon may fail and this failure may be difficult to detect. | Software design | Recall initiated the recall on July 1, 2009. US consignee recall was initiated on July 7, 2009. All affected consignees have been notified via e-mail communication followed by a field service engineer visit. Customers with questions or with need for additional assistance regarding the notification, please contact the firm's Intuitive Surgical Customer Service Department at 1-800-876-1310 toll free. |  | 782 | 782 units | Product was distributed to 551 consignees throughout the US to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, WV AND WY, as well as to Puerto Rico, and to Austria, Belgium, Bulgaria, Czech Republic, Denmark, Finland, France, Germany, Greece, Ireland, Italy, Netherlands, Norway,Qatar, Romania, Spain, Sweden, Switzerland, Turkey, the United Kingdom, Argentina, Australia, Brazil, Canada, China, India, Israel, Japan, Korea, Mexico, Russia, Saudi Arabia, Singapore and South Korea. | Terminated | February 11, 2011 | 576 | 782
Z-1235-2010 | 54284 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD-LEGACYÂż 1 Model 6400, Ambulatory Infusion Pump, |  |  | December 23, 2009 | 2009 | Smiths Medical ASD, Inc. | The CADD-LegacyÂż 1 Model 6400 Ambulatory Infusion Pump (Serial Number 370969) was shipped without documentation that the required accuracy testing had been performed during the manufacturing process. If the subject infusion pump did not go through the accuracy testing, and if it does not meet accuracy specifications, there is a potential for over or under-delivery. | Employee error | Smiths Medical telephoned the customer on December 23, 2009, to return the CADD-LegacyÂż 1 Ambulatory Infusion Pump, serial number 370969, since there was no record of the completed accuracy testing for the infusion pump. |  | 1 | 1 | MI | Terminated | December 10, 2011 | 717 | 1
Z-1232-2010 | 55063 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion, Product Number: 715000. Radionuclide radiation therapy system. Elekta, Inc. Norcross, GA. |  |  | November 06, 2009 | 2009 | Elekta, Inc. | There may be a situation where the Frame Adapter might lock the stereotactic Frame in the wrong position. | Software design | Elekta, Inc. issued a Field Change Order 200 091 titled "Update Frame Adapter with New Frame Fixation Axle" dated November 6, 2009. Consignees were given instructions to prevent further issues related to the affected device.For further information, contact Elekta, Inc. at 1-770-300-9725. |  | 32 | 32 units | Nationwide Distribution -- AZ, CA, FL, IL, KY, MI, MN, MS, MO, NY, NC, OH, OR, PA, SC, TX, UT and VA. | Terminated | December 14, 2010 | 403 | 32
Z-1231-2010 | 54886 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion. Radionuclide radiation therapy system. Elekta, Inc. Norcross, GA. |  |  | May 30, 2007 | 2007 | Elekta, Inc. | There has been an issue with "Image Fushing" where low precision calculation caused images to become inaccurate. | Software design | Elekta, Inc. issued a Field Change Order, Reference Number 0000946 dated March 14, 2008. Consignees were informed of the affected product and instructed on necessary steps to take.For further information, contact Elekta, Inc. at 1-770-300-9725. |  | 18 | 18 units | Nationwide Distribution -- AR, AZ, CA, FLA, IL, MI, MN, MO, MS, NJ, OH, OR, PA, UT and VA. | Terminated | December 28, 2010 | 1308 | 18
Z-1237-2010 | 54999 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis Systems operating with A100 Generator |  |  | March 04, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Due to a contact problem at the low voltage power supply of the generator controller, in rare cases where a user has invoked radiation, it may occur that no radiation is released. Under certain circumstances, radiation is possible by repeated release of the x-ray footswitch or button.To remedy the issue described above, a revised cabling has been installed in all potentially affected syste | Nonconforming Material/Component | The recalling firm issued a Customer Safety Advisory to affected customers via Update Instructions AX006/10/S. The letter informs customers of the potential issues and provides additional information on the implemented resolutions. For additional information, please contact the firm at 888-826-9702. |  | 1373 | 1373 units | The product was shipped to medical facilities nationwide. | Terminated | July 05, 2012 | 854 | 1373
Z-1230-2010 | 54564 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion, Radionuclide radiation therapy system. Article Number 715000, Elekta, Inc. Norcross, GA 30092. |  |  | November 09, 2008 | 2008 | Elekta, Inc. | Radiation unit doors could close too fast on emergency exit. | Software design | Field Change Order 200 070 "MCU CS SW 8.3" dated September 11, 2008 was released. Consignees were informed of the necessary steps to avoid further issues relating to the affected product.For further information, contact Elekta, Inc. at 1-770-300-9725. |  | 19 | 19 units | Nationwide Distribution -- AR, AZ, CA, FL, IL, MI, MN, MS, MO, NJ, OH, OR, PA, UT and VA. | Terminated | December 28, 2010 | 779 | 19
Z-1238-2010 | 53993 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Vertier Surgical Table, Catalog numbers 0788100000, 0788100000S, 0788100001, and 0788100001S |  |  | November 13, 2009 | 2009 | Stryker Communications Corp | In certain positions, product is susceptible to fluid instrusion which could possibly result in short circuit or unsolicited movement. | Other | Firm notified all consignees by letter sent 11/13/2009. Contact the firm at 972-410-7100 with questions relating to this recall. |  | 206 | 206 | United States (GA, AR, ND, KY, CO, CA, SD, PA, OH, TX, SC, NY, MS, IL, KS, NJ, OR, WA, MN, VA, IL, OK, WY, IN, ID, UT, MO, FL, and DC) and Canada. | Terminated | August 11, 2010 | 271 | 206
Z-1234-2010 | 54279 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ATAC Serum Control Kit. |  |  | December 04, 2009 | 2009 | Vital Diagnostics, Inc. | Vital Diagnostics observed a bias in the Quality Control Consensus Program between published values for ALP on the ATAC 8000 and reported customer means.The ranges were adjusted to reflect the analyzer population. | Other | Customers were mailed copies of Technical Bulletin 8100 and 5024 dated September 30, 2009 which advised of the Alkaline Phosphatase QC value revisions for the Envoy 500 and the ATAC 8000. If you have any questions, please contact Vital Diagnostics Technical Support at 1-800-345-2822. |  | 802 | 802 Kits | Nationwide. | Terminated | July 25, 2012 | 964 | 1673
Z-1251-2010 | 54144 | No | DXX | Steerable catheter control system. | Cardiovascular | Cardiovascular | 510(k) | Artisan brand Extend Control Catheter; An Steerable Guide Catheter; Component of Sensei X Robotic Catheter System; Model Number: 07546; Distributed and Manufactured by Hansen Medical, Inc, Mountain View, CA |  |  | October 19, 2009 | 2009 | Hansen Medical Inc | Certain Artisan Extend Catheters may develop a leak in the flush assembly of the catheter which has the potential to cause flush fluid leakage and subsequent decrease in distal flush; inability to aspirate air from flush lines, and/or introduction of air into the catheter. | Other | Recall Initiation Date: October 19,2009 (FedEx shipments on October 19, for October 20 delivery) via letters explaining the reason for the recall and requesting the product be returned. Customers with questions are instructed to call the company at 650-404-5940. |  |  |  | Nationwide to 12 states. | Terminated | February 09, 2011 | 478
Z-1250-2010 | 54540 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | MaxTorque Cannulated Screw Model # MSD-010-55-060S |  |  | April 07, 2009 | 2009 | Orthohelix Surgical Designs Inc | A pinhead sized disc was found in the cannulation of a MSD-010-55-60S screw. | Other | On April 7, 2009 the firm sent letters to their customers with detailed instructions on how to correct the problem. Pictures were included for clarification.For questions please contact either Jim Bragg via phone at 330-869-9562 ext 211 or Kristin Wolff at ext 209. |  | 45 | 45 pieces | Product was delivered to: CA, CO, GA, NM, NY, OH, VA &amp; WA. | Terminated | October 20, 2010 | 561 | 45
Z-1244-2010 | 54517 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Powerheart 9300E automated external defibrillator. |  |  | February 03, 2010 | 2010 | Cardiac Science Corporation | A component of the AED device may fail during a rescue attempt and the AED will not be able to deliver therapy. | Nonconforming Material/Component | The firm began initial notification on 2/3/2010 by sending letters to all affected consignees via USPS certified mail (domestic) and FedEx (international). Also on 2/3/2010 the firm issued a global press release and posted recall information on their website.Customers are instructed to remove their affected units from service, and have been given instructions for returning their affected devices. Replacement AED(s) will be sent as they are available. For questions, please contact the firm's Cardiac Science Support at 1.888.402.2484 or. AED195@cardiacscience.com to arrange for replacement(s). |  | 2640 | 2640 | Distributed to health care and emergency medical personnel and US Gov't/Military in the US and the following 44 countries:AUSTRALIA, BELGIUM, BELIZE, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, DENMARK, EGYPT, FRANCE, GERMANY, GRENADA, GUATEMALA, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, MALAYSIA, MALTA, MEXICO, NETHERLANDS, NORWAY, POLAND, PORTUGAL, PUERTO RICO, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, TAIWAN, THAILAND, TRINIDAD, AND GREAT BRITAIN. | Terminated | December 08, 2011 | 673 | 10340
Z-1252-2010 | 54934 | No | NSX |  |  | General Hospital | Contact ODE | CareSuite" - Critical Care Manager, PACU Manager and Anesthesia Manager |  |  | March 08, 2010 | 2010 | Picis Inc. | Under specific timing of conditions and in configuration with 3rd party infusion pumps, an error within the clinical application causes the manual documentation of clinical orders to be changed to an automatic update status. At this point, the order documentation becomes frozen at the last dose and/or rate and the user is unable to modify, stop or delete it. This error could impact the fluid bala | Software design | Picis issued letter notification beginning on March 8, 2010 with software correction information. The letter communicates details of the problem, along with recommendations for mitigating the potential to experience this error condition. Use of the product should be discontinued until the update described in the letter is performed. Customers with questions should call the company at 781-557-3000. |  | 23 | 23 | NJCanada, Australia, Austria,Denmark, Finland, Germany Norway,Portugal, Sweden, France,Netherlands, Spain. | Terminated | July 01, 2013 | 1211 | 23
Z-1254-2010 | 54618 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Qwix 4.3mm Stabilization Screws - 34mm |  |  | February 08, 2010 | 2010 | Integra LifeSciences Corporation | The screws were etched incorrectly as 32mm instead of the correct 34mm length. | Labeling mix-ups | Consignees were notified via conference calls and/or FedEx mailing. Recall aknowledgement forms were provided with written notification. This requires consignees to review their inventory and to indicate any affected lots . This form requires a signature and must be returned. |  |  |  | Nationwide distribution - Alabama, Arizona, Colorado, Connecticut, Maryland, Michigan, New England, New Jersey, North Carolina, Ohio, South Carolina, Nebraska, and Utah. | Terminated | October 27, 2011 | 626
Z-1255-2010 | 54702 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Foundation Knee System, Patella Trial, 26 mm, part number 802-01-019, Lot number 53944551 |  |  | February 22, 2010 | 2010 | Encore Medical, Lp | One lot of product was incorrectly manufactured. Pegs are larger than specification. | Process control | Firm notified consignees by e-mail on 2/22/2010. |  | 2 | 2 units | United States (FL and ID). | Terminated | July 22, 2010 | 150 | 2
Z-1242-2010 | 54891 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PROMOS Inclination set, SIZE 26mm, REF 42 186, S&N; 75006725, QTY: (1), STERILE R, Smith &amp; Nephew Orthopaedics AG, CH-Rotkreuz-Switzerland. |  |  | January 14, 2010 | 2010 | Smith &amp; Nephew Inc | Firm received complaints of Inclination Set Screws fracturing. | Device Design | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall overnight mail on January 23, 2009. All affected Smith &amp; Nephew International Distributors were notified via email and telephone on January 23, 2009. They were instructed to immediately quarantine the product for return to Stericycle, Inc., 2670 Executive Dr. Suite AIndianapolis, IN 46241. Direct questions to Smith &amp; Nephew, Inc. by calling 1-901-399-6153. |  | 362 | 362 units | Worldwide Distribution: United States, Australia, Austria, Belgium, China, Denmark, France, Germany, Greece, Great Britain, Italy, Netherlands, Poland, Portugal, Spain, Sweden and Switzerland. | Terminated | April 10, 2012 | 817 | 3025
Z-1256-2010 | 54982 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | NicoletOne vEEG System, VIASYS Healthcare, Inc., Madison, Wisconsin 53711. |  |  | March 08, 2010 | 2010 | CareFusion | CareFusion NeuroCare is voluntarily implementing a field correction to delete certain NicoletOne software protocols, which users may interpret incorrectly when conducting patient evaluations. This may result in the misinterpretation of the left versus right side of the brain.Users are being provided with instructions for removingthese protocols and verifying that the deletion has beenprop | Software design | Consignees were sent a CareFusion "Urgent Medical Device Correction" letter dated March 3, 2010. The letter described the problem and product involved. The letter included "Action to be taken by Users," "Instructions for deleting protocols," and "Photos-Amplifier label and Headbox insert to be discarded". It also included a "Verification Form" to be completed by recipient.For questions, notify CareFusion-NeuroCare, Attention: Sue Niesen, at (800) 356-0007 ext. 5129, or via email at susan.niesen@carefusion.com |  | 6131 | 2,984 US; 3,147 OUS | All States in the United States including DC.UNITED ARAB EMIRATES, ALBANIA, ARGENTINA, AUSTRIA, AUSTRALIA, AZERBAIJAN, BANGLADESH, BELGIUM, BULGARIA, BAHRAIN, CANADA, SWITZERLAND, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECHOSLOVAKIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, ESTONIA, EGYPT, SPAIN, ETHIOPIA, FINLAND, FRANCE, GREAT BRITAIN, GREECE, GUAM, GERMANY, HONG KONG, HUNGARY, CRO ATIA, INDONESIA, IRELAND, ISRAEL, INDIA, ICELAND, ITALY, JORDAN, JAPAN, KOREA SOUTH, KWAI T, KAZAKHSTAN, LEBANON, LIECHTENSTEIN, LATVIA, MOROCCO, MALDOVA, MACEDONIA, MALTA, MAURITIUS, MEXICO, MALAYSIA, NIGERIA, NETHERLANDS, NORWAY, PANAMA, PERU, PHILIPPINES, PAKISTAN, POLAND, PUERTO RICO, PALESTINIAN TERRITORY, QATAR, ROMANIA, SERBIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVEN IA, SLOVAK REPUBLIC, THAILAND, TURKEY, TAIWAN, URUGUAY, UKRAINE, VENEZUELA, YEMEN, VI ET NAAM, YUGOSLAVIA, SOUTH AFRICA, | Terminated | July 05, 2012 | 850 | 6131
Z-1257-2010 | 53869 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | AC power cords manufactured by Electri-Cord with tall bridge design. Used across all Philips product lines, as well as Respironics/Philips products, Dixtel owned products and Dixtal (formerly Novametrix or Respironics/Novametrix products. Philips HeartStart MRx Defibrillator/Monitor model number M3535A, ID number 861288. |  |  | November 19, 2009 | 2009 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock. | Nonconforming Material/Component | Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside plug. The issue affects medical device power cords equipped with a plug that has a prong and ground-pin insert design and a black plastic bridge connecting the terminal prongs on the plug ("taller bridge" design). Medical device power cords that do not have a black bridge connecting the terminal prongs on the plug are not affected. Urgent Medical Device Correction letters were sent to Distributors describing the problem and giving them instructions for notifying customers and actions to be taken. Customers are to visually inspect the affected cords, complete the Customer Reply Power Cord Order Form, and follow the instructions to dispose of affected cords. Further information or support concerning this issue can be obtained by contacting a Philips representative at 1-800-722-9377. |  | 16705 | 16,705 | Worldwide Distribution -- UNITED STATES, ALGERIA, ARGENTINA, Aruba, AUSTRALIA, BAHRAIN, BANGLADESH, BELGIUM, BRAZIL, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, Denmark, ECUADOR, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAMAICA, JAPAN, KUWAIT, LEBANON, LIBYAN ARAB JAMAHIRIYA, MALAYSIA, MAURITIUS, MEXICO, NETHERLANDS, NETHERLANDS ANTILLES, NEW ZEALAND, NICARAGUA, NORWAY, PAKISTAN, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, RUSSIA, SAINT LUCIA, SAIPAN, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIET NAM, and VIRGIN ISLANDS. | Open, Classified |  |  | 38473
Z-1310-2010 | 54823 | No | NAD | Occlusive wound dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Aplicare Advantage Dressing Change System, Central Line Dressing Change, Sterile |  |  | March 03, 2010 | 2010 | Aplicare, Inc. | Incorrect expiration date on outer kit label-extended. | Error in labeling | Aplicare notified customers by telephone on March 3, 2010 followed with "Product Recall Notice" letter issued by e-mail to the distributor level. Product will be labeled with correct expiration date. Direct questions to your sales representative or customer service by calling 1-800-760-3236. |  | 103 | 103 cases (20 each) | United States (CT, RI, NJ and OH). | Terminated | March 16, 2012 | 744 | 103
Z-1313-2010 | 54851 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Datascope AS3000 Anesthesia System; |  |  | March 05, 2010 | 2010 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | An issue affecting the caster mount on the AS3000 System have been identified, specifically reports of the caster mount breaking, causing the wheel to fall off and the unit to tip. | Other | An "URGENT PRODUCT FIELD CORRECTIVE ACTION" letter dated March 5, 2010 was sent certified mail to all facilities/customers that have the AS3000 units. The letter described the product, problem and corrective Action/remediation by the firm. A Mindray Service Representative will contact the customers to arrange for the caster mounts to be replaced on the AS3000 units in their facilities at no cost.If you have any questions, please contact Karen Maine of the Mindray DS USA, Inc. Marketing Department at 201-995-8612 or k.maine@mindray.com. |  | 479 | 372 units in US, 107 internationally. | Worldwide Distribution: USA and Mexico, Columbia, Turks &amp; Caicos Islands, BWI; Ecuador, Bolivia, Peru, Dominican Republic, Netherlands Antilles, Chile.SA | Terminated | April 25, 2012 | 782 | 479
Z-1315-2010 | 54981 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | KODAK T-MAT G/RA Film, Catalog/REF # 151 0023, 24x30 cm, 5-100 Sheet Boxes FOLHAS/Hojas, Made in U.S.A. by Carestream Health, Inc. 150 Verona Street, Rochester, NY 14608 --- Device Listing # E420904 |  |  | March 12, 2010 | 2010 | Carestream Health, Inc. | Product mix-up. The firm received a customer complaint indicating that the customer had opened a box of film and although it was supposed to be T-MAT L/RA, it was performing like T-MAT G/RA. | Mixed-up of materials/components | A "MEDICAL DEVICE RECALL (VOLUNTARY)" letters and Response Forms (dated 3/12/10) were sent to the customers on 3/12/10 via Certified Mail, Return Receipt Requested. The foreign sales offices are being notified via e-mail (started on 3/12/10) so that they in turn can notify the foreign consignees of the recall action and report back to Carestream Health. The letter describes the products, problem and action to be taken by the customers. The customers should identify and remove from inventory shelves any product that is listed in the recall, complete and return the attached Product Information Form whether they have affected product or not by fax to Carestream Health Customer Service 1-800-445-9967. Carestream Health will organize the return through designated carriers and will pay for transportation; and supply replacement stock in a timely manner.If you have any questions, concerns, or product technical questions, please call or fax Customer Service or Technical support at 1-800-328-2910. |  | 875 | 875 boxes | Worldwide distribution: USA and Puerto Rico, Australia, Canada, Chile, El Salvador, Honduras, and New Zealand | Terminated | July 07, 2010 | 117 | 1435
Z-1312-2010 | 54846 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis Systems Operating with VB22, VB23, VB30, and VB31. Model number 5917054, 7008605, 7412807, 7413078, 7555357, 7555365, 7555373, and 7728392. |  |  | February 16, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Communication problems can occur which may result in unavailability of Bypass Fluoro. | Device Design | The recalling firm issued a Customer Safety Advisory Notice dated 2/16/10 to inform their customers of the problem. The letter identified the affected product and stated that the problem had been solved by the service technician by the time that customers would receive the letter. The letter also explained when the problem occurred and the possible risks associated with it. In addition, the firm stated what will be done to avoid the problem. |  | 640 | 640 units | Nationwide Distribution | Terminated | July 05, 2012 | 870 | 640
Z-1316-2010 | 53479 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT System Software v5.00 and v5.10, List Number 5F48-22 and 5F48-23; Part Number 7-204717-01 and 7-204717-02 when used with PM Algorithm LLS Brd, part number 7-94255-01 and PM LLS kit, U10 Chip and Ejectors, part number 7-200788-01. |  |  | September 11, 2009 | 2009 | Abbott Laboratories, Inc | ARCHITECT System Software v5.00 and v5.10 are incompatible with ARCHITECT LLS Board. Software patch re-establishes correct parameters. | Software design | Technical Service Bulletins dated September 11, 2009 were issued to affected consignees. The correction was issued via the Technical Service Bulletins (TSB) to provide a software patch to re-establish the correct parameters on the LLS board. Direct questions to Abbott Laboratories, Inc. by calling 1-972-518-6000. |  | 115 | 11 in the USA; 104 worldwide | Worldwide. | Terminated | January 25, 2011 | 501 | 115
Z-1311-2010 | 54809 | No | KLC | Dental chair and accessories. | Dental | Dental | 510(k) | Pelton &amp; Crane Spirit 3000 Dental Chair. |  |  | February 09, 2010 | 2010 | Pelton &amp; Crane Company | The Dental Chair lacks a valid Device History Record. | Other | An "URGENT MEDICAL DEVICE RECALL" letter dated February 9, 2010, was delivered in person to the customer. The letter describe the product, problem and action to be taken by the customer. The customer was requested to provide the name and location of the consignee who received the chair so that Pelton &amp; Crane can contact the end user to arrange to replace the affected chair with a new dental chair, and to complete the enclosed Recall Acknowledgement/Return Form and fax to (704) 587-7204. A monthly status report will be submitted to FDA.Please contact Pelton &amp; Crane Customer Care at 1-800-659-6560 or call Mr. Frank Ray directly at 704-587-7227 to let us know where the device is currently stored. |  | 1 | 1 unit | Nationwide distribution: NC | Terminated | August 11, 2010 | 183 | 1
Z-1317-2010 | 54892 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator Monitor (Models M3535A, M3536A and M3536J) used with Philips 14.8-Volt Lithium Ion Battery (ModeI M3538A). |  |  | March 04, 2010 | 2010 | Philips Healthcare Inc. | Batteries older than 2 years have a higher potential for intermittent electrical contact, may contribute to an interruption in monitoring or therapy if the monitor/defibrillator is operating on battery. | Component design/selection | An "Urgent Medical Device Correction" letter with a Field Safety Notice was issued on March 4, 2010 to affected customers. The Instructions for Use Addendum/Application Note is included with the communication describing when to replace the unit's 14.8-Volt Lithium Ion Battery (Product #M3538A) Number M3538A in the the Philips HeartStart MRx Monitor/Defibrillator. Should you have any questions or concerns about the Device Correction, please contact your local Philips representative. |  | 37674 | 37,674 units | Worldwide: UNITED STATES, CANADA, AFGHANISTAN, ALBANIA, ,ARGENTINA, ARUBA,AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM BOSNIA&HERZEGOVINA;,BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CHILE, CHINA, COLOMBIA, CROATIA,CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, FRANCE,GERMANY, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA,IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KUWAIT, LATVIA, LEBANON,LIBYAN ARAB JAMAHIRI, LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, NAMIBIA,NETHERLANDS, NETHERLANDS ANTILLES, NETHERLANDS ANTILLES, NEWCALEDONIA, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PALAU,PANAMA, PAPUA NEW GUINEA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTORICO, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE,SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN,SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY,UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIETNAM and YEMEN. | Completed |  |  | 37674
Z-1319-2010 | 54375 | No | HWR | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | BIOMET MICROFIXATION, PATIENT DISTRACTOR DRIVER 10.87cm, (4.28")(L), RADEL, STAINLESS STEEL, NON-STERILE PRODUCT. |  |  | January 15, 2010 | 2010 | Biomet Microfixation, Inc. | The patient Distractor Driver, used in Osteotomy procedures, may not ratchet and advance the distraction device as intended. If the distraction device does not advance, bone consolidation may occur and a second surgery may be necessary. | Device Design | Consignees will be notified by letter on 01/15/2010 be either fax, email or mail. |  | 73 | 73 | US ( AZ,CA,DC,FL,GA,ID,IL,IN,MO,NC,TX,UT), Japan, Finland, Africa | Terminated | May 14, 2010 | 119 | 73
Z-1320-2010 | 52175 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | ThermoFisher Scientific Microm STP 420D automated tissue processor. |  |  | June 03, 2009 | 2009 | Thermo Fisher Scientific | The firm has received complaints regarding specimen integrity (higher rate of specimen fragmentation) as well as complaints that the chamber may stop rotating prematurely, resulting in sample deterioration. | Other | Consignees were notified of the problem by letter, on or about 6/3/09, updating the instructions for use. The units were subsequently upgraded on or before 9/8/09. For additional questions and information, please contact ThermoFisher's Technical Support department , Subject: PT Module Advisory Notice. |  | 16 | 16 | Hawaii, Kentucky, Louisiana, Michigan, Missouri, Nebraska, New York, North Carolina, Ohio, Pennsylvania, Tennessee, Texas and West Virginia. | Terminated | June 29, 2011 | 756 | 16
Z-0762-2010 | 54169 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Reflex Hybrid Screw Extractor; |  |  | October 12, 2009 | 2009 | Stryker Spine | The threaded tip of the Reflex Hybrid Screw Extractor has been found to fracture the head of the screw, causing problems in removing the screw and possible damage to the screw and/or plate while attempting to remove the screw. | Employee error | An "Urgent Product Recall" letter was sent to Stryker Branches/Agencies on October 12, 2009 and to surgeons who have used the device on October 26, 2009. The recall letter described the affected products, issue and potential hazards. Consignees were instructed to examine their inventory and hospital locations for the affected products and return to Stryker Spine as indicated in the recall letter. Customers were also instructed to complete the enclosed Customer Response form and fax to 1-201-760-8370.Direct questions about the recall to Stryker Spine by calling 1-201-760-8206. |  | 1791 | 1,791 | Worldwide Distribution. | Terminated | September 06, 2012 | 1060 | 2774
Z-1170-2010 | 54087 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | LM Aculight Corporation Perseus Pulsed Fiber Laser, Model number: ML36L-1541. Additional Model numbers with non-defective type"A" fiber: M22L-1540, M27L-1541, M281-1540, M30PL-1541, &amp; M33L-1541. Lockheed Martin Aculight, Bothell, WA, 98021. |  |  | May 01, 2009 | 2009 | Lockheed-Martin Aculight | Defect of optical fiber - The Perseus device with type "B" fiber has been found to emit potentially hazardous laser radiation. | Radiation Control for Health and Safety Act | A Radiation Warning Notification was sent to all customers. The notification stated the purpose and problem. In addition, a Notification of Potential Perseus Laser Safety Issue letter was sent to the only customer who received the defective type "B" fiber product. The letter identified the affected product and stated that the product was not considered eye-safe. The firm recommended that the customer cease using the laser with the Part Number: 35-001466-032 and Corresponding Unit Serial Numbers: 00001,00002, and 00003. The customer is to return the affected product for additional testing and warranty repair. Questions should be directed to Rick Whitty on 425-482-1100 or at Richard.j.whitty@lmco.com. |  | 3 | 3 | Worldwide Distribution -- USA, including states of PA, FL, MD, &amp; CO and the countries of Austria and Canada. | Terminated | April 30, 2012 | 1095 | 3
Z-1318-2010 | 54729 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems Accu Tnl Reagent Kit when used in conjunction with UniCel Dxl Access Immunoassay Systems; Reagent Part Numbers: A78803 &amp; 33340; Calibrator Part Number: 33345 |  |  | February 06, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports that different results have been obtained usingthe same patient samples on Access/Access 2 and UniCel Dxl platforms. Values obtained withUniCel Dxl systems have been demonstrated to have a positive bias compared to valuesobtained with Access or Access 2 systems. | Other | An "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated February 5, 2010, was sent to the customers. The letter describes the product, problem and action to be taken by the customers. The customers are recommended to immediately discontinue using AccuTnI Reagents and Calibrators on the UniCel DxI system until further notice. If this is not possible, the customers are directed to notify their stakeholders of the bias until further notice. The customers should complete, fax (786-639-4000) or mail the enclosed RESPONSE FORM.If consignees need assistance or have any questions regarding the notification, they were instructed to contact TechnicalSupport at 1-800-854-3633 in the United States and Canada. |  | 196071 | 196,071 kits | Worldwide distribution: USA and Canada | Terminated | May 09, 2012 | 823 | 196071
Z-1321-2010 | 54930 | No | NYI | Gene expression profiling test system for breast cancer prognosis. | Immunology | Immunology | 510(k) | Agendia MammaPrint, Catalog Numbers MP01, MP US version 03 |  |  | March 03, 2010 | 2010 | Agendia Inc | Over a period of about 6 months in about 15% of results provided Agendia has been "over reporting" the chance of metastasis risk being 29% risk of recurrence instead of 10% risk.This over reporting of risk did not cause direct injury to patients. However, if used by physicians as the sole basis for decision-making, affected patients may have received unnecessary therapy. | Other | Communication with physicians was initiated by the firm's Chief Medical Officer on March 3, 2010. The firm's representatives will meet with treating physicians in person while having the corrected results in hand to inform each in person of the action, to allow a decision of the best course of action. Physicians will be provided with background information on the issue at hand and corrected results.As follow up, the physician would receive a corrected report per patient by FedEx in order to register the reception of corrected information. Because each user received a personal visit, no other contact information was planned or provided by Agendia. |  | 107 | 107 | Nationwide. | Terminated | June 18, 2010 | 107 | 107
Z-1322-2010 | 54373 | Yes | FKP | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter 1550 Single Patient System Hemodialysis Machines; Baxter Healthcare Corporation, Deerfield, IL 60015; product codes 5M5538, 5M5538R, 5M5575, 5M5551 and 5M5551R. |  |  | January 29, 2010 | 2010 | Baxter Healthcare Renal Div | Some of the 1550 Hemodialysis Instruments may have replacement circuit boards with incorrect versions of software. | Software change control | An "Urgent Device Correction" letter dated January 29, 2010 was sent to the direct accounts via first class mail on the same date, to the attention of the Hemodialysis Administrator. The letter informed the accounts that subsequent to a prior device field correction in 2003-2004, some 1550 hemodialysis instruments may have replacement circuit boards with incorrect versions of software. The accounts were instructed to confirm the software version in their systems and replacement boards. Instructions for checking the software version were attached to the letter. A complementary software upgrade will be made available to the account if their 1550 instrument and/or spare circuit board inventory has a software version other than 3.0A or 4.1B. The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter and listing the serial numbers and software versions of those instruments that do nit have software 3.0A or 4.1B.Home patients were not notified directly, the information was sent to their Dialysis Center. The 1550 operator's manual was revised during the 2003 device field correction. The manual correctly warns the operator to discontinue treatment if the 1550 experiences a FL04, FL07, or FLOg fault code.To make arrangements for the software upgrade or questions regarding the communication were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 3.For additional questions, contact the company Monday through Friday from 8:00 AM to 5:00 PM CST at (Toll Free): 800-422-9837. |  | 443 | 443 boards | Worldwide Distribution -- United States - nationwide and Puerto Rico, Argentina, Brazil, Canada, Chile, China, Colombia, Ecuador, Guatemala, Honduras, Mexico, Panama, Philippines, Romania, Turkey and Venezuela. | Terminated | February 03, 2011 | 370 | 443
Z-1323-2010 | 55137 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | The Bigger Better-Bladder, Catalog Number BBB38 (with 3/8 inch ID tubing), Sterile, Single use only --- Manufactured for: Circulatory Technology Inc., 21 Singworth Street, Oyster Bay, 11771 (516) 624-2424 --- 510 (k) 981284. |  |  | June 03, 2009 | 2009 | Circulatory Technology Inc | During an FDA inspection, it was discovered that the firm had conducted a silent recall of its Bigger Better Bladder device after receiving a complaint from a hospital that the device collapsed during a procedure thereby obstructing flow. | Process control | The firm, Circulatory Technology, Inc., sent an email to Bigger Better-Bladder end user accounts on 6/03/09 and again on 6/13/09. The firm followed up with phone calls to the user customers. The emails and calls discussed the product, problem and the actions to be taken by the customers. The customers are ask to examine their stock of Bigger Better-Bladder for the affected units and then visually examine each unit to assure that adhesive has been applied and when the tube is bent away from cap, no separation is visible along the joint between the two according to the picture provided. The customers should return for full credit or a replacement of any product that has a separation of the seal.Please feel free to contact Mr. Yehuda Tamari at 516-624-2424 with any questions. |  | 144 | 144 units | Arizona, Arkansas, California, Florida, Indiana, Iowa, Maine, Michigan, Nebraska, New York, North Carolina, Pennsylvania, Ohio, Texas, Utah, Washington DC, and Germany. | Terminated | May 14, 2010 | 345 | 144
Z-1236-2010 | 55001 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife. |  |  | October 01, 2007 | 2007 | Elekta, Inc. | After updating the LGK actuator in spare part 810361, the old sleigh became obsolete due to causing insufficient locking of the helmet in combination with the new actuator. | Other | Field Change Order 0000203 "Replacement of obsolete actuator sleigh on Helmet Changer", was sent to all customers that were affected. Direct questions to Elekta, Inc. by calling 1-770-670-2548. |  | 5 | 5 units | United States (CA, TN, TX and WA) and Taiwan. | Terminated | December 28, 2010 | 1184 | 5
Z-1391-2010 | 54079 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Lifeshield, latex-free Secondary I.V. Set, convertible pin, 34 inch, piggyback with OPTION-LOK, non-DEHP, 15 drops/mL, a sterile Rx single use device, Hospira Inc, Lake Forest, IL USA. Made in Costa Rica, list 12667-28. |  |  | November 24, 2009 | 2009 | Hospira Inc. | Reports of difficulty in adjusting the flow rate, which could result in flow rate variances. | Component design/selection | Hospira sent the Urgent Device Recall letters dated 11/25/09 to their direct accounts of the affected products and requested these consignees to notify their customers of the recall. The accounts were informed that Hospira ceased manufacturing the I.V. gravity and pump administration sets with the CAIR clamp manufactured by Borla S.P.A. of Italy in March 2008, due to customer reports related to a difference in the zone of control used to adjust the flow compared to the original CAIR clamp design. To avoid the possibility of flow rate variances, Hospira is recalling all remaining sets with the Borla designed CAIR clamp. The accounts were requested to quarantine all inventory of the affected lots, complete the enclosed reply form indicating the amount of product being returned, fax a copy of the completed form to Stericycle at 1-877-377-5114, and include a copy with the returned goods to Stericycle at 2670 Executive Drive, Suite A, Indianapolis, IN 46241. |  | 308064 | 308,064 sets | Worldwide Distribution -- U.S., Puerto Rico, Dominican Republic, Bahamas, Virgin Islands, Argentina, Australia, Bahrain, Belgium, Canada, China, Denmark, Ecuador, EI Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Ireland, Italy, Jordan, Kuwait, Lebanon, Libya, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Peru, Philippines, Poland, Portugal, San Marino, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uruguay, and United Arab Emirates. | Terminated | January 29, 2012 | 796 | 35723417
Z-1399-2010 | 54542 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Quick Connect Adapter Model #CSS-QA-Adapter MaxTorque(R) Maxtorque Quick Connect Lot #s: 1014001 &amp; 1101011, QTY 1 NON STERILE MAT:17-4PH |  |  | October 16, 2009 | 2009 | Orthohelix Surgical Designs Inc | The adapters are able to come apart when a retaining ring in the proximal end of the adapter can dislodge from the retaining ring groove and the two ends of the adapter fall apart and are unable to be reassembled without special tools and knowledge of the instrument.Also the adapters have a wobble between the two ends. This is caused when the distal and proximal ends of the adapter assembly were | Other | Two "IMPORTANT RECALL NOTICE" letters, one dated October 16, 2009, and the second letter dated November 19, 2009, were sent to the customers. The letters described the product, problem and action to be taken by the customers. The customer are requested to discontinue use of the adapters and return them to the firm, Orthohelix Surgical Designs. The firm will replace the adapters as soon as possible. Should you have any questions about the systems please contact Kristin Wolff at 330-247-1445 or Jim Bragg at 330-247-1444. |  | 127 | 127 pieces | Nationwide distribution: CA, CO, FL, GA, IL, KY, MD, MI, MO, MS, NC, NJ, NM, NY, OH, OR, PA, SD, TN, TX, VA, WA , and WI. | Terminated | April 11, 2012 | 908 | 127
Z-1397-2010 | 54175 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | SmarTemp Probe Covers for use with the DPM3 Monitor; Rx only, Mindray DS USA, Inc. Mahway, NJ 07430. Packaged in boxes of 20, Model ST010 also distributed in boxes of 2000 (10 individual boxes of 200); SmarTemp Module Kit (contains one box of 20 probes) and SmarTemp Temperature Probe Kits (contains one box of 20 probes). |  |  | December 31, 2009 | 2009 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | SmarTemp Disposable Probe covers may contain microscopic holes in the molded seam that would allow liquids or body fluids inside the cover and potentially contaminate the temperature probe. | Process control | Mindray DS USA, Inc. issued an "Urgent Product Recall" notification dated December 31, 2009 by certified mail. Consignees were informed of the affected product and requested to discard all remaining stock. Replacement product will be provided by the firm at no cost.For further information, contact Mindray DS USA, Inc. at 1-201-995-8000. |  | 679520 | Total: 33,976 boxes of 20 (boxes of 20 packaged into boxes of 200) in US, 801 boxes OUS | Worldwide Distribution -- United States, Canada and Australia.Nationwide distribution to hospitals, medical facilities and distributors as well as to Canada and Australia. | Terminated | December 09, 2011 | 708 | 679520
Z-1395-2010 | 55040 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity Web Diagnostic 1.0(WebDX) software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | February 18, 2010 | 2010 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity PACS-IW and PACS Web Diagnostic software related to use of built-in reconstruction functions which may result in incorrect anatomic orientation markers. | Software design | An "Urgent Medical Device Correction" letters dated February 18, 2010 to their Centricity PACS-IW software version 3.x customers and Centricity PACS Web Diagnostic software version 3.7.4.x customers, The letters advised the users of the patient safety issues associated with the use of the software when using built-in reconstruction functions, which may result in incorrect anatomic orientation markers. The separate letters provided the users with specific safety instructions for that software version to follow until the software is updated. The customers were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479.Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. |  | 36 | 36 units | Worldwide distribution: USA and Australia, Austria, Belgium, Canada, Finland, France, Germany, Hungary, India, Ireland, Israel, Italy, Japan, Korean, Kuwait, Latvia, Lebanon, Libya, Lithuania, Malaysia, Netherlands, Philippines, Poland, Qatar, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 670
Z-1007-2010 | 51256 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACSCount Controls Kit, An Automated Differential Cell Counter, Model/Catalog Number: P/N: 340166, |  |  | January 29, 2009 | 2009 | BD Biosciences, Systems &amp; Reagents | The product is mislabeled, in that, zero/low controls were mislabeled as medium/high, and the medium/high controls were labeled zero/low. | Package design/selection | On 1/29/09, the firm initiated the recall and its notification was via Important Product Information letters identifying the affected product and explaining the reason for the recall. Customers who experienced error codes with the affected product were to contact their local BD Biosciences representative or distributor to receive replacement BD FACSCount controls. |  | 17 | 17 units | Nationwide Distribution -- Including states of WA, MO, NV, CA, NY, PA, AL, and OR. | Terminated | March 08, 2011 | 768 | 17
Z-1404-2010 | 53031 | No | NVJ | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Renessa RF System, model PR0918, manufactured by Novasys Medical, Newark, CA |  |  | July 27, 2009 | 2009 | Novasys Medical Inc | Instructions for Use (IFU) revision-- Because of complaints received by the firm of worsened stress incontinence symptoms 3-4 weeks after treatment, the instructions for use were revised to emphasize potential side effects. Although Novasys has not, to date, identified a device malfunction associated with the symptoms discussed above, the investigation over time shows that there is a correlation | Labeling design | A Field Action notification letter was sent to all current customers (physicians) along with the revised instructions for use, and a small supply of the revised patient information brochures. The company also updated physician information brochures and patient/physician DVDs to include the information. This was shipped to physicians in August/September of 2009. Additional copies of the Patient Brochures can be ordered by contacting Novasys Customer Service at 866-784-4777. Questions should be addressed to the firm at the same telephone number. |  | 72 | As of 6/12/09 there were 436 Probes at customer accounts. | US nationwide only. (TX, FL, CA, LA, KY, MO, GA,OR WA, HI, AZ, OK AL, TN) . | Terminated | January 08, 2011 | 530 | 72
Z-1400-2010 | 54004 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | V-Probe, Endocare Cryo V Probe model # Cryo 206V. |  |  | December 05, 2008 | 2008 | Endocare, a Wholly Owned Subsidiary of HealthTronics, Inc. | Endocare received a report of a patient death that occurred during a prostate cryoablation procedure. An autopsy attributed the cause of death to an air embolism. The safety notice for the cryoprobes was issued to refocus attention on the Indications and Safety Review included with the Physician Training Materials and to reinforce additional safety information. | Other | The recall was initiated with Endocare sending a notification letter to identified customers on December 5, 2008 and offering an optional replacement unit. The initial correspondence was sent out via an e-mail if the correspondence was returned undeliverable or an e-mail address was not available it was then sent Federal Express to those recipients.The notification letter explained the reason for recall and requested providers refocus their attention on the Indications &amp; Safety Review information contained in the Endocare's Physician Training Materials. Physicians were instructed to contact the firm at 1-949-450-5412 to discuss any questions they may have about the information in the notification letter. |  | 22741 | 22,741 units total for all products | Nationwide. | Terminated | October 21, 2010 | 685 | 22741
Z-1407-2010 | 54783 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Product is a pull cord type monitor, bed patient alarm used to detect patient/resident movement. An adjustable cord is attached to both the alarm and the patient. When the patient moves beyond the length of the cord the pin is pulled and the alarm activated. |  |  | February 12, 2010 | 2010 | AliMed Corporation | Alimed became aware of issue via internal receiving inspection activities on 1/20/09. Firm states that subsequent to this observation, Customer Dissatisfaction inquiry #100192 was generated on 1/22/10. A single unit was received back by AliMed and 0 n/2/1/20 from a rehabilitation center that indicated that it had received a unit and it did not function upon its initial use. Failure of the alarm | Device Design | Letter sent to customers stating that some units may have been subjected to excessive heat during the manufacturing process causing internal distortion to a key alarm component. The letter also request that the Health Care professional review records and notify all users to whom the product may have been issued and if any product has been issued outside of their facility, to forward this correspondence to them. Firm has enclosed a form for return of the product and issuance of a new product. For questions, please contact the firm using the information provided in the letter. |  | 30 | 30 units | All product was distributed domestically with the exception of one unit sold in Canada. AliMed sells and distributes this product to medical health professionals and distributors as well as directly to end users | Terminated | July 25, 2011 | 528 | 30
Z-1403-2010 | 55037 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity Web Diagnostic (WebDX) software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | February 11, 2010 | 2010 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity PACS-IW and PACS Web Diagnostic software when dragging-and-dropping images from the Navigator causing flipping of the images that may impact patient safety. | Software design (manufacturing process) | GE Healthcare issued "Urgent Medical Device Correction" notifications dated February 11, 2010. The initial notification advised users of the patient safety issues associated with the affected product. Separate letters provided the users with specific safety instructions. Accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479.For additional information, contact the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. |  | 13 | 13 units | Worldwide Distribution -- United States (California, Illinois, Indiana, Minnesota, Nevada, New York and Wisconsin), Austria, Canada, Finland, France, Germany, Ireland, Lebanon, Japan, Philippines, Saudi Arabia, Singapore, Switzerland, Turkey, United Arab Emirates and the United Kingdom | Terminated | November 15, 2010 | 277 | 42
Z-1409-2010 | 55190 | No | LGW |  |  | Neurology | PMA | Bionic NAVIGATOR Clinician Programmer Model Number SC-7150-4 |  |  | March 19, 2010 | 2010 | Boston Scientific Neuromodulation Corporation | Toshiba has determined that certain Sony battery packs installed in Toshiba portable computers can overheat, posing fire hazard to customers. Based on Toshiba's investigation with Sony Corporation, they concluded that the battery cells in specific manufacturing lots could be affected. | Other | In a letter dated March 19, 2010 Boston Scientific informed consignees that it determined that certain Sony battery packs installed in Toshiba portable computers can overheat, posing fire hazard to customers, and telling them that they may be in possession of an affected product. They were advised that a Boston Scientific representative will contact them to arrange for an inspection of the Clinician Programmer (SC-7150-4).Until they have determined if the battery pack is affected, consignees were told they can only use their Toshiba portable computer with the original AC adapter power cord supplied with that computer, and were asked to remove the battery pack until its status is determined.Included with the letter was a Reply Verification Tracking Form which consignees were asked to complete the form following the instructions on the form and return to Boston Scientific in the enclosed postage paid return envelope or fax it to (877) 856-3546 by March 31, 2010.If they experience any issues they were told to cease using the Clinician Programmer and contact Boston Scientific's 24-hour Patient Care Hotline at (866) 360-4747.Customers can also contact the Field Action team by phone at (866) 684-7504 or by fax at (877) 856-3546 if they had any questions. |  | 62 | 62 units | Domestically to AR, CA, CO, DC, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NY, OH, OR, PA, PR, SC, TX, VA, WIInternationally: Italy | Terminated | February 29, 2012 | 712 | 62
Z-1408-2010 | 55177 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | The BleaseSirius and BleaseFocus Anaesthesia Systems. |  |  | March 26, 2010 | 2010 | Blease Medical Equipment, Ltd. | Potential for delivery of a lower than set concentration of anaesthetic to a patient due to a sticking valve. The lower shuttle may become jammed, not permitting flow of anaesthesia gas through the vapouriser. The vapouriser back bar mounting valves will require exchange only on the systems identified in this matter, no other systems are affected. | Other | On 2/26/2010 notification letters were sent to all international consignees. The 2 domestic consignees were sent notification letters via FEDEX on 3/26/2010. The domestic consignees were also contacted by phone on 3/26/2010.The notification letter is titled URGENT MEDICAL DEVICE CORRECTION. It advises users of the issue and provides additional information which would permit the continued use of the Blease system until replacement back bar valves can be fitted by the recalling firm's service representatives.For users who choose to continue to use the device, the manufacturer recommends the use of an Agent Monitor and user education; also checks that vapourisers have been mounted correctly and the system is leak free, and testing in advance of a procedure. For additional questions, US Customers are directed to call Global Technical Support - +1 425-657-7200 ext: 5089.Non US Customers, call Global Technical Support - +44 1494 784422. |  | 40 | 40 units | Distributed to 2 U.S. hospitals and to customers in the following countries:Panama, Poland, Australia, Singapore, United Kingdom, Bangladesh, Nicaragua, Vietnam, Argentina, Finland, and Columbia. | Terminated | August 23, 2011 | 515 | 40
Z-1405-2010 | 55101 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI Medical Linear Accelerator. Model Numbers R3.5, R4.0 and R4.2. Elekta, Inc. Norcross, GA 30092. |  |  | December 12, 2008 | 2008 | Elekta, Inc. | It is possible that two (or more) CT reference data sets are necessary for one patient, for example, if you have to scan a patient during breath hold and then again during free breathing. | Software design | Elekta, Inc. issued an "Important Notice" dated December 12, 2008. Customers were informed of the affected device and instructed on proper steps until an update is provided.For further information, contact Elekta, Inc. at 1-770-300-9725. |  | 76 | 76 units | Nationwide Distribution -- AL, AZ, CO, CT, FL, GA, ID, KS, KY, MA, MI, MN, MS, NE, NY, OH, OK, OR, PA, TN, TX, VA, WA and WV. | Terminated | December 28, 2010 | 746 | 76
Z-1406-2010 | 55100 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PrecisePLAN 2.10 Treatment Planning System |  |  | March 15, 2006 | 2006 | Elekta, Inc. | Error in dose rates which were calculated at exactly twice the expected values. | Software design | As described in a prior related submission, Phase I - "Important Notice A285, Scrolling Selected Frame Dose Calculator Error" dated March 15, 2006 was distributed to affected customers to inform them of the potential and what steps should be taken.Phase 2 -- Software release PrecisePLAN 2.11 was released July 11, 2006 to correct this problem. All customers with PrecisePLAN 2.10 will be upgraded to the revised software PrecisePLAN 2.11 through Field Change Order #772 00 540 052 dated June 5, 2006. Phase 3 -- All customers with PrecisePLAN 2.10 will be upgraded to the revised software PrecisePLAN 2.11 through Field Change Order #772 00 540 052 dated June 5, 2006. The company is directing questions to the local Electa representative or to their offices in the U.K. at telephone 44 (0)-1293-654200. |  | 40 | 40 units | Worldwide: IL, NY, PA, TN, Canada, China, Czech Republic, Germany, India, Israel, Italy, Mexico, Poland, Russia, Spain, Switzerland | Terminated | December 30, 2010 | 1751 | 40
Z-1201-2010 | 54478 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Howmedica Restoration ADM Trial Cup Holder; |  |  | January 21, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The Restoration ADM Trial Cup Holder may not have been assembled correctly. | Other | An "URGENT PRODUCTION OPERATION" letter dated January 21, 2010, was sent to customers via FedEx. The letter described the product, problem and actions to be taken by customer. The customers should follow the instructions on the enclosed PRODUCT BULLETIN. The customer should check their internal inventory immediately and quarantine all affected devices, identify if the device is correctly assembled and complete the Product Correction Acknowledgement Form and fax it to (201) 831-6069. If you have any questions, feel free to contact Rich Wolyn, Manager, Divisional Regulatory Reporting at (201) 831-5158. |  | 5 | 5 units | Nationwide distribution: AZ, CA and OR | Terminated | June 20, 2012 | 881 | 5
Z-1421-2010 | 54547 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Cordis, AVANTI+ Introducer, STERILE EO, For one use only. |  |  | January 28, 2010 | 2010 | Cordis Corporation | Sterility compromised: Cordis found a pin hole in one of the packaging tray cavities.On January 28, 2010 Cordis Corporation initiated a recall on their AVANTI + Introducer. During routine inspection, a visual irregularity was detected in packing trays. | Package design/selection | Cordis initiated the recall January 28, 2010. (The packets were sent overnight on January 28, 2010, with first receipt by customers January 29, 2010.) Immediate corrective action was taken through an inspection and removal of all affected products. |  | 1850 | 1850 units | Nationwide. | Terminated | July 27, 2012 | 911 | 1850
Z-1413-2010 | 54640 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | 3M Micropore Single Use Rolls, Surgical Tape, REF 1530S-1, 1 in x 1-1/2 yd. |  |  | January 29, 2010 | 2010 | 3M Company / Medical Division | Reduced adhesion: Micropore Surgical Tape for use as the primary securement device for 1) dialysis needles used to access A-V fistulas or grafts; or 2) pressure dressings used after dialysis treatments. If the tape does not perform, the dialysis needle may loosen and could then dislodge. In the interim, please offer 3M Micropore Surgical Tape, 1 inch x 10 yard rolls (CAT #1530-1) as a substitutio | Nonconforming Material/Component | An Urgent Medical Device Recall letter was sent to 3M customers beginning 01/29/2010. The letter described the issue, asked distributors to check existing inventory, quarantine affected lots. Distributors were asked to notify their customers and end-users (3M has provided example letters and return forms to facilitate). End users can return product either to the distributor or contact 3M directly. A Product Recall Form is asked to be faxed back to 3M.A second 3M "Urgent Medical Device Recall" letter dated February 22, 2010 was sent to customers. This letter informed the cusotmers that 3M was expanding the recall to include bulk-packaged rolls of Micropore Surgical Tape. The letter addressed to "Valued Kit manufacturer" described the problem and the product, provided instrucitions for "Action Requried by the Kit Manufacturer", "Action Requried by the Kit Manufacturer to the End-User", and "Action Required for Future Tape Orders from 3M". A Product Recall Form is aked to be faxed back to 3M. |  | 16582600 | 165,826 boxes, 100 rolls per box (16,582,600 rolls) | Nationwide distribution, and Puerto Rico. | Terminated | February 04, 2012 | 736 | 20445231
Z-1422-2010 | 54769 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra, Licox Ref 1P2.P; |  |  | February 18, 2010 | 2010 | Integra LifeSciences Corp. | One probe that had failed a product release test was inadvertently packaged and released into inventory in Integra's distribution center. | Process control | An Integra NeuroSpecialist was sent to retrieve the probe on February 18, 2010, followed by a formal letter. |  | 1 | 1 | One hospital in Omaha, NE. | Terminated | April 20, 2010 | 61 | 1
Z-1423-2010 | 55108 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS model number 10496279 |  |  | February 16, 2010 | 2010 | Siemens Medical Solutions USA, Inc | images can be overwritten | Software design | A "Customer (Safety) Advisory notice" dated February 16, 2010, was sent all affected customers via Update Instructions IM019/10/S. This notice describes the product, problem and the action to be taken by the customers. The customers should immediately instruct their employees of the advisor notice and maintain awareness. If the customer sold the device/equipment and is no longer in possession of it, they should forward the notice and inform Siemens of the new owner. A software update to correct this issue will be distributed and installed by a Siemens Service Engineer on all affected systems under Update Instructions IM020/10/S.If you have any further questions regarding this issue, do not hesitate to contact your local SIEMENS Uptime Service Center (michael.mib.braun@siemens.com). |  | 6 | 6 units | Nationwide distribution: CA, NC, NY, PA, and TX | Terminated | July 05, 2012 | 870 | 6
Z-1412-2010 | 54648 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Eclipse Treatment Planning System, with v client versions 7.3 (build 7.3.10), 8.0 (builds 7.5.x), 8.1 (builds 8.1.x) and 8.5( builds 8.2.x). For radiation therapy. |  |  | December 09, 2009 | 2009 | Varian Medical Systems Oncology Systems | Potential for misdelivery of treatment-- When preparing a treatment plan involving a split, the device may calculate excessive or insufficient monitor units for the treatment plan. | Software design | Recall initiated on 12/09/2009. Varian issued a notification letter to its consignees, including a description and corrective action step. The firm is in the process of preparing a software fix for the issue. |  | 3177 | 3177 unit | Nationwide distribution and Mexico. | Terminated | February 24, 2012 | 807 | 3177
Z-1426-2010 | 54858 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | REF 822 IMUBIND(R) Plasma PAI-1 ELISA |  |  | March 05, 2010 | 2010 | American Diagnostica, Inc. | Unacceptable microwell to microwell variation in the section of the plate that is typically used for the calibrators | Process control | Three emails "REF 822 lot 102401 needs to be recalled" dated March 1, 2010 was sent to three customers. The emails described the product, problem and action to be taken by the customers. The customers should send the kit to the firm (SEKISUI american diagnostica inc.) attention of Leigh Ayres, Director, Regulatoy Affairs and Quality Assurance.If you have any question, please call 1-203-602-7777. |  | 25 | 25 kits | Worldwide distribution: USA and UK, Taiwan and Korea. | Open, Classified |  |  | 25
Z-1410-2010 | 54894 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Biphasic LIFEPAK 15 Monitor/Defibrillator. Manufactured by Physio-Control Inc., a division of Medtronic, Inc. Redmond, Washington USA. |  |  | March 04, 2010 | 2010 | Physio Control, Inc. | Potential for the device to power off then on by itself, or to power off by itself and requiring the operator to turn it back on, or the device doesn't turn off. | Device Design | Physio-Control Inc. issued an "Urgent Medical Device Correction" notification dated March 2010 to domestic consignees. Foreign consignees were notified by overseas representatives. Consignees were advised to keep the affected device in service and to test the units in accordance with operating instructions. They were further advised that a local service representative will call and schedule a service visit within 60 days. Consignees were asked to forward the notification to all of their sites and to notify the recalling firm if they no longer own the device.For further information, contact Physio-Control Inc. Technical Support at 1-800-442-1142, Option 5, 6:00 am to 4:00 pm (Pacific), Monday-Friday. |  | 3609 | 3,609 | Worldwide Distribution -- United States, Netherlands, Germany, Canada, Australia, and Hong Kong. | Terminated | January 03, 2012 | 670 | 3609
Z-1425-2010 | 55142 | No | GZB | Implanted spinal cord stimulator for pain relief. | Neurology | Neurology | 510(k) | Swift-Lock (TM) Anchor Accessory Kit, Directions For Use Model 1192, ANS, A St. Jude Medical Company. |  |  | February 22, 2010 | 2010 | Advanced Neuromodulation Systems, Inc | Directions For Use of product incorrectly stated lead compatibility information. | Other | An "Urgent Medical Device Correction" letter dated March 8, 2010, was sent to the customers (note: Firm verbally informed consignees on February 22, 2010.). The letter describes the product, problem and action to be taken by the firm -St. Jude Medical. The firm's Sales Representatives will contact the customers to update any DFUs in their possession.If you need additional copies of the insert or more information and if you have any questions or concerns, please contact your St. Jude Medical Sales Representative at 1-651-490-4470. |  | 1120 | 1,120 units (218 went to doctors, remaining 902 were distributed to sales representatives) | Nationwide distribution: AZ, CA, CO, DC, FL, GA, IL, IN, MI, MO, MS, NC, NM, NJ, NV, NY, OH, OK, OR, PA, SC, TN, WA AND WV. | Terminated | July 22, 2010 | 150 | 1120
Z-1428-2010 | 55144 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Leksell SurgiPlan |  |  | May 23, 2005 | 2005 | Elekta, Inc. | Images from a different patient can be imported without a warning for the user. | Other | Phase I - A technical note for LGP and SPS describing the risk for importing wrong image files and including recommendations on how to reduce this risk has been written and distributed to users of both Leksell GammaPlan and Leksell SurgiPlan on May 23, 2005 via FCO DOC 05-0001. Phase 2 - A maintenance release was created of the Scanner Interface (SIF) to correct the problem. This was distributed to Leksell SurgiPlan users on November 30, 2005 via FCO0000153-1. Additional questions are directed to the firm to Product Manager lrobert.aslund AT elekta.com, or After Sales gustaf.piehl AT elekta.com.. |  | 57 | 57 units | Nationwide, Canada and Mexico. | Terminated | December 14, 2010 | 2031 | 57
Z-1430-2010 | 54541 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Mini Max Torque 2.6 mm diameter Drill Bit, Model #HCS-072-26 &amp; 3.8 mm Countersink, Model # HCS-070-38. |  |  | September 03, 2009 | 2009 | Orthohelix Surgical Designs Inc | The epoxy is flaking off the drill bits and countersinks. | Nonconforming Material/Component | On 9/3/2009 the firm sent letters to their customers requesting return of all recalled product in stock and for those products to be replaced with the ones listed in the letter. Questions should be directed to 330-869-0625. |  | 74 | 74 pieces | Nationwide Distribution -- CO, OH, PA, SD, TX &amp; VA. | Terminated | February 28, 2012 | 908 | 74
Z-1429-2010 | 55148 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife C |  |  | August 01, 2006 | 2006 | Elekta, Inc. | Due to a material change in a critical component of the linear actuators the quality of said part was found to be inadequate. | Other | Phase 1 - Technical Notes Leksell Gamma Knife C version 1.2 (dated August 2006) "Potential Problem with Error Handling" was sent to all users. Phase 2 - Mandatory Field Action 0000748, "Helmet Changer Actuator, MTB and MCU SW Upgrade Kit was released December 3, 2007. This Field Change Order includes two mod kits and a software upgrade to provide a permanent solution to the issues. Phase 3 - Technical Note 200 068, "Safety issue with helmet host actuator", released on May 15, 2008. This Technical Note includes a warning to customers of a potential quality problem with the helmet hoist actuator that is installed in the Leksell Gamma Knife. Phase 4 - Mandatory Field Action 200 067, "Investigation and Correction of Helmet Changer Actuators", released on May 15, 2008. This FCO includes a list of user machines which must be monitored for quality issues and provides a solution for replacement of parts should a deficiency be detected. Phase 5 - Mandatory Field Action 200 075, "Helmet Changer Actuator, MTB and MCU SW upgrade, release two", released on September 16, 2008. This FCO applies to machines within specific serial number ranges. Due to a material change in FCO-0000748 and baselines (conducted without notice by the manufacturer) the quality of said part was found to be inadequate. After reversing the material change, FCO 200 075 is released to update the installed base with the new type of actuator and to finish the remaining sites never or partially updated by the now superseded FCO-0000748.For further information, contact Elekta R&D; Jonas Hallstensson onphone no. +46 709 843 553 or via e-mail jonas.hallstensson AT elekta.com. |  | 51 | 51 units | AL, CA, CO, CT, FL, GA, IL, MD, MA, MI, MN, MO, NE, NY, NC, OH, OK, OR, PA, TN, TX, VA, WA | Terminated | December 14, 2010 | 1596 | 51
Z-1431-2010 | 54975 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Unilect ECG Electrodes |  |  | March 11, 2010 | 2010 | ConvaTec | In an emergency situation where defibrillation is necessary and electrodes are already in use, the electrodes may malfunction and incorrectly report lack of cardiac activity. | Process design | The one distributor was notified personally on March 9, 2010. All remaining stock was put on hold, and a "URGENT : MEDICAL DEVICE RECALL" letter dated March 11, 2010, was sent to customers. The letter describes the product, problem and action to be taken by customers. All customers were asked to immediately stop distributing and quarantine all recalled lots. The customers were ask to perform a count of recalled product currently in inventory, complete and return the Recall Response Form even if no recalled product is in inventory, send a copy of response to distributor in order to receive reimbursement, and to call 1-800-582-6514 for instructions on how to return the recalled product.If you have any questions, please contact Convatec at 908-904-2425. |  | 567000 | 567,000 units | Nationwide distribution: CA | Terminated | June 21, 2011 | 467 | 567000
Z-1411-2010 | 55099 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI |  |  | September 19, 2007 | 2007 | Elekta, Inc. | Firm became aware of a failure in the position detection system which could result in incorrect positioning of the detector in the longitudinal or lateral direction. | Device Design | The following Important Notices were sent to customers: 1) Important Notice A305-Possible Synergy XVI panel position error (dated September 19, 2007) was sent to customers with versions R3.5, R4.0 and R4.1., 2) Important Notice A306-Possible Synergy XVI panel positioning error (dated November 8, 2007) was released to encompass customers with version R4.2., and 3) Important Notice A328 dated March 10, 2009, and sent again on April 29, 2009, included a modification kit to address the original issue and an expanded range of serial numbers affected. This Important notice A328 entitled, Elekta Synergy kV Detector Position Readback Modification Kit supersedes A305 and A306 for customers who have serial numbers within the range specified. The letter describes the product and the problem and advises the customer to follow any instructions or recommendations covered in this Notice.If you have any questions, please contact Elekta at 770-670-2548. |  | 166 | 166 units | Worldwide distribution: USA and Puerto Rico, Canada and Mexico. | Terminated | August 09, 2012 | 1786 | 166
Z-1432-2010 | 55103 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Precise Treatment Table |  |  | January 11, 2007 | 2007 | Elekta, Inc. | In some instances, the bolts which are supplied with the table have sheared off, resulting in inability to operate the powered table. | Other | Mandatory Field Change Order (FCO) 528 053 "Check for zinc plated Precise Table 'Z' mechanism fixation bolts", dated 1/11/2007 was distributed to all affected customers. The Change Order identified the product, stated the reason for the FCO, and the procedure to check and resolve the problem identified. |  | 185 | 185 units | Worldwide Distribution -- USA, Puerto Rico, Argentina, Australia, Austria, Bangladesh, Belgium, Canada, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, India, Indonesia, Iraq, Israel, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Pakistan, Peru, Poland, Russia, Serbia, Spain, Sudan, Sweden, Taiwan, Thailand, Turkey &amp; United Kingdom. | Terminated | December 28, 2010 | 1447 | 185
Z-1427-2010 | 54562 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | InterGard Knitted Collagen Coated Vascular prostheses, Catalog number: IGK1206. InterVascular SAS, France. |  |  | December 09, 2009 | 2009 | Intervascular S.A.S. | Packaging was mislabeled. | Labeling mix-ups | InterVascular SAS issued a "Recall Notification - Product Recall RC 020" notification dated January 29, 2009 to the sole consignee. The consignee had previously been notified by phone and all affected product had been segregated.For further information, contact InterVascular SAS at +33 442 08 77 84. |  | 9 | 9 | Nationwide Distribution - Arizona. | Terminated | April 27, 2010 | 139 | 9
Z-1437-2010 | 55252 | No | GWE | Evoked response photic stimulator. | Neurology | Neurology | 510(k) | Diopsys Inc., Enfant Pediatric VEP Vision Testing System, |  |  | April 02, 2007 | 2007 | Diopsys, Inc. | The computer based system may lose test counts when the time in the device was changed during a changeover from standard time to daylight savings time in March 2007. | Software design | Field Service personnel were sent an e-mail notification on 4/2/2007 which included a TZedit file to be downloaded. E-mails instructed reps to install the patch on their customers devices immediately. Field correction was completed 7/31/2007. |  | 175 | 175 units | Nationwide distribution. | Terminated | April 27, 2010 | 1121 | 175
Z-1436-2010 | 55064 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iView GT Electronic Portal Imaging Device, used with radiation therapy treatments. |  |  | July 31, 2009 | 2009 | Elekta, Inc. | Image problems that have an effect on patient position corrections and on the accuracy of the radiation treatment. | Software design | Elekta sent IMPORTANT NOTICE number A329, dated 31 July, 2009 to all customers. This corresponds with Field Change Order - Documentation 764 00 551 060 "An Important Notice (A329) has been issued for replacement of iViewGT R3.4 Service Pack 2 (SP2) installation CD". Mandatory Field Change Order 764 00 551 059 "Replacement of iViewGT R3.4 installer disk"supersedes FCO 764 00 551 053. |  | 121 | 121 lots | Nationwide &amp; Puerto Rico | Terminated | December 14, 2010 | 501 | 121
Z-1433-2010 | 55171 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKNEE SYSTEM, 3D Knee Tibial Insert-DCM, Sterile, Right, size 4 x 11mm. Encore Medical, LP. Austin, TX 78758-5445. |  |  | March 19, 2010 | 2010 | Encore Medical, Lp | Affected product had incorrect color label which is a secondary visual indicator for sizing. | Packaging process control | Encore Medical, LP notified consignees by e-mail dated March 19, 2010. Consignees were requested to return the affected product to the firm.For further information, contact Encore Medical, LP at 1-888-838-3668. |  | 7 | 7 units | Worldwide Distribution -- United States (UT, IA) and Germany. | Terminated | July 22, 2010 | 125 | 7
Z-1438-2010 | 55180 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 2.0/2.4 Cannulated Screw Driver Bit, part #210-24-003 |  |  | April 14, 2009 | 2009 | Trilliant Surgical LTD | Firm received several complaints of field failures, e.g., "broke during surgery" and the firm initiated a design change. A communication to sales staff 04/14/2009 instructed them to recover and reconcile these products from user physicians. | Other | Firm notified sales force by e-mail on 4/14/2009 of the need to recover and reconcile this product . Additional documentation is not available from the firm. Questions may be directed directly to the firm at 800-495-2919. |  | 32 | 32 units | WI, TX, MO, NM, MN, OK, MA, CA, GA, ND, NE, IA, NY, LA, and MD | Terminated | May 03, 2010 | 384 | 32
Z-1440-2010 | 54686 | No | DXY | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular | PMA | Medtronic EnRhythm, P1501DR, Dual Chamber Rate Responsive Pacemaker with RapidRead Telemetry. The EnRhythm pacing system is an implantable medical device that monitors, detects, and treats atrial tachyarrhythmia episodes. It also provides bradycardia pacing and monitoring of ventricular tachycardia (VT) episodes. Manufactured at: Juncos, Puerto Rico, USA. Medtronic, Inc., Minneapolis, MN 55432 USA. |  |  | February 11, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic is informing doctors of two specific battery issues with EnRhythm pacemakers that will be addressed by a Medtronic software update available mid-2010. If the software update referenced above is not implemented, there will be a potential risk of loss of device functionality or risk for loss of therapy at or near ERI in a small number of devices. The software will eliminate this issue, fo | Nonconforming Material/Component | Consignees were given a Medtronic "Important : Medical Device Correction" letter addressed to "Dear Doctor" and dated February 2010. In the letter, Medtronic is informing doctors of two specific battery issues with EnRhythm pacemakers that will be addressed by a Medtronic software update available mid-2010. The letter included recommendations including that Medtronic recommends physicians continue to use the ERI notification to determine time for device replacement. At this time, no other action, reprogramming or change in the frequency of follow-up is recommended. For additional information or assistance, please contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 120061 | approximately 120,061 worldwide | All States in the USA including Guam, Puerto Rico and DC.OUS: Customers in the following countries: Andorra, Argentina, Aruba, Australia, Austria, Bangladesh, Belarus, Belgium, Brazil, Brunei, Darussalam, Canada, Cayman Islands, Chile, China, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guam, Hong Kong, Hungary, India, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kenya, Kuwait, Latvia, Lebanon, Luxembourg, Malaysia, Malta, Mexico, Morocco, Netherlands, New Caledonia, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Vatican City, Venezuela, Viet Nam, Virgin Islands | Terminated | January 08, 2013 | 1062 | 120061
Z-1424-2010 | 53608 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Cardiac Science, Automated External Defibrillator (AED) under the following brand names (with respective model numbers): Powerheart AED G3: 9300A, 9300C, 9300D, 9300E, 9300P, 9390A/E; CardioVive: 92532 and 92533; Nihon Kohden: 9200G, and 9231; and Responder: 2019198 and 2023440. Cardiac Science, Bothell, WA 98021. |  |  | October 12, 2009 | 2009 | Cardiac Science Corporation | Certain models of Powerheart G3 AEDs may not have received adequate electrical safety testing prior to being shipped from the factory. Affected models are: Powerheart 9300A, 9300C, 9300D, 9300E, 9300P, 9390A, 9390E, CardioVive 92532, 92533, NK 9200G, 9231, and Responder 2019198 and 2023440. The affected devices were manufactured between October 29, 2003 and February 10, 2009. | Process design | Cardiac Science issued "Urgent - Medical Device Corrective Action" notices dated October 2009. Customers were instructed to identify the affected product and arrange for product replacement with the firm. Where possible, end user information was requested for further follow-up by the firm.For further information, contact Cardiac Science at 1-877-402-4599 or by email at AED155@cardiacscience. com. |  | 1171 | 1171 (US: 569, OUS: 602) | Worldwide Distribution -- Distributed to the United States, Australia, Belgium, Canada, Chile, China, Denmark, Egypt, France, Germany, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Korea, Malaysia, Marshall Islands, Nigeria, Norway, Philippines, Poland, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Taiwan, and the United Kingdom. | Terminated | February 13, 2012 | 854 | 1171
Z-1441-2010 | 55098 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy |  |  | March 18, 2008 | 2008 | Elekta, Inc. | Installation of a touchguard to prevent injury to patients or staff by stopping machine movements in the case of accidental collision. | Finished device change control | An Important Notice A319 "Instructions for Use for Elekta Synergy kV Source Touchguard" dated 3/18/08 is in distribution to all affected customers. This letter identified the product and asked customers to file the Notice in the User Notice section of the appropriate User Manual. It also asked them to follow the instructions and the reason for the Notice. Questions should be directed to a local Elekta representative. Field Change Order (FCO) 514 259 "kV Source touchguard for systems with touchguard ready source covers" dated 10/17/08 was sent to all affected customers. FCO 514 260 "kV Source touchguard for systems with the original source covers" dated 10/17/08 was sent to all affected customers. |  | 184 | 184 units | Nationwide Distribution -- Including Puerto Rico. | Terminated | December 14, 2010 | 1001 | 184
Z-1443-2010 | 55049 | No | HWB | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Literature for Implant Extraction Set - Implant Extraction Guide Module One &amp; Two. |  |  | March 02, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Literature did not include the design change to a smaller cannulation diameter. | Labeling Change Control | An Urgent Correction - Operative Technique letter was sent to direct accounts, Hospital Risk Management, Chief of Orthopaedics and Surgeons that use the Gamma3 system. Notification letters were sent via Federal Express on March 2, 2010. The letter identified the product and described the issue. It also discussed the potential hazards and risk mitigation. Customers are to contact Customer Service at 1-866-OR-ASSIST or any of the other firm representatives listed in the letter. |  | 790 | 790 unit | Nationwide Distribution | Terminated | August 06, 2012 | 888 | 790
Z-1134-2010 | 54631 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell GammaPlan. |  |  | October 23, 2008 | 2008 | Elekta, Inc. | Although the co-registration looks good during the verification step in the co-registration dialog, the obtain transformation may include an error that depends on the voxel sizes and acquisition parameters of the co-registered image studies. | Other | Important Notice A325 "Error in ImageMerge" dated November 21, 2008 was sent to affected customers and field service engineer as notification of the problem. The notice advises users not to use ImageMerge except if it is judged that the transformation error has no clinical significance. Information is provided in the release to help the user to make an informed decision on this issue. Field Change Order 200 084, "NTPS 8.3 SP1 for NTPS 8.0, 8.2 and 8.3" released January 28, 2009 was sent to affected customers as notification of the fix to this issue. This notice includes instructions to assist Elekta field service personnel in successfully upgrading the affected systems. Once the upgrade is in place, the issue will be corrected. |  | 72 | 72 units | Nationwide Distribution -- (AL, AZ, AR, CA, CO, FL, GA, HI, IL,KY, MA, MI, MS, NE, NV, NV, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN TX and UT). | Terminated | December 14, 2010 | 782 | 72
Z-1442-2010 | 54995 | No | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Instruction for Use for Symphony Double Pumping Kit; Medela, Inc., P.O. Box 660, 1101 Corporate Drive, McHenry, IL 60051-0660; the instructions were packaged with the following non-sterile single user accessory kits for use with Symphony breast pumps: |  |  | March 15, 2010 | 2010 | Medela Inc | The non-sterile devices were erroneously packaged with instructions stating the devices are sterile and do not require cleaning before first use. | Labeling Change Control | An "IMPORTANT:DEVICE RECALL" letter dated March 12, 2010 to the the customers on March 15, 2010, via certified mail, return receipt requested. The letters were tailored to each customer listing the affected shipments sent to them, and informing them of the labeling insert problem. The customers were requested to cease distribution of the affected kits and return them to the firm (Medela) for replacement, and if applicable, to notify their retail customers. The customers were instructed to call Medela Customer Service at 877-320-2301 for assistance in processing returns or for further clarification on the recall. |  | 13834 | 13,834 kits | Nationwide distribution | Terminated | December 10, 2010 | 270 | 13834
Z-1434-2010 | 55182 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Orthofix US XCaliber Meta-Diaphyseal Complete Kit, Sterile. 99-91600, manufactured by Orthofix SRL, Via delle Nazioni, 9, 27012 Bussolengo (VR) Italy. |  |  | February 22, 2010 | 2010 | Orthofix Srl | During standard quality controls, firm identified a potential for breakage of the XCaliber clamp when locking the cam. | Nonconforming Material/Component | Orthofix issued an "Urgent: Field Safety Notice" dated February 22, 2010. Consignees were informed of the affected product and instructed to cease use and to return the product to the firm. |  | 9 | 9 | Nationwide Distribution -- TX, ND, NY, FL and CA. | Terminated | May 05, 2010 | 72 | 34
Z-1439-2010 | 54889 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | InterGard Knitted Collagen Coated Vascular prostheses marketed under 510(k) # K964625; |  |  | December 09, 2009 | 2009 | Intervascular S.A.S. | The external lid of Intervascular's prosthesis has a cut in it which may compromise package integrity and therefore sterility. | Other | The firm sent recall letters to all customers on December 9, 12009. Calls had been made prior to the letter being sent. Because all product was recovered, no additional questions are anticipated. |  | 5 | 5 | US Distribution went to one distributor in Arizona. Also Chile, France, Germany, Hong Kong, Poland and Turkey. | Terminated | April 29, 2010 | 141 | 5
Z-1531-2010 | 52777 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Violet Ink sterilization indicators Violet Ink (Class 6) and Violet Ink (Bowie Dick). Product numbers: CI 144, BD 115, PL 315, PL319, PL232, PL236, PL237, PL 238. |  |  | July 15, 2009 | 2009 | Steritec Products Mfc. Co., Inc. | This action notifies affected SteriTec customers as well as private label distributors of the necessity to modify the shelf life of their sterilization indicator products. There is insufficient data on file to support the shelf life dating on the product labeling. | Other | Notify accounts by e-mail and letter, offering to exchange product for product with revised expiration date. Products exhausted by proper expiration date need not be returned or relabeled.Questions are directed to the firm at (303) 660-4201. |  | 2656 | 2656 cases | Nationwide distribution: IL, MI, NY, OH, VA | Terminated | June 28, 2010 | 348 | 2719
Z-1528-2010 | 55016 | No | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Baxter RenAmin (Amino Acid) Injection; Pharmacy Bulk Package Not for Direct Infusion, Rx Only, 250 mL type II glass bottles; Baxter Healthcare Corporation, Clintec Nutrition Division, Deerfield, IL 60015 USA, Made in USA; product code 2A6222. |  |  | February 15, 2010 | 2010 | Baxter Healthcare Corp. | Routine stability testing of the evacuated container product, along with trending of results, indicates that the pH will exceed the upper specification limit prior to the expiration date of this product. Baxter decided to remove all potentially affected products even though there have been no out-of-specification pH results at the 0, 6 and 12 month intervals. The recall was initiated on 2/15/10. | Nonconforming Material/Component | Urgent Product Recall letters dated 2/15/10 were mailed on the same date via first class mail to the affected direct accounts to the attention of the Director of Pharmacy, the Director of Materials Management, and the Director of Risk Management. The accounts were informed of the potential for the pH of the products to exceed the upper specification limits for pH prior to the product expiration date. The accounts were requested discontinue use and segregate the affected lots of products from inventory and call Baxter Healthcare Center for Service at 1-888-229-0001, Monday through Friday, 7:00 AM to 6:00 PM CT, to arrange for the return of the affected product for credit.The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter and the amount of each product being returned. If the accounts further distributed the products, they were requested to forward the communication to those accounts. Questions regarding the communication are directed to the Medical Information Services at Baxter at 1-800-933-0303. |  | 4164 | 4,164 units | United States -- Nationwide and Puerto Rico, Canada and Saudi Arabia | Terminated | January 29, 2012 | 713 | 148032
Z-1525-2010 | 55175 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | PowerPort isp M.R.I Implanted Port without Suture Plugs with attachable 6F Chronoflex Polyurethane Open-ended Single -Lumen Venous Catheter, BARD Access Systems Assembled in Mexico |  |  | March 18, 2010 | 2010 | Bard Access Systems | 8F PowerPort stems were mixed with 6F PowerPorts and packaged. | Mixed-up of materials/components | The firm, Bard Access Systems, Inc. (BAS) sent a "URGENT: Powerport isp M.R.I with 6F Chronoflex Catheter RECALL" letter dated March 18, 2010, to the customers. The letter describes the product, problem and actions to be taken by customers. The customers were instructed to immediately examine their inventory and quarantine all products subject to this recall and notify any customers they may have distributed the product to; DO NOT USE THIS PRODUCT OR FURTHER DISTRIBUTE ANY OF THE IMPLICATED LOTS; remove and return the affected lots to BAS, and complete and Fax the Reply Form and Inventory Reconciliation Form, if unable to Fax call and report information. Only unused products, in its original packaging, should be returned to BAS.Should you have any questions or require assistance in this matter, please contact BAS Customer Service by calling our toll-free number 1-800-290-1689. |  | 200 | 200 units total (100 units of each lot, see code information) | Nationwide distribution: AZ, CA, DC, FL, MA, MD, NC, NE, NV, NY, PA and SC. | Terminated | October 01, 2010 | 197 | 200
Z-1533-2010 | 54850 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife C |  |  | September 01, 2006 | 2006 | Elekta, Inc. | Leksell Gamma Knife C 1.2 with MCU version 4.0.0.6 improperly adjusted trolley for 4mm helmet caused jamming when helmet changer attempted to lock helmet. Lock did not fully engage and could still be raised. When the lock does not engage, then when the helmet changer rises, the helmet could disengage from the helmet changer and fall back onto the helmet trolley. | Other | Phase 1 - Field Change Order FCO-DOC-06-0003 "Technical Note - Potential risk of dropping the collimator helmet" (dated Aug. 28, 2006) was sent to affected customers as notification of the problem. Phase 2 - Technical Notes Leksell Gamma Knife C version 1.2 (dated July 2006) "Important Notice - Potential risk of dropping the collimator helmet" was sent to all Elekta Service Engineers as notification of the problem.The firm planned to create an upgrade kit for the effected Leksell Gamma Knife C1.2 systems. A service representative from Elekta will contact each site to decide on a suitable implementation schedule. |  | 88 | 88 units | Nationwide, Austria, Belgium, Canada, China, Croatia, Italy, Japan, Korea, Norway, Spain, Sweden, Taiwan and Turkey. | Terminated | December 14, 2010 | 1565 | 88
Z-0820-2010 | 52635 | No | MDS |  |  | Clinical Chemistry | PMA | Freestyle Navigator Continuous Glucose Monitor System, ART 16467 Rev A, ABB SYSTEM KIT(ORANGE)PACK NAV, Abbott Item Number 70684-01, distributed by Abbott Diabetes Care Inc, Alameda, CA. |  |  | April 14, 2009 | 2009 | Abbott Diabetes Care, Inc. | Potential failure or inaccurate readings--the device's plastic housing near the battery compartment may crack, and allow moisture to enter. | Process control | Notification letters were issued to FreeStyle Navigator customers via certified mail and notification was posted on the Abbott Diabetes Care's Website. Customers were notified of the potential issue and actions to take: 1) if moisture is observed in the FreeStyle Navigator transmitter now or at any time, discontinue use of the transmitter and call Abbott Diabetes Care Customer Care at 1-800-418-6427 and 2) if a loss of connection is encountered or if your results are not reliable or are inconsistent with how you feel, consult your User's Guide for troubleshooting guidelines. If the issue is not resolved, call Abbott Diabetes Care Customer Care at 1-800-418-6421. For outside-of-US accounts, the firm's regional offices will distribute notices to identified customers in the appropriate language. |  | 5449 | 5,449 of total units | Worldwide Distribution -- United States (OH, FL, NJ, CA, MS, MN, MA, and NY), France, Australia, the Netherlands, Sweden, the UK, Israel, Ireland and Germany. | Terminated | March 31, 2011 | 716 | 5449
Z-1535-2010 | 53836 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista CHEM 3 calibrator |  |  | October 23, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | patient results are depressed | Other | An "Urgent Field Safety Notice" dated October 2009 was issued by Siemens to all customers. The notice described the product, problem and actions to be taken by the customers. The customers were instructed to recalibrate the AMON method (AMON Flex reagent cartridge, K3019) with Vista Chem 3 Claibrator lot 9ED058 using the revised bottle values provided in the notice and to please consult with their laboratory director regarding re-evaluation of previous results. The customers should forward this notification to anyone to whom they may have distributed this product.If you have any technical questions regarding this information, please contact Siemens Technical Solutions Center at 1-800-441-9250. |  | 1753 | 1753 cartons | Worldwide distribution: USA and Australia, Canada, Japan, New Zealand, South Korea, and Switzerland. | Terminated | June 22, 2010 | 242 | 1753
Z-1534-2010 | 54264 | No | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | Hitachi Oasis Magnetic Resonance Scanners Software V1.0C |  |  | September 01, 2009 | 2009 | Hitachi Medical Systems America Inc | The device has a software error that could potentially lead to a misdiagnosis. | Software design (manufacturing process) | Sites have already been corrected. Software updated from V1.0C to V2.0B |  | 43 | 43 | Nationwide distribution: AZ, CA, CT, DE, FL, GA, IA, IL, IN, LA, MD, MI, MO, NC, NY, OH, OK, OR, PA, RI, SC, TX, VA, WA, WI, and WY. | Terminated | February 24, 2012 | 906 | 43
Z-1536-2010 | 55251 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bard Ventrio Small Oval Hernia Patch 8cm X 12cm (3.1" x 4.7"), Warwick, RI |  |  | April 08, 2010 | 2010 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: Product labeled as Small Oval (3.1" x 4.7") may contain Small Circle (3.0") | Labeling mix-ups | An "URGENT RECALL: BARD VENTRIO SMALL OVAL HERNIA PATCH (LOT DATB0015" letter dated April 8, 2010, was issued by the firm (Davol) to all customers via FedEx overnight with signature required. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to: please examine their inventory and immediately discontinue use of this lot; contact the Davol Customer Service Department at 1-800-556-6275 for instructions on how to return product to firm; complete and fax the EFFECTIVENESS CHECK sheet to 1-401-825-8753, and forward a copy of this letter, if they are a distributor, to their customers who have received the product and lot number. All products returned will be held at the Davol, Warwick location.Please contact your Davol Territory Manager or C.R. Bard's Medical Service Support Department at 1-800-562-0027 if you have any other questions. |  | 198 | 198 units | Nationwide distribution | Terminated | October 27, 2011 | 567 | 198
Z-1524-2010 | 55191 | No | NEG |  |  | Orthopedic | PMA | Ascension PyroCarbon PIP Total Joint - Trials; Ascension PIP, sizes 10P, 20P, 30P and 40P. Ascension Orthopedics, 8700 Cameron Road, Suite 100, Austin, Texas 78754; TEL: 512 836-5001. |  |  | June 17, 2009 | 2009 | Ascension Orthopedics, Inc | Trial product was larger than actual implant. Trial product was re-designed to more closely match the implant. | Device Design | Ascension Orthopedics, Inc. issued e-mail notification to distributors and sales representatives beginning June 2009. Consignees were instructed return affected product to the firm for replacement.For further information, contact Ascension Orthopedics, Inc. at 1-512-836-5001 extension 1549. |  | 129 | 129 units (64 US, 65 OUS) | Worldwide Distribution - Worldwide Distribution - United States, United Kingdom, France, Australia, Belgium, Czech Republic, Canada, Germany, The Netherlands and South Africa. | Terminated | September 27, 2010 | 467 | 129
Z-1537-2010 | 55169 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Medtronic HMS PLUS, Hemostasis Management System. Medtronic, Inc., Cardiovascular Division, Moundsview, MN 55112. |  |  | March 08, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | The heparin assay controls may run longer than normal and in some cases, fail to give an acceptable result. | Process change control | Medtronic issued an "Urgent Medical Device Notification" letter dated March 8, 2010 addressed to "Dear Valued Customer" and was sent to three locations at each customer site: Risk Management, Cardiovascular Surgery and Clinical Laboratory-Point of Care. The letter described the product and the issue, provided instructions for control and patient testing. The customer was requested to return an enclosed certificate as proof of notification.For further information, contact your Medtronic Sales Representative or call 1-763-391-9915. |  | 544 | 508 US, 36 OUS | Worldwide Distribution -- United States, International West (countries included in Western Europe and the Middle East), Australia and Canada. | Terminated | December 31, 2011 | 663 | 544
Z-1538-2010 | 54091 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Percutaneous Sheath Introducer (PSI) Kit with Integral Hemostasis Valve/Side Port for use with 7.5 - 8 Fr. Catheters, Arrow International, Inc., 2400 Bernville Road, Reading, PA 19605 |  |  | December 15, 2009 | 2009 | Arrow International Inc | Some central venous catheter sets had the incorrect lid stock. | Packaging | Consignees were notified by an Urgent Medical Device Recall letter on/about 12/15/2009. The letter identified the affected product and stated the reason for the recall. Customers were instructed to immediately check their stock and cease use, distribution, and quarantine affected product. They are also to contact Arrow's Customer Service Department at 800-523-8446 for a Return Authorization Number. Customers are to return any affected product along with the Advisory Acknowledgement and Stock Status Form. Arrow will either provide replacement with a similar product or credit customers' accounts. Customer Service can be contacted with questions or for clarification. |  | 295 | 295 kits | Nationwide Distribution -- MA. | Terminated | July 29, 2011 | 591 | 295
Z-1221-2010 | 54165 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Shutter HX SP5, a component of the Leica Point Scanning Confocal Microscope Models TCS SP5 and TCS SP5 X. |  |  | November 12, 2009 | 2009 | Leica Microsystems, Inc. | A manufacturing defect was identified that could potentially result in the laser shutter not closing as intended which could result in the shutter not completely closing disrupting the laser beam safely. | Radiation Control for Health and Safety Act | A Service Note was sent to service organizations on November 12, 2009. The communication asked customers to check if they must exchange their shutters and send their planned exchange date by 2009-11-20 to clsm.qc@leica-microsystems.com. Customers were also asked to check the serial numbers of the one-wire chips of the shutters. Then lastly, the were asked to exchange the shutters and confirm the exchange. |  | 23 | 15 microscopes and 8 spare part shutters | Worldwide Distribution -- USA, Sweden, Japan, Hong Kong, Italy, the Netherlands, Singapore, UK, and Spain. | Terminated | May 11, 2010 | 180 | 23
Z-1540-2010 | 55286 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AEM Disposable Electrode, Spatula, 45cm, REF ES0301-45, Qty 10 each Sterile R, Manufactured for: Encision Inc., Boulder, CO 80301, made in Taiwan. |  |  | April 05, 2010 | 2010 | Encision, Inc. | Electrosurgical electrode tip insulation may crack and break off during use. | Nonconforming Material/Component | A "Confirmation of Product Recall" letter dated April 5, 2010, was faxed by the firm (Encision, Inc.) to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to either return or destroy the affected electrodes and to contact any customers or sales representatives to whom they distributed the product and confirm they have been removed from use and/or they will pick them up for return. The customers is to complete form and fax to the attention of RA/QA at 1-303-444-2693 or return with the items. The number to call to arrange return and/or if you have any questions is 1-800-998-0986. |  | 20 | 20 units | Worldwide distribution: USA including states of CA, CT, IA, IN, KS, KY, MI, OK, TX, WV, and countries of New Zealand and The Netherlands. | Terminated | May 21, 2010 | 46 | 1140
Z-1517-2010 | 55354 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Shiley Specialized Tracheostomy Tube, M8FEN, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. |  |  | April 13, 2010 | 2010 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation. | Process design | Covidien issued an "Urgent Medical Device Recall" notification dated April 13, 2010. Consignees were alerted to the affected device. If an affected product is already in use, the firm recommends that the tube be replaced as soon as the patient's physician determines it is clinically appropriate. All affected product is to be returned to the firm. For further Information, contact Covidien at 1-800-635-5267. |  | 20 | 20 units | Worldwide Distribution -- United States (including Guam and Puerto Rico), Australia, Belgium, Brazil, Canada, Chile, Colombia, Honduras, Japan, Mexico, Panama, Singapore and Uruguay. | Terminated | January 18, 2012 | 645 | 559545
Z-1550-2010 | 55179 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Inion 1.5 x 6mm CPS(R) Baby Screw, 1.5mm Inion CPS(R) Baby 5+E, REF SCR 1221, 5/pack, STERILE/R Stryker Craniomaxillofacial, Portage, MI. |  |  | March 05, 2010 | 2010 | Stryker Craniomaxillofacial Division | During a sampling inspection, 2 lots of screws did not meet shear and torsion testing specifications. | Nonconforming Material/Component | The firm, Stryker, issued two "URGENT PRODUCT RECALL NOTIFICATION" letters, dated March 5, 2010 and March 19, 2010, to customers. The letters describe the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory for the catalogue and lot numbers, return the enclosed confirmation indicating the location of the affected product, return any remaining product to Stryker at the noted address, and provide a copy of this notification to any site that they may have distributed the product to.If you have any further questions, contact Rob Yamashita, Sr. Regulatory Affairs Rep at (269) 323-4258 or email: rob.yamashita@stryker.com |  | 184 | 184 packs (5/pack) | Nationwide distribution: CA, CO, CT, FL, GA, IL, LA, MA, MD, ME, MO, MT, NC, NH, NJ, NM, NY, OH, PA, RI, SC, TX and WI. | Terminated | June 06, 2012 | 824 | 184
Z-1547-2010 | 55115 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ZOLL E Series Defibrillator/Pacemaker/Monitors, BLS Model, with software versions 4.XX, 6.XX and 7.0X. Zoll Medical Corporation, Chelmsford, MA. |  |  | March 24, 2010 | 2010 | ZOLL Medical Corporation, World Wide Headquarters | Device issued Shock Advised message but failed to auto-charge the defibrillator. | Software design | Zoll Medical Corporation issued an "Urgent Medical Device Correction" notification dated March 26, 2010 via certified mail to consignees who have received affected product. Customers will be contacted and given the option to have a software upgrade kit to update their devices or return devices to the firm to have the device upgraded.For further information, contact Zoll Medical Corporation 24/7 Technical Support at 1-800-348-9011 or 1-978-421-9460. |  | 1383 | 1383 units | Worldwide Distribution -- United States, Australia, Germany, Libyan Arab Jamahiriya, Malaysia, Singapore, United Arab Emirates and the United Kingdom. | Terminated | January 23, 2014 | 1401 | 1383
Z-1548-2010 | 55147 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 7.5mm Ti Solid Humeral Nail 230 mm Sterile Catalog number 462.723S |  |  | February 25, 2010 | 2010 | Synthes USA (HQ), Inc. | potential compatibility issue | Device Design | The recalling firm, Synthes, issued an "URGENT: MEDICAL DEVICE RECALL" letter dated February 25, 2010, return receipt requested, to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to immediately cease use of the products, examine their inventory for the part/lot combinations and remove the recalled devices from their shelves, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number and return the recalled product with the letter; however, if the customers do not have the identified product , they should complete the attached verification section and return document to Synthes by Fax at 610-719-5120 or scan/email at lewis.lynne@synthes.com. If you should have any questions, please call 800-620-7025 x 5452 or 610-719-5452 or contact your Synthes Trauma Sales Consultant. |  | 1 | 1 unit | Nationwide distribution: LA, MA, ME, NC, PA and VA. | Terminated | October 20, 2014 | 1698 | 6
Z-1551-2010 | 54566 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Peloris Rapid Tissue Processor; Leica Biosystems Melbourne Pty Ltd, Mount Waverly, 3149 VIC, Australia; catalog number 26.0001 and 26.0005 |  |  | February 09, 2010 | 2010 | Leica Microsystems, Inc. | Confusion during manually changing reagents, general users running unvalidated protocols, and wax valve leakage have contributed to tissue damage during processing, leading to rebiopsies. | Other | Leica Microsystems sent Medical Device Correction letters dated 2/9/10 to the affected customers via first class mail on the same date, informing them of the potential for tissue damage during processing and outlining the software upgrade being made to minimize the risk of tissue damage. Customers were informed that their local Leica Microsystems representative will contact them to arrange the software upgrade and to provide training on the use of the new software. The letters also provided the necessary steps required to avoid tissue damage during processing while they are awaiting the software upgrade. The accounts were asked to complete the attached Acknowledgement Form, listing the serial numbers of the units at their facility and fax it back to Leica at 1-847-607-3147.Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. |  | 574 | 574 units | Worldwide distribution -- U.S. (nationwide), Canada, Australia, Europe and Asia. | Terminated | June 30, 2011 | 506 | 574
Z-1560-2010 | 54239 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | One Touch Ping Insulin Pump using ezManager Max diabetes management software |  |  | December 18, 2009 | 2009 | Animas Corporation | Software malfunction which disallows users to download, view and print information from their pump (including blood glucose, insulin delivery, carbohydrate and pump events). Additionally, users are unable to customize the built-in food database and pump settings using the software. | Software design | The recalling firm telephoned the patients/users to inform them of the problem. A letter dated 12/22/09 was also sent to users. Users will have a replacement unit sent to them so that they will be able to fully utilize the pump and the software. As a courtesy a letter dated 12/22/09 was issued to the health care provider/prescribing physician for the affected users. |  | 21 | 21 pumps | The pumps were provided to patients/users in CA, CT, FL, GA, ID, MA, MO, NJ, OH, OR, PA, TX, and WA | Terminated | June 22, 2010 | 186 | 21
Z-1553-2010 | 55216 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator, Models E102 and E110. Sterilized using ethylene oxide. Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 66112-5798, USA. The product is indicated for patients who are at risk for sudden cardiac death due to ventricular arrhythmias and may require pacing support. COGNIS devices also provide bi-ventricular pacing. |  |  | March 15, 2010 | 2010 | Boston Scientific CRM Corp | Boston Scientific Corporation has stopped shipment and is retrieving unimplanted devices from US hospital shelves of all its implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) after determining that some manufacturing process changes were not submitted for approval to the US Food &amp; Drug Administration. | PMA | BSC issues press release dated 15 March 2010. FDA press release posted 18 March 2010. On 15 March 2010, BSC Sales Representatives begin removal of devices from health care facilities stock. Starting 22 March 2010 BSC sent a physician letter dated 22 March 2010 to implanting and following physicians via FedEx. The Urgent Medical Device Recall letter described the issue of why devices were being removed from hospital shelves. The letter did NOT recommend any changes to normal follow-up procedures for patients implanted with an ICD or CRT-D device. The letter included ALL potential US models. To be complete and mitigate the possibility of retrieval errors, the ICD and CRT-D US models that may be on customer's shelves were included in the letter. The devices subject to retrieval, however, were those non-implanted devices with domestic customers.For questions, please contact the local representative (1-800-227-3422) for pick up and return or contact Boston Scientific Technical Services at 1.800.CARDIAC (227.3422). |  | 1330 | 1330 | nationwide within US only (OUS distribution not affected). | Terminated | April 05, 2012 | 752 | 2490
Z-1561-2010 | 55117 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | PM1000 Flow Selector - An accessory to an emergency ventilator (resuscitator). The selector is attached to a flowmeter and is used to direct the outlet flow from any standard flowmeter to any one of the three outlets it provides. |  |  | March 05, 2010 | 2010 | Precision Medical, Inc. | The on/off outlet is mislabeled so that when a user/operator turns selector to "on", it actually may be turned "off" | Labeling mix-ups | The recalling firm issued Customer Alert Product Recall letters on 3/5/10 to inform users of the problem and how to identify affected product. |  | 19518 | 19,518 | The products were shipped to distributor and medical facilities nationwide. The products were also shipped to Taiwan, Germany, Hong Kong, Canada | Terminated | December 13, 2010 | 283 | 19518
Z-1574-2010 | 54083 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics, Scorpio Ceramic 4:1 Guide Size 3; |  |  | December 03, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The ceramic guide rails, within the XCelerate 4:1 Ceramic Cutting Blocks, may fracture and displace from the block. | Device Design | Urgent Product Recall Letters were sent on December 3, 2009 via Federal Express to Stryker Branches/Agencies, Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. Foreign accounts were notitied on December 7, 2009. Questions or comments should be addressed to Colleen O'Meara at 201-972-2100. |  | 9833 | 9833 Scorpio Cutting Blocks, total. | Worldwide Distribution: USA, Sweden, Ireland, Canada, Hong Kong, Germany, Spain, India, Australia, Italy, New Zealand, and United Kingdom. | Terminated | September 14, 2012 | 1016 | 9833
Z-1590-2010 | 54683 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | An Imaging device for medical procedures, to provide surgeons or operating room personnel the ability to capture still images and to record video during the surgical procedures in various file formats. Images can be transferred to external media. |  |  | October 23, 2009 | 2009 | Stryker Endoscopy | Video device is not able to assign the correct time/date stamp. | Device Design | Stryker Endoscopy has notified the 2865 domestic consignees affected by the device correction. The notification was mailed to affected consignees and their corresponding sales representatives. It included two mailings of the Urgent: Device Correction notice, SDC HD 3.0z correction software, and a SDC HD upgrade dongle. 100% of consignees have been contacted and notified of the device correction needed. The effectiveness of the correction has been tracked via acknowledgement cards mailed, faxed, or emailed. The local Stryker Sales Representative will install a software upgrade to fix this issue. Additional questions should be directed to the firm's Customer Service at (800)624-4422 or (408)754-2039. |  | 12896 | 12896 devices | US government facilities and other US domestic consignees(nationwide), and to foreign countries | Terminated | February 17, 2011 | 482 | 12896
Z-1552-2010 | 49383 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Intramedullary Fixation Rod, common name Fibula Rod, labeled IM Rod System Manufactured or Distributed by: ACUMED Hillsboro, OR |  |  | August 21, 2008 | 2008 | Acumed LLC | Products improperly handled after cleaning, prior to sanitation and may be contaminated with inorganic material. | Other | On 8/21/2008 the firm initiated telephone contact with consignees, followed by an e-mail, and a letter. Consignees were instructed to return product to the firm. Consignees were provided with a response form to complete and return to the firm. For additional information, please contact the firm directly at telephone 888-627-9957. |  | 36 | 36 devices | Product was distributed to 6 consignees, 1 consignee located each in MA, CA, NM, TX, and the United Kingdom and South Africa. | Terminated | May 14, 2010 | 631 | 36
Z-1562-2010 | 55189 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ConMed Goldline Hand Controlled With Needle Electrode, Rocker Switch, Safety Holster and 10 FT Cable, REF 130301A, Sterile R, Assembled in Mexico, ConMed Corporation, 525 French Rd., Utica, NY 13502. |  |  | March 29, 2010 | 2010 | ConMed Electrosurgery | Electrosurgical pencil may remain activated or self activate after power switch is released. | Device Design | Customers were notified by letter on March 25, 2010, and were instructed to return all affected devices and to forward the recall information to anyone that may have received the devices to the end user. More information is available by contacting the firm directly at 800-552-0138. |  | 1400 | 1400 units (US) | Nationwide, including 2 VA facilities in GA and 1 military facility in UT. Foreign distribution to Australia, Belgium, Canada, China, Dominican Republic, El Salvador, Japan, Korea, Mexico, Puerto Rico, Saudi Arabia, Spain, Taiwan. | Terminated | September 27, 2010 | 182 | 683440
Z-1591-2010 | 55446 | No | LYJ |  |  | Neurology | PMA | VNS TherapyÂż DemipulseÂż Generator, Model 103 and |  |  | April 06, 2010 | 2010 | Cyberonics, Inc | Battery life projection is inaccurate. | Software design | Safety Alert sent April 6, 2010 to all physicians of record for Model 103 or Model 104 generators and to all hospitals with non-implanted affected product, that explains the event, risks presented and recommended actions to take. The generator can be reprogrammed by the physician the next time the patient is seen in the clinic.The issue can also be corrected during the initial implant procedure with the instructions provided.Additional questions should be directed to the firm; United States customers can contact Clinical Technical Support at 866-882-8804. |  | 18722 | Model103: 8109 units; Model 104: 1722 units; Programmers: 8891 | Worldwide distribution. | Terminated | August 20, 2012 | 867 | 18722
Z-1592-2010 | 54680 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Micro Introducer Kit 5F, Model number: KIT-014-29, &amp; Customer Model number: 7204. Vascular Solutions, 6464 Sycamore Court, Minneapolis, MN 53369. |  |  | February 08, 2010 | 2010 | Galt Medical Corp | Potential problem with pouch seal thus affecting sterile product. | Packaging process control | Firm notified distributor by telephone 2/8/2010. Firm and distributor partnered to notify customers by an Urgent: Medical Device Recall - Lot Specific letter, dated 2/10/2010, beginning 2/16/2010. The letter identified the affected product and the reason for the recall. The letter asked customers to immediately check their inventory remove any affected kits and place them in a secure location. Customers are to complete the Product Recall Inventory Form and fax it to the Customer Service Department. Customers are also to contact Customer Service for a Return Authorization Number at 1-866-240-6001. Affected products will be replaced upon receipt. Questions or concerns should be directed to a local Sales Representative or Jill Eisenzimmer, Product Manager, at 763-656-4300. |  | 400 | 400 units | Nationwide Distribution -- MI, PA, IL, &amp; Washington, D.C. | Terminated | March 21, 2011 | 406 | 400
Z-1527-2010 | 55355 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE, 1009-9002-000 Avance Anesthesia Machine and Monitor |  |  | March 12, 2010 | 2010 | Ge Healthcare, Llc | Unanticipated shut-down: A specific lot of GE Healthcare Aisys and Avance machines control board wiring harness have a defect, which can cause the machine to unexpectedly shut down, terminating ventilation, anesthetic delivery, and potentially patient monitoring. | Process design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated March 12, 2010. The letter was addressed to Chief of Anesthesia, Health Care Administrator/Risk Manager, Director of Biomedical/Clinical Engineering.The letter listed the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. |  | 18 | 18 (4 USA, 14 OUS) | Worldwide distribution: USA (CA, FL, TN), CANADA, DENMARK, FINLAND, FRANCE, ITALY, SAUDI ARABIA, SOUTH AFRICA, SPAIN, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | January 30, 2012 | 689 | 33
Z-0350-2010 | 52852 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Orthopediatrics IM Femoral Nail - 8mm x 34 cm, Rt, non-sterile, Orthopediatrics, Warsaw, IN; Part # 10-1500-036. |  |  | July 21, 2009 | 2009 | Orthopediatrics Corp | The potential exists for the nails to fracture during insertion of the nail. | Device Design | Distributor consignees were notified by email dated July 28, 2009 and requested to return the recalled products and instrument sets. If there are questions regarding this issue, please contact Gary Barnett at 574-268-6379. |  | 23 | 23 of all products | Nationwide. | Terminated | October 25, 2010 | 461 | 23
Z-1595-2010 | 55305 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Basic Core Pack # OR1805, Sterile, Tri State Hospital Supply Corporation Manufacturers of Centurion Healthcare Products. |  |  | April 09, 2010 | 2010 | Centurion Medical Products | The firm is conducting a sub recall of kits containing ConMed Goldline Rocker Switch Pencils. The pencils activation switch may remain in the "on" position after the switch has been depressed with excessive force, thereby continuously activating the pencil. | Nonconforming Material/Component | Centurion sent out an "Urgent Product Recall Notice" to one customer on 4/9/2010. The customer was instructed to discontinue use of the product. The customer was also instructed to return any instock product . |  | 74 | 74 kits | TX | Terminated | October 25, 2010 | 199 | 74
Z-1594-2010 | 55213 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics, Accolade Distal Centralizer 13 mm Medium; Catalog No.: 1059-4513. For use with Accolade hip prosthesis. |  |  | March 31, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | One lot of 13 mm medium distal ring centralizers and 14 mm large distal ring centralizer have been labeled incorrectly. The size 14 Large centralizers have been labeled as size 13 Medium centralizers and size 13 Medium centralizers have been labeled as size 14 Large centraliizers. | Process control | Notification letters and Product Accountability Forms have been sent to Stryker Orthopaedics Branch and Agencies via Fed Ex on March 31, 2010 with return receipt.. Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons were sent notificatio letters and Product Recall Acknowledgment Forms on April 1, 2010. with return reciept. |  | 15 | 15 | Nationwide distribution. | Terminated | July 09, 2010 | 100 | 32
Z-1598-2010 | 55303 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 ProtÂżgÂż EverFlex Self-Expanding Biliary Stent System, Sterile EO, (REF: PRB35-06-150-080, REF PRB35-06-120-120 and PRB35-06-150-120), 4.5 - 5.5mm, Expiration: 2011-10-01, Rx only. ev3 Inc., 4500 Nathan Lane North, Plymouth, MN 55442-2929. |  |  | March 31, 2010 | 2010 | Ev3, Inc. | ev3 is conducting a voluntary recall on three lots of the ProtÂżgÂż EverFlex" Biliary Stent System. They determined that a lot of 120mm ProtÂżgÂż EverFlex" Biliary Stent system may contain a 150mm length Self-Expanding stent and two lots o 150mm ProtÂżgÂż EverFlex" Biliary Stent System may contain a 120 mm length Self-Expanding stent. | Mixed-up of materials/components | ev3 issued "Medical Device Recall" notifications dated March 31, 2010 addressed to "Dear Risk Manager or Cath Lab Manager". Each letter addressed the affected product. Users were instructed to locate and remove from use the affected product and a sales representative will be in contact shortly to arrange for return.For further information, contact ev3 Customer Service at 1-800-716-6700. |  | 19 | 19 | Nationwide Distribution -- CA, FL, IL, LA, MA, MI, NH, PA, TN, TX, UT, VA and WA. | Terminated | March 08, 2012 | 708 | 19
Z-1601-2010 | 55199 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Distal Centralizer, Exeter II Hip System. Catalog number - 0920-2-920. |  |  | March 29, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker became aware that device labels have incorrectly printed manufacture dates and thus incorrect expiration dates on the product label. | Error in labeling | Urgent Product Recall letters and Product Accountability Forms have been sent via Fed Ex on March 29, 2010 with return receipt. Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeon were also sent Notification letters and Product Recall Acknowledgment Form on March 29, 2010 with return receipt. The letters explained the issue, the potential hazards, and risk mitigation. The firm asked customers to examine their inventory and hospital locations to identify the affected products. All affected product should be retrieved and returned to their branch or agency warehouse for reconciliation. The attached Product Accountability form should be used to reconcile all product(s) and it should be faxed back to 201-831-6069. |  | 24 | 24 units | Nationwide Distribution -- FL, MN, &amp; OR. | Terminated | September 10, 2012 | 896 | 24
Z-1600-2010 | 55222 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea 5.1. Vitrea fX 3.0, Vitrea Enterprise Suite 1.2 with the licensing of CT lung or CT Colon. |  |  | February 23, 2010 | 2010 | Vital Images, Inc. | Vital images, Inc is issuing a recall notice on their medical device software in the use of CT Lung and CT Colon in Vitrea 5.1, Vitrea fX 3.0 or Vitre Enterprise Suite 1.2. because incorrect measurements may result. Questions or comments are directed to the firm's Customer Support at support@vitalimages.com or at 1-800-208-3005. | Software design | Consignees were sent a Vital "Urgent Medical Device Software Recall Notice" letter dated 23 February 2010. The letter was addressed to "Dear Customer." The letter described the problem, product affected and provided recommendation to their customers. Questions or comments are directed to the firm's Customer Support at support@vitalimages.com or at 1-800-208-3005. |  | 95 | 95 | AZ, AR, CA, COT, CT, FL, GA, IN, IA, KY, LA, MD, MA, MI, MN, NV, NJ, NY, ND, OH, PA, SC, TX, VA, WA, WV, WI.AUSTRALIA, SPAIN, FRANCE, BELGIUM, NETHERLANDS, PERU, NIGERIA, SWEDEN, CANADA, SWITZERLAND, EGYPT, CHINA, GERMANY, PANAMA, TAIWAN, HONG KONG, VENEZUELA, SOUTH AFRICA, CHILE, MEXICO, KOREA, COLOMBIA. | Terminated | June 07, 2012 | 835 | 95
Z-1597-2010 | 55254 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD Visitec EdgeAhead MVR Knife, BD Opthalmic Systems, Waltham, MA. |  |  | April 06, 2010 | 2010 | Becton Dickinson and Company | A field action was initiated due to a recent customer complaint from Japan there is a labeling mixup and that some of the BD Visitec EdgeAhead MVR Knife 0.90mm (20G), REF 585230, shelf packs may contain BD Visitec EdgeAhead Slit Knife 2.3mm (40 degree), REF 8009918. | Packaging process control | BD Medical issued a "Voluntary Advisory Notice" dated April 6, 2010. Consignees were informed of the affected product and asked to complete the enclosed Customer Response Form indicating whether the product will be returned to the firm.For further information, contact BD Medical Opthalmic Systems at 1-781-906-7952 or via email at ahusebo@bd.com. |  | 1170 | 1170 (US 810, OUS 360) | Worldwide Distribution -- United States, Canada and Japan. | Terminated | August 06, 2012 | 853 | 1170
Z-1599-2010 | 55503 | No | LGW |  |  | Neurology | PMA | Eon Mini Implantable Pulse Generator (IPG). Advanced Neuromodulation Systems, Inc. Plano, TX. |  |  | April 23, 2010 | 2010 | Advanced Neuromodulation Systems, Inc | Due to a missed inspection step during manufacturing, a small number of devices may experience a gradual or sudden loss of charging capability. | Manufacturing material removal | St. Jude Medical issued an "Important Medical Device Information" notification dated April 2010. Consignees were informed of the affected product and instructed on necessary steps. Any returned product should be sent to the firm.For further information, contact your St. Jude Medical Neuromodulation Division Sales Representative or call 1-972-309-2154. |  | 7 | 7 devices | Nationwide Distribution -- GA, NC, CA, OK, SC and NJ. | Terminated | July 27, 2015 | 1921 | 7
Z-1603-2010 | 53197 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Stryker Endoscopy PrepStrap Shoulder Sleeve, Model Number 3910900200, Manufactured by Stryker Endoscopy, San Jose, CA |  |  | July 27, 2009 | 2009 | Stryker Endoscopy | Latex-- Certain lots of Product contain undeclared latex. | Nonconforming Material/Component | An Urgent: Device Removal notice, dated 7/27/2009 was sent to the sole distributor, Stryker Orthopaedics in New Jersey. The distributor will execute the sub-recall. The firm intends to scrap all returned product at that facility, who will provide a record of disposition. The distributor will notify consignees via a mail packet that will be sent via FedEx overnight mail. Consignees will be instructed to return all product to the distributor, with credit provided per policy. |  | 842 | 842 units | Nationwide Distribution | Terminated | December 08, 2010 | 499 | 842
Z-1602-2010 | 55184 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Polarus(R) Humeral Fixation System, size: 3.5mm x 25.0mm Cortical Screw, QTY 1, Ref: HCO3250-S Acumed, LLC Hillsboro, OR |  |  | May 19, 2009 | 2009 | Acumed LLC | Mix-up of parts - Product labeled cortical bone screw actually contains Acutract fixation screw | Employee error | On May 18, 2009, the firm, ACUMED, began calling customers. On May 19, 2009, the ACUMED sent e-mails and a "URGENT NOTICE: DEVICE RECALL" Initial Notice letter to all customers. The letter describes the product, problem and actions to be taken by customers. The customers were instructed to identify the product having the lot code(s) and quarantine them, return all products from the lot code(s) to Acumed immediately, to please notify the customers of this recall if you further distributed this product, and complete and return the Acumed Product Recall Effectiveness Form.Please contact Acumed Customer Service at 1-888-627-9957 for any questions regarding this notification. |  | 83 | 83 | Worldwide distribution: USA including FL, GA, MI, MN, NY, TX, VA and countries of United Kingdom, Herzeles, South Korea, South Africa, China, and Milan | Terminated | May 17, 2010 | 363 | 83
Z-1606-2010 | 55317 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | EXXCEL Soft Standard Wall EPTFE Vascular Graft; D:8 mm; L: 50 cm straight. |  |  | March 26, 2010 | 2010 | Maquet Cardiovascular, LLC | This is an expansion of Recall # Z-866-2010, adding an additional two lots to the recall of EXXCEL soft Stand Wall EPTFE Vascular Grafts due to a labeling mix up. | Labeling mix-ups | The firm, MAQUET Cardiovascular, LLC, sent a "DEVICE REMOVAL IMMEDIATE ACTION REQUIRED" letter dated March 23, 2010, to all customers via Federal Express. The letter describes the product, problem and action to be taken by customers. The customers were instructed to examine their stock immediately to determine if they have any of the affected product from the lot, discontinue dispensing the lot and complete and return by fax (973-709-6511) the enclosed Field Action Response Form whether or not they have any affected products.Please contact your local MAQUET cardiovascular sales representative or Customer Service at 1-888-880-2874 for the exchange. |  | 22 | 22 units | Worldwide distribution: USA including states of TX, NJ, SC, MO, ID, TN, FL, CA, PA and Canada and Germany. | Terminated | April 23, 2012 | 759 | 23
Z-1605-2010 | 55498 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Universal Flat Panel Yoke, a component of the Stryker Flat Panel Arm. The yoke is intended to hold a flat panel monitor. |  |  | April 07, 2010 | 2010 | Stryker Communications Corp | Some of the pitch shafts were not properly welded which results in the monitor becoming unlevel. | Nonconforming Material/Component | An Urgent: Medical Device Recall letter dated April 19, 2010 was sent to customers. It was recommended that customers continue use of the product unless the monitor becomes unlevel. The customer was notified that a Stryker Representative would contact their facility to coordinate a service call to inspect the affected units. Questions are directed to the firm via phone at 972-410-7617 or via email at |  | 1498 | 1498 | Worldwide --US and Canada, Australia, China (Hong Kong), Latin America (Mexico, Chile, Ecuador, Costa Rica) . | Terminated | March 30, 2011 | 357 | 1498
Z-1604-2010 | 54890 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell GammaPlan 8.0 image fusion |  |  | January 07, 2008 | 2008 | Elekta, Inc. | The precision of the calculation used to create the fused study in LGP 8.0 is too low and should not be used until the system is upgraded to LGP 8.2. | Other | Phase 1 - Important Notice and Technical Note #200 058 - "Accuracy problems in Leksell GammaPlan 8.0 image fusion", dated December 10.2007 and Field Change Order #200 058, "Technical Note: Accuracy problems in Leksell Gamma Plan 8.0 image fusion", dated January 7, 2008 were sent to affected customers and all field service engineers as notification of the problem. The notice describes the product, problem and actions to be taken by customers. The firm recommended that the customers not use the fusion function until their Leksell GammaPlan(R) is upgraded to version 8.2.For further question, do not hesitate to contact your Elekta support respresentative at |  | 66 | 66 units | Nationwide distribution: AL, AR, CA, CO, CT, FL, GA, IL MD, MA, MI, MN, MS, MO, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, and WA. | Terminated | December 28, 2010 | 1086 | 66
Z-1608-2010 | 55428 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coated VICRYL RAPIDE (polyglactin 910) Braided Coated Synthetic Absorbable Suture, Undyed. |  |  | April 20, 2010 | 2010 | Ethicon, Inc. | Channels in the packaging seal that can lead to suture degradation and loss of packaging integrity. Recall Z-0021-2009 was initiated in 2008. Ethicon had decided to recall all sutures product during the may to June 2007 timeframe. | Other | "Expansion of 2008 Voluntary Recall" notification letters dated April 20, 2010 were sent by second day UPS to customers, distributors, Materials Managers, Director of Surgical Services, and Directors of Risk Management. The letters state that the expansion includes 37 lots of products. The letters describe the product, problem and action t to be taken by customers. The customers were instructed to immediately discontinue use of any of the affected product and to remove from all inventory and returned to Stericycle. If they are not returning product, they should mail the enclosed business reply card to Stericycle. The customers should notify their end customers of this recall. If their end customers have the affected product they should first call 1-888-345-0483 to get appropriate shipping labels. THIS RECALL EXPANSION ONLY IMPACTS THE VICRYL RAPID SUTURE. THIS DOES NOT INCLUDE THE VICRYL* ADN VICRYL* PLUS PRODUCTS.Please don't hesitate to contact Christie Bielinski directly at 908-218-2553 or via email at cbielinsi@its.jnj.com should you have any questions or require additional information. |  | 8275666752 | 118332 eaches were sent to the US Distribution Center, of which 69936 eaches were shipped to primary consignees. | Worldwide distribution: USA and Austria, Belgium, Canada, Chile, China, Czech Republic, Egypt, Germany, Greece, Hong Kong, Hungary, Israel, Italy, Japan, Korea, Malaysia, Middle East, Netherlands Antilles, New Zealand, Poland, Portugal, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom | Terminated | July 19, 2011 | 455 | 8275666752
Z-1609-2010 | 55280 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Odyssey |  |  | March 23, 2010 | 2010 | PerMedics, Inc. | Anomaly 1, Invalid Characters in Patient ID Error:Currently, Odyssey considers a valid ID to be comprised of numbers and/or letters. If a patient ID in a study contains a character other than a letter or number (including spaces). Odyssey removes the invalid character when moving it to a patient folder. Due to this, it is possible that two different patient studies could be listed within Odyssey | Other | A letter dated March 23, 2010 was mailed to customers acknowledging the anomalies and informed the reader that the mandatory version correcting the software will be provided by May 19, 2010. Customers with version 4.4 will receive 4.4.04:t hose with version 4.5 will receive 4.5.04; and those with version 4.6 will receive 4.6.01 containing the corrections to these issues. In the meantime, the following recommendations are provided. For Anomaly I . Invalid Characters in Patient ID Error, it 1s recommended that the users do not create patient IDS that contain characters other than letters or numbers. For Anomaly 2.0dysse.v Muchitle Library Collimator Type Error, it is recommended that the customer contact PerMedics ifany modifications to the library have been performed. The primary view should also be examined and used to detect and avoid the conditions causing the anomaly. Finally, independent field shape and monitor unit checks should be performed as part of a routine quality assurance program. For Anomaly 3, OOdyssey MLC Block Error, it is recommended to avoid using the MLC Block feature until the mandatory version correcting this anomaly is installed. If you need to use this feature, then it is recommended that a physical block tray is used during treatment for each beam that contains an Odyssey MLC Block. If customers believe a patient treatment has been affected by any of these issues, they should notify Customer Support as soon as possible at 877-4PerMed (877-473-7633) or support@permedics.com |  | 16 | 16 | Worldwide distribution: USA including NY, GA, CA, MA, and countries of Brazil, Hungary, China, India, Italy, Mexico. | Terminated | July 30, 2012 | 860 | 16
Z-1596-2010 | 55165 | No | LOL | Hepatitis A virus (HAV) serological assays. | Microbiology | Microbiology | 510(k) | ARCHITECT HAVAB-M Reagent |  |  | December 30, 2009 | 2009 | Abbott Diagnostics International, Ltd. | There is a potential to generate a false grayzone or reactive HAVAB-M result when the ARCHITECT AUSAB assay precedes the ARCHITECT HAVAB-M assay. | Pending | Direct accounts (hospitals, private laboratories, and other institutions) were sent a Product Correction letter, dated 12/30/2009. The letter identified the affected product, explained the issue and patient impact, and listed several necessary actions. Customers were instructed to follow one of 4 different alternatives to protect the integrity of test results. The letter is to be retained for laboratory records and a copy should be sent to other laboratories that kits were forwarded to. Questions regarding this information should be directed to Customer Service at 1-877-4ABBOTT. |  | 1326 | 1,326 | Nationwide Distribution -- Including Puerto Rico. | Terminated | August 31, 2011 | 609 | 1326
Z-1610-2010 | 55215 | No | GAG | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Disposable Skin Stapler - 35 Wide; Surgical Skin Stapler; Rx, sterile, single use skin stapler, individually packaged in a thermo-formed tray with a Tyvek lid, 5 trays per box; Distributed by Cardinal Health, Dublin, OH 43017, Made in China; catalog number FS-35W. |  |  | April 01, 2010 | 2010 | Cardinal Health | The sterility of the staplers is questionable. | Packaging process control | Cardinal Health issued an "Urgent: Product Recall" letter dated April 1, 2010 to all customers who may have received affected product. Customers were instructed to immediately quarantine any affected inventory, return the notification acknowledgement form and return the product to the firm.For further information, contact Cardinal Health Customer Service at 1-888-444-5440. |  | 1638 | 1,638 units | Nationwide Distribution | Terminated | August 17, 2011 | 503 | 1638
Z-1612-2010 | 55508 | No | HWX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 5.5MM Bone Tap, Modular; Manufactured by: Blackstone Medical, Inc. 1720 Bray Central Drive, McKinney, TX 75069. |  |  | April 21, 2010 | 2010 | Blackstone Medical, Inc. | Certain lots of the non-cannulated Modular Bone Taps used with the Firebird Spinal Fixation Systems may contain a manufacturing non conformance resulting in the taps creating a single lead screw thread rather than the required dual lead screw thread. | Nonconforming Material/Component | Orthofix issued an "Urgent Product Recall" notification dated April 21, 2010 by USPS certified mail to hospitals and distributors. Consignees were asked to return all affected product to the firm.For additional information, contact the Orthofix Regulatory Affairs Department at 1-973-406-2847. |  | 94 | 94 | Nationwide and to one location each in Australia and Spain. | Terminated | August 10, 2010 | 111 | 465
Z-1616-2010 | 55555 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi AIRIS, AIRIS II, AIRIS Elite MRI Scanner Compensation Coil. Part numbers: MR-FC-2EO, MR-FC-3EO |  |  | April 19, 2010 | 2010 | Hitachi Medical Systems America Inc | Hitachi determined that the mounting method for a Compensation Coil did not include an appropriate factor of safety, so it is possible that the coil could fall, contacting the patient table and potentially putting a patient at risk. | Other | On 4/15/2010 the firm sent Device Correction Notice letters to their customers identifying the affected product and informing them of the potential problem. The letter stated that the firm will change the coil mounting method used at the customer's facility, if necessary, at no cost to the customer. A service rep will contact them to schedule a visit to correct the problem. |  | 56 | 56 devices | Nationwide Distribution -- AL, AR, AZ, CA, CT, DC, DE, FL, ID, IL, IN, KY, LA, MD, MO, NJ, NY, OH, OK, OR, RI, TN, TX, VA &amp; WA. | Terminated | September 21, 2012 | 886 | 56
Z-1617-2010 | 55504 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport e360 Ventilator |  |  | April 14, 2010 | 2010 | Newport Medical Instruments Inc | The recall was initiated due to a recent increase in complaints related to the early failure of a low pressure transducer which is on Printed Circuit Board (PCB) 2104A in the Newport e360ventilator. Ventilators manufactured between September and December 2009 areaffected by this field correction. All affected ventilators require the replacement of the PCB2104A board to avoid early failure of t | Other | The firm, Newport Medical, sent a "Field Correction Notice" dated April 14, 2010, to all customers. The notice describes the product, problem and actions to be taken by the customers. The customers were instructed to: 1) provide your affected customers with a copy of this letter and advise them to pay close attention to pressure errors which can be an early warning sign of the pressure transducer failure; 2) Schedule their ventilators for replacement of the PCB2104A board as quickly as possible to avoid any unanticipated disruption in their use of their product;3) Contact our Technical Support Department (800.451.3111 for US customers, 714.427.5811x500 for international customers) to request the replacement kits you will need; and4) Complete the attached form and return it to Newport Medical as you complete your PCB2104A replacements. This will allow us to ensure that all affected ventilators have been addressed and close this field action.For questions, customers were instructed to contact Newport Medical Technical Support Department (800.451.3111 for US customers, 714.427.5811x500 for international customers). |  | 990 | 990 | Worldwide distribution: USA including states of CA, KY, NM, NY, UT and internationally to: Armenia, Bolivia, Brazil, Bulgaria, China, Colombia, Ecuador, Egypt, Greece, Hong Kong, India, Malaysia, Mauritius, Mexico, Nicaragua, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Russia, Saudi Arabia, Serbia, South Africa, Sri Lanka, Syria, Taiwan, Thailand, Tunisia, Turkey, UAE, Ukraine, Vietnam. | Terminated | August 18, 2011 | 491 | 990
Z-1628-2010 | 53212 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN 1700, list numbers 03H5301 and 03H53-03, manufactured by Abbott Hematology, Santa Clara, CA |  |  | June 08, 2009 | 2009 | Abbott Laboratories | Use of Syringe LN 04H36-01 with the Die-cast syringe drive may result in leakage, leading to delay in reporting patient results. | Nonconforming Material/Component | A Product Correction Letter and Customer Rely Form were sent to all currently active Cell -Dyn Sapphire customer. New customers are being notified with a Product Information Letter.For questions regarding this information, U,S, customers should call Customer Support at 1-877-4ABBOTT. Customers outside the U.S., please contact your local hematology customer support representative. |  | 274 | 274 | Worldwide distribution: United States, Canada, Mexico, Anguilla, Australia, Argentina, Belize, Bermuda, Brazil, China, Chile, Columbia, Costa Rica, Ecuador, El Salvador, Germany, Guatemala, Haiti, Honduras, Hong Kong, Japan, Paraguay, Peru, Puerto Rico, Singapore, South Korea, Uruguay, and Venezuela. | Terminated | December 03, 2010 | 543 | 2090
Z-1627-2010 | 52527 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Endotec Prothesis Ankle, Titanium Alloy, Sterilized, BP Ankle System. |  |  | June 05, 2009 | 2009 | Endotec, Inc. | Endotec is recalling B-P Hip, B-P Shoulder, and B-P Knee Replacement System devices due to sterility issues involving validation, process variations and test results. | Other | Recall Letters were sent to each distributor on 06/07/2009. Questions are directed to Endotec at (407) 822-0021. |  | 156 | 156 Pieces | Worldwide distribution -- US in OH, TX, CA, FL, NJ, TN; and International Distribution: Korea, Australia, Switzerland, The Netherlands, Spain.Domestic: OH, TX, CA, FL, NJ, TN. | Terminated | October 22, 2013 | 1600 | 3477
Z-1620-2010 | 49126 | No | LOM |  |  | Microbiology | PMA | Bio-Rad's Monolisa (TM) Anti-HBs EIA test kit, Labeled as, "MONOLISA TM Anti-HBs EIA 25220; |  |  | August 11, 2008 | 2008 | Bio-Rad Laboratories | The package insert provided with this kit contains a typographical error, giving incorrect directions to prepare a Working Conjugate Solution. If these directions are followed, it would result in a 1:6 dilution and invalidate the assay. | Packaging process control | All domestic customers were notified via an Urgent Product Correction customer letter sent by DHL or FedEx on 8/11/08. Enclosed with the letter was a corrected package insert. The Canadian subsidiary was notified via e-mail on 8/11/08. The letter advised customers of the correct wording, and advised that a new corrected package insert was enclosed and to make note of the correction in the laboratory procedures. Questions should be directed to Bio-Rad Laboratories Technical Support at 1-800-224-6723. |  | 795 | 795 kits | Worldwide Distribution -- USA and Canada. | Terminated | May 20, 2010 | 647 | 795
Z-1619-2010 | 55304 | No | GWL | Physiological signal amplifier. | Neurology | Neurology | 510(k) | Inserter Wand Holder, Non-Sterile, Serial # 4273-0000, Blackrock Microsystems, Salt Lake City, UT / or / Cyberkinetics Neurotechnology Systems, Inc., Salt Lake City, UT. |  |  | March 26, 2010 | 2010 | I2s Micro Implantable Systems Llc | Component of cranial array may not allow sufficient tightening to maintain inserter position. | Device Design | The firm, Blackrock Microsystems, sent a "RECALL" letter dated March 26, 2010, to all customers. The letter describes the product, problem and actions to be taken by the customers. The customers were instructed to discontinue use of all affected devices and promptly return any inventory, perform a physical account of your inventory for the recalled product and email a list of all serial numbers to Andrew Gotshalk at agotshalk@blackrockmicro.com, and to please email Andrew Gotshalk even if you do not have any inserter wand holders or have destroyed or removed from use the inserter wand holders. Blackrock will rework or replace all Inserter Wand Holders that is returned under this recall. If you have any questions regarding this recall please call 801-582-5533 x 275. |  | 22 | 22 units | Nationwide distribution: CA, MA, NY, RI, and UT. | Terminated | May 25, 2010 | 60 | 22
Z-1633-2010 | 55241 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Leica Wireless Footswitch, Type B used in conjunction with Leica Surgical Microscopes - Models M620, M844/820, M525 and M720; There are two wireless models: |  |  | March 29, 2010 | 2010 | Leica Microsystems, Inc. | The pedals of the Footswitch Type B stick and the Zoom/Focus moves to the end position. | Other | Leica Microsystems sent Medical Device Correction letters dated 3/29/10 to the affected customers via first class mail on the same date, informing them of the potential for injury to the patient's eye if the cross pedal of the footswitch sticks. They were informed that their local Leica Microsystems representative will contact them to arrange for the replacement of the cross pedal. The letters also provided the interim necessary steps required to avoid the risk of patient injury while they are awaiting the cross pedal replacement, i.e. do not use the affected footswitches with a surgical microscope that has a BIOM accessory. The accounts were requested to complete the attached Acknowledgement Form, listing the serial numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. A company contact was provided in the letter for questions. |  | 3 | 3 units | Worldwide distribution -- U.S. (nationwide), Canada, Australia, Europe and Asia | Terminated | February 15, 2011 | 323 | 57
Z-1631-2010 | 53322 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific, LATITUDEÂż Communicator, Distributed by Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 55112. Models 6476 (USA) and 6468 (OUS) |  |  | September 10, 2009 | 2009 | Boston Scientific CRM Corp | Boston Scientific has determined that alert conditions for a limited subset of out-of-range Daily Measurements related to leads may not be recognized by Model 6476 (United States) or Model 6468 (Europe) LATITUDE in-home communicators. If a communicator fails to recognize a valid alert situation, notification for this or subsequent alerts for the same test will not be displayed for physician review | Software design | Consignees were sent on 9/10/09 a Boston Scientific "Urgent Medical Device Information" letter dated September 9, 2009. The letter was addressed to "Dear Doctor". The letter provided the information related to the LATITUDE Patient Management system recall as to Background, Description, Rate of Occurrence, Patient management, Devices Affected and Further Information. Patients were also sent a Boston Scientific letter dated September 9, 2009. The letter addressed Background information, Why are we communicating, What You Should Do and Questions. |  | 25656 | 25,656 (USA), 390 (OUS) | All States in the US including Puerto RicoUnited Kingdom, Switzerland, Netherlands, Italy, Germany, France, Finland, Denmark, Belgium, Austria | Terminated | April 05, 2012 | 938 | 25656
Z-1639-2010 | 54174 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Churchill Medical Systems Bag Access Device w/Smartsite Needleless Injection Site, Sterile, Non pyrogenic, Latex Free Product Code: KSA-140. Churchill Medical Systems, Inc. Dover, NH 03820. |  |  | October 20, 2009 | 2009 | Churchill Medical Systems, Inc. | Bond between the spike and needle may leak, break or detach resulting in leakage or spillage of chemotherapeutic agents | Process change control | Churchill Medical Systems issued an "Urgent: Product Advisory Customer Notification" dated October 20, 2009 requesting return of affected product and notification to end users. For further information, contact Churchill Medical Systems at 1-603-743-5988 or via fax at 1-603-743-6328. |  | 283900 | 283,900 each | Distributed to Florida. | Terminated | March 28, 2013 | 1255 | 283900
Z-1636-2010 | 55436 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ADVANCE DURAMER TIBIAL INSERT, REF: KIMP-312L, SIZE 3 LEFT MEDIAL PIVOT, THICKNESS 12MM, 1 EACH, STERILE EO, Wright Medical Technology, 4677 Airline Road, Arlington, TN 38002. |  |  | November 24, 2009 | 2009 | Wright Medical Technology Inc | The left medial pivot and right medial pivot inserts were co-mingled | Packaging process control | Wright Medical Technology, Inc. issued an "Urgent: Medical Device Voluntary Recall" notification dated November 24, 2009. The letters were addressed to distributors, hospital administrators, and surgeons. Consignees were asked to return all affected product in addition to a completed Response Form to the firm.For further information, contact Wright Medical Technology, Inc. at 1-800-874-5630. |  | 23 | 23 units | Worldwide Distribution -- United States, The Netherlands, Italy, Republic of Korea and Japan. | Terminated | October 21, 2010 | 331 | 70
Z-1635-2010 | 55407 | No | MOZ |  |  | Orthopedic | PMA | Euflexxa (1% sodium hyaluronate). Pre-filled syringe, Rx only, 3 syringes per carton; Manufactured by Ferring Pharmaceuticals In, Parsippany, NJ 07054; |  |  | April 16, 2010 | 2010 | Ferring Pharmaceuticals Inc | Out of Specification results for molecular weight or viscosity through the established expiration date. | Process control | Ferring Pharmaceuticals issued a "Medical Device Recall" notification dated April 16, 2010 by Federal Express to all distributors, retail pharmacies and physicians who received the product. Consignees were asked to stop distribution, return affected product to the firm and fax a completed Product Inventory Form to the firm.For further information, contact Ferring Pharmaceuticals at 1-973-796-1600 or via fax at 1-973-796-1660. |  | 88128 | 29,376 cartons with 3 syringes per carton. | Nationwide distribution. | Terminated | July 27, 2010 | 102 | 88128
Z-1645-2010 | 55477 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Ceiling Suspensions (Surgical Light and Flat Panel Suspensions); Catalog Numbers: 0682000863, 06820001009, 0682001010, 0682001076, 0682000864, 0682000857, 0682000858, 0682001077, 0682000865, 0682001011, 0682001078, 0682000866, 0682001012, 0682001080, 0682000867, 0682000959, 0682001079, 0682000850, 0682000851, 0682000852, 0682001269, 0682001567, 0682001569, 0682001570, 06820012691, 06820015671, 06820015691, &amp; 0680015701. |  |  | April 15, 2010 | 2010 | Stryker Communications Corp | Installation records are incomplete. | Employee error | An Urgent: Medical Device Product Correction customer notification letter, dated 4/15/2010, has been sent out to all affected accounts via regular mail in the form of a certified, return request receipt letter. Accounts will receive a follow-up phone call from a Stryker Representative to schedule a re-inspection of equipment to see if it mets established critical specifications or not. The firm recommends that customers continue using all Stryker Communications products listed in the notice. Questions can be directed to Katie Blackwell at 972-410-7660 or via email at StrykerCommunicationsRA@stryker.com. |  |  |  | Nationwide Distribution | Terminated | January 24, 2011 | 284
Z-1638-2010 | 52291 | No | HTS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | BD Beaver Arthro-Lok Pointed Tip 4mm Blade, REF 376780. Becton Dickinson and Company, Waltham, MA 02454. |  |  | June 12, 2009 | 2009 | Becton Dickinson and Company | Mislabeled: Arthro-Lok Pointed Tip 4mm Blade curving right instead of left. | Packaging process control | Becton Dickinson issued an "Urgent: Field Safety Notice" beginning June 25, 2009 by telephone and letter fax to remove and discard affected product from their possession. Customer verification responses were requested.For further information, contact Becton Dickinson and Company at 1-781-906-7900. |  | 354 | 354 units | Worldwide Distribution -- United States, Japan and Belgium. | Terminated | January 26, 2012 | 958 | 354
Z-1618-2010 | 55235 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator Monitor, Model M3535A |  |  | April 07, 2010 | 2010 | Philips Healthcare Inc. | ECG signal processing chip may fail causing failure to deliver therapy in AED mode or the failure to indicate poor electrical contact (very high impedance) with the patient in Manual Defib mode. | Component design/selection | The firm, Philips Healthcare, sent an "URGENT voluntary Medical Device Correction" Field Safety Notice, letter dated April 7, 2009, to all customers via UPS. The letter describes the product, problem and the actions to be taken by customers, with precautions to follow if the device is used before upgrade is installed. The customers were instructed to check their serial numbers for the affected product, if the device is affected, it may remain in service prior to correction using the following precautions 1) if using AED, have a back up AED; 2) if using external defibrillator paddles, disregard paddles and manually ensure good paddlet-to-patient contact, and 3) if monitoring, be prepared to switch back to the primary ECG that was initially obtained for a given patient if other leads are lost. Customers will be contacted and assisted by Philips field service engineers to install the software upgrade which corrects the problem.Should you have any questions or concerns about this Device Correction, please contact your local Philips representative at 1-800-722-9377. |  | 1379 | 1379 units | Worldwide distribution:USA and Canada, Argentina, Australia, Austria, Belgium, Brazil, China, Finland, France, Germany, Iceland, India, Ireland, Italy, Japan, Korea, Republic of Lithuania, Mexico, Netherlands, New Zealand, Norway, Poland, South Africa, Spain, Switzerland, Thailand, Trinidad, and Tobago, Ukraine, United Arab Emirates, United Kingdom. | Terminated | February 20, 2013 | 1050 | 1379
Z-1651-2010 | 55079 | No | OPU |  |  | Microbiology | EUA - Emergency Use Authorization | Cepheid Xpert Flu A Panel (GXFLUPANEL-10), manufactured by Cepheid, Sunnyvale, CA. |  |  | February 19, 2010 | 2010 | Cepheid | Firm has had an increasing number of customer complaints regarding higher than expected Indeterminate Test Results due to "Invalid" test reports from our system when they have been conducting their own internal validation studies using banked frozen repository specimens. | Other | Cepheid issued a notification to consignees regarding the affected product. Additional instructions were provided to consignees in order to avoid further "Indeterminate Results".For further information, contact Cepheid at 1-888-336-2743. |  | 1803 | 1803 kits | Worldwide Distribution -- United States, France, Belgium, Germany, Norway Austria, Denmark, Spain, Taiwan and Hong Kong. | Terminated | March 23, 2011 | 397 | 2747
Z-1663-2010 | 54268 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | Bausch &amp; Lomb Millennium Microsurgical System. |  |  | January 15, 2010 | 2010 | Bausch &amp; Lomb Inc | The AC power cords used on these products are subject to an FDA Alert -- Safety Investigation of Certain Medical Device Power Cords: Initial Communication, Date Issued: October 19, 2009. | Nonconforming Material/Component | A "URGENT: POWER CORD RECALL" letter and copies of the FDA Alert dated Oct. 19, 2009 were sent to all customers on or about Jan. 15, 2010, by Stericycle, a recall notification service. The letter describes the product, problem and actions to be taken by the customers as well as FDA's recommendations and actions. The customers were made aware of the FDA Alert concerning the AC power cords and notified that these power cords were used in the devices. They were requested to inspect the power cords for damage and notify the recalling firm if damage is noted or if they have experienced sparking, charring, or firing from the power cords and/or cracking or failure of the prongs. In the absence of those conditions, customers were given the option of returning the power cord for replacement and/or during the regularly scheduled servicing of the device. If you would like a new power cord, please contact Customer Service at 1-800-338-2020 for a free replacement. |  |  | Quantity never provided by firm | Worldwide distribution: United States &amp; Puerto Rico; Venezuela, India, Trinidad and Tobago, Colombia, Sri Lanka, Hong Kong, Vietnam, and Peru. | Terminated | May 17, 2013 | 1218
Z-1681-2010 | 54473 | No | ESD | Hearing Aid. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | InSound Medical, Inc, Lyric in-canal hearing aid. |  |  | January 19, 2010 | 2010 | InSound Medical, Inc. | Manufacturing error could result in electrolyte leakage from the product's battery. | Process design | InSound sent to its consignees a "Dear Lyric Provider" letter of explanation on 2/1/2010. |  | 2019 | 2019 units | Distribution only in US ( nationwide). No foreign consignees involved. | Terminated | August 02, 2011 | 560 | 2019
Z-1657-2010 | 55176 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CoulterÂż LH 500 Series System. Beckman Coulter, Inc. Brea, CA. |  |  | February 04, 2010 | 2010 | Beckman Coulter Inc. | Issue 1: The system will substitute or omit the characters # @ [\] ' {I) when the system is configured in a language other than English or Chinese. Example: Using the character @, a substitution in the string '123@456' would result in '123#456' and an omission would result in '123456'Issue 2: The system will omit the character ~ regardless of language.Impact: There is a potential for Specimen | Software design | Beckman Coulter, Inc. issued "Urgent: Product Corrective Action" notifications dated February 12, 2010. Consignees were informed of the affected product and instructed on proper actions. The firm also asked that the enclosed response forms be returned.For further information, contact your local Beckman Coulter Representative or call 1-800-526-7694. |  | 1520 | 1520 units | Distributed Nationwide and Canada. | Terminated | September 02, 2014 | 1671 | 1576
Z-1654-2010 | 55300 | No | LNR |  |  | Gastroenterology/Urology | PMA | Therakos CellEx Photopheresis System; |  |  | March 29, 2010 | 2010 | Therakos Inc | Higher than normal blood leaks on THERAKOS CELLEX Procedural Kits that are intended for use with the THERAKOS CELLEX Photopheresis System | Nonconforming Material/Component | On March 29, 2010, Therakos Photopheresis sent Important Product Correction Notification letters by FedEx to all direct end use and distributor customers who received lot X120 in addition to phone contact. On March 31, 2010 those who received lot X123 were sent letters by Fed Ex as well as phone contact.. The one Canadian customer was notified by phone on March 31, 2010 followed by a Fed Ex letter. On April 12, consignees who were shipped Lot X119 were contacted by phone, followed by a hard copy letter.On May 7, 2010 all direct end use and distributor consignees were notified of the additional lot being added to the recall. Questions may be directed to Therakos Customer Technical Services at 1-877-865-6850. |  | 840 | 280 units in US (1 sales unit contains 3 kits each) | Nationwide Distribution, and Canada and Germany | Terminated | July 18, 2013 | 1207 | 840
Z-1653-2010 | 55002 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Blackstone Medical Inc./ Orthofix Spinal Implants, Hallmark Anterior Cervical Plate System; Plate Bender Part number 62-0070. |  |  | March 15, 2010 | 2010 | Blackstone Medical, Inc. | The Hallmark Anterior Cervical Plate Bender may have a non conformity causing difficulty accepting 3-4 level plates or to not accept 3-4 level plates that are intended to be additionally contoured. | Nonconforming Material/Component | Orthofix send letters of notification of planned product removal to distributors who keep the instrumentation set on 3/29/2010 by Fed Ex. Letters of planned product removal explain the reason for the planned removal and the specific product involved. Replacement product will not be available until July 2010, however, the firm elected to notify customers of the potential problem and the date that formal notification and replacement Plate Benders will be available. Letters state that if there is any problem with the Plate Benders, questions can be directed to David Thompson in Marketing at 972-529-3024 or Jose Charvet in Product Development at 973-406-2824. |  | 112 | 98 units domestically, 14 units internationally. | Nationwide distribution, including Puerto Rico, and to Australia. | Terminated | September 09, 2011 | 543 | 112
Z-1659-2010 | 55445 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 4100 IQ: GE Innova 4100 / 4100 IQ., Cardiovascular Imaging System (The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | March 24, 2010 | 2010 | Ge Healthcare, Llc | A potential software failure related to the Worklist functionality has been identified on a subset of Innova 2100IQ, 3100IQ and 4100IQ systems. When using the Worklist function to import patient data to Innova, there are two scenarios thatwill cause the next exam to fail and the inability to recall acquired sequences. | Software design | Consignees were sent on 3/24/10, a GE Healthcare " Urgent Medical Device Correction" letter dated March 23, 2010. The letter included the Safety Issue, Affected Product Details, Product Correction and Contact Information. |  | 28 | 28 | Worldwide distribution: USA, Puerto Rico, VENEZUELA, UNITED KINGDOM, TURKEY, THAILAND, TAIWAN, PROVINCE OF CHINA, SWEDEN, SPAIN, RUSSIAN FEDERATION, ROMANIA, POLAND, PHILIPPINES, NORWAY, NETHERLANDS, MOROCCO, REPUBLIC OF KOREA, KAZAKSTAN, JAPAN, ITALY, ISLAMIC REPUBLIC OF IRAN, INDONESIA, INDIA, GERMANY, FRANCE, FINLAND, EGYPT, CHINA, CHILE, CANADA, BRAZIL, BELGIUM, BELARUS, AUSTRALIA, and ARGENTINA. | Terminated | May 21, 2013 | 1154 | 195
Z-1656-2010 | 55194 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number 629029. Beckman Coulter, Inc. Brea, CA. |  |  | March 02, 2010 | 2010 | Beckman Coulter Inc. | Issue 1: The Diagnostic procedure, Simulate-CDNR Cycle, is not functioning properly. Impact: There is a potential for inadequate sample preparation of samples following the use of this Diagnostic function impacting WBC, Differential, Reticulocyte and/or NRBC results.Issue 2: Use of the Auto Prune function causes the deletion of Quality Control files.Impact: If the last run for a control lot | Software design | Beckman Coulter, Inc. issued an "Urgent: Product Corrective Action" notification dated March 10, 2010. Consignees were informed of the affected product and instructed on proper steps to take and requested to return the enclosed acknowledgement form. For further information, contact your local Beckman Coulter Representative or call 1-800-526-7694. |  | 63 | 63 units | Nationwide Distribution (AZ, CA, CO, FL, GA, IA, IL, IN, KS, LA, MO, NC, NJ, NV, NY, OH, PA, SC, TN, TX, UT, WA, and WI) and Canada. | Terminated | May 23, 2012 | 813 | 63
Z-1655-2010 | 55361 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta XiO Radiation Treatment Planning System |  |  | April 15, 2010 | 2010 | Computerized Medical Systems Inc | Under certain circumstances, XiO will over-estimate dose in the region where the backup jaw is inside the MLC opening | Software design | Elekta issued an undated "Important Safety Notice" to all affected sites via e-mail on 4/20/2010. The notice was also sent to the sites on 4/6/10 and 4/12/10 (for 4 sites) before the firm made the decision to recall. The letter explains the correct functioning of the system, the problem, when and why it occurs, the clinical impact, and the work around. Although not stated in the Safety Notice, a self-addressed postcard was enclosed for the customer to return as evidence they received and understood the notice. Sites that do not return the card will be contacted by phone. |  | 34 | The release is being used by 34 customers that also have an Elekta linear accelerator | Worldwide Distribution: USA, Canada, Australia, Spain, Italy, the Netherlands, United Kingdom, Taiwan, Malaysia, France, Ireland, and Israel. | Terminated | June 07, 2012 | 784 | 34
Z-1680-2010 | 55462 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029. Beckman Coulter, Inc. Brea, CA. |  |  | December 08, 2009 | 2009 | Beckman Coulter Inc. | The Corrective Action was initiated after Beckman Coulter confirmed three issues related to host transmissions with the UniCel DxH 800. This information only affects instruments interfaced to a host system. | Software design | Beckman Coulter, Inc. issued an "Urgent: Product Corrective Action" notification dated December 8, 2009. Consignees were informed of the affected product and provided temporary solutions until further support will be made available.For additional information, contact your local Beckman Coulter Representative or call 800-526-7694. |  | 42 | 42 units (40 in US) | Distributed Nationwide and Canada. | Terminated | June 04, 2012 | 909 | 42
Z-1683-2010 | 53668 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Responder 2000 Defibrillator/Monitor |  |  | October 15, 2009 | 2009 | Cardiac Science Corporation | The Operators Manual may not have adequate information to clearly explain some device behaviors. Pressing and holding down the Charge or Shock button on the Responder 2000 for more than 5 seconds causes a "Button Stuck" error. With the "Button Stuck" error, the device will automatically exit Manual mode and enter Monitor mode. The user cannot apply therapy while the device is in Monitor mode. | Other | The firm, Cardiac Science, sent an "Urgent-Medical Device Corrective Action" letter dated October 15, 2009, was sent to customers. The letter describes the product, problem and action to be taken by the customers. Cardiac Science requested GE to provide the end users with the Operator's Manual addendum. The customers are instructed to place the Operator's Manual addendum with their Responder 2000 Operator's Manual for future reference.If you have any questions about your Responder 2000, please contact Cardiac Science technical support at +44.161.926.0011 or our EU Representative, Ludger Moller, MDSS GmbH at +49.511.6262.8630. |  | 1838 | 1838 devices | Worldwide distribution: USA and Australia, Brazil, Columbia, Costa Rica, Denmark, Finland, France, Germany, Italy, Korea, Mexico, New Zealand, Panama, Paraguay, Peru, Singapore, South Africa, Spain, and Uruguay. | Terminated | December 07, 2011 | 783 | 1838
Z-1682-2010 | 55385 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow 4FR Peripherally Inserted Central Catheter (PICCs) Kits. Product numbers: ASK-01451-BMH; ASK-05041-CHC; ASK-05041-CHC1; ASK-05041-CHM; ASK-05041-HMC; ASK-05041-QV; ASK-07041-BMH; CN-05041-LW; MSO-01451-UCL; MTO-01451-RH; PL-05041; PR-04041-HPX; PR-05041; PR-05041-HP; PR- 05041-HPX; PR-05041-LW; PR-05041-MW; PR-05041-T; PR-05042; PR-05541-HPX; PR-07041-PTSP; and RJ-01451-W. |  |  | March 29, 2010 | 2010 | Arrow International Inc | Tight fit in sheath. Arrow has received complaints of the PICC catheter fitting too tightly in the "kit supplied" peelable sheath. In some cases, clinicians have been unable to insert the catheter through the peelable sheath during the insertion procedure. If the catheter will not pass through the sheath, the clinician is required to place a guide wire back into the sheath, remove the sheath, a | Nonconforming Material/Component | Arrow International issued an Urgent Medical Device Recall letter dated 3/29/10 to inform their customers of the problem and the need to return the product. |  | 6719 | 6719 | Worldwide distribution: USA, Australia, Brazil, Canada, New Zealand, Hong Kong, Malaysia, and Chile. | Terminated | September 27, 2012 | 913 | 6719
Z-1668-2010 | 55365 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Mallory Head Calcar Femoral, 34 mm resection/porous coated, 13.0 x 170 mm,, Sterile, REF 11-104904, Biomet Orthopaedics, Warsaw, IN. |  |  | March 10, 2010 | 2010 | Biomet, Inc. | The firm has found that some lots of Mallory Head Revision Stems had flute geometry that was out of specification. The flute geometry was too deep. | Process change control | Biomet Orthopedics issued an "Urgent Medical Device Recall Notice" dated March 10, 2010. The letter went to distributors and users of the recalled device. The letter instructed the consignees to locate the recalled product and return the product to the firm.For further information, contact Biomet at 1-800-348-9500 or 1-574-372-3983, Monday through Friday, 8:00 AM to 5:00 PM. |  | 6 | 6 | Worldwide Distribution -- United States (OH, IN, FL, KY, NV and CA), Mexico, Chile, China, Japan and The Netherlands. | Terminated | October 18, 2010 | 222 | 81
Z-1704-2010 | 54168 | No | GWJ | Evoked response auditory stimulator. | Neurology | Neurology | 510(k) | Intelligent Hearing Systems Corporation, SmartEP with Vestibular Evoked Myogenic Potentials Version 3.96. Model M811016 |  |  | October 21, 2008 | 2008 | Intelligent Hearing Systems Corp. | This was a marketing correction mandated by the United States Food and Drug Administration (USFDA). USFDA notified Intelligent Hearing Systems that 501(k) premarket notification application must be submitted and cleared in order to market and sell devices with Vestibular Evoked Myogenic Potential (VEMP) specific functions. The Food and Drug Administration considers that the VEMP testing functions | Device Design | Intelligent Hearing Systems sent a letter "Information Regarding the Discontinued Use of the SmartEP Evoked Response System for Vestibular Evoked Myogenic Potentials (VEMP) Testing" dated September 30, 2008 to customers and distributors. Each customer will receive an upgrade and VEMP removal kit which will remove the VEMP feature. |  | 50 | 50 | Nationwide distribution. | Terminated | September 21, 2011 | 1065 | 50
Z-1686-2010 | 55529 | binder, medical, therapeutic | MDR | Therapeutic medical binder. | General Hospital | General Hospital | 510(K) Exempt | Juxta-Lite compression wrap XL Short 28cm Length with Anklet Catalog 23026017 |  |  | April 20, 2010 | 2010 | Circaid Medical Products Inc | Some systems in these lots have Built-In Pressure System lines with the incorrect amount of spacing. Some systems in these lots had I inch space between each line with no tension applied. The system should have 1.25 inches of space. The correct spacing (1.25") is necessary to achieve the measured tension ranges. The incorrect spacing may result in the inability to achieve the correct tension level | Process control | CircAid Medical Products Important Device Product Recall letter dated Apr 20, 2010 was mailed to consignees describing problem, listing affected product and distribution dates. Customers were instructed to locate and return these Juxta-Lite Systems from stock and notify their customers A Retum Authorization number and contact name was provided.this listing for product returns. |  | 372 | 372 total, all sizes | Worldwide distribution: USA, Japan, Canada, Netherlands | Terminated | January 06, 2011 | 261 | 372
Z-1705-2010 | 55538 | No | HQD | Rigid gas permeable contact lens. | Ophthalmic | Ophthalmic | 510(k) | Boston ES Rigid Gas Permeable Contact lens manufactured with Fluoro Silicone Acrylate Rigid Gas Permeable Contact Lens Material. Bausch &amp; Lomb, Inc. Wilmington, MA 01887. |  |  | April 20, 2010 | 2010 | Bausch &amp; Lomb Inc | The amount of D&C; #6 Dye added to the formulation exceeded specified amounts. | Process control | Bausch &amp; Lomb issued an "Urgent Voluntary Medical Device Recall" notification dated April 20, 2010. Consignees were informed of the issue and asked to identify and return all affected product to the firm.For further information, contact Bausch &amp; Lomb, Inc. at 1-585-338-5477. |  | 2575 | 2575 buttons | Nationwide Distribution -- TX, IN and WI. | Terminated |  |  | 2575
Z-1706-2010 | 55570 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | AIR OPTIX (lotrafilcon B) for Astigmatism Soft Contact Lenses. CIBA Vision, Corp. Duluth, GA 30097. |  |  | April 27, 2010 | 2010 | Ciba Vision Corporation | The lenses inside the package do not match the prescription information for power labeled on the primary package. | Packaging process control | CIBA Vision, Corp. notified consignees of the affected product via letter beginning April 15, 2010. Users were asked to return the product to the firm.For further information, contact your CIBA Vision sales representative or call 1-877-542-5928. |  | 5040 | Lot #8717082, total 5040 lenses and Lot #8717066 total 5394 lenses | Nationwide Distribution -- AZ, AL, CA, FL, GA, IA, IL, MA, MO, NC, NJ, NY, OH, OR, PA TN, VA and WI. | Terminated | October 22, 2010 | 178 | 5040
Z-1707-2010 | 55715 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | PULSESPRAY INFUSION SYSTEM, 90 cm x 20 cm infusion length, Catalog No./Ref 12500402, STERILE, EO. AngioDynamics, 603 Queensbury Avenue, Queensbury, NY 12804. |  |  | May 17, 2010 | 2010 | Angiodynamics, Inc. | This lot was packaged with the wrong occluding ball guidewire. The PulseSpray Catheter is 90 cm in length and requires a 109 cm occluding ball guidewire, however this lot was packaged with a 60 cm occluding ball guidewire. | Process control | Angiodynamics issued an "Urgent Medical Device Recall" notification (dated 5/11/10) to user accounts via Certified Mail. Consignees were asked to identify, segregate and return all affected product to the firm in addition to sending in a completed Reply Form via fax.For further information, contact Angiodynamics Customer Service at 1-800-772-6446. |  | 20 | 20 units | Worldwide Distribution -- United States (AR, NY, PA, MI, MN, LA and WA), Spain, Austria and Denmark. | Terminated | July 25, 2012 | 800 | 20
Z-1703-2010 | 55744 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | S1813OB 1.8X13mm Sendax MDI Mini Dental Implant O-Ball Prosthetic Head |  |  | May 14, 2010 | 2010 | 3m Imtec Corporation | Product has Incorrect size Labeling. Actual length is 13mm. One label is correct, and one states 10mm. | Packaging process control | IMTEC notified consignees by phone on May 14, 2010 and by Certified Letter dated May 14, 2010. |  | 189 | 189 Implants | Nationwide. | Terminated | July 14, 2010 | 61 | 189
Z-1732-2010 | 55647 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis systems with Motor Controller Unit. Siemens Medical Solutions USA, Inc. Malvern, PA 19355. |  |  | April 12, 2010 | 2010 | Siemens Medical Solutions USA, Inc | sporadic malfunctions of the Artis Systems with Motor Controller caused by tolerances on electronic parts that might increase during the lifetime of the system. | Device Design | Siemens issued an "Urgent Field Safety Notice" to consignees describing the affected device instructing users to contact the recalling firm's service organization to schedule maintenance.For further information, contact Siemens Medical Solutions USA, Inc. at 1-610-448-3237. |  | 83 | 83 units | Nationwide Distribution | Terminated | April 11, 2013 | 1095 | 83
Z-1714-2010 | 54200 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb Stellaris Optimized Stability Vacuum Pack, Model BL5112, 6 packs/box, RX, Sterile. The firm name on the label is Bausch &amp; Lomb Incorporated, Rochester, NY. |  |  | January 08, 2010 | 2010 | Bausch &amp; Lomb Inc | Users may incorrectly assemble the plastic needle wrench to the phacoemulsification needle resulting in the generation of plastic particulate | Component design/selection | The recalling firm, Bausch &amp; Lomb, issued an "URGENT: PHACO NEEDLE WRENCH CORRECTION" letter dated January 7, 2010, via first class mail to all customers. The letter describes the product, problem and actions to be taken by the customers. The letter also provides instructions and pictorial representations demonstrating proper seating of the needle to the wrench as well as a description of the appropriate force to tighten the needle to the handpiece. The customers were instructed to please read and follow the attached instructions entitled (Potential for Particulate Matter in the eye and proper assembly and use of needle wrench) and fill-out and return/fax at 949-716-8362, the Bausch &amp; Lomb Phacoemulsification Needle Wrench Particulate correction form. If you have any questions or concerns please contact, Chris Willis or Stephen Haight at 1-800-387-3284. |  | 14685 | 14,685/6-pack boxes | Worldwide distribution: USA and Germany, France, Italy, Norway, Sweden, Finland, The Netherlands, Switzerland, U.K., Austria, Luxembourg, Ireland, Hong Kong, New Zealand, India, Australia, and Japan. | Terminated | April 02, 2012 | 815 | 72120
Z-1735-2010 | 55723 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Voiceware clip software update to certain Horizon Medical Imaging Systems. McKesson Medical Imaging Company. Richmond, Canada V6X 3G5. |  |  | May 07, 2010 | 2010 | Mckesson Medical Immaging | Voice Clip issue: When the user is recording voice clips in Horizon Rad Station, there is a potential for voice clips to be saved to the incorrect study of the same patient. | Software design | McKesson Medical Imaging Company issued an "Advisory Notice" dated May 7, 2010. Consignees were notified of the affected product and advised to contact McKesson Customer Support department to review their sites current configuration and install the appropriate product software update(s). For further information, contact McKesson Medical Imaging Customer Support at 1-800-663-2533. On June 19, 2010 the recall was amended to include the following actions to follow: Instead of highlighting the anchor study in the Patient Portfolio Sidebar before saving t he voice clip in the Patient Portfolio Sidebar, McKesson recommends that you save the voice clip to the desired study first, before browsing through other documents. A software update that prevents the problem from occurring is currently under development for all affected releases and will be available in approximately three months. Consignees will be notified when the software update is ready to be installed on their system. |  | 100 | 100 units | Distributed Nationwide (AK, AL, AR, AZ, CA, CO, FL, GA, IN, IA, IL, KY, LA, MA, MD, MI, NC, ND, NE, NM, NY, NV, OH, OR, PA, RI, SC, TN, TX, WY, VA and WV) and Canada | Terminated | August 15, 2012 | 831 | 100
Z-1734-2010 | 55713 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS BENEPHIT(TM) CV Infusion System, 35 cm, Targeted Renal Therapy Infusion system, Catalog/Model/REF Number 60035, STERILE, EO, Intended for Single Use Only. AngioDynamics, Incorporated 603 Queensbury, NY 12804. |  |  | April 30, 2010 | 2010 | Angiodynamics, Inc. | Benephit Infusion Catheters may exhibit a hole or tear in the primary sterile barrier packaging. The product is intended to be sterile. Breaches in the sterile packaging may result in non-sterile product, which could potentially cause a patient infection which is a serious medical condition. | Packaging change control | Angiodynamics, Inc. issued an "Urgent Medical Device Recall" notification (dated 4/28/10). Customers were advised of the affected product and instructed to return all recalled product to the firm.For further information, contact Angiodynamics, Inc. Customer Service at 1-800-772-6446. |  | 22 | 22 | Distributed Nationwide and to Italy | Terminated | July 25, 2012 | 817 | 763
Z-1731-2010 | 55661 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Vessel Sizing Angiographic Catheter, REF 7602-2M65, 5F PIG-VSC-2B, UHF Pigtail, Performa, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095. |  |  | May 05, 2010 | 2010 | Merit Medical Systems, Inc. | Marker bands on angiographic catheters may not be properly attached, allowing movement of the bands. | Process control | Merit Medical Systems, Inc. began notifying sales representatives beginning May 5, 2010 via phone and instructed to contact all of their affected accounts. All affected product was to be returned to the firm.For further information, contact Merit Medical at 1-713-757-5067 or 1-801-208-4468. |  | 96 | 96 units | Nationwide Distribution -- GA, IL, MO, NY, PA and WI. | Terminated | June 01, 2010 | 27 | 96
Z-1736-2010 | 55659 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | EVA 450 Patient Lift Battery Charger. |  |  | April 28, 2010 | 2010 | Romedic, Inc. | Incorrect battery chargers were supplied with the EVA 450 Patient Lifts. Charger unit stops working and will not charge lift. New chargers were supplied to customers. | Other | The recalling firm sent out a letter with instructions as follows: 1) Locate EVA 450 Patient Lift(s), 2) Fill out Customer Response Form and fax to 888-691-1091, 3) Replace battery charger(s) with new battery charger(s) supplied with this notification; NEW battery charger to be used according to EVA 450 Instructions, 4) Repackage replaced battery charger(s) in original packing in which shipment was recieved, 5) Apply supplied sticker "EVA 450 Battery Charger RECALL" on outside of box, 6) Apply supplied UPS return label on outside of box and contact UPS for pickup and mail to the following Address: RoMedic, Inc., 300 West Chestnut St., Ephrata, PA 17522; Attn: EVA 450 Battery Charger Recall.Additional questions are directed to the company at phone: 888-625-4343. |  | 76 | 76 units | Nationwide distribution. | Terminated | September 30, 2010 | 155 | 76
Z-1737-2010 | 55234 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ARTERIAL LINE KIT Arterial Line Kit #2 20ga X 3" REF/REORDER NO. 194284 Labeled Sterile, Rx. Argon Medical Devices, Inc. 1445 Flat Creek Road Athens, TX 75751 800-927-4669. |  |  | March 29, 2010 | 2010 | Argon Medical Devices, Inc | There are 1-2 places on some of the multi-purpose kit blisters/trays in which there are holes and/or the material is so thin that holes are easily created. This would potentially cause the sterility of the products to be compromised. | Package design/selection | Argon Medical Devices issued an "Urgent Medical Device Recall" notification via certified mail dated March 29, 2010 to distributors and end users. A Response Form was attached. Consignees were informed of the affected product and asked to quarantine and return all unused product to the firm.For further information, contact Argon Medical Devices at 1-903-677-9375. |  | 10 | 10 kits | Distributed to United States (IL, AL, NY, TX, OK, MI, NC, CA, OH, VA, KY, CO, LA, WA, MO, AND IA) and Canada. | Terminated | July 29, 2010 | 122 | 199455425
Z-1747-2010 | 55387 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage 2.2 workstations with software version 22_01 and 22_02 |  |  | February 09, 2010 | 2010 | Ge Healthcare, Llc | GE Healthcare has recently become aware of inaccurate Scale factor annotation provided by SenoAdvantage 2.2 workstations that may impact patient safety.When using the Print Image or Print Screen feature on SenoAdvantage 2.2, the Scale factor annotation printed on hard copy images is inaccurate. Secondary Capture (SCPT) images displayed on the screen and Secondary Capture hard copy images may a | Software design | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated February 8, 2010 to customers. The letter described the Safety Issue, Affected product details, Safety Instructions to prevent further inaccurate scale factor annotation, Product Correction and Contact Information. GE Healthcare will correct all affected workstations by providing a software upgrade and an updated Operator Manual. The GE service representative will contact customer to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800-437-1171, Japan: -0120-055-919. For other countries, please contact your local GE Healthcare Service Representative. |  | 73 | 73 devices | Worldwide distribution: USA including states of CA, FL, GA, MD, NJ, NM, NY, NC, OK, TX and countries of THAILAND, TAIWAN, SWEDEN SPAIN, SAUDI ARABIA, RUSSIAN FEDERATION, PORTUGAL, PAKISTAN, MEXICO, JAPAN, ITALY,GREECE, FRANCE, COLOMBIA, CHILE, CANADA, BELGIUM, and BAHRAIN. | Terminated | December 18, 2011 | 677 | 73
Z-1746-2010 | 55613 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Trestle Anterior Cervical Plate 1-level assembly, Part Number: 61001-026. Alphatec Spine, Carlsbad, CA 92008. |  |  | August 25, 2009 | 2009 | Alphatec Spine, Inc. | The recall was initiated after the firm discovered that the affected lot was labeled incorrectly in that the packaginglabel indicates that the part number of the lot is 61002-026, which corresponds to a two-level Trestle Anterior Cervical Plate implant. | Labeling mix-ups | Alphatec Spine initiated a recall communication via telephone beginning August 25, 2009. Customers were given an explanation of the problem and instructed to check their inventory for the affected product. The firm indicated that they would arrange to have Federal Express contact customers for pick-up of the affected product(s) and replacement product would be sent out immediately.For further information, contact Alphatec Spine at 1-800-922-1356. |  | 13 | 13 units | Nationwide Distribution -- AL, CA, FL, TN, TX and UT. | Terminated | June 07, 2010 | 286 | 13
Z-1748-2010 | 53856 | No | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | PLAC Test Reagent Kit, immuno-turbidimetric method, Part Number 90107, 100 test kit, manufactured by DiaDexus, South San Francisco, CA. |  |  | October 29, 2009 | 2009 | DiaDexus, Inc | Inaccurate results-- some samples may return with falsely elevated results potentially as a result of an interfering substance. Some patients may be subjected to additional lifestyle changes or intensified pharmaceutical therapy; such as statins, as a result of the falsely elevated result. | Other | Notification letter issued to consignees on 10/29/2009. Customers were instructed to not report PLAC test results for samples that are greater than 360 ng/mL, and pursue confirmatory testing by contacting DiaDexus. Customers with questions please call the company technical support line toll free at 1-877-752-2837 for assistance.UPDATE: On April 30, the firm suspended marketing of the PLAC TIA method, as the limited range no longer covered the clinical data/reference range data cited in the product insert. The firm sent out a letter notifying customers on May 10, 2010. The firm set up a sample send out program, so that customers can send samples to labs running the ELISA method PLAC test. Arrangements were made so that labs could be enabled to run a CLIA complex assay, for customer credit programs for un-used products, and a fax back response form for response tracking. |  |  | all lots | Product was distributed to 123 consignees throughout the US. | Terminated | January 26, 2011 | 454
Z-1754-2010 | 55527 | No | CAH | Breathing circuit bacterial filter. | Anesthesiology | General Hospital | 510(k) | Xenon Diagnostic Circuit Kits labeled as "Xenon Diagnostic Circuit", "Xenon Circuit", and Needleless Xenon Diagnostic Circuit" Manufactured for RAF MEDICAL, INC., (Model number's RAF2306, RAF2304, RAF2380, RAF2300, RAF2370, RAF2372, RAF2350, RAF2341, RAF2302, RAF2340 |  |  | April 16, 2010 | 2010 | Continental Medical Labs,Inc | CML, Inc. was notified by A-M Systems in a letter dated March 31, 2010 that they were recalling a component of their Xenon Kit (ViroMax viral and bacterial filter) because the filters may have a leak at the seam where the two halves of the filter are bonded together. | Release of Material/Component prior to receiving test results | The firm, A-M Systems, sent two "URGENT: Notification of Product Recall/Correction" letters one dated March 31, 2010 and the other dated April 7, 2010 to customers. The letters described the product, problem and action to be taken by customers. Also enclosed was a list of the affected products that was sold to the consignee. The customer was instructed to cease further distribution or use of any remaining products from the listed lots, immediately notify any customers who received products from the lots, provide A-M Systems with the number of items distributed and still remaining in their inventory, once the number is provided the customer should destroy or discard any remaining product in their possession from these lot numbers.Please feel free to contact Robert Thompson by phone 360-683-8300 x302 or email: rthompson@a-msystems.com with any questions, comments or concerns. |  | 5520 | 5520 | Nationwide distribution: WI | Terminated | April 05, 2012 | 720 | 5520
Z-1755-2010 | 55693 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | The Baylis LumbarCool Pain Management Kit consisting of: 1 LumbarCool Pain Management Probe, 3 LumbarCool Pain Management introducers, and 1 Pain Management Tube Kit. Model/Catalogue Number: LUK-17-150-4 |  |  | May 12, 2010 | 2010 | Baylis Medical Corp * | Name of the device reflected on the product packing sleeve is incorrect. | Error in labeling | The firm, Baylis Medical, notified US Distributor by email and sent a "URGENT Field Corrective Action Notice" dated May 12, 2010, to customer. The notifications described the product, problem and action to be taken by customer. The customers were given the follow options: Option 1: In order to prevent confusion in identification of the product before use, please immediately remove and dispose the package sleeve that covers the tray packaging of the LumbarCool Pain Management Kit.Option 2: Alternatively, you may exchange your kit from Lot LKFA160310 with a correctly labeled replacement LumbarCool Pain Management Kit by contacting Kimberly-Clark Health Care sales representative. The customers were ask to fill out the Field Corrective Action Acknowledgment and return it.Any question please call Baylis Sales Representative at 905-602-4875 and/or contact your local sales representative for any further information you may require. |  | 58 | 58 | Nationwide distribution: TX | Terminated | October 04, 2010 | 145 | 58
Z-1757-2010 | 51812 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Treadmill models TM55, TM65, and CR60 (brand names include Quinton, Burdick, and Cardiac Science). |  |  | March 30, 2009 | 2009 | Cardiac Science Corporation | 28 stress and rehab treadmills manufactured between September 2006 and January 2009 may not have received adequate electrical safety testing prior to being shipped from the factory. If these components break down in the field, a patient or operator could be exposed to excess touch currents on the surface of the product. Because safety testing may not have allowed for complete evaluation of the el | Other | On April 3, 2009 the recalling firms notified consignees by certified letter to domestic customers and via FedEx to international customers with delivery confirmation signature required. Customers were advised to remove treadmill from service until it is tested by a Field Service Engineer.For customers with questions or those that have not already scheduled a service call with Cardiac Science Field Service Engineers, please contact the firm at 1.800.426.0337, option 1. |  | 28 | 28 treadmills | Distributed domestically (including Puerto Rico) to 20 end-user accounts and 2 distributor accounts;Distributed internationally to: Canada, Mexico, Portugal, and Saudi Arabia. | Terminated | January 10, 2012 | 1016 | 28
Z-1758-2010 | 55166 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COULTER PrepPlus, Part Number: 286600 |  |  | January 21, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports of potential dilution of reagent vials loaded onto the Preppies or PrepPlus 2 systems. Affected systems have a syringe pump with part number A46748 (see illustrations below) which cause the probe to prematurely dispense Isoflow prior to arriving at the wash station.This condition only occurs with worklist panels which include both | Other | A Product Corrective Action (PCA) letter with attached PCA Response Form was sent on the week of January 25, 2010 to all customers who purchased the pipetting and diluting system: COULTER PrepPlus, or COULTER PrepPlus 2. The letter provides the customers with an explanation of the problem identified and an instruction on proper placing antibody reagent when using a worklist panel which involves both blood specimens and antibody reagents. Customers were also instructed to complete and return the PCA Response Form in addition to sharing the information contained in the PCA with their laboratory staff. If consignees had any questions they were told to contact their local Beckman Coulter Representative. Questions can be directed to the company's Customer Service office at (800) 526-7694. |  | 193 | 193 units in total | Domestically and Canada. | Terminated | April 04, 2012 | 804 | 193
Z-1766-2010 | 55701 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DURAMAX (TM) Hemodialysis Catheters, Stacked Tip 40 cm Straight Basic Kit, Catalog Number/REF 10302807, 510 (k) 080400, Sterile --- Manufactured in USA by AngioDynamics, Inc. 603 Queensbury Avenue, Queensbury, NY 12804 |  |  | January 08, 2010 | 2010 | Angiodynamics, Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test. | Other | Customers were sent URGENT MEDICAL DEVICE RECALL LETTERS (dated 1/07/10) via Certified Mail, Return Receipt Requested. Customers were instructed to return all recalled product. A return authorization number was assigned to each customer in the recall notification letter and a UPS account was provided for return freight. A reply form was provided in the recall notification letter. Questions were directed to the company sales representatives or to Customer Service at 800-772-6446. |  | 208 | 208 | Nationwide and St. Thomas, VI | Terminated | July 25, 2012 | 929 | 2950
Z-1756-2010 | 55807 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | Fancy i Color cosmetic lens, Contents: A pair of sterile soft color contact lens immersed in buffered isotonic saline solution, Size: B.C. 8.6mm DIA. 14.5mm. Dreamcon Co. LTD. Korea. |  |  | May 17, 2010 | 2010 | Kim's Trading Inc | Recall is being conducted because the recalled color cosmetic contact lenses are unapproved medical devices. | PMA | Kim's Trading Inc. verbally contacted the distributor to provide notification of the recall. All affected product was to be properly disposed.For further information, contact Kim's Trading Inc. at 1-484-895-3773. |  | 300 | 300 pairs of contact lens | All product was distributed to one distributor in the Bethlehem PA area. | Terminated | August 11, 2010 | 86 | 300
Z-1772-2010 | 55740 | No | FCL | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Boston Scientific Radial Jaw 4 Large Capacity w/Needle Biopsy Forceps |  |  | May 06, 2010 | 2010 | Boston Scientific Corporation | Clevis may detach from the coil preventing the device to open and/or close. Delays in the procedure may result. | Process control | Boston Scientific issued Recall notification letters via Federal Express Priority mail on May 6, 2010. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. In the letter, customers with questions are provided with a phone and email at the firm to contact. |  | 320 | 320 units | Nationwide; Panama. | Terminated | September 20, 2011 | 502 | 320
Z-1760-2010 | 55620 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Back-Up Meier Steerable Guidewires, Sterilized using ethylene oxide. |  |  | April 19, 2010 | 2010 | Boston Scientific Corporation | Boston Scientific Corporation is conducting a Medical Device Recall of certain lots/batches of its Back-up Meier Steerable Guidewires. Through their internal inspection process, they identified that the polytetrafluoroethylene (PTFE) coating on the gold plated distal coil of the Back-up Meier Steerable Guidewires of the identified lots/batches have the potential for PTFE delamination. Inconsist | Nonconforming Material/Component | Consignees were sent a Boston Scientific "Urgent Medical Device Recall" letter dated April 19, 2010. The letter was addressed to Risk Manger/Field Action Contact. The letter described the problem and product affected and recommended that to immediately discontinue use of and segregate recalled product. They also requested the consignee to complete and return the Account Reply Verification Tracking Form.Customers with questions are directed to the firm's Field Action Coordinator at (763-) 494-1417. |  | 885 | US 885 single units (177 boxes), OUS 630 single units (126 boxes) | AL, AZ, AR, CA, CT, FL, GA, IL, IN, IA, KS, KY, ME, MD, MA, MS, MO, NJ, NM, NY, NC, OH, OR, PA, SC, SD, TN TX, VT, VA, WA.SWITZERLAND, GERMANY, FINLAND, FRANCE, ITALY, JORDAN, POLAND, SWEDEN, CANADA.Complete consignee list not available at this time. | Terminated | December 10, 2011 | 600 | 885
Z-1775-2010 | 54643 | No | IMK | Stair-climbing wheelchair. | Physical Medicine | Physical Medicine | 510(k) | iBOT 4000 Mobility System Battery pack. |  |  | February 22, 2010 | 2010 | Independence Technology LLC | A thermal breaker in the battery pack has the potential to open due to overheating and disable one third of the cells in the affected battery pack. | Device Design | The company sent Medical Device Field Correction letters to all customers on February 22, 2010 by certified mail. |  | 192 | 192 devices | Nationwide and to the United Kingdom. | Terminated | October 01, 2010 | 221 | 192
Z-1780-2010 | 53611 | No | OCL | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | St. Jude Medical Epicor UltraWand LP Handheld Ablation Device, model UW-LP-2, Sterilized with Ethylene Oxide. Single-use. The Ultra Wand LP device is designed for use with the Epicor Cardiac Ablation System. St. Jude Medical 240 Santa Ana Court, Sunnyvale, CA 94085 USA |  |  | June 15, 2009 | 2009 | St. Jude Medical Atrial Fibrillation Division Inc | The firm states that overheating of cardiac tissue is caused by inadequate cooling. The potential exists when the UltraWand is run using off-label cooling methods (gravity saline flow instead of pump) that cooling flow to one or both cells can be decreased based on pressure applied by the physician to the device. To reflect this information to the user, the Atrial Fibrillation Division (AFD) has | Other | St. Jude Medical issued a Market Bulletin, dated 06/15/2009, to their Field Representatives who were instructed to contact customers to make them aware that changes were made to the Instructions for Use (IFU) by SJM and that the customer may notice an extra IFU included in the shipping package. All such worldwide regulatory approvals will be needed before a revised IFU can eventually be included directly within the device package (as opposed to the shipping box). Existing instructional materials including the Epicor LP System Set-up ACS hang sheet and the EpicorLP In-service slides are being revised to reflect the UltraWand LP IFU changes The forthcoming Epicor LP In-service video is also being updated to reflect these changes. An Important Medical Device Safety Information letter, dated 11/06/2009 was sent via FedEx beginning 11/10/2009 to all US customers and will also shipped to all OUS customers beginning 11/13/2009. A copy of the new IFU will be sent with the letter which is being sent to all customers that have purchased or consigned the SJM UltraWand LP product since its commercial release in July of 2008. The customer notification includes the SJM Marketing managers and director for questions on the action. |  | 1813 | 1813 (1155 US, 658 OUS) | AL, AZ, CA, CO, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI. OUS to include Australia, Canada, Japan, Korea, Singapore | Terminated | March 02, 2012 | 991 | 1813
Z-1777-2010 | 55638 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MoblVac Portable Wound Care Pump; an AC/DC powered portable aspirator that provides negative pressure for the application of wound drainage; Ohio Medical Corporation, Gurnee, IL 60031; Models 769600 (with Truseal), 759600 (refurbished), 769600 Rental, 769601 (rental), 769602 (evaluation pump with Truseal) and 769610 (sales unit with case). |  |  | May 10, 2010 | 2010 | Ohio Medical Corporation | Ohio Medical has received a report of an explosive event with a MoblVac Portable Woundcare Pump housing where parts were separated while in an enclosed nylon transport bag during use. | Other | Ohio Medical sent Urgent Medical Device Recall letters dated 5/7/10 to the direct accounts on 5/10/10, with follow-up telephone calls on 5/11/10. The accounts were informed of the one report of an explosion of the pump, and were requested to take the pumps out of service and return the pumps to Ohio Medical. The accounts were requested to complete and fax back to Ohio Medical the enclosed acknowledgement form indicating the number of pumps being returned. Any questions were directed to Anita Christ at 847-855-6270. Follow-up letters dated 5/13/10 were sent to the accounts on 5/14/10, providing instructions for the return of the pumps and bags. Shipping boxes will be sent to the accounts based on the number of pumps reported on the acknowledgment form. |  | 375 | 375 units | Nationwide and internationally to Qatar and Oman | Terminated | August 19, 2011 | 466 | 593
Z-1776-2010 | 55505 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | The Alternans Sensor Kit Assembly is a sealed kit that contains the components necessary for conducting a Microvolt Twave Alternans test. |  |  | April 29, 2010 | 2010 | Cambridge Heart, Inc. | Some kits contained a component with expiration date prior to the date of expiry on the outside of the kit. The user receives feedback of potential unreliable results at several stages during the test process, so any risk of unreliable results is low. | Process control | The Alternate Sensor Kit Assembly P/N 203270209 may contain a pouch of standard ECG electrodes with an expiry date which is earlier than the expiration date on the outside label of the kit. All customers were sent a letter dated April 23, 2010 with information that included instructions for returning the kits. Questions are directed to the company's Quality Manager and Recall Coordinator at phone number 978-654-7625. |  | 2740 | 2740 | Nationwide | Terminated | September 20, 2011 | 509 | 2740
Z-1779-2010 | 55637 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RAYHACK DRILL GUIDE ANGLED, REF 4010-0206, 1 EACH, Rx ONLY, NON-STERILE, Wright Medical EMEA Krijgsman 11, 1186 Amstelveen, The Netherlands, Wright Medical Technology, 4677 Airline Road, Arlington, TN 38002 |  |  | April 08, 2010 | 2010 | Wright Medical Technology Inc | Drill guides from 2 lots may not seat properly on the compression block. | Other | The firm, Wright Medical Technology, Inc., sent an "URGENT: Medical Device Voluntary Recall Letter" on April 8, 2010 via FedEx to customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to please examine their inventory of the products and check it against the information provided, and to immediately return any affected products to Wright as soon as possible. The customers were also instructed to contact Customer service at 800-238-7117 for return instructions and replacement inventory. In addition, the customers were to immediately complete and return by fax at 901-867-7401 the attached response confirming receipt of the notice.If you have any questions regarding this matter, please contact Debby Daurer at 800-874-5630. |  | 13 | 13 units | Worldwide: USA and The Netherlands | Terminated | May 17, 2011 | 404 | 13
Z-1647-2010 | 54884 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bard Mesh Monofilament Knitted Polypropylene; Contents: size 2" x 4", 3 units per box, Davol Inc., A Subsidiary of C.R. Bard, Inc. 100 Sockanossett Crossroad Cranston, RI 02920. |  |  | March 05, 2010 | 2010 | Ram Medical Inc | Counterfeit product mixed with authentic product. The problem was recognized by the mismatching of expiration dates and the subtle differences in packaging. | Nonconforming Material/Component | Ram Medical Inc., contacted Customers by phone beginning March 4, 2010 and issued "Urgent Device Recall" notifications beginning March 8, 2010. Further notifications were issued to expand the scope of the recall by the firm. Each notification described the affected product and asked the User to complete a Recall Acknowledgement Form and return it to the firm. All affected product should be returned to the firm.For further information, contact Ram Medical, Inc. at 1-973-633-0400 Monday through Friday, 9:00 AM to 5:00 PM EST. |  | 306 | 102 boxes of 3 units each | Nationwide Distribution -- CO, IL , NY, MO and FL. | Terminated | June 06, 2016 | 2285 | 3720
Z-1784-2010 | 55256 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Modular Foot System 3.5mm x 22mm Locking Screw M/N: MFT-021-35-22. OrthoHelix Surgical Designs, Inc. Medina, OH 44256. |  |  | March 23, 2010 | 2010 | Orthohelix Surgical Designs Inc | On 3/9/10 it was discovered that the recalling firm packaged MFT-021-35-24 Screws from lot 0964092 as MFT-021-35-22 lot 0964081. | Employee error | OrthoHelix Surgical Designs, Inc. issued an "Important Field Correction Notice" dated March 23, 2010 via FedEx overnight. They instructed their customers to inventory their stock for affected product. All affected product will be replaced by the firm.For further information, contact OrthoHelix Surgical Designs, Inc. at 1-330-247-1444. |  | 62 | 62 pieces | Nationwide Distribution -- AZ, CA, CO, FL, GA, ID, IL, MD, NC, NJ, NM, NY, OH, OK, PA, SC, TX, UT, VA and WA. | Terminated | September 10, 2012 | 902 | 62
Z-1783-2010 | 54630 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova StatStrip Glucose Test Strips, Catalog Number: 42214. Nova Biomedical Corporation. Waltham, MA 02454. |  |  | February 12, 2010 | 2010 | Nova Biomedical Corporation | Glucose strips report low glucose results | Nonconforming Material/Component | Nova Biomedical notified it's domestic Consignees by a Customer Advisory Notice by phone and follow up by email and/or fax dated February 12, 2010. The firm will replace all affected inventory at customer sites via overnight shipment of replacement Test Strips to customer site. For further information, contact your local Nova Support Department or call 1-800-545-6682. |  | 392 | 392 Cases (1800 strips each) | Worldwide Distribution -- United States (CO, MA, MN and RI), United Kingdom, Germany, France and Italy. | Terminated | July 22, 2011 | 525 | 392
Z-1782-2010 | 55497 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ReviveR AED semiautomatic external defibrillator used with the DBP-2800 Battery Pack. Manufactured by Defibtech, LLC. Seymour, CT 06483. |  |  | May 21, 2010 | 2010 | Defibtech LLC | The AED when used with an affected battery pack, may falsely detect an error condition, cancel charge and not provide therapy | Device Design | Defibtech, LLC issued an "Urgent Medical Device Safety Information and Correction" notification dated May 20, 2010 by email and hard copy via US Mail. The notification instructed Consignees on proper identification of affected product and steps to take for product replacement.For further information, contact your distributor or Defibtech Technical Support at techsupport@defibtech.com or call 1-877-453-4507 or 1-203-453-4507, 8:30 AM to 5:00 PM Monday - Friday, EST. |  |  |  | Worldwide Distribution -- United States, Canada, Australia, Jordan, Spain, Israel, UK, Poland, Italy, Netherlands, Malaysia and Ireland. | Open, Classified
Z-1785-2010 | 55572 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Passport V Monitor: Mindray, North America. Mindray DS USA, Inc. Mahwah, NJ 07430. |  |  | April 28, 2010 | 2010 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | An issue with the Passport V Monitor has been identified where the Arrhythmia Analysis, ST Analysis and Invasive Blood Pressure functions are not operational. | Software design | Mindray issued Corrective action letters directly to all customers beginning April 28, 2010 by certified mail, return receipt requested. consignees were informed of the affected product.For further information, contact Mindray DS USA, Inc. at 1-201-995-8391. |  | 41 | 24 US, 17 International | Worldwide Distribution -- United States (OH, TX, AZ, NY, MO, IL, FL and TN), Spain, Indonesia, Brazil, Italy, Columbia and British West Indies. | Terminated | June 22, 2011 | 420 | 41
Z-1788-2010 | 54276 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Triage Total Controls 5, Level 1: Five 0.25 mL tubes. Storage Temp &lt;-20C. |  |  | February 06, 2009 | 2009 | Biosite Inc | Incidence of low recovery results has exceeded post market surveillance thresholds of customer complaints in the field. | Other | The firm, Biosite Incorporated, sent an "URGENT MEDICAL DEVICE RECALL" dated February 6, 2009, to all customers. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to immediately discontinue all use and/or sale of the listed products, discard all affected product in accordance with their local regulations, complete and fax the enclosed Verification Form within 10 days to confirm their receipt of the notice, to indicate the number of kits discarded from their inventory and the product replacement will be sent to them. Contact Inverness Medical Technical Service at 1-877-441-7440 and select option 2 from the phone prompt and a representative will confirm the number of kits discarded from your inventory and arrange for replacement product to be sent to you.Should you have any questions about the information contained in this notification, please contact (877) 441-7440 or email: Recalls@biosite.com |  |  |  | Nationwide distribution | Terminated | September 01, 2011 | 937
Z-1801-2010 | 55352 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE" Monitor B850, Software version 1.0.1.13 and 1.0.2.1 |  |  | May 06, 2010 | 2010 | Ge Healthcare, Llc | GE Healthcare has become aware of a safety issue associated with the CARESCAPE" MonitorB850 when monitoring ECG. The safety issue includes the potential for a delay in treatment. | Software design | The customers were sent on May 6, 2010, a GE Healthcare "Urgent Medical Device Correction" letter dated April 22, 2010. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, Director of Biomedical Engineering. The letter listed the Safety Issue, Affected Product Details, Product Correction, Safety Instructions and Contact Information. The customers were instructed to immediately discontinue use of the "Update Lead Set" function until the device software has been upgraded.If you have any questions regarding this medical device correction or identification of affected items, please contact Technical Support at 1-800-558-7044, or your local service representative. |  | 798 | 798 units. (229 USA, 569 OUS) | Worldwide distribution: USA including CO, MA, NC, WI, and countries of UNITED KINGDOM, SWITZERLAND, SWEDEN, SPAIN, SINGAPORE, PORTUGAL, NORWAY NETHERLANDS, ITALY, INDIA, GERMANY, FRANCE, FINLAND, BELGIUM, and AUSTRALIA. | Terminated | June 05, 2012 | 761 | 798
Z-1798-2010 | 54652 | No | JPI | Hematocrit measuring device. | Hematology | Hematology | 510(K) Exempt | i-STAT E3+ cartridges |  |  | February 05, 2010 | 2010 | Abbott Point of Care Inc. | Abbott has determined that there is a dependence of the i-STAT Hct assay on the vertical pitch angle of the i-STAT Analyzer during the blood analysis cycle. | Other | The firm, Abbott Point of Care, sent an "Urgent Field Safety Notice Product Recall" letter dated January 2010 to the customers. Safety Alert /Product Recall letters were also sent by Federal Express commencing on February 5, 2010. The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to replace the i-STAT Hematocrit/HCT and Calculated Hemoglobin/HB /Cartridge and Test Information (CT) Sheet found in their i-STAT System manual with those included in the mailing , and to provide a copy of the letter another facility, if they have forwarded any cartridges. If you should have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020 or your Abbott Point of Care representative. |  | 40446 | 40,446 total cartridges, all varieties | Worldwide distribution: USA and Argentina; Austria; Bahamas; Bolivia; Brazil; Chile; Hong Kong; China; China; Columbia; Costa Rica; Denmark; Dominican Republic; Ecuador; El Salvador; Finland; Georgia; Greece; Honduras; India; India; Ireland; Israel; Italy; Japan; Kuwait; Mauritius; Mexico; Norway; Oman; Panama; Paraguay; Peru; Poland; Portugal; Qatar; Russian Federation; Saudi Arabia; Singapore; South Africa; South Korea; Spain; Switzerland; Taiwan; Thailand; Trinidad and Tobago; Turkey; United Arab Emirates; Uruguay; Yemen; Netherlands Antilles; Venezuela; Australia; Belgium; France; Germany; Italy; New Zealand; Spain; Sweden; and The Netherlands. | Terminated | July 25, 2012 | 901 | 40446
Z-1786-2010 | 55404 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | directCHECK Whole Blood Control; |  |  | April 15, 2010 | 2010 | International Technidyne Corporation | Indication not cleared or approved. The acceptable performance range published in the directCHECK Microcoagulation Quality Control package insert is incorrect. | Error in labeling | Notification letters were sent vis Federal Express on April 15, 2010 to all distributors and end users to return unused product or to replace the package insert. Please direct questions ITC Technical Support at phone 800-631-5945 Ext. 4707. |  | 94 | 94 boxes | Distributors and medical facilities in AZ, CA, IL, KY, MA, MD, ME, MO, PA, WA. Foreign distribution to Australia, Canada, Switzerland, Germany, Spain, Quatar, and Saudi Arabia. | Terminated | September 23, 2010 | 161 | 94
Z-1800-2010 | 55674 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | InviroSnapÂż Retractable Safety Syringes, Rx Only, Single use, Sterile, Inviro Medical Devices, Inc., Duluth, GA 30096 USA. |  |  | May 10, 2010 | 2010 | Inviro Medical Devices, Inc | Syringe Defects. Group 1 syringes: The plunger pull out force was below specification. Group 2 syringes: The needle length is out of specification. | Other | Consignees were notified by letter on or about 05/10/2010. Consignees were instructed to review their inventory of InviroSnap Safety Syringes for the lot numbers and remove them from inventory and to return the affected product to Inviro Medical. They were instructed to pass on the information to anyone who uses or orders Inviro Snap Safety Syringes at their facility and to ensure a copy of the letter is provided to any other organizations to which affected devices have been transferred. A Response Card was included to be completed and faxed back to Inviro Medical Devices even if there is no product in inventory to verify receipt of notification. Questions regarding the recall were directed to the company's Customer Service number at (866) 914-2259. All product from Group 1 will be scrapped. Product from Group 2 will be retained in Quarantine until such time as the product can be exported the permission of the importing countries Ministry of Health. |  | 623000 | 623,000 units | CA, FL, IL, MI, MO, MN, NC, OH, TX; and Canada | Terminated | December 08, 2011 | 577 | 623000
Z-1787-2010 | 55697 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | Difco (TM) Decarboxylase Base Moeller, packaged in 500 gm bottles, labeled in part ***Becton Dickinson Microbiology Systems, Becton Dickinson, Sparks, MD 21152*** |  |  | December 09, 2009 | 2009 | Becton Dickinson &amp; Co. | Diagnostic test reagents may not meet quality control specifications. | Other | The recalling firm notified distributors by letter flagged as "Urgent Product Recall" on 12/09/09. The letter instructed distributors to discontinue distribution and discard product. Distributors were asked to provide customer lists and the recaller subsequently notified the end users to discontinue and discard identified lots for replacement. Customers were asked to return a response form.If further assistance is needed regarding replacements, please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. |  | 209 | 209 bottles | The product was sold to distributors in the US and to affiliates internationally. A complete list is not currently available. | Terminated | June 14, 2010 | 187 | 209
Z-1799-2010 | 55658 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DT2 Mammo, Size 8 x 10 inch, Ref No. E2RN6. Agfa Corporation. Greenville, SC 29601. |  |  | April 22, 2010 | 2010 | AGFA Corp. | The RF tags on the white protection sheets of the media stacks were incorrect. | Nonconforming Material/Component | AGFA Healthcare issued an "Urgent Safety Notice" dated April 22, 2010 requesting a check of warehouse inventory and a request for a sub-recall, if necessary. All affected product should be returned to the firm for replacement.For further information, contact AGFA Healthcare at 1-877-777-2432. |  | 180 | 180 boxes | Worldwide Distribution -- United States (AL, CT, FL, KS, MA, MI, NJ, NM, NY, PA and WV), British Columbia and Guatemala. | Terminated | February 01, 2011 | 285 | 180
Z-1802-2010 | 53126 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medtronic Racer Biliary Stent System, Model number XD518YF, UPC code 00721902758164 in a box containing sterilized pouch with product, manufactured by Medtronic Vascular, Sunnyvale, CA. |  |  | July 30, 2009 | 2009 | Medtronic CardioVascular | Mislabeling-- The incorrect Instructions for use may have been included with this product. | Mixed-up of materials/components | The firm, Medtronic, Inc., sent an "URGENT- VOLUNTARY MEDICAL DEVICE RECALL" dated July 30, 2009, to custumers (Sales Representatives, to health care provider, patient, and risk manager, hospital administrator). The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to immediately segregate any Racer Biliary units from model number XD518YF (L#0000705315) in their possession to prevent inadvertent use. Medtronic, Inc., representative will contact the customer to review the labeling and retrieve the device should it be in their inventory. The customers are being ask to cooperate by providing the representative with access to the product.If you have any questions, please contact 707-566-1548. |  | 13 | 13 | Worldwide distribution: USA and Greater China. | Terminated | January 31, 2011 | 550 | 13
Z-1803-2010 | 55624 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | The SwitchPoint Infinity 2 (SPI2) and SwitchPoint Element (SPE), an integrated voice, video, data router and teleconferencing interface for the operating room. |  |  | April 30, 2010 | 2010 | Stryker Communications Corp | Users may experience issues with the functionality of the SwitchPoint Infinity 2 and SwitchPoint Element that may be exhibited as a "Page Not Found" error message, missing drop down options, or a frozen touch panel. | Component design/selection | The firm, Stryker, sent an "URGENT: Medical Device Correction" letter dated May 4, 2010, to all customers. The letter describes the product, problem and action to be taken by the customers. The firm recommended that the customer continue to use their SwitchPoint Infinity 2 or SwitchPoint Element. The customer is instructed to refer to their user manual, in the event they experience any failures; drive on module failure and/or product failure during a procedure. A Stryker representative will contact the customers facility to coordinate a service event to replace the existing component with a new component at no cost. Should the customer experience any of the aforementioned failures and require urgent technical support, please call Stryker Technical Support at 1-866-841-5663.If you have any questions regarding this letter or the implementation of this correction, please contact Shanta McKinzie at (972) 410-7319 or via email at StrykerCommunicationsRA@stryker.com. |  | 1118 | 1,118 devices | Worldwide distribution: USA and countries of Australia, EMEA, Canada, China, S. Africa, India, Japan, Latin America, Korea, Mexico, New Zealand, Taiwan, and Signapore. | Terminated | April 25, 2011 | 360 | 1118
Z-1805-2010 | 55698 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Calcium Sulfate Hemihydrate 0.5 gram kits. |  |  | May 14, 2010 | 2010 | Ace Surgical Supply Co., Inc. | Fast set 4% potassium sulfate solution may not be sterile. For customers who have used the Calcium Sulfate Hemihydrate kit without the Fast Set solution, then these potential risks identified do not apply. Risks of adverse health consequences with the Fast Set 4% potassium sulfate solution include tissue inflammation and infection. | Process control | Ace Surgical has become aware through internal review of possibility sterility issues with one of the components included in the Calcium Hemihydrate 0.5 and 1.0 gram kits.Customers with questions should contact ACE Surgical using the recall hot line at (800) 441-3100 Ext. 204. |  | 4133 | 4133 | Class 2 recall - Worldwide distribution . Item Number 505-2001: South Africa, Australia, Vietnam, Ontario, Canada, Italy, Lebanon, Petaling Jaya, Malaysia, Portugal, and West Indies.Item Number 505-2002: Australia, Bermuda, Canada, Lebanon, Petaling Jaya, Malaysia, and Portugal. | Terminated | February 02, 2012 | 629 | 12063
Z-1804-2010 | 55438 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Medullary Tubes |  |  | April 01, 2010 | 2010 | Synthes USA (HQ), Inc. | After repeated sterilizations product may become discolored and break including while in the medullary canal. Synthes has determined that the Medullary Tube will be a single use item that is to be discarded after one sterilization cycle. | Other | The recalling firm issued an Urgent medical Device Recall letter dated 4/9/10 to all trauma hospital accounts informing them of the problem and the need to return the products. All sales consultants were notified via email and letter.Questions are directed to the company at 800-620-7025 x 5452 or 610-719-5452, or contact your Synthes Trauma Sales Consultant. |  | 18505 | 18,505 tubes | The product was shipped to medical facilities nationwide. | Terminated | March 25, 2015 | 1819 | 18505
Z-1808-2010 | 55724 | No | MXG | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Harris Medical UV Lamp; the device does not have a label identifying the name and place of business or manufacturer, packer or distributor. It is packaged with a "Harris UV Lamp Training Manual" |  |  | May 13, 2010 | 2010 | Harris Medical Resources | The Harris UV Lamp was promoted for sub-lingual treatment of HIV/AIDS; Hepatitis B; Hepatitis C; Malaria, Typhoid, Amoeba Infections and other blood diseases. The product was marketed without a 510(k), or an approved application for premarket approval (PMA) or an approved application for an investigational device exemption (IDE). Questions are directed to the firm at 773-583-7500. | PMA | Urgent Device Recall letters were sent registered, certified mail, return receipt requested, to the direct accounts on 5/13/10 - 5/18/10. The accounts were informed that Harris could make no claims as to the safety or efficacy of the Harris Medical UV Lamp in their possession and could recommend no medical indications for use because the necessary regulatory filings have not been completed and reviewed by the FDA, and that Harris Medical has no approved investigation clinical trials underway for the Harris Medical UV Lamp. The accounts were requested to check their inventory and quarantine any Harris Medical UV Lamps they may have in their possession and return them to Harris Medical using the enclosed pre-paid UPS shipping label. If the accounts further distributed the device, they were instructed to sub-recall them to the user level. The accounts were also requested to complete and return via mail or fax the enclosed return response form acknowledging the receipt and understanding of the recall letter and indicating the disposition of the devices. Any questions were directed to Thomas Perez at the company, at 773-583-7500. |  | 20 | 20 units | Florida, Hawaii, Ohio, Texas and internationally to Canada, Dominican Republic, Israel, Kenya and Peru | Terminated | December 10, 2010 | 211 | 20
Z-1807-2010 | 55410 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ulnar Shorting Saw Blade is labeled in part: "Congruent Bone Plates, SIZE Ulnar Shortening Saw Blade - Hub Style S, REF 80-0499-S, Material: Stainless Steel, STERILE, Rx Only, ACUMED, Hillsboro, OR, USA". |  |  | April 14, 2010 | 2010 | Acumed LLC | The saw blades may not cut aggressively enough for their intended application therefore creating a risk that the saw could generate excessive heat, which may contribute to bone necrosis and/or non-union. | Other | On April 14, 2010, Acumed began contacting consignees by telephone. On April 15, 2010, Acumed sent out the URGENT NOTICE: DEVICE RECALL Initial Notification letter and Recalled Product Status Form to their consignees. The letter describes the recall product as Ulnar Shorting Saw Blade-Hub Style S, Part no. 80-0499-S and Lot Codes. The Initial Notification letter instructs consignees to return the product back to Acumed immediately and notify their sub-accounts (customers) if products have been further distributed. The consignees also are asked to fill out and send back the Recalled Product Status Form .Customers should call the Customer Service at 1-888-627-9957 or e-mail to RecallNotification@acumed.net for any questions regarding the recall. |  | 718 | 718 units | Worldwide distribution: USA and to the following countries: Spain, Great Britain, South Africa, Puerto Rico, Canada, Turkey, Ireland, Finland, Belgium, Australia, Germany, and Sweden. | Terminated | June 28, 2010 | 75 | 718
Z-1811-2010 | 55649 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CS100 Intra-Aortic Balloon Pump. |  |  | May 10, 2010 | 2010 | Datascope Corporation | Certain MAQUET/Datascope Intra-Aortic Balloon Pumps (IABP) have the potential to experience display related issues. Although the IABP continues to deliver therapy to the patient, display related issues may cause the user to be unable to view the IABP on the information screen. | Other | Medical Device Recall Letters were sent to Hospital Administrators beginning May 10, 2010 by overnight mail.Customers with questions are directed to contact your local Sales/Service Representative or the company representative at 973-244-6314. |  | 161 | 161 total devices | Class 2 - Worldwide distribution -- US and Distribution to Australia, Belarus, Belgium, Canada, Egypt, England, Germany, Hong Kong, India, Iran, Israel, Japan, Kuwait, Latvia, Lebanon, Serbia, Singapore, Spain, Switzerland, Turkey, Venezuela. | Terminated | October 27, 2010 | 170 | 161
Z-1810-2010 | 55603 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Cardiac Science Powerheart 9390A automated external defibrillator. This fully automatic model does not require the user to press a shock button in order for the device to deliver therapy (when appropriate). |  |  | December 18, 2009 | 2009 | Cardiac Science Corporation | The K302 relay switch of the G3 series AED product line may experience an early failure. If the K302 contacts are open at the time of rescue, there is a possible delay in therapy or failure to deliver therapy. | Other | Customer notification letters were mailed to US end users and distributors (final consignees) between December 18 and 23, 2009 by USPS certified mail, return receipt requested.Notification letters were sent by FedEx to Canadian distributors (final consignees) on December 23, 2009. Secondary consignee (distributor) letters were mailed on December 23, 2009.Letters advise user to check the status indicator on front of AED to determine if it is rescue ready. Users are advised to keep the AED in service and to perform the Daily and Monthly Scheduled Maintenance as outlined in the Operator and Service Manual.Questions about this communication should be addressed to 1.877.901.1788 toll free from inside the US, or to email |  | 332 | 332 units | Domestic: 323 units to distributors and end users.Canada: 9 units to 2 distributors. | Terminated | December 08, 2011 | 720 | 332
Z-1816-2010 | 55535 | No | GHC | Erythrocyte sedimentation rate test. | Hematology | Hematology | 510(K) Exempt | Excyte Slider Barcode Tube Adapter used with Excyte Vacuum or Non-Vacuum ESR Tubes. Single Use. Product Number: EX-I0552. Manufactured by Vital Diagnostics S.R.L., Italy. |  |  | April 12, 2010 | 2010 | Vital Diagnostics, Inc. | The tube may fall out or separate from the Slider adapter | Component change control | Vital Diagnostics issued recall letter dated April 14, 2010 via Certified Mail in an envelop marked "Urgent Recall Information Inside", requesting accounts to discontinue use and destroy recalled product. A Certificate of Destruction and Refund Form was included to direct customers. A second Distributor Notice was sent by email to the distributors with Important Recall Information in the title. This was accompanied with the same Technical Bulletin, Certificate of Destruction and Refund Form.For further information, contact Vital Diagnostics at 1-800-345-2822. |  | 4723 | 4,723 units | Distributed Nationwide and Canada. | Terminated | September 13, 2010 | 154 | 4723
Z-1815-2010 | 55051 | No | LRJ | General purpose disinfectants. | General Hospital | General Hospital | 510(K) Exempt | CaviBleach Wipes; a disinfectant medical device, active ingredient Sodium hypochlorite 0.525%; a single 9" x 9" towelette, pre-moistened with 0.525% - 0.656% sodium hypochlorite, packaged in a flexible poly foil pouch, 50 wipes per dispenser box, 8 boxes per case; Manufactured for Metrex Research Corporation, 28210 Wick Road, Romulus, MI 48174. |  |  | December 02, 2009 | 2009 | Medtrol, Inc. | The disinfectant wipes were found out of specifications for the disinfectant activity prior to the expiration date. | Process control | Medtrol telephoned Metrex on December 2, 2009, informing them that Medtrol was requesting the recall of the affected product.Metrex sent Urgent Medical Device Recall letters dated December 3, 2009 to their accounts, requesting them to examine their inventory for the affected product, and return any found for credit. They also requested their accounts to sub-recall the product to the retail level. Metrex also requested their accounts to complete and return the enclosed reply sheet, acknowledging receipt of the recall letter and indicating the amount of product being returned. For further information, contact Metrex Customer Care Center at 1-800-841-1428. |  | 150 | 150 cases | Nationwide Distribution -- Including state of Michigan. | Terminated | June 23, 2010 | 203 | 150
Z-1814-2010 | 55405 | No | DPZ | Ear oximeter. | Cardiovascular | Anesthesiology | 510(k) | DATEX-OHMEDA TRUSAT PULSE OXIMETER AND ACCESSORIES, MODEL 3500 |  |  | April 29, 2010 | 2010 | Ge Healthcare, Llc | GE Healthcare has recently become aware that theTruSat Pulse Oximeter with external power supply doesnot comply with current safety standards. It is possiblethat the external power supply could leak electricalcurrent. This current could cause third degree burns oran abnormal heart rhythm that, if untreated, could leadto death. | Component design/selection | Consignees were sent on 4/29/10 a GE Healthcare "Urgent Medical Device Correction" letter dated 4/22/10. The letter was addressed to Healthcare Administrator/Risk Manager, Chief of Nursing, Director of Biomedical Engineering, and Home Healthcare Provider. The letter included the Safety Issue, Affected Product Details, Product Correction, Safety Instructions and Contact Information. |  | 28442 | 28,442 units | Worldwide distribution: USA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BANGLADESH,BELGIUM, BOLIVIA, BRAZIL, CANADA, INDIAN OCEAN TERRITORY, CHILE COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON LIBYAN ARAB JAMAHIRYA, LITHUANIA, MACEDONIA, MALAYSIA, MALTA, MEXICO, MOZAMBIQUE, NEPAL, NETHERLANDS, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAK IA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIETNAM, ZAMBIA. | Terminated | December 03, 2013 | 1314 | 28442
Z-1890-2010 | 55676 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | TAUT ADAPt Balloon Open Access Port And Syringe 10/12mm x 100mm, REF 41244, Rx Only, Sterile, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | May 12, 2010 | 2010 | Teleflex Medical | Balloon does not hold air or water once inflated. | Other | The firm, Teleflex Medical, sent an "Urgent Medical Device Recall" letter dated May 10, 2010 to customers The letter described the product, problem and action to be taken by the customers. They were instructed to return all of the affected product to Teleflex. Each distributor was instructed to forward the letter to their customers to retrieve relevant product from those locations. A second mailing to non-responding customers will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers will be made approximately four to six weeks after the second mailing or as an alternative, they may be contacted by phone, email, or fax. A record of these contacts will be maintained. Affected products received by Teleflex Medical will be destroyed upon receipt. Records will be maintained accordingly. Teleflex Medical will communicate to FDA on the status of recall activities.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 366 | 366 eaches | Worldwide distribution: USA including states of AR, CO, FL, GA, IL, IN, KY, MI, MN, MO, NC, NY, OR, PA, WI, WV and country of France. | Terminated | October 27, 2010 | 168 | 366
Z-1891-2010 | 52206 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE 2000/2500 Chemiluminescent Substrate Module |  |  | May 15, 2009 | 2009 | Siemens Healthcare Diagnostics Inc. | These substrate lots have shown a tendency to introduce a shift in both control and patient results with multiple methods. | Device Design | An Urgent Field Safety Notice dated May 2009 was sent to consignees of record for the previously mentioned lots. The Notice advised the following:These substrate lots have shown a tendency to introduce a shift in both control and patient results with multiple methods. ActionsIf you are using one of these lots with your instrument, please take the following actions:1. Upon receipt of a non-affected substrate lot, discontinue using IMMULITE 2000/2500 Chemiluminescent Substrate. Prior to receiving a non-affected substrate lot, you can continue to use the affected substrate if you follow the Interim Procedure.2. Discard all bottles of the affected Substrate lots remaining in your inventory once non-affected Substrate lots are available.3. The first time ONLY that you switch from one of the affected substrate lots in Table 1 to a non-affected lot, you should follow the procedure and steps below during instrument Startup:a. Empty the substrate reservoir completely before adding the new lot.b. Prime the instrument to cleanse the lines of the previous lot of substrate.c. Add the new lot of substrate to the reservoir and prime the instrument again.*d. Adjust all methods run on that instrument.*Note: Prime the substrate line as instructed on Page 83, Priming the Substrate Probe , of the IMMULITE 2000/2500 Operators Manual Rev. A.4. Discuss the contents of this letter with your laboratory director.5. Complete the Completion Notification Form that accompanies this notice and return it as directed on the form. Submitting this form indicates that you have received the information in the Urgent Field Safety Notice and have either discarded all lots of the affected IMMULITE 2000/2500 Substrate, or are following the Interim Procedure while awaiting the receipt of a non-affected Substrate lot and will discard the affected lots as soon as the non-affectedlots are received in your laboratory.Interim ProcedureSiemens recognizes that you |  | 41573 | 41,573 | Nationwide Distribution | Terminated | January 18, 2011 | 613 | 41573
Z-1899-2010 | 55597 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology Hemo Monitoring System |  |  | January 27, 2010 | 2010 | McKesson Provider Technologies - Medical Imaging Group | McKesson has identified a potential hardware configuration problem in some Horizon Cardiology Hemo systems which may cause delay and/or loss of the patient's physiological parameters on the Horizon Cardiology Hemo screen and on the patient record. | Software design | The firm, McKesson Israel Ltd., sent notification letter dated January 27, 2010 to customers. The letter describes the product, problem and action to be taken by customers and the firm. McKesson's will be contacting the customer to coordinate replacement of the sound component of the Horizon Cardiology Hemo system and will provide the additional sound adapter for all affected systems at your facility at not cost. The customers were instructed to pleasel immediately conduct the following instructions to disable sound operations-1) Open McKesson Management Console and locate the Windsurfer&gt;Monitor Settings; 2) Uncheck "Enable QRS beep" in the sounds section; 3) Locate Windsufer&gt;Hardware Settings in McKesson Management Console, and 4) Uncheck "Enable QRS beep" in the Analog Output, QRS Beep Section.If you have any questions, please do not hesitate to contact Ronen Gans directly at (973) 355-9900. |  | 22 | 22 devices | Nationwide distribution: CA, CT, FL, LA, NC, TN, WA, and WV. | Terminated | June 25, 2010 | 149 | 22
Z-1896-2010 | 55939 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Trilogy 100 Ventilator, Model/Catalog Number 1054260, U1054260 (Domestic). Philips Respironics, Monroeville, PA 15146. |  |  | June 10, 2010 | 2010 | Respironics, Inc. | It is possible that the battery could electrically short, resulting in a potential for fire or a thermal event. | Nonconforming Material/Component | Philips Respironics has contacted all of its consignees by phone and instructed them to remove the potentially affected external battery packs from their inventory and any that may be on their patients/end users. These consignees have been instructed to quarantine batteries they have removed away from flammable materials. Product return/replacement instructions were provided.For further information, contact Philips Respironics at 1-877-387-3311. |  | 23 | 23 units | Worldwide Distribution -- United States (MT, IL, IA, NY, NC, TX, CA and FL), France, South Korea and Turkey. | Terminated | November 18, 2010 | 161 | 23
Z-1900-2010 | 55673 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | GENESIS II TIBIAL DRILL GUIDE, REF 71440238, SIZE 5, QTY: (1), NON-STERILE, Smith &amp; Nephew, Inc., Memphis, TN 38116, USA |  |  | April 20, 2010 | 2010 | Smith &amp; Nephew Inc | Markings for the drill guides are on the incorrect side: left is printed on right side and right is printed on the left side | Process control | All affected Smith &amp; Nephew consignees (International distributors) were notified of problem and the recall via email on 04/20/2010 and an Urgent: Medical Device Recall letter, dated 4/16/2010. The letter stated the reason for the recall and distributors were requested to immediately cease distribution and use of the affected products. They were also asked to examine their stock and stock of their accounts and quarantine any affected product. The Verification Section of the letter should be completed and returned even if there is no affected stock. Anyone the product was distributed to should also be notified and also return the product. Questions should be directed to Kelly Knight at 901-399-5441. |  | 12 | 12 units | International Distribution Only -- Belgium, Spain, Germany, Australia, and Brazil. | Terminated | December 07, 2011 | 596 | 12
Z-1897-2010 | 55857 | No | HLF | Ophthalmic contact lens radius measuring device. | Ophthalmic | Ophthalmic | 510(K) Exempt | Lenstar LS 900, Biomicroscope, Slit-Lamp, AC-powered. HAAG-STREIT AG, Switzerland. |  |  | May 19, 2010 | 2010 | Haag-Streit USA Inc | During the creation of the Lenstar LS 900 software, the ability to perform a manual adjustment of the retinal thickness measurement was not disabled for products targeted for US distribution. The current software installed allows the user to manually alter the retinal thickness measurement, which could modify the automatic axial eye length measurement resulting in the selection of an incorrect in | PMA | HAAG - Streit USA, Inc. issued notification letters to their customers informing them of the problem. They instructed their customer to refrain from adjusting the gates for the measurement of retinal thickness until after the system has been updated.For further information, contact HAAG - Streit USA, Inc. at 1-513-336-7255. |  | 45 | 45 units | Nationwide Distribution -- Including AL, AZ, CA, FL, ID, KY, LA, MA, MI, MN, MO, NC, NJ, NY, OH, OR, PA, SC, VA, VT, WA and WV. | Terminated | February 24, 2012 | 646 | 45
Z-1892-2010 | 55651 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging Version Lines VB20, VB30, VB35 in combination with a RIS. |  |  | April 06, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a potential malfunction and hazard to patients when using Syngo Imaging version lines VB20, VB30, VB35A in combination with a RIS that is violating DICOM Standard by creating non-unique study instance UIDs. | Device Design | The recalling firm sent a Customer Safety Advisory Notice, dated 04-06-2010, to customers explaining when the malfunction occurs and the potential risks, steps to avoid potential risk of the issue by ensuring that the RIS works according to DICOM Standard and does not create double study instance UIDs. Siemens Service Reps will contact the consignees in a timely manner to check if there are symptoms for the reported issues. The issue will be resolved with a Software-Update for each affected version, except for version VB20. All customers on VB20 will be upgraded. |  | 3 | 3 units | Nationwide Distribution -- Including CT, NM, &amp; WI. | Terminated | July 05, 2012 | 821 | 3
Z-1817-2010 | 55784 | No | JCT | Tracheal prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ultraflex Tracheobronchial Stent System. Uncovered - 8/20 Catalog Number: M00568920. Single Use, Rx only. Boston Scientific, Natick, MA 01760. |  |  | May 17, 2010 | 2010 | Boston Scientific Corporation | Suture related device failures during deployment and repositioning of the stent | Device Design | Boston Scientific issued an "Urgent Medical Device Recall - Immediate Action Required" notification dated May 17, 2010 to U.S. Distributors stating that use and distribution of products should cease immediately and return all affected product to the firm for replacement. Distributors were asked to complete and return reply a verification tracking form.For further information, contact Boston Scientific at 1-866-868-4004. |  |  |  | Worldwide Distribution - Only US distributed product is affected by recall | Terminated | October 20, 2011 | 521
Z-1901-2010 | 51498 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT II Hemoglobin Testing System with CDM 3.5 Software Version 3.5; Model Number: 270-2000; Product is manufactured and distributed by Bio-Rad Laboratories, Diagnostics Group, Hercules, CA. |  |  | January 22, 2009 | 2009 | Bio-Rad Laboratories Inc | Sample test results for the first run appears to be correct but all subsequent results during that run were assigned incorrect barcode and tube position. | Software design | Bio-Rad Laboratories, Inc. issued and "Urgent: Medical Device Correction" dated January 22, 2009 to all affected consignees by FED EX, followed up by telephone calls for those not responding. Consignees were instructed on proper steps to take in dealing with the affected product. Further communication will follow to provide a solution.For further information, contact Bio-Rad Laboratories at 1-510-724-7000. |  |  |  | Worldwide Distribution -- United States, Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Czech Republic, France, Germany, Hong Kong, Hungary, India, Italy, Mexico, New Zealand, Poland, Singapore, South Korea, Sweden, Taiwan, Thailand and the United Kingdom. | Terminated | March 22, 2011 | 789
Z-1913-2010 | 55765 | No | OCK |  |  | Gastroenterology/Urology | PMA | FemSoft Insert, Female Urethral Insert, Sterile (R), Rx Only. Assorted Sizes Multi-Pack, REF 70200. |  |  | April 27, 2010 | 2010 | Rochester Medical Corp | Rochester Medical is voluntarily recalling specific lots of FemSoft urethral inserts manufactured and distributed in 2009 and 2010. Testing has revealed a small number of packages which contain a small void in the package seal, and it has been determined that product sterility cannot be guaranteed. If product sterility has been compromised, there is a potential for increased risk of urinary trac | Packaging | Consignees were contacted by telephone on 5/3/10 and if not reached, were sent a Rochester Medical "Urgent Medical Device Recall" letter on 5/6/10 dated April 26, 2010. The letter was addressed to Rochester Medical Corporation Customers. The letter described the problem and product involved; recommended consignees to record the quantity and lot numbers (s) on the enclosed postcard and return it to them; and to immediately destroy the product with the lots affected. In the case of distributors, the telephone call was followed by a letter sent/dated April 27, 2010 to each distributor. It advised distributors to isolate any remaining inventory with the specified lot numbers and to complete the disposition report. Letters included a customer letter for dissemination to their retail level customers. Retail customers were asked to segregate and destroy the affected products and record the action. The company will replace the destroyed product. The distributors started their contact to consignees on 4/28/10. Please contact Rochester Medical at 1-800-243-3315, or the firm's Customer Service Department at 1-507-533-9309 for additional information or for product return &amp; replacement instructions. |  | 15400 | 15,400 devices | AL, AK, AZ, CA, CO, DC, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, WY, WI,.AUSTRIA, BANGLADESH, CANADA, CYPRUS, ETHIOPIA, FRANCE, GERMANY, GREECE, INDIA, JAPAN, KOREA, NETHERLANDS, P.R. CHINA, SAUDI ARABIA, SCOTLAND, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TURKEY, UAE, UNITED KINGDOM. | Terminated | December 28, 2011 | 610 | 105980
Z-1905-2010 | 55495 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Unicel DxH 800 Coulter Cellular Analysis System, |  |  | July 28, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed an issue with the UniCel DxH 800 during in-house testing. It wasdetermined that there is a potential for a false increase in platelet results, due to micro-bubbles. This situation could present a risk of falsely increased platelet results in patient samples as diluent containers are depleted. | Other | Beginning on July 29, 2009 Customers were contacted by Customer Technical Support (CTS) personnel and provided with the Product Corrective Action (PCA) letter. The PCA letter informed customers of the issue, the impact of the issue and actions/resolutions required by Beckman Coulter Service representative ad the affected consignee.The consignees were informed that their Beckman Coulter Service representative has made temporary changes to the Supplies setup. In addition to the Supplies setup change, consignees were instructed that one of the following workflow options must be implemented to ensure no results are affected. Please continue to follow one of the options listed below until a permanent fix is available. These options were given as: Option A: Monitor Diluent cycles or Option B: Unable to monitor Diluent cycles. The consignees were instructed that a long term solution to eliminate this risk is being developed and will be provided as soon as available. The consignees were instructed to please share this information with your laboratory staff. If they have any questions regarding this Product Corrective Action, please call 800-526-7694 in the United States or contact your local Beckman Coulter Representative. |  | 7 | 7 | Nationwide to: AZ, LA, IL, NY, PA &amp; TN | Terminated | April 04, 2012 | 981 | 7
Z-1904-2010 | 55953 | No | IZO | Diagnostic x-ray high voltage generator. | Radiology | Radiology | 510(K) Exempt | Hydradjust IV DR Urological Table Sedecal Generator (p/n 750732 - 64W generator &amp; p/n 750733 80W generator). Mallinckrodt, Inc. Cincinnati, OH 45237. |  |  | June 14, 2010 | 2010 | Mallinckrodt Inc | Covidien received complaints of the Sedecal Generator failing and producing smoke at customer sites. The failure modes are located in different parts of the generator and will lead to the generator becoming inoperative. | Process control | Covidien Issued an "Urgent Device Correction" notification dated June 14, 2010 informing their customers of the potential problem. Consignees were instructed that Field Service Engineers will be in contact arrange a service visit to perform corrections on the affected units.For further information, contact Covidien Product Monitoring at 1-800-778-7898. |  | 144 | 144 units | Worldwide Distribution -- United States, Canada and China. | Terminated | February 25, 2012 | 621 | 144
Z-1914-2010 | 53698 | No | CIB | Angiotensin I and renin test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GammaCoat Plasma Renin Activity 125 I RIA Kit, REF: CA-1533 (100 test kit ), CA-1553 ( 500 test kit), DiaSorin, Stillwater, Minnesota 55082-0285. |  |  | October 12, 2009 | 2009 | Diasorin Inc. | The kit control provided with the GammaCoat Plasma Renin Activity kit was labeled with the incorrect range. | Nonconforming Material/Component | Consignees were Emailed or faxed a "DiaSorin Customer Notification" letter dated October 21, 2009. The letter was addressed to "Dear Valued Customer". The letter described the Product name and the problem and provided the correct range for the Renin Activity Control provided with the kit. The letter also stated that replacement product for any previous failed runs. Customers are to complete the bottom portion of the letter and return it to the firm. Questions or concerns should be directed to DiaSorin Inc., Technical Services at 1-800-328-1482 or 651-439-9710. |  | 289 | 289 kits | Worldwide Distribution -- USA, including states of CA, CT, IN, LA, MI, NC, OH, TN, UT, and WA and countries of CANADA, BELGIUM, BRAZIL, INDIA, NORWAY, PORTUGAL, SPAIN, and SWEDEN. | Terminated | April 17, 2012 | 918 | 289
Z-1916-2010 | 55579 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Automated External Defibrillator(AED) of brand name "Philips HeartStart HS1 OnSite" (Model number M5066A). |  |  | April 30, 2010 | 2010 | Philips Medical Systems | Seventeen AEDs failed production line testing (Final Acceptance Test). | Nonconforming Material/Component | The firm notified U.S. distributors on 4/30/2010 via phone call and e-mail to obtain consignee location information. A website providing information was activated on 4/30/2010. On 5/3/2010 Urgent Medical Device Recall letters, dated April 29, 2010 were sent to direct customers in the US. The letters explained the reason for the recall and Customers were asked to verify that they have the product subject to recall. Replacement AEDs will be provided at no charge to the customer. Field Safety Notices were attached to the letters and provided a list of actions for customers to take. Customers should check the serial number on their devices against the list provided, contact Philips Healthcare to arrange for replacement, and ensure that necessary staff are aware of the recall. For further information and support, Philips Healthcare should be contacted at 1-800-263-3342. |  | 205 | 205 | Worldwide Distribution -- USA, Canada, Japan, United Arab Emirates, Austria, Switzerland, Germany, Spain, France, Italy, and Norway. | Terminated | December 10, 2012 | 955 | 471
Z-1922-2010 | 55552 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 3100/ 3100IQ system with InnovaIQ Table Option |  |  | June 03, 2010 | 2010 | Ge Healthcare, Llc | GE Healthcare has become aware of a potential loss of gantry and table motions affecting Innova 2100IQ, 3100, 3100IQ, 4100 and 4100IQ systems with InnovaIQ Table that may impact patient safety. | Device Design | Consignees were sent on 6/3/10 a GE Healthcare "Urgent Medical device Correction" letter dated May 12, 2010. The letter was addressed to Hospital Administrators/ Risk Mangers, Managers of Radiology/Cardiology and Radiologists/ Cardiologists. The letter included the Safety Issue, Affected Product details, Safety Instructions, Product Correction and Contact Information. Customers are to contact a service representative immediately if they notice any cable damage. |  | 13 | 13 | Worldwide Distribution -- USA, including states of AL, AZ, CA, CO, FL, GA, IL, IN, KY, MA, MI, MS, MO, MT, NE, NJ, NY, NC, OH, OR, PA, TN, TX, VT, WA, WV, and WI and countries of UNITED KINGDOM, SWEDEN, REPUBLIC OF KOREA, NORWAY, NEW ZEALAND, ITALY, ISRAEL, IRELAND, INDIA, HONG KONG, GERMANY, FRANCE, FINLAND, CZECH REPUBLIC, and AUSTRALIA. | Terminated | September 28, 2015 | 1943 | 92
Z-1919-2010 | 55818 | No | HPS | Cryophthalmic unit. | Ophthalmic | Ophthalmic | 510(k) | MIRA CR4055 25mm Finger Probe |  |  | April 08, 2010 | 2010 | Mira, Inc. | Devices distributed without an approved 510(k) | PMA | Mira contacted accounts by email between 4/08/10-05/21/10 requesting return of inventory. |  | 5 | 5 units | Worldwide Distribution -- USA, including the state of NY and the countries of Israel and Italy. | Open, Classified |  |  | 5
Z-1941-2010 | 55850 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STERIS 4085, 5085 &amp; 5085SRT Surgical Tables |  |  | May 14, 2010 | 2010 | Steris Corporation | STERIS has learned that a limited quantity of manifold assemblies were manufactured by our supplier with incorrect washers and springs in the valves of the manifold. This supplier error has the potential to cause: a) an inability to articulate/move a table section when commanded by the operator; b) two different articulations occurring when a single articulation is commanded, or c) that the table | Nonconforming Material/Component | On May 14, 2010 the firm sent FCL letters to their customers notifying them of the problem. They informed their customers that a service representative will contact them to arrange correction of the unit. |  | 79 | 79 units | Worldwide distribution: USA, Australia, India, Korea &amp; New Zealand | Terminated | February 24, 2012 | 651 | 79
Z-1940-2010 | 55654 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Biosciences, BD FACS Sample Prep Assistant II (SPA II), blood sample preparer for flow cytometry analysis |  |  | February 19, 2010 | 2010 | BD Biosciences, Systems &amp; Reagents | Samples prepared on the BD SPA II, SN T0076 produced inaccurate sample results due to inaccurate sample dispensed as a result of the cap piercing probe | Nonconforming Material/Component | The firm filed an MDR and implemented a field action consisting of a February 2010 notification letter and product replacement of affected probes. A CAPA for the probe design has been implemented. Effectivity is being tracked via mail and follow up phone call. |  | 349 | 349 units | Worldwide distribution: USA, Australia, Belgium, China (Hong Kong), Colombia, Jamaica, Korea, Mexico, Singapore and Uruguay. | Terminated | May 18, 2011 | 453 | 349
Z-1943-2010 | 54366 | No | JEY | Bone plate. | Dental | Dental | 510(k) | BIOMET MICROFIXATION REF: 91-6103 "1.5MM" SYSTEM HIGH TORQUE (HT), SD, X-DR, SCREW 1.5MM X 3.5MM (WXL) TITANIUM LOT 829760, ***NON STERILE PRODUCT***. |  |  | October 05, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling 91-6103 1.5x3.5mm HT SD X-drive screw Lot Number 829760. This package may contain 91-6704 1.5x4mm ThinFlap SD X-drive screw. | Process design | Biomet Microfixation notified each consignee by email and fax on 10/05/2009. Each consignee was asked to use the included Response Form to provide product disposition. |  | 91 | 91 units | US distribution: CA, CO, MI, and MO. | Terminated | December 15, 2010 | 436 | 91
Z-1942-2010 | 55801 | No | DAK | Complement components immunological test system. | Immunology | Immunology | 510(k) | Inova Diagnostics BINDAZYME Human C1q binding Circulating Immune Complex EIA Kit, Part#: MK021 |  |  | October 28, 2009 | 2009 | Inova Diagnostics Incorporated | The recall was initiated after the binding site notified Inova Diagnostics, inc. that the affected Bindazyme Human C1q binding Circulating Immune Complex EIA Kits may contain an ELISA plate labeled as "tissue Transglutaminase Coated Wells" instead of "C1q Coated Wells". | Process control | Notification letters from INOVA Diagnostics Incorporated were mailed to each of the consignees via Certified mail on October 30, 2009.The notification letters inform the consignees of the reason for recall and provided instructions for final disposition of the recalled kits. In addition, a FAX back Attachment was attached to the recall notice for customers to respond to the recall. |  | 700 | 700 kits | Worldwide distribution: USA, Germany, Italy, Belgium, Romania, Portugal, Argentina, Spain, Sweden, and India. | Terminated | June 30, 2010 | 245 | 700
Z-1937-2010 | 55885 | No | FKG | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | C.A.P.D. Tubing Exchange Tray, HT80, Sterile, Micropore Surgical Tape, Single Use Rolls, Tri-State Hospital Supply Company Manufacturers of Centurion Healthcare Products Howell, MI. |  |  | March 16, 2010 | 2010 | Centurion Medical Products | The 3M micropore tape is under recall due to the tape not performing and needles becoming loose and dislodged. | Other | The firm sent Urgent Recall Product Notices to its two customers dated 3/16/2010. The letter explained that the tape may not adhere properly. The consignees were instructed to remove any kits containing the recalled tape from their inventory, if the tape was used to secure dialysis needles or pressure dressings. If the tape was used for other purposes, the kits could still be used. The recall letter included a reply form for customers to complete and email or fax back. Upon receipt of the form, a Centurion Representative will contact them to retrieve affected product. Customers should contact Matthew Price at 517-546-5400 ext. 1135, if further info is needed. |  | 24 | 24 kits | Nationwide Distribution -- AR and OH. | Terminated | February 09, 2011 | 330 | 3312
Z-1935-2010 | 55796 | No | IKN | Diagnostic electromyograph. | Physical Medicine | Neurology | 510(k) | Lumbar Matrix Scan, Diagnostic Electromyography System consisting of the CERSR Electromyography System and Electrode Array; the system consists of two devices. The first is a system of capital equipment referred to as CERSR Electromyography System, which includes a cart. The second device is disposable electrodes arrays, which are single use devices, used with the capital equipment that is adhered to the patient for monitoring the bioelectric signals produced by the muscles. The disposable is referred to as the Electrode Array. CERSR Electromyography System is packaged in a reusable crate maintained by SMI. The Electrode Array is packaged in a sealed foil pouch. Both devices are non-sterile and the packaging is for protection from damage during shipping. |  |  | January 29, 2010 | 2010 | SpineMatrix Inc | A full design review of the CERSR system was conducted in 7/2009. The audit revealed numerous design verification tests that either did not pass or did not meet the design input requirement. One requirement found that it was possible for the physician to think they were reviewing one patients scan, but were actually reviewing a scan from a different patient. This event did not happen; however, | Device Design | On 1/29/2010 all customers were provided a Voluntary Product Recall Notification letter that was sent via U.S. Postal Service Certified Mail and they were contacted by phone to arrange for pick up of the devices. The letter identified the affected product and described the problem. Customers are to check their inventory for the affected product and contact SpineMatrix Customer Service to receive instruction as to how to return the affected product, if the firm did not already contact the customer. Customers must complete the verification form and return it. Questions should be directed to Customer Service. |  | 1166 | CERSR - 14 units, Electrode Arrays - 1,152 units | Nationwide Distribution -- CA, OH, OR &amp; TX. | Terminated | September 19, 2012 | 964 | 1166
Z-1939-2010 | 55749 | Yes | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | Hospital Grade Power Cord, UL, Part # 66319, 055476, manufactured by Electri-Cord Manufacturing Co. beginning in February 2003. The cords are a detachable unit, supplied with the instruments listed below: |  |  | April 21, 2010 | 2010 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | The potential risks during the use of this power cord include electrical shock, delay in setup and therapy, interruption of therapy, device failure, and fires. Although Medtronic has received no reports of adverse events related to the affected cord, a failure may potentially lead to serious adverse health consequences, depending on the device, and therapy being interrupted. This Field Notificatio | Nonconforming Material/Component | Consignees were sent on 4/21/10 a Medtronic "Urgent Medical Device Recall Notice" letter dated April 19, 2010. The letter was addressed to their customers and described the problem and the product involved. They recommended to discontinue use of the AC power cords and to destroy them. Requested consignees to fill out the attached Medtronic Recall Certificate (FCA 1005), and return it to Medtronics. |  | 699 | 619 USA, 80 OUS | Worldwide distribution: USA, Latin America, Middle East, Japan, and Canada. | Terminated | December 24, 2011 | 612 | 699
Z-1903-2010 | 55853 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK 20 defibrillator/monitor. |  |  | May 27, 2010 | 2010 | Physio Control, Inc. | A failure on the power supply assembly can result in either "No DC power" or "No DC or AC power." A failure of DC (battery) power can result in a delay of defibrillation therapy if no AC (line) power is available and the device will not operate. | Material/Component Contamination | Firm began mailing Urgent Medical Device Correction notifications to affected customers on May 26, 2010 regarding the power supply to the device. All affected power supplies will be updated; this will be accomplished in two phases. Devices at a higher risk will be updated first in phase 1. The remaining lower risk devices will be updated afterward in Phase 2. Customers are advised to keep the defibrillators in service, and to follow recommended daily Operator Checklist steps, while service updates are scheduled. Customers with questions should contact the firm at 800-442-1142 between 6 am and 4 pm Pacific Time, Monday - Friday |  | 43001 | 43,001 units | Worldwide distribution. In US, distributed domestically including to gov't accounts; and internationally to the following countries: ANGUILLA, ARGENTINA, AUSTRALIA, BAHAMAS, BAHRAIN, BARBADOS, BOLIVIA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, CYPRUS, DOMINICAN REPUBLIC, EGYPT, EL SALVADOR, GERMANY, GUAM, GUATEMALA, HONG KONG, INDIA, IRAN, ISRAEL, JAMAICA, JORDAN, KUWAIT, LEBANON, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PUERTO RICO, SAUDI ARABIA, SOUTH AFRICA, SYRIA, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIRGIN ISLANDS, and YEMEN. | Terminated | October 21, 2014 | 1608 | 43001
Z-1887-2010 | 55786 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Navilyst Medical, Inc., Vaxcel(TM) Port Plastic Standard Port with 9F Polyurethane Catheter (1.8 mm/2.8 mm), REF Catalog No. 45-360, UPN Product No. M001453600, STERILE, Rx ONLY, For single use only --- Implantable Vascular Access System --- (kit contains Huber needle) -- previously distributed by Boston Scientific Corporation. |  |  | May 13, 2010 | 2010 | Navilyst Medical, Inc | FDA testing determined that non-coring needles included in the Vaxcel Port and Vaxcel Port with PASV Valve may result in coring and resulting septum damage which could lead to patient injury. | Nonconforming Material/Component | Navilyst Medical, Inc. is recalling Vaxcel Ports and Vaxcel Ports with PASV Valve with valid expiration dates worldwide. Navilyst sent "Urgent Medical Device Recall - Immediate Action Required" letter/recall notification packages on May 13, 2010, via Federal Express to 471 domestic consignees; on May 14, 2010, via e-mail to 14 international consignees and to Boston Scientific Corporation (prior owner of the affected products) via facsimile and Federal Express. On June 3, 2010, Navilyst Medical sent Recall Notification Packages on behalf of Boston Scientific Corp. via Federal Express to 182 domestic consignees. The letter instructs the customer to immediately discontinue use of and segregate recalled product, complete and return the reply verification tracking form to fax# 518-742-4450 and package and return the recalled product to the address provided.If you have any question, please call Navilyst Medical, Inc., Customer Service at 1-800-833-9973 or Michael Duerr at 1-518-742-4571. |  |  |  | Worldwide distribution:USA and countries of Belgium, Brazil, Canada, Colombia, Croatia, Germany, Greece, Hong Kong, Italy, Saudi Arabia, South Korea, Spain, Tunisia, and United Kingdom. | Terminated | August 03, 2012 | 813
Z-1898-2010 | 55670 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Cepheid Xpert MRSA/SA Blood Culture Test for use with the GeneXpert Dx System, manufactured by Cepheid, Sunnyvale, CA. |  |  | April 05, 2010 | 2010 | Cepheid | The firm is receiving reports of inaccurate results, which may result in delay of care for seriously ill patients. | Nonconforming Material/Component | The firm, Cepheid, issued a technical bulletin to customers on April 5, 2010 and an updated to the April 5 bulletin dated April 9, 2010 instructing them to perform a reflexive repeat test on any blood culture results initially determined as having SA.If further information on this please contact your Cepheid Sales Representative 1-888-336-2743 or www.CEPHEID.COM. |  | 5510 | 5,510 units | Worldwide distribution: USA and Europe, Taiwan, and Australia. | Terminated | May 25, 2011 | 415 | 5510
Z-1947-2010 | 55186 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Tabs Professional Monitor with Wire Bracket and Push Button Reset, Model 25223, Stanley Senior Technologies, Lincoln, NE. |  |  | March 26, 2010 | 2010 | Stanley Security Solutions, Inc. | Product might not indicate low battery alarms or exit alarms when powered by a DC power supply. | Component change control | The firm made phone calls beginning 3/26/10 to customers who purchased the 12-volt power supply informing them of the recall and instructing them to determine the number of units that will need to be returned for replacement. Important Safety Notice letters dated 3/25/10 were issued via regular mail on 3/26/10 and also providing the reason for the recall, instructions, and customer service contact information to arrange for return of the units. |  | 45 | 45 units | Nationwide Distribution | Terminated | September 16, 2011 | 539 | 374
Z-1952-2010 | 55776 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Artificial ligament fixation device # 7PE ZIPLOOP Extended Toggle, Sterile, REF 904755, Biomet Sports Medicine Warsaw, IN. |  |  | March 31, 2010 | 2010 | Biomet, Inc. | The firm has received complaints that the saddle is too long. The saddle is supposed to be between 7/8 and 1 1/8 inches long. The saddle in the complaints was up to 1.9 inches long. | Process control | Biomet Sports Medicine issued an "Urgent Medical Device Recall Notice" dated March 31, 2010 to user facilities and distributors. A second notification containing updated information was issued June 2, 2010. Consignees were instructed to discontinue use of the product and return it to the firm along with a completed Response Form. For further information, contact Biomet at 1-574-371-3755 or 1-574-371-3756, Monday through Friday, 8AM to 5PM. |  | 2466 | 2466 both products | Worldwide Distribution -- United States, Canada, Australia, The Netherlands, Belgium, Italy, France, New Zealand, Austria, United Kingdom, Chile, Greece, Poland, Denmark, Turkey, Portugal, South Africa and Spain. | Terminated | August 06, 2012 | 859 | 2466
Z-1951-2010 | 55771 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | OptifluxÂż F180NR Dialyzers, Catalog Number: 0500318E. Manufactured by Fresenius Medical Care North America, Ogden, UT 84404. |  |  | May 14, 2010 | 2010 | Fresenius Medical Care North America | Hemodialyzer may leak at the header/end cap | Process control | Fresenius Medical notified customers by telephone script beginning May 14, 2010 and follow up letter advising users to examine inventory for recalled lot, discontinue use and contact Fresenius Customer Care 1-800-325-5188 , ext. 6610 for instructions on how to return the recalled product. |  | 34152 | 34,152 units | Nationwide Distribution | Terminated | September 13, 2011 | 487 | 34152
Z-1954-2010 | 55894 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Stellant CT Dual Syringe Kits, Catalog Number SDS-CTP-QFT and SDS-CTP-SPK |  |  | May 24, 2010 | 2010 | Medrad Inc | Kits contain a T connector that may be susceptible to cracking or breaking or leaking. This could cause product leakage or, less likely, patient injury. | Other | The recalling firm issued an Urgent Medical Device Field Correction on 5/24/10 requesting customers to check their inventory for the affected lots and complete the fax back form with the amount of product in stock. The manufacturer plans to make arrangements to provide individually packaged low-pressure connector tubing to use with the product in the event that cracking or breaking and leaking of the T connector is observed. This requires vigilance by the customers over the life of all of the devices.For questions concerning this recall, please contact the company 1-800-633-7231, select option 2, and indicate that you are calling about the Stellant CT Dual Syringe Kit field correction. |  | 1382889 | 1,382,889 kits | nationwide and Worldwide distribution includes AU, BE, BG, BM, BR, CA, CH, CI, CL, CN, CO, CS, DE, DK, EC, EG, FI, FR, GB, GT, HN, IN, IT, JM, JP, KW, KY, MU, MX, MY, NL, OM, PA, PH, PK, PR, SE, SG, TW, YE, ZA. | Terminated | October 13, 2010 | 142 | 1382889
Z-1889-2010 | 55824 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | 22ga x 1" Right Angle Huber Needle |  |  | May 26, 2010 | 2010 | Multi-Med, Inc. | FDA sample determined the needles to be coring. | Pending | The firm, Multi Med, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated May 26, 2010 to their sole consignee Navilyst Medical, Inc., via email and overnight courier with signature notification. The letter describes the product, problem and action to be taken by customer. The customers was instructed to immediately examine their inventory and quarantine product subject to recall, if the product was distributed to others customers, identify and notify them at once and include recall notification letter,. Recall has been initiated due to FDAs sample analysis that these needles may core implanted port septums.If you have any questions, call me, Sue Starkey, at 603-357-8733 or via email at astarkey@multimedinc.com. |  | 5000 | 5,000 | Worldwide distribution | Terminated | May 23, 2017 | 2554 | 10169
Z-1949-2010 | 55825 | No | NBE |  |  | General &amp; Plastic Surgery | PMA | Coseal Surgical Sealant 4mL, Product Codes: 934071 &amp; 934074. |  |  | January 14, 2010 | 2010 | Baxter Bioscience | The recall is being conducted as a precautionary measure due to out of specification results at the 18 month and 21 month time points (out of the 24 month licensed shelf life to expiration) for the percent thiol substitution during a Coseal Stability study. The parameter that is out of specification is an indicator of possible failure for the product to gel appropriately. | Process control | The recall notification was initiated on January 14, 2010 with the firm sending a Product Recall letter and attached Customer Reply Form to the affected consignees. The recall letter informed the consignees of the reason for recall and requested them to immediately stop the usage/distribution of the lots noted above and to quarantine any units you may still have in your inventory. The letter provided instruction for product return.The letter also requested the consignees to please forward the recall information as appropriate to other services/facilities/departments that they may have provided this product to. Consignees were instructed to complete theCustomer Reply Form, and fax it as soon as possible to 888-345-4261. Consignees with questions were instructed to contact Stericycle at 877-870-4486. |  | 15262 | 15,262 | Nationwide Distribution | Terminated | December 22, 2010 | 342 | 26440
Z-1895-2010 | 55854 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira brand Symbiq Two-Channel Infuser; List Number: 16027; Distributed by: Hospira, Inc. Morgan Hill, CA and Manufactured by: Hospira Inc, La Aurora, Heredia, Costa Rica. |  |  | April 09, 2010 | 2010 | Hospira Inc | Hospira has received numerous customer complaints of the Symbiq infusion pump's failure to detect air-in-line conditions during operation. | Device Design | Hospira issued a "Urgent: Medical Device Recall Notification" dated June 11, 2010 to update a previously issued Clinical Bulletin. Customers were informed of the issue and instructed on proper mitigation steps to further avoid the issue associated with the affected product. Consignees were asked to complete and return the Reply Form.For further information, contact Hospira Technical Support Operations at 1-800-332-4002, 6AM to 4PM PST, Monday through Friday. |  | 28485 | 28,485 total | Worldwide Distribution -- United States (AL, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NH, NJ, NY, OH, OK, PA, SC, SD, TN, TX, VA and WA), and Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Puerto Rico, Singapore and United Arab Emirates. | Terminated | December 03, 2013 | 1334 | 28485
Z-1956-2010 | 54371 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | BIOMET 3i GS300 GOLD SLOTTED SCREW-3mm, STERILE, Rx Only. Made in USA. Lots 868517, 865348, and 868515. |  |  | October 06, 2009 | 2009 | Biomet 3i, LLC | Biomet 3i, Palm Beach Gardens, FL is recalling their Gold Slotted Retaining Screw 3MM, Model Number GS300. Product is Out Of Specification - product was manufactured with a material not specified for use in this product. | Other | Each US customer was contacted by telephone using a prepared phone script. Following telephone contact, Biomet 3i performed a follow-up of each call with a faxed letter. International customers were first notified by e-mail then by fax letter. Each customer was instructed to check their respective inventory for the affected product and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory. They were also instructed to return the affected product to Biomet 3i.Questions or concerns are directed to Kelly Taylor at the firm at 561-776-6700. |  | 495 | 495 screws | Product was distributed to 3 distributors, 14 direct consumer/users. Product was distributed in Japan, Spain, Switzerland, Uruguay, Paraguay, Mexico, France, Netherlands, Australia, and in the US in KY, DC, IL, RI, MA, NJ and TX. | Terminated | August 30, 2010 | 328 | 495
Z-1955-2010 | 55774 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | SSH-880CV Aplio Artida System |  |  | September 16, 2009 | 2009 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) is initiating a field correction on the SSH-880CV Aplio Artida System due to a problem related to a wrong measurement value display in using with DICOM viewer.When an image is transferred from this system in DICOM transfer and a measurement is performed on a Doppler image in the DICOM viewer, measured values such as VTI (Velocity Time Integral) and PI/RI | Software design | The software will be replaced. To accomplish this, Toshiba America Medical Systems (TAMS) will notify consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS for retention.For questions or concerns, please contact your Toshiba representative or the firm at (714) 730-5000. |  | 22 | 22 | Nationwide ((CA, GA, IA, LA, MA, MD, MN, MT, PA, PR, OH, OR, TN). | Terminated | March 21, 2012 | 917 | 22
Z-1971-2010 | 55316 | No | LGC | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | Sure-Vue HSV-2 Rapid Test, Catalogue No: 23-046395. |  |  | April 14, 2010 | 2010 | Biokit U.S.A. Inc. | Negative serum was giving a positive result Herpes Simplex Type 2 (HSV-2). It is possible that misdiagnosis (i.e., positive) could be made. | Nonconforming Material/Component | Biokit USA notified US Customers by fax on 2/18/10 of the reason for the kit recall and requested to discard any remaining units. Accounts were advised to retest patients or review test results. Questions are directed to the firm at 800-926-3353. |  | 8 | 8 kits | US (CA, FL, GA, IN, KY, MA, NY, VA) and UK | Terminated | January 26, 2012 | 652 | 200
Z-1960-2010 | 55710 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker 10 MM AVS TL Trial Spacer; Catalog number: 48389410. |  |  | May 13, 2010 | 2010 | Stryker Spine | Stryker Spine has received reports of breakage of the distal end of the AVS TL Spacer Trial during the trialling step of the AVS TL procedure. | Process control | Urgent Product Recall letters, dated May 13, 2010 were sent via Federal Express to all direct accounts (branches and hospitals) with a separate letter to surgeons. The letter identified the affected product and also described the issue, potential hazards, and risk mitigations. The letter also asked customers to quarantine any affected product and return to Stryker Spine. A Customer Response Form was provided and customers should contact their Stryker Spine Sales Rep for additional information on replacement product. Questions should be directed to Tiffani Rogers on 201-760-8206. |  | 193 | 193 units | Nationwide Distribution. | Terminated | September 06, 2012 | 847 | 2110
Z-1969-2010 | 55721 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Starion Instruments TLS3 35C Thermal Ligating Shears, Electrosurgical cutting and coagulation device and accessories, manufactured by Starion Instruments Corp, Sunnyvale, CA. |  |  | February 06, 2008 | 2008 | Starion Instruments | Some jaws exhibited lack of ductility and could break if overloaded or used improperly by placing too much tissue in the bottom of the "v" of the jaw. | Nonconforming Material/Component | The firm called domestic consignees but no letter was sent to domestic cosignees and a notification letter was sent to international consignees.The firm called domestic consignees and sent a notification letter to international consignees. No records were kept of phone calls and a copy of the letter sent to international consignees was not obtained |  |  |  | Worldwide distribution | Terminated | November 27, 2010 | 1025
Z-1972-2010 | 55750 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Starion Instruments Real Hand ThermaSeal Thermal Ligating Shears, electrosurgical cutting and coagulation device and accessories, manufactured by Starion Instruments, Sunnyvale, CA |  |  | June 26, 2008 | 2008 | Starion Instruments | Due to an improper fixture revision, device may be missing a screw at the jaw assembly, resulting in jaw components falling off during use. | Process control | The firm submitted an MDR but determined that there was no risk to health. E-mails were sent on June 26, 2008, to consignees, describing the product, the problem, and action to be taken by consignees. The consignees were ask to provide a list of all unused ThermaSeal devices and their locations, to return the products to Starion from the lots noted and to provide a copy of the attached letter to their sales force and customers if they choose.If you have any questions, please call (408) 522-5200 ext. 326 or (408) 350-9934. |  | 217 | 217 units | Worldwide distribution. | Terminated | October 27, 2010 | 853 | 217
Z-1973-2010 | 55419 | No | HXS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | TORQ Sternal Closure Device, SINGLE-USE |  |  | March 31, 2010 | 2010 | KARDIUM INC. | A component of the device is installed incorrectly which can result in the component falling out of the device and into the patient during use. | Other | The firm, Kardium, informed it's US distributor on March 30, 2010 by e-mail. The distributor was instructed to return all unused product to Kardium. The units remaining at hospitals will be returned to Kardium by the distributor sale reps. If you have any questions, please contact (604) 248-8891. |  | 420 | 420 units | Worldwide distribution: USA and Canada. | Terminated | October 07, 2010 | 190 | 420
Z-1974-2010 | 55598 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology Hemo Component Schiller Argus Pro PB-1000 Monitor unit. |  |  | April 15, 2010 | 2010 | McKesson Provider Technologies - Medical Imaging Group | McKesson has recently discovered that certain Horizon Cardiology Hemo systems were running specific Schiller Argus-Pro PB 1000 firmware versions that might cause a delay in the display of physiological parameters on the Horizon Cariology Hemo Real Time Monitor Screen. | Software change control | The firm, McKesson, sent an "Advisory Notice" letter dated April 15, 2010, to all customers. The letter describes the problem, products affected, and actions to be taken by customers. The customers were instructed not to use the Emergency Replacement Hemo kit until they receive a replacement Schiller Argus Pro PB-1000 and they have contacted McKesson Support department at 1-866-777-0202.If you have any questions regarding this notice, please call our Customer Support department at 1-0866-777-0202 and ask to speak to your support Manager. |  | 494 | 26 units; 19 in operational use, 8 as emergency replacement units. | Nationwide distribution: CA, CO, MD, NH, NT, NY, TN, and TX . | Terminated | July 21, 2010 | 97 | 494
Z-1995-2010 | 55206 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems enhanced AccuTnl Reagent Kit; Part Number: A78803 |  |  | March 11, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because as a follow up to PCA # 12605 dated December 16, 2009, Beckman Coulter has confirmed additional reports that each of the following assays may produce lower than expected results when run in conjunction with Access AccuTnl (REF A78803):(1) Access Testosterone (REF 33560);(2) Access Thyroid Uptake (REF 33810);(3) Access Total T4 (REF 33800).This issue can occ | Other | The firm, Beckman Coulter, issued a "URGENT: PRODUCT CORRECTIVE ACTION" (PCA) letter with attached Customer Response form dated March 24, 2010 to all customers. The letter describes the issue, impact, action and resolution. In the PCA beckman Coulter recommends the following mitigation to reduce the risk of generating erroneous results: (1) Verify the APF software on the immunoassay system. If the updated APF software listed in the table above was already loaded , skip the remainder of this section and proceed to the Resolution section;(2) If the updated APF software has not yet been loaded, but is available, load the software appropriate for the system, skip the remainder of this section and proceed to the Resolution section;(3) If the updated APF software is not immediately available, please contact Technical Support at 1-800-854-3633 in the United States and Canada. Consignees outside the United States and Canada, contact their local Beckman Coulter Representative and follow the instructions below until the software has been loaded on their system. (a) Access and Access 2 users should batch process Access Testosterone, Thyroid Uptake, or Total T4 samples; (b) For UniCel DxI users, assign Access AccuTnI (REF A78803) to a different reagent pipettor(s) than Access Testosterone, Thyroid Uptake, or Total T4;(c) The Access Testosterone, Thyroid Uptake, or Total T4 results should be interpreted in light of the total clinical presentation of the patient, including: symptoms, clinical history, data from additional tests, and other appropriate information. Abnormally low Access Testosterone, Thyroid Uptake, or Total T4 results from samples tested in conjunction with Access AccuTnI (REF A78803) should be reviewed in the context of other diagnostic tests.Customers were instructed that after loading the updated APF software, the Access AccuTnI assay (REF A78803, test code TropI, test ID 275) must be recalibrated. Recalibrate of the Access Testosterone, T |  | 32407 | 32,407 units | Worldwide distribution: USA and countries of Algeria, Australia, Bangladesh, Belgium, Bulgaria, Canada, Chile, Czech Republic, France, Germany, Greece, Hong Kong, Ireland, Israel, Italy, Lebanon, Libyan Arab Jamahiriya, Monaco, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Poland, Puerto Rico, Romania, Russian Federation, South Africa, Spain, Swaziland, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | May 04, 2012 | 785 | 32407
Z-2020-2010 | 54688 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Angiotech Galactography Kit. |  |  | September 01, 2009 | 2009 | Medical Device Technologies, Inc. | Medical Device Technologies doing business as Angiotech has conducted a recall on Angiotech Breast Localization Needles, Soft Tissue Biopsy Needles, Vascular Access Needles, Bone Biopsy Needles, Access Needles, Galactography Kit, Stabilization Needles and Needle Guides for packaging integrity. | Other | Angiotech sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated September 1, 2009, to all customers. on September 2, 2009 by UPS 2nd day delivery with an assigned tracking number. The letter described the product, problem and action to be taken by customers. The customers were instructed to please review your current inventory and segregate the affected lot numbers, complete the attached Customer Acknowledgement Form and return via fax to our Quality Assurance department at 1-352-338-0662 or 1-800-333-0440. If you are a distributor of this product, you are responsible for notifying any customer you may have sold the product to.If you have any questions concerning this notification, please contact Shannon Brooks at (352) 338-0440 ext.355 or Katrenia Williams at ext. 353. |  | 2630 | 2,630 | Worldwide distribution: USA and countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BERMUDA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, ESTONIA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, GREECE, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PERU, PHILIPPINES, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, and UNITED ARAB EMIRATES. | Terminated | February 25, 2011 | 542 | 305698
Z-1999-2010 | 55790 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Microstream O2/CO2 NASAL FILTERLINE ADULT 25UN |  |  | May 26, 2010 | 2010 | Oridion Medical 1987 Ltd. | The Microstream Filterline may not be recognized by the monitor when connected. No CO2 readings will be displayed and no alarms may sound for unmonitored patients. Replacing the filterline may delay treatment. | Process design | Oridion notified customers by letter on 5/26/10 and requested that they remove the FilterLines from the related production lots from the field and an immediate replacement with FilterLines from subsequent production lots. Questions are directed to the responsible person at the supplier named in the letter, or to Oridion Capnography at 888-674-3466. |  | 2 | 2 units | AZ, CA, FL, HI, LA, , MA, MO, NJ, OH, PA, WA, WI, | Terminated | June 21, 2012 | 757 | 21212327
Z-1934-2010 | 55648 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Constellation Vision System, Model: Constellation T, Catalogue Number: 8065751558 |  |  | April 30, 2010 | 2010 | Alcon Research LTD dba Alcon Laboratories, Inc. | Alcon identified system performance and machine settings that may impact the infusion performance of the CONSTELLATION Vision System, causing unexpected shut down, unresponsive touch screens, and fluidics issues. | Software design | The firm, Alcon, sent a "URGENT: MEDICAL DEVICE RECALL" letter dated July 2, 2010, along with training material to all customers. The letter informed the customers of the problem and indicated that Alcon Surgical Sales representative will be contacting them in the near future to follow-up on this training material and to answer any questions you may have. Alcon informed the customers that this action will not require removal of the CONSTELLATION Vision System from their facility. The letter also informed the customers that their Alcon Technical Customer Service Representative will be contacting them when we are authorized to provide you with the software and hardware updates for your system.For assistance and questions regarding the CONSTELLATION Vision System, customers were told they may contact their Alcon Surgical Sales Representative or Alcon Technical Consumer Affairs at 1-800-561-6466, Option 8. |  | 824 | 824 units total | Worldwide distribution. | Terminated | September 19, 2011 | 507 | 824
Z-1998-2010 | 55675 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | O-Arm Imaging System |  |  | May 10, 2010 | 2010 | Medtronic Navigation, Inc. | The possibility that the real time image display can freeze up during use of a 2D continuous fluoroscopy. | Device Design | The firm, Medtronic, Inc., sent an "Urgent Field Safety Notice" letter dated May 10, 2010, to all custmers. The letter describes the product, problem and actions to be taken by customer. The firm recommended that the users of the O-Arm Imaging System reveiw the information provided to ensure patient safety remains uncompromised during use of the product. The customers were instructed to discontinue use and contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or (702) 890-3200, if the image is not updating as expected. A Medtronic Navigation representative will be contacting you within the coming months to schedule a time to evaluate and service the potentially affected components within your system. A service visit will also be scheduled starting November 2010 to update your system with a software based solution that will prevent this issue from occuring.If you have any questions about this notification, please contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or (702) 890-3200. |  | 221 | 221 | Worldwide distribution: USA and countries of Germany, Singapore, South Africa, Poland, Libya, Belgium, Switzerland, Turkey, Italy, Denmark, France, Spain, Finland, Czech Republic, Ireland, Slovakia, the NetherLands, Austria, UAE, Kuwait, Israel, Australia, Sweden, Canada, South Korea, Japan and India. | Terminated | October 07, 2011 | 515 | 221
Z-1997-2010 | 55544 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | PVT-681MV Endocavitary Transducer used with the Aplio XG and Xario XG Ultrasound Systems. |  |  | January 05, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on the PVT-681MV Endocavitary Transducer used with the Aplio XG and Xario XG Ultrasound Systems due to an incorrect sweep direction setting. Images acquired using the endo-cavitary transducer PVT-681MV and cross-sectional images created from these volume images are reversedhorizontally on the Aplio XG (SSA-790A) and the Xario XG ( | Other | Toshiba America Medical Systems (TAMS) will issue notification to consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. matter. Questions regarding this letter are directed to the firm at (800) 421-1968 or to your local Toshiba representative. |  | 27 | 27 | Nationwide | Terminated | February 29, 2012 | 785 | 27
Z-1996-2010 | 55895 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | i-STAT PT/INR Cartridge |  |  | May 05, 2010 | 2010 | Abbott Point of Care Inc. | A high mean bias (12-16%) was documented in 8 lots of i-STAT PT/INR cartridges. This may lead a clinician to believe that a patient is adequately anticoagulated when they are not. Inadequate anticoagluation therapy may be given and issues with clotting and/or thrombosis could occur. No reports of injury have been received. | Other | Urgent recall notifications were sent on May 13 2010. Customers were instructed to discontinue use and to return the identified lots. Additional questions are directed to Abbott Point of Care Technical Support at 800-366Âż8020, Option 1 or your Abbott Point of Care representative. |  | 164072448 | 44976 in US, 3648 in rest of world. | Worldwide distribution: Argentina, Hong Kong, China, Finland, Greece, India, Portugal, Russian Federation, Saudi Arabia, United Arab Emirates, Italy, Spain, Australia, Germany, UK, Sweden, The Netherlands, France, Canada. | Terminated | September 22, 2011 | 505 | 164072448
Z-2025-2010 | 55867 | No | GNC | Salmonella spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Difco Salmonella 0 Group A Antigen, catalog #240731, packaged in 5 mL glass vials, labeled in part ***Becton, Dickinson and Company, Sparks, MD 21152*** |  |  | May 19, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic reagent to aid in the diagnosis of Salmonellosis may exhibit decreased or no reactivity. If proper controls are run with the lot in question, it would become apparent that the antigen is non reactive and patient specimen results should not be reported. If control testing is not performed a salmonella-infected patient could go undiagnosed, with possible adverse effects. | Nonconforming Material/Component | The recalling firm notified customers by fax and e-mail on 05/19/10. The notification advised that the referenced lot number of Salmonella 0 Group A Antigen may exhibit decreased or no reactivity when testing the patient's serum directly for homologous antibodies by either a slide or tube agglutination test to aid in the diagnosis of Salmonellosis. The recalling firm also advised that they would issue replacements for the referenced lot number under Catalog number 240731, which would not be available until approximately July 15, 2010. Consignees were requested to return enclosed response form.Customers with questions about replacements should contact BD customer service at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. |  | 93 | 93 units | Product was sold to distributors nationwide and to affiliates in Belgium and Korea. | Terminated | July 21, 2010 | 63 | 93
Z-2027-2010 | 54565 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OneTouch SureStep and SureStep Pro Test strips, for use with the OneTouch SureStep Blood Glucose Meters, in 100 count and 50 count packaging distributed by LifeScan Inc., Milpitas, CA, Manufactured by LifeScan Puerto Rico |  |  | February 24, 2010 | 2010 | Lifescan Inc | Products may provide inaccurate results at readings over 400 mg/dL | Other | The firm, Lifescan, Inc., issued two "Urgent Medical Device Recall" letters dated February 24, 2010 to all customers for the OneTouch SureStep Test Strips and the SureStep Pro Test Strips. The letters describes the product, problem and action to be taken by the customers. The customers were instructed to look at your test strip vials for the recalled lots and check the strips or if you have test strips from the lot, call LifeScan Customer Service at 800-574-6139 between 5 a.m. to 7 p.m. Pacific Time, seven days a week or visit www.SureStep.com. Replacement product will be shipped to you immediately at no charge. The customers were instructed to continue to test your blood glucose, while wafting for replacement. Also, the customers are asked to continue use of the meter with precautions if results over 400 units are observed. For more information, please visit www.SureStep.com or call our Healthcare Professional Line at 888-455-5627. |  | 1014350 | 1,014,350 | Worldwide distribution: USA and countries of Europe, Latin America and Asia. | Terminated | December 08, 2010 | 287 | 9640700
Z-2026-2010 | 54521 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ADVANTAGE Turbo 2-Button Shaver. Rx Only, Made In USA, NON STERILE. |  |  | October 29, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec is recalling Advantage, 2-Button Turbo Handpiece Catalog Number D9924. These products may self-activate due to moisture intrusion into the housing body. This issue has been isolated to products manufactured between June 7, 2004 and June 1, 2006. | Other | ConMed Linvatech notified affected customers by letter sent by FedEx next day delivery. Customers were asked to complete a response form and indicated if they have the product or have transferred the product. A Service Request number will be assigned for the returned product.Additional questions are directed to the firm's Customer Service department at 800-535-8536. |  | 2373 | 2,373 | Worldwide distribution. Products were distributed domestically throughout the U.S. Also United Arab Emirates, Argentina, Australia, Belgium, Switzerland, Chile, China, Germany, Denmark, Spain, France, United Kingdom, Hong Kong, India, Italy, Japan, Lithuania, Morocco, Netherlands, New Zealand, Singapore, Turkey, South Africa. | Terminated | August 19, 2010 | 294 | 2373
Z-2031-2010 | 55575 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | DePuy ASR 100 Acetabular Implant, 44 mm, 999800744, sterile, DePuy International, Ltd, Leeds, United Kingdom. |  |  | March 05, 2010 | 2010 | Depuy Orthopaedics, Inc. | Notification to clinicians of new revision rate data/information regarding the use of the device. As part of the post-market surveillance, DePuy is continually evaluating data from a variety of sources including national joint replacement registries, published literature, company sponsored clinical trials, internal complaints data and unpublished clinical research reports. As part of the post-m | Other | DePuy Orthopaedics issued a Field Safety Notice to share the new data with surgeons and to reiterate the importance of correct implant positioning and patient selection in achieving optimal implant performance and survivorshop. For additional information regarding this matter, please contact the DePuy Scientific Information Office at 888 554 2482. |  | 388 | 388 | Nationwide | Terminated | May 10, 2012 | 797 | 37834
Z-2030-2010 | 55900 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic U-CLIP Double Arm Approximation Device, model V100D U-Clip, Catalog Number V100DRFRN2. Sterilized using Irradiation. The U-Clip consists of specially designed vascular clips optionally connected to needles via flexible members. These clips may or may not include coils surrounding a Nitinol wire core. This design allows precise placement of clips prior to closure and facilitation from standard medical and implantable grade materials. Accessories include Clip delivery and removal devices as well as vessel cutting and retention tools. |  |  | April 02, 2010 | 2010 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | Medtronic received reports involving the V100D U-Clip where during handling, the release mechanism has become separated from the needle assembly. These instances of separation have occurred when pulling the clip through prosthetic material, or when grasping and releasing the clip using robotic instrumentation. While this issue does not impact performance of the clip once implanted, separation of | Other | The firm, Medtronic CardioVascular, sent an "Urgent Medical Device Recall Notice", dated April 2, 2010, to Risk Managers of each affected account. The letter described the issue, identified affected product, and the action to be taken by customers. The customers were ask to immediately discontinue use of affected devices and to quarantine all unused V100D U-Clip inventory. The letter stated that a Medtronic representative will contact the account to arrange return of the V100D U-Clips devices and that credit will be issued upon return of unused devices. Medtronic has stopped manufacture of the product therefore replacements will not be available. Devices can be returned to Medtronic CardioVascular, 7611 Northland Drive, Brooklyn Park, MN 55428. The customers were ask to please fill out the Medtronic Recall Certificate and fax it to 651-367-2620.If you need additional information, please contact CardioVascular LifeLine Technical Services at 877-526-7890, or you local Medtronic sales representative. |  | 22946 | 22,946 ( 20,714 US, 2,232 OUS) | Worldwide distribution: USA including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and countries of Australia, Belgium, Canada, China, France, Germany, Greece, Hong Kong, Hungary, Japan, Korea, Netherlands, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, United Kingdom, and Venezuela. | Terminated | December 17, 2011 | 624 | 22946
Z-2029-2010 | 56013 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic StealthStation S7 System |  |  | June 14, 2010 | 2010 | Medtronic Navigation, Inc | Excessive wear at the power cord retractor or incorrect assembly of the power plug may cause sparking, intermittent loss of power, or electrical shock. | Device Design | Customers were notified by letter on Jun 14, 2010, advising them of the problem. Instructions were provided to continue using existing retractable power cords until they could be replaced with non-retractable cords by Medtronic, beginning on Aug 17. Further information is available from the company by calling 800-595-9709. |  | 370 | 370 units | Nationwide, including 3 VA facilities. No military distribution. Foreign accounts in Australia, Austria, Belgium, Canada, Chile, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Latvia, Libya, Lithuania, Luxembourg, Malaysia, Namibia, Netherlands, New Zealand, Poland, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, United Arab Emirates. | Terminated | February 18, 2011 | 249 | 370
Z-2062-2010 | 56004 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Hemashield Double Velour Cardiovascular Fabrics. |  |  | May 24, 2010 | 2010 | Maquet Cardiovascular, LLC | Various Hemashield products may develop glycerol droplets on the inner blister of the package if not stored under proper conditions, which could lead to the prolongation of surgery and potentially compromise patient safety. | Package design/selection | Letters were sent by Federal Express commencing July 12, 2010.Distribution has to be obtained from Boston Scientific.Questions are directed to the local Sales/Service Representative, or to the firm's Customer Service at 1-888-880-2874. |  |  | xx | Nationwide Distribution. | Terminated | July 14, 2011 | 416
Z-2075-2010 | 56139 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Electrical Venous Occluder, Catalog 12-80-00, for use with Sorin HLM C5 and HLM S5. |  |  | June 18, 2010 | 2010 | Sorin Group USA, Inc. | Although accessories were designed for use with two models, the HLM C5 and the HLM S5, the submission that included references to the Mast Roller Pump, B-Care5, and the Electrical Venous Line Occluder accessories covered the new HLM S5 system. The three accessories were not cleared for use with the HLM C5. | PMA | The two customers who had received the suspect device accessories were identified and were contacted by letter on 6/18/2010. They were told to immediately discontinue use of the accessories and arrange for return of the products. They were also told that they could continue to use the CD5 Machine without the recalled accessories. A Product Recall Response Form was to completed by each consignee. Questions are directed to the firm at 800-221-7943. |  | 1 | 1 unit | CA, PA, TN. | Terminated | July 21, 2010 | 33 | 3
Z-2072-2010 | 56056 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art SystemÂż, ver. 4.0.0, 4.0.1, 4.0.2, ThomoTherapy Incorporated 1240 Deming Way, Madison, WI 53717. The TomoTherapy HI-ART SystemÂż is intended to be used as an integrated system for the planning and delivery of intensity modulated radiation therapy (IMRT). The HI-ART System provides precise delivery of radiation to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. |  |  | June 23, 2010 | 2010 | TomoTherapy Incorporated | An issue has been identified, as a result of input received from one TomoTherapy Hi-Art Treatment System customer.With a particular workflow, the Planning Station's Draft and/or final Plan Report can provide incorrect dose statistics. | Software design | Consignees were sent on 6/23/10 a TomoTherapy "Urgent Field Safety Notice Medical device Correction" letter dated June 22, 1010. The letter addressed the Issue, Product Affected, Recommended Action and Resolution.Prior to generating a draft or final Plan Report for a given plan, the plan must be closed and reopened at least once. This issue will be remedied by Hi-Art software version 4.0.3, scheduled for release in late 2010.Customers with questions were provided with a contact in the letter at the TomoTherapy Customer Interaction Center. The firm can be contacted at 608-824-2800. |  | 29 | 29 | CA, CO, IL, LA, MI, MO, NY, ND, OH, OK, PA,VA, WI, BELGIUM CANADA, KOREA, MALAYSIA, SWITZERLAND, ITALY, GREAT BRITAIN, FRANCE. | Terminated | December 17, 2011 | 542 | 29
Z-2071-2010 | 56059 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Fathom"-14 Steerable Guidewire, Peripheral Vascular Use, Pre-Shaped 300 cm x 10 cm, REF 50-815, UPN M001508150 Sterilized with ethylene oxide gas, Rx Only, Made in 2405 Orion Circle, West Valley City, UT 84119 USA. |  |  | June 03, 2010 | 2010 | Boston Scientific Corporation | Boston Scientific is initiating a Medical Device Recall regarding specific lots of the FathomTM-14 Steerable Guidewire. For lubricity, the distal portion of these guidewires is coated with hydrophilic polymer and the proximal portion is coated with polytetrafluoroethylene (PTFE). Through internal inspection process, the firm determined that Fathom guidewires from affected lots may exhibit a P | Process design | Consignees were sent on 6/3/10 an "Urgent Medical Device Recall" letter dated 3 June 2010. The letter was addressed to Risk Manger/Interventional Radiology Manger/Material Manger. The letter described the problem and the product involved. The letter included the "Customer Instructions" which stated to Immediately discontinue use of an segregate recalled product and, requested customers to return the Reply Verification Tracking Form. Questions are directed to the firm at 510-624-2563. |  | 78 | 78 units | AL, AR, CA, CO, CT, FL, GA, ID, IL, KS, KY, LA, MD, MA, MS, NV, NY, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT,VA, WA, WT. | Terminated | December 10, 2011 | 555 | 745
Z-2080-2010 | 56084 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTISY-100-HC-PERC6, G13166. |  |  | April 21, 2010 | 2010 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly. | Nonconforming Material/Component | Cook Medical sent an Urgent Product Recall letter dated April 21, 2010, to hospitals and other user facilities receiving product direct from Cook as well as distribution centers receiving direct shipments from Cook. Consignees were directed to cease distribution of product and notify their customers immediately. |  | 28819 | 28819 total, all products | Worldwide distribution: USA, Canada, Mexico, Australia, Chile, Costa Rica, Denmark, Dominican Republic, Guatemala, India, Malaysia, Panama, Peru, Puerto Rico, Republic of Korea, Trinidad and Tobago, and Uruguay. | Terminated | January 23, 2012 | 642 | 28819
Z-2100-2010 | 56158 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | 19" Barco MFCD 1219 (Touch Screen) + Low Profile Stand and Rack with LCD Arm |  |  | June 30, 2010 | 2010 | AGFA Corp. | Monitor fell from the mounting bracket used to support the monitor. | Other | The firm, AGFA Healthcare sent an 'URGENT SAFETY NOTICE" dated June 30, 2010, via FED-EX to customers. The letter describes product, problem and actions to be taken by the firm and customers. AGFA will provide a safety check to the customers facility, have a service representative visit , in the near future, their site to inspect the monitor/s and attach a warning label on each monitor checked for preventative action. The customers are ask to distribute the information within their facility to all individuals who need to be aware, to complete the attached URGENT SAFETY NOTICE FEEDBACK FORM as soon as possible and return via fax at 864-421-1664. If you have any questions about this matter, please feel free to contact me via email at Debbie.norris@agfa.com or call me at 864-421-1754. |  | 805 | 805 units | Worldwide distribution: USA and countries of Puerto Rico and Canada. | Terminated | April 28, 2011 | 302 | 805
Z-2109-2010 | 56130 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MICROTEK MEDICAL, INC, STACKHOUSE LASER RESISTANT SUCTION PROBE, Size: 1/4" x 9" (0.64cm x 22.9cm), REF 24014, QTY 10 EA, STERILE EO, 2015-05, Microtek Medical, Inc. , Columbus, MS 39702, www.microtekmed.com, CE0044, EC REP Microtek Medical B.V. , Zutphen, The Netherlands, Manufactured in the U.S.A. Suction probe for a smoke evacuation device |  |  | May 25, 2010 | 2010 | Microtek Medical Inc | Product labeled as sterile was distributed without sterilization. | Other | The firm notified its consignees of the problem and requested return of the product by telephone on 05/25/2010. They followed with a letter dated on 06/03/2010 re: Product Removal stating the reason for the recall and confirming the receipt of the recalled product. For further information please contact (662) 244-3220 |  | 20 | 2 cases of 10 units | Nationwide distribution including the states of Montana and Georgia | Terminated | August 12, 2010 | 79 | 20
Z-2111-2010 | 56085 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Temno Coaxial Chiba Fine Needle Aspiration; a coaxial introducer needle (20G x 15 cm) with a Chiba biopsy needle (22G x 20 cm); sterile, for single use only sets, 5 sets per case; CareFusion, McGaw Park, IL 60085 USA, Made in Dominican Republic; REF CCH2220 |  |  | June 25, 2010 | 2010 | Carefusion 2200 Inc | The introducer needle is only 10 cm in length instead of the required 15 cm. | Mixed-up of materials/components | Cardinal Health, on behalf of CareFusion, telephoned the customers on June 15, 2010 and sent a follow-up "URGENT: PRODUCT RECALL" letter on the same date via UPS next day air. The letter describes product, problem, and action taken by customers. The customers were instructed to immediately stop using affected product, return the enclosed acknowledgment form via fax to 847-686-9101 (note: Cardinal Health is required to confirm receipt of this notification from the customers), notify any customers to whom they may have distributed product to, contact CareFusion Customer Service at 800-653-6827 to arrange for credit and return of the affected product, and send the affected product to CareFusion, 1240 Waukegan Rd, Waukegan, IL 60085, Attn: Scott Fitzer, after obtaining an RMA/RGA number from CareFusion.If you have any questions, please contact 800-292-9332. |  | 65 | 65 units | Nationwide distribution: California, Florida, Georgia, Illinois, Kentucky, Massachusetts, New Jersey, New York and Pennsylvania. | Terminated | January 24, 2011 | 213 | 65
Z-2107-2010 | 56169 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | GEM Check Coag CNTRL Normal APTT Whole Blood Controls Package Insert: |  |  | July 08, 2010 | 2010 | International Technidyne Corporation | Certain lots of GEM Check Coag Whole Blood Controls contain an error in the package insert where the incorrect Acceptable Performance Ranges were published. | Error in labeling | Urgent: Medical Device Correction Notices, dated July 1, 2010, were sent by Federal Express on July 8, 2010 to all first consignee end users. The letters identified the affected product and asked customers to review and follow all instructions as outlined in the ITC in the communication. The info should be shared with laboratory staff and the notification should be retained as part of the laboratory Quality System documentation. The enclosed response form should be completed and returned to ITC within 5 days. |  | 3040 | 95 kits C9GC5001; 32 kits G9GC5003; 46 kits J9GC5004; 5 kits C0GC5001; Each kit contains 15 vials. | Worldwide Distribution -- USA, including states of NY, TN, CA, AR, IL and TX and countries of Australia, Italy, Spain, Germany, Chile, Spain, Malta, Belgium, Austria, and Switzerland. | Terminated | September 23, 2010 | 77 | 3141
Z-2110-2010 | 56114 | No | KFF | Heparin assay. | Hematology | Hematology | 510(k) | RQCHRT HRT Control Level 1 &amp; 2; |  |  | June 24, 2010 | 2010 | International Technidyne Corp. | An ingredient in the RQCHRT Whole Blood Controls for use with Hemochron Kaolin and Celite HRT Test tubes may become less potent over time and result in excessive High LQC Failures. | Other | ITC sent "URGENT MEDICAL DEVICE RECALL" letters dated June 24, 2010, via Federal Express to Laboratory Director, Clinical Laboratory Manager, or Point of Care Coordinator consignees/customers. The letter described the product, problem and action to be taken by customers. The customers were instructed to check their inventory for the product, if they have boxes or loose vials to please stop using the product, remove it from inventory and contact ITC Technical Support to arrange return. They were also instructed to complete the attached Customer Account Tracking Form, whether or not they have remaining inventory, and return it by fax, email or mail.If you have any questions, please call ITC Technical Support at 800-631-5945 (US) or 732-548-5700 (international), Extension 4707 or email techsupport@itcmed.com. |  | 114 | 38 in US, 3 Out of US | Worldwide distribution: USA including states of CA, NC, TX, FL, MD, OK, NY, and NJ and country of Italy. | Terminated | September 23, 2010 | 91 | 114
Z-2112-2010 | 56090 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Results Manager/Results Manager Administration Tool |  |  | June 11, 2010 | 2010 | AGFA Corp. | Agfa Service technician had incorrectly modified a report template at one site resulting in erroneous anatomic segment locators. | Other | A letter was sent to the consignee on 6/11/2010 indicating what steps would be taken by Agfa concerning the problem. The letter also included a feedback form which was to be returned to Agfa. The recall was expanded and a "URGENT SAFETY NOTICE" was sent to each additional consignee on January 25, 2012. The letter described the potential issue and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood has been requested from the additional consignees. |  | 358 | 358 | AK, AZ, CA, CO, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, MS, MT, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, VT, VA, WA, WV, WI and Canada | Terminated | July 02, 2013 | 1117 | 358
Z-2113-2010 | 56057 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | BleaseSirius Anaesthesia System, gas machine. Blease Beech House, Chiltern Court Asheridge Road, Chesham, Bucks HP5 2PX, England |  |  | June 24, 2010 | 2010 | Spacelabs Healthcare, Llc | Potential for the touch screen on the ventilator portion of the device to go blank when touched. Possibility that this failure may also cause mechanical ventilation to stop. | Other | On 06/24/2010, Spacelabs began mailing the URGENT-MEDICAL DEVICE CORRECTION letter to their US consignees and e-mailed the same letter to their foreign consignees.The letter describes the touch screen of the BleaseSirius Anesthesia machine goes blank when touched and the possibility that this failure may cause the mechanical ventilation to stop.The consignees are asked to weigh the benefits versus the risk when deciding to use the device or not until the firm upgrades the device. The letter also instructs the consignees to notify their staff about the possibility of the display failure to prevent any adverse event that may occur. The letter states that their field service engineer will contact consignees to schedule a time to visit and upgrade all affected devices at no cost.The US consignees are advised to call 1-800-522-7025, Select 2, and the foreign consignees are advised to call 1-425-657-7200 x5089 for technical support or questions regarding the recall. |  | 80 | 80 units total | Worldwide Distribution -- USA, including states of IA, GA, LA, NC, TN, and TX and countries of AUSTRALIA, CHILE, CHINA, INDIA, ITALY, PHILIPPINES, MEXICO, POLAND, SINGAPORE, THAILAND, TURKEY, UNITED KINGDOM, and VENEZUELA. | Terminated | December 03, 2010 | 162 | 80
Z-2115-2010 | 56052 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art SystemÂż, Version 4.0.0 &amp; 4.0.1. |  |  | February 22, 2010 | 2010 | TomoTherapy Incorporated | In some cases, the patient's diagnostic CT image is narrower than the Hi-Art radiotherapy couch image. | Software design | Consignees were sent on 2/22/10 a TomoTherapy Incorporated "Urgent Medical Device Correction Field Safety Notice" letter dated Monday, February 22, 2010. The letter was addressed to TomoTherapy Customer and identified the the affected product, and included a description of the Issue, Product Affected, Recommendation Actions and Resolution. The issue will be remedied by software versions 4.0.2 and later. Questions should be directed towards TomoTherapy Customer Interaction Center by e-mail or telephone. |  | 7 | 7 | Worldwide Distribution -- MI, MO, VA, and WI and country of CANADA. | Terminated | February 10, 2011 | 353 | 7
Z-2060-2010 | 55224 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Terumo HX2 Temperature Management System 120V, 50/60 Hz catalog number 809810, Terumo Cardiovascular Systems Corp., Ann Arbor MI. |  |  | June 11, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The HX2 Temperature Management System mixing valve may fail resulting in an unusable water channel and then the machine would go into standby mode. | Device Design | The firm, TERUMO Cardiovascular Systems, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 11, 2010, to the consignees. The letter describes the product, reason for correction, what to do in the event of failure, potential hazard, correction, affected population and action to be taken by the customers. The customers are instructed to review the correction notice, assure that all users are aware of this notice, and complete, sign and fax the attached Customer Response Form to Terumo CVS Customer Service at 800-292-6551. The Terumo CVS Customer Service Dept. will contact you upon receipt of the Customer Response Form to schedule the replacement. The firm plans to inspect and correct all units in the field within 5 months.We encourage you to contact Terumo CVS Customer Service at 1-800-521-2818, Monday-Friday, 8 AM-6 PM EST, with any questions or concerns. |  | 100 | 100 | Nationwide Distribution: AR, CA, CO, CT, FL, HI, IA, IL, KS, LA, MD, MI, MO, MS, NY, OH, PA, TX, UT, WA and WI. | Terminated | April 13, 2011 | 306 | 100
Z-2067-2010 | 55868 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. |  |  | May 26, 2010 | 2010 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity Laboratory software where results from interfaced external reference laboratories are not displayed or stored correctly. This happens when the test is set up in Centricity Laboratory as a decimal (numeric) type result and the test result sent from the reference lab includes a comma in the result field. | Software design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated May 26, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the patient safety issue involving results from interfaced external reference laboratories that are not displayed or stored correctly if an unexpected alpha character is encountered in a result message that is expected to have a numeric value. This happens when the test is set up in Centricity Laboratory as a decimal (numeric) type result and the result sent from the reference lab includes a comma in the result field. The letters provided the users with specific safety instructions to follow until the software is updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to the their GE Service Representative or the GE Helpdesk at 888-778-3375. |  | 37 | 37 units | Worldwide distribution: California, Georgia, Michigan, Minnesota, New York, Ohio, Oregon, Tennessee and Texas, and internationally to Australia, Canada, China, England, India, Malaysia, New Zealand, Qatar, Scotland and Singapore. | Terminated | July 15, 2011 | 415 | 37
Z-2093-2010 | 55578 | No | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | The "RS-FBG Full Back Conductive Garment" with electrode attachments for use with interferential and muscle stimulation with a Transcutaneous Electrical Nerve Stimulation device (TENS) and/or muscle stimulator. |  |  | May 17, 2010 | 2010 | International Rehabilitative Sciences, Inc. | Device distributed with labeling and IFU outside of their 510(k) clearance. | Other | The firm, RS Medical, mailed "URGENT NOTICE OF VOLUNTARY CORRECTIVE ACTION" letters dated May 14, 2010, to customers on May 17, 2010 and letters to physicians/consignees on May 19, 2010. The letters included an updated RS-FBG Operation Manual and stated what the update covered; i.e. that the device is indicated for use on the "middle and upper back". The customers were ask to please read your new manual and replace your original RS-FBG Full Back Conductive Garment Operation Manual with this updated version and discard your old RS-FBG Garment Operation Manual. RS Medical stated that your RS-FBG Garment is not affected by this Corrective Action, only the Operation Manual. RS Medical informed the customers that you may continue to use the RS-FBG Full Back Garment with your RS Medical Stimulator as prescribed by your healthcare provider.If you have any questions or concerns, please call our Toll Free Customer Service number at 1-866-417-1293, email us at RSFBG@rsmedical.com or visit our special web page at www.rsmedical.com/rsfbginfo. |  | 190073 | 108134 units, 81,939 subject to recall | Nationwide distribution. | Terminated | August 15, 2012 | 821 | 190073
Z-2094-2010 | 56033 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Rolko Trapeze handles, part 148671, for use with Care Assist bed P1170 and Hill Rom 1000 bed P1160, Hill Rom Batesville, IN. |  |  | June 11, 2010 | 2010 | Hill-Rom, Inc. | The firm received a complaint regarding the plastic housing used to support the trapeze handle. The firm investigated and found that the manufacturer of the trapeze handle didn't correctly process the plastic component. Therefore the plastic had the potential to break prematurely. | Process control | The firm, Hill-Rom, sent an "Urgent Field Safety Notice" letter dated June 11, 2010, to its customers. The letter describes the product, problem and action to be taken by the customers. Hill-Rom provided the quantity of handles needed to replace handles customers received during the affected period. The customers were instructed to immediately remove and replace all Trapeze Handles that were delivered during the affected period; to replace the Trapeze Handle Assemblies (with a manufacturing date from January 2007 to January 2009) used on their Patient Helper with the enclosed unit(s); complete and return the Customer Response Form as soon as possible, or by July 26, 2010; advise the firm immediately if they have relocated the beds and/or the accessories to another location, and forward the letter to any other facility personnel they deem appropriate. If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. |  | 319 | 319 | Worldwide distribution: USA and Canada. | Terminated | February 18, 2011 | 252 | 319
Z-2101-2010 | 56161 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Merit Laureate Hydrophylic Guide Wire, Angled, Made In Ireland, CAT No: LWSTDA35180, 180cm, Sterile EO, Merit Medical Systems, Inc., South Jordan, UT 84095. |  |  | June 28, 2010 | 2010 | Merit Medical Systems, Inc. | Guidewires have decreased lubricity at the proximal end, inhibiting proper operation. | Process control | The firm, Merit Medical System, Inc., notified the Sales representatives by e-mail on June 28, 2010, and instructed them to contact all of their accounts and to advise them to discontinue use and return any unused products to Merit. A Product Recall Notice was to be completed by each facility under the direction of the rep and all activities were to be completed by July 9, 2010. Questions should be directed to Penny Goldman, Senior Product Manager at 801-208-4198 or Jackie Preece, Manager, Customer Service at 801-208-4365. |  | 437 | 437 units | Nationwide distribution: CA, CO, FL, GA, IL, IN, KS, MA, MD, MI, MN, NC, NH, NJ, NY, PA, and SD. | Terminated | October 19, 2010 | 113 | 964
Z-2095-2010 | 55465 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Evolis-Femoral Sizer Posterior Reference |  |  | September 11, 2009 | 2009 | Medacta Usa Inc | Medacta USA Inc is initiating a recall on the Femoral Sizer, an orthopedic manual surgical instrument, because the "L" and "R" markings are mislabeled which would cause incorrect usage. | Process design | The firm, Medacta USA Inc, sent an "Urgent Safety Notification" letter dated August 5, 2009, to the customers. The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to check their inventory for devices listed, if found, to immediately sequester them and DO NOT use them in surgery. The customers were ask to please pack the affected femoral sizer(s) posterior reference for shipment to Medacta USA. Medacta informed the customers that Medacta USA personnel will contact them by telephone to provide them with an RMA number and the reworked femoral sizer(s) posterior reference will be returned as soon as they are available.If you have any questions, please feel free to contact the Medacta USA office at 805-437-7085 or contact Ms. Heather Neely directly at 805-886-1069. |  | 5 | 5 | Worldwide distribution: USA including states of Texas and Idaho and country of France. | Terminated | February 29, 2012 | 901 | 5
Z-2118-2010 | 55330 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Howmedica Osteonics Corp, |  |  | April 05, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | There is the potential for the posts on the Distal Resection Guide Stand to disassociate from the main body part. | Device Design | Stryker Urgent Product Recall letters dated April 8, 2010, were sent by Federal Express to Stryker Branches/agencies, Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. All hospitals and surgeons that used the product in the last year were notified. |  | 1981 | 676 in US, 1305 internationally. | Worldwide distribution. | Terminated | August 28, 2012 | 876 | 1981
Z-2117-2010 | 55226 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Dimension Vista LOCI Reaction Vessels, catalog number KS855. For in vitro diagnostic use. |  |  | March 15, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Vessel pick up errors and vessel jams when the vessel is being loaded. The error halts the instrument and disrupts processing. Vessel flange is thin. | Process design | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice dated March 2010 informing users of the potential problem, to discard any remaining product, and if needed, to request a replacement lot. |  | 5544 | 5544 bags | Worldwide distribution: USA, Canada, Europe, Japan, Korea, and New Zealand. | Terminated | August 03, 2010 | 141 | 5544
Z-2119-2010 | 55357 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | X-Celerate Instrumentation Punch Thru Tibial Base Plate, Catalog numbers: 3750-0003. 3750-0004, 3750-0005, 3740-0006, 3750-0007, 3750-0009, 3750-0011, 3750-0013; |  |  | April 05, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The locating pins on the Punch Thru Tibial Baseplates could break or disassociate from the main body of the instrument. | Device Design | Stryker Urgent Device Correction letters dated April 5, 2010, were sent to all Stryker branches/agencies, hospital Risk Management and surgeons who have used the product by Federal Express. Questions on the bulletin should be directed to Scott Solis, Product Manager ,Knee Marketing at 201-831-5665. |  | 8276 | 8276 | Worldwide distribution: USA, China, Taiwan, India, Australia, Korea, Singapore, Germany, France, Spain, Italy, Switzerland, Poland, Romania, South Africa, United Kingdom, Netherlands, Sweden, Turkey, Japan, Chile, Brazil, Colombia, Mexico, Panama, Argentina, and Uruguay. | Terminated | August 28, 2012 | 876 | 8276
Z-2164-2010 | 55605 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | MIKA Speedblock, Size 8-12. |  |  | April 26, 2010 | 2010 | Encore Medical, Lp | A crack or complete fracture at the anterior chamfer cut slot may occur on the size 8, 10, and 12 MIKA Speedblocks. | Component design/selection | An "Urgent Field Safety Notice" Letter dated May 20, 2010, was sent to the customers. The letter described the product, problem and action to be taken by customer. The customers were instructed to locate, inspect and verify the quantity still in their possession, complete and return the field safety notice response page to Beth Onderlinde at beth.onderlinde@djosurgical.com or fax to 512-834-6313 by June 7, 2010, (completion of form is mandatory), and identify and inform customers that they may have distributed/sold product to, of the notice. The customer was also instructed to please inspect the listed part numbers prior to use and again after use, if a crack is found in the speedblock, contact your Customer Service Representative and be sure to utilize the femoral/tibial impactor (801-01-043) for impaction of the speedblock and reiterate to your physicians that the MIKA mini slaphammer (800-02-290) should be used for speedblock removal. The product will be replaced as it becomes available.If you have any questions, call Director of Commercial Logistics at (512) 834-6330 or International Customer Service at (512) 834-6275. You may also contact our EU Authorized Representative, MDSS GmbH, via email at info@mdss.com or via phone at +49 (0) 511 6262 8630. |  | 796 | 729 and 67 specials | Worldwide distribution. | Terminated | December 15, 2011 | 598 | 796
Z-2140-2010 | 55931 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Exacta-Mix EVA Container, Sterile R, 2000 mL, Order No. REF: 140, Baxa Corporation. |  |  | May 26, 2010 | 2010 | Baxa Corporation | Fill port cap on TPN bags may become detached prior to use, thus compromising the sterility of the fluid pathway. | Process design | Consignees were notified by e-mail and Baxa Bag Product Notice letter on 05/26/2010 with instructions to not use any bags if the fill port cap is detached. Photos of a detached cap and properly attached cap were included. Defective bags could be returned to Baxa for credit. Information available at 800-678-2292. |  | 3306700 | 3,306,700 bags | Worldwide distribution, including USA, Argentina, Australia, Bermuda, Brazil, Canada, China, Colombia, Hong Kong, Malaysia, Mexico, New Zealand, Pakistan, Panama, South Korea, Taiwan, United Kingdom, and Uruguay. | Terminated | August 04, 2011 | 435 | 12612066
Z-2133-2010 | 55621 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BARD DORADO PTA BALLOON DILATATION CATHETERS, PRODUCT CODE # DR135712 |  |  | May 04, 2010 | 2010 | Bard Peripheral Vascular Inc | The affected DORADO PTA Balloon Dilatation catheters contain catheter materials that, under certain conditions, are subject to premature breakage allowing the balloon to partially or completely separate from the catheter shaft. | Nonconforming Material/Component | The Bard Peripheral Vascular "URGENT MEDICAL DEVICE RECALL NOTIFICATION" with attached "RECALL AND EFFECTIVENESS CHECK FORM" was sent to the affected consignees beginning May 7, 2010 via FedEx. The recall notice informed the consignees of the reason for recall, the products affected and instructed them not to use or further distribute any affected product.The consignees were also instructed to call the firm's Recall Coordinator Silvia Carrillo at 1-800-321-4254 Option #2 Ext 2727 (M-F 7am to 4pm MST) or email at silvia.carrillo@crbard.com to facilitate the expedient return of the affected products. Additionally they were asked to fax the Recall and Effectiveness Check Form to BPV at 1-800-994-6772. |  | 72 | 72 units | Nationwide | Terminated | February 01, 2012 | 638 | 3279
Z-2120-2010 | 55367 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Howmedica Osteonics Corp., |  |  | April 05, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | One of the two posts on the Anterior Resection Guide could disassociate from the main body of the part which has been determined to be caused by a fracture weld. | Device Design | Stryker Urgent Product Recall letters dated April 5, 2010, were sent out via Federal Express to Stryker branches/agencies, Hospital Risk Management, hospital Chief of Orthopedics and Surgeons. All hospitals and surgeons that used the product in the last year were notified. Questions and/or concerns should be directed to Rich Wolyn at 201-972-2100. |  | 1974 | 682 units in US, 1292 units internationally. | Worldwide distribution: USA, Ireland, Germany, Sweden, The Netherlands, France, Italy, Switzerland, South Africa, United Kingdom, Japan, Argentina, Mexico, Miramar (city), China, Taiwan, India, Australia, Korea, New Zealand, Malaysia, Singapore, and Canada. | Terminated | August 28, 2012 | 876 | 1974
Z-2116-2010 | 55874 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | PowerPark is an accessory to the V or H Universal Stand and comprised of two parts: the stand mounted assembly (P12822) and a docking assembly (P12834). The stand mounted assembly is mounted to the V or H Universal stand and the docking assembly is plugged into a wall circuit. It provides a docking and charging station for the V and H Universal Stands. |  |  | May 27, 2010 | 2010 | Sonosite, Inc. | PowerPark system has a potential for loss of power at the corresponding wall outlet and/or circuit due to poor alignment. | Nonconforming Material/Component | On 05/27/10, SonoSite began contacting their consignees by telephone. On 06/01/10, SonoSite sent out the URGENT-MEDICAL DEVICE RECALL letter to their consignees. The letter describes the recall product as SonoSite PowerPark. Customers should call the Technical Support at 1-877-657-8118 or e-mail to service@sonosite.com for any questions regarding the recall.The letter instructs consignees to stop using the PowerPark immediately and return the product back to SonoSite in exchange for a full refund. The consignees will be contacted by the firm's Sales Representative and/or Clinical Applications Specialist to schedule an appointment to retrieve the device and deliver a refund. |  | 26 | 26 | Worldwide Distribution -- USA, including states of GA, NM, PA, TX, VA, and WA and country of Australia. | Terminated | April 18, 2011 | 326 | 26
Z-2151-2010 | 56037 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Procedure Kit Custom |  |  | June 08, 2010 | 2010 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid | Equipment maintenance | Terumo, Ashland, MA notified consignees on 6/8/10, byphone, fax, email and mail to advise distributors and end-users of the required recall actions. Follow-up communication sent by priority overnightmail on 6/9/10, with proof of delivery, specifying the awareness to the issue and necessary correctionactivities should a unit be discovered. Phase 2 initiated on 7/1/10 to sale rep to remove recalled product and perform a product reconciliation to ensure that all affected product is identified, returned to Terumo CVS, and that replacement product is sent immediately. |  | 48 | 48 packs | Distribution USA and Canada | Terminated | May 07, 2012 | 699 | 589
Z-2121-2010 | 56063 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MEDFUSIONÂż SYRINGE INFUSION PUMP, MODELS 3010 AND 3010a WITH SOFTWARE VERSIONS 2.0.2, 2.0.3, AND 2.0.4 MedfusionÂż Syringe Infusion Pumps are designed for use in critical care areas for the administration of fluids requiring precisely controlled infusion rates. |  |  | June 22, 2010 | 2010 | Smiths Medical ASD, Inc. | Smiths Medical is conducting a correction of MedfusionÂż Syringe Infusion Pumps, Models 3010 and 3010a, with obsolete software versions 2.0.2, 2.0.3, and 2.0.4. Smiths Medical has found that pumps with obsolete software may continue to run past the set volume limit. This can result in an over-delivery  if the syringe is over-filled. Smiths Medical has received no reports of serious patient in | Software design | Consignees were sent on 6/22/10 a Smith Medical "Urgent Medical Device Correction Notice', dated June 11, 2010. The letter was addressed to Dear Risk/ Safety Managers, Distributors, Clinical/ Biomedical Engineering, Clinicians, Critical Care Physicians, Pediatricians, Pediatric Intensivists, Neonatologists; and Nursing, Pharmacy,and Anesthesia Professionals. The letter described the problem and product involved. It provided Advise on Action to be Taken by the User which included to 1) Determine the software version of their Medfusion pump 2. Complete and return the Confirmation Form 3. Return the completed Software Update Test Form 4. If unable or unwilling to perform the software upgrade they provided information for returning the pumps to Smiths Medical for the upgrade. For further information please contact Smiths Medical at 651-628-7280. |  | 4816 | 4816 | Worldwide Distribution: AL, AK, AZ, AR, CA,CO, CT, DC, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, ME, MD, MI, MN, MS, MO, MT, NE, NV,NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI. and countries of CANADA, MEXICO, NEW ZEALAND, BRASIL, HONG KONG, and GREAT BRITAIN. | Terminated | April 17, 2012 | 665 | 4816
Z-2172-2010 | 55871 | No | NRP | Sirolimus test system. | Toxicology | Toxicology | 510(k) | Architect Sirolimus Reagent Pack; 100 test pack containing 1 bottle (8.0 mL) Anti-Sirolimus Coated Microparticles, 1 bottle (8.0 mL) Sirolimus Acridinium-labeled Conjugate in Citrate Buffer, and 1 bottle (10 mL) Assay Diluent containing saline; Abbott Laboratories, Abbott Diagnostics Division, Abbott Park, IL 60064 USA; list 01L76-25. |  |  | May 27, 2010 | 2010 | Abbott Laboratories | The Architect Sirolimus assay lot 80162M100 is exhibiting higher than normal reports of barcode read errors. | Equipment maintenance | Abbott sent a "Product Recall" letter dated May 27, 2010, to all customers receiving the affected lot of reagent. The customers were informed of the increased frequency of barcode read errors due to the print quality of the barcode label. The following error codes have been observed for some kits within the lot: 4000 - Unable to read reagent bar code in position (x) on (y) carousel, 0208 - Reagent carousel scan error, and 2011 - Extra reagent bottle detected in position (x) on (y) carousel. The customers were requested to determine if they are currently using or have inventory of lot 80162M100, discontinue use of and destroy any remaining inventory of the lot, and retain a copy of the letter for their laboratory records. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. Customers outside the U.S., please contact your local Customer Service representative. |  | 875 | 875 kits | Worldwide distribution: USA including Puerto Rico, and countries of Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, New Zealand, Pakistan, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Taiwan, Thailand, Turkey, United Kingdom, Venezuela and Vietnam. | Terminated | July 15, 2011 | 414 | 875
Z-2166-2010 | 55507 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | The Well@Home Monitor by Patient Care technologies |  |  | February 02, 2009 | 2009 | Zoe Medical Incorporated | Touch panel may become unresponsive. | Component change control | Zoe Medical notified the 2 distributors by e-mail on January 26, 2009 and Feb 6, 2009 of the recall, requesting return of the units for correction. |  | 62 | 62 | Nationwide Distribution -- MA and MD. | Terminated | August 03, 2010 | 547 | 62
Z-2171-2010 | 55557 | No | NRP | Sirolimus test system. | Toxicology | Toxicology | 510(k) | IMx Sirolimus Reagent Pack; 100 test pack containing 1 bottle (9.7 mL) Anti-Sirolimus Antibody Coated Microparticles, 1 bottle (9.7 mL) Sirolimus Alkaline Phosphatase Conjugate, and 1 bottle (10 mL) 4-Methylumbelliferyl Phosphate, 1.2 mM; Produced by Axis-Shield Diagnostics, Ltd., Dundee, UK for Abbott Diagnostics Division, Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA; list 5C91-21 |  |  | April 30, 2010 | 2010 | Abbott Laboratories | IMx Sirolimus reagent lot 802873106 has exhibited an increased frequency of calibration errors. | Other | Abbott sent a "Product Recall" letter dated April 30, 2010, to all customers receiving the affected lot of reagent. The accounts were informed of the increased frequency of calibration errors, including: 157 Check 1 Out of Range, 161 Check 5 Out of Range, and 173 Mode 1 deviation Factor High. The accounts were requested to determine if they are currently using or have inventory of lot 802873106, discontinue use of and destroy any remaining inventory of the lot, and retain a copy of the letter for their laboratory records. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. Customers outside of the U.S., please contact your local Customer Service Representative. |  | 466 | 466 kits | Worldwide distribution: USA and countries including Puerto Rico, Argentina, Austria, Bahrain, Brazil, Canada, Chile, China, Croatia, Cyprus, Egypt, France, Germany, Greece, Hungary, India, Israel, Italy, Japan, Jordan, Lebanon, Lithuania, Netherlands, Norway, Philippines, Poland, Russia, Panama, Saudi Arabia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE and Venezuela. | Terminated | May 13, 2011 | 378 | 466
Z-2170-2010 | 56018 | bed, manual | FNJ | Manual adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom Basic Care Bed, P1441A (Non-electric), Hill Rom Batesville, IN. |  |  | September 16, 2009 | 2009 | Hill-Rom, Inc. | The flexing of the junction of the sleep deck's head and mid sections can create excess pressure on the pivot pin E-ring, which could cause the E-ring to come off. | Device Design | The firm sent an Urgent Field Safety Notice to its customers dated 9/16/2009. The letter identified the affected product, explained the reason for the recall, actions to be taken, and a contact reference. Customers are to locate the affected product, examine the pivot points and ensure the E-ring is present. If the E-ring is missing, then the affected bed is to be taken out of service. Customers' distributors or agents will contact them with a permanent fix. Questions should be directed towards Hill-Rom Technical Support. |  | 962 | 962 beds | Worldwide Distribution -- Argentina, Bahrain, Bhutan, Brazil, Bulgaria, Chili, China, Columbia, Costa Rica, Ecuador, Egypt, Eritrea, France, Hong Kong, India, Indonesia, Iran, Iraq, Jordan, Kenya, Kuwait, Lebanon, Libya, Malaysia, Mexico, Morocco, Oman, Pakistan, Palestine, Panama, Papua New Guinea, Paraguay, Peru, Philippines, Poland, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Syria, Taiwan, Thailand, Trinidad &amp; Tobago, Turkey, United Arab Emirates, United States, Uruguay, Venezuela, and Vietnam. | Terminated | March 16, 2012 | 912 | 11056
Z-2165-2010 | 56055 | No | GWJ | Evoked response auditory stimulator. | Neurology | Neurology | 510(k) | SonaMed Clarity System II; evoked response auditory stimulator; SonaMed Corp., Waltham, MA 02154 |  |  | June 04, 2010 | 2010 | Natus Medical Inc | There is insufficient evidence to support any claims for safety and efficacy on the SonaMed Clarity devices. | PMA | Natus sent Urgent Medical Device Recall letters dated 6/4/10 to all known SonaMed Clarity device accounts, to the attention of the Director of Maternal Child Health and the Risk Manager. The letters referenced the letters previously sent to them by Natus in October 2008 and April 2010 informing them of Natus' concerns about whether the Clarity device could be lawfully used to comply with FDA-related requirements and other pertinent professional guidelines/standards for newborn infant hearing screening, and that they concluded that the Clarity devices should no longer be used. Natus recognized the need for the accounts to continue screening without interruption, and are providing on loan at no-charge handheld screener(s) to facilities that do not have alternate screening devices. The accounts were requested to contact Natus Technical Service group at 800-272-8075 to discuss this loaner device option. Natus requested that the accounts stop using the SonaMed Clarity devices; disable the screening function of their Clarity system(s) by returning to Natus, in the enclosed pre-addressed return envelope, either of the two probes (TPI-830- 8-pin ear probe, or the TPI-815- 7-pin ear probe) they may possess and have used in screening, making sure to return all 7- or 8-pin probes for all of your Clarity systems; and complete the attached "Device Recall Reply Form" and return it in the enclosed self-addressed stamped envelope. If you have any questions concerning any aspect of this recall notice, please call Natus Technical Support at 1-800-272-8075. |  | 284 | 284 units | Nationwide Distribution | Terminated | June 22, 2011 | 383 | 284
Z-2173-2010 | 55775 | No | KWH | Toe joint polymer constrained prosthesis. | Orthopedic | Orthopedic | 510(k) | Sgarlato Labs GAIT (Great Toe Implant Technique) Implant-- Large, Product Number 30-100-03 Large, Manufactured by Sgarlato Med LLC, San Jose, CA |  |  | March 01, 2009 | 2009 | Sgarlato Med, LLC | Products with an expiration date of August, 2009 were labeled with an expiration date of May 2013. Additionally, the outer packages were labeled with an incorrect lot number. The outer and inner packages did not match. | Labeling mix-ups | Recall initiated in March 2009 via telephone calls. |  | 27 | 27 units | Nationwide Distribution | Terminated | August 19, 2011 | 901 | 27
Z-2174-2010 | 55813 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD GeneOhm Cdiff Assay, Catalog 441401, 200 tests/kit, labeled, in part ***BD Diagnostics, 2555 Boul. Du Parc-Technologique, Quebec, Canada G1P4S5*** |  |  | March 05, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic test kits do not perform to specifications. | Nonconforming Material/Component | The recalling firm notified accounts by an Urgent Product Recall letter on 03/05/2020. The letter identified the affected product and explained the reason for recall. It also advised customers to discontinue use of the affected product and discard any remaining inventory and fax back an acknowledgement form. The letter stated that BD will provide replacements for any discarded product. Further assistance can be obtained by contacting BD Technical Service at 1-888-436-3646. |  | 31 | 31 kits | Nationwide Distribution -- IL, IN, PA, WA, CA, VA, NJ, and KS. | Terminated | October 15, 2010 | 224 | 31
Z-2175-2010 | 55908 | No | GKH | Automated blood cell diluting apparatus. | Hematology | Hematology | 510(K) Exempt | CellTracks Auto Prep System, Circulating Tumor Cell Kit (epithelial, IVD |  |  | May 19, 2010 | 2010 | Veridex, LLC | Complaints of patient sample carryover. | Process control | On 19 May 2010, all consignees were notified, US and foreign, via Urgent Product Correction Notification letters. The letter identified the affected product and explained the reason for recall. Customers were asked to follow the required actions provided in the letter, including completing and returning the attached Confirmation of Receipt form. Questions should be directed to Customer Technical Services at 1-877-837-4339. |  | 188 | 101 Domestically, 87 Internationally | Worldwide distribution -- USA, China, Japan, England, France, Germany, Italy, Spain and Canada. | Terminated | May 08, 2014 | 1450 | 188
Z-2177-2010 | 56049 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Straumann Impression set for WN solid abutment H 4.00mm POM/ AI. Ref : 040.380 Kit Y4021. |  |  | June 15, 2010 | 2010 | Straumann Manufacturing Inc. | The label on the tubular packaging for the Straumann Impression set f. WN solid abutment has the incorrect description of RN solid abutment instead of a WN. | Labeling mix-ups | Field Safety Notice, dated June 15, 2010, issued to consignees states that the 040.380 Impression set f. WN solid abutment Lot Y4021 which you purchased is incorrectly labeled Impression set f. RN solid abutment. Customers were asked to return the affected products and completed Customer Confirmation Form for replacements. Questions should be directed towards Bernard McDonald at 978-747-2514. |  | 163 | 163 | Nationwide Distribution | Terminated | September 22, 2011 | 464 | 163
Z-2191-2010 | 49374 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | SonoSite TEE, 8-3 MHz Transducer labeled: "Tee Transducer... SonoSite, Inc. 21919 30th Drive SE Bothell, WA 98021". |  |  | August 08, 2008 | 2008 | Sonosite, Inc. | The TEE transesophageal transducer, when connected to a SonoSite M-Turbo Ultrasound System, may exhibit erratic operation of the Scanplane Orientation Indicator on the system display. The indicator may move erratically and not accurately represent the orientation of the TEE transducer array. | Other | On August 21, 2008, SonoSite began calling 15 consignees who have purchased the TEE/8-3 MHz Transducer and an M-Turbo 1.1 system. The firm provided instruction to discontinue use the system to 11 consignees and left voice messages to the remaining four consignees.On August 21, 2008, SonoSite sent out the URGENT-MEDICAL DEVICE CORRECTION letter and the Urgent Field Safety Notice to all affected consignees in the US and in foreign countries.The URGENT-MEDICAL DEVICE CORRECTION letter describes the recall product as TEE/8-5 MHz Transducer, SonoSite Part Numbers P05183 and P06841. The letter informed the consignees that the TEE transesophageal transducers may exhibit erratic operation of the Scanplane Orientation Indicator on the system display when it connected to the SonoSite M-Turbo Ultrasound System. The letter instructs consignees to discontinue use of the TEE/8-3MHz Transducer with the M-Turbo (or Micromaxx system) immediately, while the firm is working to create a software solution to resolve the issue and offers customers a software upgrade within 6-8 weeks.The consignees are also advised to forward the recall notice to others who need to be aware of the issue or where the potentially affected devices have been distributed.Customers should call the Technical Support at 1-1-425-951-1330 or 1-877-657-8118 from Monday to Friday, 6 am-5 pm PST or e-mail to service@sonosite.com for any questions regarding the recall. |  | 357 | 357 units | Worldwide distribution: USA and countries including Argentina, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Finland, France, Germany, India, Iran, Ireland, Israel, Italy, Korea, Netherlands, New Zealand, Poland, Slovenia, South Africa, Spain, Switzerland, and Thailand. | Terminated | August 24, 2011 | 1111 | 357
Z-2183-2010 | 55139 | No | INK | Powered wheeled stretcher. | Physical Medicine | Physical Medicine | 510(k) | Stryker Atlas Transport Stretcher with Zoom, Model 660Z, Stryker Medical, Portage, MI |  |  | March 19, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The retractable cord stretchers were fitted with power cords, manufactured by Electri-Cord, whose prongs have the potential to fracture inside the molded section of the plug, causing melting and charring. | Component design/selection | The firm has sent Urgent - Medical Device Recall letters, dated March 19, 2010, to all direct accounts notifying them of the recall. The letter identified the affected product and described the reason for recall. Customers were asked to locate the stretchers identified in the letter and visually inspect each cord. If the plug shows signs of excessive wear, etc. the stretcher is to be taken out of service immediately and the plug removed. Customers are to call the number provided in order to be placed on the priority replacement list. If the cord does not show signs of damage then the stretcher can remain in service until a replacement cord is available. The enclosed post card should be completed and returned to the firm. A copy of the notice should be forwarded to any users that stretchers have been sold or loaned to. The firm should be provided with the serial numbers of any stretchers that have been disposed of. Questions should be directed towards Stryker at 800-STRYKER. |  | 61 | 61 stretchers | Worldwide Distribution -- USA, Canada, Brazil, Japan, Singapore, Spain, and the Far East. | Terminated | May 02, 2011 | 409 | 4915
Z-2188-2010 | 56091 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | CLAWÂż HEX SCREW CHARLOTTE " F&A; SYSTEM CONTENTS: 1 EACH Foot and Ankle Screw ... IMPLANT MATERIAL: Stainless Steel Ref: 4013-3520 ... Size 3.5MM, Length 20 MM, Use with ClavaÂż 3.5MM - Non-Sterile. Implantation: used in Food and Ankle surgeries. |  |  | June 07, 2010 | 2010 | Wright Medical Technology Inc | One lot of CLAW Hex Foot and Ankle Screws was incorrectly packaged containing Multi-use Compression Screws. | Packaging process control | The firm initiated their recall to their sales representatives (distributors) via telephone on 06/07/2010 and followed with a letter dated June 15, 2010 sent via Fed Ex on 06/28/2010. The hospital involved was also notified by letter on 06/29/2010. The letter described the product, the problem,and the action to be taken by the customers. The customers were instructed to immediately return any affected product to Wright as soon as possible, and to complete an attached response form confirming receipt of the notice and return a copy by fax to 901-867-7401. Customers may contact customer service for return instructions and replacement inventory at 800-238-7117.If you have any questions regarding this matter, contact Debby Daurer at 800-874-5630. |  | 8 | 8 units | Nationwide Distribution in States of FL, PA, KS, IN, MO, and CA | Terminated | November 23, 2011 | 534 | 8
Z-2190-2010 | 56115 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Corporation, QuickCat Extraction Catheter, REF 60090-01, Sterile R, Working Length 145cm, for embolus and thrombus removal from arterial vessels. |  |  | July 01, 2010 | 2010 | Spectranetics Corporation | Guidewire lumen may be occluded, preventing guidewire insertion into the catheter. | Process control | Consignees were notified by letter on Jul 1, 2010, and told to not use affected lots and to contact Spectranetics to arrange a product exchange. More information is available at 800-231-0978. |  | 5080 | 5080 units (1778 US, 3302 Non-US) | Worldwide distribution: USA, Argentina, Austria, Belgium, Brazil, Chile, Czech Republic, Finland, France, French Polynesia, Germany, Hungary, Israel, Italy, Netherlands, Poland, Spain, Sweden, Switzerland, and Venezuela. | Terminated | August 20, 2010 | 50 | 5080
Z-2192-2010 | 56163 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | REF 87100 87k Arthroscopy Tubing Set. For use with the 87k Arthroscopy Pump (87000). STERILE EO. Rx ONLY. Warning: Do not use this tubing set with any Small Joint Cannula System. |  |  | July 06, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec Largo, FL is recalling 87K Arthroscopy Tubing Set, Product Number 87100, Lots 0904151 to and including 1005201. Product may have a breach in the seal that could potentially compromise the sterility of the contents. | Process control | ConMed Linvatec Corporation sent an "URGENT-MEDICAL DEVICE RECALL NOTIFICATION" letter dated July 6, 2010, via FedEx next day delivery to all customers. The letter described the product, problem and action to be taken by customers within 30 days of receipt of the letter. The customers were asked to PLEASE DISCONTINUE USE OF THE IDENTIFIED PRODUCT WITH LOT NUMBERS BEGINNING WITH 0904151 THROUGH AND INCLUDING 1005201; to immediately check their inventory for the affected lots number and Do Not Use if they have any of the products; to complete and return a Reply Form and to notify their customers if the product was further distributed and return all unused affected product with lot codes starting 0904151 to and including 1005201, to ConMed Corporation, 525 French Road, Utica, NY 13502. On receipt, ConMed Linvatec will issue credit to customers.Should you have any additional questions, please contact our Customer Service department at 800-535-8536 or fax at 727-319-5701 or email: Custerserv1@linvatec.com |  | 97830 | 97,830 | Worldwide distribution: USA and Canada. | Terminated | February 24, 2011 | 233 | 97830
Z-2193-2010 | 56321 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Crossfire Console, model number 475-000-000, Manufactured by Stryker Endoscopy San Jose |  |  | June 28, 2010 | 2010 | Stryker Endoscopy | RF energy or powered shaver handpieces may continue to operate when not intended when used with the iSwitch Wireless Foot switch. | Software design | Stryker Endoscopy sent an "Urgent: Device Correction" letter dated June 28, 2010, via FedEX Mailer to US customers. A letter dated July 9, 2010, was sent to International sites.The letter describes the product, problem, and action to be taken by customers. The Customers were instructed to follow the included instructions and use the USB dongle to install the correction and to complete and sign the enclosed "acknowledgement of receipt" form,upon completion of the software correction, and fax to (480) 754-8378 or scan in and email to crossfireconsole@stryker.com. The firm noted that-It is important to send the self addressed confirmation forms back to Stryker RAQA. If you have any questions regarding this letter, please contact us by phone at 1-800-624-4422 or via email at crossfireconsole@stryker.com. |  | 989 | 989 units | Worldwide distribution: USA and countries including Australia, France, Germany, Greece, South Africa, Iberia, Italy, New Zealand, Poland, UK, Switzerland, Canada, China, Japan and the EMEA. | Terminated | March 21, 2011 | 266 | 989
Z-2194-2010 | 48998 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | The Philips EasyDiagnost Eleva with DRF. Philips Medical Systems multifunctional x-ray system. |  |  | July 09, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Potential for bright artifacts to appear on the patient image from a previous exposure. | Other | On July 9, 2008, Philip Medical System began sending out the URGENT-Device Correction letter to their consignees. This letter describes the recalled product as EasyDiagnost Eleva with DRF system with SW release 1.0.The firm informs the consignees that under special circumstances there is a small possibility that bright artifacts appear on a patient image from a previous exposure.The consignees are advised to take care about the X-ray field limitation. If the detector is not affected by direct radiation, no problem will occur. If direct radiation reaches the detector, the collimation should not be too tight, such that the area of direct radiation is larger than 1.2 inch x 1.2 inch (or 30 mm x 30 mm). The minimum area is needed to detect and prevent memory artifacts.The firm will upgrade any affected products with the Field Change Order (FCO)70600031 and it will be implemented before the end of December 2008. The upgrade software is 1.0.3. The correction would be free of charge and the consignees should contact Philips Call Center at 800-772-9377, #5, #2. |  | 7 | 7 | Product was distributed in AZ, CA, FL and WA. | Terminated | October 27, 2010 | 840 | 7
Z-2195-2010 | 55787 | No | JQX | Nephelometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMAGE Immunochemistry Systems Buffer 1 (BUF1), Part Number: 447650 |  |  | May 21, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall of IMMAGE Immunochemistry Systems Buffer 1 (BUF1) due to received reports of increased recovery of control andpatient samples using the lots of IMMAGE BUF1. Some lots of BUF1 have been reported to cause high shifts in recovery of control or patient samples. Quality Control failures may result after replacing BUF1 with a different lot. | Nonconforming Material/Component | Beckman Coulter will send a Product Corrective Action (PCA) letter on the week of April 19, 2010 to the affected customers. The letter provided the customers with an explanation of the problem and a work around. |  | 1309 | 1309 | Worldwide: U.S., Canada, Australia, Austria, Belgium, China, Czech Republic, Denmark, France,Germany,Greece,Hong Kong, Iceland,Ireland,Italy,Japan,Liechtenstein,Macao,Malaysia,Netherlands,New Zealand,Norway,Oman,Poland,Qatar,Russian Federation,Singapore,Slovakia,South Africa,Spain,Sweden,Switzerland,Taiwan,Tunisia,Turkey andUnited Kingdom, | Terminated | May 22, 2012 | 732 | 1309
Z-2202-2010 | 56092 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Intellivue Patient Monitor Model MP70 (M8007A) |  |  | June 28, 2010 | 2010 | Philips Healthcare Inc. | Additional Instruction for Use for Intellivue MP60/MP70 to prevent monitors from becoming loose and detaching from the mounting arm. | Labeling design | Philips issued on 6/28/10 an Urgent Medical Device Correction notification/Field Safety Notice. Included with the Urgent Medical Device Correction notification/Field Safety Notice is an updated instructions for use and service instructions for these monitors that remind users to follow themounting arm manufacturer's instructions for use when repositioning the monitor and instructing them toperform a periodic check that the monitor is not loose on the arm. |  | 82019 | 82,019 units | Worldwide distribution: USA, Canada, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, Bosnia, BRAZIL, BULGARIA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRI, LITHUANIA, LUXEMBOURG, MACAU, MALAYSIA, MAURITIUS, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA &amp; MONTE NEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, United Arab Emirates, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, VIRGIN ISLANDS (BR), and YEMEN. | Completed |  |  | 93548
Z-2204-2010 | 55909 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Vanguard CR femoral 57.5 mm left interlok, Sterile, REF 183022, Knee prosthesis component. |  |  | April 20, 2010 | 2010 | Biomet, Inc. | A knee implant labeled as left was opened and the package actually contained a right knee implant. | Process control | On 4/13, 14 and 15 Biomet made phone calls to the consignees to inform them of the recall. The firm sent Urgent Medical Device Recall Notices dated 4/20/2010. The customers were instructed to immediately locate the recalled products and return them to the manufacturer. |  | 5 | 5 | 1 unit in US. 3 units in Biomet distributor inventory, 1 unit in Japan | Terminated | January 19, 2011 | 274 | 5
Z-2200-2010 | 56117 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | International Technidyne Corporation (ITC) directCheck Whole Blood Quality Controls Plasma, Coagulation Control. Catalog number: DCJLR-N; in vitro diagnostic |  |  | June 25, 2010 | 2010 | International Technidyne Corp. | The package insert for the directCheck Microcoagulation Quality Control contains an incorrect acceptable performance range regarding ACT-LR cuvettes | Nonconforming Material/Component | ITC sent Urgent Medical Device Correction letters on June 25, 2010 by Federal Express to hospital Risk Managers and Lab Directors, Clinical Laboratory Manager or Point-of-Care Coordinators. |  | 122 | 122 boxes in US, 1 out of US. | Nationwide and one customer in Belgium. | Terminated | September 23, 2010 | 90 | 122
Z-2196-2010 | 56157 | No | NMF | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | ASCENT HEALTHCARE SOLUTIONS. Reprocessed Femoral Compression Device, Model 11165. |  |  | June 23, 2010 | 2010 | Ascent Healthcare Solutions, Inc. | The FemoStop Devices may fail to inflate or hold pressure. A separation between the dome and arch base prevents the device from maintaining pressure when inflated. | Reprocessing Controls | Ascent sent an urgent medical device voluntary recall notification Letter and Response Form to each consignee on June 25, 2010. Consignees were requested to immediately examine their inventory of Ascent Reprocessed FemoStop devices with lot numbers located on an attached list; remove and quarantine all affected devices; complete the attached Recall Effectiveness Check Form; ship all affected product back to Ascent using the address indicated in the letter, and fax a completed Recall Effectiveness Check Form to 480-763-5350. |  | 175 | 175 devices | Devices were distributed to 8 consignees (end users/hospitals) located in VA, CA, FL, MI, WA. | Terminated | August 11, 2010 | 49 | 175
Z-2197-2010 | 56187 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | CERCLAGE WIRE, 0.6 MM DIA. X 175 MM LENGTH, SST, REF 71163006, QTY: (6), STERILE R, CE 008, 2020/07, Smith &amp; Nephew, Inc., Memphis, TN 38116, EC REP: Smith &amp; Nephew Orthopaedics GmbH, Tuttlingen, Germany |  |  | June 28, 2010 | 2010 | Smith &amp; Nephew Inc | Plastic trays which encase these products have burrs on the hinged packaging which can puncture the single barrier sterile pouch, compromising sterility. | Process control | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall was sent via overnight mail on 01/23/2009. All affected Smith &amp; Nephew International Distributors were notified via email and telephone on 01/23/2009. They were instructed to immediately quarantine the product for return. |  | 1 | 1 unit | KY, TX, DC, NC, CA, AZ, GA and OH; 3 International: Australia, Malaysia, and Switzerland | Terminated | January 23, 2012 | 574 | 43
Z-2205-2010 | 56201 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Pulmonetic Systems P/N 19189-001 Patient Circuit SPU W/O PEEP 15 FT 5/Pkg. |  |  | July 02, 2010 | 2010 | CareFusion 203, Inc. | CareFusion is recalling the P/N 19189-001 Patient Circuit (SPU W/O PEEP 15ft) because some of the circuit sense lines may have been reversed during the manufacturing process. If the sense lines have been reversed, immediately upon connection of the patient circuit to the ventilator and turning the ventilator on, the ventilator will begin to autocycle (with increased pressure delivered) with both | Process control | Customers were sent a CareFusion "Urgent Medical Device Recall Notification" dated July 2, 1010. The letter was addressed to Respiratory Care Manger. The letter described the product , problem and advised consignees to remove the affected product from inventory and promptly return it to CareFusion for replacement.Please contact CareFusion Technical Support at 800-754-1914, option 2, upon receipt of this notification to arrange for prompt return and replacement of any the patient circuits in your inventory. |  | 200 | 200 boxes (5 each) | Nationwide distribution: AL, CA, FL, GA, HI, IL, IN, KS, MA,MI, MN, MT, NJ, NY, NC, OR, PA, RI, TN, TX, UT, VA, WA, and WI. | Terminated | December 17, 2011 | 533 | 200
Z-2210-2010 | 48992 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Motorized Height Ceiling Suspensions (MCS) and motorized height movement (actuator) for X-ray systems. |  |  | June 26, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Monitor ceiling carriage may separate from the monitor ceiling suspension. | Other | On 07/02/08, the firm sent a "Urgent - Device Correction" letter to all their consignees. The letter describes a product as X-Ray system with Monitor Ceiling Suspension and motorized height movement. The consignees are informed that the fixation bolts supporting the monitor ceiling carriage may come loose creating a situation which may allow the monitor ceiling carriage to separate from motorized height Monitor Ceiling Suspension and that would cause serious injury to patient and user.The firm has issued a mandatory Field Correction Action to correct the problem and the firm's technicians will inspect all ceiling suspension actuators for loose screws. If loose screws are found the actuator will be replaced.Consignees should call Philips Call Center at 800-722-9377 #5, #4, #1, and Reference FCO72200126. |  | 1111 | 1111 units | United States and Canada. | Terminated | December 02, 2010 | 889 | 1111
Z-2206-2010 | 54370 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | BIOMET MICROFIXATION "1.5MM" SYSTEM HIGH TORQUE (HT) CROSS-DRIVE SCREW, QTY 1, 1.5 X 4MM TITANIUM. REF 91-1504 1 COUNT and REF 95-1504 - 5 PACK. |  |  | March 17, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation has received a complaint from a customer who reports finding a 3.5mm length screw in one 5 pack of 4mm length screws. | Other | Biomet Microfixation is recalling their 1.5x4mm HT X-drive Screws Lot Number 148530 after receiving a customer reported finding one 3.5MM screw in a package labeled as containing 4.0MM screws. Lot 148530 was packaged into one count packages 91-1504 and 5 count packages 95-1504. The 1 and 5 count packages were distributed between 12/10/2009 and 02/01/2010. Beginning on 03/17/2010 customers identified as receiving Lot 148530 HT X-drive Screws were notified of the recall by a combination of an Urgent Medical Device Recall Notice letter, e-mail, fax and telephone calls. Customers were instructed to inspect their inventory and to return any affected products. They were instructed to complete the attached Inventory Reconciliation sheet and to fax it to Biomet Microfixation. Customers were also instructed to contact any customers to whom they may have further distributed the product. Questions should be directed to Christy Cain at 1-800-874-7711 or 904-741-4400, ext. 468. |  | 1247 | 1247 screws | Worldwide Distribution -- USA, including states of FL, GA, IN, LA, MD, MI, MO, NC, NY, OK, TX, and VA and countries of Philippines, Japan, Brazil, and Canada. | Terminated | August 11, 2010 | 147 | 1247
Z-2208-2010 | 55678 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems AFP QC, Part Number: 33219 |  |  | April 21, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed an issue with the Access AFP QC kits (REF 33219) which contain three levels of control material identified as QC1, QC2, and QC3, Beckman Coulter has confirmed that vials of QC1 in the kit lots identified above contain microbial contamination. Customers may see that the results from a contaminated vial of QC1 are within the established refere | Other | A Product Corrective Action (PCA) letter was senton the week of April 19, 2010 to all customers whopurchased the Access Immunoassay Systems AFPQC material, lot numbers referenced above. Theletter provided the customers with an explanation ofthe problem identified and instructed the customersto (1) Please discard all remaining inventory of the three Access AFP QC kit lots identified above,including all vials of QC1, QC2, and QC3.(2) Please review your ac and patient results obtained while using the affected control lots. Interpret Access AFP test results in light of the clinical presentation of the patient, including: symptoms,clinical history, data from other tests and other appropriate information.(3) Please share this information with your laboratory staff and retain this notification as part of yourlaboratory Quality System documentation. (4) Complete and return the enclosed response formwithin 10 days so we can be assured that you have received this important notification.For customers that need assistance or have any questions regarding this notification, they were instructed to contact Technical Support at 1-800-854-3633 in the United States and Canada. |  | 166 | 166 | Nationwide and Canada | Terminated | May 07, 2012 | 747 | 166
Z-2207-2010 | 56031 | No | NQK |  |  | General &amp; Plastic Surgery | 510(k) | Monarch Labs Medical Maggots with the accessory LeFlap dressing (previously known as Creature Comforts II), 250-500 per vial, Batch #: MM-090406 with CCII-0901. |  |  | May 03, 2009 | 2009 | Monarch Labs, LLC | Monarch Labs initiated the field notification regarding their Medical Maggots from Batch MM090406 with the accessory LeFlap dressing due a customer complaint from a physician of a patient developing cellulitis after using maggots from the affected batch. | Other | The customer recall communication was initiated on May 3, 2009 with Monarch Labs notifying all consignees with a via a faxed Recall Letter with fax-back attachment of the initial compliant they had received and the possible infection that may occur if instructions for use are not followed. All therapists, who used maggots from that batch were instructed to contact the firm immediately if they have experienced any adverse events, such as infection, cellulitis, or any other problem. If not, they were asked to simply return the fax-back attachment , noting "no problem" . |  | 14 | 14 vials | Nationwide Distribution -- CA, IA, KY, IL, TX, SC &amp; UT. | Terminated | December 13, 2012 | 1320 | 14
Z-2209-2010 | 56118 | Yes | NVZ |  |  | Cardiovascular | PMA | Medtronic Carelink Monitor Model 2490C. The Medtronic CareLink Monitor Model 2490C is designed to automatically gather information from the implanted heart device. This automatic wireless communication between the heart device and the monitor takes place at times scheduled by the doctor or clinic. The gathered heart device information is sent over a telephone line to the clinic using a toll-free number. The monitor can also send wireless heart device alerts to the clinic automatically. |  |  | June 11, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic has identified that a subset of 2490C CareLink Monitors recently received an incorrect software update. Patients reported that their monitor was making a "ticking sound" and they were unable to transmit device data via CareLink. These monitors are now non-functional. No patient injuries have been reported as a result of this issue. | Software change control | Starting on 6/11/2010, Medtronic began sending Australian and Canadian consignees "Medical Device Recall" letters dated 6/10/2010. The letters describe the problem and product involved. The letters instructs consignees to order replacement monitors, notify affected patients, and instruct each patient to perform a manual transmission upon receipt of the new monitor. Medtronic will provide affected patients with a prepaid return kit to faciliate the return of the monitor to Medtronic.Australian consignees should contact Scott Dunlop at 61-7 3025-3547. Canadian consignees should contact Medtronic Patient Services at 1-800-268-5346. |  | 37 | 37 Foreign, 0 Domestic | International Distribution Only: Australia and Canada. | Terminated | December 24, 2011 | 561 | 37
Z-2212-2010 | 56131 | No | JDL | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | DYNASTYÂż A-CLASSÂż POLY LINER, REF DLXP-LD36, 1 EACH, Rx ONLY, STERILE EO, GROUP D, I.D. 36 mm, LIP 15o, LINER STD, Wright Medical Technology, 4677 Airline Road, Arlington, TN 38002. Product is used in total hip arthroplasty. |  |  | June 07, 2010 | 2010 | Wright Medical Technology Inc | The inner and outer product packaging was missing the following information: shelf life, translations, manufacturing date, CE marking and also listed the incorrect sterilization method. | Software design (manufacturing process) | On 06/28/2010, Wright Medical Technology sent a letter to consignees. Consignees were instructed to examine their inventory for recalled product and return it to the firm. Consignees should contact customer service at 800-238-7117 for return instructions and replacement inventory. Questions concerning the recall should be directed to Debby Daurer at 800-874-5630. |  | 6 | 6 units | USA distirbution only, in states of WI, WV, IN, IA, and GA | Terminated | May 31, 2011 | 358 | 6
Z-2216-2010 | 55364 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trabecular Metal acetabular revision system column buttress left posterior/right anterior, 00-4894-300-00, sterile, Zimmer Inc., Parsippany, NJ. |  |  | April 06, 2010 | 2010 | Zimmer Inc. | The firm has determined that some of the packages could develop a breach of the seal in the sterile barrier pouches. | Packaging | The firm sent an "Urgent Device Removal" letter dated 4/6/2010. The consignees were instructed to inventory the products on hand and return the recalled product to the manufacturer. More information is available at 1-800-613-6131. |  | 1972 | 1,972 all products | US, Germany, Brazil, Singapore, Canada, Australia, England, India, and Switzerland. | Terminated | October 21, 2011 | 563 | 1972
Z-2211-2010 | 55502 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DATEX-OHMEDA NETWORK AND ICENTRAL '05, SALES REVISION 4.3 ( Datek-Ohmeda Network and iCentral (including iCentral Client)) The Datex-Ohmeda Network and iCentral (including iCentral Client) transfers information between networked Datex-Ohmeda devices in the Monitor Network as well as between Monitor Network and networked devices in GE Unity NetworkTM. It also allows information transfer between several iCentrals. Within one Monitor Network it allows a networked device to display, store, print and otherwise process information received from other networked devices. The iCentral maintains the network connections between the Datex-Ohmeda bedside monitors and other networked devices in Monitor Network. Network connections consist of hardwired network cables and/or Wireless LAN (WLAN) connections. Furthermore, it coordinates the transfer of information between devices in the Datex-Ohmeda Network, between the Datex-Ohmeda Network and devices in GE Unity Network as well as between the Datex-Ohmeda Network and Hospital Information Systems (HIS). The iCentral can be used for remote monitor management, storing, printing, viewing, reviewing or otherwise processing of information from several bedside monitors or other networked devices. The iCentral Client can be used for remote monitor management, printing, viewing, reviewing or otherwise processing of information from several bedside monitors or other networked devices. The Datex-Ohmeda Network will be used for patients in the hospital and it is meant for continuous use. The device is for use by qualified personnel only. |  |  | June 25, 2010 | 2010 | Ge Healthcare, Llc | The iCentral or iCentral client v5.0.2 with HP xw 4400 and 4600 may spontaneously shutdown, causing monitoring loss at the central monitoring station. Monitoring loss may result in prolonged delay of treatment of potentially lethal arrhythmia. | Software design | GE Healthcare sent a "Urgent Medical Device Correction letter" dated 6/18/10 on 6/25/10 to consignees. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. The letter instructs consignees to follow specific detailed instructions for safe use of the device pending provision of a software upgrade by the firm. Consignees will be contacted by a GE Healthcare representative to schedule the software upgrade. |  | 2962 | 2962 | AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MI, MN, MS, MO, NE, NH, NH, NY, NC, ND, OH, OK PA, RI, SC, TN, TX VT, VA, WA, WI.ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BARBADOS, BELGIUM, BERMUDA, BOLIVIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, GUYANA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA , KUWAIT, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, NAMIBIA, NETHERLANDS, NEW ZEALAND,NICARAGUA, NIGERIA, NORWAY, PHILIPPINES, POLAND, PORTUGAL QATAR, RUSSIAN FEDERATION SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN THAILAND TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN. | Terminated | August 14, 2012 | 781 | 2962
Z-2213-2010 | 56189 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DSD Application Guide (MG04-0001 revision M) for the Pentax Universal Video Scope Hookup (DSD-110-HU0109) only when used with the DSD automated endoscope reprocessor (AER) and the Pentax EG-3630U endoscope. |  |  | June 15, 2010 | 2010 | Minntech Corp | The DSD hook up Application Guide incorrectly recommends use of the Medivators DSD-110-HU0109 endoscope hookup to connect the Pentax EG-3630U Ultrasound Endoscope to a Medivators DSD endoscope reprocessor or Medivators Scope Buddy Endoscope Flushing Aid. This may result in inadequate disinfection during reprocessing, and subsequent infection risk to the patient. The recall was due to the misprinte | Labeling False and Misleading | Consignees were sent a Medivators Reprocessing Systems "Urgent Medical Device Recall" letter date June 17, 2010. The letter was addressed to "Dear Medivators Customers". The letter described the problem and product involved. Recommended consignees to immediately examine the Pentax EG-3630U Ultrasound Endoscope hookups and inform all staff of the notice to ensure that the DSD-11--HU01109 is not used with the Pentax endoscope.Also requested consignees to complete and return the enclosed response form. |  | 185 | 185 | AZ, AR, CA, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MD, MA, MI, MN, MS, MO, NH, NJ, NY, NC, ND, OH, PA, SC, TN, TX, UT, VA, WA, WV, WI, WY. | Terminated | December 17, 2011 | 550 | 185
Z-2219-2010 | 56166 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Hi-Art System, H-0000-0003 |  |  | December 01, 2009 | 2009 | TomoTherapy Incorporated | An issue was identified with the TomoTherapy HI-Art System. In the event a patient or DQA plan has a moved image, roll adjustments applied during registration will be incorrect. The Planning Station Plan Settings and DQA Setup tabs allow for images to be moved during planning. During registration when roll is applied on moved images, the Operator Station incorrectly rolls the image about | Software design | Consignee was sent a TomoTherapy "Urgent Medical device Correction Field Safety Notice' dated December 01, 2009. The letter described the Issue, Product Affected, Recommended Action and Resolution. |  | 1 | 1 | VA | Terminated | December 17, 2011 | 746 | 1
Z-2180-2010 | 56222 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical, Engage TR Introducer, 6 F - ACT (2.25 mm), 25 cm length, .035" Max Guidewire O.D., Rx, Sterile EO, REF C408513, 100017549. |  |  | June 28, 2010 | 2010 | St. Jude Medical Cardiovascular Division | Potential for a partial or complete separation of the Shaft (sheath) from the hub or a material break in the hub assembly just distal to the hemostasis seal. If either of these were to occur during use, it is likely that fluids would leak around the Introducer hub and strain relief. | Process design | Consignees were sent "Urgent Medical Device Recall Notice" dated June 24, 2010 from St. Jude Medical . The letter was addressed to Cath Lab Manger/ Risk Manager. The letter described the problem and the product involved. Advised consignees to discontinue use of the product and remove from inventory. Sales Representatives were also sent a St. Jude Medical Memorandum "Field Action Notice" dated June 24, 2010 and a reconciliation form. The memorandum described the problem and product. Advised Sales Rep to visit the affected accounts and to performed the assigned tasks.A St. Jude Medical sales representative will be contacting for removal and replacement of the affected product. For additional information, please contact St. Jude Medical at 651-491-4470. |  | 15 | 5 US, 10 OUS | Worldwide Distribution including USA, Italy, Denmark, France, Switzerland, England, Finland, Austria, Spain and Estonia. | Terminated | December 27, 2011 | 547 | 5120
Z-2220-2010 | 56246 | No | KYG | Ocular surgery irrigation device. | Ophthalmic | Ophthalmic | 510(K) Exempt | CORNEA COAT Hydroxypropylmethylcellulose 1.5ml Syringe. STERILE, Rx ONLY. |  |  | May 12, 2010 | 2010 | Insight Instruments, Inc. | Insight Instruments, Inc., Stuart, FL is recalling CORNEA COAT 2.5% Hydroxypropylmethylcellulose solution in 1.5 ml Syringe due to a potential weakness of the peel pouch seals. Products affected are Item Numbers 600100.000 sold individually and 40201.000 thru 10206.000 sold as components of BIOM Pack. | Packaging | All consignees were notified of the recall by letter sent via first class mail. Letter contained a Return Response Form. |  | 16128 | 16,128 | Worldwide Distribution -- USA, including Puerto Rico, and Frankfurt, Germany. | Terminated | February 16, 2012 | 645 | 66853
Z-1634-2010 | 55077 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCIÂż AQUA+Âż Flex, Hygroscopic Condenser Humidifier, HCH Tidal Volume Range 300-1500mL, ICU/Anesthesia Adult/Child, Catalog number 1570, Distributed by Teleflex Medical. The Aqua+ Flex is a passive humidifier indicated for use on patients under anesthesia and intensive care to effectively warm and humidify the inspired air. |  |  | March 19, 2010 | 2010 | Teleflex Medical | The patient end of the 22 cm connector on the flex tube may not fit securely within the endotracheal tube (ET) connector and therefore disconnect from the endotracheal tube. A disconnect in ventilator dependent patients without prompt response to the alarm could lead to hypoxia, organ failure, or cardio-respiratory arrest. | Device Design | Teleflex Medical sent an Urgent Medical Device Recall letter on March 22, 2010 to consignees. They were instructed to: 1) Immediately discontinue use and quarantine any products labeled as AQUA+ FLEX, Catalog number 1570 with the lot numbers provided. 2) Complete and fax the Recall Acknowledgment Form to 866-804-9881. 3) After contact by a Customer Service Representative, who will provide them with a Return Goods Authorization Number, return the product to Teleflex Medical for credit. Letters addressed to distributors instructed them to communicate the recall to any of their customers who have received the product using the customer letter template and acknowledgement form provided by Teleflex Medical. Their customers were to return the product and acknowledgment form to the distributor who would consolidate and return the product to Teleflex Medical. Affected products received by Teleflex Medial will be destroyed upon receipt. For additional instructions or information, contact the firm at 1-866-246-6990. |  | 3841 | 3841 units | Worldwide distribution in the states of CA, FL, GA, KY, KS, IL, MA, MD, MI, MN, NC, NJ, NM, NY, OH, OK, OR, PA, RI, TN, TX, VA, and WI and to the country of Japan. | Terminated | November 28, 2011 | 619 | 3841
Z-2222-2010 | 56256 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | SIEMENS AXIOM Luminos TF and Sireskop SD Solid State x-ray imager (flat panel/digital imager), Fluoroscopic x-ray diagnostic system. |  |  | July 13, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Fluid may enter the multi-functional handle (Opti Grip) when using the Axiom Luminos TF or Sireskop SD system, causing unintended reaction or movement, increased radiation or useless images. | Device Design | Siemens Healthcare Sector sent an "Urgent Field Safety Notice" to consignees describing the issue and informing them to take specific precautions when performing interventional procedures to avoid fluids penetrating the system. The multi-functional handle (Opti Grip) will be replaced through implementation of Update Instructions AX030/10/S. The modified version of the handle resists fluids from entering the handle and prevents malfunction. |  | 1296 | 1296 | Nationwide distribution: all 50 US states and PR. | Terminated | September 29, 2012 | 809 | 1296
Z-2223-2010 | 55804 | No | DSA | Patient transducer and electrode cable (including connector). | Cardiovascular | Cardiovascular | 510(k) | Disposable Surgical Cable, Product Code: FL-601-97 (100), Single use, Rx Only, Remington Medical, Inc., 6830 Meadowridge Court, Alpharetta, GA 30005 |  |  | May 19, 2010 | 2010 | Remington Medical Inc. | The Safe Connect component may separate from the cable, causing the cable to become unplugged from the EPG. | Process design | Consignees were notified on/about 05/19/2010. They were instructed to examine their inventory and quarantine product subject to the recall. If they distributed the product, they were advised to identify their customers and notify them of the recall (via the recall letter). A Return Authorization was provided. More information can be obtained by calling 770.888.8520. |  | 1488 | 1488 units | AZ, CA, CO, CT, DE, FL, GA, IL, IN, LA, MD, MI, MO, MS, NC, NM, NY, OH, OK, OR, PA, SC, TX, TN, TX, VA, WA, WI and Canada. | Terminated | November 29, 2010 | 194 | 1488
Z-2176-2010 | 56290 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | Ikaria INOMAX DS Drug Delivery System, model 10003. Ikaria 6 Route 173, Clinton, NJ 08809 USA. |  |  | July 21, 2010 | 2010 | Ikaria Holdings | The pressure switch in the INOMAX DS drug-delivery system may fail, which may interrupt or delay the administration of INOMAX (nitric oxide) for inhalation to patients. Delay in administration of INOMAX therapy may cause: -worsening of systemic oxygenation indices (i.e., hypoxemia, especially manifested as decreased arterial oxygenation saturation). -hypotension, and/or -increase in pulm | Component design/selection | Ikaria sent an Urgent Medical Device Recall letter to customers dated July 21, 2010. This letter is also posted on the Inomax webpage. The letter described the issue, how to recognize it, what to do in the event of a suspected leak, and recommended maintaining back-up cylinders and depressurization of the regulator when not in use. The replacement plan for all INOMAX DS drug-delivery systems was described. Customers can contact Ikaria Customer Care at 1-877-566-9466. |  | 2507 | 2507 | Worldwide Distribution to include: US Nationwide and Countries of Canada and Malaysia | Terminated | December 27, 2011 | 524 | 2507
Z-2227-2010 | 54546 | Yes | FKP | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Main Power Supply Assembly (50 and 60 Hz) and AC Power cord assembly for hemodialysis machines and Granuflo 1 mixers, Distributed by Fresenius USA Inc., Walnut Creek, CA |  |  | November 20, 2009 | 2009 | Fresenius Medical Care North America | Power cords may fail if the prongs crack or fail at the plug, and may become a fire/burn/shock hazard or delay/interrupt therapy. | Component design/selection | An urgent device recall notification was sent to all consignees. Customers are asked to inspect cords for defects, and are refered to preventive maintenance requirements in the manual. Customers who identify the defective cords will be instructed to discontinue use as soon as possible without jeopardizing patient care. Customers are instructed to monitor cords regularly, and pay attention to machines that are frequently plugged or unplugged. Personnel are instructed to not use the cord to unplug the device from the wall (yanking the plug from the wall) and to refrain from rolling over the cord or plug. |  | 10267 | 10267 | Product distributed to 5404 customers throughout the US and the world. | Terminated | February 07, 2011 | 444 | 72142
Z-2229-2010 | 56193 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SpeedPass Disposable Suture Lariat with Nitinol Wire 25 Degree Left, QTY. 1, Biomet Sports Medicine, Warsaw, IN 46581. Sterile. Part number 904010. |  |  | June 18, 2010 | 2010 | Biomet, Inc. | The nitinol wire may be fractured or may have the potential to fracture making the suture lariat unusable and possibly delaying surgery. | Component design/selection | The firm notified all 29 consignees by An Urgent Medical Device Recall Notice letter on 6/18/2010. The letter identified the affected product and explained the reason for recall. Customers are asked to locate and remove the affected products, follow the instructions on the enclosed "FAX Back Response Form," and fax a copy of the form back prior to return of the affected product. Additionally, if product has been further distributed, then hospital personnel must be notified and given the "Dear Biomet Customer" letter. Questions should be directed to Mary Johnson at 800-348-9500 or 574-372-3983. |  | 301 | 301 units total | Worldwide Distribution -- USA, including the states of FL, IL, NY, CO, MD, CA, GA, AL, PA, OH, SD, KY, TX, ID, WA, NC, AK, &amp; MN, and Columbia and The Netherlands. | Terminated | August 06, 2012 | 780 | 301
Z-2233-2010 | 56262 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis MP/Artis dMP and Artis zee Multipurpose. Angiographic x-ray system. |  |  | July 13, 2010 | 2010 | Siemens Medical Solutions USA, Inc | The Axiom Artis MP, Artis dMP or Artis zee Multipurpose C-arm gearbox can become damaged. If the gearbox is damaged, the C-arm may tilt away when being repositioned and collide with nearby objects in its range of motion. | Other | Siemens has issued an "Urgent Field Safety Notice: Customer Safety Information to consignees with Update Instructions AX033/10/S. This letter informed customers of the issue, and provides an Addendum to the User Manual. The Update Instructions describe the process for installation of an additional clamping mechanism provided by Siemens into the existing system, which will be activated if the gearbox is damaged. If activated, the clamping machanism will prevent the C-arm from tilting away, stop regular operation, and alert the operator to check the system and repair the gearbox damage. |  | 131 | 131 units | Nationwide distribution: States of AL, AZ, CA, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NY, OH, OK, PA, PR, RI, SC, TN, TX, UT, VA, WA, WI, and WY. | Terminated | July 05, 2012 | 723 | 131
Z-2235-2010 | 56225 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | medCOMP(R) MRCTI66841 6.6F Plastic Dignity Low Profile CT Port with Pre-Attached Chronoflex(R) Polyurethane Catheter - 1 Unit. Catalog #MRCTI66841. |  |  | July 02, 2010 | 2010 | Medical Components, Inc dba MedComp | Recalled product was packaged with the incorrect port and does not match the label. The kit is labeled as a 6.6F "Dignity" Low Profile CT Port. The kit contains a 6.6F "Pro-Fuse" Low Profile CT Port. | Employee error | Medcomp sales reps and foreign distributor were notified via email on 7/1/10. Product is to be returned to Medcomp. |  | 47 | 47 units | Worldwide Distribution -- USA, including states of NC and TX and country of Israel. | Terminated | October 13, 2010 | 103 | 47
Z-2234-2010 | 56138 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Accessory Power Pack; a 4 outlet power strip with a custom pole clamp bracket to attach to the IV pole with the infusion pump; Distributed by: Hospira, Inc., 275 N. Field Drive, Lake Forest, IL 60045; part number 39003-04-01, Modo #HSP1700, Model Tripp-Lite PS-415-HG |  |  | July 02, 2010 | 2010 | Hospira Inc. | The Power Pack Accessory is not certified to UL 60601 standard or Subject 1363A for relocatable power taps intended for use with medical equipment. Potential problems could include electrical shock, electrical short or electrical fire. | Device Design | Hospira sent Urgent Device Recall letters dated 7/1/10 to their direct accounts on 7/2/10 via Federal Express 2 day delivery, informing them of the potential for electrical shock, electrical short, or electrical fire, since the power packs were not certified to UL 60601 standard or Subject 1363A of UL standard 1363 for relocatable power taps intended for use with medical equipment. The accounts were instructed to examine their inventory for the affected product and quarantine it immediately, and Stericycle at 1-800-321-7388 to obtain a return kit for returning the affected product. If they distributed the affected product further, they were instructed to sub-recall the product from their sub-accounts. They were also requested to complete the enclosed reply form and return it to Hospira by fax at 1-262-577-6921, indicating the amount of product placed on hold for return, if they further distributed the product and if they notified their retail customers. For medical inquiries, the accounts were directed to contact Hospira Medical Communications at 1-800-615-0187. |  | 1766 | 1,766 units | Worldwide Distribution -- US and Canada. | Terminated | July 20, 2011 | 383 | 1766
Z-2236-2010 | 55663 | No | FDI | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus SD Snares, Reusable Electrosurgical Snares. Models: SD-9L/U-1, SD-11L/U-1, SD-12L/U-1, SD-13L/U-1, SD-13L/U-1, &amp; SD-16L/U-1. Snare tube model numbers: MAJ-13 and MAJ-15. |  |  | May 03, 2010 | 2010 | Olympus America Inc. | An investigation of consumer complaints revealed the inability of the snare wire to be retracted into the snare tube after deployment of the snare. | Other | Firm initiated this recall following investigation of customer complaints regarding the inability of the snare wire to be retracted into the snare tube. Customers were notified via certified mail, an Urgent: Medical: Device Recall letter issued on May 4, 2010. The letter identified the affected product and the reason for the recall. Customers were instructed to inspect their inventory and identify the affected product. Customers are to remove from use any identified inventory, quarantine, and discard them. Customers should contact the firm to obtain replacement products. Questions should be directed towards Laura Storms-Tyler at 484-896-5688 or laura.storms-tyler@olympus.com. |  | 108 | 108 units | Nationwide Distribution -- AR, CA, CO, CT, DC, FL, GA, IN, KS, LA, MD, MO, NC, NY, OH, PA, SC, TN, and TX. | Terminated | May 20, 2013 | 1113 | 108
Z-2218-2010 | 54687 | No | NWX |  |  | Cardiovascular | PMA | AngioSculpt Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter, cutting/scoring, Easy Exchange (EX) delivery system, Manufactured by AngioScore, Fremont, CA. |  |  | December 04, 2009 | 2009 | AngioScore Inc. | Separations of the PTCA catheter shaft at the guide wire exit port have occurred during treatment of coronary artery stenosis. Fragments of the catheter may become lodged in coronary arteries, which resulting in serious patient outcome and require emergency coronary bypass surgery and medical intervention to remove the retained device fragments. | Process design | AngioScore sent an "Urgent: Medical Device Voluntary Recall" dated December 4, 2009 addressed to "Catheterization Lab Manager" at consignees. The Recall notification letter instructed consignees to inspect their inventory to detect product from the involved lot, and segregate it from their inventory. Consignees were asked to record all affected product and fax return forms to the firm. AngioScore Customer Service will contact consignees and provide instructions for return and replacement of affected product.. Consignees can contact AngioScore at 877-264-4692. |  | 3870 | 3870 units | Nationwide distribution. | Terminated | June 21, 2011 | 564 | 3870
Z-2237-2010 | 49012 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz Dual series Computed Tomography (CT) Scanner System, X-Ray System. Part Number: 989605651321. |  |  | July 21, 2008 | 2008 | Neuisys, LLC | The reference lines for image generated for a Surview scan may appear in the incorrect position in the "FILM" display mode. | Software design | On 07/21/2008, NEUISYS Imaging Systems Solutions distributor in Greensboro, NC began sending the URGENT DEVICE CORRECTION letter on behalf of Philips and NeuSoft Medical Systems (PNMS), Co., Ltd in China.The letter informs the consignees that there is potential safety problem in the NeuViz Dual series CT scanners due to the malfunction of "Incorrect Reference Line in the Film". When a "Head first, Couch out" or a "Feet first, Couch in" Surview scan is started then paused midstream, the image generated is what the doctor expects, but the reference lines appear in the incorrect position in "FILM" display mode. The acquired images accurately represent the exact positions planned on the original view. The incorrect/unmatched reference line in film may result in incorrect diagnosis.The consignees are advised not to stop or pause during Surview acquisition while PNMS in China will provide the updated software patch to all affected systems free of charge.The Field Change Order #FCOP-08-CT-NMS-008 SP4 will be implemented by NEUISYS Imaging Systems Solutions in Greensboro, NC., who is a distributor for PNMS. Consignees are advised to e-mail Service Support Department of PNMS at helpdesk@pnms.neusoft.com regarding the recall. |  | 19 | 19 units in the US | Nationwide Distribution | Terminated | August 23, 2010 | 763 | 19
Z-2239-2010 | 55043 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The Starion Instruments Thermal Ligating Shears 35C Model 102-138D and 23C Model numbers 102-137D allow the surgeon to position the distal jaws of the instrument around the region of tissue to be cut/cauterized. The surgeon squeezes the handle and depresses the switch, heating elements activates in the jaws. The heat is conducted to the tissue between the jaws to provide cutting/cauterization. For the simultaneous cutting and cauterization of soft tissue during surgery. Cutting of natural or synthetic, non-metallic, sutures during surgery. |  |  | March 09, 2010 | 2010 | Starion Instruments | The Starion Instruments Thermal Ligating Shears (STLS) 35C and the Starion Thermal Ligating Shears (STLS) 23C may turn on and remain on without the finger button being depressed. The devices may also work intermittently, incorrectly indicating the device is 'on". | Device Design | Starion distributed a "STLS Device Recall - Important Medical Device Information" letter dated March 9, 2010 to consignees. The letter identified the affected model and lot numbers. Consignees were asked to complete a inventory status form and fax to Starion. the letter included a shipping labe with a pre-assigned return authorization number. Starion will ship replacement after receipt of the inventory status form and returned product.. |  | 228 | 228 units | Worldwide Distribution: USA states of CA, OH, MI, KS and MO and countries of Italy, China, Kuwait, and Germany. | Terminated | November 27, 2010 | 263 | 228
Z-2238-2010 | 55921 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 2000, manufactured by GE Medical Systems, SCS |  |  | June 10, 2010 | 2010 | GE Healthcare, LLC | A hardware component (potentiometer and encoder)mismatch in Innova 2000 digital fluoroscopic imaging systems may cause a Source to Image Distance (SID) failure, which will cause the imager to fail to detect the distance between the imager and the patient. This will result in loss of the x-ray image if the recovery process is not performed. The error message to use the Emergency Back-Out Button for | Other | GE Healthcare sent a GE "Urgent Medical Device Correction" dated June 9, 2010letter to consignees. The letter was addressed to Hospital Administrators / Risk Mangers; Mangers of Radiology/Cardiology and Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product Details; Safety Instruction; Product Correction and Contact Information. The letter detailed the method for identifying SID failure and the use of the emergency back-out buttoms to recover the X-ray image. A GE service representative will affix a label to the device on-site to remind operators how to recover from loss of X-ray from SID failure. Consignees may contact GE at 800 437-1171, option 4. |  | 885 | 885 | Worldwide distribution: States of AL, AZ, AR, CA, CO, CT,DE, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV,NH, NJ, NM, NH,NC, ND, OH, OK OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, and WI, and countries of VENEZUELA, UNITED KINGDOM, TURKEY, TUNISIA, THAILAND, TAIWAN, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN SPAIN, SOUTH AFRICA, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, POLAND, PAKISTAN, NORWAY, NEW ZEALAND, NETHERLANDS, MOROCCO, MEXICO, MARTINIQUE, LITHUANIA, LEBANON, JORDAN, JAPAN, ITALY, ISRAEL, ISLAMIC REPUBLIC OF IRAN, IRELAND, INDONESIA, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, EGYPT , ECUADOR, COLOMBIA, CHINA, CANADA, BULGARIA, BRAZIL, BELGIUM, BELARUS, AUSTRALIA, and ALGERIA. | Terminated | April 01, 2014 | 1391 | 885
Z-2240-2010 | 56258 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | PoleStar FlexiCoil band D/E, revisions A and B, Catalog # 9733063, |  |  | July 01, 2010 | 2010 | Odin Medical Technologies Ltd. | Electronic box on headstrap may overheat, causing burn to patient. | Device Design | Customers were notified by letter on 07/07/2010, advising them of the overheating problem and alerting them that a service rep will be visiting to inspect and upgrade any affected FlexiCoils. Questions were referred to 800-595-9709. |  | 560 | 560 units | CA, CO, DC, IL, IN, MD, MI, MN, MO, NE, NJ, NV, NY, OH, OR, WA, and WI. Foreign distribution to Austria, Belgium, Canada, China, France, Germany, Israel, Italy, Korea, Netherlands, Poland, Spain, Russia, Switzerland and Turkey. | Terminated | November 09, 2010 | 131 | 560
Z-2241-2010 | 55153 | No | NRU | Dental chair and accessories. | Dental | Dental | 510(K) Exempt | The magnetic headrest is designed to fit the following dental chairs (THE HEADREST IS THE RECALLED ITEM, NOT THE CHAIRS): |  |  | March 23, 2010 | 2010 | Dental Equipment LLC | The magnetic headrest on certain dental chairs may affect the function/programming of some implantable pacemakers or defibrillators if the implanted device is programmed to respond to a magnet. | Other | DCI Equipment sent an "Urgent Field Safety Alert" letter dated March 15, 2010 to consignees advising them of the potential problem. The letter instructed that patients with implantable pacemakers or defibrillators should avoid dental chairs with the magnetic headrest. A warning label was sent with the letter for insertion into the product Use and Care manual for operator warning. Consignees were asked to complete a Field Safety Alert Acknowledgement/Return Form and fax to the firm. Consignees were requested to notify customers.Consignees may contact the firm at 503 537-3602. |  | 108 | 108 | Worldwide distribution in the USA and Canada. | Terminated | January 04, 2012 | 652 | 108
Z-2243-2010 | 56291 | No | NUK |  |  | Clinical Chemistry | Contact ODE | Human Cardiac Specific Fatty Acid-Binding Protein (H-FABP) ELISA TEST KIT Product is an In Vitro Device (IVD) categorized as a Cardiac Marker, manufactured and distributed by BioCheck, Inc., Foster City, CA. Enzyme Immunoassay for the quantitative determination of Human Cardiac Fatty Acid-Binding Protein (H-FABP) concentration in serum, citrate-plasma, EDTA-plasma, and heparin-plasma. |  |  | July 15, 2010 | 2010 | BioCheck Inc | Products were marketed without 510(k) for diagnostic use. Unapproved device. | PMA | BioCheck sent an "Urgent Device Recall" letter dated July 15, 2010 to consignees identifying the products as unapproved and requesting prompt return to the firm. Consignees are asked to sub-recall. The letter requested a reponse form be completed and returned to the firm.Diagnostic claims have been removed from the website. Returned product will be quarantined by the firm pending determination of marketing status and labeling revision. |  | 187 | 187 kits | Worldwide Distribution - USA, including the states of CA, MI, NJ, and PA and countries of Canada, India, S. Korea, and UAE. | Terminated | December 13, 2010 | 151 | 187
Z-2242-2010 | 56285 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Surgi-Tip Transducer Cover, Ref. 610-834, Sterile (20.3 x 244cm) polyethylene telescopically-folded cover with attached (6cm) NeoGuard tip. The latex-free surgi tipped covers are utilized to cover ultrasound transducers and prevent cross-contamination from occurring with equipment and patients. |  |  | July 20, 2010 | 2010 | Civco Medical Instruments Inc | Polyethylene material did not have proper treatment applied during manufacturing. Untreated polyethylene material does not maintain a secure seal with the adhesive tape used to secure the tip to the cover. | Nonconforming Material/Component | CIVCO sent a "Urgent - Medical Device Correctdion" letter date July 19, 2010 to consignees describing the reason for recall and requesting they dispose of the affect lot number of product. A response form was enclosed for completion and return to the firm, indicating the amount of product in inventory and requesting a Return Authorization Number. The form was to be returned via FAX or e-mail. Customers may contact CIVO at 800-445-6741. |  | 22 | 22 boxes | Nationwide distribution: the states of OH, MD, NE, VA, IA, MI, WA, and WI. | Terminated | April 13, 2012 | 633 | 22
Z-2244-2010 | 56200 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKnee Baseplates, Size 8, Left, Part # 333-01-108, Lot# A1000003. |  |  | June 28, 2010 | 2010 | Encore Medical, Lp | One lot of product has an incorrect color label. The color label is a secondary visual indicator of sizing. | Labeling mix-ups | Firm notified their consignees by letter on June 28, 2010. |  | 18 | 18 | Florida, Missouri, and Italy | Terminated | October 04, 2010 | 98 | 18
Z-2245-2010 | 54701 | No | HWQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF 8570 MENISCAL REPAIR NEEDLE, STERILE, Rx ONLY. |  |  | January 28, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | This recall was due to the possibility that the products may have a breach in the poly-Tyvek pouch that could potentially compromise the sterility of the contents. | Material/Component Contamination | Each consignee will be notified of the recall by an Urgent Medical Device Recall Notification letter sent by FedEx or an equivalent method. The letter identified the affected product, explained the reason for recall, and health risk. Consignees were instructed to examine their inventory and segregate the affected product for return. Customers are to complete and return the Reply Form and to notify their customers if the product was further distributed. Consignees are instructed to return any remaining product to ConMed Linvatec. Questions should be directed to the Customer Service department at 800-535-8536. |  | 35451 | 35,451 total | Worldwide Distribution -- United Arab Emirates, Austria, Australia, Belgium, Canada, Switzerland, Germany, Denmark, Spain, Finland, France, United Kingdom, Hong Kong, Ireland, Israel, India, Italy, Republic of Korea, Netherlands, Portugal, Serbia, Singapore, Slovenia, El Salvador, Thailand, Taiwan, South Africa, and throughout the U.S. | Terminated | December 15, 2010 | 321 | 35451
Z-2267-2010 | 55156 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom Care Assist Bed, Model P1170. |  |  | April 13, 2009 | 2009 | Hill-Rom, Inc. | Pinch hazard. Potential for shearing action between the side rail arm and the bracket used to secure the side rail to the bed. | Device Design | Hill-Rom sent a customer notification letter dated 4/9/2009. A second notice was sent to all customers on 7/9/2009. The letter included instructions and a Customer Response Form, along with the required parts. Customers were instructed to do the correction on their affected beds and return the completed customer response form within 45 days. |  | 39971 | 39,971 both products | Worldwide distribution: USA, Afghanistan, Algeria, American Samoa, Argentina, Australia, Bermuda, Brazil, Canada, Chili, China, Colombia, Costa Rica, Cuba, Ecuador, Egypt, Hong Kong, India, Indonesia, Israel, Italy, Japan, Jordan, Kuwait, Kyrgyzstan, Lebanon, Libya, Malaysia, Mexico, Oman, Panama, Paraguay, Philippines, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Korea, Syria, Taiwan, Thailand, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Yemen. | Terminated | June 06, 2011 | 784 | 39971
Z-2266-2010 | 56273 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Philips M1026BT Gas Modules with manifolds |  |  | July 21, 2010 | 2010 | Philips Healthcare Inc. | The manifold seal in the Philips M1657B Water Trap used on the Philips M1013A IntelliVue G1, M1019A, IntelliVue G5, and Anesthesia Gas Modules M1026B/BT may be unintentionally removed when changing the traps. A missing seal may result in leakage and may cause incorrect monitor readings of inspiratory or expiratory gas levels. | Device Design | Philips Healthcare issued an "Urgent Device Correction" letter dated July 14, 2010 with an addendum to the Instructions for Use including additional instructions for the replacement of the water trap for the Philips M1013A IntelliVue G1, M1019A IntelliVue G5, M1026B/BT Anesthesia Gas Modules. This addendum is being sent along with the Urgent Medical Device Correction notice to all affected customers.Customers will be requested to incorporate the information contained in the Urgent Medical Device Correction notice and the Addendum to the Instructions for Use with the existing labeling (Instructions for Use) provided with the product. Customers can contact Philips at 1-800-722-9377 for further information.. |  | 367 | 367 units | Worldwide Distribution: United States, and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Bulgaria, Chile ,China ,Colombia, Czech Republic, Estonia, Finland, France, Germany, Hong Kong, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Republic of Korea, Latvia, Libyan Arab Jamahiriya, Luxembourg, Malaysia, Mexico, Mozambique, Netherlands, New Zealand, Norway, Oman, Pakistan, occupied Palestinian territory, Peru, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Viet Nam. | Terminated | March 09, 2012 | 597 | 4675
Z-2278-2010 | 52030 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ABC Bend-A-Beam Handpiece REF 134009 Handcontrol Malleable Handpiece, 9" (22.8 cm) QTY: 1 STERILE CONMED CORPORATION 525 FRENCH RD UTICA, NY 13205 ASSEMBLED IN MEXICO (Pouch label) |  |  | December 18, 2008 | 2008 | ConMed Electrosurgery | The internal electrode can protrude in front of ceramic insulator at the tip of the ABC Handpiece used with Conmed Electrosurgical units. | Process design | An Urgent Device Recall notification was issued 12/18/2008, which explained the problem/risk associated with the device, the discontinuance of use and the return of the device. If you have any questions, please contact 800-552-0138, ext. 5225. |  | 233 | Dom: 233 cs (2330 units); Int'l: 275 cs (2750 units) | Nationwide and International | Terminated | October 22, 2010 | 673 | 5527
Z-2270-2010 | 50226 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Intesys Clinical Suite (ICS) G2 92810 Smart Disclosure. The Smart Disclosure product collects patient waveforms, alarms, vital signs and 12-lead reports. A 24-hour waveform acquisition is standard, with an upgrade option to 72 hours. |  |  | October 17, 2008 | 2008 | Spacelabs Healthcare, Incorporated | Potential for device to print out the wrong waveform data for the time indicated in the printout or display. | Device Design | On 07/17/2008, Spacelabs Healthcare began sending the MEDICAL DEVICE CORRECTION letter to their consignees. The letter informs the consignees that there is an issue involving ICS G2 Smart Disclosure (Model 92810) which causes the display or printing of an incorrect waveform for the time indicated in the printout or display.The consignees are advised whether to discontinue to use the device until the firm upgrades the system. If the consignees decide to continue to use the device, then they should discontinue any use of the "Refresh" button. When they print or view waveform reports,t hey should utilize the Saved Event view functionality of the product rather than Arrhythmia, Alarm or Waveform Views. If the printing from the Arrhythmia, Alarm, or Waveform views, the consignees must visually inspect the print out and verify that the waveform contains what you expected to see printed.The consignees could contact Spacelabs Healthcare at 1-800-522-7052 or 1-425-657-7200, x5089 for any questions about the corrective actions of the affected device. The firm's field service engineer will make a schedule a mutually agreeable time with the consignees to update the ICS G2 system at free of charge. |  | 135 | 135 units | Worldwide Distribution -- US, Turkey, Chile, Brazil, Netherlands, Puerto Rico, Switzerland, Australia, Czech Republic, Canada, Italy, Singapore, France, Germany, and Israel. | Terminated | September 29, 2010 | 712 | 135
Z-2269-2010 | 54810 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | ESSENTIAL MEDICAL SUPPLY, INC. ENDURANCE HD, BUILT TO LAST, HEAVY DUTY FOUR WHEEL WALKER for users up to 500 pounds. |  |  | December 21, 2009 | 2009 | Essential Medical Supply, Inc. | Essential Medical Supply, Inc. is recalling W1801B and W1801R Endurance HD Walkers due to recent reports and tests that show the front caster may detach from the frame. The potential failures include a collapse of the plastic fork or shearing of the stud attaching the caster to the frame. | Other | All Consignees were mailed an Urgent: Product Recall W1801B and W1801R Endurance HD Walkers notification on December 30, 2009. All notifications were sent first class mail in a 9" X 6" envelope. Response cards were included. The letter identified the affected product and the reason for recall. Customers were asked to examine their inventory for affected product and to discontinue further use or distribution of the product. If the product was further distributed, customers should contact their customers, advise them of the recall and have them return the units to them. Customers should follow the instructions provided for return and replacement of units. |  | 2954 | 2,954 units | Nationwide Distribution | Terminated | February 21, 2012 | 792 | 2954
Z-2274-2010 | 55933 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Arrow Percutaneous Sheath Introducer Kit with Integral Hemostasis Valve/Side Port for use with 7.5 Fr. Catheters; Catalog number AJ-09801 |  |  | June 07, 2010 | 2010 | Arrow International Inc | Sterility not assured. The stopcock handle position was inadvertently changed from an open to a closed position, which prevents sterilization of the closed portion of the device. Sterility of the affected product cannot be guaranteed. | Process change control | June 7, 2010 Notification letter to customer instructed them to return the product to Arrow. Arrow will destroy the returned product. |  | 3000 | 3000 kits | All affected product was distributed to one foreign consignee: Arrow Japan Ltd. | Terminated | September 29, 2012 | 845 | 3000
Z-1977-2010 | 55855 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | FLEXIMA Biliary Stent System, 7Fr-10cm, M00539220, Sterile EO, Boston Scientific, Spencer, IN. The Flexima Biliary Stent System is intended for delivery of the stent to the biliary tract for drainage of the bile ducts, for splinting of a bile duct during healing, or for providing bile duct patency in a stricture or past a stone. |  |  | April 27, 2010 | 2010 | Boston Scientific Corporation | Flexima Biliary Stent System Devices were packaged with Tyvek tray lids that had incorrect adhesive, resulting in sterile barrier seal defects. The breach of the sterility barrier could lead to the device being non-sterile. | Packaging | Boston Scientific sent an "Urgent Medical Device Recall - Immediate Action Required" letter date April 27, 2010 to all consignees. The letter included instructions to immediately discontinue use and segregate the affected product for return, and complete and return by fax a Reply Verification Tracking Form. Customers with affected product were instructed to contact customer service at 866-868-4004 to obtain a Return Authorization (RGA) number, and package and ship the product to Boston Scientific. |  | 1406 | 1406 | Worldwide Distribution: United States and Puerto Rico, and countries of Arab Emirates, Argentina, Austria, Belgium, Canada, China, Columbia, Croatia, El Salvador, Estonia, Finland, France, Great Britain, Germany, Greece, Hungary, Iran, Israel, Italy, Japan, Korea, Latvia, Lebanon, Malaysia, Mexico, the Netherlands, Norway, Poland, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Ukraine, and Uruguay. | Terminated | March 23, 2011 | 330 | 13227
Z-2275-2010 | 56141 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE MONITOR B850 |  |  | July 16, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware of four potential safety issues associated with the CARESCAPE" Monitor B850:1. Loss of user input when using certain displays or the USB remote control.2. Potential for a delay in treatment or missed alarm when the alarm light is used as a primary alarm source and audible alarms are turned down or off.3. The CARESCAPE Monitor B850 will not use user entered hemo | Software design | Consignees were sent on 7/16/10, a GE "Urgent Medical Device Correction" letter dated July 12, 2010. The letter was addressed to Healthcare Administrator/Risk Manger, Chief of Nursing, Director of Biomedical Engineering. They stated the Safety Issues, Affected Product Details, Product Correction, Safety Instructions and Contact Information. If you have any questions, please contact Technical Support at 1-800-558-7044. |  | 875 | 875 | CA, CO, FL, GA, KY, LA, MA, MI, NE, NJ, NY, NC, OH, OR, PA, TX VA, and WI, Foreign: UNITED KINGDOM, TURKEY, SWITZERLAND, SPAIN , SLOVENIA, SINGAPORE, SAUDI ARABIA, PORTUGAL NETHERLAND, ITALY, GERMANY, FRANCE,FINLAND, DENMARK, CZECH REPUBLIC, BELGIUM, AUSTRIA, and AUSTRALIA. | Terminated | September 18, 2013 | 1160 | 875
Z-2271-2010 | 56276 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MedtronicÂż EverestÂż 20 Disposable Inflation Device, AC2200. |  |  | July 16, 2010 | 2010 | Medtronic, Inc. | Product sterility may be compromised due to breach of package seal integrity. | Packaging process control | Medtronic issued a Urgent Medical Device Recall Notice to US customer on 7/15/2010 and to foreign accounts advising of the issue and requesting that they segregate recalled product. A Medtronic Sales Rep will visit their facility to account for all affected product and initiate a credit. Foreign accounts will return product to their respective distribution centers. |  | 2875 | 2,875 units | Worldwide Distribution -- USA, including the state of MI and the countries of Austria, Belgium, Czech Republic, Finland, France, Germany, Greece, Ireland, Israel, Italy, Lithuania, Pakistan, Poland, Portugal, Slovakia, Spain, Sweden, Switzerland, Ukraine, United Arab Emirates, and United Kingdom. | Terminated | June 07, 2011 | 326 | 4897
Z-2287-2010 | 54474 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Deknatel BondekÂż Plus Polyglycolic Absorbable Surgical Sutures, Bon+VI R 0 HRX27 1N (70 cm), Teleflex Medical. |  |  | February 01, 2010 | 2010 | Telefelx Medical | Labeling error; The Bondek Suture product inside the box are of a different catalog number than labeled on the outer box. | Labeling mix-ups | Teleflex Medical issued consignees a "Urgent Medical Device Recall" letter dated February 1, 2010. Distributors were instructed to forward the letter to their consignees to retrieve relevant product from those locations. The letter described the packaging issue and potential impact. Customers were instructed to immediately discontinue use and quarantine any identified products, complete and fax the Recall Acknowledgement Form to 866-804-9881, in order to receive a Return Goods Authorization (RGA) Number, and then return the product to Teleflex Medical for replacement. Affected products received by Teleflex Medical will be destroyed upon receipt. |  | 6936 | 17 sales units (24 packs per sales unit/408 units) | Worldwide Distribution: USA and Germany | Terminated | April 19, 2011 | 442 | 6936
Z-2288-2010 | 55772 | board, bed | FPS | Bed board. | General Hospital | General Hospital | 510(K) Exempt | ReliaMed Deluxe Portable Adjusting Table. Product is a medical table used on patients for massage and chiropractic treatment. |  |  | April 23, 2010 | 2010 | MDC Acquisition Co. | Two complaints were received of the vinyl-covered, padded, plywood table surface breaking near the foot end. | Component design/selection | "Urgent: Medical Device Recall" letters were sent to consignees from Milliken Medical and Meyer Distributing Company in April 2010 describing the issue. They instructed their customers to immediately stop use of the table. customers were instructed to return the tables to the firm with an enclosed Recall Form for future replacement, or for replacement with a comparable table from another menufacturer at no cost, or for refund of the original cost of the device. Consignees may contact Meyer Distributing Company at 800-472-4221 and Milliken Medical at 800-532-1356. |  | 265 | 265 Tables | Worldwide distribution: USA, including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA and WI, and the country of Canada. | Terminated | February 16, 2012 | 664 | 265
Z-2284-2010 | 55870 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity Delta/Gamma X XL Patient Monitor. Portable monitor used to monitor vital signs of neonatal, pediatric, and adult patients for bedside and transport applications. |  |  | June 04, 2010 | 2010 | Draeger Medical Systems, Inc. | Monitor keys may become inoperative or activate spontaneously, causing a membrane switch panel malfunction. This may discharge a patient automatically. | Device Design | Draeger Medical issued "Urgent-Medical Device Recall Notification" letters to affected consignees. The Delta/Gamma X XL monitor keyboards will be updated with new material at each affected facility. If the keys on affected monitors become inoperative or activate without user interaction, the customer is instructed to remove the monitor from service and contact Draeger at 978-379-8573. A Technical Service Bulletin (TSB) will be distributed globally, instructing subsidiaries to inform international customers. |  | 800 | 800 units | Worldwide Distribution - United States, including the states of AL, AR, CA, CO, DC, FL, GA, IA, IL, IN, KS, LA, MA, ME, MN, MO, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SD, TN, TX, UT, VA, WA, WI, WO, and W, and the countries of United Arab Emirates, Angola, Argentina, Austria, Australia, Bosnia, Belgium, Bulgaria, Benin, Brazil, Canada, Switzerland, Chile, China, Columbia, Costa Rica, Serbia, Czech Republic, Germany, Egypt, Spain, Finland, France, Great Britain, Greece, Hong Kong, Coratia, Hungary, Ireland, Israel, India, Iraq, Italy, Japan, Jordan, Kenya, South Korea, Kuwait, Kazakhstan, Lebanon,Lithuania, Latvia, Libya, Morocco, Macedonia, Martinique, Mexico, Netherlands, Norway, Nepal, Peru, Poland, Puerto Rico, Portugal, Qatar, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovenia, Slovak Republic, Turkmenistan, Turkey, Ukraine, Uruguay, Venezuela, Mayotte and South Africa. The government accounts are Bay Pines VAMC, Bay Pines, FL; Naval Medical Center, Portsmouth, VA; and Walter Reed Army Medical Center, Washington, DC. | Terminated | July 30, 2012 | 787 | 9475
Z-2291-2010 | 56032 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VERSYSÂż HIP SYSTEM |  |  | July 05, 2010 | 2010 | Zimmer Inc. | Zimmer, Inc. has determined that there is a potential for a package to be labeled as a 28 mm femoralhead but contain a 32 mm femoral head. In addition, there is a potential for the patient record label tobe incorrect. | Other | The firm, Zimmer, Inc., sent an "URGENT: DEVICE REMOVAL" letter dated July 9, 2010, to all consignees. The letter described the issue, clinical implications, affected product, and required actions. The consignees were instructed to carefully review the letter, ensure users of the VerSys Hip System Femoral Head device at their faculty have been informed of this removal and perform a physical count of all affected product at their facility, record the data on the Inventory Return Certification Form and/or the User Facility/HCP Form and return the form via fax at (574) 372-4265 or email: aimee.wood@zimmer.com. Zimmer Inc. will conduct a removal of all remaining (unused) units of the subject lots. For shipping assistance, questions or assistance in notifying your accounts about the device removal, please contact Zimmer, Inc. at 1-800-613-6131 or 574-372-4463. |  | 103 | 103 | Worldwide distribution: USA and countries including Bolivia, Canada, New Zealand, Taiwan, Thailand, Singapore, and India. | Terminated | January 12, 2011 | 191 | 103
Z-2292-2010 | 56329 | No | HPZ | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | Sterile Bio-Eye Hydroxyapatite Orbital Implant and Conformer |  |  | June 07, 2010 | 2010 | Integrated Orbital Implants Inc | Integrated Orbital Implants, Inc. has initiated a recall of it's Sterile Bio-Eye Hydroxyapatite Orbital Implant and Conformer due to a labeling issue in which the expiration date of the product was omitted. | Error in labeling | Customers were contacted by telephone and email |  | 40 | 40 units | USA (MT, AL, IN, IA, NY, KY, IL, MI, and FL), Canada, Germany, and Turkey | Terminated | August 31, 2010 | 85 | 40
Z-2285-2010 | 56542 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Suros ATEC 0912-20 Breast Biopsy and Excision System Handpiece, REF ATEC0912-12, sterile, Hologic, Inc. 6100 Technology Center Drive Indianapolis, IN. |  |  | August 11, 2010 | 2010 | Hologic, Inc. | The firm received a customer complaint that the a device labeled as ATEC 0912-12 was actually ATEC 0912-20. | Packaging | The firm, Hologic, Inc., sent out "Urgent Medical Device Recall" letters by FedEx on August 11, 2010, to all customers. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to check their inventories and to quarantine any of the recalled product; if the product was distributed further, identify the users and notify them at once of this product recall; segregate all products from this lot and quarantine them in preparation for return to Hologic, and complete and return the enclosed RECALL RESPONSE FORM by fax at 877-574-3255 or mail. Note: Replacement product cannot be sent until Hologic receives your response form.If you have any questions, please contact Hologic directly at 317-344-7630. |  | 990 | 990 | Worldwide distribution: USA and countries including Canada, Belgium, Algeria, Germany, London, Lebanon and Italy. | Terminated | April 13, 2011 | 245 | 990
Z-2289-2010 | 56328 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer ITST" INTERTROCHANTERIC/SUBTROCHANTERIC FIXATION LOCKING BOLT EXTRACTOR - NONSTERILE |  |  | July 05, 2010 | 2010 | Zimmer Inc. | Instruments were manufactured with the incorrect grade of steel, increasing the potential for fracture during use. | Nonconforming Material/Component | Zimmer sent "Urgent: Device Removal" letters dated July 8, 2010 to consignees identifying the potential for fracture of the device. Consignees were instructed to inventory affected product and record on the Inventory Return Certification Form, to document users of the affected device on the User Facility/HCP Form, and return both forms by fax or email. Affected product is to be returned to Zimmer with a copy of the Inventory Return Certification Form. Returned product will be destroyed. Each of the affected consignees will be provided with a new instrument, manufactured with the correct 17-4 stainless steel.Consignees can contact Zimmer at 800-613-6131 or 574-372-4463 |  | 20 | 20 individual units | Worldwide Distribution in the countries of US, Canada, Spain, Germany, and Great Britain. | Terminated | February 25, 2011 | 235 | 20
Z-2309-2010 | 55877 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Mitek VAPR Temperature Control Electrode (integrated |  |  | June 07, 2010 | 2010 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Product does not meet the packaging specifications for pouch seal integrity which could lead to a breach in sterility | Packaging process control | The firm, Depuy Mitek, sent an "Urgent Voluntary Product Recall" letter dated June 7, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to not use any of the identified recalled products; to isolate all inventory of the affected product; complete the Business Reply Form, and return the affected products along with the form to Depuy Mitek; ATTN: Recall Coordinator, 50 Scotland Boulevard, Bridgewater, MA 02324 or fax completed form to 1-508-828-3750. Note the IMPORTANT PRODUCT INFORMATION: Any product that contains a round green sticker has been 100% inspected, meets specifications and is acceptable for use. Please Do Not Return any product that is labeled with the green dots because this product is not affected by the recall.If you have any questions or concerns with regard to this program, please contact your local Depuy Mitek Representative or DePuy Mitek Customer Service directly at 1-800-382-4682. |  | 1414 | 1414 units | Worldwide distribution: USA and countries including Canada, Argentina, Australia, Bahrain, Belgium, Brazil, Chile, China, Columbia, Costa Rica, Czech Republic, Ecuador, Egypt, France, Germany, Greece, Guatemala, Hong Kong, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Jordan, Kazakhastan, Kuwait, Lebanon, Maylasia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Palestine, Panama, Peru, Poland, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Portugal, Sweden, Switzerland, Syria, Thailand ,Turkey, UK, Ukraine, UAE, Venezuela, Vietnam, and Yemen. | Terminated | July 12, 2012 | 766 | 813657
Z-2294-2010 | 56281 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Greatbatch OptiSeal Valved PTFE Peelable Introducer, Model 1000093-001, 1000093-002, 1000093-003, 1000093-004, 1000093-005, 1000093-006, 1000093-007. Sterilized using Ethylene Oxide. Greatbatch Medical, 2300 Berkshire Lane North, Minneapolis, MN 55441 USA. The OptiSeal Valved PTFE Peelable Introducer is designed to reduce blood loss and air intake by providing hemostatic sealing at venous pressures. The device is offered in two configurations, one without a sideport and one with a sideport attached to a segment of extension tubing terminating in a 3-way stopcock. |  |  | July 02, 2010 | 2010 | Greatbatch Medical | The outer sterile barrier (pouch) packaging for the 12 Fr device is not intact. No breaches in the inner sterile barrier (tray/tyvek lid) have been identified. The OptiSeal introducer and accessories sealed in the tray are sterile. Infection may result if a sterile field is contaminated by the outer surface of the inner tray. | Packaging | Greatbatch Medical issued an "Urgent Medical Device Recall letter dated July 12, 2010 via email to consignees. The letter described the problem and affected device. Consignees were instructed to immediately discontinue distribution of the product and return the unused units to Greatbatch Medical and to notify their customers to immediately return unused units. For shipping assistance, questions or assistance in notifying accounts, contact Shannon Springer at 763 951-8244 or Kimberly Briggs, 716 759-5702. |  | 216 | 211 5-pack boxes | Worldwide Distribution: United States, including the states of MN and CA. | Terminated | December 17, 2011 | 533 | 216
Z-2293-2010 | 55612 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator/Monitor with Q-CPR |  |  | May 06, 2010 | 2010 | Philips Healthcare Inc. | Reduced sensitivity in the force measurement sensor can result in inaccurate feedback to the user on how well he or she is releasing pressure after each chest compression. This can result in incomplete release of chest compression, which will reduce CPR effectiveness. | Device Design | Philips issued an "Urgent Medical Device Recall notification/Field Safety Notice" dated April 2010 to all consignees. The letter described the issue, how to identify affected devices, and required actions. The consignees were instructed to remove affected devices from service. Philips will contact consignees to arrange on-site replacement of affected meters with meters containing revised labeling and calibration software. The new meters will indicate that meter recalibration is required after approximately 50,000 compression cycles. Consignees returning devices to Philips for routine recalibration will receive free replacement. All CPR meters will be replaced with re-designed meters when they become available.For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377. |  | 72 | 72 units | Worldwide distribution in the countries of USA, Canada, Australia, China, Germany, Iceland, India, Libyan Arab Jamahiri, Netherlands, Saudi Arabia, Spain, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | February 19, 2013 | 1020 | 72
Z-2296-2010 | 56420 | No | GDI | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | V. MuellerÂż Peanut Sponges, Catalog number 23275-470, Sterile, For Single Use Only, Size: 3/8, X-Ray Detectable In Holders, Distributed by Cardinal Health, McGaw Park, IL 60085 USA. There are various uses including as gauze/sponge internal, x-ray detectable and as dissectors. |  |  | March 01, 2010 | 2010 | He, Inc | Sterile product packaging contains weak seals, which may result in compromised product sterility. | Packaging | Harod Enterprises, Inc. (HE) issued a "Notice of Recall" letter to Cardinal Health/V. Mueller dated March 1, 2010 identifying the issue and affected product. Cardinal Health/V. Mueller will issue a subrecall letter to all their customers (end users) instructing them to discontinue use of the affected lots and return for replacement. and will track all responses and removal of the product from their private label customers. HE will repackage and/or replace all returned product.Consignees can contact Harod Enterprises, Inc. at 1-706-228-5165. |  | 55100 | 551 cases of 100 packages per case | Nationwide: including state of IL. | Terminated | November 02, 2011 | 611 | 55100
Z-2295-2010 | 56365 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK TWO AST-N085 REF 22 225 For In Vitro Diagnostic Use - each test card is packaged into a foil pouch with 20 pouches per box. The firm name on the label is bioMerieux, Inc., Durham, NC and bioMerieux, sa, France.This is an Export Only test kit that is not marketed in the U.S. |  |  | July 29, 2010 | 2010 | Biomerieux Inc | Foil pouch was punctured by drug susceptibility cards during packaging, allowing moisture into pouch, resulting in inaccurate and/or incorrect drug susceptibility results. | Packaging | BioMerieux issued an Urgent Product Removal Notice, dated July 29, 2010 to consignees, identifying the reason for the recall and the affected lots. Customers were instructed to discontinue use of and destroy any remaining inventory of the affected lot, and to notify their customers who received the product. All consignees were instructed to complete and return the Customer Acknowledgement and Product Replacement Form. These cards will be used to distribute replacement cards. Consignees can contact BioMerieux at 1-314- 731-8667. |  | 25940 | 1297/ 20 card boxes | Worldwide Distribution: Country of Canada only. No product was distributed in the US. | Terminated | November 02, 2010 | 96 | 25940
Z-2310-2010 | 56218 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Genzyme Biosurgery Essentials Kit, Item Number 82004 |  |  | July 13, 2010 | 2010 | Genzyme Corporation | Revised Labeling of Carticel Essentials Kit clarifiesthe non-sterile packaging of the outer clear plastic tray which should not be opened in the sterile field. | Labeling design | Genzyme notiifed physicians with the Dear Doctor letter on July 13, 2010 via UPS Ground delivery. The letter informed healthcare workers of the recent incidents and advise them to heed the Essentials Kit label and exercise proper handling of the Kit.. A fax return form will be provided for the physician to sign and return to Genzyme to acknowledge receipt of the notification. ContactGenzyme Medical Information at 800-453-6948, option #3 for any additional information or contact your Genzyme sales representative with any questions. |  | 3132 | 3132 kits | Nationwide | Terminated | January 30, 2012 | 566 | 3132
Z-2312-2010 | 50214 | Yes | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | Pediatric Paddles attachments used with LIFEPAK 9, LIFEPAK 9P, LIFEPAK 12 and LIFEPAK 20/20e Defibrillator/Monitors. |  |  | November 17, 2008 | 2008 | Physio Control, Inc. | Sterilization efficacy could not be confirmed for sterilization methods on External Sterilizable Paddles and Pediatric Paddle attachments. | Labeling False and Misleading | On 11/14/08, Physio Control began sending the Urgent User Information letter via certified mail to their customers. The letter notified the customers about the modification to the cleaning and sterilization instruction of the External Sterilization Defibrillation Paddles and Pediatric Paddle attachments used with LIFEPAK 9, LIFEPAK 9P, LIFEPAK 12, and LIFEPAK 20/20e Defibrillator/Monitors The customers are informed to discontinue use of the gas sterilization methods listed as 12/88 Ethylene Oxide (EO) or 100% Ethylene Oxide (OE). They are notified that they can continue to use the sterilization method such as STERRAD per existing Operating Instructions, but limit sterilization cycles from 100 to 50 cycles.The customers are advised that prior to and after each use, they should examine cables and connector for damage or signs of wear such as loose connections, damaged pins, exposed wires and cable connector corrosion and examine paddles for scratched or pitted surfaces and stop using product if damage has occurred. Within 45 days, the firm will update the Operating Instruction that will include revised Sterilization Guidelines on a website www.physio-control.com/products/product-support. Customers can call Technical Support at 1-800-442-1142, #5 or they can visit the firm's website at www.physio-control-notices.com/extpaddles regarding the recall. |  | 1689 | 1689 units | Devices were distributed to medical facilities through out the US and WORLD WIDE to Canada, Germany, India, Netherlands, Australia, Hong Kong, Japan, Brazil, Panama, Russia, Korea, Malaysia, Venezuela, and United Arab Emirates(UAE) . | Terminated | September 28, 2010 | 680 | 39874
Z-2334-2010 | 53427 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Total Knee System Primary Tibial Baseplate; |  |  | September 01, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | A lot for lot mix up occurred between the Triathlon Total Knee System Primary Tibial Baseplate # 4 Cemented, Catalog # 5520-B-400, Lot SH8NT and the #2 Cemented, Catalog # 5520-B-200, Lot SH7RP. | Mixed-up of materials/components | The firm, Stryker Orthopaedics, sent an "URGENT PRODUCT RECALL" letter dated September 1, 2009, via FedEx to all customers. The letter described the product, problem and action to be taken by customers. The customers were instructed to examine their inventory and hospital locations to identify the affected product; retrieve and return all affected products via Inter-Org Transfer and ship to Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, 07430; Attn: Regulatory Reporting, and attach the fluorescent orange PRODUCT REMEDIATION sticker and mark the outer box with the words "PRODUCT RECALL; complete and fax the attached Product Accountability Form and spreadsheet to (201) 831-6069 within 5 days of receipt of the notice. If you have any questions, please contact the Manager, Divisional Regulatory Reporting at (201) 831-5970. |  | 6 | Lot Code: SH7RP - 6 in the US. Lot Code SH8NT was not received in the US. | Nationwide distribution: AL, AZ, CA, MT, NJ and TX. | Terminated | August 27, 2012 | 1091 | 6
Z-2313-2010 | 56284 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | VP 2000 Processor Heated Reagent Basins; automated slide stainer; Abbott Molecular Inc., Des Plaines, IL 60018; The basins are sold under the the following part numbers: |  |  | July 20, 2010 | 2010 | Abbott Molecular | The VP2000 Processor Heated Reagent Basins crack along the bottom seam, have non-uniform coating of interior corners, may be warped, and some may have incorrect dimensions. | Reprocessing Controls | Abbott sent Urgent Product Recall letters dated 7/20/10 to the VP 2000 Processor customers on same date via FedEx, informing them that the heated reagent basins could crack along the bottom seam, have non-uniform coating of interior corners, warping of the aluminum housing or incorrect dimensions. These deficiencies could result in leakage or the inability to place the basins into the VP 2000 Processor, resulting in possible exposure to chemical reagents, a slippery floor hazard, and/or delay or failure to start a protocol run. An Abbott Molecular customer service representative will call them to coordinate replacement of all of their heated reagent basins before 10/31/10. Customers were to complete and return the Customer Reply form provided. Any questions were directed to their local Abbott Molecular Representative.*** A second recall letter was sent to all customers on 12/21/10 explaining that the VP 2000 heated reagent basins previously inspected and determined to be acceptable could develop a leak, and that they would be contacted by an Abbott Molecular Customer Service Representative or Field Service representative in early 2011 to coordinate replacement of all of their VP 2000 heated reagent basins. *** |  | 219 | 219 basins | Worldwide Distribution -- Australia, Belgium, France, Germany, India, Italy, Japan, Pakistan, Poland, Russia, Singapore, South Korea, Switzerland, Taiwan, Thailand, United Kingdom, Uruguay, and the United States. | Terminated | April 19, 2012 | 639 | 219
Z-2315-2010 | 56064 | No | OAH | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | E.faecalis/OE PNA FISH |  |  | June 08, 2010 | 2010 | AdvanDx, Inc. | Some strains of streptococus anginosus cross-reach with E. faecalis/OE PNA FISH to produce a false positive green signal. | Device Design | The firm, AdvanDx, sent a letter dated June 8, 2010, and email notification on June 11, 2010, to all customers. The letter described the product, problem and action to be taken by customers. AdvanDx notified the customers that there was a change to the package insert under "limitations" to include the statement "some strains of Streptococcus anginosus produce a false green positive fluorescence due to sequence similarities and enclosed a package insert for all E. facealis/OE PNA FISH (KT003) kits. The customers were instructed to replace the Package Insert (revision B) included with any kits in their inventory with the enclosed Package Insert (revision C). Note: All shipments of E. faecalis/OE PNA FISH, as of June 14, 2010, will contain this new package insert.If you have any questions or requests for further information, please contact AdvanDx Technical Support at US (Toll Free): 1-800-376-0009 or email techsupport@advandx.com. |  | 541 | 541 units | Worldwide distribution: USA and countries including Danmark, Finland and Sverige (European). | Terminated | June 06, 2011 | 363 | 541
Z-2335-2010 | 53811 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Abbott Spine PathFinder, Cannulated Poly Screwdriver II - MOD, non-sterile, Manufactured by Abbott Spine, Inc., Austin, TX; Ref. 1160-2. |  |  | January 12, 2009 | 2009 | Zimmer Inc. | The screwdriver may fracture, resulting in surgical delays and/or fragments going into the surgical site. | Other | The sales force was provided with new instructions for use on 1/12/09 and again in 05/09. On April 6, 2010 the firm sent out an Urgent Medical Device Correction to it's field representatives for distribution to it's customers. The field representatives were to distribute the forms and provide a list of the customers that recieved the recall notice.For questions or assistance, please contact Zimmer Spine at 800-326-0635. |  | 691 | 691 | Worldwide. | Terminated | December 22, 2010 | 709 | 691
Z-2321-2010 | 56331 | No | ESW | Esophageal prosthesis. | General &amp; Plastic Surgery | Gastroenterology/Urology | 510(k) | Boston Scientific Ultraflex Covered Esophageal Stent System - Proximal Release (18/23/15/12) |  |  | July 15, 2010 | 2010 | Boston Scientific Corporation | Units have excessive fraying of the deployment suture thread, as a result, these devices may experience a suture break during deployment. | Process control | Boston Scientific issued "Urgent Medical Device Recall - Immediate Action Required" letters dated July 15,2010 to consignees. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. Foreign accounts received notification translated into the local language and with changes that are applicable to local regulations.Boston Scientific can be contacted at 1-508-683-4427. |  | 10 | 10 units | Worldwide Distribution: USA, and the countries of: Austria, Germany, Finland, France, UK, Ireland, Iran, Italy, the Netherlands, Portugual, Spain, and Sweden | Terminated | October 14, 2011 | 456 | 118
Z-2314-2010 | 56402 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | PERI-LOC (R), 3.5 MM ANTEROLATERAL DISTAL TIBIA PLATE, 10 HOLE, RIGHT - 148 MM, REF 71800710, QTY: (1), STERILE R, CE 0086, Smith &amp; Nephew, Inc., Memphis, TN 38116, USA, EC REP: Smith &amp; Nephew Orthopaedics Gmbh, Tuttlingen, Germany. Adult and pediatric patients, as indicated for pelvic, small, and long bone fracture fixation. |  |  | July 20, 2010 | 2010 | Smith &amp; Nephew Inc | Inner packaging of sterile product was not sealed. Lack of sterility assurance. | Packaging | Smith &amp; Nephew issued an "Urgent:: Medical Device Recall" letter dated July 20, 2010 to consignees (International distributors), identifying the issue and affected product. Consignees were requested to identify and quarantine affected stock, and cease distribution. Consignees were instructed to contact Smith and Nephew for a Return Authorization number. The Return Authorization number and an inventory of affected devices are to be recorded on the Verification Section of the notification letter and faxed to the firm. Affected stock is to be sent to Smith &amp; Nephew. Consignees were requested to notify anyone to whom the product was distributed to also contact the firm for return of the product. Smith &amp; Nephew can be contacted at 1-901-399-5441. |  | 10 | 10 units | Worldwide Distribution: INTERNATIONAL ONLY: Countries of China, UK, Italy, Australia, and South Africa | Terminated | November 17, 2011 | 485 | 10
Z-2339-2010 | 52830 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba Aquilion ONE CT SystemTSX-301 A/2D |  |  | February 03, 2009 | 2009 | Toshiba American Medical Systems Inc | ECG Gated Dynamic Volume Prospective scan may stop if preset eXam Plan is confirmed and CTA On is changed to CTA OFF, due to exposure control PWB. | Software design | Toshiba America Medical System Inc (TAMS) sent an Urgent: Medical Device Correction dated January 14, 2009 to their consignees. The letter described the issue and provided instructions to test the device to determine if it is an affected device. If the device is affected, contact the Toshiba service representative at 800 521-1968. Consignees were requested to complete and fax a response form to Toshiba. |  | 16 | 16 | Nationwide Distribution: including the states of AL, CA, FL, GA, KY, MA, MD, MT, NV, NY, OH, PA, and Puerto Rico. | Terminated | September 02, 2011 | 941 | 16
Z-2342-2010 | 54265 | bed, manual | FNJ | Manual adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Eye surgery stretcher with crank fowler, model 1079 Stryker Medical 3800 E. Centre Ave Portage MI 49002. |  |  | February 26, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The fowler of the bed is raised by a crank. The crank consists of a steel screw that moves through a brass nut. Over time the screw can wear the threads on the brass nut. The screw and the nut can become misaligned, causing a loud squealing noise when the screw comes in contact with the drive tube. In addition the fowler can become inoperable over time with excessive wear of the brass nut. T | Device Design | The firm, Stryker Medical, sent an "Urgent Medical Device Correction" letter dated February 26, 2010, to all customers. The letter describes the product, problem and action to be taken by customers. The letter states that Stryker field service tech will contact the firm and perform device corrections within four months. The customers were instructed to locate the recalled stretchers and verify that there is no grinding noise when activating the fowler. If the fowler exhibited a grinding noise, the customers were instructed to remove the stretcher from service and contact the recalling firm at 800-STRYKER, option 3. Note: Do not put the stretcher back into service until it is repaired by our field service representative. The customers were also instructed to return the enclosed post card to confirm receipt of this notice; inform any new users and advise the firm of new locations if any of the listed stretchers were on loan or sold, and advise the firm of the obsolescence by providing the serial numbers if any of the listed stretchers were disposed of. If you have any urgent questions or concerns, contact (269) 324-6884. Normal hours are Monday-Friday 8am-5pm(EST). |  | 729 | 729 | Worldwide distribution: USA and countries including Australia, Canada, Chile, China, France, Germany, Greece, Spain, Italy, Japan, South Africa, Taiwan, Korea, Singapore, and UK. | Terminated | April 21, 2011 | 419 | 3732
Z-2337-2010 | 55699 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER LH 500 Series System |  |  | February 12, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter identified four (4) issues with the COULTER LH 500 Series System and the Unicel DxH 800 Coulter Cellular Analysis System:Issue 1:There is a potential for misidentification to occur when the system is configured in languages other than English or Chinese. The characters #,@,[,\,],`,{,¦,}, or ~ are substituted or omitted in other languages. The issu | Software design | The customer recall notification was initiated when Product Corrrective Actions letters with attached Customer Response form were sent on the week of February 15, 2010 to all customers who purchased the COULTER LH 500 Series System and the Unicel DxH 800 Coulter Cellular Analysis System. The letters will provided the customers with an explanation of the problem identified and a work around.Customers with questions regarding the Product Corrective Action were instructed to call 800-526-7694 in the UnitedStates or Canada, or contact your local Beckman Coulter Representative. |  | 1520 | 1373 units in US and 147 units in Canada | Throughout the United States and Canada | Terminated | September 10, 2010 | 210 | 3152
Z-2348-2010 | 56323 | No | NQP | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | NFlex Stabilization System 150mm |  |  | February 09, 2009 | 2009 | Synthes USA (HQ), Inc. | package integrity compromised | Equipment maintenance | The firm, SYNTHES, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 9, 2009 and an updated letter dated February 12, 2009, with updated expiration information to customers. The letters described the product, problem and action to be taken by customer. The customers were instructed to call Synthes at 1-800-479-6329 if they have any of the identified devices to obtain a Return Authorization Number; complete the enclosed Verification Section attachment and return this letter with the recalled product to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If they do not have the identified product ,complete the Verification Section attachment and return the document to Synthes by Fax 610-719-5120 or scan/email: lewis.lynne@synthes.com.If you have any questions, please call 800-620-7025 x 5454 or 610-719-5454 or contact your Synthes Spine Sales Consultant. |  |  |  | Worldwide distribution: USA including states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NJ, NM, NV, NY, OH, OR, PA, SC, TN , TX, UT, VA, WA, WI, and WV and countries including Denmark, Germany, Korea and Switzerland. | Terminated | September 15, 2010 | 583
Z-2355-2010 | 56214 | No | JNL | Thyroid autoantibody immunological test system. | Immunology | Immunology | 510(k) | Beckman Coulter Access Thyroglobulin Antibody II Reagent, Part Number: A32898. The Access Thyroglobulin Antibody II (TgAb) assay intended use is for the quantitative determination of thyroglobulin antibody levels in human serum and plasma using the Access Immunoassay Systems. |  |  | June 29, 2010 | 2010 | Beckman Coulter Inc. | The affected Access Thyroglobulin Antibody II reagent may produce inaccurate results due to a change in the thyroglobulin that was used in the kits. | Other | Beckman Coulter sent an "Urgent Product Corrective Action" letter dated June 29, 2010, with attached Customer Response form to the affected customers. The letter identified the affected product and issue.Customers were instructed to:(1) Discontinue use of Access Thyroglobulin Antibody II reagent lot numbers 916271, 917073, and 918433.(2) Review results in the context of the total clinical presentation of the patient, and rerun questionable results using another reagent.(3) Complete and return a response form within 10 days.(4) For replacement contact Beckman Coulter at (US) 1-800-526-3821, (Canada) 1-800-463-7828. Customer Technical Support can be contacted at 1-800-854- 3633. |  | 1071 | 1071 units | Worldwide distribution : United States, including states of CA, CO, FL, GA. IA, IL, KY, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI, and country of Canada. | Terminated | May 04, 2012 | 675 | 1071
Z-2340-2010 | 56364 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Long Nail Kit R2.0, Ti RIGHT 011x360 mm x 125 degrees; |  |  | July 23, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The nail has no thread for the set screw hindering insertion. | Process control | Urgent Product Recall letters were sent on July 23, 2010 by Federal Express to Stryker branches. A corrected letter was sent on July 26, 2010 to correct an error in the size of the screw. Additionally, letters were sent on July 26, 2010 to Hospital Risk Managers, Chief of Orthopaedics and Surgeons. The letter identified the affected product, explained the issue, discussed the potential hazard, and the risk mitigation. Customers are to examine their inventory, retrieve all affected product and return it; then reconcile all products on the Product Recall Acknowledgement Form and fax it back. Questions should be directed towards Richard Wolyn at 201-972-2100. |  | 12 | 12 units | Nationwide Distribution -- NJ, NY, DE, IN, WI, IL, PA and NM. | Terminated | August 06, 2012 | 745 | 12
Z-2350-2010 | 56309 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN 3200 CS and SL, Hematology Analyzer manufactured by Abbott Laboratories, Santa Clara, CA. Waste Reservoir Assembly Part number 8921174802. |  |  | April 16, 2010 | 2010 | Abbott Laboratories | Epoxy holding the two parts of the waste reservoir together fails, potentially resulting in biohazardous fluid leaks. | Nonconforming Material/Component | Product Recall letters were sent on 04/16/2010. The letters identified the affected product, explained the reason for recall, the patient impact/user safety, stated the necessary actions, and contact information. Users are to discontinue use until the Waste Chamber(s) have been replaced and follow the instructions in the Hazards section. In addition, appropriate protective equipment needs to be worn along with following their recommended laboratory procedures dealing with spills. Users are to contact their local Abbott Customer Support rep. US customers should be direct their questions to Customer Support. |  | 69 | 69 units total. | Worldwide Distribution -- CA, CT, ID, IN, and MA; and, countries of Australia, Brazil, Canada, China, Colombia, Mexico, Puerto Rico, Singapore, and Uruguay. | Terminated | December 06, 2010 | 234 | 69
Z-2354-2010 | 56489 | No | EFA | Dental handpiece and accessories. | Dental | Dental | 510(k) | Titan 3 5K Low-Speed Motor - Star Titan 3, 5,000 rpm- Lube Free |  |  | August 11, 2010 | 2010 | Dental EZ Stardental Division | The housing of 5K motor is labeled as a 20K motor. | Labeling mix-ups | DentalEZ group/Star Dental Division sent a "Medical Device Recall" letter to consignees on 8/11/10, informing them that the device's 5K motor was mislabeled as a 20K motor. Customers are to stop using of the device, and return to the manufacturer or to the distributor using supplied shipping labels. Replacement devices will be sent to customers at no charge. Consignees are to complete and return a form indicating whether they have any affected devices, and whether customers have been contacted about this recall. Customers can contact DentalEZ at 866-383-4636 . |  | 41 | 41 units | Nationwide Distribution: States of IN and TX | Terminated | December 13, 2010 | 124 | 41
Z-2376-2010 | 54625 | No | GXZ | Needle electrode. | Neurology | Neurology | 510(k) | 21G x 150mm Stimulation Catheter Continuous Nerve Block Tray, REF 554-021-150, Rx, Sterile, Distributed by: Stryker Instruments, Kalamazoo, MI. |  |  | February 01, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | Defects include, but not limited to: Not sterile, not properly labeled, drugs and swabs may be ineffective, and catheter tips may break in patients. | Process control | Stryker Instruments issued an "Urgent Medical Device Recall Notification" letter to consignee risk managers, directors and material managers on February 2, 2010, identifying the affected devices and product issues. The consignees were instructed to:1) Forward the recall notice and reply form to all affected locations. 2) Review all inventory and quarantine affected product.3) Destroy all affected product using attached instructions.4) Complete and return the Business Reply Form to the firm by fax.Stryker Instruments can be contacted at 1-9800-800-4236 ext 4354 |  | 18826 | 18,826 all products | Nationwide Distribution: United States, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, and UT. | Terminated | March 28, 2011 | 420 | 18826
Z-2380-2010 | 56561 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Vital Signs Inc. enFlow Disposable IV Fluid/Blood Warmer Cartridge. Medical point-of-care device designed to warm blood and IV fluids. |  |  | July 15, 2010 | 2010 | Vital Signs Inc | The enFlow Disposable IV Fluid/Blood Warmer Cartridge's male luer-lock fitting may be oversized and may not fit properly with standard female luer fittings. | Equipment maintenance | GE Healthcare issued an "Urgent Medical Device Correction" letter, dated July 15, 2010, to consignees, identifying the reason for the recall and the affected lots. Customers were instructed to discontinue use of and quarantine the affected devices, and contact Vital Signs Inc. Customer Service to arrange for return and replacement of the affected devices. Consignees were instructed to complete and return the Confirmation Form. Consignees can contact Vital Signs Inc. Customer Service at 1-973-956-5300 . |  | 406596 | 385,820 domestic - 20,776 international | Worldwide Distribution; USA (all states), and to the countries of Australia, Ecuador, Finland, India, Israel, Latvia, Lebanon, Malaysia, Mexico, Morocco, New Zealand, Panama, Philippines, Qatar, Saudi Arabia, Singapore, Switzerland, the Netherlands, and Vietnam. | Terminated | April 27, 2015 | 1747 | 406596
Z-2384-2010 | 53819 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Whatman Body Fluid Collection Paper- 903 California Newborn Screening Card Ref 110534 |  |  | October 29, 2009 | 2009 | Whatman, Inc. | Neonate newborn screening card demographic and filter portion may be detached | Process control | The firm, Whatman, issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 29, 2009, to the customer. The letter described the product, problem and action to be taken by the customer. The letter instructed the customer to: inspect the card and ensure the 903 Paper Portion is firmly and completely fixed to the Demographic portio, prior to recording patient demographic details on the card and obtaining a patient sample; discard the card, obtain a new card and inspect as above, if the 903 Paper Portion of the card partially or completely detaches from the Demographic portion of the card; and to please staple the card stub at two points as specified in the illustration, to prevent any opportunity for 903 paper detachment from the demographic portions.If Whatman customers/users have any questions in relation to this customer notification, please contact, either Product Manager at +1 732 885 6570,email: lili.lee@ge.com or Quality Assurance at +1 207 459 7557, email: heidi.kuell@ge.com. |  | 488000 | 488,000 cards | Nationwide distribution: CA | Open, Classified |  |  | 488000
Z-2382-2010 | 56315 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Impella 5.0 Circulatory Support System |  |  | July 05, 2010 | 2010 | Abiomed, Inc. | Pump may separate at the motor housing while device is explanted | Process control | ABIOMED issued the recall notification letter on July 5 , 2010 to the one affected customer (France) and requested that they segregate and return any unused pumps. The pumps in Canada were already used.If you have any questions, contact +49 241 8860-0 or e-mail europe@abiomed.com. |  | 10 | 10 units | International only: Canada and France | Terminated | September 10, 2010 | 67 | 10
Z-2383-2010 | 56488 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | 3M Red Dot Monitoring Electrode, Catalog #2560 &amp; Novaplus Universal Monitoring Electrode #V2560, used for ECG monitoring. |  |  | July 27, 2010 | 2010 | 3M Company / Medical Division | The sensing element of the 3M Red Dot Monitoring Electrode, Catalog # 2560 and Novaplus Universal Monitoring Electrode #V2560 may have corrosion. In an emergency situation where defibrillation is necessary and electrodes are already in use, the electrodes may malfunction and provide a delayed or no trace response. | Process control | 3M sent an "Urgent Medical Device Recall" letters and a "Distributor Product Recall Form" dated July 29, 2010 to customers and distributors of the 3M and Novaplus electrodes. The letters described the 3M Red Dot and Novaplus Universal products, problems and actions to be taken by customers. Customers were instructed to examine their inventory and quarantine the affected lots pending return to the firm. After completing and returning the enclosed Customer Product Recall Form to 3M, customers will be issued a Return Authorization Number and provided with instructions for return of the affected devices. Distributors were instructed to notify customers to the end user level and have them return affected lots to the distributor for forwarding to 3M, or forward the customer letter to end-users to instruct them on how to contact 3M for recovery of the device. The 3M Customer Helpline can be contacted at: 1-800-228-3957. |  | 553150 | 11,063 cases (20 pouches/case, 50 electrodes/pouch) | Nationwide Distribution: States of CA, FL, IL, KS, KY, MI, MO, NE, NJ, NY, OH, PA, RI, SD, TN, VA, and WI. | Terminated | April 19, 2012 | 632 | 553150
Z-2385-2010 | 54460 | No | JMO | Iron-binding capacity test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension IBCT Flex Reagent Cartridge. The total (IBCT) method used on the Dimension system is an in vitro diagnostic test intended to quantitatively measure Total Iron Binding capacity in human serum and plasma. |  |  | December 16, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Test produces falsely elevated IBCT results and Abnormal Reaction test report messages on heparinized plasma samples. | Nonconforming Material/Component | Siemens issued an "Urgent Field Safety Notice" letter dated December 2009 to users identifying the issue and affected product. Customers were instructed not to use heparinized plasma samples for IBCT testing. IBCT Instructions for Use (IFU) has been updated to reflect that heparinized plasma is no longer an acceptable sample type. Beginning with lot ED0341, the new IFU will be packaged in the carton with the IBCT FlexÂż reagent cartridges. Consignees were requested to forward the notification letter to anyone to whom the product was distributed. The Siemens Technical Solutions Center can be contacted at 800-441-9250. |  |  | unknown - all lots involved | Worldwide Distribution: USA (all states) and the countries of Canada, Switzerland, New Zealand, Australia, Japan, and South Korea. | Terminated | October 25, 2010 | 313
Z-2386-2010 | 54624 | No | LAN | Digoxin test system. | Toxicology | Toxicology | 510(k) | ONLINE TDM N-acetylprocainamide, Part number 04642520190 N- acetylprocainamide reagent Roche Diagnostics Corporation Indianapolis, IN. |  |  | February 05, 2010 | 2010 | Roche Diagnostics Operations, Inc. | The N-acetylprocainamide reagent lot 14456200 shows an incorrect expiration date of 7/31/2011. The correct expiration date for the lot is 1/31/2011. | Error in labeling | Roche issued an "Urgent Medical Device Removal" letter dated February 2010 to consignees identifying the issue and affected product. Customers were requested to complete the attached form and return in to the firm to obtain replacement product. After receipt of replacement reagent, the customer should discard the affected product product. If unaffected lots are not available, customers may use the affected lot until replacement product is received. Use of the affected product before the corrected expiration date of January 31, 2011 should not produce erroneous test results. Roche Diagnostics Technical Support can be contacted at 1-800-428-2336. |  | 40 | 40 kits | Nationwide Distribution: USA states of AL, CT, NY, PA, and TX. | Terminated | November 01, 2010 | 269 | 40
Z-2332-2010 | 56162 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | BagEasy manual resuscitation devices are used for patients requiring complete or intermittent ventilatory support. It can provide positive pressure ventilation or oxygen for the spontaneously breathing patient with a face mask or through an artificial airway. |  |  | June 22, 2010 | 2010 | Westmed Inc | This recall was initiated after Westmed, Inc. became aware of a potential for disconnection at the patient port retention ring assembly of the Bag Easy manual resuscitation device. | Nonconforming Material/Component | The firm, Westmed, sent a letter dated June 21, 2010, to all customers. The letter describes the product, problem and action to be taken by customer. The customers were instructed to cease distribution immediately, notify any facilities that they may have distributed the product identified of the recall, call Customer Service Manager at 1-800-975-7987 x 1223 and/or email kpeterson@westmedinc.com to receive Return Authorization to return identified product, and complete and return the enclosed Bag Easy Recall Reconciliation Checklist.If you have any question, contact Customer Service at 1-800-975-7987 or 1-520-294-7987. |  | 22145 | 22145 | Worldwide distribution. | Terminated | March 22, 2011 | 273 | 24879
Z-2381-2010 | 56523 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Inc. Symbiq One- and Two- Channel Infuser, List Numbers 16026 and 16027 respectively. |  |  | March 26, 2010 | 2010 | Hospira Inc | If the administrator set is removed prior to the cassette carriage fully opening without closing the slide/roller clamp, unrestricted flow and delivery may result. | Device Design | Hospitra Urgent Device Correction Notification mailed on 03/26/2010 via federal express. Tracking done by delivery tracking. Distribution of product continues and has not stopped. Additional notification was issued in 09/2010. A software correction is planned. |  | 28485 | 28485 units | Worldwide distribution: USA, Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Puerto Rico, Singapore and the UAE. | Terminated | March 30, 2012 | 735 | 28485
Z-2387-2010 | 54743 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista Chemistry 2 Calibrator, Catalog number KC120. |  |  | February 26, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | There is incorrect barcode information, which can potentially cause the Alert Message "Calibrator Insert Missing" to be displayed. | Employee error | The recalling firm issued an Urgent field Safety Notice, dated February 2010, to all affected customers informing them of the problem. Customers have been told to run calibrators from cups when calibrating or request a replacement lot. The notification should be forwarded to that this product was further distributed to. Siemens Technical Solutions Center should be contacted at 800-441-9250, if they are any technical questions. |  | 1310 | 1310 cartons | Worldwide Distribution -- USA, Austria, Germany, Netherlands, Italy, France, Denmark, Portugal, Canada, South Korea, Malaysia, Australia, and New Zealand. | Terminated | October 13, 2010 | 229 | 1310
Z-2336-2010 | 55852 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Mayfield Infinity XR2 Base Unit, model A-2079, manufactured by Integra LifeSciences Corporation. The MAYFIELD Infinity XR2 Base Unit (A-2079) is an operating room table accessory intended to be used to support a patient during diagnostic examination and/or surgical procedures where a rigid support between surgical table and headrest, or skull clamp is necessary, positional freedom is required and where x-ray imaging modalities will be used. |  |  | May 19, 2010 | 2010 | Integra LifeSciences Corporation | Storage case incorrectly labeled. The labels were for the Infinity XR2 Skull Clamp storage containers but were inadvertently placed on the Infinity XR2 Base Unit storage cases. | Labeling mix-ups | The firm, Integra Life Sciences Corporation,contacted their customers via telephone and Field Correction Letter on May 19, 2010 to notify their customers of the problem. The firm indicated that their sales representatives will visit the consignees affected, correct the case labeles according to the specific instructions indicated, revise their account tracebility records to reflect the labeling change, then, sign and return the Acknowledgement of Reciept and Return Form for confirmation. For questions regarding the correction, please contact Integra Life Sciences Corporation at (877) 44-1114 or (513) 533-7923. |  | 13 | 13 Units | Product was shipped to the following states: IN, MO, NJ, TX and VA. Product was also sold in Belgium, Canada and New Zealand. | Terminated | October 27, 2011 | 526 | 13
Z-2393-2010 | 56333 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Invacare AC-Powered Adjustable Hospital Bed. |  |  | October 16, 2006 | 2006 | Invacare Corporation | Invacare Corporation, Florida issued a device correction in October 2006 for the IVC Bed Foot Actuator used on the IVC Bed Series, Models 5890 IVC and 5490 IVC. The actuators were not in the fully extended position and could not be mounted onto the beds. All affected consignees were notified of this correction in October 2006. | Nonconforming Material/Component | Invacare Corporation issued a notice of an Urgent Field Correction to Consignees on 10/18/2006. Consignees were instructed to check the serial number of their units to see if they were potentially affected. If so, the firm provided inspection and adjustment instructions of the foot actuator to prevent a likelihood of injury. After adjustment, the consignees were requested to complete and return the enclosed tracking sheet acknowledging receipt of the recall notice and correction of the bed(s). For questions or assistance regarding this correction, please contact the Customer Relationship Management Group at (800) 333-6900, or via e-mail at CRMGroup@Invacare.com. |  | 731 | 731 | Product was distributed throughout the U.S. | Terminated | September 13, 2010 | 1428 | 731
Z-2389-2010 | 56424 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SYNCHRON Systems Modular Creatinine (CREm) Reagent Kit, Part Number: 472525 |  |  | August 02, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports that black / brown particulates were observed in some bottles of the Creatinine Alkaline Buffer (Buffer lot M911427 and M911518) in modular Creatinine (CREm) reagent kit lots M911428 and M911519. Beckman Coulter have had reports of erroneous results when using the lots of Creatinine Alkaline Buffer listed above. Singlecreatinine d | Pending | The firm, Beckman Coulter, sent a "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 3, 2010, to all customers. The letter describes the product, problem, and action to be taken by customer. The customers were instructed to: (1) Please discontinue use of Creatinine (CREm) reagent kit lots M911428 and M911519 and discard appropriately.(2) Please contact Beckman Coulter at (800) 526-3821 in the United States or (800) 463-7828 in Canada to discussCREm kit M911428 and M911519 replacement options. Outside the United States and Canada, please contactyour local Beckman Coulter Representative.(3) Please complete and return the enclosed response form within 10 days so that we may maintain our records.(4) Please share this information with their laboratory staff and retain this notification as part of their Quality Systemdocumentation. If you have forwarded the affected products listed above to another laboratory, please provide a copy of this letter to them.If customers have any questions regarding this Product Corrective Action letter, they were instructed to please contact Beckman Coulter's Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/Customer Support/SupportRequests, or call (800) 854-3633 in the United States and Canada. |  | 3106 | 3106 units | Worldwide Distribution - USA including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, KS, LA, MA, ME, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, and WY and country of Canada | Terminated | July 02, 2012 | 700 | 3106
Z-2390-2010 | 56464 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira brand Symbiq One-Channel Infuser |  |  | February 22, 2010 | 2010 | Hospira Inc | The clamp may not secure to the IV pole and the pump may slip or fall off. The pump may fall on a person or pull the tubing out of the IV container. | Device Design | Hospira issued an "Urgent: Device Field Correction" Letter dated February 22, 2010 to all consignees, identifying the reason for the recall and affected product. Customers were informed that they will be contacted by a Hospira representative to replace the affected pole clamps. Consignees can contact Hospira Technical Support Operations at 1-800-241-4002. |  |  |  | Worldwide Distribution: USA (all states and Puerto Rico) and the countries of Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Singapore. and the UAE. | Terminated | April 01, 2011 | 403
Z-2392-2010 | 56133 | No | GEN | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BIOMET MICROFIXATION REF: 02-0461 REES INSULATED FORCEPS, 21CM, (8.25") (L), STAINLESS STEEL, NYLON, ***NON-STERILE PRODUCT***. The Rees Insulated Forceps are used during surgery to grasp, dissect and coagulate selected tissue. |  |  | June 25, 2010 | 2010 | Biomet Microfixation, Inc. | The forceps insulation may have cracks or breaks in the insulation that may cause burns to patients when used. | Manufacturing material removal | Biomet Microfixation issued a letter dated July 2010 to their consignees. Each consignee was ask to discontinue use of the affected product and return the product back to the company for credit. Biomet asked their customers to complete a "Inventory Reconciliation" form and fax the completed form to Biomet Microfixation. If the consignee has further distributed the product, the consignee is asked to contact their customers. Biomet Microfixation can be contacted at 904 741-4400, ext 468. |  | 130 | 130 units | Worldwide Distribution: USA including states of CA, FL, GA, ID, IL, IN, MA, MD, MS, NC, NH, NJ, NV, NY, OH, PA, TX, VA, WA, and WI, and countries of Norway, Canada, and Mexico. | Terminated | June 21, 2011 | 361 | 130
Z-2388-2010 | 55359 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GLUCm used on Synchron LX and UniCel DxC Synchron Clinical Systems, Part Number: 472500 ( ynchron LX20, LX20 PRO, LXI 725, UniCel DxC 600, UniCel DxC 600 PRO, UniCel DxC 600i, UniCel DxC 800, UniCel DxC 800 PRO, UniCel DxC 880i, UniCel DxC 680i, UniCel DxC 660i, and UniCel DxC860i ), |  |  | April 15, 2010 | 2010 | Beckman Coulter Inc. | UniCel DxC and Synchron LX instruments produced inaccurate (low) GLUCm results. False low results could affect or delay diagnosis or treatment. | Other | Beckman Coulter issued an "Urgent: Product Corrective Action" (PCA) letter dated April 2010 to all consignees, identifying the reason for the recall, the affected lots, and the potential for inaccurate GLUCm results due to bubble retention and short samples. Customers were instructed to:1) establish specific protocols for samples from defined patient populations or sample types, which could include running specimens using GLUCm in duplicate or using cartridge glucose.2) To reduce the possibility of inaccurate results, ensure that all elements are properly installed and secured, the sample probe is adequately tightened, and proper sample volume and integrity are maintained. A copy of the PCA letter will also be included in all new instrument ship kits.Consignees were instructed to complete and return an enclosed response form. Consignees can contact Beckman Coulter at 1-973-956-5300 . |  | 84186 | 84,186 units (79,071 in US &amp; 5,115 in Canada) | Worldwide Distribution: USA and Canada | Terminated | December 31, 2013 | 1356 | 84186
Z-2398-2010 | 55700 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clinical Chemistry Phosphorus assay |  |  | May 04, 2010 | 2010 | Abbott Laboratories Inc. | Product may give inaccurate results and aspiration errors with reagent configuration, including results outside of linear range &amp; calibration failure. | Other | Abbott Laboratories issued a "Product Recall: Immediate Action Required" letter dated May 4, 2010 to consignees. The letter described the product, problem and action to be taken by the customer. The letter instructed the customer to:1) Identify affected product in inventory2) Discontinue use and destroy any remaining inventory of the affected lot: 7D71-213) Complete and return an enclosed Customer Reply Form4) Order alternative kit sizes for Clinical Chemistry Phosphorus: LN 7D71-31 or 7D71-225) For AEROSET Systems, manually update parameters as described in the letter.6) For ARCHITECT cSystems, implement alternative kit sizes without system changes.7) Provide a copy of the letter to other laboratories that have been distributed the affected product.US customers can contact Abbott Customer Service at 1-877-4ABBOTT. Customers outside the U&gt;S&gt; are instructed to contact their local area Customer Service. |  | 3136 | 3136 | Worldwide Distribution: US, including Puerto Rico, and countries of Germany, New Zealand, Trinidad/Tobago, Cayman Island, Hong Kong, Chile, Curacao, Jamaica, Australia, and Bahamas. | Terminated | April 05, 2012 | 702 | 3136
Z-2395-2010 | 55478 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas 2020 Insulin Infusion Pumps |  |  | March 18, 2010 | 2010 | Animas Corporation | sensor could be damaged | Equipment maintenance | An Email was sent by the firm, Animas, to US patients and health care providers where email addresses were on file. Animas sent two "URGENT: MEDICAL DEVICE CORRECTION " letters one dated March 18, 2010, to all US customers, HCPs and distributors of One Touch Ping Glucose Management System and the Animas 2020 Insulin Infusion Pump and the other letter to all US patients, HCPs and distributors of Animas 2020 Insulin Infusion Pump, Animas IR 1200 and Animas IR 1250 (1200 plus). Animas also notified all global distributors and affiliates and sent the letter to customers and health care professional in those regions. The letters describe the product, problem and action to be taken by the customers. The customers were instructed to follow their User Guide instructions; Pay attention to all pump warnings; NEVER rewind the pump motor, load the insulin cartridge, or prime their infusion set while connected, and fill out the enclosed reply card and send it back to Animas as soon as possible.Should you have any questions or concerns, please call the Animas Customer Support Department at 1-866-949-1525 or call toll free number on back of pump at 1-877-937-7867. If you are traveling outside of the U.S., please call 610-644-8990 and press option 1. |  |  |  | Worldwide distribution: USA including states of CA,FL,LA, MA, MI, NY, OH, PA,TX, UT, VA and WA, and countries including Canada, United Kingdom, Finland, Germany, France, Hungary, Puerto Rico, Australia, New Zealand, Israel, Sweden, Czech Republic, Spain, Mexico and Norway. | Terminated | October 27, 2010 | 223
Z-2349-2010 | 56330 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Picasso/Picasso Lite lasers |  |  | April 15, 2010 | 2010 | AMD Lasers, LLC | Product is noncompliant with the Federal laser product performance standard in certain requirements. The ON/OFF key on some of the Picasso/Picasso Lite lasers could be removed with force while the key was in the ON position. | Radiation Control for Health and Safety Act | Letters, dated April 30, 2010, addressed to Dental Professionals were sent identifying the affected product and reason for the recall. The letter asked customers to check their units and, if applicable, AMD Lasers would replace the key assembly on their lasers. The service would be provided free of charge including round trip shipping charges to the corporate office. Customers are to call to schedule the service and obtain a RMA number and shipping instructions. Questions should be directed to 1-866-999-2635 or support@amdlasers.com. |  | 3000 | 3,000 units | Worldwide Distribution | Terminated | April 21, 2011 | 371 | 3000
Z-2400-2010 | 55610 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus CS Acute Care Troponin I TestPak ( Troponin I assay) |  |  | April 13, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | False elevated Stratus CS CTNI results | Process control | Recalling firm issued " Urgent Field Safety Notice" dated April 2010, informing users of the problem and to discuss with laboratory directors regarding the need to review previous test results, conduct patient follow up, and/or repeat testing for tests conducted on or after Nov 1, 2009. |  | 7500 | 7500 cartons | Nationwide Distribution | Terminated | October 25, 2010 | 195 | 7500
Z-2399-2010 | 56271 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Peregrine 20 GA Endo Ocular Adjustable Laser Probe |  |  | July 19, 2010 | 2010 | Peregrine Surgical Ltd | Product did not have proper 510(k) clearance | Other | A recall letter is being sent to all affected accounts with instructions to return the recalled product to the recalling firm. |  | 132 | 132 boxes | Product was distributed to hospitals and eye surgery centers in Florida, Maryland, Missouri, North Carolina, Ohio, Pennsylvania, Texas. | Terminated | September 30, 2010 | 73 | 132
Z-2401-2010 | 56268 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet HLM Tubing Pack; Bioline Coating |  |  | July 13, 2010 | 2010 | Maquet Inc. | The color of the tape on the arterial and the venous line were placed incorrectly, causing the connection of the arterial and the venous lines to be interchanged. | Process control | Recall notification letters were sent by Federal Express on July 13, 2010 to affected account and also an e-mail to the Maquet field representative who services the account. |  | 5 | 5 units | Distribution: Pennsylvania | Terminated | August 10, 2011 | 393 | 5
Z-2114-2010 | 56370 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Microwave Ovens, Model Number DG-2 |  |  | May 12, 2010 | 2010 | Duke Manufacturing Co. | The oven did not meet Federal Performance Standard for certain requirements. | Radiation Control for Health and Safety Act | Letters were sent to Duke Manufacturing Speed Oven Owners. The letters identified the affected product and stated the reason for recall. The letter also states that the firm instituted a repair program on all DG-2 Speed Ovens. The firm will repair of replace all affected products at no charge to the customer. Customer Service personnel will contact customers directly and coordinate the schedule of the required repairs. |  | 62 | 62 units | Nationwide Distribution | Terminated | April 17, 2012 | 706 | 62
Z-2402-2010 | 56336 | cover, mattress (medical purposes) | FMW | Mattress cover for medical purposes. | General Hospital | General Hospital | 510(K) Exempt | CareGuard Mattress Models CG10180 and CG10180A used with IVC and ValueCare Series of manual, semi-electric and electric beds. |  |  | October 11, 2007 | 2007 | Invacare Corporation | In October 2007 Invacare Corporation, Florida recalled CareGuard Mattresses, Models CG10180 and CG10180A, associated with IVC Bed Series Full Electric Beds, Semi-Electric Beds and Manual Beds and the ValueCare Bed Series Semi Electric Beds, when used with Invacare bed rails. Affected bed packages are BED24IVC, BED25IVC, BED24CA, and BED25CA. The mattresses when used in combination with the IVC a | Device Design | Consignees were notified of the recall by letter dated October 15, 2007 and entitled "CareGuard Mattresses Provided Since September 2006 In combination with IVC Bed Series Full Electric Beds, Semi-Electric Beds and Manual Beds, and the Value Care Bed Series Semi Electric Bed, and Full and Half-Length Rails". Depending on the mattress usage the consignees were instructed to either: A: 1) Move the bed rail cross bar from the 5th and 6th spring location on the head section of the bed to the 6th and 7th spring location on the head section of the bed in accordance with the enclosed instructions and document your actions on the enclosed tracking sheet. 2) Return of the Product Tracking Report indicating disposition of the recalled product and acknowledgement of receipt of the recall package. 3) The tracking sheet must be sent to Invacare by January 15, 2008 in order for you to receive the $10 credit. or B) 1) contact Customer Service and request part No. 6022, which is a kit that contains a 2" convoluted foam pad, warning label and installation instructions. Testing has shown that the addition of the foam pad on a new mattress with Invacare half rails / bed combination meets the Guidance requirements." 2) When you have completed placing the label and foam topper on all of your affected mattresses, return the tracking sheet to us and we will credit your account $10 for each mattress you have listed. the tracking sheet must be sent to Invacare by January 15, 2008 in order for you to receive the $10 credi |  | 19143 | 19143 | Nationwide Distribution | Terminated | June 08, 2015 | 2797 | 19143
Z-2403-2010 | 56335 | cover, mattress (medical purposes) | FMW | Mattress cover for medical purposes. | General Hospital | General Hospital | 510(K) Exempt | Mattress Model 5185 and 5185XL. |  |  | August 29, 2007 | 2007 | Invacare Corporation | In September 2007 Invacare Corporation recalled the innerspring mattresses associated with Invacare's IVC Bed Series Full Electric Beds, Semi-Electric Beds and Manual Beds and the ValueCare Bed Series Semi Electric Bed. These mattresses were manufactured between June 28, 2007 and July 12, 2007. These mattresses when used in combination with the IVC and ValueCare Series of manual, semi-electric a | Device Design | Consignees of the product were notified by letter dated 09/05/2007. Consignees were instructed to destroy the affected mattresses and to return the completed confirmation of the destruction form to Invacare. |  | 4896 | 4896 | Nationwide Distribution | Terminated | September 13, 2010 | 1111 | 4896
Z-2405-2010 | 46217 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | EasyDiagnost Classic has a radiographic tilting table. This system has a cassette film option and it does not have a Digital Spot Imaging (DSI) system. |  |  | January 04, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Potential for the tower (holding the serial changer and/or image intensifier) to slide unexpectedly and potentially injure a patient or user in its path. | Other | On 1/4/08, the firm issued the URGENT PRODUCT CORRECTION letter dated 1/3/08 to their consignees via certified mail, return receipt requested. The letter informs consignees the issue of the safety switch and counter weight steel band on the Diagnost 55/66/76Plus, Easy Diagnost, and Easy Diagnost Eleva X-ray systems. However, the Easy Diagnost Eleva X-ray system was distributed and applied only to foreign consignees outside US.If the consignees hear the primary steel band clicking/rattling noise while tilting the table to 90 degrees table or positioning the serial changer, the consignees are instructed to tilt the table base towards the 0 degree tilting position and do not use the tilting drive movement until the mandatory action by a Philips service representative has been performed. The Field Change Order (FCO) would be implemented on 01/23/08 and expected to be completed by July 2008. The firm's representatives will visit sites and will inspect all systems where the switch has been replaced or adjusted for proper operation of the safety sensor and switch. Two steel bands will be inspected and replaced if any abnormity is noted.In the mean time, consignees should contact Philips Call Center at 800-722-9377, #5, #2 and reference the FCO 70600029 with any questions. |  | 1 | 1 unit | Nationwide Distribution | Terminated | September 13, 2010 | 983 | 38
Z-2407-2010 | 56145 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension Tacrolimus Flex Reagent Cartridge, manufactured by Siemens Healthcare Diagnostics Inc., Newark, Delaware. |  |  | June 15, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Low patient sample recovery results in all affected lots, with possibility of inappropriate increase in tacrolimus dosage based on false low result. | Other | Siemens issued an "Urgent Field Safety Notice" dated June, 2010 to all affected customers. Customers were requested to immediately discontinue use and discard any remaining inventory of the TACR Flex reagent lots DC0306, CA0320, DA0334 and GA0348. The letter contains the following instructions: 1) Immediately discontinue use of the affected lots; 2) Discard all remaining inventory of the affected lots; 3) Contact Siemens Healthcare Diagnostics Technical Solutions Center for replacement product; 4) Complete and return the Field Correction Effectiveness Check by fax to 302 631-8467;4) Siemens is recommending that the content of the letter should be discussed with the laboratory director. If unexpectedly low results are obtained on a patient sample, caution should be exercised before decisions regarding therapeutic dosage adjustments are made. As stated in the package insert, "Confirmation of unexpected or atypical results by an alternative methodology is recommended prior to any adjustments in tacrolimus dosage. Also individual tacrolimus values cannot be used as the sole indicator for making changes in the treatment regimen. Each patient should be thoroughly evaluated clinically before treatment adjustments are made."Siemens Technical Solutions Center may be contacted at 800-441-9250 |  | 4580 | 4580 Cartons | Worldwide Distribution to USA states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, KS, LA, MI, MO, MS, NC, NE, NH, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA and WI, and countries of Austria, Belgium, Canary Islands, Egypt, France, Germany, Great Britain, Hungary, Italy, The Netherlands, Poland, Portugal, Romania, Saudi Arabia, Slovakia, Spain, Sweden, United Arab Emirates, Argentina, Brazil, Canada, Columbia, El Salvador, Venezuela, Australia, Hong Kong, India, Japan, Malaysia, New Zealand, South Korea, Singapore, Taiwan and China. | Terminated | December 16, 2010 | 184 | 4580
Z-2420-2010 | 55923 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Stryker Bur Guard REF 2296-301, Stryker Instruments, Kalamazoo, MI (USA) |  |  | June 02, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | The reason for the field action is that the inner diameter of the Bur Guards may be out of specification. If there is excess material on the inside of the Bur Guard at the location where the guard tapers, it may come into contact with the Nose Cone of the drill. This could result in an injury or burn to the patient and/or user. | Device Design | Stryker Instruments sent an Urgent Medical Device Recall Notification dated June 10, 2010, to their customers. The letter described the product the problem and the action to be taken by the customers. The customers were instructed to:1) immediately review this recall notification.2) forward this letter and the attached Business Reply Form (BRF) to all affected locations. 3) Immediately check and quarantine any affected product found.4) Destroy all affected product by breaking off at least 3 of the tabs on the bur guard and then dispose of the affected product in normal trash.5) Complete and sign the BRF and fax to Stryker Instruments Regulatory Department, 866-521-2762.NOTE: Upon receipt of the BRF, a replacement bur guard will be sent to the customer's account for each affected bur guard destroyed.For questions regarding this recall please contact Stryker Instruments:1-800-800-4236, ext. 3808jennifer.mars@stryker.com |  | 10982 | 10982 | Worldwide Distribution: US AND COUNTRIES OF SCANDINAVIA, AUSTRALIA, CANADA, CHILE, CHINA, BRAZIL, FRANCE, GERMANY, GREECE, SOUTH AFRICA, ITALY, JAPAN, MEXICO, SWITZERLAND, PORTUGAL, SINGAPORE and UK. | Terminated | April 21, 2011 | 323 | 10982
Z-2414-2010 | 55876 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen complete knee solution MIS total knee procedure stemmed tibial component fixed bearing precoat size 7 sterile, REF 00-5950-057-01, Zimmer Inc. Warsaw, IN. |  |  | April 26, 2010 | 2010 | Zimmer Inc. | The firm has received complaints of loosening of the implanted device requiring revision surgery. There have been 114 MDRs filed all reported that the device loosened and the patient required additional surgery to replace the device. | Other | The firm, Zimmer, Inc., sent an "URGENT DEVICE CORRECTION" letter dated April 26, 2010, to all customers. The letter described the product, problem and action to be taken by customers. The customers were instructed to destroy or disregard all previous versions of the surgical technique, review the attached revised surgical technique to familiarize yourself with the modifications and warnings, and complete and return the Health Care Professional Dcoumentation Form via fax to Zimmer, Inc. at (574) 372-4265. (Note: Zimmer made a modification to the surgical technique and instructions for use. They added a warning to fully cement and pressurize the anterior and posterior surfaces of the tibial component, and to strongly recommend the use of a drop down stem extension. This change was completed and approved on April 23, 2010.) For questions please contact Zimmer at 1-877-946-2761. |  | 68384 | 68,384 all products | Worldwide distribution: USA including states of AZ, IL, IN, KY, MO, NE, NM, NY, OH, TX, VA, WA, and WI; and countries of Austria, Belgium, France, Germany, Italy, Netherlands, Spain, Sweden, Switzerland, UK, Ausralia, China, Hong Kong, India, Japan, Korea, Malaysia, Nicaragua, Singapore, Tawain and Thailand. | Terminated | March 28, 2011 | 336 | 68384
Z-2422-2010 | 47477 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | DigitalDiagnost X-ray System is a stationary Diagnostic X-ray using Digital Solid State Imaging Detector. It is identified with Part Number 4512 201 04752 Software Version 1.5.1 and Part Number 4512 201 04753 Software Version 1.5.2. |  |  | December 13, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | When using DICOM print functionality, the printed images may contain data for the wrong patient. | Software design | Philips Healthcare issued an "URGENT PRODUCT CORRECTION " letter dated March 10, 2008 to their consignees. The letter described the product, problem and action to be taken by the customer. The letter instructed the customer to: 1) Validate that image and patient information are correct on hard copy prints.2) Reprint images that have mixed patient information, and re-check that image and patient information are correct.3) Do not use printed images if the information is incorrect.The firm will issue the Field Change Order (FCO) 71200025 and the field service engineers will visit each consignee and conduct a software upgrade to solves the printing problem. The consignees may contact the Philips Call Center at 1-800-772-9377, #5 #2 and reference "FCO 71200025". |  | 119 | 119 units | Nationwide Distribution: including the states of AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, NC, ND, NE, NJ, NM, NY, OH, OK, PA, SC, TN, TX, UT, VT, WA, and WI. | Terminated | September 14, 2010 | 1006 | 119
Z-2421-2010 | 56422 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Sapphire Probe Detector PCB Assembly (also called the Aspiration Bottom Sensor) used on the Cell-Dyn Sapphire Hematology Analyzer, Manufactured by Abbott Laboratories, Santa Clara, CA |  |  | April 05, 2010 | 2010 | Abbott Laboratories | Under certain conditions in closed mode of operations, the Aspiration Probe Detector PCB assembly may fail to operate properly and provide inaccurate results. | Other | Abbott Diagnostics Division of Santa Clara, CA initiated a field correction of the CELL-DYN Sapphire Probe Detector Assembly used on the CELL-DYN Sapphire Hemotology Analyzer, on April 5, 2010. The correction is being addressed through field service representatives and a mandatory Technical Service Bulletin (TSB). The TSB requires inspection, testing, and potential replacement of the affected assembly. The firm will not be issuing a seperate notification. For questions regarding this correction please contact (408) 982-4800, or fax (408) 982-4863. |  | 926 | 926 units | Product was distributed to 926 retailers throughout the US and to argentina, Australia, Brazil, Canada, Chile, China, Columbia, Germany, Hong King, Japan, Mexico, New Zealand, Panama, Paraguay, Peru, Singapore, South Korea, Taiwan and Venezuela. | Terminated | May 17, 2011 | 407 | 926
Z-2424-2010 | 56320 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | NeoMarkers Rabbit Monoclonal anti-Human Cyclin D1 Antibody (Clone SP4), immunochemistry Reagent and Kit, For In Vitro Diagnostic use, manufactured by Lab Vision Products, Fremont, CA, a subsidiary of ThermoFisher Scientific |  |  | March 05, 2010 | 2010 | Lab Vision Corporation | Two lots may be contaminated with Ki-67 antibody, which could cause abnormal staining patterns, and produce false positive results. | Material/Component Contamination | Thermo Fisher Scientific issued an "Advisory Notice" letter dated March 5, 2010 to all domestic customers and international distributors. The letter described the product, problem and action to be taken. Customers were instructed to:1) discard the affected lots, and2) retest specimens using a replacement lot supplied by the firm, and 3) complete and return the Confirmation of Receipt Form via fax to 1-510-979-5489International distributors were instructed to:1) place the affected lots and replace with a replacement lot supplied by the firm, and 2) provide customers with copies of the customer letter, and3) complete and return the Confirmation of Receipt Form via fax to 1-510-979-5489For questions, contact Thermo Fisher Scientific at 1-800-828-1628 or 1-510-771-1595. |  | 103 | 103 units | Worldwide Distribution: USA, including the states of CA, SC, IL NY, WI, PA, MN, AL, WA, MA, and KS, and the coutnries of Austria, Australia, Denmark, Finland, France, Germany, Hong Kong, Italy, Japan, Malaysia, the Netherlands, Norway, Pakistan, Poland, Portugal, Russia, the UK, and Turkey. | Terminated | March 21, 2011 | 381 | 103
Z-2423-2010 | 55763 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Inc., Dimension Vista Intelligent Laboratory System with Alpha Pinnacle 10kVA UPS External Battery Pack. Part numbers: 070200001, 070200011, 070200021; Dimension Vista 1500 and 3000T instruments. |  |  | May 05, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Defective batteries used in the Alpha Pinnacle P10 UPS External Battery Pack may overheat and cause a strong sulfuric odor to be released and a corrosive sulfuric acid solution leak to occur. | Nonconforming Material/Component | Siemens issued an Urgent Field Safety Notice in May 2010 to customers with information on what to do if the battery leaks, overheats, or vents an odor. Firm also issued a Support Bulletin to field service personnel to inspect and replace affected units. |  | 228 | 228 units | Worldwide distribution: USA, Australia, Austria, Belgium, Canada, France, Germany, Italy, Japan, Malaysia, Netherlands, New Zealand, Portugal, Slovenia, Spain, and Switzerland. | Terminated | October 25, 2010 | 173 | 228
Z-2419-2010 | 54682 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista ECREA Flex Reagent Cartridge |  |  | February 12, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | The ECREA test results may be falsely elevated or depressed. | Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notice dated February 2010 to all affected customers. The letter described the product, the problem, and the action to be taken by the customers. The customers were instructed to immediately discontinue the use of and discard the ECREA Flex reagent cartridge lot 09224AA. Customers are to contact their Siemens Customer Service Center at 800-241-0420 to report their additional no-charge replacement needs, and complete and return the enclosed Field Correction Effectiveness Check Form via Fax at 302-631-8467.For any technical questions regarding this Recall contact Siemens Technical Solutions Center at 800-441-9250. |  | 324 | 324 | Worldwide Distribution - USA, including the states of AZ, CA, FL, IA, IL, MT, NC, NY, OR, PA, TX, and WI and countries of Belgium, Italy, Canada, Germany and the Netherlands. | Terminated | October 13, 2010 | 243 | 324
Z-2430-2010 | 49489 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Vista Nuclear Magnetic Resonance Imaging System, Picker International, Inc. World Headquarters 595 Miner Road Cleveland, Ohio 44143, U.S.A.. |  |  | September 19, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Water may collect and freeze in the vent pipe. This can rupture the magnet vessel during a cryogen transfer, and release helium into the MRI suite. | Other | Philips issued an Urgent Correction letter dated 09/19/2008 to their consignees, identifying the reason for the recall and the affected products. The firm issued the Field Change Order (FCO) 78100232. Field service engineers will visit each consignee and conduct an inspection of the magnets venting system of Edge, Vista, Eclipse, Polaris, Infinion, and Panorama 0.6T MRI systems. The consignees should contact the Philips Call Center at 1-800-772-9377, #5 #2 and reference "FCO 78100232". |  | 16 | 16 units | Nationwide Distribution: USA., including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY. There is no distribution to foreign consignees. | Terminated | September 17, 2010 | 728 | 374
Z-2428-2010 | 56564 | No | GDS | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Safen' simple Ostomy Appliance Deodorant (8 oz , 1 oz package sizes) |  |  | August 16, 2010 | 2010 | BioMed Laboratories | The product may contain the mold Aspergillus ssp. | Material/Component Contamination | Biomed Laboratories, Inc. of Garland, TX is conducting a recall of the Reliamed Odor Eliminator Drops, Securi-T Ostomy Deodorant, and the Safe n' Simple (Sn'S) Ostomy Appliance Deodorant. Lots being recalled were distributed between June 9, 2009 though June 28, 2010. Biomed will work with their 3 consignees to inform sub-distributors and down to any firm that keeps the product in stock to have the product removed from further distribution and returned. If there are any questions or concerns, please contact Biomed Laboratories, Inc. at (972) 282-8008 or email them at biomedlabs.net. |  | 22253 | 22,253 (18,060 8oz bottles. 4,193 1oz bottles) | OH, FL, MI | Terminated | January 24, 2011 | 161 | 72259
Z-2435-2010 | 54668 | No | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | The Provent Nasal Cannula is a nasal appliance device placed in both nostrils intended to be used for transmission of respiratory airflow signals between the Provent Professional Sleep Apnea Therapy device and physiologic recorders. |  |  | December 03, 2009 | 2009 | Ventus Medical, Inc. | The catalog number printed on two Nasal Cannula Diagnostic Kits (MRS0231) was different that the catalog number printed on the bulk pack shipper box and the other Nasal Cannula Kits in the same shipper (MRS0208). No devices were used on any patient . | Labeling mix-ups | Ventus Medical, Inc. of Belmont, CA has notified all customers who received distribution from the affected manufacturing lot (1880783) to immediately return the product in question for replacement. If there are any questions regarding this correction, please contact the firm at (650) 632-4165. |  | 60 | 60 | Product was distributed in the US through a single distribution supplier. | Terminated | March 11, 2011 | 463 | 60
Z-2436-2010 | 56626 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | The Gyrus ACMI Pleatman Sac Tissue Removal System facilitates removal of specimens while minimizing contamination of the abdominal cavity or wound tract. The system contains one (1) Sac and one (1) obturator in a sterile, ready -to -use, single- use procedure tray. |  |  | August 03, 2010 | 2010 | Gyrus Acmi, Incorporated | Gyrus Medical, Inc. has issued a Urgent Medical Device Recall for the Gyrus ACMI product 004942-903 - Pleatman SacÂż Tissue Removal System, sold as lot 186540FC with an expiration date of 06/2015. As a result of a manufacturing error in the production of the trays used in this lot, it is possible that a crack could form that would compromise the package integrity, and possibly the sterility of th | Packaging process control | Gyrus Medical, Inc. of Maple Grove, MN issued an Urgent Medical Device Recall drawing attention to a possible package integrity compromise that could affect the product sterility. USA consignees were contacted via telephone on 8/3/10. The firm directed consignees to inspect their stock to confirm if any of the affected lot of product remains in their possession. Use of remaining affected product should cease immediately. Affected products should be return to Gyrus ACMI. For questions or concerns, consignees should contact the firm at (888) 524-7266. |  | 2 | 2 cases (20 devices) | UT, PA, Canada | Terminated | August 23, 2012 | 751 | 2
Z-2437-2010 | 53044 | No | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | SYNCHRON Gentamicin (GEN) Reagent |  |  | February 19, 2009 | 2009 | Beckman Coulter Inc. | The reagent produces false low test results. Beckman Coulter has confirmed low recovery of values at the low end of the therapeutic range (&lt;1.5 ug/mL, &lt;3.13 umol/L) with SYNCHRON Systems Gentamicin Reagent (GEN) lot M804552. A false low result can cause a patient to receive more medication than necessary. | Other | Beckman Coulter sent an Urgent Product Corrective Action (PCA) letter dated February 27, 2009 to customers identifying the affected product and issue prompting the recall. The letter asked them to discontinue use and discard the reagent lot, and contact their Customer Service Representative at (800) 526-3821 in the United States or (800) 463-7828 in Canada or contact their local Beckman Coulter Representative to discuss reagent replacement options. They were also asked to complete and return the enclosed response form within 10 days.Customers were directed to call the Customer Support Center (Hotline) at (800) 854-3633 in the US and Canada if they had questions in regards to the PCA letter. Outside the US and Canada they were told to contact their local Beckman Coulter Representative. |  | 1547 | 1,547 Kits | Worldwide Distribution - USA, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, and WV, and countries of Canada, Belgium, Denmark, France, Germany, Ireland, Lithuania, Netherlands, Spain, Switzerland, United Kingdom, New Zealand, Singapore, Egypt, South Africa, Taiwan, and United Arab Imarets. | Terminated | July 16, 2012 | 1243 | 1547
Z-2438-2010 | 55427 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Integrated Data Management System - Version 7.1, 7.2 and 7.3; |  |  | April 09, 2010 | 2010 | International Technidyne Corporation | A tabulation error in the Integrated Data Management Systems (IDMS) data management software that results in incorrect Electronic Quality Control (EQC) Summary Reports: The Operator and the Device Performance Summary Reports. The IDMS software is not used for the clinical management of individual patients. | Software design | ITC sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 10, 2010 to Distributors and First Consignee End Users identifying the affected product and affected summary reports, the reason for recall, and the actions to be taken by the distributors and customers. New Errata Instructions were enclosed with each letter and are to be put in the idms user guide for reference. The new instructions are to be followed when running reports for the ESV or TVT tests run by the Hemochron Response. The enclosed Customer Account Tracking Form is to be completed and returned to ITC indicating that the letter has been received and appropriate questions answered as to the use of the software system. Distributors are asked to check inventory for recalled product. If affected product is found, shipment should be stopped and placed on hold. Distributors should notify customers of the correction and provide them with a copy of the correction letter and updated instructions. This should be done by traceable mail. Distributors should return the account tracking form and return affected software disks for replacement.Questions should be directed to ITC Technical Support at 800-631-5945. |  | 5544 | 84 total idms software systems; 66 in US, 18 outside of US. | Worldwide distribution: USA including the states of AL, AZ, CA, FL, GA, IL, KS, LA, MD, MI, NC, ND, NJ, NM, NY, OH, PA, TN , TX, UT, VA, WA, and WI, and the countries of Italy, Turkey, Canada, France, Belgium, Germany, UAE, and New Zealand. | Terminated | May 25, 2011 | 411 | 5544
Z-2441-2010 | 55163 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Unicel Synchron Clinical Systems |  |  | January 18, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated due to " A loose Aliquot Probe tube fitting in the Closed Tube Aliquotter (CTA), may result in erroneous Access immunoassay results or Synchron chemistry results".This may result in dilution of the sample with Wash Buffer; patient results can potentially be affected for Access immunoassays or Synchron chemistries. Access Quality Control may also be affected if run throug | Other | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated January 18, 2010, to all affected customers. The letter provided the customer with an explanation of the problem identified, the affected product, and instructions for the customer. The letter requested customers complete the following actions for :(1) CTA Aliquot Probe Inspection and Maintenance:a. Check daily for signs of leaks or damage to the tube fitting by looking and feeling for liquid at the fitting, the end of the probe and the surrounding areas;b. The CTA Aliquot Probe tube fitting should be checked to ensure that it is connected tightly to the Aliquot Probe. Without loosening the tube fitting, ensure it is securely tightened;c. When replacing the CTA Aliquot Probe, ensure that the Aliquot Probe tube fitting is securely tightened, prime the CTA, inspect for leaks, and run Access Quality Control through the CTA;d. If a leaking Aliquot Probe is suspected, review all sample results for both Access and Synchron in light of the total clinical presentation of the patient including: symptoms, clinical history, data from additional tests, and other appropriate information;e. As this issue may result in dilution of the sample with Wash Buffer; it may be appropriate to discard and redraw any affected sample tubes run on the CTA. The letter also informed customers that Beckman Coulter is currently evaluating potential solutions for this issue and requested they share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded the affected products to another laboratory, they were asked to provide a copy of this letter to them.They were also instructed to complete and return the enclosed response form within 10 days. Customers were told to call Customer Technical Support Center (Hotline) at (800) 854-3633 in the United States and Canada with any questions. |  | 672 | 672 units for both systems | Worldwide Distribution - USA: including the states of AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, AND WV and the country of Canada. | Terminated | August 13, 2012 | 938 | 672
Z-2440-2010 | 56390 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TRABECULAR METAL Âż FEMORAL CONE AUGMENT - RIGHT, Part Number: 00-5450-012-32. |  |  | July 05, 2010 | 2010 | Zimmer Inc. | Zimmer, Inc. has determined that there is a potential for the packages to be mislabeled. Specifically, the left femoral cone augment component may be labeled as a right and the right femoral cone augment component may be labeled as a left. | Labeling mix-ups | Zimmer, Inc. sent Urgent: Device Removal letters, dated July 9, 2010, identifying the affected product and stating the issue. Specifically, the left femoral cone augment component may be labeled as a right and the right femoral cone augment component may be labeled as a left. The letter also discussed the clinical implications along with required actions. Customers are to perform a physical count of all affected product and record the data on the Inventory Return Certification Form included with the letter. In addition, customers should complete and return the Inventory Certification and User Facility/HCP Documentation Forms. Affected product along with the Inventory Return Certification Form should be returned. Zimmer, Inc. should be contacted at 1-800-613-6131 if there are any questions or need for additional information. |  | 50 | 50 total | Worldwide Distribution -- USA, Canada, Singapore, Australia, and Switzerland, Spain, UK, Sweden, and Israel. | Terminated | December 22, 2010 | 170 | 50
Z-2446-2010 | 56183 | No | IKT | Diagnostic electromyograph needle electrode. | Physical Medicine | Neurology | 510(k) | Disposable Detachable Monopolar Needle Electrode for electromyograph model number 110-725, for Physician use, manufactured by Chalgren Enterprises Inc., Gilroy, CA. |  |  | April 12, 2007 | 2007 | Chalgren Enterprises Inc | Sterility Compromised-- Investigation found that individual packaging seals were separating, which may compromise sterility. | Packaging | Chalgren Enterprises, Inc. removed all of the affected inventory from inventory and had all affected inventory retrieved from distribution. |  |  |  | Nationwide Distribution: USA states of CA, IL, NJ, OH and TX. | Terminated | March 24, 2011 | 1442
Z-2445-2010 | 56460 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | Primary (inner) Package Label: ACUVUE OASYS BRAND CONTACT LENSES for ASTIGMATISM with HYDRACLEAR Plus (senofilcon A) Johnson &amp; Johnson Vision Care, Inc. Rx Only, STERILE. Lot B006WBS1 -9.00D -1.25 CYL 80 , 2013/05 AXIS. (Correct labeling) |  |  | July 16, 2010 | 2010 | Vistakon | ACUVUE OASYS Brand Contact Lenses for ASTIGMATISM may be mislabeled on the outer (secondary) packaging carton. The outer (secondary) cartons labeled as Lot B006X641 -8.50D -1.25cyl. 100 Axis may actually contain primary packages from Lot B006WBS1 -9.00D -1.25cyl. 80 Axis. The inner (primary) package labeling is correct. Product was only distributed to Eye Care Professionals for use in proper f | Error in labeling | Vistakon issued an Important Product Notification dated July 28, 2010 to customers identifying the affected product, the labeling issue, and actions to be taken by the customer. Vistakon issued a separate Important Product Notification dated July 28, 2010 to distribuitors identifying the affected product, the labeling issue, and actions to be taken by the distributor. Customers are instructed to check their inventory for affected products and return a business reply card. If replacement -8.50 lenses are needed, the customer can contact Customer Service at 1-800-874-5178.Distributors are instructed to check their inventory for affected products and return them using an enclosed return label. Replacement product will be sent by the firm. Customer Service can be contacted at 1-800-874-5178.Accounts will be instructed to go to their inventory and look for the B006X641, -8.50 lot and open the secondary carton and verify the product inside. If it is the -9.00 product they can keep the product and re-slot it in the proper place within the kit (product is ok, nothing wrong with primary package or lens inside) and order a replacement -8.50 if necessary. Business reply cards will be supplied for accounts to return to third party indicating if they had any mix/mislabeled product. |  | 2123 | 2,123 lenses | Worldwide Distribution: USA including all states, District of Columbia and Puerto Rico, and the countries of Canada, Bahamas, Honduras, and United Arab Emirates. | Terminated | March 28, 2011 | 255 | 2123
Z-2447-2010 | 56327 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | PSD Gel that's provided with Peri-Strips Dry Staple Line Reinforcement with Apex Processing (PSD) and Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSD-V) products. Sterile, Rx only, SynovisÂż Surgical Innovations. |  |  | June 30, 2010 | 2010 | Synovis Surgical Inovation Div. of | SynovisÂż Surgical Innovations is voluntarily recalling PSD Gel that's provided with Peri-Strips Dry Staple Line Reinforcement with Apex Processing (PSD) and Peri-Strips Dry with Veritas Collagen Matrix Staple line Reinforcement (PSD-V) products. Internal investigation have shown there is a potential for small defects in the PSD gel tube pouch which may compromise the sterility of the outside of t | Packaging | Consignees were sent on 6/30/10 a Recall Notification letter dated June 30, 2010. The letter was addressed to "Risk Manager". The letter described the problem and product involved. They listed both the device and gel lot numbers since it is possible that some hospitals have separated the pouched gel tubes from the PSD/PSDV products. They advised consignees to quarantine the affected lots and to complete and fax the attached "Product Return/Replacement" form to Synovis Surgical Innovations. |  | 5384 | 3,989 gel tubes and 1,395 PSD &amp; PSDV devices. | Worldwide Distribution -- USA, including states of AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL,IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, and WI and countries of AUSTRALIA, GERMANY, GREECE, HONG KONG, NETHERLANDS, PUERTO RICO, ROMANIA, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | December 17, 2011 | 535 | 5384
Z-2443-2010 | 56121 | No | DTE | External pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | Oscor PACE 203H, Dual-Chamber External Pulse Generator |  |  | June 30, 2010 | 2010 | Osypka Medical Inc | The recall was initiated after Osypka Medical received two (2) complaints from their Japanese distributorregarding an incompatibility of the PACE 203H with some new types of 9 V batteries.The investigation revealed that PACE 203H of a certain manufacturing period malfunction(i.e., not stimulate) when a type of 9 V batteries with higher than usual open circuit voltage (OCV) is used. Failure t | Device Design | OSYPKA Medical sent "Urgent Medical Device Recall" letters dated June 30, 2010 and July 7, 2010 to the direct customers. The recall notice informed the customers of the Reason for Recall, the Potential Hazard, the Temporary Fix of the Problem, the Permanent Fix of the Problem and Recommended Actions to be taken by the customer. The customers were instructed to contact Osypka Medical via telephone (858) 454-0021 or e-mail to m.osypka@osypkamed.com to verify that the Model 3085 and/or PACE 203H is subject to this product removal, obtain a return-to-manufacturer-authorization (RMA) number ,and then return the affected devices to Osypka Medical.If you have any questions call (858) 454-0021 or e-mail to mail@osypkamed.com |  | 205 | 205 units (US Only) | Nationwide Distribution: USA including states of CA &amp; FL | Terminated | August 10, 2012 | 772 | 217
Z-2461-2010 | 49594 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Nuclear Magnetic Resonance Imaging System (MRI). Trace item 9896-030-14351 -1.5T Sense XL Torso coil 12nc on coil: 4535-671-4188x. |  |  | September 30, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | The cables of the coil may become too hot and cause burns when placed too close to the patient. While using the 1.5T Sense XL Torso coil, unwanted radiofrequencies (RF) interaction with the patient can occur. Especially in combination with high SAR levels and when the RF cables of the coil are too close to the patient, this RF interaction has resulted in RF burns to patients. | Other | Philips Healthcare issued URGENT-Device Correction Notices dated September 30, 2008 to their consignees. The notice identified the recalled device as 1.5 Intera and 1.5T Achieva MRI systems, described the risk to patients, and the actions to be taken by customers..The firm has issued Field Change Order (FCO) 78100242. Field service representatives will visit each consignee and conduct a hardware upgrade, applying a sleeve over the cable that will ensure proper spacing between the cable and patient, reducing the chance of RF interaction and burns.The consignees should contact the Philips Call Center at 1-800-772-9377, #5 #3, #1 and reference "FCO 78100242 - PATIENT CABLE SLEEVE FOR 1.5T SENSE XL TORSO COIL". |  | 47 | 47 units total (Achieva 1.5T and Intera 1.5T) | Nationwide Distribution: Throughout USA. No distribution to foreign consignees. | Terminated | September 30, 2010 | 730 | 47
Z-2464-2010 | 50066 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips BV Endura Mobile X-Ray system with the extended rotation option. |  |  | September 10, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | The C-arm rotation brake does not always function properly, and may not hold the C-arm position if the C-arm is not positioned in its maximum rotation and if it is accidentally bumped by medical staff or patients. Unwanted C-arm movement may cause the C-arm to collide with a patient. The brakes function properly if the C-arm is rotated to its maximum positions. | Other | Philips Medical Systems issued an Urgent Product Correction letter dated September 30, 2008 to consignees. The letter identified the affected device, what the problem is and under what circumstances it can occur, and the actions that should be taken by the customer or user in order to prevent risks for patients or users. Philips will contacdt customers to implement Field Corrective Action (FCO) #71800026), which describes a mechanical brake modification which will be installed by Philips. Until the Field Corrective Action has been completed, users should ensure that the C-arm is placed in the maximum extended position, or ensure that the C-arm is not touched during critical procedures if the C-arm is not in a maximum position.The Philips Call Center can be contacted at 800-722-9377 |  | 71 | 71 units | Nationwide Distribution: Throughout the US. | Terminated | January 07, 2011 | 849 | 1420
Z-2476-2010 | 50114 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Allura XPER FD20 Angiographic X-ray system. This system with Velara Generator can be identified as type number 9890 000 PRODUCTS 6501 * (number located on the generator cabinet). |  |  | October 23, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | A high voltage reliability issue resulting in the loss of the x-ray generation and this can cause an intermittent stop of fluoro acquisition and exposure during an examination without warning and hence no image on the display. | Other | The firm, Philips, sent "URGENT-Device Correction" letters dated November 10, 2008, to their customers and informed them about the reliability issue in the Allura XPER and Integris X-ray systems with the High Voltage Velera (HV) Generator. This reliability issue can lead to a system error causing intermittent loss of fluro and/or exposure acquisition, potentially resulting in the loss of the X-ray image on the display.The customers are advised if they experience High Voltage Generator reliability issue and the HV Generator stops working generating an O3EA error, a cold restart is required and reboot the system when the X-Ray image disappears.The firm will implement multiple mandatory Field Change Orders (FCOs) to solve the Generator reliability by updating the software for each X-ray system.The customers could contact Philips Care Center at 800-722-9377, #5, #4, #1 for any questions regarding this recall or 425-487-7469. |  | 838 | 838 units in US | Worldwide distribution. | Terminated | August 10, 2011 | 1021 | 1681
Z-2448-2010 | 56093 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | MONOPLANE X-ray Imaging Systems include the Allura XPER FD10 F/C, Allura XPER FD20, Integris H5000/Allura 9 F/C, Integris CV, and Allura 12-15 Monoplane. |  |  | February 04, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | A lock-up might occur inside the Velara Generator due to intermittent shorts inside the system. The problem could only resolved by a warm or cold system restart. In exceptional cases the power converters of the KVmA unit in the generator are damaged and the system can not be restarted. | Software design | On 02/04/2010, Philips Healthcare began sending out the URGENT-Device Correction Notices to their consignees. The firm issued two notifications, one for the Allura Integris and Xper FD10C, FD10F, FD10F/C and FD20 Monoplane X-Ray imaging systems with the Velara Generator. The second notification is for the Allura Integris Xper FD10 and FD20 Biplane X-ray imaging systems with the Velara Generator.This notice informs the consignees that a system locks up during a critical interventional case and it might impose a serious injury to patients. This lock up issue could only be resolved by a warm or cold system restart. A warm restart would take about 90 seconds or a cold restart would take about four minutes, and the device could not be used during this time. If the system does not work after the restart, the consignees should contact the Philips Healthcare Customer Service Representative at 1-800-772-9377 #5 (then follow prompts). A service representative will visit each consignee and conduct a software and hardware upgrade. |  | 1470 | 1470 units | Nationwide Distribution | Terminated | September 23, 2011 | 596 | 1737
Z-2425-2010 | 56190 | No | HAO | Nonpowered neurosurgical instrument. | Neurology | Neurology | 510(K) Exempt | Integra NeuroSciences Implants, NeuroBalloon Catheter, |  |  | July 02, 2010 | 2010 | Integra LifeSciences Corp | Certain NeuroBalloon Catheters have the potential to improperly inflate or deflate under certain conditions. | Process design | All customers and Integra Sales Specialists were notified of the recall on July 2, 2010 by fax, e-mail and letter send by Fed Ex overnight delivery. |  | 338 | 338 units | Nationwide, Canada, Taiwan and to Integra, Australia. | Terminated | May 29, 2012 | 697 | 338
Z-2465-2010 | 56179 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Spider Flex Monopolar Hook (Spider Surgical Instruments), TransEnterix, Inc., Durham, NC |  |  | June 22, 2010 | 2010 | Transenterix, INC | The hook end effecter from a monopolar device may become dislodged and fall into the patient when in use during laproscopic surgery. Retrieval from the patient may involve extended anesthesia time. | Other | TransEnterix issued a "Field Action" letter dated June 22, 2010 to consignees, identifying the affected product and actions to be taken by the customer. Customers were instructed to return the product to TransEnterix, Inc. using the address label and return information provided. The firm will replace the product at no charge. TransEnterix Customer Care can be contacted at 1-888-879-4111. |  | 46 | 46 boxes (6 units per box) | Nationwide Distribution in the states of AZ, FL, GA, NC, NY, OH, OR, SC and TX. | Terminated | January 31, 2011 | 223 | 46
Z-2466-2010 | 56303 | No | HEB | Unscented menstrual tampon. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Life Brand Regular Absorbency 40 Count Unscented Plastic Applicator Tampons |  |  | March 25, 2010 | 2010 | First Quality Hygienic Inc | Life Brand Regular Absorbency 40 Count Unscented Plastic Applicator Tampons package contained Super Plus Unscented Plastic Applicator Tampons. | Labeling mix-ups | Consignees were notified via telephone and letter of the need to withdraw product from the market and requested to return all product to manufacturer. |  | 456300 | 195 Cases/2340 Cartons | Canada | Terminated | October 06, 2010 | 195 | 456300
Z-2460-2010 | 56506 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Heavyweight Gantry Scanner, M/N 728306, S/N: 100016, 100018 -100025 |  |  | April 14, 2010 | 2010 | Philips Medical Systems (Cleveland) Inc | The gantry covers of the iCT have a latch at the base of the system which holds the cover in the closed position. During the course of maintenance the CT Field Service Engineer (FSE) will activate this latching mechanism which in turn has a potential to penetrate the insulation on cabling which carries 277 VAC. If the latch does come into contact with the cabling there is a potential that the FS | Other | Philips Healthcare sent an URGENT - Field Safety Notice letter, dated April 14, 2010, to their customers notifying them of the product, the problem, and the action to be taken. The letter asked them to identify the affected product by checking for the serial number tag (s/n 10016  10025), ensure that the electrical power was shut-down before opening or closing the gantry cover to prevent potential of an electrical shock, and that a Field Service Engineer would contact the facility to make arrangements for the implementation of a Field Change Order (FCO) before December to correct the problem. For further information or support concerning this issue, the customer was instructed to contact their Philips representative: Customer Care Center (1-800-722-9377, option 5 or their local Philips Healthcare Office). |  | 9 | 9 units | Worldwide Distribution - USA, including the states: DC, GA, IL, IN, NY and OH; and the countries of Germany and the Netherlands. | Terminated | September 24, 2012 | 894 | 9
Z-2450-2010 | 56563 | No | GAZ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Medi-Vac Non-Conductive Suction Tube with Maxi-Grip Connectors, Inner Diameter: 5 mm (3/16"), Length: 1.8 m (6 ft.); Sterile, for single use only, 50 tubes per case; Cardinal Health, McGaw Park, IL 60085 USA, Made in Mexico; catalog N56A and NEX56A (European label) |  |  | August 16, 2010 | 2010 | Cardinal Health | The packaging seals may not be intact, resulting in compromised sterility for the listed Medi-Vac Non-Conductive Tubing. | Other | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated August 16, 2010 to consignees, identifying the reason for the recall, the affected lots, and actions to be taken. Cardinal has inspected all packages in their warehouses and removed the affected products. All packages that have been inspected have a 3/4" green dot on the front of each unit and on the case carton label." Cardinal Health end users were instructed to segregate and destroy their affected product and request credit or replacement" Distributor customers were instructed to destroy their product and request credit or replacement" Distributors were instructed to conduct sub-recalls and return their inventory to Cardinal Health" To arrange for credit or replacement of the product destroyed, the accounts were instructed to call customer service at 800-964-5227 or 888-444-5440." All customers were requested to complete the enclosed Recall Acknowledgment Form acknowledging receipt of the letter and the number of tubing units destroyed, and fax it to 847-689-9101.Cardinal Health can be contacted at 800-292-9332.Cardinal sent recall letters dated 8/26/10 to their PreSource kit/module customers via UPS next day air, informing them that the Medi-Vac Non-conductive Suction Tubing listed may have packaging seals that are compromised. The accounts were informed that Cardinal has inspected tubing packages in their warehouses and removed the affected products. All packages that have been inspected will have a 3/4" green dot on the front of each package. Enclosed with the recall letter were recall labels to affix to each affected Presource pack/module in the account's inventory. These labels instruct the user NOT to use the affected tubing component and to remove and discard the component upon opening the kit. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of packs/modules on hand. Distributo |  |  |  | Worldwide Distribution, including USA. | Terminated | November 28, 2011 | 469
Z-2470-2010 | 56667 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DRIVER STAR #8 CANNULATED DART-FIRE (TM), COMPRESSION SCREW, SIZE: #8, REF: DSDS-0008, 1 EACH, Rx ONLY, NON-STERILE, EC REP Wright Medical EMEA, Krijgsman 11, 1186 DM Amstelveen, The Netherlands, Wright Medical Technology, Inc., 5677 Airline Road, Arlington, TN 38002 |  |  | August 26, 2010 | 2010 | Wright Medical Technology Inc | After autoclaving, the epoxy identifier (colored) band bled and there is concern that it may flake. | Process design | The firm initiated their recall the product by letter on 08/26/2010. The letters were addressed to distributors and hospital administrators with a fax back response form. Please contact Customer Service at 800-238-7117, if you have any questions. |  | 13 | 13 units | TX, PA, TN, VA, GA, and Spain | Terminated | August 02, 2012 | 707 | 13
Z-2467-2010 | 56686 | No | ING | Wheelchair elevator. | Physical Medicine | Physical Medicine | 510(k) | Braun wheelchair lift, model NL, Braun Corporation, Winamac, IN |  |  | August 15, 2007 | 2007 | The Braun Corporation | The wheelchair lifts threshold warning system may not operate and the wheelchair user can fall from the vehicle and be seriously injured. | Device Design | The Braun Corporation issued a Service Bulletin 34273 entitled "Lift Threshold Sensitivity Bulletin 34273 on October 2007. Braun Corporation also sent a "Safety Recall" letter to the Braun lift owner and letters to all customers on or about August 15, 2007. The letters described the product, problem and the action to be taken by the customer. The customers were instructed to notify the National Highway Traffic Safety Administration (NHTSA) and conduct a safety recall if they used recalled lifts as original equipment in a system or vehicle that they built; ensure that any associated customers are notified to discontinue the use of their vehicles until they schedule a repair, and verify the threshold warning system is within the NHTSA specification. The customers were informed to contact the NHTSA with questions by calling the Vehicle Safety Hotline at 1-888-327-4236 or go to http//www.safercar.gov. If they have any problem obtaining the needed repair or wish to schedule service with an approved Braun technician, contact the Braun Corporation Product Support at 1-800-946-7513. |  | 8356 | 8,356 | Worldwide distribution | Terminated | May 12, 2011 | 1366 | 12940
Z-2482-2010 | 55360 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECTÂż i1000SR" Processing Module System For In Vitro Diagnostic Use, Abbott Laboratories Abbott Park, IL |  |  | October 21, 2008 | 2008 | Abbott Laboratories, Inc | During manufacturing, two components of the ARCHITECT i1000SR were reversed which could potentially cause falsely elevated patient results due to a sample-reagent and reagent-sample carryover. | Mixed-up of materials/components | ABBOTT issued Technical Service Bulletins to Field Service Representatives who then inspected every unit . |  | 404 | 404 units distributed worldwide (including US) | Worldwide distribution:USA including states of: CO, DC, FL, GA, IL, LA, MA, MN, NY, OH, OK, TX, UT, and WV and countries including: Australia, Brazil, Bogota, Canada, China WFOTC, Columbia, Curacao (NETH ANT), Germany, Hong Kong, Japan, Mexico, Singapore, and Taiwan. | Terminated | April 13, 2012 | 1270 | 404
Z-2480-2010 | 55911 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, Catalog number 801763, 100/120V System 1 base, Terumo Cardiovascular System, Ann Arbor MI. |  |  | September 10, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The design document states that the oxygen sensor would operate for 300,000 hours but the supplier of the oxygen sensor indicated that it has a 100,000 hour warranty. The oxygen sensor is a consumable part that slowly depletes over time. At the end of its service life, a depleted oxygen sensor results in a gas system calibration failure or a message to the user prompting service for the gas syst | Nonconforming Material/Component | Terumo issued a URGENT MEDICAL DEVICE CORRECTION letter dated September 10, 2010, identifying the affected product, the issue prompting the correction, and what actions are to be taken by the customer. The firm included the addendum to the operators manual with the recall letter. The addendum recommended replacement of the oxygen sensor every six months instead of every 12 months. The addendum provided additional user instructions in the event of a calibration failure of the sensor. Customers are to complete and return the enclosed Customer Response Form to the firm. Terumo CVS Customer Service can be contacted by 1-800-521-2818 Monday through Friday, 8 AM  6 PM EST. |  | 1340 | 1340 | Worldwide Distribution: USA, including the states of AL,AZ, CA, CO, CT, DC, FL, GA, HI, IA, IN, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WI, and WV, and the countries of Canada, Indonesia, China, Chile, UAE, Japan, Singapore, UK, Malaysia, Taiwan, Belgium, Korea, Phiippines and Thailand. | Terminated | December 20, 2011 | 466 | 2655
Z-2462-2010 | 56560 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GCX Mountable Downloader-Recharger; |  |  | August 20, 2010 | 2010 | Abbott Point of Care Inc. | Reliability issues associated with broken connection pins within the downloader that mate with the pins from the analyzer. | Device Design | Abbott Point of Care, Inc. sent Urgent Recall Notices, dated August 20, 2010, by Federal Express to two Clarion Health locations.The letter identified the product, the issue, and the actions to be taken by the firm and the customers. 1) An Abbott Point of Care representative will contact the customers to plan the replacement of affected product at their facility.2) Replacement will occur in batches of approximately 50 units at a time.3) Abbott will provide personnel to perform the replacements to help expedite the replacement.4) The customer is required to acknowledge that all units have been replaced by signing the form provided with the Urgent Recall Notice.. If you have any questions regarding this information please contact Abbott Point of Care Technical Support at 800-366-8020, Option 1 or your Abbott Point of Care representative. |  | 277 | 277 | Nationwide Distribution - USA, including the state of Indiana | Terminated | July 25, 2012 | 705 | 277
Z-2478-2010 | 56676 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Fabius GS Anesthesia Machine, OR/Anesthesia, Draeger Medical, Inc., 3135 Quarry Road, Telford, PA 18969 |  |  | August 26, 2010 | 2010 | Draeger Medical Systems, Inc. | potential failures to Electri-Cord Manufacturing AC power cord | Other | Draeger Medical, Inc., sent an Urgent - Medical Device Recall letter dated August 2010 to all customers. The letter described the product, the problem, and actions to be taken by the customer. Customers were instructed to:1) Immediately take the cord out of service if the plug has bent or cracked prongs, burned plastic, or shows signs of excessive wear.2) Monitor the affected cords regularly and be mindful of excessive wear and tear, misuse, or abuse until the affected cord is replaced. 3) Customers were also asked to complete and return the enclosed Customer Power Cord Form via mail: Draeger Medical, Inc., 3135 Quarry Road, Telford, PA 18969; fax: 1-215-721-5808; or e-mail: kathy.benedetto@draeger.com.4) Disable the cords by cutting off the plug and dispose of the cord in accordance with their facility's disposal procedure when affected power cords are removed.NOTE: The firm will send replacement power cords free of charge. Draeger will only replace power cords originally provided by Draeger (identified with the white flag/part number 4117266.If you have any questions regarding this letter please call 1-800-543-5047 (press 1 at the first prompt, 2 at the second prompt and then 32349#). |  | 1973 | up to 1973 products (GS &amp; Tiro) | Worldwide Distribution: USA and countries of Brazil, Canada, Columbia, Costa Rica, Cuba, Ecuador, Japan, Lebanon, Nicaragua, Mexico, Panama, Philippines, Saudi Arabia, Taiwan, and Venezuela | Terminated | August 23, 2011 | 362 | 1973
Z-2483-2010 | 56498 | No | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sterile double breast personal accessory convenience kit for the Symphony Breast pump; the kit connects to a hospital grade breast pump for the removal, collection, and storage of mother's breast milk; Manufactured exclusively by and for Medela, Inc., P.O. Box 660, 1101 Corporate Drive, McHenry, IL 60051-0660; article # 67099S |  |  | August 06, 2010 | 2010 | Medela Inc | The product sterility may be compromised due to possible pinholes in the rigid tray of the package. | Packaging | Medela sent Important Device Recall letters dated August 6, 2010 to all direct accounts, identifying the affected product, the labeling issue, and actions to be taken by the customer. The accounts were informed that the sterility of the affected products may be compromised due to holes in the packaging. They were instructed to stop use, sale and distribution of the kits and return their stocks to Stericycle using the enclosed packing slip and prepaid UPS Return Service shipping label. The accounts were requested to complete and return the enclosed postage paid reply card, indicating the amount of product being returned. Distributors were requested to notify their customers. Questions about the recall process may be directed to Stericycle at 1-877-247-9970. Information on the recall is posted on Medela's website www.medelabreastfeedingus.com |  | 581296 | 581,296 kits | Worldwide Distribution: Nationwide throughout USA, including Guam, and to the countries of Barbados, Bermuda, British West Indies, Canada, Panama, Switzerland, Trinidad and Tobago, and Uganda. | Terminated | September 07, 2011 | 397 | 6924825
Z-2487-2010 | 56423 | [6 empty] | Ear Candles made from either Paraffin or Beeswax, in plain, herbal scented, or lavender scented varieties, manufactured by Wally's Natural Products, Auburn, CA. |  |  | February 17, 2010 | 2010 | Wally's Natural Products | Product is an unapproved medical device. Use as directed may result in accidental damage to the eardrum, leading to injuries up to loss of hearing. | Other | The firm conducted communications via phone, mail and website, stating that they have stopped manufacturing and distributing ear candles, and asks consumers to return for a refund. |  |  | Not Specified. All | Nationwide Distribution | Terminated | March 28, 2011 | 404
Z-2501-2010 | 54745 | No | LCX | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UCG Beta Slide Monoclonal II is a latex agglutination inhibition slide test for the detection of hCG in urine. This test is used for the diagnosis of pregnancy. |  |  | February 12, 2010 | 2010 | Inverness Medical Professional Diagnostics | Specific lots no longer meet their stated limit of detection claims of 2000 mIU/mL for the UCG-Slide test and 500 mIU/mL for the UCG-Beta Slide Monoclonal II test. | Other | The recall was initiated on February 12, 2010. Letters were sent by Fed Ex on 2/12/2010. |  | 585 | 585 kits | Distributed Nationwide and to Australia | Terminated | May 14, 2012 | 822 | 1637
Z-2503-2010 | 56441 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | GATOR. REF 9263A, 4.2mm STERLING, STERILE Rx Only. |  |  | August 05, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Largo, FL is recalling ConMed Linvatec 9263A,4.2mm Sterling Blade, Gator Meniscus Cutter, Lot Number 178026 for breech in the sterile barrier. | Packaging process control | Recall Notifications will be sent to all direct consignees by FedEx, email, Fax, etc. If the account no longer has the product in their possession, they will be asked to respond as such. A Return Material Authorizations Number will be assigned for the returned product as required. Product returned will be placed into quarantine and will be dispositioned in accordance with ConMed Linvatec procedures. 100% of the consignees will be contacted. The reply form is the mechanism to verify notification effectiveness to all consignees. If a reply is not received within approximately 30 business days of the first notification ConMed Linvatec will make two additional contacts by phone and / or email, as necessary, to insure all consignees have received notification and have taken appropriate action. If a reply is not received following the third contact by email or phone, no further contacts will be made to the consignee. |  | 174 | 174 | Products were distributed to ConMed Linvatec consignees, sales representatives and affiliates within the US and Australia. US distribution was to WA, AZ, PA, KS, NM, UT, ID, IL. Products were distributed to ConMed Linvatec Australia. | Terminated | December 14, 2010 | 131 | 174
Z-2499-2010 | 56406 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Liquid Mouse Monoclonal Antibody Thyroid Transcription Factor-1; 1 mL liquid tissue culture supernatant in a 1 mL vial; an in vitro diagnostic reagent for the qualitative identification by light microscopy of thyroid transcription factor-1 molecules in paraffin-embedded tissue by immunohistochemical staining; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom; Catalog No.: NCL-L-TTF-1 |  |  | August 02, 2010 | 2010 | Leica Microsystems, Inc. | In combination with a sensitive polymer detection system, the TTF-1 antibodies may stain additional tumor types other than is currently indicated on the instructions for use. | Device Design | Leica Microsystems sent Medical Device Field Correction letters dated 7/28/10 to the affected customers via first class mail on the same date, informing them that in combination with a sensitive polymer detection system, the TTF-1 antibodies may stain additional tumor types other than is currently indicated on the Instructions for Use. Currently they state that lung and thyroid tumors may stain, but Leicas investigation further characterized the expression of the TTF-1 antigen and it may also be present in other tumors, e.g. thymomas. Customers who have received the affected lot numbers of the three products were requested to refer to the new Instructions for Use to review the amendments made to the Results Expected via the following links: http://bsd.leica-microsystems.com/pdfs/products/ttf-1-ce.pdf; http://bsd.leica-microsystems.com/pdfs/products/ttf-1-l-ce.pdf; and http://bsd.leica-microsystems.com/pdfs/products/pa0364.pdf; The accounts were asked to complete the attached Acknowledgement Form, acknowledging the receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. |  | 860 | 860 units Total for all three products | Nationwide and Canada | Terminated | January 29, 2012 | 545 | 860
Z-2502-2010 | 56418 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GreenLight HPS fiber optics, Model Number 10-2090, manufactured by AMS (American Medical Systems), San Jose, CA. Potentially Mislabeled as GreenLight PV 10-2079 |  |  | January 22, 2010 | 2010 | AMS Innovative Center - San Jose | Product may be packed with incorrect external or internal labels and/or missing product insert. | Labeling mix-ups | Notification issued via letter with instructions and response sheet through certified mail on 01/12/2010. |  | 11668 | 11668 units | Product was distributed to 148 medical facilities throughout the US. and 6 consignees in Great Britain, Canada and Australia. | Terminated | January 25, 2011 | 368 | 11668
Z-2504-2010 | 56591 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Modular Microplasty Cup inserter 3/8" thread , REF 31-400600 |  |  | July 28, 2010 | 2010 | Biomet, Inc. | There was a clerical error in order entry which resulted in the recalling firm recieving spring components from a supplier made of "music wire" The recalling frim conducted tests to determine the effects of theis change in materials. The testing showed that the wire would oxidize and there would be early failure of the device. | Employee error | The firm send URGENT MEDICAL DEVICE RECALL NOTICES to its distributors on 7/28/2010. The notices instructed the distributors to discontinue use of the product and return it to Biomet. |  | 331 | 331 total for both products | US, Australia, Canada, Mexico, Chile, Japan. | Terminated | October 01, 2012 | 796 | 331
Z-2496-2010 | 56363 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Two-Lumen Central Venous Catheterization Kit with Blue FlexTip ARROWg+ard Blue Catheter |  |  | July 29, 2010 | 2010 | Arrow International Inc | Incorrectly assembled 18 Ga x 2 1/2 Introducer Needle | Mixed-up of materials/components | The recalling firm issued an Urgent Medical Device Recall letter dated 7/23/2010 along with a a recall acknowledgement and stock status form. |  | 5 | 5 | The product was shipped to medical centers and hospitals nationwide and to Australia, New Zealand, and Europe. | Terminated | July 05, 2012 | 707 | 113
Z-2582-2010 | 49599 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Product is labeled in part: "CP-SM397A-02...Mono-Dox Violet Monofilament Polydioxanone Synthetic Absorbable Sterile Suture. Do Not Resterilize.... 4/0 (1.5 metric) 30" (75cm) Rx Only...LOT L0535440 EXP 2009-11...4/0 M397A...Mono-Dox CUTTING NFS-2 NEEDLE...Manufactured by: CP Medical, Portland, OR USA". |  |  | September 17, 2008 | 2008 | C P Medical | Seal integrity of inner foil pouch of the Mono-Dox Synthetic Absorbable Suture may be compromised. | Vendor change control | On 09/17/2008, the firm's Sale Representative began using the Telephone Script and calling all their customers.On 09/19/2008, the firm began sending the Recall Notice, dated 9/17/2008, letter to their customers. The letter states that the recalled product is Polydioxanone Suture, Code M397, Lot Number L0535440, Expiration Date: November 2009.The customers are advised to discontinue selling this product and return any existing stock to CP Medical. The firm also requests their consignees to notify their sub-accounts about this recall notice if the product is further distributed.The customers are instructed to provide return all accounted quantity and provide quantify information and location of any product that they are unable to recover for accountability purposes. Customers could contact CP Medical Customer Service Representative at 1-800-950-2763 for return goods authorization number and shipping account number. |  | 20208 | 1684/12 pcs. Boxes (1628/12 pcs. Boxes to veterinary market, 56/12 pcs. Boxes to human suture market) | Worldwide Distribution -- US, Australia, and Taiwan. | Terminated | September 28, 2010 | 741 | 20208
Z-2580-2010 | 56046 | No | OCL | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | Cardima Surgical Ablation Probe (SAP) Electrosurgical cutting and coagulation device, Model Number 01-242208, manufactured by Cardima Inc, Fremont, CA |  |  | May 20, 2010 | 2010 | Cardima Inc | Sterility compromised-- The sterile pouch in which the product is packaged may be compromised, which could result in patient infection. | Device Design | Cardima issued a Product Recall Notification dated May 18, 2010 to customers, identifying the affected product, the sterility issue, and the actions to be taken by the customer. Custormers are to quarantine and return the affected product to the firm. The firm will replace returned product as soon as it becomes available.Cardima Customer Service can be contacted at 800 354-0102. |  | 28 | 28 units | Worldwide Distribution to Distributors in USA (Pittsburgh, PA) and the countries of Ireland and Belgium, and customers in the states of CA, OH, MI, MO, AND WA, and the countries of Ireland and the Netherlands. | Terminated | March 23, 2011 | 307 | 28
Z-2584-2010 | 56379 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | WECK IMA/ENT Blade Electrode Âż (1.9 cm), REF 809335, Rx Only Sterile, Teleflex Medical, RTP, NC |  |  | July 30, 2010 | 2010 | Telefelx Medical | Complaints were received regarding peeling and melting of the insulation of the cautery tip of the IMA/ENT blade electrodes. | Other | Teleflex Medical sent an "Urgent Medical Device Recall" letter dated July 30, 2010, to all customers. The letter described the product, problem and actions to be taken by customers. The customers were instructed to immediately discontinue use, quarantine any affected products and return all of the affected product to Teleflex Medical; forward the letter to their customers to retrieve product from those locations and complete and return the Recall Acknowledgement Form via fax at 1-866-804-9881. A second mailing to non-responding customers on the consignee list will be mailed approximately four to six seeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the second mailing or as an alternative, they may be contacted by phone, e-mail, or fax. Note: A record of these contacts will be maintained. Affected products received by Teleflex Medical will be destroyed upon receipt. Records will be maintained accordingly.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 190344 | 190,344 eaches | Worldwide distribution: USA and countries of Canada, Australia and Singapore. | Terminated | April 21, 2011 | 265 | 190344
Z-2579-2010 | 56199 | Yes | NIK |  |  | Cardiovascular | PMA | Sorin Paradym CRT-D Model 8750 and Sorin Group Paradym sonR CRT-D Model 8770. |  |  | June 07, 2010 | 2010 | ela Medical Llc | The Paradym CRT 8750 and CRT sonR 8770 device models have a software anomaly, which will cause the device to lose the ability to sense/pace and to deliver further therapy. The software anomaly could only occurs under a specific sequence of events. | Software design | Sorin Group issued a "Medical Device Correction" letter dated June 7, 2010 to consignees. The letter described the product, the problem, and actions to be taken by customers. A Sorin Group representative will contact customers to update the software. Prior to receiving the software update, customers should program the exercise AV delay to 40 ms, which will trigger a reset in the event of an anomaly. This will restore normal operation after shock delivery. If the anolmaly occurs on an implanted unit, a RV coil or SVC coil continuity test should be started, which will ensure that the device recovers proper operation.Corin Group can be contacted at 1-615-788-1807. |  | 215 | 215 | Nationwide Distribution in the states of AZ, CA, FL, IA, LA, MA, MS,MO, NJ, NY, NC, OH, OK, PA, SC, TN, TX, and VA. | Terminated | March 29, 2012 | 661 | 215
Z-2585-2010 | 54561 | No | FZW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CONCEPT GRAFIX TENDON STRIPPER. REF TS8850. 7.0MM, QTY 1. Rx Only, Non-Sterile, Made In USA. |  |  | January 29, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | There is a possibility that the device may exhibit Hydrogen Embrittlement in the material and are subject to potential breakage. | Process control | Recall Notification Letters were sent to customers on January 29, 2010. Customers were requested to complete and return a Customer Response Form indicating the number of products available for return and number distributed. Each distributor was asked to notifiy their customers if the product was further distributed. |  | 149 | 149 | Distributed throughout U.S. and Internationally to Arab Emirates, Belgium, Canada, China, Columbia, Cyprus, Germany, Denmark, Egypt, Spain, France, United Kingdom, Greece, Hong Kong, India, Italy, Japan, Republic of Korea, Lithuania, Morocco, Malaysia, Poland, Portugal, Saudi Arabia, El Salvador, Thailand, Taiwan and Viet Nam. Domestic Distribution to CA, FL, GA, IL, IN, KS, MI, MT, NY, OH, PA, TX, and WA. | Terminated | January 12, 2011 | 348 | 149
Z-2588-2010 | 55708 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Integra Suction Reservoir Kit for use with Integra NeuroSciences Subdural Drainage Catheter; |  |  | May 07, 2010 | 2010 | Integra LifeSciences Corp. | The sterility of the product may be compromised. Integra LifeSciences Corp. received a recall notification from Saint-Gobain Performance Plastics, France, contract manufacturer of the Integra Suction Reservoir with Anti-Reflux Valve, that the product's sterility may be affected by a sterilization issue discovered by the supplier. | Other | Integra sent Urgent Product Recall letters dated May 7, 2010 to customers identifying the affected product, the sterility issue, and actions to be taken by the customer. Customers are instructed to quarantine and return all affected inventory to Integra's distribution center in Sparks, NV, using a Returned Material Authorization Number received from Integra LIfeSciences Customer Service at 800 654-2873. All Customers are requested to return the attached "Recall Acknowledgement. and Return Form" Integra LIfeSciences can be contacted at 609-936-2495. |  | 313 | 313 units | Worldwide Distribution: USA, including the states of AK, AL, AR, AZ, CA, CO, FL, GA, IL, MD, MI, MO, NC, NH, NJ, NY, OH, PA, RI, SC, TX, UT, VA, WI, and the countries of Australia, Chile, Costa Rica, and Sri Lanka. | Terminated | August 27, 2012 | 843 | 313
Z-2609-2010 | 56160 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Engenex Non Acute Care Suction Pump for use with Engenex Negative Pressure Wound Therapy |  |  | July 09, 2010 | 2010 | Convatec Inc. | Reduced vacuum capacity has been reported with a small number of units caused by electrostatic discharge (ESD) or static electricity. | Device Design | Recall notification letters were sent to 46 customers, acute care facilities and durable medical device distributors on July 9, 2010, by Federal Express. |  | 1190 | 1190 units | Nationwide distribution. | Terminated | June 07, 2011 | 333 | 1190
Z-2586-2010 | 56507 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Epic II Critical Care Bed, Model 2030/2031, Stryker Medical, Kalamazoo, MI. |  |  | May 14, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The Fowler mechanism may malfunction, preventing it from being lowered into a flat position. | Component change control | Stryker Medical sent an Urgent Medical Device Alert letter dated May 14, 2010 to their customers. The letter identified the product, the problem, and the actions customers should take. The letter instructed users to do the following:" Return the postcard included with the letter to confirm receipt of the notification." Locate the recalled beds, verify the Fowler mechanism is functioning properly and does not display uneven motion when raising and lowering, and is not unusually noisy. " If the Fowler mechanism is not operating correctly, they were to call Stryker at 1-800-STRYKER, option 1, Reference RA-2010-014 to obtain service for the bed. " If any beds have been loaned or sold, the customer is to forward a copy of the notice to the new users and advise the firm of the new location." Provide the firm with the serial numbers for any beds that were disposed of." Retain the revised maintenance manual included with the notification.If the customer had any urgent questions or concerns, they were instructed to call the firm at 800-STRYKER, option 1, reference RA-2010-014, Monday-Friday 8am to 5pm (EST). |  | 9134 | 9,134 | Worldwide Distribution - USA (all states), and including the countries of ARGENTINA, BRAZIL, CHILE, CHINA, CANADA, MEXICO, GREECE, JAPAN, KOREA, PUERTO RICO, SPAIN, POLAND, TAIWAN, SINGAPORE, INDIA, SWITZERLAND, UK and MALAYSIA. | Terminated | June 13, 2011 | 395 | 14060
Z-2608-2010 | 56709 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Langston Dual Lumen Catheters, 6F, REF 5540, Sterile EO, RX only, Vascular Solutions, Inc, 6464 Sycamore Court, Minneapolis, MN 55369, 110 cm length, |  |  | September 02, 2010 | 2010 | Vascular Solutions, Inc. | Investigation of a recent Device Experience Report has made us aware of a potential problem with our 6F Langston dual lumen pigtail catheter (Model 5540) with the following lot numbers: 548110, 548227, 548364, 548430, 548554 &amp; 548723.Recently, it was reported that pouches containing Langston Pigtail catheter (Model 5540) were missing a seal, posing a risk of product contamination. It is possib | Process control | Consignees were sent on 9/2/10, a Vascular Solutions "Urgent Medical Device Field Action" letter dated September 1, 2010. The letter was addressed to Cath Lab Manager. The letter descried the product and the problem. Provided instructions to examine pouches of opened boxes to see if they have not been sealed. If unsealed, they recommended to remove pouches from the current inventory. For unopened boxes they recommended to remove the boxes from the inventory. Requested consignees to complete the Field Action Customer Inventory Form included with the letter. For any questions, please contact the firm at 763-656-6032. |  | 1128 | 1,128 boxes ( 5pk) | AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS,MO, NE, NY,NH, NJ, NM, NY,NC, OH, OK, OR, A, PR, SC, TN, TX, UT, VA, WA, WV, WI, and WY.SINGAPORE, MEXICO, CANADA, SPAIN, AUSTRIA, GERMANY, ITALY, ENGLAND, and SWITZERLAND. | Terminated | April 24, 2012 | 600 | 1128
Z-2589-2010 | 56543 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Apligraf is supplied as a living, bi-layered skin substitute. |  |  | August 02, 2010 | 2010 | Organogenesis, Inc. | Product is contaminated with a yeast identified as Pseudozyma antarctica. | Process control | Organogenesis notified accounts by telephone, e-mail or fax on 8/2/010 to remove from use and return recalled units. If unit was applied to patient then monitor for adverse events. The Notification Letter dated August 6, 2010, described the issue and provided instructions for customers. Customers are to place the affected unit on hold for its return to the firm. The firm will contact customers to make arrangements for its return. Replacements have already been sent. If the affected unit has already been applied, the patients should be monitored closely for any potential adverse events. An Organogenesis Medical Affairs rep will be in contact regarding patient follow up. Customers are to complete the attached form and fax it back to the number provided. Questions should be directed to the Medical Information Center at 1-888-432-5232. |  | 32 | 32 units | Nationwide Distribution -- AL, AZ, CA, CO, IL NC, PA, TN, and VA. | Terminated | March 10, 2011 | 220 | 32
Z-2598-2010 | 56522 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 8 MM DISPOSABLE CORING TREPHINE WITH PROTECTION SLEEVE, REF 900728, Sterile, BIOMET Sports Medicine, Warsaw, IN 46581. |  |  | July 01, 2010 | 2010 | Biomet, Inc. | The affected instruments are discolored and/or have a sticky residue on the end. | Packaging change control | The firm sent out "Urgent Medical Device Recall" notices dated 7/1/2010. Letters were sent to distributors, customers and OR managers. The letters stated that the devices should be located, quarantined and shipped back to the recalling firm. On 7/15/2010, the firm added one lot of one additional product to the recall (product 900788 lot 780470). The firm sent letters to physicians that had used that lot of product.Questions relating to this recall should be directed to 574-371-3755. |  | 100 | 100 | US, Canada, Australia, Chile, Argentina, Netherlands, Costa Rica and Japan. | Terminated | August 20, 2012 | 781 | 601
Z-2591-2010 | 56695 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump2 250 mL PainPump with Luer Lock Tubing Set, sterile, REF 525-350, Stryker Instruments, Kalamazoo, MI. |  |  | September 02, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | All of the failures lead to the patient receiving less medication than intended. Certain lots of the pain pump have a high failure rate associated with a cracked port housing which leads to internal or external leaks. Some lots have a high failure rate associated with E2 errors. | Nonconforming Material/Component | The firm sent our Urgent Medical Device Recall Notices dated 7/12/2010 to Risk Managers and Materials Managers. The recall notice listed specific lots of PainPump2 and PainPump2 BlockAid that were under recall. Consignees were told to check stock for the recalled product, quarantine the product, and contact Stryker to arrange return shipment of the recalled product.On 9/2/2010 the firm sent out a second Urgent Medical Device Recall Notice. This notice expanded the scope of the recall to include all products distributed within the past 12 months. The shelf life of the product is 1 year -so the recall was for all product currently in the supply chain. The letter again directed consignees to locate the recalled products and contact Stryker to arrange return shipment of the recalled product. The firm sent an additional letter dated 9/1/2010 to its customers. This letter stated that Stryker will no longer be offering any of their PainPump2 products. |  | 54568 | 54,568 all products | Worldwide Distribution -- US, Canada, and Australia. | Terminated | April 21, 2011 | 231 | 54568
Z-2607-2010 | 56485 | No | CFL | Human growth hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IGFBP-3 Immunoradiometric Assay (IRMA) Kit, Part |  |  | August 09, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed customer reports that IGFBP-3 measurementsusing IGFBP-3 IRMA kit (REF DSL-6600) lots identified above may result in IGFBP-3 levels that are not in agreement with simultaneously measured levels ofIGF-1. The investigation confirmed that above mentioned lots of IGFBP-3 kit provide results elevated by 84% on average. Erroneously elevated IGF | Pending | The recall communication was initiated with a Product Corrective Action (PCA) letter that was sent on August 09, 2010 to allcustomers who received the IGFBP-3 IRMA kits, lot numbers referenced. The letter will provide the customers with an explanation of the problem identified and an action to be taken.Customers were instructed to:(1) Discontinue using all lots identified above and discard in accordance withlocal regulations.(2) To continue testing for IGFBP-3 levels, find an alternate method.(3) Immediately notify all parties of interest (e.g., laboratory staff, clinicians) ofthe possibility of erroneously elevated IGFBP-3 results for patient samples tested with these lots, and advise all parties that clinical decision making and patient safety may be impacted.(4) Review patient IGFBP-3 results for the lots listed above, and interpret them in the light of additional tests (e.g., IGF-1) and the patient's clinical status and history.Customers were also asked to share the recall information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. Complete and return the enclosed response form within 10 days. For any questions regarding this notification, customers were to contact the DSLTechnical Support Center at 1-800-231-7970 option 4 in the United States, or contact your local Beckman Coulter Representative. |  | 63 | 63 | Nationwide in US and Canada | Terminated | August 13, 2012 | 735 | 63
Z-2596-2010 | 56366 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MODEL P8400 HANDLE WITH INTEGRATED CABLE, FOR USE WITH MODEL P8400 CONTROLLER, REF 7013-8050, QTY (1), STERILE EO, HEMOSTATIX MEDICAL TECHNOLOGIES, LLC, BARTLETT, TN 38133 USA, CE 0473, EU REP : QFI LTD, STRATFORD, LONDON E15 2ST UK |  |  | July 27, 2010 | 2010 | Hemostatix Medical Technologies LLC | Lack of sterility assurance | Other | On July 27, 2010, the firm, Hemostatix Medical Technologies, LLC, notified their international distributor in Japan of the recall by email. Hemostatix sent an "URGENT: MEDICAL DEVICE RECALL" letter dated July 29, 2010, to U.S. customers delivered via FedEx Priority Overnight. All non-responding customers were to be notified by telephone after 2 weeks. The other international customer, Italy, were to be notified by email on July 30, 2010. The letter describes the product, problem and action to be taken by customers. The customers were instructed to immediately check their inventory and remove any affected product found; ensure that satellite facilities/clinics receive the recall notice; complete and return the Recall Response Form via fax to Hemostatix's international fax number at 901-261-0017, and contact the Customer Care Department's recall hotline at 901-261-0012 for further instructions, if they do have affected stock.If you have any questions or concerns, please feel free to contact our Customer Care Department at: 1-800-722-5735 in the USA or 1-901-261-0012 outside the USA. |  | 40 | 40 units | Worldwide distribution: USA including CA, GA, LA, MA, MD, MO, NC, OH, OR, PA, SC, TX, VA, and WI; and countries of Japan and Italy. | Terminated | November 29, 2010 | 125 | 259
Z-2606-2010 | 56322 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Centrifugal Pump with X-Coating, packaged within custom ordered Cardiovascular Procedure Kits, P/N 3ZZ164275X, 3ZZ164275, CEOT028, CEOT032; and as single sterile units with product codes CV-164275, 164275 and 164275X. |  |  | March 15, 2010 | 2010 | Terumo Cardiovascular Corporation | Medical device component may malfunction and require user intervention during surgical bypass procedures. | Other | The recalling firm notified end users, perfusion groups, and distributors by phone, fax, email and mail on 03/15/10 to advise of the required recall actions with instructions for labeling and returning the involved product from the customer level. The firm expanded their recall to include additional lots on 05/17/10. If you have any questions, please contact Customer Service at 800-521-2818. |  | 28257 | 28257 pumps | The product was sold to medical device distributors and to direct accounts nationwide and to a subsidiary in Japan. | Terminated | November 04, 2011 | 599 | 28257
Z-2624-2010 | 50095 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GMRS Extension Piece, 60 mm; Sterile |  |  | May 05, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | Mislabeled package: Package labeled as GMRS Extension Piece 60 mm, lot ZYCOC1 actually contains a GMRS Extension Piece 70 mm extension piece. | Labeling mix-ups | Stryker issued Important Market WIthdrawal letters dated May 5, 2005 to customers and distributors. The letters identified the affected product, the labeling issue, and actions to be taken by the customer or distributor. Customers were instructed to identify and quarantine affected product, which will be reconciled and returned by the Stryker sales representative. Distributors were instructed to identify affected product, reconcile all product on an attached Product Accountability Form to be faxed to Stryker at 201 831-4825 within two days of receipt of the letter. The affected product was to be returned using the enclosed pre-paid mailing label. Contact the Stryker sales representative to re-order product that is being returned.Stryker can be contacted at 201 831-5825. |  | 5 | 5 units | Nationwide Distribution: USA, including the states of AZ, FL, MD, and PA.G | Terminated | October 01, 2010 | 1975 | 5
Z-2621-2010 | 55188 | No | CGR | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems CORTISOL, Part Number: 33600 |  |  | February 17, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated by Beckman Coulter after the firm confirmed customer reports that the affected lot of Access Cortisol reagent (REF 33600) identified above may produce suppressed patient results. Beckman Coulter also confirmed customer reports that the two lots of Access Progesterone reagent (REF 33550) identified above may produce suppressed patient results and an increased incidence of I | Other | Two (2) Product Corrective Action (PCA) letters with attached Customer Response form were sent on the week of February 15, 2010 to the affected customers. The letters provide the customers with an explanation of the problem identified and instruct the customers to discontinue use of Access Progesterone and Access Cortisol Reagents, lot numbers referenced. For product replacements consignees were instructed :(1) In the United States, please contact Customer Service at 1-800-526-3821; (2) In Canada, please contact Customer Service at 1-800-463-7828;(3) Outside of the United States and Canada, contact your local Beckman Coulter Representative.If consignees need assistance or have any questions regarding this notification, they were instructed to contact TechnicalSupport at 1-800-854-3633 in the United States and Canada. Outside the United States and Canada, contact your local Beckman Coulter Representative. |  | 7282 | 7282 units | United States &amp; Canada | Terminated | April 19, 2012 | 792 | 13694
Z-2618-2010 | 54617 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Jazzy Select with XLR Port Harness |  |  | February 01, 2010 | 2010 | Pride Mobility Products Corp | Improper set-up of the wire harness can allow wires to be pulled from their intended position and become caught under the motor mounting bracket located under the powered wheelchair. This can cause the wire connector and/or the wire itself to be pulled or pinched which can lead to a short or open circuit. | Other | The recalling firm issued recall notifications with response cards to their dealers. Included with the notification was fasteners and an example letter for customer notification of the problem. Dealers were instructed to contact the end users to make arrangements with the customer/end user to verify that the devices were properly set up. If the dealers' inspection revealed that additional fasteners were required they dealer would install them. The recalling firm also issued a recall letter to consumers. The consumer recall letter included instructions that if the correction was completed by anyone other than the wheelchair provider, then the consumer should schedule an appointment with the wheelchair provider to ensure the correction was completed properly. |  |  |  | The product was shipped to dealers nationwide who further shipped the product to end users. The product was also shipped to distributors in Canada, Australia and United Kingdom | Terminated | November 20, 2014 | 1753
Z-2622-2010 | 56110 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Stryker Scope Warmer, disposable. Endoscopic accessory composed of polyvinyl Chloride (PVC) and Sodium Acetate, Model Number 502-360-000, manufactured by Stryker Endoscopy, San Jose, CA. |  |  | June 21, 2010 | 2010 | Stryker Endoscopy | Outer packaging seals have inconsistencies that have resulted in compromised seals and may result in compromised sterility. | Process control | Stryker issued an Urgent Device Removal letter, dated June 21, 2010, to customers which identified the affected product, the sterility issue, and the actions to be taken by the customer. Customers were instructed to identify and quarantine any affected product in their inventory, and contact Stryker Customer Service at 800 624-4422 and request an RMA#. The RMA# is to be written on the outside of the package used to return products, using an enclosed pre-paid shipping label. All cusotmers are to complete and sign an enclosed "acknowledgement of receipt" and fax to (408) 754-8278.Stryker can be contacted at 800-624-4422. |  | 191855 | 191855 units | Worldwide Distribution: USA, Including the states of AK, AL, AR, CA, CO,CT, DE, DC, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MO, MT, NE, NJ, NM, NH, NV, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY, and the countries of Australia, Brazil, Canada, Chile, China,France, Germany, Iberia, India, Italy, Netherlands, New Zealand, Japan, Korea, Latin America, Polska, Singapore, South Africa, Switzerland, Thailand, and United Kingdom. | Terminated | December 16, 2010 | 178 | 191855
Z-2533-2010 | 56548 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | All Med Medical 22ga x 1.5" Standard Bore Winged Administration Set |  |  | August 30, 2010 | 2010 | Multi-Med, Inc. | Huber needles may core and result in port leakage or emboli being flushed our of the port. | Device Design | Multi-Med issued an Urgent: Medical Device Recall notification on 8/30/10 via fax and followed by hard copy letter. Users were advised of the problem and asked to immediately examine their inventory and quarantine the affected product. They were to also notify their customers of the recall. The firm stated in the letter that they would contact customers to determine the number of product that they had and the amount they expected from the field. The firm will coordinate a plan for the return of the affected product. Questions should be directed to Sue Starkey at 603-357-8733. |  | 672 | 672 units | Worldwide Distribution -- USA and Mexico. | Terminated | January 31, 2013 | 885 | 546873
Z-2623-2010 | 56261 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Strattice Reconstructive Tissue Matrix for Stoma Reinforcement. |  |  | July 12, 2010 | 2010 | LifeCell Corporation | Mislabeled -One lot of product labeled as 8 cm x 8 cm were actually cut using a 6 cm x 6 cm template. | Employee error | LifeCell issued Recall Notice letters dated July 12, 2010 to customers, identifying the affected product, the labeling issue, and actions to be taken by the customer. Two versions of the recall notification were sent, one for implanted units and one for those not yet implanted. Both letters stated that the mislabeling has no impact of the safety or effectiveness of the medical device. Replacement units were offered for all affected units. Those with unused units were to contact LifeCell Customer Solutions for return and replacement of the unit at 1-866-423-2433.Those with implanted units may contact LifeCell Customer Solutions at 1-866-423-2433 with any additional questions or concerns. |  | 32 | 32 units | Nationwide Distribution: USA, including the states of AZ, DC, FL, GA, ID, IL, KY, LA, MO, NC, OH, OK, OR, TX, WA, and WI. | Terminated | October 01, 2010 | 81 | 32
Z-2626-2010 | 56646 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips M8102A IntelliVue MP2 |  |  | August 31, 2010 | 2010 | Philips Healthcare Inc. | Speakers on the Philips IntelliVue X2 and MP2 patient monitors may fail, causing absence of an audible alarm and delaying patient treatment. | Process design | Philips issued a URGENT - MEDICAL DEVICE letter dated August 26, 2010, to each affected customer. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue and gives instructions on actions to take while awaiting their replacement speaker assemblies. The speaker replacement assembly is expected to be available in 8 weeks.Affected products may continue to e used while awaiting replacement of the speaker, provided that the user routinely checks for the display of the "Speaker Malfunct." INOP at power-up.If you experience this INOP or there is no sound from your X2/MP2 speaker while it is not connected to another audible alarm annunciation system, remove the device from use and contact your local Philips service representative at 1-800-722-9377. |  | 2832 | 2,832 UNITS | Worldwide Distribution - USA and the countries of : ARGENTINA AUSTRALIA AUSTRIA BAHRAIN BANGLADESH BELGIUM BRAZIL CHINA CZECH REPUBLIC DENMARK EGYPT ESTONIA FINLAND FRANCE GERMANY GREECE HONG KONG HUNGARY ICELAND INDIA INDONESIA IRAQ IRELAND ISRAEL ITALY JAPAN JORDAN KENYA KUWAIT LATVIA LIBYAN ARAB JAMAHIRI LITHUANIA MALAYSIA MAURITIUS MEXICO MOROCCO NAMIBIA NETHERLANDS NEW ZEALAND NORWAY OMAN PAKISTAN PALESTINE PARAGUAY POLAND PORTUGAL QATAR ROMANIA RUSSIA SAUDI ARABIA SINGAPORE SLOVAKIA SLOVENIA SOUTH AFRICA SOUTH KOREA SPAIN SRI LANKA SWEDEN SWITZERLAND SYRIAN ARAB REPUBLIC TAIWAN THAILAND TUNISIA TURKEY TURKMENISTAN UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN and VIET NAM | Completed |  |  | 22232
Z-2627-2010 | 56505 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Gemstar Pump Set, Latex-Free Extension Set with Integral Pressure-Activated Anti-Siphon Valve and Y-Extension with Backcheck Valve-SL, Nonvented, 110 inch; Rx, sterile, 24 individually packaged sets per case; Hospira, Inc., Lake Forest, IL 60045 USA, Made in Costa Rica; list number 12640-01 |  |  | August 12, 2010 | 2010 | Hospira Inc. | Underdelivery of infusion during clinical use at low rate settings (less than 10 mL/hour). Patient may not receive the intended amount of medication. | Nonconforming Material/Component | Hospira issued Urgent Device Recall letters dated August 12, 2010 to Hospira's direct accounts, informing them of the potential for under delivery under low rate settings using the affected lots of pump sets. The accounts were instructed to examine their inventory for the affected product, quarantine it immediately, and complete the attached reply form indicating the amount of product being returned, faxing it to Stericycle at 1-877-523-9110. If they distributed the affected product further, they were instructed to sub-recall the product from their sub-accounts. Accounts may contact Hospira Customer Care at 1-877-946-7747 or their Hospira Account Manager to obtain replacement product. For medical inquiries, the accounts were directed to contact Hospira Medical Communications at 1-800-615-0187. Customers may contact Stericycle, which is managing this recall, at 877-274-7163 for further information concerning this recall. |  | 22680 | 22,680 sets | Worldwide Distribution: USA, (all states), and the countries of Canada, Costa Rica and Taiwan. | Terminated | June 21, 2012 | 679 | 340728
Z-2617-2010 | 56413 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Cardiovascular Procedure Kit |  |  | June 08, 2010 | 2010 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid. | Equipment maintenance | Consignees were contacted on June 8, 2010, by phone, fax, email and priority mail to advise distributors and end-users of the potential occluded vent port defect and provide instructions for the replacement product recall. Consignees were mailed follow-up communications by priority overnight mail on June 9, 2010. |  | 60 | 60 kits | The products were distributed nationwide. | Terminated | February 29, 2012 | 631 | 3979
Z-2590-2010 | 55173 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC 880i, 860i and 600i Systems; Software Part Number: A84500, A86646; Software |  |  | February 25, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall on their UniCel DxC Clinical Systems (880i, 860i,680i and 660i Systems with serial number &lt;2760 that have software version 4.9.01 installed and have the VME Motorola ICS board) because they may exhibit issues with stat sample rack loading, causing a possible delay in results. | Other | Beckman Coulter sent an URGENT: FIELD SAFETY NOTICE letter, dated March 25, 2010, to all customers. The letter identified the affected product, the problem, and the actions customers take.Customers were instructed to:if the UniCel DxC system serial number was less than 2760 the customer should contact their local representative to arrange for an upgrade to software version 4.9 build 02.Complete and return the enclosed response form within 10 days of the notice.Share the information with their laboratory staff and retain the notification as part of the Quality System documentation. If affected products were forwarded to another laboratory, provide a copy of the notification letter to them.Customers may contact the Customer Support Center (Hotline) at 1-800-854-3633. For questions or comments please contact the firm at (714) 961-3634 |  | 22 | 22 | Class II Recall - USA, including the states of CA, IN, GA, NC, NY, OH, SC, and WI | Terminated | July 02, 2012 | 858 | 22
Z-2653-2010 | 56247 | No | DEW | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | SYNCHRON System(s) Immunoglobulin A (Ig-A) Reagent, Part Number: 467920 |  |  | June 28, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter Inc. confirmed reports of high recovery of immunoglobulin A (IgA) in proficiency surveys or patient samples using Synchron Ig-A reagent lots M812164 or M902345. Using these reagent lots, some samples deficient in IgA have yielded results in the normal range.Other samples with normal levels of IgA have yielded falsely high results. | Other | A Product Corrective Action (PCA) letter was sent on the week of June 28, 2010 to the affected customers. The letter provided the customers with an explanation of the problem identified and advised customers to discontinue use of the affected lots of the SYNCHRON System(s) Immunoglobulin A (Ig-A) Reagent.Customers were instructed to:(1) Discontinue use of Ig-A reagent lots M812164 and M902345. Contact your Customer Service Representative at (800) 526-3821 in the United States or (800) 463-7828 in Canada or contact your local Beckman Coulter Representative to discuss reagent replacement options.(2) You may need to re-evaluate samples for IgA that were tested with either of the listed lots of Ig-A reagent.(3) Please complete and return the enclosed response form within 10 days so that we may maintain our records.Customers were also instructed to share this information with your laboratory staff and retain this notification as part of yourQuality System documentation. If you have forwarded the affected products listed above to another laboratory, please provide a copy of this letter to them. If they have any questions regarding this Product Corrective Action letter, customers were told to contact our Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/Customer Support/Support Requests, or call (800) 854-3633 in theUnited States and Canada. |  | 852 | 852 units | US Nationwide &amp; Canada | Terminated | July 31, 2012 | 764 | 852
Z-2634-2010 | 55680 | No | LJX | Leukocyte peroxidase test. | Hematology | Hematology | 510(K) Exempt | Hypoguard, DiaScreen 4pH, Reagent Strips for Urinalysis, 100 strips, Cat. No. D11400. |  |  | March 24, 2010 | 2010 | ARKRAY USA INC. | This recall has been initiated due to a labeling issue with the Hypoguard DiaScreen 4-way pH 100 strips. The front and back panel of the bottle carton lists five reagents when it should only list four reagents. The Ketone reagent should not be included on the label. The test strips contained in the bottle and all other labeling are correct and properly align with 4pH part number D11400 listed on | Labeling False and Misleading | Distributors were sent a Arkray "URGENT: Voluntary Medical Device RECALL " letter dated March 24, 2010. The letter described the problem and the product being recalled. The letter requested the distributors to immediately examine their inventory , quarantine and return the product subject to the recall. If the distributors had further distributed the recalled product, then they were requested to notify the customers of the recall. Arkray enclosed a customer recall letter that the distributors were to use for notifying their customers.Distributors were also contacted by phone on March 24, 2010, and were informed of the recall issue, the product involved, instructed to quarantine the product and to notify their customers. |  | 5504 | 5504 | FL, NJ, MI, NV, FL, OH, GA, WI, NY, KY, LA, CA, CT, TN, ILISRAEL | Terminated | March 08, 2012 | 715 | 5504
Z-2636-2010 | 56220 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Glenoid Head Inserter (RSP Inserter/Impactor), Device Part #804-03-041, All Lots. |  |  | July 12, 2010 | 2010 | Encore Medical, Lp | Complaints indicate product's distal threaded tip may fracture under certain circumstances. | Device Design | DJO Surgical sent an "Urgent Field Safety Notice" Letter dated July 12, 2010, to all of their consignees. The letter described the product, the problem, and the action consignees should take. Consignees were instructed to: Contact their Customer Service Representative for an R-RPR (RMS) number and return all RSP Glenoid Head Inserters, 804-03-041, under the assigned number to receive credit. Complete and return the Field Safety Notice Response to the firm. Identify and inform any customers who were sold or received the product of the Field Safety Notice Any questions consignees should call the firm at (512) 834-6330. |  | 104 | 104 | Worldwide Distribution - USA, including the states of AR, AZ, CA, FL, GA, ID, IL, IN, LA, MA, ME, MI, MO, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, WA, and WI; and the countries of Puerto Rico, Germany, and Saudi Arabia. | Terminated | June 30, 2011 | 353 | 104
Z-2651-2010 | 56304 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GYRUS ACMI, dissector PlasmaKnife" , REF 7035-3005, Rx only, STERILE R, CE 0344, 161416-CA, MANUFACTURED FOR; GYRUS ACMI, INC., 136 TURNPIKE ROAD, SOUTHBOROUGH, MA 01722-2104, USA, EC REP, GYRUS MEDICAL LIMITED, FORTRAN ROAD, ST. MELLONS, CARDIFF, CF3 0LT, UK |  |  | August 06, 2010 | 2010 | Gyrus Acmi, Incorporated | Lack of sterility assurance due to compromised package seals. | Packaging | The firm initiated their recall to their direct accounts by Priority Mail with delivery confirmation on 08/06/2010. The letter requested return of the product and was extended to the retail level. |  | 4244 | 4244 units | WorldwideIn the United States: AL, AK, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, ME, MI, MN, MO, MS, NC, NE, NJ, NV, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV and the District of Columbia. Foreign: UK, TAIWAN, SOUTH KOREA, GERMANY, CANADA, AUSTRALIA, SPAIN, ITALY, ISRAEL, NORWAY, LITHUANIA, SWEDEN, SWITZERLAND, SAUDI ARABIA, TURKEY, and GREAT BRITAIN | Terminated | January 26, 2012 | 538 | 4244
Z-2638-2010 | 56463 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra CUSA Ultrasonics NXT Ultra Surgical Aspirator System |  |  | July 29, 2010 | 2010 | Integra LifeSciences Corp. | A ground fault malfunction at the rectifier retaining screw may cause the aspiration pump in the CUSA NXT Service module to run at maximum speed, and continue to run until the entire system is turned off. This may result in excess aspiration that requires medical intervention. | Device Design | Integra issued Urgent: Product Recall Notification letters dated July 29, 2010 to all customers on July 29, 2010 identifying the affected products, issue prompting the recall, and actions to be taken by the customer. Letters were also forwarded to all Integra Account Reps.Customers were advised to discontinue use of and quarantine the affected service modules. The Surgical Aspirator may be operated without the Service Module, utilizing the hospitals wall aspiration. Company representatives will contact customers to arrange product return and repair of the affected service modules. Customers are to complete and return an enclosed Recall Acknowledgement and Return Form . The firm can be contacted at 609 936-2485. |  | 208 | 26 units in US, 8 in Europe, 1 in Asia (both types of units) | Worldwide Distribution: USA , including the states of AZ, CA, CO, FL, GA, IA, KY, MN, MO, NV, NY, OH, OR, RI, TX, VA, WI, and the countries of Germany, Hungary, Indonesia, Italy, Netherlands, and Switzerland. | Terminated | May 15, 2012 | 656 | 234
Z-2641-2010 | 56188 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS Panels Test Strips, CHOL +HDL+GLU Panel for CardioChek PA brand analyzers, 15 tests 1 MEMo Chip, manufacturer Polymer Technology Systems, Inc., 7736 Zionsville Rd., Indianapolis, IN 46268. |  |  | July 06, 2010 | 2010 | Polymer Technology Systems, Inc. | The test strips appear to under-recover Cholesterol, HDL and glucose. | Other | Distributors and direct customers were notified by letter on July 2, 2010. They were told to discontinue use and dispose of the affected lot of test strips. The firm also asked customers to complete an acknowledgement and confirmation document . |  | 5805 | 5805 | US accounts in: CA, MA, FL, NY, KY, VA, UT, MN, OH, and MD. Direct foreign accounts in: Belgium, England, Sweden, and Italy. | Terminated | May 11, 2011 | 309 | 5805
Z-2635-2010 | 56292 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EnsealÂż TRIO Product Code ETRIO335H, Lot Code: G4TC14, Exp. Date 2012/05 |  |  | July 09, 2010 | 2010 | Ethicon Endo-Surgery Inc | The weld that connects the electrode to the active rod was separated. This connection is critical to the proper function of these devices. If the weld is separated, the product will not deliver energy to the jaws, resulting in the potential for a cut/transection without a seal. Ethicon Endo-Surgery identified this issue during a routine finished goods quality inspection. | Other | Ethicon Endo-Surgery issued an Urgent: Device Recall, Event 2193 letter dated July 19, 2010 identifying the affected device, the issue prompting the recall, and actions to be taken by consignees.Consignees were instructed:Do not use the affected product.Identify and quarantine affected product from inventory.Complete the Business Reply form and fax to Stericycle at 877-597-9580 within three daysTo return product, place it in a box with a copy of the Business Reply form and mail to Stericycle using the pre-printed UPS shipping label. Replacement product should be received by the customer within 60 days of return of affected product. Customers can contact Stericycle about this recall at 877-243-7314 and referencing Event 2193. |  | 396 | 396 instruments | Nationwide Distribution to the following states: AZ, AL, AZ, CA, CO, CT, FL, GA, IN, KY, MA, MD, MI, MO, MS, NC, NV, NY, OH, PA, TN, TX, WA, and WV. | Terminated | February 14, 2011 | 220 | 396
Z-2652-2010 | 56481 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Endo Ocular Directional Laser Probe, Lumenis Inc. Salt Lake City, UT STERILE EO |  |  | July 27, 2010 | 2010 | Lumenis, Inc. | Firm did not have proper 510(K) clearance. | Component change control | Lumenis sent letters to all customers on 7/27/10 instructing them to review their inventory, isolate affected product, complete a recall verification form and return all affected product to them. |  | 2500 | 250 probes (25 boxes, each containing 10 probes) | FL, IL, MD, CO and TX | Terminated | October 05, 2010 | 70 | 2500
Z-2654-2010 | 49387 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Laparoscopic Ultrasound Transducer (LAP Transducer) is used on the SonoSite MicroMaxx ultrasound system. The LAP Transducer label is labeled in part: "MicroMaxx LAP...Manufactured for SonoSite, Inc., Bothell, WA 98021". |  |  | November 02, 2007 | 2007 | Sonosite, Inc. | Sterilization of the SonoSite LAP surgical Transducer with the Sterrad NX System may result in damage to the transducer. | Process control | On November 2, 2007, SonoSite sent out the Important Information Regarding Sterrad NX Incompatibility letter and the Technical Bulletin 005/07 Rev A to their consignees. The letter advises the consignees that SonoSite has recently found that sterilization of the SonoSite SLT and LAP surgical ultrasound transducers with the Sterrad NX system may result in damage to the transducer.The Technical Bulletin 005/07 Rev B dated November 2, 2007 informs the consignees that the damage has been observed primarily on the circuit board in the transducer connector, and cable and laparoscope damage may also be occurring. All damages have occurred after processing using the Sterrad NX systems in its "Advanced" cycle and the potential damage from the use of "normal" cycle is being evaluated by the firm.Consignees should call the Technical Support at 1-425-951-1330 or 1-877-657-8118 or e-mail to service@sonosite.com for any questions regarding the recall. |  | 16 | 16 units | Devices were distributed to medical centers and hospitals in AR, CA, IN, KS, MA, MD, MI, NC, NE, NJ, NV, OH, SC, TN, TX, WI, Washington DC and to FOREIGN COUNTRIES Australia, China, France, Germany, Greece, Hong Kong, India, Iraq, Italy, Korea, Republic of, Libyan Arab Jamahiriya, Mexico, Spain, United Arab Emirates, United Kingdom, and United States. | Terminated | October 05, 2010 | 1068 | 108
Z-2639-2010 | 54382 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Ultra Test Strips, Part Number 020-245-32, lot number 2964512, Manufactured by LifeScan Inc., Milpitas, CA. |  |  | January 07, 2010 | 2010 | Lifescan Inc | One Touch Ultra Test Strip products may actually contain a different product, One Touch Select strips. | Mixed-up of materials/components | Notification was issued by the firm during the first week of February. This letter asked consignees to share the information with patients, and are asking users to contact the firm and to examine their stock for the recalled product. |  | 38524 | 9631 cartons, 4 vials per carton | Product was distributed to one wholesale dealer and one distributor and 2 medical facilities in CA, OH and PA, as well as the UK. | Terminated | March 14, 2011 | 431 | 38524
Z-2658-2010 | 55196 | No | CEC | Free thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems Free T4, Part Number: 33880 |  |  | March 24, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter determined that kit boxes for Access Free T4 reagent (REF 33880), lot number 916293 may contain reagent packs labeled as Access Vitamin B12, (REF 33000), lot 916277. Using an Access Vitamin B12 reagent pack, lot 916277, filled with Access Free T4 reagents may generate incorrect patient results. | Other | A Product Corrective Action (PCA) letter was sent on March 24, 2010 to the affected customers. The letter provides the customers with an explanation of the problem identified and instructs them to discard the FT4 kit, Lot 916293 in their stock if the FT4 packs within the kit packaging are labeled asVitamin B12 Lot 916277.The PCA requested customers:(1) Inspect all Access Vitamin B12 reagent packs that have been removed from their kit boxes, including reagent packs currently loaded on your system and discard any reagent packslabeled as Access Vitamin B12, lot 916277;(2) Inspect your Access Free T4 reagent kit inventory for any reagent packs labeled as Access Vitamin B12, lot 916277 and discard any reagent packs labeled as Access Vitamin B12, lot916277; (3) Review your historical quality control to ensure that assay performance for Access Vitamin B12 lot 916277 has been within your laboratory specifications.In addition customers were instructed to share this information with their laboratory staff and retain this notification as part of your laboratory Quality System documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that you have received this important notification.If customers need assistance or have any questions regarding the notification, they were instructed to contact Customer Technical Support Center at 1-800-854-3633 in the United States and Canada. |  | 10826 | 10826 | Nationwide In the US &amp; Canada | Terminated | May 07, 2012 | 775 | 10826
Z-2656-2010 | 56565 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | AVI Sterile Impad Rigid Sole Foot Cover, Right Foot |  |  | August 11, 2010 | 2010 | Covidien LP | Covidien issued an Urgent Product Recall of the Sterile ImPad Rigid Sole Foot Cover. The product is indicated for prevention of blood pooling in the limb. 120 cases of each Lot code 6066 (right foot) and 6067 (left foot) were shipped to the distributor before sterilization. The customer is instructed to quarantine the product and return it to the manufacturer for replacement. If there are any | Release of Material/Component prior to receiving test results | Covidien issued an Urgent Product Recall of the Sterile ImPad Rigid Sole Foot Cover. The product is indicated for prevention of blood pooling in the limb. 120 cases of each Lot code 6066 (right foot) and 6067 (left foot) were shipped to the distributor before sterilization. The customer is instructed to quarantine the product and return it to the manufacturer for replacement. If there are any questions please contact, Covidien Customer Service at (800) 962-9888, Option 1. |  | 1984 | 1984 units (496 cases) | Nationwide | Terminated | June 12, 2012 | 671 | 6840
Z-2642-2010 | 56593 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Camera Head OTV-S7H-1D-L08E and Camera Head OTV-S7H-1D-F08E (designed to be used with Olympus endoscopic equipment) |  |  | August 20, 2010 | 2010 | Olympus America Inc. | Mislabeled. Endoscopic body-type camera heads labeled as cardiac-type. | Labeling mix-ups | Asked consignees to return camera heads to correct the labeling. |  | 217 | 217 camera heads | nationwide; camera heads were serviced and returned to consignees in March 2003 - September 2007 timeframe | Terminated | September 27, 2012 | 769 | 217
z-2640-2010 | 56652 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Sonopet Ultrasonic Surgical System, Model UST-2001, Manufacturer Miwatec Co. LTD, Tokyo, Japan. |  |  | July 15, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments acquired the Sonopet Ultrasonic Surgical System from Mutoh co. Ltd on 3/29/2010. Stryker determined that the Sonopet consode did not meet the IEC-60601-1 standard. Stryker decided to recall the Sonopet consode. | Device Design | The firm to send URGENT MEDICAL DEVICE RECALL NOTIFICATIONS to customers dated 7/15/2010. The customers were instructed to continue to use the device. Stryker technicians were dispatched to replace a board with a new board. |  | 1161 | 1161 | Wordwide | Terminated | October 19, 2011 | 461 | 1161
Z-2583-2010 | 56343 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | SynFix-LR Spinal Implant System (Implant Holder for SynFix-LR and surgical technique brochure) |  |  | July 23, 2010 | 2010 | Synthes USA (HQ), Inc. | A potential for the implant holder tip to break off and lodge in the SynFix-LR implant, which pose a potential risk to the patient. | Error in labeling | Synthes Spine Hospital sent an ' URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated August 20, 2010 to all customers. The letter described the labeling issue and actions to be taken by the customers. The customers were instructed to review the supplementary labeling provided; contact their Synthes Spine Sales Consultant to arrange for provision of the Holding Forceps (p/n 388.401) prior to each SynFix-LR procedure; and complete and return the enclosed Verification Section form to Synthes via fax at 610-251-9005 or email: FieldAction@Synthes.com.If you have any questions, please call 1-800-620-7025 x 3500 or contact your Synthes Spine Sales Consultant. |  |  |  | Worldwide Distribution: USA and the countries of Canada, Japan, Switzerland, Australia, Mexico, Malaysia, New Zealand, and Singapore. | Terminated | December 18, 2014 | 1609
Z-0003-2011 | 56411 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | BVI 9600 Bladder Volume Instrument, Aorta Scan Mode |  |  | August 23, 2010 | 2010 | Verathon, Inc. | The BladderScan BVI 9600 with AortaScan devices may experience the loss of a factory installed software configuration file which results in a loss of the aorta measurement function. | Software design | Verathon sent an URGENT Medical Device Recall Notification letter dated August 12, 2010 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Verathon provided an addendum to the User Manual with instructions on how to reset the instrument. Customers were advised to:"Add the enclosed addendum to their current Users Manual."Update the application software of the device to v2.7.0 or higher. If the customers had questions or needed support for the software update process, they were to contact Verathon Medical Customer Care at (425) 867-1348 or (800) 331-2313 orcservice@verathon.com. |  | 351 | 351 units | Worldwide Distribution - USA and the countries of Australia, Canada, China, Colombia, European Union, United Kingdom, Saudi Arabia, and Singapore. | Terminated | September 08, 2011 | 381 | 684
Z-0014-2011 | 50100 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Duracon Total Knee Distal Femoral Locking Screw for 5 mm Spacer; Product No. 6632-5-600. |  |  | October 05, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that a box labeled as Duracon Total Knee Distal Femoral Locking Screws for 5 mm spacer, may actually contain a 10 mm screw. | Labeling mix-ups | Important Market Withdrawal Letters were sent via Federal Express October 5, 2006. The letters identified the affected product and the reason for recall. Customers were asked to examine their inventory to identify the product; and retrieve all affected product lots and return it their branch or agency warehouse for reconciliation. All product should be reconciled on the Product Accountability Form and the form should be faxed back. Customers are to contact a Customer Service Rep to re-order the product that is being returned. Questions can be directed to Rita Intorrella on 201-831-5825. |  | 7 | 7 units | Nationwide Distribution -- Including Ohio. | Terminated | October 07, 2010 | 1463 | 7
Z-0012-2011 | 49138 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Essenta DR is a digital multifunctional X-ray systems, suitable for all routine radiographic exams, including specialist areas like trauma or pediatric work, excluding mammography. |  |  | August 15, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | When a patient is selected in the patient list and examination tab pressed with background query running at the same time, the wrong patient may be selected. | Other | Philip Medical System sent an "URGENT-Device Correction" letter dated August 13, 2008 to their customers. The letter describes the recalled product as Essenta DR Patient and Image Data Software (SW) release 1.0. The firm informs the customers that when patient is selected in patient list and examination tab pressed with background query running at the same time, there is a risk that the wrong patient is selected on the Essenta DR Computed Radiography System.Customers are advised to verify the selected patient's information before releasing an exposure while the firm works on upgrading the Software. All affected units will receive a software upgrade to correct the issue.Customers should contact Philips Customer Care Service at 800-772-9377, #5, #4, #1 and reference FCO 712000031 for any questions or support concerning the recall. |  | 23 | 23 units | Nationwide distribution: CA, GA, FL, KS, MN, NY, OH, OR, SC, TN, TX, UT, and WA. | Terminated | October 07, 2010 | 783 | 23
Z-0013-2011 | 56026 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Summit Broach Handle, NON-STERILE, MFG: DEPUY ORTHOPAEDICS, INC., PO BOX 988 WARSAW, IN 46581, 1 800 366 8143 |  |  | June 15, 2010 | 2010 | Depuy Orthopaedics, Inc. | Firm discovered that some of their straight broach handles are experiencing early failure due to a weakened handle locking mechanism (Leaf Spring). | Process control | On June 3, 2010, an "Alert: Femoral Broach Handle Repair" went out to the entire sales force/customers via "Sales Mail" and posted to the DePuy Internal website. The Alert stated that "We have discovered that some of our straight broach handles are experiencing early failure due to a weakened handle locking mechanism (Leaf Spring).As you perform routine maintenance on your sets, if you have any of these broaches stamped with the lot numbers listed, it is recommended you have them sent in to be reworked with a new spring device."Send the broach handles to: DePuy OrthopaedicsAttn: Repairs700 Orthopaedic DrWarsaw, IN 46581Please include return shipping informationIf you have any questions regarding this information or need additional information, call 574-372-7333. |  | 8358 | 8358 | Nationwide distribution: AR, AZ, CA, FL, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, NC, NV, NY, OH, PA, TN, VA, and WI. | Terminated | November 30, 2011 | 533 | 8358
Z-0011-2011 | 56728 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | MAQUET HLM Tubing set with Bioline Coating (BEQ-TOP 21400 with 1/2 x 1/2 HG 0286 Connector; Device Part Number BEQ-TOP 21400 |  |  | September 07, 2010 | 2010 | Maquet Inc. | Tubing packs may contain defective 1/2 inch x 1/2 inch tubing connectors. Leaks at the shank of the connector may result in air aspiration in the venous line. | Device Design | Maquet Getigne Group issued Product Recall/Urgent Medical Device Corrective Action letters dated September 7, 2010 to customers, identifying the affected product, the issue prompting the recall, and actions to be taken by customers.Customers are to check their inventory for affected product. Customers with affected product are to contact Laura Olilver at 201-995-8849 for an RMA number and shipment instructions. All firms are to coomplete and return the Corrective Action Response Form.Maquet Customer Service can bre contacted at 800-777-4222. |  | 22 | 22 items | Worldwide Distribution: USA State of IL and country of the Netherlands. | Terminated | September 25, 2012 | 749 | 22
Z-0010-2011 | 56599 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Datex-Ohmeda Amingo Anesthesia System, manufactured by GE Healthcare, P.O. Box 900, FIN-00031 GE, Finland |  |  | August 30, 2010 | 2010 | GE Healthcare, LLC | Users may inadvertently leave a test plug in the inspiratory flow sensor after completing Low P Leak Check on the SCGO (Switched Common Gas Outlet) configuration of Avance, Amingo, and Aisys anesthesia machines that may impact patient safety.If the test plug remains in place and therapy is initiated, the machine will not provide gas to the patient. This would inhibit mechanical or manual ventila | Employee error | GE Healthcare issued an " Urgent Medical Device Correction" letter dated August 30, 2010 to consignees. The letter was addressed to Chief of Anesthesia, Health Care Administrator/Risk Manager and Director of Biomedical/ Clinical Engineering. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Instructions.Customers were instructed to:1. Ensure the plug is removed at the completion of the Low P Leak check.2. Review the Preoperative Checkout procedure listed in section 4 of the User Reference Manual with appropriate personnel. Ensure that the Preoperative Checkout procedure is completed in its entirety prior to each case, or as required.All facilities will be contacted by their local GE GE Field Technical Service Team to schedule installation of a software update and a revised test plug for your machines configured with SCGO. Customers are to contact their local GE Healthcare Customer or Technical Support Representative if they need additional information. |  | 131 | 131 | Worldwide Distribution: USA: All States, including DC, except Delaware, and the countries of YEMEN, VIETNAM, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SYRIA, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI, ARABIA, RUSSIA, PORTUGAL POLAND, PERU, PANAMA, NORWAY, NETHERLAND, MACEDONIA, LITHUANIA, LEBANON, LATVIA, KUWAIT, KOREA, KENYA, JORDAN, JAPAN, ITALY, IRELAND, INDIA, HUNGARIA, HONG KONG, HONDURAS, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CZECH REPUBLIC, CROATIA, COLOMBIA, CHINA, CHILE, CANADA, BULGARIA, BRAZIL, BERMUDA, BELGIUM, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | September 13, 2012 | 745 | 7711
Z-0021-2011 | 56707 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | The BABIX holder is an accessory used with the BuckyDiagnost X-ray system. It is used to hold the baby either in a secured or immobilized standing position in front of the X-ray detector during the examination. |  |  | September 09, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Potential of broken strands in the main steel cable supporting the accessory holder used with BuckyDiagnost and DigitalDiagnost X-ray systems. | Equipment maintenance | Philips issued URGENT-Device Correction Notice letters dated September 9, 2010 to their consignees identifying the affected products and action to be taken by customers. The letter states that Babix holder used with Philips BuckyDiagnost and DigitalDiagnost X-Ray System - VS Wall Stand Risk of steel cable fracture in the support holder. The consignees are informed that there is a potential of broken strands at the main steel cable within the holder which cannot be recognized by the user. The firm will issue the Field Change Order (FCO) and the Service Engineer will replace the Babix holder.Consignees can call Philips Call Center at 800-722-9377, #5 (then follow prompts) and reference to FCO 70400038 for the BuckyDiagnost X-ray system and FCO 71200046 for the DigitalDiagnost X-ray system. |  | 4 | 4 units | Nationwide Distribution in the states of CA, GA, and TN. | Terminated | July 22, 2011 | 316 | 5
Z-0023-2011 | 56663 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BD 30G1/2 PrecisionGlideÂż Needle;Catalog/Lot Number: 305106. |  |  | August 24, 2010 | 2010 | Becton Dickinson &amp; Company | There have been reports of blocked/clogged needles for the 30 gauge PrecisionGlide Needles | Pending | BD sent out Product Correction letters to all end users of the product on 8/24/10. Letters were mailed by BD via UPS second day air. The letter stated the issue and said that customers should assess the patency prior to injection. The letter also stated that a BD 30 gauge PrecisionGlide Needle from another lot can be used in place of a clogged needle. If customers require replacement needles or further assistance, they can contact BD at 1-888-237-2762. Finally, the letter states that BD has implemented corrective actions to reduce the potential for further recurrence of the blockage issue. |  | 2599000 | 2,599,000 units | Worldwide Distribution -- USA and Canada. | Terminated | August 25, 2011 | 366 | 2599000
Z-0020-2011 | 56536 | No | GBS | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | HydraGlide Silicone Chest Tube Catheter, 32 Fr Straight |  |  | August 10, 2010 | 2010 | Atrium Medical Corporation | Mislabeled: One lot of 32Fr Straight Catheters packaged and labeled as 24Fr Straight Catheters. | Packaging process control | Atrium Medical isued Product Recall Notice letters dated August 10, 2010 to customers, identifying the affected product and actions to be taken by customers. Customers were instructed to remove the affected products from inventory and return them to Atrium, and complete and fax a recall notice acknowledgement to Atriuim at 603-821-1420. Customers can contact Atrium at 800 528-7486, x5350 (Joe DePaolo) or x5351 (Karen Hall) for further information. |  | 10 | 1 case, 10/case | Nationwide Distribution: USA, including the states of AK, CT, OH, and UT. | Terminated | April 22, 2016 | 2082 | 110
Z-0017-2011 | 56630 | No | KRG | Pacemaker programmers. | Cardiovascular | Cardiovascular | PMA | Guidant, Vitality HE ICD, Model: T180DDR, Sterile Lot, Guidant Corporation, Cardiac Rhythm Management, 4100 Hamline Avenue North, St. Paul, MN 55112-5798. |  |  | July 22, 2010 | 2010 | Boston Scientific CRM Corp | Exposure to a magnet (typically in a clinic/hospital environment) may cause a magnetic reed switch to become permanently stuck in a closed position in certain Boston Scientific CRT-Ds and ICDs (CONTAK RENEWALÂż 3 / 3 RF / 4 / 4 RF and VITALITYÂż HE implantable defibrillators). This prevents delivery of programmed tachy therapy, and use of the Patient Triggered Monitor feature. This issue may require | Nonconforming Material/Component | Boston Scientific Issued Urgent Medical Device Correction and Physician Device Advisory Notice letters dated July, 2010 to customers identifying the affected devices, the issue prompting the recall, clinical implications, and actions to be taken by customers.Customers were instructed:1) Physicians should continue routine follow-up sessions, and patients should be reminded to contact their clinic or go to the hospital emergency room immediately if they hear tones/beeps from their device. 2) In a hospital/clinic/surgery setting, if tones are heard upon magnet application but do not cease upon magnet removal, the device should be interrogated with a programmer and checked per normal standard of care. 3) In the United States, use of the LATITUDE remote patient monitoring system may help identify loss of daily measurements and thereby facilitate timely detection of a stuck reed switch. [NOTE 1: A pop-up message and/or LATITUDE alert do not appear for missing Daily Measurements. NOTE 2: Daily Measurements may be unavailable for reasons other than a stuck switch.] 4) Prophylactic explant is NOT recommended.5) If a stuck magnetic switch is confirmed, program the Enable Magnet Use feature to Off, which ensures that programmed therapy will be provided to treat tachyarrhythmias. However, if Enable Magnet Use is programmed Off, a magnet will no longer inhibit tachy therapy, and the Patient Triggered Monitor feature will no longer be available. Do not routinely program Enable Magnet Use to Off in the absence of a confirmed stuck magnetic reed switch because the benefits of magnet use to disable tachy therapy in emergent situations outweigh the probability of patient harm associated with a stuck reed switch.In the US, Boston Scientific Cardiac Rhythm Management can be contacted at 800-227-3422, and in Europe, Boston Scientific Europe S.A. can be contacted at +32 2 416 7222. |  | 5821 | 5821 | Worldwide Distribution: All States in the continental USA including PR and DC , but excluding VT, the countries of Canada, Mexico, and in Europe.ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BERMUDA, BRITISH VIRGIN ISLANDS, CANADA, CAYMAN ISLANDS, CHILE, COSTA RICA, CYPRUS, CZECH REPUBLIC, CZECHOSLOVAKIA, DENMARK, EGYPT, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, LUXEMBOURG, MALAYSIA, MEXICO, NETHERLANDS,NEW ZEALAND, NORWAY, POLAND, PORTUGAL , QATAR, SAINT KITTS &amp; NEOIS, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM,VENEZUELA, URUGUAY, ZIMBABWE, IRAN. | Terminated | April 03, 2012 | 621 | 38377
Z-0018-2011 | 56600 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Oxylog 3000, Emergency and Transport Ventilator; Catalog #2M86300; Drager Medical Inc., 3135 Quarry Road, Telford, PA 18969. |  |  | August 19, 2010 | 2010 | Draeger Medical Systems, Inc. | Patients received insufficient ventilation. The Oxylog 3000 Instructions are inadequate relating to dead space volume. | Error in labeling | The firm sent Important Safety Notice Letter, dated August 2010, and Instructions for Use supplement to consignees. The letter explained the issue with the affected product. The letter asks customers to attach the Instructions for Use amendment sheet supplied with the letter to their corresponding Instructions for Use. The firm also states that a they will provide a dedicated pediatric breathing circuit for the Oxylog 3000. Questions regarding the letter should be directed to Michael Kelhart at 1-800-543-5047. Questions regarding the use of the device should be directed to DragerService Technical Support at 1-800-543-5047 or your local Drager Sales Representative. |  | 78 | 78 potentially afftected units | Nationwide Distribution | Terminated | November 18, 2010 | 91 | 78
Z-0015-2011 | 56219 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Exodus(TM) 8F (2.7 mm)/25 cm Standard Loop Multipurpose Drainage Catheter with Hydrophilic Coating, REF Catalog No. 10002, UPN Product No. H965100201, Rx ONLY, STERILE/EO --- Made in USA: 2301 Centennial Blvd., Jeffersonville, IN 47130 --- Manufactured for: Navilyst Medical, Inc. 26 Forest Street, Marlborough, MA 01752 USA --- Each box unit contains 5 catheters. --- The Exodus Standard Loop Multipurpose Drainage Catheter consists of a radiopaque polyurethane catheter with a locking pigtail distal end and is coated with a hydrophilic coating. Catheter markings are provided in centimeters on the shaft to indicate the distance form the distal tip. The catheter is provided with a trocar, a metal stiffening cannula and a plastic cannula. |  |  | May 24, 2010 | 2010 | Navilyst Medical, Inc | One box containing five catheters did not complete the sterilization process prior to distribution. | Employee error | On/about 5/24/10, NMI telephoned the consignee. The recall notification package was sent by NMI to the sole consignee on 6/11/10, followed by a telephone call by NMI Quality staff. On 6/15/10, the recall notification package was sent via FedEx to the sole consignee. Subsequently the consignee was provided a return goods authorization (RGA) number for return to NMI. The letter described the product, problem and action to be taken by the consignee/customer. The consignee/customer was instructed to immediately discontinue use of and segregate recalled product; remove the affected lot from their inventory; complete and return the Reply Verification Tracking Form via fax to the QA Manager; Navilyst Medical, Inc., at 1-518-742-4414; to please call the Navilyst Medical Customer Service at 1-800-833-9973 between 8:30 a.m. and 7:00 p.m. (Eastern Standard Time) to obtain a Return Authorization (RGA) Number, and package/ship the recalled product to Navilyst Medical, Inc, Distribution Center, Attn: QA Returns, 10 Glens Falls Technical Park, Glens Falls, NY 12801. If you have any questions, call Navilyst Medical Customer Service at 1-800-833-9973. Subsequent to the initial notification letter, NMI identified that a typo was made to the lot number on the letter. NMI telephoned the consignee and sent a revised letter via fax on 6/25/10 indicating the correct lot number. |  | 1 | 1 box unit (contains 5 catheters) | Nationwide distribution: Ohio only | Terminated | October 07, 2010 | 136 | 1
Z-0034-2011 | 50215 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The device's marketing brochure states: "IntelliVue M3150 Information Center...Philips Medical Systems". |  |  | November 21, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Potential for Internal Sound Amplifier to malfunction which could delay treatment. | Component change control | On 11/20/08, Philips began sending the URGENT-Medical Device Correction letter to their consignees via e-mails. On 11/21/08, the firm sent out the recall letter via United State Postal Service.The letter states the affected device as Philips IntelliVue Information Centers, running on the HP rp5700 PC platform equipped with an active speaker solutions using a DB-15 to 3.5 mm cable set.The consignees are advised to ensure that the audible alarm system is clear, free of noise and distortion and adjustable to different sound levels which is explained in p. 5-47 "Adjusting Alarm Tone Volume" of IntelliVue of the Instruction for Use, Release K Manual, supplied with the hardware, Philips Part Number 453564062301.Consignees can call Philips Healthcare Customer Care Service Center at 1-800-722-9377, #3, #1 and reference FCO 99862012. |  | 424 | 424 units in US. | Nationwide Distribution | Terminated | October 12, 2010 | 690 | 424
Z-0027-2011 | 54527 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | SpeedLink 5.5 Transverse Connector Medium 43-51mm, non-sterile, 1302-55M |  |  | January 01, 2009 | 2009 | Zimmer Inc. | The firm received complaints that the device was damaged or fractured while the surgeon was bending it prior to implantation. There were also complaints that the cam may dislodge during tightening. The firm added additional instructions for use to reduce the probability of these two types of events. | Other | The firm provided a revised surgical technique document to the Zimmer spine sales force in January, 2009. The firm will provide an Urgent Medical Device Correction notification to speedlink users through the sales force. |  | 23205 | 23,205 Units for ALL PRODUCTS IN THIS RECALL | The devices are contained in implant cases stocked by Zimmer spine and made available to the Zimmer sales force for surgeries. They are not in general distribution. | Terminated | January 05, 2011 | 734 | 23205
Z-0025-2011 | 56767 | No | MJG | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Tru-Core II Disposable Automatic Biopsy Instrument and Tru-Core II URO Disposable Automatic Biopsy Instrument. The Tru-Core Disposable Automatic Biopsy Instrument is used to obtain multiple core samples from soft tissue such as liver, kidney, prostate, breast, ect. The recalling lot's for the product are as follows: |  |  | July 02, 2010 | 2010 | Medical Device Technologies, Inc. | Medical Device Technologies, Inc., dba Angiotech, Gaineville, Florida is recalling Tru-Core II Disposable Biopsy Devices due to compromised packaging integrity. The Tru-Core Disposable Automatic Biopsy Instrument is used to obtain multiple core samples from soft tissue such as liver, kidney, prostate, breast, ect. The recall is for the following Model Numbers: 763114100X, 763114160X, 763114 | Other | Angiotech sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated July 2, 2010 and response card to all customers including distributors. The letter described the product, problem and action to be taken by the customers. The customers were instructed to review their current inventory and segregate the affected lot numbers; arrange for return and replacement of the affected product, and complete and return the Customer Acknowledgment Form via fax to the Quality Assurance department at 1-352-338-0662 or 1-800-333-0440.If you have any questions concerning this notification, contact Shannon Brooks at (352) 338-0440 ext 355 or Katrenia Williams at ext 353. |  | 11123 | 11,123 boxes (111,230 units) | Worldwide distribution: USA and countries including: Argentina, Australia, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Israel, Japan, New Zealand, Peru, South Africa, , Uruguay, and Venezuela. | Terminated | June 28, 2011 | 361 | 11123
Z-0024-2011 | 56645 | No | KRL | Cardiopulmonary bypass bubble detector. | Cardiovascular | Cardiovascular | 510(k) | Air bubble detection cable assembly, catalog numbers 78-8067-7375-6 and 149892, Terumo Cardiovascular System, Ann Arbor, MI |  |  | October 15, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | Malfunction of the air bubble detection cable assembly causes a false air alarm or a check sensor message during system set-up. In some instances, users were not able to reset the alarm to allow for further air bubble detection. . | Other | The firm sent URGENT MEDICAL DEVICE CORRECTION notices dated September 17, 2010. The notices told consignees that a Terumo representative would visit, inspect the cables and replace the recalled cables. if the customer has a service contract, the visits would occur at the next scheduled preventative maintenance visit. Customers without a service contract will be contacted by Terumo representative to schedule an appointment.Customers can contact Terumo Customer Service at 800-521-2818. |  | 5000 | 5,000 estimated | Worldwide Distribution: USA states and territories of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, WA, WI, WV, and WY, and the countries of Australia, Belgium, Canada, Columbia, Chile, Costa Rica, Germany, Dominican Republic, Ecuador, Hong Kong, Indonesia, Japan, Korea, Mexico, Malaysia, Panama, Philippines, Saudi Arabia, Singapore, South Africa, Thailand, Taiwan, UAE and Vietnam. | Terminated | June 29, 2012 | 623 | 5000
Z-0040-2011 | 55602 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith &amp; Nephew Aimer Endofemoral 3mm Offset (Blue) |  |  | April 29, 2010 | 2010 | Smith &amp; Nephew, Inc. Endoscopy Division | Endofemoral Aimer tip may separate due to failure of the soldered joint | Component design/selection | Smith &amp; Nephew sent a "Medical Device Recall Information" letter dated April 29, 2010, with an attached "Urgent-Product Recall 1st Notification - Urgent" response form via Federal Express on April 29, 2010, to all customers. An e-mail was sent to the contacts in the identified OUS countries alerting them of the recall. ( Note: On June 24, 2010, the firm expanded the recall to include 4 additional catalog numbers 7220l713F, 72201714F, 72201715F, 722017l6F that are refurbished and one additional lot number to catalog Number: 72201714.) The letter described the product, problem and action to be taken by customers. The customers were instructed to immediately check stock for affected product; immediately remove the items from stock and set aside for quarantine; complete and return the attached Urgent-Product Recall 1st Notification - Urgent" response form, even if no product is being returned, via fax at +1-508-261-3636 attention Cindy Burns or email a PDF to Cindy.Burns@smith-nephew.com; call for a return Authorization number at 1-800-343-5717 Option #3, and return product to Smith &amp; Nephew, Inc., Endoscopy Division as soon as possible. If you have any questions or require additional information, please contact Cindy Burns at 1-508-261-3655. |  | 577 | 577 units | Worldwide distribution: USA and countries including: Australia, Austria, Belgium, Brazil, Canada, Costa Rica, Denmark. Dubal, Finland, France. Germany, Greece, Hong Kong, India, Italy, Japan, Malaysia, Netherlands. Norway, Portugal, Puer To Rico, Singapore. South Africa, Spain. Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela. | Terminated | May 09, 2012 | 741 | 3607
Z-0036-2011 | 55802 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | HLD SYSTEM 520 is a two-tank Medical Device Cleaning and High Level Disinfection Systems (Washer/Pasteurizers). Each unit has a metal serial tag that is riveted to the enclosure of the unit and it is labeled in part: "HLD SYSTEMS 520***FOR SALES.SERVICE.SUPPLIES CALL***Cenorin*** Kent, WA 98032***P:800.426.1042***F: 253.395.2650***www.cenorin.com". |  |  | June 07, 2010 | 2010 | Cenorin | The water level sensor in the Medical Device Cleaning and High Level Disinfection (Washer/Pasteurizer) is sensitive to externally induced vibration noise, and it produces a false signal that the water level in the tank is full while it is only 75% full. It could present a health risk of biological hazard because the medical devices are not being adequately cleaned or pasteurized with enough wat | Component design/selection | On June 7, 2010, firm, CENORIN, sent "URGENT: MEDICAL DEVICE RECALL" Initial Notification letter with Recall Response Form, Series 500 Water Level Check Instructions, and the Water Level Switch Upgrade Kit Instruction (service kit P/N # 082395-00) to their consignees. The firm also sent these documents in Spanish language to their foreign consignees. The letter describes the recall product as HLD Systems Medical Device Cleaning and High Level Disinfection Series 500 Washer/Pasteurizer-Model 540, Model 540HT and Model 520. Customers are advised to call the Customer Service at 1-800-426-1402 or e-mail the firm.The Initial Notification letter instructs consignees to conduct and quarantine the recall products and notify their sub-accounts (customers) if products have been further distributed. While waiting for a replacement kit, the consignees are asked to follow the Series 500 Water Level Check Instructions to perform a test to ensure the water tank is completely filled up the expected water level. The firm also advised the consignees to add a cautionary label on the product to avoid any sharp or bump impact that may create any externally induced vibration noise to the device.The firm is developing a replacement kit that should be ready to send out to their consignees by June 30, 2010. A response card would be also sent out with a new kit, and it is used to verify that the kit has been successfully implemented. If you have any questions, please email jbennett@cenorin.com or call 253-395-2400 or toll free: 1-800-426-1042. |  | 9 | 9 units | Worldwide distribution:USA and countries including: Costa Rica, Ecuador, Japan, Mexico, and Taiwan. | Terminated | May 07, 2012 | 700 | 93
Z-0001-2011 | 56425 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2M8151, 2M8161, and 2M9161. |  |  | August 04, 2010 | 2010 | Baxter Healthcare Corp. | The FDA sent a letter to Baxter on April 30, 2010, ordering the company to recall and destroy all models of its Colleague Volumetric Infusion Pumps currently in use in the United States. FDA determined that this action is necessary, as Baxter has failed to adequately correct, within a reasonable timeframe, the deficiencies in the Colleague pumps still in use. This recall is a consolidation of al | Device Design | Baxter Healthcare Corporation sent "Urgent Product Recall" letters dated August 4, 2010 to all Colleague Infusion Pump customers in the U.S., to the attention of the Hospital Administrator, Chief Executive Officer, Chief Operating Officer, Chief Medical Officer and Chief Nursing Officer, referencing Baxter's May 3, 2010 announcement to recall all Colleague pumps from the U.S. market pursuant to an order under its existing consent decree with the FDA. The letters provided the customers with important information regarding the actions they must take prior to November 14, 2010 in order to continue using their Colleague pumps during the transition period. Customers were instructed to complete the attached Certificate of Medical Necessity Form and return it to Baxter as soon as possible, but no later than November 14, 2010, by scanning and e-mailing to Baxter at FCA@Baxter.com. Any questions were directed to their local Baxter Sales Representative or by calling the Center for One Baxter 1-800-422-9837, between 8AM and 5 PM Central Standard Time (CST).*** On July 20, 2011 Baxter sent recall letters to those Colleague customers who had not completed a decision form regarding the status and replacement of their Colleague Pumps. The accounts were asked to inform their Baxter sales representative as soon as possible of their pump decision, and to complete the attached customer reply form and return it to Baxter via fax or scanned e-mail. *** |  | 221097 | 221,097 pumps | Nationwide, including Puerto Rico | Terminated | November 26, 2012 | 845 | 302872
Z-0006-2011 | 56257 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CareFusion Alaris PC Unit Model 8015 (formerly Medley |  |  | July 13, 2010 | 2010 | CareFusion Corporation | Under certain wireless network conditions a communication error can occur, which freezes the PC Unit screen, which may result in a delay of therapy. A delay of therapy may result in serious injury and/or death. CareFusion identified this potential risk associated with the Alaris PC unit model 8015 only (PC Unit) manufactured between December-2008 and September-2009. | Component design/selection | CareFusion sent the customer notification letter to all customers using the affected Alaris PC Unit beginning August 24, 2010. The customer notification letter is addressed to the Directors of Nursing, Risk Management, and Biomedical Engineering. Customers will be required to confirm receipt of the notification by returning the Recall Response Card to CareFusion by postage-paid, self-addressed mail, fax or email. The recall notice informed the customers of the reason for the notice, the units affected, the potential risk, a step-by-step description of the issue and the required actions for users.The customers were also told that CareFusion will contact their facility through phone or in person within 60 days to initiate implementation of the corrective action. The corrective action will require a hardware update to all affected units. CareFusion does not require that you return your devices. In the interim, if you observe this issue, remove the device from service and contact CareFusion Recall Center immediately.Customers were also told that any adverse reactions experienced with the use of this product, and/or quality problems should also be reported to the FDA's MedWatch Program. The recall notice provided contact information for CareFusion's Recall Center, Customer Advocacy and Technical Support units. |  | 17081 | 17,081 units (17,050 in US) | Nationwide and worldwide to: Canada, Netherlands, New Zealand and Singapore | Terminated | December 12, 2011 | 517 | 17081
Z-0050-2011 | 56680 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Scorpio/ NRG 4in1 Cutting Block |  |  | August 23, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker s has received reports in which the Cutting Blocks have fractured on impaction during surgery. | Device Design | Stryker Orthopaedics sent an "URGENT PRODUCT RECALL" letter dated August 30, 2010, to all consignees/customers (branches/agencies, hospital risk management, hospital Chief of Orthopaedics, surgery) including foreign countries. The letter described the product, problem, and action to be taken by customers. The customers were instructed to complete and return the attached PRODUCT RECALL ACKNOWLEDGMENT FORM via fax to 201-831-6069.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance (201) 972-2100. |  | 839 | 839 (total) | Worldwide distribution: USA and countries including: Asia, Canada, Europe, Middle East and Africa | Terminated | August 30, 2012 | 738 | 839
Z-0058-2011 | 56694 | No | DSQ |  |  | Cardiovascular | PMA | HeartMate II System Controller, a component of the HeartMate II Left Ventricular Device System (LVAS), distributed both as stand alone devices and as components of the LVAS Implant kit. Manufactured/Distributed by Thoratec Corporation, Pleasanton, CA |  |  | September 01, 2010 | 2010 | Thoratec Corp | Connection failure-- product may malfunction because of bent pins or broken wires in power cord connection | Device Design | Thoratec Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated August 30, 2010, to all customers. The letter described the affected product, description and symptoms of problem, and immediate action to be taken by the customers. The customers were instructed to:1) contact all ongoing patients using the affected Systems Controllers (serial numbers starting with "EPC") to determine if they are experiencing any unexplained intermittent alarms or have any bent pins in the connectors between the System Controller and the power sources, if so, have them return to the hospital so that their power leads can be examined. If further assistance is needed with the evaluation, please contact Thoratec Technical Service (800-456-1477).2) review the attached addendum to the HeartMate II labeling with all VAD clinicians and ongoing HeartMate II LVAS patients, and3) complete and sign the attached Acknowledgment Form via fax to Thoratec Regulatory Affairs at (952) 847-8571.If you have any questions, please contact 925-847-8600. |  | 8965 | 8965 units | Worldwide distribution: USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX , UT, VA, WA, and WI; and countries including: Austria, Belgium, Canada, Czech Republic, Denmark, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Lithuania, Netherlands, Singapore, Slovenia, Sweden, Switzerland, and the UK. | Terminated | April 07, 2011 | 218 | 8965
Z-0060-2011 | 56651 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Trabecular metal Modular Acetabular system 56 mm Porous shell with cluster holes, Catalog 00-6202-056-22, Zimmer Inc., Warsaw, IN. |  |  | August 18, 2010 | 2010 | Zimmer Inc. | The threads on the device were not completely machined and therefore would not mate with athe corresponding cup positioner/shell inserter. | Other | The firm, Zimmer, made telephone calls using a phone script document entitled "Urgent: Medical Device Recall" Phone Script Call Communication, to the customers that received this product starting on 8/18/2010. The telephone calls were to determine the status of the recalled product. A Zimmer Field Action Strategy notification dated August 9, 2010 was also provided. The notifications/phone script described the product, problem and actions to be taken by the customers. The customers were instructed to remove the 00-6202-056-22, lot 61091653 product from their inventory and return to Zimmer Product Service Department, Attn: Product Service, 1777 West Center Street, Warsaw, IN 45680. Note: If the product had been used, no action was necessary. The returned product will be evaluated to determine how many of the devices had the defect.If you have any questions, please call (574) 371-8852. |  | 24 | 24 | Worldwide distribution:USA including states of: MI, IN, WI, MO, CA, VA, MD, TX, NC, PA, OR, WA, FL, GA, NY, ME, AZ, and UT and country including Canada. | Terminated | February 07, 2011 | 173 | 24
Z-0059-2011 | 56589 | No | HXO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Closed knot pusher REF 902813, Biomet Sports Medicine, Warsaw, IN. |  |  | August 02, 2010 | 2010 | Biomet, Inc. | The firm has received four complaints that the closed knot pusher is missing a radius which contributes to sutures being cut during surgery. | Device Design | Biomet Sports Medicine sent URGENT MEDICAL DEVICE RECALL NOTICES dated 8/2/2010 to its customers and distributors. The notices stated that the use of the device should be discontinued and the products returned to Biomet.Healthcare professionals and consumers were to report serious adverse events (side effects) or product quality problems with the use of this product to the FDA's MedWatch Adverse Event Reporting program either online at www.fda.gov/MedWatch/report.htm, by regular mail to MedWatch, 5600 Fishers Lane, Rockville, MD 20852-9787 or by phone at (800)FDA-1088 or Fax (800)FDA-0178.Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8AM to 5PM. |  | 101 | 101 | Worldwide Distribution - USA, including FL, NY, VA, CA, GA, AL, PA, MI, LA, KS, SD, KY, TX, AR, and WV, and the countries of Belgium, Italy, Argentina, Poland, Netherlands, Turkey and Portugal. | Terminated | August 20, 2012 | 749 | 101
Z-0071-2011 | 55157 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Continuum Acetabular Systems Shell Inserter Adapter with Rotational Control, 00-8790-003-00, Zimmer Inc., Warsaw, IN. |  |  | March 19, 2010 | 2010 | Zimmer Inc. | The firm has received reports of users encountering difficulties in usage leading to surgical delays or minor damage to a threaded surface of the shell implant. | Device Design | The firm sent Urgent: Device Removal letters dated 3/15/2010. The letter identified the affected product, described the issue, discussed the clinical implications, and provided a list of required actions. The consignees were instructed to perform an inventory of the recalled product and report the amount on hand to Zimmer by completing the Inventory Return Certification Form. After Zimmer received the report, they would send out new product to the consignee. After new product was received the consignees were instructed to return the recalled product to Zimmer. The firm expects to have all product replaced by 4/30/2010. If there are any questions or additional info needed, consignees should contact Zimmer, Inc. at 1-800-613-6131 or 574-372-4463. |  | 1125 | 1125 all products | Worldwide Distribution -- US, Singapore, Australia, Portugal, Germany, Switzerland, Spain, Finland, United Kingdom, Israel, Italy and Sweden. | Terminated | January 12, 2011 | 299 | 1125
Z-0068-2011 | 56719 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Architect Tacrolimus Reagent Kit; an in-vitro diagnostic chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of tacrolimus in human whole blood; Abbott Laboratories, Abbott Diagnostics Division, Abbott Park, IL 60064 USA; list 1L77-25 |  |  | September 07, 2010 | 2010 | Abbott Laboratories | There are reagent barcode read errors when using reagent lot 86599M500 due to the print quality of the barcode. | Other | Abbott Laboratories sent a "Product Recall Immediate Action Required" letter dated September 7, 2010, to all customers. The letter described product, problem and action to be taken by customers. The customers were instructed to determine if they are currently using and/or have inventory of lot 86599M500; discontinue use of and destroy any remaining inventory; inform and provide copies of letter to other laboratories if appropriate; retain the letter for laboratory records and complete and return the Customer Reply Immediate Action Required form via fax at 1-800-777-0051 or email QAGCO@abbott.com. IMPORTANT NOTE: If you do not currently have replacement material than you may continue to use this reagent lot until your laboratory receives replacement material. Upon receipt of the replacement material, discontinue use of and discard lot 86599M500 and switch to the replacement material. Contact Abbott Customer service for replacement material (1-877-4ABBOTT, Option 1).If you or any of the health care providers you serve have any questions regarding this information, U.S. customers should call Customers Service at 1-877-4ABBOTT. Customers outside of the U.S., please contact your local Customer Service representative. |  | 1995 | 1,995 kits | Worldwide distribution: USA and countries including Puerto Rico, Argentina, Australia, Austria, Belarus, Belgium, Bosnia, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Egypt, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Lithuania, Montenegro, Netherlands, New Zealand, Norway, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom and Vietnam. | Terminated | May 13, 2011 | 248 | 1995
Z-0070-2011 | 56699 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Buretrol Solution Set, 83" (2.1 m), 150 mL Burette Drip Chamber Filter Valve, 2 Injection Sites, Male Luer Lock Adapter, Large Bore 4-Way Stopcock with Rotating Male Luer Lock Extension Set, 38" (97 cm), Vol. 5.8 mL, Injection Site, Male Luer Lock Adapter, 60 DPM; a sterile Rx IV fluid pathway device; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 3C0166 |  |  | September 03, 2010 | 2010 | Baxter Healthcare Corp. | There are leaks in the Drip Chamber Check Valve Subassembly of the Solution Set. A leak in a sterile pathway could allow microbial entry, which could lead to a blood stream infection. | Employee error | Baxter sent an "Urgent Product Recall" letter dated September 2, 2010, via first class mail to the direct account, to the attention of the Director of Materials Management, with a copy sent to the Director of Nursing, on the same date, informing them that leaks were detected in the Drip Chamber Check Valve Subassembly used in the affected lot of Interlink System Buretrol Solution Sets. A leak in a sterile pathway could allow microbial entry, which could lead to a blood stream infection. The account was requested to examine their inventory in all locations to determine if they have any of the affected lot and remove the affected product from inventory; call Baxter Healthcare Center for Service at 1-888-229-0001 to return the affected product, request a credit, or request replacement product; complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 948 | 948 units | Nationwide distribution: Wisconsin only. | Terminated | February 25, 2011 | 175 | 948
Z-0069-2011 | 56278 | No | LNX |  |  | General Hospital | 510(k) | Hospital Information System Results Reporting Interface Versions 4.0.2.38.1, 4.0.2.39.2, 4.02.40.1, and 4.0.2.41.1. |  |  | October 05, 2007 | 2007 | SCC Soft Computer | SCC Soft Computer, Clearwater, FL issued a Correction for Hospital Information System Results Reporting Interface Versions 4.0.2.38.1, 4.0.2.39.2, 4.0.2.40.1, and 4.0.2.41.1 on 10/05/2007. A software anomaly in the result reporting interface caused result comments entered after an empty line | Software design | SCC Soft Computer sent a RISK-TO-HEALTH NOTIFICATION TASK VERBIAGE dated October 5, 2007, through their Task Management System to all affected clients on October 17, 2007. The notice identified the product, the problem and the action clients should take to correct the problem. Clients were to acknowledge receipt of the task notice and to identify if the Workflow or functionality is not used and if they Agree to use alternative working solutions. Each client was asked to grant permission to load the required software to correct the issue.For any questions regarding this recall call (727) 789-0100. |  | 35 | 35 | Worldwide Distribution - USA, including AL, CO, FL, GA, IN, KS, KY, MA, MO, NC, NH, NJ, NY, OH, PA, TN, and WA and the country of Canada | Terminated | October 19, 2010 | 1110 | 35
Z-0066-2011 | 56756 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Control Syringe, 12ml, CAT No: CCS201E Version B, Sterile R, Merit Medical Systems, Inc. |  |  | September 09, 2010 | 2010 | Merit Medical Systems, Inc. | Small holes in the packaging may render the product non-sterile. | Employee error | Merit Medical sent an Urgent Product Recall Notice dated September 9, 2010, to all sales representatives. The notice identified the product, the problem, and the action to be taken by the sales representatives and their customers. Sales representatives were instructed to contact all customers via telephone and instruct them to immediately discontinue use and quarantine any unused product. Recall Notification Forms were to be completed and signed by the site representative and the customer and faxed to Merit customer service at 801-208-3378. If customers were returning affected products they were to place the completed original Recall Notification Form in the box with the products to be returned, write the RGA# on the shipping box and return by Federal Express 2nd Day air to: RGA DepartmentMerit Medical Systems, Inc.1600 West Merit ParkwaySouth Jordan, UT 84095-2416For any questions regarding this recall call 801-316-4822 or 801-208-4365. |  | 199 | 199 units | Worldwide Distribution - USA, including IL, PA, WI and WV and the countries China, Greece, Hong Kong, Japan, New Zealand and Spain. | Terminated | March 21, 2011 | 193 | 16574
Z-0067-2011 | 56275 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI Releases 3.17 through 4.1.0. |  |  | March 28, 2007 | 2007 | SCC Soft Computer | SCC Soft Computer, Clearwater, FL issued a Correction for SoftPath GUI Releases 3.17 and 4.1 in March 2007 for clients with high volume output of final reports who use the Batch Print for Final Reports in the Scheduler, there is a potential for some reports not printing. This issue was corrected for the one affected user in October 2003. | Software design | SCC Soft Computer sent a SAFETY NOTIFICATION TASK VERBIAGE dated January 5, 2004, to one client. The Notice identified the product, the problem, and the action to be taken by the client. SCC recommended the client set the Search Date Range to a small enough range to assure that the number of cases scanned during the job would be less than 10,000.The client was asked to acknowledge receipt of the task and to indicate if the Workflow or functionality was used/not used and Agree to use the alternative working solution. For any questions regarding this recall call (727) 780-0100, ext. 4235. |  | 84 | 84 | Nationwide - USA, including the states of CA, CO, CT, FL, GA, IL, IN, LA, MD, MI, MO, MS, NC, NH, NJ, NV, NY, OH, OK, PA, SD, TX, WI, and WV | Terminated | October 19, 2010 | 1301 | 84
Z-0075-2011 | 54550 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Pathfinder End Screw Extender Sleeve, REF 1165-10, Rx, non-sterile, sold under label Abbott Spine, Inc. Austin TX (prior to 2008) and Zimmer Spine Austin TX (7/2009 to present) |  |  | April 13, 2010 | 2010 | Zimmer Inc. | The tip of the PathFinder End Screw Extender Sleeve has fractured during use. In some cases, this instrument fracture led to surgical delay or additional surgical steps necessary to remove the fragments of the broken tip. | Device Design | Zimmer Spine issued an "Urgent Medical DeviceCorrection" letter through the Zimmer Spine sales force to be provided to PathFinder system users in February 2010. the letter identified the affected product, safety issue, and actions to be taken by customers. Customers will receive replacement devices with a design change for the recalled devices, as they are available from the manufacturer. |  | 1471 | 1471 (888 subject to removal) | Worldwide Distribution: USA including states of AZ, CA, CO, FL, GA, LA, MI, MN, MO, NC, NH, NY, OH, OK, PA, TN, TX, WA, and WI, and the countries of Australia, Belgium, Brazil, Canada, China, Colombia, Cyprus, France, Germany, Greece, Italy, Japan, Mexico, Thailand, Turkey, UK and Switzerland. | Terminated | February 10, 2012 | 668 | 1471
Z-0078-2011 | 55150 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyRA Cuvette Segment 1, P/N 007000-001 |  |  | March 29, 2010 | 2010 | Medica Corporation | Chemistry analyzers may report Calcium results below the sample's actual concentration. Reported results may be ten to fifteen percent lower than expected sample results.The defect is a shortened path length at specific cuvette positions of those segments manufactured using molding cavity 1. The shortened path length was determined by measuring the cuvette dimensions with a Co-ordinate Measuri | Other | Medica issued an Urgent Product Recall letter dated March 29, 2010 to customers. Users were requested to cease using and discard the affected product, and complete and return a Response Form in order to obtain replacement product.Customers can contact Medica Technical Support at 877-777-5895. |  | 14926 | 14,926 | Worldwide Distribution: USA, and the countries of Bulgaria, Canada, India, Mexico, Panama, Phillipines, Russia, Turkey, New Zealand, South Africa, United Arab Emirates, and Venezuela. | Terminated | June 07, 2011 | 435 | 14926
Z-0079-2011 | 56805 | No | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | R3 (TM) 22 MM I.D., 52 MM O.D., CONSTRAINED LINER, UHMWPE, CO-CR, TI-6AL-4V, REF 71339152, STERILE EO, QTY 1, Smith &amp; Nephew, Inc., Hip prosthesis acetabular component. |  |  | August 13, 2010 | 2010 | Smith &amp; Nephew Inc | The firm received reports of a femoral head popping out of the liner intraoperatively. | Other | Smith &amp; Nephew notified all distributors via e-mail and sent a letter dated August 13, 2010 to all hospitals notifying them of the product, the problem, and the action to be taken. 1) Distributors were to immediately cease distribution and use of the product, examine their stock and the stock of the accounts they supplied if applicable and remove any affected devices from all inventories and place them into quarantine to prevent accidental distribution. 2) Identify customers that received these products and immediately notify them of the recall.3) Carry out a physical count and record the data on the Verification Section included in the letter.4) Contact their inventory specialist for an RA# and record this on the Verification Section.5) Return the Verfication Section even if they do not have the recalled product.6) Return the recalled product to: Smith &amp; Nephew, Inc. Attn: Return Goods 3303 Holmes Road Memphis, TN 38118Smith &amp; Nephew, Inc. sent a second notice via e-mail on 09/03/2010 to non-responding distributors and by Fed Ex to the hospitals 09/21/2010.For any questions regarding this recall customers were to call (901) 399-5441. |  | 118 | 118 units | Worlwide Distribution, USA including AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, MA, MI, MN, MO, NC, NJ, NY, OH, OK, PA, TN, TX, WA, and WI and the countries of Australia and Portugal | Terminated | September 21, 2011 | 404 | 716
Z-0074-2011 | 56274 | No | LNX |  |  | General Hospital | 510(k) | Soft Query Results Reporting Versions 2.2.0.6.3 and 2.2.0.7.1. |  |  | November 30, 2004 | 2004 | SCC Soft Computer | Soft Query Results Reporting Versions 2.2.0.6.3 and 2.2.0.7.1 HIS may display incorrect collection data for a procedure or test result. Clients using HBO Star RR interface to query for results on a canceled test may see the previous occurrence of the test with the current date and time displayed as the collection date and time instead of the correct collection date and time sent by the SCC interf | Software design | Client Notification Task Verbiage was sent through the Task Management System on December 3, 2003, identifying the software issue, the clinical impact, and corrective action. Customers were informed that the HIS may display incorrect collection data. Customers received the SQRR Service Pack with revised software, which will will not produce inaccurate information if there are no verified results. Customers were to acknowledge receipt of the task notification and indicate whether they would like the corrected software.SCC Soft Computer proprietary Task Management System is used to manage client communication. Clients use this system to request services, receive and provide software updates for any issue.Customers can contact SCC Soft Computer at 727-789-0100. |  | 5 | 5 | Nationwide Distribution: including states of GA, LA, and NH. | Terminated | October 19, 2010 | 2149 | 5
Z-0073-2011 | 56789 | No | EZQ | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | ActiveLife One-Piece Drainable Pouch with Stomahesive Skin Barrier; |  |  | September 17, 2010 | 2010 | Convatec Inc. | Mislabeled: the market units were incorrectly labeled as # 125339 ActiveLife Drainable Pouch 25 mm; the product inside was actually ActiveLife Pouch Small Tan 32 mm, ICC # 125333. | Labeling mix-ups | ConvaTec, Inc. sent a Returned Goods Authorization form by e-mail dated September 3, 2010, to one customer. The form identified the product, the problem, and the action the customer should take. The customer was instructed to detach the label and affix it to the outside of the box, copy the RGA # and retain a copy for their records, enclose the completed form and return with the product, For questions call (908) 904-2675. |  | 1 | 1 case (20 market units) | USA, including IL | Terminated | October 19, 2010 | 32 | 1
Z-0076-2011 | 56736 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Sequencer Verify &amp; Record System, for use with radiotherapy treatment machines, Class IIb, manufactured by Impac Medical Systems, Sunnyvale, CA. |  |  | September 02, 2010 | 2010 | Impac Medical Systems Inc | Software issue-- matching criteria can be configured that permits an invalid patient match, resulting in overwriting of patient records with information from a different patient. | Software design | Elekta Impac Software issued a Important Safety Notice letter dated September 2010 to all customers identifying the software issue and clinical impact. Customer Support representatives will contact each site and modify site configuration or upgrade existing software to the version containing the corrective fix. Customers can contact Impac Software at 800 488-4672 |  | 797 | 797 units | Worldwide Distribution: throughout the US and the countries of Australia, Austria, Belgium, Canada, China, Hong Kong, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Ireland, Israel, Italy, Japan, Jordan, Morocco, Netherlands, New Zealand, Poland, Portugal, Singapore, Spain, Sweden, Switzerland, and UK. | Terminated | June 24, 2011 | 295 | 797
Z-0077-2011 | 56729 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | DDC-6-AA Data Card, an accessory to the Sentry DDU-1 00 semiautomatic external defibrillator (AED). |  |  | August 27, 2010 | 2010 | Defibtech, LLC | Data cards may include an incorrect format which may cause the AED to not function correctly. | Mixed-up of materials/components | Defibtech Initially notified customers by telephone on 8/27/10. Customers were notified of the product, the problem and the action to be taken. Customers were instructed to locate affected data cards and send them back to Defibtech for evaluation. If the data card was inserted into the AED, then the customer was instructed to check the AED software version. If the software version was at 2.005, the customer was instructed to keep the AED out of service until Defibtech could evaluate the returned data card. (Any other software version indicated that the device was not affected by a faulty data card). If the software version was any other version other than 2.005, then the customer was instructed to keep the AED in service unless the original service inquiry required it to remain out of service.If the data card was not inserted into the AED, then the customer was instructed to return the data card for evaluation.For any questions customers were to contact Customer Service at (866) 333-4241 from 8:30-5pm EST, M-F. |  | 38 | 38 | USA - Nationwide Distribution and Puerto Rico | Terminated | October 05, 2011 | 404 | 38
Z-0088-2011 | 56851 | No | LNX |  |  | General Hospital | 510(k) | SoftMic Versions 4.0.4.7, 4.5.1.3. |  |  | July 19, 2010 | 2010 | SCC Soft Computer | Softmic system may delay or omit reporting of clinically significant results including organisms and drug sensitivities. Incomplete culture results may appear complete. SCC Soft Computer is issuing a Correction for SoftMic Versions 4.0.4.7-4.5.1.3 software due to this potential hazard. | Software design | SCC Soft Computer issued Risk-to-Health notifications to clients on July 19, 2010 via proprietary communication software. Corrective actions will be via programming correction via hot fix (software modification) or such so that Global variable is not used. Clients are asked to acknowledge receipt of this task and update the task with permission to load the required correction. |  | 35 | 35 | Worldwide Distribution: USA, including the states of CA, CT, DE, FL, IA, KY, MA, MA, MI, MN, NJ, NY, OH, PA, TX, VA, and WI, and the country of Canada. | Terminated | July 16, 2012 | 728 | 35
Z-0090-2011 | 56280 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI version 3.17.6.3.1. |  |  | March 11, 2004 | 2004 | SCC Soft Computer | For clients using the SoftPath GUI version 3.17.6.3.1 Results Reporting interface with parsed text and having the auto correct and auto formatting functions active in MS Word, there is a possibility for some ANSI characters/symbols to be translated incorrectly to ASCII. SCC Soft Computer initiated a Correction in March 2004 for all affected clients. | Software design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System on March 11, 2007. SCC recommended turning off the autocorrect and autoformatting functions in MS Word if custom configuration parses text to an ASCII file (HIS) for the Result Reporting interface. If this function was turned on at the time of creating the canned message in the setup file, edit and re-save all canned messages in SoftPath. Client was asked to acknowledge receipt of the task and grant permission to load the required software correction. |  | 1 | 1 | Nationwide distribution: One hospital in NY. | Terminated | November 03, 2010 | 2428 | 1
Z-0087-2011 | 56502 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Manta Ray: Anterior Cervical Plate (ACP) System, 510 (k) #K080690. The labeling for the plate component reads in part: 'REF: 22-'**** MAT'L: Titanium per ASTM F-136, Lot: W****, Description: Plate - 40 mm, Qty: 1.' The labeling for the screw component reads in part: 'REF: 22-**** MAT'L: Titanium per ASTM F-136 LOT W**** Description: Screw-Fixed Tapping 4x10mm QTY: 1. |  |  | July 29, 2010 | 2010 | Theken Spine Llc | The Manta Ray Cervical System was subject to corrective action by the firm after two (2) customer complaints were received alleging that screws were 'backing out' of the device during surgical procedures. | Other | Integra sent a Technical Bulletin dated July 29, 2010 via Fedex to their customers. The bulletin identified the product, the problem, and action to be taken by the user/customers (surgeons) of the two (2) reported instances of screws backing out of the plates. The Bulletin provided instructions outlining the proper surgical techniques for implantation of screws using the Manta Ray ACP System. The bulletin also included an attached acknowledgement form which the user/customers were asked to complete and FAX to (609) 275-5363 or E-mail back to the E-mail address provided. The form verified that each customer: 1) received the technical bulletin; and 2) has read and understood the proper technique for surgical placement of the screws as outlined in the bulletin.For technical support customers were to contact the Customer Service Department at (866) 942-8698. |  | 1081 | 1081- includes totals distributed for both plates &amp; screws | Nationwide Distribution - USA including AL, AZ, CA, CT, FL, GA, HI, IL, ID, IN, KS, LA, MI, MD, MN, MA, MD, MS, NE, NY, NC, NJ, NV, NM, OH, OR, OK, PA, TN, TX, UT, WI, and WA.WA. | Terminated | September 21, 2012 | 785 | 1081
Z-0091-2011 | 56808 | No | FHN | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Ten Shooter Saeed Multi-Band Ligator, Rx Only, Wilson Cook Medical 4900 Bethania Station Road, Winston-Salem, NC 27105 |  |  | September 03, 2010 | 2010 | Cook Endoscopy | A section of the ligator barrel may separate from the barrels friction fit adapter that attaches to the patent end of the endoscope. This could result in a section of the ligator barrel detaching inside the patients gastrointestinal tract. | Other | COOK Endoscopy sent an URGENT: PRODUCT RECALL dated September 3, 2010, via USPS Certified Mail and by email to all consignees. The letter identified the products, the problem, and the action to be taken by the consignees. All consignees were instructed to quarantine any affected product and return them to Cook Medical. Distributors were instructed to contact their customers and coordinate the return of the affected product to Cook Medical. Within 2 weeks of the initial recall notification, Level B effectiveness checks (40% of the total number of consignees) will be conducted. These customers will receive a recall effectiveness survey. All non-responding consignees will receive a total of 3 recall notifications (2 follow-up requests for a response in addition to the initial recall notification. Monthly status reports regarding this recall will be provided to FDA. All devices returned as part of this recall will be destroyed by the recalling firm upon receipt of authorization to do so by FDA. Forty percent (40%) of total customers affected received a second notice on 9/17/2010 as part of the effectiveness checks. As of 9/21/2010, 38% of all affected customers have responded.For questions contact the Customer Relations Department at CustomerRelationsNA@cookmedical.com or call 1-800-457-4500, press 4 then enter extension 2146. |  | 2011 | 2011 units (total including both 6 &amp; 10 Shooters) | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, WA, WI, WV, and WY; and the countries of Algeria, Australia, Austria, Belgium, Canada, Chile, Cyprus, Denmark, France, Georgia, Germany, Greece, Hong Kong, Ireland, Israel, Italy, Jordan, Korea, Mexico, Malaysia, Morocco, Netherlands, New Zealand, Peru, Poland, Portugal, South Africa, Slovakia, Spain, Switzerland, United Arab Emirates, UK and Venezuela. | Terminated | March 31, 2011 | 209 | 2012
Z-0089-2011 | 56277 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI Versions 3.17 and 4.1. |  |  | December 30, 2003 | 2003 | SCC Soft Computer | SoftPath Cases with Word Perfect documents that were not signed out in ASCII are printing in GUI as a converted document from Word Perfect to PDF document instead of the current MS Word Document. This affects clients upgrading from ASCII to SoftPath GUI 3.17.7.6 Version. All affected clients have been notified and corrections have been made. | Software design | SCC Soft Computer issued Risk-to-Health notifications with Client Notification Verbiage to affected clients using SCC Soft Computer's proprietary Task Management System. beginning on December 30, 2003. Clients were instructed to sign out all Word Perfect documents before upgrading to GUI. The correction for this anomaly is to upgrade to the latest release of version 3.17.7 or 3.17.8. Clients were asked to acknowledge receipt of this Task and indicate whether they would like to take the correction at this time. When the client decides to accept the correction a new task will be created to initiate the patch using the Task Management System. |  | 5 | 5 | Nationwide Distribution to the states of IN, NC, NH, and NJ. | Terminated | November 03, 2010 | 2500 | 5
Z-0107-2011 | 55195 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active I-PTH Enzyme-Linked Immunosorbent Assay(ELISA), Part Number: DSL-10-8000 |  |  | March 04, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed that the stated stability of the DSL-10-8000 I-PTH ELISA assay maynot be achieved for lot numbers: 891198, 990980 and 991448.The kit lot numbers listed above passed the stability testing performed at 6 months, but failed the testing performed at 12 months. Therefore the affected I-PTH ELISA kit lots may produce falsely elevated patient r | Other | Beckman Coulter sent an URGENT: PRODUCT CORRECTIVE ACTION letter dated March 17, 2010 with attached Customer Response form to all customers who received Active I-PTH ELISA kits with the affected lot numbers. The letter identified the product, the problem, and the actions to be taken.Customers were instructed to:(1) Discontinue use of all materials provided with I-PTH ELISA kit lot numbers 891198, 990980 and 991448;(2) Review historical external quality control results to ensure the assay performance of these lots was within laboratory specifications. Retesting may be necessary if your laboratory uses only the control materials provided with the I-PTH ELISA kit;(3) Abnormally elevated patient results that are obtained with these lot numbers and that are inconsistent with the patient's clinical status should be reviewed in the context of other clinicaldiagnostic tests; (4) I-PTH ELISA kit lot numbers 891198, 990980 and 991448 will no longer be shipped.Alternate lots will be provided for current and future replenishment orders or requests.in addition customers were instructed to share this information with your laboratory staff and retain this notification as part of yourlaboratory Quality System documentation. Complete and return the enclosed response form within 10 days.For questions regarding this recall contact (800) 231-7970. |  | 70 | 70 kits | Nationwide Distribution including the states of CA, CT, IA, IL, MI, NY, SC, and TX | Terminated | April 27, 2012 | 785 | 70
Z-0099-2011 | 56282 | No | LNX |  |  | General Hospital | 510(k) | SoftWeb version 4.0.4.10.1. |  |  | June 08, 2006 | 2006 | SCC Soft Computer | If a case number is re-used in SoftWeb with SoftPath, the SoftPath case is linked to the original MRN in SoftWeb. This causes SoftPath cases being linked to the wrong patient in SoftWeb. Incorrect results on the patient with original MRN and no results on the patient with the second MRN (and the reused case number) are displayed in SoftWeb Results Query. SCC Soft Computer issued a Correction for | Software design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System on June 8, 2006. SCC recommended SoftPath/SoftWeb users NOT re-use SoftPath case numbers when cases are cancelled. The generic option setting "Specimen Registration - Standard Settings - Minimum Time to Reactivate Case Number" should be set to "0" so the user is not prompted to reuse the same case number. Client was asked to acknowledge receipt of the task and grant permission to load the required software correction. |  | 1 | 1 | Nationwide Distribution: 1 hospital in Alaska. | Terminated | November 03, 2010 | 1609 | 1
Z-0098-2011 | 56687 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump2 blockAid 400 mL Painpump for continuous Nerve Block, REF 575-100, Sterile, Rx only, Stryker Instruments, Kalamazoo, MI. |  |  | August 16, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | The instructions for use (IFU) has been revised for health care professionals. The indication for use has been changed to remove site specific pain management. The firm has revised the warning to not use the pump around joint spaces. A warning has been added regarding the use of the pump with anticoagulants in epidural applications. A special reminder to users to always evaluate the potential | Other | Stryker Instruments sent Urgent Medical Device Recall Notification letters dated 8/16/2010 to healthcare professionals. The recall letters outlined the changes for the instructions for the use of these products. The letters informed the healthcare professional to review the new IFU, to distribute the new IFU to all healthcare professionals that use the device, and to complete and return the business reply form. The letters included a complete IFU that was to be used with any remaining stock of devices. Customers should dispose of the old version of the IFU. There is no product being returned for this recall. |  | 11976 | 11,976 all products | US, Canada and Australia. | Terminated | April 21, 2011 | 248 | 11976
Z-0103-2011 | 56675 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | LIBERTY Automated Peritoneal Dialysis System Line Cycler Set with stay safe connectors (Dual Patient Connector) PIN: 050-87212, manufactured for Fresenius Medical Care, Waltham MA by Erika de Reynoza, Reynosa Tampa, Mexico. |  |  | September 03, 2010 | 2010 | Fresenius Medical Care Holdings, Inc. | Cycle set cassettes may leak due to holes/scratches in the film, contaminating dialysate and potentially resulting in peritonitis. | Nonconforming Material/Component | Fresenius Medical Care issued an Urgent FMCNA LIBERTY Cycler Set Recall letter on September 3, 2010 identifying the specific affected part numbers and product lots and requesting that customers identify and discontinue the use of the affected lots immediately. Customers were requested to return a receipt of notification. Customers may arrange return unused product to FMCNA by contacting customer service at 1-800-323-5188. |  | 363265 | 363,265 units | Nationwide Distribution: Throughout USA. | Completed |  |  | 370588
Z-0108-2011 | 56802 | No | JEY | Bone plate. | Dental | Dental | 510(k) | 2.0/1.0 4 hole long plate (titanium plate ).Lot 120150. |  |  | August 31, 2010 | 2010 | Biomet Microfixation, Inc. | Mislabeled size: The 2.0/1.0 4 hole long plate, Lot 120150 is labeled as 01-9205 (2.0/1.0 4 Hole Long Plate), however, the package contains a 01-9204 (2.0/1.0 4 Hole Regular Plate). | Other | Biomet Microfixation sent URGENT MEDICAL DEVICE RECALL NOTICES dated August 31, 2010 to its customers and distributors, identifying the affected product and actions to be taken by customers. Customers were requested to check their inventory for product and return any affected product to Biomet Microfixation. Customer questions or concerns should be directed to Laura Sabo, VP Global QA/RA at 904-741-4500. |  | 56 | 56 | Worldwide Distribution: USA including states of Hawaii and Kansas, and the countries of Australia, Denmark, Italy, Jordan, and South Africa. | Terminated | November 19, 2010 | 80 | 56
Z-0093-2011 | 56279 | No | HBA | Neurosurgical paddie. | Neurology | Neurology | 510(k) | Codman 5/16" Round Surgical Patties |  |  | July 16, 2010 | 2010 | Codman &amp; Shurtleff, Inc. | Surgical Pattie x-ray stripe (radio-opaque marker) became detached | Process control | Codman notified customers by letter dated 7/16/10 issued July 23 through July 26, 2010 via Fed'x overnight advising users that the x-ray marker on Codman Surgical Patties may become detached. Customers are instructed to return product from affected lots. An error in the original letter was discovered and a corrected letter dated July 28, 2010 issued. If there are any questions, please contact Codman at (800) 225-0460. |  | 6730 | 673 boxes (20 pouches of 10 patties per box) | NationwideForeign: Canada, Singapore, Belgium, Mexico, Madrid, Dubai, Tokyo, Brazil, New Zealand, Egypt, South Africa, Singapore, | Terminated | July 09, 2012 | 724 | 53090
Z-0102-2011 | 56272 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI version 4.2.3.4 2008-06-30 (Manuf Date). |  |  | January 26, 2009 | 2009 | SCC Soft Computer | For SoftPath GUI versions 4.2.3.4, 4.2.4.1, 403.0.3, and 4.3.0.4., the Previous Patient History and some QA Management reports may have missing or incorrect diagnosis text. However, the patient final report, HIS resulting and Query display are all correct. All affected clients have been provided corrected software.. | Software design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to affected clients using SCC Soft Computer's proprietary Task Management System on January 26, 2009. SCC issued a software upgrade with corrective patch electronically to all clients. Client was asked to acknowledge receipt of the task and grant permission to load the required software.SCC will contact each client to coordinate the running of the utility and the loading and validation of the correct patch. |  | 52 | 52 | Worldwide Distribution: Throughout USA and Canada. | Terminated | November 03, 2010 | 646 | 52
Z-0106-2011 | 56346 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand TrueBeam and TrueBeam STx, |  |  | May 24, 2010 | 2010 | Varian Medical Systems Oncology Systems | Due to a software anomaly, guidance-based couch shift values may not be applied as expected when the operator selects "Apply Shift" and presses Motion Enable buttons. This can result in the actual patient position differing from that position indicated by the imaging application affecting poistion accuracy and precision. | Software design | Varian Medical Systems issued an Urgent Medical Device Correction letter date dated May 24, 2010 to users identifying the affected device, the issue prompting the correction, and the actions to be taken by the user.Customers were asked to set couch tolerance values to small numbers, and for radiosurgical treatment tolerance values to be set to zero. A Varian representative will contact customers to schedule installation of a software version correcting the issue. US Customers can contact Varian at 1-888-827-4265,. European customers can contact Varian at +41 41 749 8844. |  |  |  | Worldwide Distribution: Two consignees, one in the US and one in Switzerland. | Terminated | July 20, 2011 | 422
Z-0101-2011 | 56682 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | NOVAPLUS (TM), SUCTION CANISTER, 1200cc, RIGID, WITH ELBOW, Cat. No. V71-3001, Rx only, NON-STERILE, Made exclusively for Novation by DeRoyal, 200 DeBusk Lane, Powell, TN 37849 |  |  | July 01, 2010 | 2010 | DeRoyal Industries Inc | Suction canisters may crack during use | Other | DeRoyal sent a Recall letter dated July 1, 2010 to all their consignees via UPS Air. The letter identified the product, the problem, and action to be taken by the consignees. Consignees were to complete the Recall Response Form by checking their inventory for DeRoyal Suction Canisters V71-3001, Lot #22622679.Indicate the number of replacement products needed on the Recall Response Form, including the address, reference PO#, etc. to be used for the replacement products and fax the form to DeRoyal at 865-362-3755 no later than July 7, 2010.Recalled products were to be returned to DeRoyal via UPS Ground Attn: Quality Control 300 DeBusk Lane Powell, TN 37849Once the recalled product was received by DeRoyal, the replacement product would be shipped per the instructions on the Recall Response Form.Distributors were to forward the recall notice to their end-users, or provide customer listing so that DeRoyal may contact them.For any questions or assistance regarding this recall call 865-362-1037. |  |  |  | Nationwide | Terminated | April 10, 2012 | 649
Z-0035-2011 | 56685 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Excelsior Disposable Syringe W/Normal Saline (0.9% sodium Chloride) |  |  | September 02, 2010 | 2010 | Excelsior Medical Corp | There is the potential for a dimensional issue in 6 ml syringe products, codes E0100-50, 10056-1000, 10056-240, 14056-240, 910056-1000, S5, to cause leakage and.or loss of sterility | Device Design | Excelsior Medical sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 2, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to discontinue using syringes from the listed lots and return the attached Customer Reply Form via fax (866) 688-3185 as confirmation that notification had been received and that proper arrangements can be made to have the product returned to the firm. The customers were also instructed to inform their customers, if they have resold products, of the recall and instruct them to return the form, and verify that they have returned all unused syringes from any of the listed lots . NOTE: The customers were informed that they may alternatively have the products destroyed and certified to Excelsior that they have done so. They must include a copy of the certificate of destruction with the attached Customer Reply Form via fax. - Excelsior issued a press release on October 14, 2010. The press release advised consumers of the recall and requested that use of the product be immediately discontinued and returned to the point of purchase.Should you have any questions regarding the recalled product, please contact Excelsior Medical Customer Service at (800) 487-4276 or if you prefer by fax (866) 688-3185. You may also contact us at recall10@excelsiormedical.com. |  | 13734640 | 13,734,640 | Worldwide distribution: USA and countries including Canada and Brazil. | Terminated | August 23, 2012 | 721 | 13734640
Z-0100-2011 | 56283 | No | LNX |  |  | General Hospital | 510(k) | SoftLab version 4.01 ASCII, ASCII Base 3.1.3.0-3.1.3.32, 3.1.4-3.1.4.27, 3.1.6-3.1.6.4, 4.0.0.0-4.0.0.9; LAB/MIC Add-on 4.0.1, 4.0.1.0-4.0.1.13, 4.0.2.0 in 2006. |  |  | August 04, 2006 | 2006 | SCC Soft Computer | For SoftLab version 4.01 ASCII, the qualifying test result does not reach the Call List because the order is locked by either another application, or two Hosparams (HISRR_RBS_sync and Auto_sync_RBS). This creates a lock on the RBS server. The problem was identified in the RBS "C" rule which sends qualifying results to the Call List for SoftLab. All affected clients received the upgrade for Soft | Software design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System in June 2006. SCC recommended SoftLab version 4.01 ASCII users change program settings to prevent locking of the RBS server. A software upgrade with corrective patch will be delivered electronically to all clients. Client was asked to acknowledge receipt of the task and grant permission to load the required software. |  | 131 | 131 | Nationwide Distribution: USA in the states of AL, AR, CA, CT, FL, GA, IA, IL, IN, KS, KY, LA, MD, ME, MI, MO, MS, NC, ND, NH, NJ, NY, OH, OK, PA, SD, TN, TX, VA, WA, WI, and WV. | Terminated | July 16, 2012 | 2173 | 131
Z-0110-2011 | 55922 | Yes | NIK |  |  | Cardiovascular | PMA | Medtronic Concerto II CRT-D, model D274TRK and Model D294TRK not available in the US. Digital implantable cardioverter defibrillator with cardiac resynchronization therapy (DDE-DDR). Sterilized using ethylene oxide. Medtronic, Inc., Minneapolis, MN 55432 USA. Manufactured in: Juncos, Puerto Rico, USA. |  |  | May 05, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic is informing you of a rare device software issue in certain ICD and CRT-D devices. A software update will be available to correct his issue. There have been 5 reports of this issue occurring with no reported patient injuries or deaths. Medtronic has identified the root cause to be a rare and specific sequence of events that must occur within a few milliseconds of each other. In the u | Software design | Medtronic, Inc. sent an IMPORTANT MEDICAL DEVICE CORRECTION LETTER, dated May 2010 to physicians beginning May 05, 2010. The letter informed doctors of the issue and informed them that a software update would be available to correct this issue. The software was to be installed on the programmer, with a Medtronic representatives assistance, so that on a patient's next in-clinic interrogation the device would receive the upgrade automatically and their device would be corrected.For any questions regarding this recall physicians were to contact their local Medtronic Representative or Medtronic Technical Services at 800-723-4636. |  | 18778 | 18,778 | Worldwide Distribution - USA, Puerto Rico, Virgin Islands, and the countries of Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Brunei Darussalam, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Saint Lucia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vatican City, Venezuela, and Viet Nam. | Terminated | March 15, 2012 | 680 | 210237
Z-0121-2011 | 56484 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatics Systems QCFIRST Custom Rule |  |  | July 23, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | the EasyLink QCFIRST custom rule does not operate as intended. | Other | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice dated July 2010 to all customers that have the EasyLink QCFIRST custom rule installed. This Urgent Field Safety Notice included information on the problem and confirmed that a Siemens representative would contact them to facilitate the removal of the rule from the EasyLink System.For questions contact the Siemens Healthcare Diagnostics Technical Solution Center at 1-800-441-9250. |  | 82 | 82 | Worldwide Distribution - USA including AZ, CA, FL, GA, IA, IL, IN, KS, KY, LA, MO, MS, NC, NH, NJ, NM, NY, OH, OR, PA, TX, VA, WA, WI, and WY and the country of Canada. | Terminated | March 19, 2013 | 970 | 82
Z-0120-2011 | 56701 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂż Complete Knee Solution Stemmed Tibial |  |  | January 26, 2010 | 2010 | Zimmer Inc. | Zimmer had determined that two lots of these tibial implants have incorrectly positioned or missing flange plugs that were not seated in the device upon receipt to the customer. A missing flange plug could lead to a delay of surgery or possibility of wear leading to more surgery. | Employee error | The firm sent an URGENT MARKET WITHDRAWAL to its sales force and distributors on 1/26/2010. The letter identifies the affected product and explains the reason for the Market Withdrawal. The letter also explains the clinical implications and instructs consignees to locate any unused devices and quarantine them immediately. A physical count should be carried out of all affected devices and recorded on the Inventory Return Certification Form that's included with the letter. A copy of the letter should be faxed to Zimmer, Inc. and the affected product should be returned along with the form to the firm. |  | 61 | 61 units | Worldwide Distribution -- US, Canada, Singapore Australia, and India. | Terminated | October 23, 2010 | 270 | 61
Z-0109-2011 | 56762 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | in2it (I) A1c Test Cartridges (Catalog 281-0001EX) and in2it (II) A1c Test Cartridges (Catalog 281-0001), manufactured by Bio-Rad Laboratories Deeside Ltd., Flintshire, UK, distributed by Bio-Rad Laboratories, Hercules, CA. |  |  | September 07, 2010 | 2010 | Bio-Rad Laboratories Inc | Test Cartridges may leak and have the potential to generate error messages and biased inaccurate results. | Component design/selection | BIO-RAD Laboratories Deeside sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated September 7, 2010, via registered mail to all subsidiaries/customers. The letter described the product, problem and actions to be taken by the customer. The customers were instructed to stop using and discard the test cartridges from the listed lots; complete and return the attached CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM via fax +44 1244 833401; record the number of affected kits of product, confirm that they have discarded the cartridges and record the method of disposal.If you have any questions, contact +44 1244 288888. |  | 4056 | 4056 | Worldwide distribution: USA and countries including: Australia, New Zealand, Taiwan, France, the UK, Singapore, Germany, Italy, Thailand, China, Korea, Sweden, Sri Lanka, Maldives, Canada and the Philippines. | Terminated | May 17, 2011 | 252 | 4056
Z-0123-2011 | 56627 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory Instrument Interface software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. |  |  | August 27, 2010 | 2010 | Ge Healthcare It | Instrument may incorrectly report the result of an antimicrobial sensitivity test. | Software design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated August 27, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the patient safety issue associated with the use of GE Centricity Laboratory instrument interface to Biomerieux's Vitek analyzer where results may be incorrectly reported. The letters provided the users with specific safety instructions to follow until the software is updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. |  | 22 | 22 units | Worldwide Distribution -- USA, including states of Georgia, Michigan, Minnesota, New York, Ohio, Oregon, Tennessee, and Texas and countries of Australia, Canada, England, India, New Zealand, Qatar, and Scotland. | Terminated | July 15, 2011 | 322 | 22
Z-0133-2011 | 56508 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System MX 8000 IDT 16; Model # 728120; Serial #'s: 2513, 2518, 3147, 3217, 3014, 3050, 3068, 3188, 3091, 3051, 3099, 3126, 3183, 3303, 2539 &amp; 3056. |  |  | July 27, 2010 | 2010 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has decided to recall the Computed Tomography X-Ray System MX 8000 IDT due to the fact that the software upgrade for the MX 8000 IDT 16 v4.2.1 may not provide the optional collimation of 4x4.5 (18mm). If the customer does not verify the increment and thickness coincide with the newly selected collimation suggested by the conversion report, there is potential for a 6mm gap to ap | Software change control | Philips Medical Systems (Cleveland), Inc. sent an URGENT - Field Safety Notice dated July 29, 2010 and August 23, 2010 to all customers. The notices identified the product, the problem, and the steps to be taken.Customers were instructed to locate and correct the error. Field Service Engineers were to visit each site to update the softwareFor further information customers were to contact their local Philips Customer Care Center at 1-800-722-9377 |  | 16 | 16 units | Product was sold in the following states: CA, DE, IN,OR, MA, NV, TN. Product was also sold in Canada and Australia | Terminated | September 21, 2012 | 787 | 16
Z-0131-2011 | 56791 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | ARRAY and POLARIS Spinal System Medium Cross Connector |  |  | September 16, 2010 | 2010 | Ebi, Llc | One of the two connector ends of the medium cross connector may not be fully secured to the rod after tightening. This may not be readily apparent to the operating surgeon, resulting in inadequate torsional stability. | Other | Biomet Spine issued an Urgent Medical Device Recall letter dated September 16, 2010 to distributors and physicians, identifying the affected device and actions to be taken.Customers and distributors were requested to identify, remove, and immediately return the affected product to Biomet, following enclosed Product Return Instructions, and complete and return the Removal Tracking and Verification Form.Customers and distributors can contact Biomet at 1-800-526-2579. |  | 82 | 82 units | Worldwide Distribution: USA and the countries of Australia, Columbia, Great Britain, The Netherlands, New Zealand, and Switzerland. | Terminated | June 01, 2012 | 624 | 82
Z-0122-2011 | 56503 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | in2it (I) and (II) System A1c Test cartridges and in2it Self Test System A1c Test Cartridges, for use with the in2it (I) and (II) self test system analyzer, manufactured by Bio-Rad, Flintshire, UK |  |  | July 26, 2010 | 2010 | Bio-Rad Laboratories Inc | Inaccurate results with high bias-- Inaccurate results may be read by the machine, creating the risk for untreated hypoglycemia at elevated altitudes. | Device Design | BIO-RAD Laboratories Deeside Ltd., sent an URGENT: MEDICAL DEVICE CORRECTION letter, dated July 26, 2010, to their subsidiaries and to all customers identifying the product, the problem, and the action to be taken. Subsidiaries were to ensure that all customers received the letter with the Customer Medical Device Correction Response Form and collect the completed form from every customer. Once forms were received from all customers they were to be faxed to BIO-RAD Deeside Regulatory Affairs at +44 (0) 1244833401. Customers were instructed to run a test and visually inspect the test cartridge to detect any residual clear liquid. If the cartridge did not run properly the test was to be rejected.Customers were to run Quality Control test with each new shipment of test cartridges. Patient's results were not to be reported if the QC failed.Customers were to complete the CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM and fax it to the number provided on the form. For questions customers were to contact their regional BIO-RAD office. |  | 90703 | 90703 | Product was distributed to Bio-Rad's US subsidiary, along with 18 other distribution subsidiaries outside the US. These distributed further to other consignees. POroduct was distributed to Mexico, Korea, Singapore, Italy, South Africa, India, Australia, Spain, Greece, France, Thailand, Germany, New Zealand, China, Hong Kong, Portugal, Czech Republic, United Kingdom, Canada, Russia, Netherlands, Taiwan, Austria, Denmark, Israel, Finland and Belgium as well as to other African and Latin American countries. | Terminated | October 25, 2011 | 456 | 90703
Z-0124-2011 | 49488 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD10 Monoplane X-ray system. |  |  | August 21, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Due to a potential geometry software lock up during a Digital Rotational Angiography (DRA) run, the stand and table movement may no longer be possible. | Software design | Philips Medical Systems North America Co. sent an URGENT - Device Correction, dated September 15, 2010, to their consignees. The letter identified the product, the problem, and the action to be taken. The consignees were informed that there was a potential geometry software lock up during a Digital Rotational Angiography (DRA) run, the stand and table movement may no longer be possible.The consignees were advised that as geometry software errors happen intermittently, no actions can be performed by the consignees to minimize the risk except for not performing DRA procedures until the firm upgrades the software.For question consignees were to contact the Philips Call Center at 800-722-9377 #5, #4, #1 and reference FCO 72200111 or FCO 72200125. |  | 47 | 47 units | Devices were distributed to hospitals and medical facilities through out the U.S. | Terminated | November 19, 2010 | 820 | 73
Z-0132-2011 | 56610 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | PD Access Vascular Access Device. Escalon Vascular Access, Inc., 2440 South 179th Street, New Berlin, WI 53146: |  |  | May 05, 2010 | 2010 | Escalon Medical Corp | The packaging of the vascular access device may potentially become damaged, compromising the sterility of the product. Use of a non-sterile product increases risk of patient infection. This recall involes: Model 78050 18 Gauge PD Access Vascular Access Device, Model 78060 Extended 18 Gauge PD Access Vascular Access Device, Model 75010 18 Gauge SmartNeedle Vascular Access Devices with Expiration Da | Package design/selection | Escalon Vascular Access issued an "Urgent: Recall" letter dated May 5, 2010 to customers. The letter identified the affected product being recalled, described the problem, and the actions to be taken.. Consignees were recommended to immediately examine their inventory, quarantine the product, notify their customers, and return product affected. They also advised them that if they have further distributed the product, to please identify the customers and notify them at once of the product recall. Replacement product can be ordered from Vascular Solutions, Inc at 888-240-6001. Customer service can be contacted at 800-676-0043. |  | 188375 | 188,375 | Worldwide Distribution: USA, including the states of AL, AR, AK, AZ, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, ME, MD.MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, MN, NY, NC, ND, OH, OK OR, PA, RI, SC, SD,TN, TX UT, VA, WA, WV, and WI, and the countries of: Australia, Austria, Bahamas, Canada, Colombia, Denmark, England, Germany, Greece, Hong Kong, Italy, the Netherlands, New Zealand, Norway, Portugal, Spain, Sweden, and Switzerland. | Terminated | March 28, 2013 | 1058 | 188375
Z-0127-2011 | 56462 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira brand Symbiq One-Channel Infuser |  |  | February 22, 2010 | 2010 | Hospira Inc | Hospira has identified motor encoder failures in the Symbiq pumping mechanism that causes the infuser to cease operation during infusion therapy, resulting in delay or interruption of infusion therapy. | Device Design | Hospira issued a Recall Notification letter dated October 4, 2010, updating the Urgent Device Field Correction letter dated February 22, 2010, ideentifying the affected products, the issue prompting the recall, and actions to be taken by customers.Customers are to examine their inventory to determine if they have any of the affected products, and recommended to remove them from service in critical care areas, and contact Hospira to schedule removal and replacement of the motor. Replace affected devices with a loaner or corrected back-up pump if available.Custoemrs can contacvtd Hospira Global Product Safety and Complaints at 1-800-441-4100, and Technical Support Operations at 1-800-241-4002. |  |  |  | Worldwide Distribution: USA and the coutnries of Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Puerto Rico, Singapore and the UAE. | Terminated | December 03, 2013 | 1380
Z-0126-2011 | 56730 | No | DNN | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | EndoTool Drug Dose Calculator-Software Version 7.2.1800.3 and 7.2.1825.5 |  |  | July 26, 2010 | 2010 | Hospira, Inc. | The EndoTool software did not always provide an audible alert when a blood glucose measurement was due. This could result in a patient not receiving an appropriate dose of insulin. | Other | Hospira issued an Urgent Device Field Correction letter dated July 23, 2010 to consignees, providing details of the potential issue and the associated patient risk. Hospira is finalizing software enhancements which will resolve this situation within the next 90 days ,and will contact each consignee when it is available to be downloaded from the Hospira website. The notification letter also included a clinical bulletin, outlining mitigation instructions for clinicians to use until the software enhancement is available. Consignees were requested to complete and return an attached Reply Form by fax.Customers can contact the Hospira EndoTool support line at 1-877-349-4582 for clinical inquiries about this recall. |  | 97 | 97 units total (96 version 7.2.1800.3 and 1 version 7.2.1825.5) | Nationwide Distribution in states of CA, FL, IL, IN, KY, KS, MI, MO, MT, NC, NJ, NM, NV, OH, SC, SD, TN, TX, VA, WA, and WY. | Terminated | February 25, 2014 | 1310 | 97
Z-0140-2011 | 49605 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips BV Pulsera R2.2 Mobile X-ray system. The Type Product Code 718094 can be found under the X-ray trolley cover on the front lower right side of the module or under the stand cover. The device's brochure is labeled in part: "Philips...Mobile C-Arm systems BV Pulsera". |  |  | September 05, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Reliability of the Stand-Trolley cable may result in loss of images on the monitor during an examination. This intermittent problem can lead to System lock-up, where no radiation is possible and images are not visible on the monitor | Component design/selection | On 10/04/2008, Philips began sending the URGENT-Device Correction BV Pulsera R2.2 and Endura R2.2 letters to their customers. On 09/03/2008, the firm issued the Field Change Order (FCO) 71800024 and the firm's representative will replace the stand-trolley cable free of charge.Customers should contact Philips Call Center at 1-800-722-9377, #5, #2, and reference "FCO 71800024" with any questions. |  | 1331 | 1331 units | Nationwide distribution. | Terminated | December 07, 2010 | 823 | 1458
Z-0135-2011 | 56657 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | ACL TOP (Base) |  |  | August 13, 2010 | 2010 | Instrumentation Laboratory Co. | Software: New data check, "Multiple Threshold Check", was added to identify and fail abnormal (bimodal) clot curves | Software design | Instrumentaion Laboratory notified the US Distributor/Canada via BCI Service Modification Notice (#10793) on August 13, 2010. The Mandatory software upgrade to Verison V4.3.0 in the United States and Canada will be conducted by BCI service representatives and tracked through the distributor, Beckman Coulter (BCI). International Affiliates/Dealers:Worldwide Regulatory Notices were e-mailed to IL international affiliates and dealers, notifying them of the mandatory software upgrade, providing issue description and a risk summary and the support. The upgrades were to be handled through the local representatives. |  | 177 | 177 units | Worldwide Distribution, USA and Canada | Terminated | May 23, 2012 | 649 | 1076
Z-0139-2011 | 56703 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista V-LYTE Standard A, Catalog No K820 |  |  | August 19, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | High potassium content in the affected lots results in the potential for low potassium recovery for Dilution Check, QC and patient samples when run on the Dimension Vista System. | Other | Siemens sent an Urgent Field Safety Notice (dated "August 2010") to customers identifying the product, the problem and the action to be taken. Customers were instructed on how to continue testing until new lots of the recalled product were available and asked to complete and fax a Field Correction Effectiveness Check to 301-631-8467 to confirm receipt of the Urgent Field Safety Notice. For technical questions or concerns customers were instructed to contact the Siemens Technical Solutions Center at 800-441-9250. |  | 3706 | 3,706 cartons | USA, Nationwide Distribution | Terminated | October 26, 2010 | 68 | 3706
Z-0143-2011 | 56790 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | The label (which is affixed to ScanPoint Communications Cradle) states: "...VERATHON ScanPoint Docking Station...Part No. 0570-0168...Bothell, WA USA..." |  |  | September 20, 2010 | 2010 | Verathon, Inc. | Incorrect programming causes QuickPrint to fail to recognize the communication Cradles. As a result, ScanPoint Communications Cradles could not communicate with ScanPoint Online Tool while using the QuickPrint interface. This will affect the transfer of exams and device calibration through QuickPrint interface. | Software design (manufacturing process) | On 09/20/10, Verathon issued an URGENT: Medical Device Accessory Recall Notification letter dated September 20, 2010 to consignees, identifying the affected product, the issue prompting the recall, and actions to be taken by consignees.The consignees are instructed to return any cradles within the serial number range between CC012000-CC012084, and the Part Number 0570-0168.The firm will provide a replacement communication cradle at no cost. Consignees can contact Verathon Customer Care at 425-867-1348 or 800-2331-2313 or send e-mail to cservice@verathon.com. |  | 70 | 70 units | Worldwide Distribution: U.S. nationwide including states AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, KY, LA, MA, MD, MI, MN, MO, ND, NY, OH, PA, SC, TN, TX, UT, WI, and Washington DC, and the countries of Australia, Canada, France, South Korea, Taiwan, The Netherlands, and United Kingdom. | Terminated | April 01, 2011 | 193 | 70
Z-0413-2011 | 56580 | No | KOX | Surgical skin degreaser or adhesive tape solvent. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Detachol adhesive remover, 4 fl oz (118 mL), Marketed by Eloquest Healthcare, Inc., a Ferndale Pharma Group company, Manufactured by Ferndale Laboratories, Inc., Ferndale, MI 48220 U.S.A. |  |  | July 07, 2010 | 2010 | Ferndale Laboratories, Inc. | There is a defect on the neck surface of the bottles, causing bottles to leak. Bottle defect may potentially cause contamination of solvent stored in bottles. | Other | Ferndale Pharma Group issued an Urgent Medical Device Recall letter dated July 7, 2010 to customers identifying the affected product. Customers were instructed to identify and quarantine all affected product and return it to Ferndale Laboratories for a credit. Customers are requested to complete and return an enclosed response form.Customers can contact Ferndale at 800-621-6003. |  | 55000 | 55,000 bottles | Nationwide Distribution: Throughout the US. | Terminated | March 02, 2011 | 238 | 55000
Z-0145-2011 | 56893 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Vertex Max Drill Guide, Catalog Number 9732355, to be used with a Medtronic Computer-Assisted Surgery System to help assist with screw placement during spine procedures. |  |  | September 24, 2010 | 2010 | Medtronic Navigation, Inc | Drill guide handle may detach unintentionally during use. | Other | Medtronic Navigation sent a "Product Correction Notification" letter dated September 22, 2010 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to set aside and do not use the Vertex Max Drill Guide 9732355, Lot Number 100526 and complete and return the Verification of Field Action Completion form via email: steve.murray@medtronic.com or fax: 720-890-3517. A Medtronic representative will contact the customer, visit their facility, remove the defective product, and provide a replacement Vertex Max Drill Guide that is correctly manufactured.For further information or if you have questions regarding use of the Vertex Max Drill Guide, please contact Medtronic Navigation Technical Services at 1-800-595-9709. |  | 8 | 8 units | Worldwide distribution: USA including states of DC, IN, PA, and VA and country of Japan. | Terminated | April 13, 2011 | 201 | 8
Z-0144-2011 | 56754 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis Systems VB23A-VB23H with Perivision functionality, Model Numbers: 5904433,5904649, 5904466, 5904441, 5904656, 7728350, and 7007755 |  |  | September 09, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Sporadic system crash while Perivision is performed. | Software design | Siemens issued an Urgent Field Safety Notice to all customers. The letter informed customers of the potential issues and provided additional information on the implemented resolution. To resolve the problem, a system update was to be installed on each affected system under Update Instructions AX054/10/S. For questions phone (610) 448-3237. |  | 286 | 286 units | USA - Nationwide Distribution | Terminated | September 29, 2012 | 751 | 286
Z-0147-2011 | 56844 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Recalled chairs include the following: |  |  | September 07, 2010 | 2010 | Winco, Inc. | Winco Mfg., LLC, Ocala, Florida is recalling their Care Cliner. Due to malfunctioning axles which become lose from the caster assembly. This chair is a multi position recliner with a steel frame used primarily in clinics, hospitals and treatment centers for recovery, dialysis, oncology, examination, infusion. Recalled chairs include the following: Care Cliner, XL Care Cliner, Drop Arm Care Cliner | Device Design | Winco sent an Urgent Device Recall letter dated October 1, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Replacement parts were sent out to the consignees and customers were asked to dispose of the defective casters. Any consignee not reponding by November 15, 2010, would receive an additional letter. The firm placed a follow up phone call to the non responding consignees after two weeks.Customers were advised to call 800-237-3377 or email customer service at customerservice@wincomfg.com with questions or concerns.For questions or concerns regarding this recall call 352-854-2929 x 110. |  | 18522 | 18,522 casters | Worldwide Distribution - USA (all states except Vermont and Nebraska) and the countries of Belgium, Canada, Chile, Ireland, Kuwait, Netherlands, Saud Arabia, Singapore and United Arab Emirates. | Terminated | May 20, 2011 | 255 | 18522
Z-0158-2011 | 56718 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Catheter Extension Set, Luer Activated Valve for I.V. Access, Male Luer Lock Adapter, Volume 1.0 mL, Length 6" (15 cm); an Rx sterile device; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product 2N9191 |  |  | September 09, 2010 | 2010 | Baxter Healthcare Corp. | During packaging testing, Baxter observed tears and pinholes in Chevron-style pouches for certain infusion disposable products. The pinholes and tears could compromise the sterile barrier properties of the pouch. | Package design/selection | Baxter Healthcare Corporation sent an Urgent Product Recall letter dated September 9, 2010, to all customers. The letter identified the product, the problem and the action to be taken by the customer. The customers were requested to examine their inventory in all locations to determine if they had any of the affected lots of the affected product codes and to remove the affected product from their inventory; they were to call the Baxter Healthcare Center for Service at 1-888-229-0001 to return the affected product, request a credit, or request replacement product; complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they could comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  |  | data not available | Worldwide Distribution - U.S., Canada, Japan, Australia, New Zealand, and Chile. | Terminated | September 07, 2011 | 363
Z-0161-2011 | 56820 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Licox Brain PMO Probe Kit; |  |  | September 15, 2010 | 2010 | Integra LifeSciences Corp. | Incorrect expiration date - The outer boxes of the IP2P Licox Complete Kit for Cerebral Pressure, Temperature and ICP Monitoring Probes, Lot 250210A are labeled with an incorrect shelf life date (10 months too long). | Employee error | Integra issued Urgent Product Recall Notification letters dated September 16, 2010 to US customers and foreign consignees, identifying the affected lots and actions to be taken by customers.Affected product is not to be further distributed and should be quarantined pending return to the firm. Customers and Distribors should contact Integra Customer Service for a Returned Material Authorization Number, and then return the product to the Distribution Center. In addition, a Recall Acknowledgement and Return Form is to be completed and returned.Customers can contact Integra at 609 936-2495. |  | 23 | 23kits | Worldwide Distribution: USA, including the states of FL, ID, IL, and NJ, and the countries of Argentina and Costa Rica. | Terminated | April 18, 2012 | 581 | 23
Z-0160-2011 | 56148 | No | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Edwards' MHANM8 Pressure Slave Cable, Part Number:693657001 |  |  | June 25, 2010 | 2010 | Edwards Lifesciences, Llc | The recall was initiated because the Edwards' MHANM8 Pressure Slave Cable was wired incorrectly, thus will not work as stated. | Process design | Edwards Lifesciences contacted customers via telephone on June 22, 2010 informing them of the product, the problem and the action they should take. Customers were asked to locate their cables and send them back overnight using the provided Edwards FedEx account number.Edwards Lifesciences indicated their personnel would follow-up with each customer until all units were returned.For any questions regarding this recall call (949) 250-2007. |  | 20 | 20 units | Worldwide Distribution - USA including CA, PA, and MI and the countries of Switzerland and Thailand | Terminated | May 24, 2011 | 333 | 20
Z-0159-2011 | 55712 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Impac Software Sequencer Verify &amp; Record System, for MOSAIQ, manufactured by Elekta Impac Software, Sunnyvale, CA. |  |  | February 19, 2010 | 2010 | Impac Medical Systems Inc | Mistreatment-- The database conversion utility used during the upgrade errantly changes the start gantry angle to zero. Any affected treatment fields, if not already treated, would start delivering from zero instead of the original prescribed value. | Process change control | ELEKTA IMPAC SOFTWARE sent a USER NOTICE dated February 19, 2010, to all customers. The notice identified the product, the problem, and the action to be taken by the customers.Customers were urged to verify the start angle of the gantry against the treatment field definition prior to delivery.Review all treatment plans and charts with conformal arcs. If a treatment plan with an errant gantry start angle was identified, customers were to cease using that plan immediately. A workaround was to be re-imported into their current version of MOSAIQ.IMPAC Support was to follow-up with each site and execute a script to detect and correct all affected treatment plans.For questions, concerns, and requests for upgrade, customers were to email support@impac.com or call their support center. |  | 19 | 19 units. | Worldwide Distribution -- USA, Australia, Belgium, France, Germany, Hong Kong, Ireland, Israel, and Italy | Terminated | March 28, 2011 | 402 | 19
Z-0167-2011 | 56975 | No | KPT | Radionuclide dose calibrator. | Radiology | Radiology | 510(k) | Captus 3000 Thyroid Uptake System with Portable Stand option. Capintec item number 5430-2130, serial numbers: 91001, 91003 through 91010. Manufactured by Capintec, Inc, Pittsburgh, PA. |  |  | September 20, 2010 | 2010 | Capintec Inc | The spring arm on portable stand may have been over tightened during assembly, which may result in failure. If the spring arm fails, it will not balance the collimator weight, and the arm will drop to the lowest position. | Employee error | Capintec issued an Alert letter to all customers identifying the affected device, informing them of the problem and what actions are to be taken by customers. Capintec will send a replacement arm for the system, which is to be installed by the customer using included instructions. The firm included a correction acknowledgement form to be completed and returned by the customer.Customers can contact Capintec concerning this recall at 201 825-9500. |  | 9 | 9 units | Worldwide Distribution: USA in the states of NC and TN, and in the countries of Pakistan, Paraguay, Portugal, Saudi Arabia, and Spain. | Terminated | July 05, 2012 | 654 | 9
Z-0162-2011 | 56826 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | IV Catheter - Optiva IV Catheter. This is found in box labeled 'Graft Jacket Xpress Flowable Soft Tissue Scaffold'; labeled by Wright. |  |  | September 13, 2010 | 2010 | LifeCell Corporation | Graft Jacket Xpress kits contain catheters that will expire or have already expired prior to the labeled expiration of the human tissue component of the product on the outer box. | Labeling mix-ups | Lifecell Corporation, Branchburg, NJ initiated a recall to their distributor Wright Medical Technology, Arlington, TN on 9/13/2010. |  | 1507 | 1507 kits | Nationwide Distribution: USA, to distributor in TN | Terminated | August 17, 2012 | 704 | 1507
Z-0163-2011 | 56852 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion with EXTEND Frame System, manufactured by Electa Instrument AB, Stockholm, Sweden. |  |  | September 17, 2010 | 2010 | Elekta, Inc. | The anterior fixation bar of the EXTEND mouthpiece frontpiece may separate from the dental impression tray. This may result in reduced treatment accuracy. | Other | Elekta issued an Important Notice: Urgent Field Safety Notice dated September 17, 2010 concerning Field Change Order 10000202003-A342 "Important Notice regarding recall of mouthpieces for EXTEND system, to all users of the EXTEND system for LGK Perfexion. the letter identified the issue and associated potential hazard, and requested the identification and return of all affected mouthpieces, and confirmation of receipt of the letter. Elekta will arrange return and replacement through local service representatives.Customers can contact Elekta concerning this recall at 770 300 9725. |  | 3 | 3 units | Nationwide Distribution: to the states of MO, VA, and WA. | Terminated | November 10, 2010 | 54 | 3
Z-0164-2011 | 56908 | No | HBT | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LeMaitre Vascular AnastoClip GC Vessel Closure System-Medium Size |  |  | September 22, 2010 | 2010 | Lemaitre Vascular, Inc. | Mislabeled size on external packaging: The contents section listed on the Tyvek lid labels of the Medium and Large sizes incorrectly identified the size of the clip as XL, while the device inside was either a Medium or Large size. | Packaging change control | LeMaitre Vascular issued an Urgent AnastoClip GC Device Field Safety Notice letter dated Sepember 28, 2010 identifying the affected devices and actions to be taken by customers. The letter contains a form that is requested to be returned to LeMaitre Vascular as record of notification and reconciliation. The letter provides an option for the notified hospital to return the affected product for replacement if unused or if they simply wish to have it corrected. Customers can contact LeMaitre Customer Service for replacement at 800 628-9470. |  |  |  | Nationwide Distribution: to the states of CA, FL, GA, IL, IN, LA, MD, MS, NJ, OK, OR, and TX. | Terminated | September 15, 2011 | 358
Z-0166-2011 | 56352 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Siemens brand 550 TxT Patient Table, Part No. 7346534, Distributed by and/or Manufactured by Siemens AG, Medical Solutions, Germany |  |  | March 31, 2010 | 2010 | Siemens Medical Solutions USA, Inc | An increase in voltage may cause a short circuit of a power transistor that may cause a smoldering and smoking effect of the electrical wiring and also an unexpected lowering of the Table. Both issues present the potential for injury to a patient or the operator. | Other | Siemens sent a Customer Information Letter on March 31, 2010, to all consignees. The letter identified the product, the problem, and the action that should be taken. Siemens implemented minor modifications to address the issues. No further action was required for the continual operation of the 550TxT Table. For questions regarding this recall call 925-602-8083. |  | 326 | 326 units | Worldwide Distribution including USA | Terminated | March 01, 2012 | 701 | 326
Z-0170-2011 | 56909 | No | LPB |  |  | Cardiovascular | PMA | Boston Scientific Peristaltic Tubing Kit, Model 2104, UPN M0043521040. For use with Boston Scientific CHILLI Cooled Ablation Catheters and Model 8005 Continuous Operation Pump. Sterilized using Ethylene Oxide. Contents: peristaltic pump fluid delivery extension tube set. Manufactured for: Boston Scientific Corporation, One Boston Scientific Place, Natick, MA 01760-1537, USA. |  |  | September 27, 2010 | 2010 | Boston Scientific Corporation | Boston Scientific is conducting a Recall on CircuCoolÂż Tubing Sets used with the Chilli IIÂż Cooled Ablation System because there may be a potential for a sterile barrier breach which could compromise device sterility and potentially the sterile field. Even though there have been no reported complaints. | Packaging | Boston Scientific Corporation sent an Urgent Medical Device Recall letter, dated September 27, 2010, to all customers via overnight Federal Express. The letter described the product, the problem, and the action the customer should take. The letter instructed customers to cease distribution of affected devices and remove them from their inventory, segregate in a secure location for return to Boston Scientific. A Reply Verification Tracking Form was to be completed and returned via email or faxed to (510) 624-2475 even if the customer did not have any of the affected product. If affected units were located the customer was to call (510) 624-2563 to obtain a Returned Good Authorization (RGA) Number and return to:Boston Scientific CorporationUS Distribution CenterBoston Scientific Marina BayCustomer Fulfillment Center500 Commander Shea Blvd.Quincy, Massachusetts 02171RGA: _________________For additional questions call (510) 624-2563. |  | 4694 | 4031 US , 663 OUS | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NY, OH, PA, RI, SC, TN, TX, VA, WA, and WV and the countries of France, Germany, Great Britain, Hungary, Israel, Italy, Netherlands, Saudi Arabia, Spain, and Taiwan | Terminated | April 17, 2012 | 568 | 4694
Z-0171-2011 | 56921 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobilett XP and Mobilett XP Hybrid, Mobile radiology device |  |  | October 04, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Units not equipped with protective plastic cover over electronics. | Device Design | Siemens Medical Solutions USA, Inc issued a mandatory field update. The letter identified the product, the problem and the action to be taken by the customer. All systems that were not yet equipped with protective cover over the D972 board would be upgraded.For questions regarding this recall call (610) 448-3237. |  | 177 | 177 units | Nationwide Distribution including AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, LA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, PA, TX, VA, WA, WI, WY AND PUERTO RICO. | Terminated | September 29, 2012 | 726 | 177
Z-0134-2011 | 56902 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medtronic Octopus Nuvo Tissue Stabilizer, model TSMICS1. The Octopus Nuvo Tissue Stabilizer is a disposable tissue stabilization system consisting of a rigid shaft, detachable tissue stabilizer headlink and a flexible articulating arm. The device is assembled within the thoracic cavity. Sterilized using ethylene oxide, non pyrogenic, disposable, single use only. |  |  | September 13, 2010 | 2010 | Medtronic Perfusion Systems | Medtronic has received reports regarding the Octopus Nuvo Tissue Stabilizer stating that during use, the collet (a component of the device that stabilizes the head link) fractured, thereby causing immediate separation of the head link from the shaft of the device. The resulting potential hazards are that fragments of the collet could fall into the patient's chest cavity and/or damage the heart ti | Component design/selection | Medtronic sent an "Urgent Medical Device Recall Notice" dated September 14, 2010, to Risk Managers at each affected account. The letter described the issue, identified affected lot numbers, asked for immediate discontinued use of the device, quarantine all unused devices and to return unused devices to Medtronic CardioVascular, 7611 Northland Drive, Brooklyn Park, MN 55428 and requested a recall certificate be faxed back to Medtronic indicated acknowledgement of the Recall Notice and quantity to be returned.If you need additional information, please contact Medtronic Cardiovascular Lifeline Technical Services at 877-526-7890, or your local Medtronic sales representative. |  | 571 | 571 units | within the US to: AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, MA, MI, MN, MS, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI and Washington D.C. OUS to include: Belgium, Canada, Germany, Italy, Poland, Switzerland, United Kingdom. | Terminated | April 16, 2012 | 581 | 571
Z-0168-2011 | 56794 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano s5/s5i imaging system Family Software version 3.2, Part Numbers 435-0602.01 to 05, including ChromaFlo and Virtual Histology, manufactured by Volcano Corporation of Rancho Cordova, CA. |  |  | June 29, 2010 | 2010 | Volcano Corporation | Software defect results in improper image brightness, which may result in introduction of artifacts, which may cause misclassification of results and affect classification of tissue types. This may result in misdiagnosis and inappropriate treatment. | Software design | Volcano issued a Voluntary Recall of Volcano s5 Family Software version 3.2 letter dated June 30, 2010, identifying the software defect and actions to be taken by Volcano. Field Service Engineers will visit each site and reinstall the previous software version, which will resolve the issue.Customers can contact Volcano at 916 281-2790. |  | 79 | 79 software kits | Worldwide Distribution: USA, and the countries of France, Germany, Japan, and the UK. | Terminated | December 22, 2010 | 176 | 79
Z-0169-2011 | 56858 | Yes | NOX |  |  | Orthopedic | Contact ODE | Bone Void Filler, sterile orthopedic product, 510 (k) #K053228. The product is delivered in a syringe which is packaged in an aluminum-backed polyethylene pouch and is distributed one (1) pouch per box. The product affected by this recall was distributed by the recalling firm under six (6) different names. They are: Form-Putty; SBvFShape; Valeo BP; X-Putty; TheriGraft Putty; and SynFlow. |  |  | August 26, 2010 | 2010 | Theken Spine Llc | The firm's quality control testing determined that a separation was occurring between the flowable carrier and the B-TriCalcium Phosphate within some of the product syringes. | Other | Theken Spine Llc sent Recall Notification Letters dated August 26, 2010, to all customers. One (1) document, titled an 'Urgent Product Recall Notification' letter was addressed: 'ATTENTION; PRODUCT RECALL COORDINATOR/RISK MANAGER.' The other letter was titled: 'RECALL NOTIFICATION.' The recall letters were accompanied by an attached 'Recall Acknowledgement and Return Form' and were sent out via Federal Express. The letters identified the product, the problem, and the action to be taken by the customer. The letters requested that the customers review their product in inventory and identify any product remaining from the affected lot numbers; part numbers; and/or item names for immediate return of these items to the recalling firm. The letters also asked that the customer complete and return the attached Recall Acknowledgement and Return form according to the instructions provided on the form. In addition, the 'RECALL NOTIFICATION' letter instructed the distributor/customers provide information concerning their sub-account customers to Theken Spine so they could issue a Recall Notification letter to these customers as well. To obtain instructions from the recalling firm concerning product return and issuance of credit, customers were instructed to call (866) 942-8698. For questions regarding this recall call (330) 825-8258. |  | 2904 | 2904 units (syringes) | Nationwide Distribution - including UT, SD, TX, AL, PA, CA, MI, OH, NJ, NH, MD, ID, FL, LA, NM, TN, IA, and OR. | Terminated | October 27, 2011 | 427 | 2904
Z-0288-2011 | 53863 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | The product is a 19" Spacelabs Medical Flat Panel Display, Model 94260-19, manufactured by ELo Touchsystems |  |  | October 23, 2009 | 2009 | Spacelabs Healthcare, Incorporated | The Spacelabs Medical Flat Panel Display, Model 94260-19 may fail prematurely and causing the screen to go blank. This inability to view data may result in delayed patient treatment. | Other | On 10/23/09, Spacelabs issued URGENT: MEDICAL DEVICE RECALL notification letters dated October 23, 2009 to their consignees.The notification states the recalled product as Spacelabs Medical 19" Flat Panel Display, Model 94260-19. The consignees are informed that there is apossibility that the internal power distribution circuit board will fail prematurely, causing the display to go blank. Audio alarms will continue to function normally.The firm's Field Service Engineer will be contacting the consignees to set up a convenient time to schedule the upgrades of the affected displays. Displays with part number ET1920L-7SKR-1-35S-SPC are not part of this recall and should not be sent in for upgrading.The firm recommends that their consignees brief their staff on the possibility of the display going blank to help prevent an adverse event from happening. The firm will be providing loaners while the affected displays are upgraded at free of charge.If the consignees have any questions about the recall, they could contact Spacelabs Medical at 1-800-522-7025 and select 2 for Technical Support . Consignees outside the United States may call 1-425-657-7200 x 5089. |  | 2560 | 2560 units (2020 domestic, 540 international) | Worldwide Distribution: USA in the states of AK, AL, AR, CA, CO, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WY, and the following countries: Australia, Austria, Canada, Czech Republic, Egypt, France, Germany, Grand Cayman, Greece, India, Israel, Italy, Kuwait, Lebanon, Mexico, Netherlands, Pakistan, Panama, Poland, Puerto Rico, Saudi Arabia, Singapore, Slovakia , South Africa, Switzerland, Trinidad and Tobago, Turkey, United Kingdom, and Venezuela. | Terminated | November 03, 2010 | 376 | 2560
Z-0308-2011 | 55242 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Panta Support Device PantaÂż Arthrodesis Nail Instrumentation Tray One unit per package. |  |  | March 01, 2010 | 2010 | Integra LifeSciences Corporation | Inability to properly align the calcaneal screw placement on a consistent basis, due to a manufacturing defect of of the targeting frame (Panta Support Device). | Nonconforming Material/Component | OIntegra issued Urgent Product Recall Notification letters dated 3/1/2010 to customers and distributors to notify them of the problem. Customers were requested to check their inventory to determine if they have an of the affected devices. If so, custo0mers should stop using the product. Custoemrs with the affected product should contact their local Sales Representative for assistance, or contact Integra at 800-654-2873 - 3 to obtain a returned material authorization number, and return the product to their local representative.All customers were requested to complete and return a Recall Acknowledgement and Return Form.. |  | 108 | 108 Units | Worldwide distribution: USA including states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SD, TX, UT, VA, WA and WI; and countries of: Australia, Austria, Belgium, Czech Republica, Denmark, Finland, France, Germany, Ireland, Italy, Portugal, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | February 16, 2012 | 717 | 108
Z-0302-2011 | 56938 | No | LTI |  |  | Gastroenterology/Urology | PMA | SAGB QuickClose Swedish Adjustable Gastric Band including Velocity" Injection Port and Applier. Product Code BD2XV, Lot #: 1040602, , 1040603, , 10270503, , 11020504, , 11030502, , 11090510, , 11090511, 1090512, 12030502, 12070502, 12070503, 12070504, 12200501, 12200502, 12200503, 12290505, 12290506, 01120603, 01200602, 01260601, 01260602, 01260603, 01260604, 02210605, 02210606, 02210607, 02280622, 02280623, 03080601, 03080602, 03080605, 03220611, 03220612, 05150601, 05150602, 05150603, 05150604, 05300613, 05300614, 06190606, 06190607, 06190608, 06190614, 06190615, 06190616, 06190617, 06190618, 06260601, 07060601, 07100605, 08140607, 08140608, 09050601, 09050602, 09050603, 09050604, 09050605, 09130601, 09130602, 09130603, 09130604, 09130605, 09180602, 09180603, 09180604, 09250602, 10110601, 10110602, 10110603, 20002024, 20002082, 20002185, 20002186, 20002193, 20002438, 20002439, 20002440, 20002441, 20002484, 20002485, 20003152, 20003153, 20003154, 20003155, 20003196, 20003362, 20003363, 20004437, 20004438, 20005244, 20005659, 20006819, 20006921, ZGNB43, ZGNB44, ZGNB45, ZGPBCC, ZGPBCD, ZGPBCJ, ZGPBCK, ZHBBBN, ZHBBBP, ZHBBBT, ZHCBB0, ZHCBB1, ZHCBB2, ZHCBB3, ZHCBB5, ZHCBB6, ZHCBB7, ZHCBB8, ZHCBB9, ZHDBF7, ZHFB76, ZHFB77, ZHFB79, ZHFBBW, ZHFBBY, ZHGBBJ, ZHGBDJ, ZHGBDK, ZHGBDL, ZHHBD6, ZHHBD7, ZHHBD8, ZHHBD9, ZHHBFB, ZHHBFC, ZHJBBJ, ZHJBBK, ZHJBBL, ZHJBBM, ZHJBBN, ZHJBBP, ZHMBDP, ZHNBC1, ZHNBC2, ZHPBBT, ZJBBBF, ZJBBDP, ZJCBB6, ZJCBDP, ZJCBGL, ZJDBBD, ZJDBBF, ZJDBBG, ZJDBBH, ZJDBBJ, ZJDBBK, ZJFBL6, ZJFBML, ZJHBCH, ZJHBK7, ZJJBC3, ZJKBC1, ZJKBHC, ZJKBK2, ZJMBB4, ZJMBN9, ZJNBBV, ZKDBM6, ZKDBM7, ZKFBDC, ZKFBDH, ZKGBDY, ZKHBBJ, ZKHBCR, ZKHBFG, ZKNBCK, ZKNBHC, ZLBBFK, ZLDBB8, ZLFBBC, ZLFBGG, ZLGBFV &amp; ZLGBG5. |  |  | September 27, 2010 | 2010 | Ethicon Endo-Surgery Inc | The Tubing Strain Relief component of the Injection Port may detach from the Locking Connector, and migrate from its original position on the Locking Connector. This may result in: Tubing Strain Relief migrating on the Tubing; Kinking of the Tubing resulting in difficulty/inability to adjust the fluid volume within the Band; or a free foreign body in the abdominal wall or abdominal (peritoneal) ca | Other | Ethicon Endo-Surgery issued an Urgent : Devoced Recall Event 2240 letter dated October 8, 2010 to all consignees identifying the affected devices and actions to be taken.Consignees were instructed:DO NOTt use product from the affected lots, but implanted devices do not need to be removed as part of this recall.Examine inventory and remove the affected product.If they DO have the affected product, complete the Business Reply Form, keeping a copy for records, and place the original form in the box with the product to be returned, affix the pre-printed label, and return to Stericycle.If they DO NOT have the affected product, complete the Business Reply Form, and fax it to Stericycle.A redesigned replacement product will be sent out to consignees within 30-60 days of product return.Ethicon Endo-Surgery can be contacted at 513 337-8208 concerning this recall. |  | 26966 | 26,966 pieces | Worldwide Distribution: USA including the following states: AK, AL, AR, AZ, CA, CO, CT, Washington DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, Puerto Rico, RI, SC, TN, TX, UT, VA, VT, WA, WI and WV, and the following countries: ARAB EMIRATES, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, EGYPT, FRANCE, GERMANY, GREAT BRITAIN, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LIBYA, MEXICO, NETHERLANDS, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, RUSSIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY and VENEZUELA. | Terminated | November 18, 2011 | 417 | 127130
Z-0130-2011 | 56735 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | AMO Total Care (1) Contact Lens Solution, Formula: 8772x; Product Code: 90275 (240 ml). |  |  | July 23, 2010 | 2010 | Abbott Medical Optics Inc (AMO) | The recall was initiated after AMO determined, through internal testing, that a limited number of the flip top caps used during production of these solutions may leak and, although unlikely, the sterility of the product may be compromised. Products that are non-sterile have the potential to cause eye infections, which may be sight threatening. | Pending | AMO prepared and issued on July 28, 2010 both a customer and distributor notification letter as part of their customer communication process. The letter informed the customers of the reason for recall, the products involved, the hazard involved and actions to be taken regarding the affected products.Customer with questions related to the COMPLETE Multi-Purpose Solution and/or Total Care (1) Contact Lens Solution return process, were instructed to contact an AMO Customer Service Representative at 1800 931 663. |  | 144 | 144 units | International Distribution Only -- United Kingdom, Lebanon, Dubai, Bahrain, Russia, Kazakhstan, Ireland, and France. | Terminated | October 13, 2011 | 447 | 34224
Z-0309-2011 | 56837 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 3.5mm LCP Periarticular Proximal Humerus Plates |  |  | September 16, 2010 | 2010 | Synthes USA (HQ), Inc. | screw can pass through plate hole | Nonconforming Material/Component | Synthes (USA) notified all Synthes Trauma Hospital accounts of record of this recall via an Urgent Medical Device Recall letter dated September 27, 2010. The notification included instructions for hospital accounts to immediately cease use of the product, examine inventory, remove the 3.5 mm LCP Periarticular Proximal Humerus Plates from their shelves and complete and return the attached Verification Section form via mail: Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132; fax: 610-215-9005 or Scan/email: FieldAction@synthes.com. If they have any of the identified devices, they were to call Synthes at 1-800-479-6329 to obtain Authorization number prior to completing Verification form.If you have any questions, please call 1-800-620-7025 x5403 or contact your Synthes Trauma Sales Consultant. |  | 2849 | 2849 units | Worldwide distribution: USA and countries including: Bermuda, Korea, Italy, Japan, Georgia, Qatar and Turkey | Terminated | April 16, 2015 | 1673 | 2849
Z-0287-2011 | 56324 | No | DTB | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | PMA | enpath medical Myopore Sutureless Myocardial Pacing Lead, Bipolar (labeled as Unipolar) Connector, , REF 511212. 54 cm. Distributed by Guidant/CPI. Contents in unopened, undamaged package have been sterilized by ethylene oxide gas. Manufactured by enpath medical, Inc., 2300 Berkshire Lane North, Minneapolis, MN 55441 USA. |  |  | May 13, 2010 | 2010 | Greatbatch Medical | The Myopore Bipolar Epicardial Leads packaging is labeled as unipolar leads. The lead body itself is correctly labeled as bipolar. This could result in a bipolar lead being implanted with a pulse generator operating in unipolar mode. | Employee error | Greatbatch Medical issued an Urgent Medical Device Recall letter dated May 13, 2010 to their distributor. The letter described the problem and product involved, and gave the action required which included:1) immediately discontinue distribution of the units and return any unused units to Greatbatch Medical.2) Reconcile medical recordsbetween BSC and Greatbatch Medical to confirm implant data3) Confirm pulse generator type (bipoloar, unipolar, or unknown) used with each serial number listedGreatbatch will work with customers to contact physicians as necessary regarding patient monitoring. Greatbatch Medical can be contacted at 763-951-8312. |  | 34 | 34 | Worldwide Distribution: to USA and Europe, through distributor in MN. | Terminated | May 10, 2012 | 728 | 34
Z-0312-2011 | 55287 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | STar Drive System, |  |  | March 30, 2010 | 2010 | FHC, Inc. | Fixation thumbscrew on the DBS depth stop adaptermay be overtightened and damage the implantable lead. | Device Design | FHC issued an "URGENT: MEDICAL DEVICE CORRECTION" letter dated March 29, 2010, to all customers. The letter described the product, problem and actions to be taken by the customers. FHC informed the customers that that they had returned to the original design of the fixation screw in the Depth Stop Adapter. FHC sent replacement Depth Stop Adapter with an original fixation screw. The customers were instructed to discard the recalled Depth Stop Adapter (s) and replace with the enclosed corrected version(s) and sign and fax the receipt notification to 207-666-8309 as proof of their receipt of notification. For assistance with the product replacement, contact the FHC's technical support line at 800-326-2905. This line is staffed 24 hours a day.Should you have any further questions regarding this matter, please contact Dorothee Sasse, product manager, or Keri Seitz, recall coordinator at 207-666-8190. |  | 7 | 7 | Worldwide distribution: USA in the states of AZ, CA, FL, ID, NC, MA, MI, MN, NV, RI, UT, VA, and WA, and countries of Columbia and South Africa. | Terminated | February 25, 2015 | 1793 | 63
Z-0316-2011 | 55748 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACS Sample Prep Assistant III; (BD SPA III), Manufactured by BD Biosciences, San Jose, CA. |  |  | February 19, 2010 | 2010 | BD Biosciences, Systems &amp; Reagents | Due to inaccurate sample dispense, use of product may result in getting incorrect results on samples prepared using the device. Having a piece of the lid partially or fully clogging the Probe during piercing a hole in the lid of tube may cause inaccurate sample dispense, resulting in incorrect test results.The piece of the lid clogging the Probe can be in the processed sample tube. It has the | Device Design | BD Biosciences sent an "Important Product Information" letter dated February 2010 and an "URGENT: Follow-Up to Letter Dated February 2010" letter dated July 2010 to customers identifying the affected device and actions to be taken: Customers were instructed to stop use of the SPA probes and to contact BD for replacement of the affected probes and assistance in the event of probe clogging. BD has arranged replacement of affected probes. BD's followup communication detailed the potenial hazards and reiterated remedial actions in the event of error messages from FACS flow cytometers.Customers with questions about the performance of theirr SPA instrument or the content of the letters may contact BD Customer Support at 1.877.232.8995, prompt 2. |  | 349 | 349 units | Worldwide distribution: USA including OH, NC, AL,TX, FL, CA, NY, and MN and the country of Belgium. | Terminated | May 18, 2011 | 453 | 349
Z-0315-2011 | 56499 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Distal Centralizer 10 mm |  |  | July 28, 2010 | 2010 | Ortho Development Corporation | The hip stem did not fit the distal centralizer because the hole was not drilled with the appropriate taper. | Nonconforming Material/Component | Ortho Development sent a Field Notification Memo dated July 28, 2010 to Japan Medical Dynamic Marketing instructing them to reconcile parts that may have been used and return unused parts to Ortho. Customers can contact Ortho Development at 1-801-619-3419. |  | 39 | 39 units | Worldwide Distribution: Japan | Terminated | November 04, 2010 | 99 | 39
Z-0319-2011 | 56668 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT BNP cartridge |  |  | January 01, 2008 | 2008 | Abbott Point Of Care Inc. | i-Stat cartridges may not fill, cartridge labels or channels may appear discolored, cartridges may have fluid visible in the cartridge channel, and cartridge packaging may be distended. | Other | Abbott Point Of Care Inc.sent notification letters with a customer reply form to all affected customer in January, 2008. All distributors outside the US who were sent shipments from the affected lots were notified of the lot numbers that were affected. Affected product is to be quarantined and returned for replacement. Customers with questions about this recall can contact Abbott Point Of Care at 609 454-9000. |  | 482500 | 482,500 (total) | Worldwide Distribution: USA and Canada. | Terminated | September 14, 2011 | 1352 | 482500
Z-0310-2011 | 56825 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | smith&nephew; EZCARE Negative Pressure Rx only. |  |  | January 29, 2010 | 2010 | Smith &amp; Nephew Inc. | Smith&Nephew; Inc., Advance Wound Management, St Petersburg, FL is recalling their Ezcare Negative Pressure Wound Therapy pumps manufactured by BlueSky Medical Group Inc., between July 17, 2007 and December 13, 2007 and distributed by Smith&Nephew; between July 18, 2007 and January 14, 2010. Affected products codes are 66800187, 66800322, 66800323, 66800324, and 66800325 with Serial Numbers startin | Other | Smith &amp; Nephew Inc. FEDEXed overnight an URGENT DEVICE RECALL letter dated January 27, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were to check their inventory and identify any affected product, fill out and return the response form by both mail and fax, and return the affected product with the prepaid Federal Express (FedEx) Ground label provided by Smith &amp; Nephew. For any questions regarding this recall contact the Customer Care Center at (800) 876-1261. |  | 2717 | 2717 units | Worldwide Distribution - US including AL, AR, AZ, CA, GA, IL, IN, MA, MD, MD, MO, NC, NJ, NY, OH, OR, PA, TN, TX, and UT and the countries of Australia, UK, South Africa, Canada, Puerto Rico, and New Zealand. | Terminated | November 10, 2010 | 285 | 2717
Z-0320-2011 | 52231 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Straumann T-Sleeve 2.8mm, height 6mm, guided |  |  | June 04, 2009 | 2009 | Straumann Manufacturing Inc. | Incorrect Inner diameter of the T-sleeve | Process control | Strauman USA LLC notified the consignee by telephone on June 4, 2009, and sent an Urgent Device Recall letter dated, June 8, 2009 as a follow-up. The letter described the product, the problem, and the action to be taken by the customer. The customer was asked to return any affected product, in the prepaid preprinted mailer pouch as soon as possible to:T-sleeve Recall, Straumann USA, LLC60 Minuteman RoadAndover, MA, 01810Customers were asked to sign and indicate on the Recall Acknowledgement Form if they would like a replacement product or a credit to their account for the product returned and fax the form to (978) 747-0023.For any questions regarding this recall call (978) 747-2575. |  | 3 | 3 units | Worldwide Distribution - Canada | Terminated | February 20, 2012 | 991 | 3
Z-0321-2011 | 56785 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Single-stage 3.5mm Healing Abutment, 4 mm, REF SYHA40, Rx Only, STERILE R, gamma irradiated, expires 2014-44, CE 0473, BIOHORIZONS, BIRMINGHAM, AL 35244. |  |  | September 07, 2010 | 2010 | BioHorizons Implant Systems Inc | Healing abutment may not fully seat on the implant. | Other | Consignees were notified by an Urgent: Medical Device Recall letter on 09/07/2010. The letter identified the affected product and gave some background regarding the issue. It also stated the possible impact and the cases that were not affected. Consignees were instructed to contact BioHorizons Returns &amp; Complaints Manager Rick Self at 205-986-1250, if they have unused product from the affected lot, for a Return Authorization Number and to receive replacements. The attached Notification Certification form is to be completed and returned via fax. |  | 396 | 396 | Worldwide Distribution -- USA, including PA, MI, NY, AL, WA, and WV, and countries of Canada and Chile. | Terminated | December 13, 2010 | 97 | 396
Z-0305-2011 | 56164 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | B. Braun addEASE Binary Connector, Catalogue #N7990, 200 units (50 units per bag with 4 bags per shipping box) |  |  | June 24, 2010 | 2010 | B. Braun Medical Inc | During this insertion, fragmentation of the PAB container stopper may occur resulting in a small amount of visible particulates in the solution. | Device Design | On July 8, 2010 B Braun sent an Urgent - Medical Device Recall letter, dated June 28, 2010, to consignees informing them of the recall and that further distribution or use of the affected product should cease immediately. The attached "Product Removal Acknowledgement" form should be used to record the number of units and full-unopened cases that customers have in their inventory. The completed form should be returned via mail or fax within 2 weeks of receipt. If consignees have any full or partial cases of these affected products, they are to please call the Customer Support Department at (800) 227-2862 to arrange for return. Canadian customers may call (800) 624-2920. International customers may call (610) 691-5400. A Customer Support Representative will provide consignees with instructions for handling the affected product. The firm will arrange for unopened cases to be returned to B. Braun Medical, Inc. for proper disposal. Individual units (out of shipping cases) may be destroyed at the consignees facility.Consignees were informed that B. Braun is currently working to resolve this issue and provide an acceptable alternative to the market. They expect this product to be available in August 2010. In the interim, B. Braun PAB containers can continue to be used safely with a standard syringe and needle in accordance with the Direction for Use. If customers have additional questions regarding with a matter, they are to call B. Braun Customer Support Department at the number listed above. Should consignees identify product where this condition has occurred, consignees were asked to report it promptly to B. Braun Medical Inc. by contacting the B. Braun Clinical and Technical support at 1-800-854-6851. |  | 33905600 | 33,905,600 units | Worldwide Distribution -- USA and Canada. | Terminated | October 19, 2011 | 482 | 36405600
Z-0307-2011 | 56750 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | WalkMed Triton Pole Mount Infusion Pump, Catalog 300000. |  |  | June 17, 2010 | 2010 | Walkmed Infusion LLC | Pump door may be in a near shut position, but unlatched, and the "Door Open" alarm may not sound. | Device Design | All customers were notified by letter on 06/17/2010, describing the affected product and advising them of the possible malfunctioning alarm in the case a user does not follow the instructions as specified in the Triton Operation Manual. The letter also reminded all users to follow the procedures in the IFUs. The firm stated that it was best if customers re-educated all nurses on the proper door closure procedure to help prevent the potential bypass issue from occurring. Attached to the letter were two pages explaining the proper process for closing the pump door. Follow-up phone calls were made to coordinate an upgrade to all affected pumps. Questions or concerns should be directed to Ross Kurz at 720-531-4933. |  | 1565 | 1565 pumps | Nationwide Distribution -- AL, CA, FL, and PA. | Terminated | March 14, 2012 | 636 | 1565
Z-0146-2011 | 56833 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Sigma Spectrum VOLUMETRIC INFUSION PUMP, with or without Master Drug Library, Part Number 35700 -- |  |  | September 15, 2010 | 2010 | Sigma International General Medical Apparatus, Llc | The infusion pumps have the potential to fail causing inaccurate flow conditions during use. These conditions range from back flow to free flow, which could result in over-infusion. | Device Design | Urgent: Medical Device Recall Notification letters and forms (dated 9/14/10) were sent to the consignees via Certified Mail on/about 9/15/10. The letters inform the consignees of the recall of the pumps since these units may fail suddenly causing inaccurate flow conditions during use, ranging from back flow to over-infusion, including free flow. Consignees are instructed to check their inventory to verify whether their serial numbers fall within the range of pumps being recalled. Due to possible bearing failure, SIGMA is taking the precautionary measure of servicing or replacing all affected Spectrum Infusion Pumps and is requiring the return of the recalled devices to its facility. SIGMA has provided the following instructions for users: (1) Do not use the affected Spectrum Infusion Pumps on patient populations where inaccurate flow, ranging from back flow to over-infusion, including free flow, could result in serious adverse health consequences or death. (2) Do not use on neonatal patients. -- If the user facility has no acceptable alternative methods for delivering therapy, the user should weigh the risks and benefits for the continued use of the affected pumps. If you must use these pumps, SIGMA recommends the following actions which may mitigate the risk. (1) for pediatric and adult patients, an add-on Buretrol may limit the amount of an over-infusion if the failure occurs. Always follow the Buretrol manufacturer's instructions. (2) consider using air-eliminating filters. Under free flow conditions pumps cannot stop fluid flow. If air is in the line with fluid, the air will also infuse. (3) Frequently observe that the drip rate (mL/hr) compares to the programmed pump flow rate. Also, make an observation of the container volume to make sure it is correct. If a discrepancy is observed, immediately remove the pump from service and treat the patient as appropriate. There are additional instructions to Asset Management Customers: SIGMA requests that each A |  | 17557 | 17,557 Infusion Pumps | Nationwide Distribution | Terminated | June 28, 2017 | 2478 | 17557
Z-0323-2011 | 55655 | No | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | The device is labeled in part: " RS Medical RS-4i Stimulator". The device is sold in a kit that includes a carrying case, connector cables, pads (electrodes), battery charger, patient document, instructional DVD, return mail bag, and the operation manual. |  |  | June 11, 2008 | 2008 | International Rehabilitative Sciences, Inc. | The charge current is leaking through two Diodes, which are supposed to block the reverse current, when the device is warm during the main battery charge operation, it then causes the lithium battery to swell around two years of use and it causes the device to be inoperable. | Other | Since June 11, 2008, RS Medical has taken the following actions when the device was sent back through a consumer complaint, or when the device was returned for routine maintenance, or through the rental service. The firm has upgraded the blocking diodes with higher temperature rated diodes, replaced the Lithium coin cell battery, and the software register bit has been changed to denote for subsequent automated testing that the diode upgrade/Lithium coin cell battery replacement has occurred. The firm reported that the device must be returned for the rechargeable NiMH battery replacement around two years of use, and the battery should not be replaced by patients.On 10/29/10, the firm began sending out URGENT NOTICE of VOLUNTARY RECALL letters to their customers. The firm will continue to send out the letters through 2011. The letter states the recall product as RS-4i Sequential Stimulator, Affected Serial Numbers 4i300000 to 4i467119. The customers are informed that "RS-4i Sequential Stimulators, only those within the affected serial number range indicated above, are susceptible to early expiration of the lithium coin cell battery that provides back-up power to the device memory when the stimulator is powered OFF. Customer experience indicates that the 3V lithium coin cell battery can be depleted as early as one third of its rated life. The RS-4i Stimulator will perform as intended until the lithium coin cell battery is fully depleted. When this happens, the Stimulator will turn OFF and fail to turn ON, even if it has been charged by battery charger connection to a standard electrical wall outlet".Customers are advised to make a schedule for a free Return and Repair of a RS-4i Sequential Stimulator and should call the Toll Free Service Line at 877-363-5753 (English) or 877-363-5754 (Spanish) from 6AM to 5PM (Pacific Time) Monday-Friday.If the customers have any other questions or to review additional information before scheduling a device |  | 166131 | 166, 131 units | Nationwide Distribution | Terminated | September 10, 2012 | 1552 | 166131
Z-0322-2011 | 56255 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator/Monitor |  |  | July 16, 2010 | 2010 | Philips Healthcare Inc. | Updated Instructions for Use- HeartStart MRx monitor/defibrillator. The update provides more detailed information on:" Default ECG Lead Behavior" Synchronized Cardioversion using External Paddles | Labeling design | Philips issued an Urgent Voluntary Medical Device Correction notification/Field Safety Notice and an Instructions for Use Addendum on July 20, 2010, to affected customers. The Notice identified the product, the problem, and the action customers should take. In the USA, the notification sent to all affected customers via UPS and confirmation of delivery would be tracked. Outside the USA, Philips representatives in the affected geographies would manage their mailings locally. Customers were instructed to review the HeartStart MRx Instructions for Use addendum and keep a copy with their equipment - Instructions for Use. Additionally, they were instructed to insure that all users are fully trained on the MRx and, in particular, ensure that they review and understand have reviewed and understand the additional detail provided in the notification.For questions regarding this recall call (800) 722-9377. |  | 41119 | 41,119 units | Worldwide Distribution - US Nationwide and the countries of: CANADAAFGHANISTANALBANIAARGENTINAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHRAINBELGIUMBOSNIA&HERZEGOVINA;BRAZILBRUNEI DARUSSALAMBULGARIACHILECHINACOLOMBIACYPRUSCZECH REPUBLICDENMARKECUADOREGYPTESTONIAFINLANDFRANCEGERMANYGREECEGUADELOUPEHONG KONGHUNGARYICELANDINDIAINDONESIAIRELANDISRAELITALYJAPANJORDANKENYAKUWAITLATVIALEBANONLIBYAN ARAB JAMAHIRILITHUANIALUXEMBOURGMACEDONIAMALAYSIAMEXICOMOROCCONAMIBIANETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANORWAYOMANPAKISTANPALESTINEPANAMAPAPUA NEW GUINEAPERUPHILIPPINESPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIASAUDI ARABIASERBIA&MONTE; NEGROSINGAPORESLOVENIASOUTH AFRICASOUTH KOREASPAINSRI LANKASWEDENSWITZERLANDTAIWANTHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYUGANDAUKRAINEUNITED ARAB EMIRATESUNITED KINGDOMVENEZUELAVIET NAMYEMEN | Terminated | September 05, 2012 | 782 | 41119
Z-0328-2011 | 56628 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker M1 Ambulance cot, Model 6100, Stryker Medical, Kalamazoo, MI. |  |  | May 15, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The firm has received 42 complaints regarding the head end lock rod, head end slide tube, brushing, and upper cross tube failures. | Device Design | Stryker Medical sent an Urgent Medical Device Alert letter dated May 15, 2010 to all customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter stated that there is a possibility for fatigue related failures of certain components of the cot and that a Stryker representative will visit the firm to correct the device. A representative from Stryker field service would contact the customer to perform the device corrections, free of charge.Immediate Action required by the customer:Locate the ambulance cots listed in this notice.Return the enclosed post card to confirm receipt of this notification.If any of the cots listed in the letter were loaned or sold the customer was to forward a copy of the notice to the new users and advise Stryker of their new location.If the customer disposed of any of the affected stretchers and they were no longer in use, customers were to advise Stryker of their obsolescence by providing the serial numbersFor any questions or comments regarding this issue, please call (800) 869-0770, ext. 6902. |  | 7314 | 7,314 | Worldlwide Distribution - US including AZ, CA, CO, FL, LA, MS, NC, NJ, OH, PA, and TN and the countries of Australia, China, France, Germany, Hong Kong, Italy, Japan, New South Wales, Poland, Portugal, Spain, Sweden, Switzerland, and the United Kingdom. | Terminated | February 07, 2012 | 633 | 7314
Z-0333-2011 | 55301 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Excel Care ES Bariatric Bed, P 610B, Hill-Rom, Batesville, IN. |  |  | March 05, 2010 | 2010 | Hill-Rom, Inc. | The firm received complaints regarding the siderails not operating correctly or becoming inoperable. The siderails could experience false latching, unintentional lowering or become inoperable. There have been reports of 5 patient falls and 2 injuries to caregivers. | Device Design | Hill-Rom sent out "Urgent Medical Device Correction" letters dated March 5, 2010. The letters described the product, problem and actions to be taken by the customers. The letters instructed the customers to locate the beds and examine the siderails for any visible damage (cracks, bending, loose parts) and conduct a physical assessment of all four siderails on each bed to determine if the sideral could be placed in all positions. If the siderails were found to be inoperable, the bed was to be taken out of service until a qualified Hill-Rom technician can make any necessary correction. The customers were instructed to contact Hill-Rom to make arrangements and inform them of all beds that they have loaned, sold, no longer use or disposed of at 800-445-3720, option 2. The letter states the firm is evaluating a design change to the siderail.If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720, option 2. |  | 1406 | 1,406 beds both models. | Worldwide distribution: USA states including: AK, FL, HI, IN, KS, LA, MS, OK, PA, TN, VT, and WA and countries including: Canada, Mexico, Angola, Australia, Saudi Arabia, Qatar, and United Emirates. | Terminated | October 06, 2014 | 1676 | 1406
Z-0329-2011 | 55840 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | Titan XL Handpiece Infrared Therapeutic Heating Lamp, Manufactured by Cutera Inc, Brisbane, CA |  |  | May 08, 2010 | 2010 | Cutera Inc | Risk of burns-- under certain conditions, devices containing a gold plated end block could produce energy above what is stated on the device setting. | Device Design | Cutera sent Certified letters sent to all consignees on May 8, 2010. The letter identified the product, the problem, and the action the consignees should take. Consigness were instructed to immediately stop using the handpiece and complete the attached Recall Tracker Form. A second letter was sent to customers who received a replacement handpiecce but did not return the affected devices. All affected products were to be exchanged for new or reworked devices.For questions contact customer service at (888) 428-8372, ext 5561. |  | 1344 | 1,344 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OK, OR, PA, OH, SC, TN, TX, UT, VA, VT, WA, WI, WV, AND WY and the countries of Argentina, Australia, Austria, Barbados, Belgium, Bermuda, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cypress, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Lle de la Reunion, Mexico, New Zealand, Phillipines, Poland, Puerto Rico, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Thailand, Turkey, UAE, and Ukraine. | Terminated | December 28, 2010 | 234 | 1344
Z-0330-2011 | 56800 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | External Compression Device; 1 unit per package |  |  | August 30, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Variations in the outer diameter of the shaft of the External Compression Device and the inner diameter of the Nail Holding Screw may make assembly of these 2 instruments not possible. | Other | Stryker sent Recall notification letters by Federal Express on September 9, 2010 to all Stryker branches/agencies and Risk Managers, Chief of Orthopaedics of affected hospitals and Surgeons on September 10, 2010, that used the product. The letter identified the product, the problem, and the action that customers should take. Customers were instructed to:Examine their inventory and hospital locations to identify the affected product and return it to their branch or agency warehouse for reconciliation.Reconcile all product on the attached Product Recall Acknowledgment Form and fax a copy to (201) 831-6069 within 5 days of receipt of notice.A Stryker representative would contact the customer with details for return of the product.For any questions regarding this recall call (201) 972-2100. |  | 11 | 11 | Nationwide Distribution including AZ, CA, CO, FL, GA, IL, KY, MD, OH, OK, and TX | Terminated | August 21, 2012 | 722 | 33
Z-0335-2011 | 57024 | No | LKK |  |  | General Hospital | PMA | Medtronic Intrathecal Catheter Spinal Segment Revision Kit, 8598A. Method of Sterilization: Ethylene Oxide. Manufacturer Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. The Medtronic Model 8598A Spinal Segment Revision Kit is used when a revision to the spinal segment of a Model 8731 or Model 8731SC catheter is required. The catheter is part of an infusion system that stores and delivers parenteral drugs to the intrathecal space. The implanted infusion system components consist of a Medtronic pump and a Model 8731 or Model 8731SC catheter. The catheter connects to the pump at the catheter port. |  |  | September 29, 2010 | 2010 | Medtronic, Inc. - Neuromodulation | Medtronic is recalling specific lots of Model 8731SC intrathecal catheter kits and Model 8598A intrathecal catheter spinal revision kits because there is the potential that the 11.4 cm introducer needles within these kits may exceed USP requirements for bacterial endotoxin. | Material/Component Contamination | Medtronica sent an "Urgent: Medical Device Recall" letter, dated September 2010, was sent to medical facilities/customers beginning September 29, 2010. The letter described the issue, identified affected product, requested that unused inventory be segregated and returned to Medtronic Neuromodulation per the enclosed Product Return Instructions form. A Recall Reply Card was also asked to be returned. Additionally an "Urgent: Medical Device Recall" letter was addressed to Healthcare Providers/customers to inform them of the recall of the Catheter and Revision Kits, identified patients by means of registration information, Informed them of the issue with the needle inside the kits, and requested they stop using the specific lots and to monitor any patients for symptoms. If you have any questions or comments, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 420 | 420 | AL, CA, CO, CT, FL, GA, HI, ID. IL, IN, KS, KY, MA, MD, MI, MN, MO, MT, NC, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, VA, WA, WI, WV, WY. OUS to include: Argentina, Austria, Belgium, Canada, Chile, Czech Republic, France, Germany, Greece, Hungary, Ireland, Israel, Italy, San Marino, Jordan, Lithuania, Netherlands, Norway, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, United Arab Emirates, United Kingdom. | Terminated | December 10, 2011 | 437 | 1743
Z-0337-2011 | 56973 | Yes | NIM |  |  | Cardiovascular | PMA | RX ACCULINK Carotid Stent System 9 X 20mm, Part Number: 1011341-20, Lot Number: 0071361 |  |  | September 10, 2010 | 2010 | Abbott Vascular | The recall was initiated because Abbott vascular has discovered through internal testing that the affected lots may not meet the firm's quality specifications for catheter shaft tensile strength. If the affected device is used it may result in acute stent deployment difficulties and subsequent intervention. | Pending | Abbott Vascular sent an URGENT DEVICE RECALL letter dated September 10, 2010, to all consignees. The letter identified the product, the problem, and the action to be taken. The Sales Representatives were to personally visit each account to deliver the recall letter and assist in identifying and returning anyunused devices to Abbott Vascular. Also, Abbott Vascular Regulatory Compliance initiated contact with the Risk Management department at each account by phone to verify their mailing address and provide a courtesy copy of the recall letter.The consignees were instructed to work with their local account representative to review their inventory, complete the Recall Effectiveness Check and return identified products to Abbott Vascular.For questions regarding this recall call (951) 914-3324. |  | 42 | 42 units total for all products | Nationwide Distribution including AL, CA, CT, DC, FL, GA, IA, IL, IN, MA, ME, MI, MN, MO, OH, OK PA, TX, VA, and WI | Terminated | November 16, 2010 | 67 | 42
Z-0343-2011 | 56976 | No | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus(R) CS Rotors, catalog number CROTOR, lot number NH42-251-09 |  |  | December 07, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Low frequency defect in the gel filling of this lot that may result in inadequate separation of cells in centrifugation. | Process control | The recalling firm, SIEMENS, issued an "Urgent Field Safety Notice" dated December 2009 to all affected customers. The notice described the product, problem and the action to be taken by the customers. The customers were instructed to review the situation with their laboratory director to determine the need to evaluate previously reported test results; immediately examine their CROTOR inventory and discard all bags that are from lot NH42-251-09 including rotors that were not stored in their original packaging (bag); contact Siemens Customer Service Center at 800-241-0420 to request additional no-charge replacements (Note: A "no-charge" order has already been placed so that one bag of replacement rotors would be recieved shortly); forward this notice to anyone they may have distributed this product, and complete and return the FIELD CORRECTION EFFECTIVENESS CHECK via fax at 302-631-8467.If you have any questions regarding this information, please contact the Siemens Technical Solution Center at 800-405-6473. |  | 330 | 330 bags | Worldwide distribution: USA including states of AK, AZ, AR, CA, FL, GA, HI, IA, IL, IN, KY, LA, MI, MN, MO, MS, MT, NC, ND, NH, NM, NY, OH, TX, UT, WA, WI, and WV and country of Canada | Terminated | November 10, 2010 | 338 | 330
Z-0344-2011 | 56863 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab/Mic ASCII Add-on Versions 4.0.1.14, 4.0.1.15, and 4.0.1.16. |  |  | February 16, 2010 | 2010 | SCC Soft Computer | When a test is being verified, the wrong (expired) ranges may be associated to that test result, causing an incorrect flag to be posted for that test result. | Software design | Clients were notified via the proprietary communication software on 02/16/2010. The Notification Verbiage identified the affected product and informed clients of the potential hazard, root cause, corrective action, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software. |  | 31 | 31 | Nationwide Distribution - US and Puerto Rico. | Terminated | July 17, 2012 | 882 | 31
Z-0341-2011 | 56751 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | The box label states: "ATF 40 Fast Start Kit 9108491...STERILE...FRESENIUS KABI...Fresenius Kabi AG 61345 Bad Hamburg Germany...Federal (USA) Law restricts this device to sale by or on order of a physician". |  |  | March 29, 2010 | 2010 | Fresenius Kabi, LLC | Using AT1 Autotransfusion Set from specific batches in combination with the auto transfusion device CATS Continuous Autotransfusion has led to insufficient suction of shed blood from the collection reservoir. The results in the failure message "alarm blood flow" and the CATS device stops in a secure mode. | Other | Fresenius Kabi sent out a Technical Notification letter dated March 29, 2010, to one distributor located in Ann Arbor, Michigan. The distributor sent the technical notification with a cover letter to each of their customer's that had ordered AT1's, ATF120's and ATF40's over the last year. The letter identified the product, the problem, and the action to be taken. Customers were notified that the autotransfusion prime and / or washing procedure can be continued after alarm blood flow by following Alternative procedures. The Technical Notification informs the distributor about the problem of the AT1 (item #9005101), but also mentions the other two products ATF120 (item #9108501) and ATF40 (item #9108491) which have no known problem.The distributor is advised to contact Fresenius Kabi, LLC for any question regarding the Technical Notification.For questions regarding this recall call (425) 242-2074. |  | 461 | 17, 444 units distributed in US | Worldwide Distribution - USA and the countries of Australia,Austria, Brazil, Canada, China, Denmark, France, Germany, Hong Kong, India, Italy, Japan, Malaysia, Netherlands, Phillipines, Portugal, Switzerland, Thailand, UK, and Vietnam. | Terminated | June 15, 2011 | 443 | 1444
Z-0285-2011 | 56760 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK 2 Gram Negative Susceptibility Card, AST-N107, REF 22 265, For In Vitro Diagnostic Use, 20 cards/box. The firm names on the label are bioMerieux, Inc., Durham, NC, and bioMerieux sa, 69280 Marcy-l'Etoile France. |  |  | September 16, 2010 | 2010 | Biomerieux Inc | AMENDED REASON FOR RECALL: The piperacillin/tazobactam drug on the antimicrobial susceptibility cards is providing false susceptible and false resistant results for E. coli, as well as an additional 7 organisms; and the same drug is providing false resistant results for K. pneumoniae. | Nonconforming Material/Component | bioMerieux issued a "Product Correction Notice" letter dated September 13, 2010, to their entire customer base (regardless of whether they ever ordered the recalled cards). The letters explained the reason for the recall and instructed the account to confirm all results for E. coli using a second method, and for K. pneumoniae, to confirm with a second method only for resistant results. These alternate test confirmations are to continue until further notice. The customers were to complete and return the enclosed Acknowledgement Form via fax at (314) 731-8689 or (919) 620-6350 which indicate that the notice has been read, the instructions provided in the "Actions Required" section were followed, and notes if they have received any reports of illness or injury related to this issue. All future shipments will include an orange multilanguage insert to go into the boxes reminding customers they must perform an alternate test for TZP/E. coli combinations and that customers must also perform an alternate method prior to reporting TZP results when a resistant result is obtained for the TZP/K. pneumoniae. Additionally, this recall letter will accompany shipments of instrument software kits for the new customers.The recalling firm issued a press release on 10/27/10 via the Associated Press notifying the public of the recall.The recalling firm expanded their recall to add 7 more organisms to the list of organism testing that were to be suppressed. As a result the recalling firm issued a letter dated 3/30/11 via Fed Ex on 4/4/11 flagged "URGENT PRODUCT CORRECTION NOTICE (EXPANDED RECALL), VITEK 2 Piperacillin/Tazobactam Test, Document #9300786." The letter describes the issues, how the issues impact the test results, describes the product changes that are going to take place, and provides instructions on required actions to be taken.The product changes will involve the eventual changeover to new cards that will be created which do not contain Piperacillin/Taz |  | 90 | Product not distributed as of 9/10/10 | Worldwide Distribution: USA (all states), and the countries of:SUBSIDIARIES - Argentina, Canada, Switzerland, Chile, Colombia, Germany, Spain, Great Britain, Hong Kong, Japan, Korea, Mexico, Poland, Brazil, Sweden, Singapore, and Thailand. DISTRIBUTORS - Dutch Antilles, Australia, Bermuda, Uruguay, Costa Rica, Dominican Republic, Ecuador, France, Guatemala, Honduras, Indonesia, Israel, Italy, Singapore, Malaysia, Nicaragua, The Netherlands, Panama, Peru, Philipines, Paraguay, El Salvador, Venice, and Viet Nam. | Terminated | May 08, 2012 | 600 | 5709147
Z-0349-2011 | 51883 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | ViewSonic 19" flat panel LCD display Model VG930M which is used as part of the Welch Allyn Acuity Central Monitoring System. The affected serial numbers are those less than QC2082880001. |  |  | May 01, 2009 | 2009 | Welch Allyn Protocol, Inc | The 19" Model VG930M ViewSonic flat panel LCD display may have faulty capacitors at certain locations in the power supply section. The faulty capacitors may cause a loss of power to the display resulting in the display screen going dark and loss of both visual and secondary patient alarms. | Other | On May 1, 2009, Welch Allyn Protocol, Inc. sent out recall notification to their consignees in via UPS and US Postal Certified Mail.The letter describes the recalled product as 19" model VGM930M ViewSonic flat panel LCD display with serial numbers less than QC2082880001. For the Welch Allyn service contract consignees/customers, the firm's Field Service Engineer will replace the LCD monitor for the consignees. The consignees could contact Technical Support at 1-800-289-2501.For the Warranty/non-Warranty contract consignees/customers-For the Welch Allyn service contract consignees, Stericycle Return Management Services (RMS) will manage all recall and recovery initiatives. Consignees/customers can call Welch Allyn Customer Care at 1-800-535-6663 for any questions or concern about this recall. |  | 637 | 637 units | Worldwide distribution: USA including states of: AK, AZ, CA, CO, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MD, MI, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SD, TX, UT, VT, WI, and WV; and countries including: AUSTRALIA, CANADA, ESTONIA, GERMANY, ICELAND, IRELAND, ITALY, JORDAN, MALAYSIA, POLAND, and RUSSIA. | Terminated | July 25, 2011 | 815 | 637
Z-0367-2011 | 55839 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Sterile Disposable Syringe Kit Recall, Catalog Number 150-FT-Q; Kit includes angiographic injector and syringe |  |  | May 17, 2010 | 2010 | Medrad Inc | Some kits have a hole in the packaging which could compromise the sterility of the components. | Packaging | MEDRAD issued an URGENT MEDICAL DEVICE RECALL letter dated May 17, 2010, to all affected customers. The letter identified the product, the problem, and the action the customer should take.Customers were instructed to identify the product with the affected lot numbers.Contact MEDRAD Customer Support at 1-800-633-7231 to make arrangements for the return of the product.Complete and sign the form attached to the letter and fax it back to MEDRAD at 412-406-0942.For any questions regarding this recall call (412) 767-2400, ext 6851. |  | 39448 | 39,448 syringe kits | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO MS, NC, ND, NE, NJ, NM, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, and WV and the countries of Belgium, China, Columbia, Germany, Denmark, France, Great Britain, Hungary, Ireland, India, Israel, Italy, Mexico, Netherlands, Sweden, and Singapore. | Terminated | November 17, 2010 | 184 | 39448
Z-0357-2011 | 53048 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific brand Neuro Renegade Hi-Flo Microcatheter, 150cm/20cm; Rx Only; |  |  | August 24, 2009 | 2009 | Boston Scientific Corporation | Sterility Loss-- The sterile barrier pouch seal (top seal) may be breached by the product, resulting in loss of sterility which may result in transmission of disease or infection. | Packaging process control | Boston Scientific issued an Urgent Medical Device Recall letter dated August 26, 2009 identifying the affected devices, the issue prompting the recall, and actions to be taken by customers.Customers were instructed to determine if they have affected product in their inventory, and if so, immediately quarantine and return to Boston Scientific.Customers can contact Boston Scientific at 510 624-1734. |  |  |  | Distribution Worldwide, including throughout the US. | Terminated | February 03, 2011 | 528
Z-0355-2011 | 56792 | device, vein stabilization | LBJ | Vein stabilizer. | General Hospital | General Hospital | 510(K) Exempt | Baxter INTERLINK System, Non-DEHP Catheter Extension Set Kit; 7.1" (18.0 cm), Vol. 0.85 mL, 1 Catheter Extension Set with injection site and male luer lock adapter with retractable collar, 1 VITAL-HOLD Medium Catheter Stabilization Device, 2 foam tape strips; an Rx sterile convenience kit for IV infusions; 50 units per case; Manufactured by an affiliate of Baxter Healthcare Corp., Deerfield, IL 60015 U.S.A., Made in Costa Rica; catalog 2N8221 |  |  | September 17, 2010 | 2010 | Baxter Healthcare Corp. | The removable paper liner covering the adhesive strip that is used to secure the catheter stabilization device to the patient is pyrogenic. The device is labeled as nonpyrogenic. | Component design/selection | Baxter issued Urgent Product Recall letters dated September 17, 2010 to the affected direct accounts. The accounts were informed that the paper lining covering the adhesive strip that is used to secure the catheter stabilization device to the patient was found to be pyrogenic. The accounts were requested to discontinue use and segregate the affected lots of products from inventory and call Baxter Healthcare Center for Service at 1-888-229-0001, Monday through Friday, 7:00 AM to 6:00 PM CT, to arrange for the return of the affected product for credit. The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter and the amount of each product being returned. If the accounts further distributed the products, they were requested to forward the communication to those accounts. Clinical questions regarding the communication are directed to the Medical Information Services at Baxter at 1-800-933-0303. Any other questions may be directed to the Center for One Baxter at 1-800-422-9837. |  | 12500 | 12,500 units | Nationwide Distribution: Throughout the US. | Terminated | October 04, 2011 | 382 | 172300
Z-0368-2011 | 56814 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE LightSpeed VCT Scanner System, models 5212920-300, 5212920-310, 5311595-10. Composed of a gantry, patient table, operator console, computer, and PDU and includes image acquisition hardware, image acquisition and reconstruction software, associated accessories and connections/interfaces to accessories. GE Healthcare, 3000 North Grandview, Waukesha, WI 53188 USA. |  |  | September 03, 2010 | 2010 | GE Healthcare, LLC | GE has become aware of a potential set of circumstances that could cause x-ray continuation during an unexpected table stop on certain LightSpeed VCT and BrightSpeed systems. | Software design | An Urgent Medical Device Correction letter dated September 2, 2010 was sent to consignees beginning September 3, 2010. This letter addressed the LightSpeed VCT systems. Another Urgent Medical Device Correction letter, dated September 24, 2010 was sent September 24, 2010 for the BrightSpeed Systems. The letters described the issue and provided safety instructions for the health care provider to be aware if the table should stop unexpectedly and the x-ray remains on to manually stop the scan until the devices software is upgraded.If you have any questions, please contact the US (800) 437-1171, Japan 0120-055-919, Korea 1544-6119, Australia/New Zealand 800 659 465, China 800-810-8188; for other countries, please contact your local GE Healthcare Services Representative. |  | 68 | 68 | Worldwide distribution: USA including states of: MI, NM, OH, TX, and Puerto Rico and countries including: Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Honduras, Hungary, India, Indonesia, Ireland, Islamic Republic of Iran, Israel, Italy, Jamaica, Japan, Kenya, Lebanon, Malaysia, Mauritania, Mexico, Morocco, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Pakistan, Paraguay, Philippines, Poland, Portugal, Republic of Korea, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Togo, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Viet Nam. | Terminated | May 15, 2012 | 620 | 1014
Z-0364-2011 | 56865 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SA BASE with SoftLab Version 3.1.4.x. |  |  | June 01, 2010 | 2010 | SCC Soft Computer | Functionality issue: For clients using the instrument to generate results on the interface menu, two different result records could be displayed under the same order and sequence numbers. This could cause incorrect results to be populated to a patient record. | Software design | Clients were notified via the proprietary communication software on 06/01/2010. The communication discussed the risk-to-health notification task, potential hazard, root cause, corrective action, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software. |  | 40 | 40 | Worldwide Distribution -- US, including states of CT, FL, IL, KY, GA, LA, IN, MS, NJ, NY, PA, OH, and TN and country of Canada. | Terminated | July 16, 2012 | 776 | 40
Z-0363-2011 | 56461 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Transport ProÂż Monitor with the CARESCAPE" Patient Data Module. |  |  | October 05, 2010 | 2010 | GE Healthcare, LLC | Transport Pro Monitor stops communication with the CARESCAPE Patient Data Module (PDM) after 414 days of continuous run time. This time will equate to different amounts of real time depending on how much the units is actually in service per day. Transport Pro contains an internal timer that is used to control the software and remind users to perform preventive maintenance. When this internal timer | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated September 17, 2010 to consignees. The letter was addressed to Hospital Administrator, Head of Biomedical Engineering and Nursing Manager. The letter described the product, Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Service representatives will updates all of the affected Transport Pros with PDM that were distributed.Customers may contact GE at 262-548-2731 about this correction. |  | 3256 | 3256 | Worldwide Distribution: Domestic in the states of AL, AZ, AR, CA,CT, CO, DC, FL, GA, ID, IL, IN, IA, KY, LA, MD, MA, ME, MI, MN, MS, MO, NE, NV, NJ, NM, NY, NC, OH, OK, PA, SC, TN, TX, VA, VT, WA, WI, and WY, and Internation in the countries of BRAZIL, BERMUDA, BELGIUM, AUSTRIA, AUSTRALIA, CANADA, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY GREECE, INDONESIA, IRELAND, ITALY, JAPAN, JORDAN, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MAURITIUS, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, REPUBLIC OF KOREA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD 7 TOBAGO, and the UNITED KINGDOM. | Terminated | May 17, 2012 | 590 | 3256
Z-0360-2011 | 56873 | No | LNX |  |  | General Hospital | 510(k) | SoftID.NET versions 2.0.0.6 - 2.0.0.19. |  |  | August 28, 2009 | 2009 | SCC Soft Computer | SCC Soft Computer issued a Correction for SoftID.NET versions 2.0.0.6 - 2.0.0.19. When using the Cancel functionality, if the collection list refreshes during the transaction and changes the position of the patient on the list, the wrong patient order can be cancelled. All affected clients have been notified. | Software design | SCC Soft Computer notified Customers on August 28, 2009, via the proprietary communication software entitled "Risk-to-Health Notification Task Verbiage" dated August 17, 2009. The notification described the product, problem and action to be taken by customer. The customers were instructed to acknowledge receipt of the notification; document a request to load the software, if they wish to receive the correction; document that the Alternative Working Solution will be used by the customers facility until the correction has been taken live.If you have any questions, contact (727) 789-0100. |  | 41 | 41 | Nationwide distribution: AR, CA, FL, IL, IN, KS,LA, MA, MI, MN, MO, NC, ND,NH, NJ, NY, OH, PA, RI, TN, TX, VA and WI. | Terminated | November 16, 2011 | 810 | 41
Z-0359-2011 | 56664 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Power supply boards on Dash 3000, 4000 and 5000 Monitors. |  |  | October 11, 2010 | 2010 | GE Healthcare, LLC | The power supply board in GE Healthcare Dash 3000, 4000 and 5000 patient monitors may fail due to insufficient solder of a component. Insufficient solder on the power supply board could result in excess EMC emissions, smoking of components and eventual failure of the power supply, tripping of the fuse and immediate shutdown resulting in a loss of monitoring that could delay recognition of a clinic | Nonconforming Material/Component | GE Healthcare issued an "Urgent Medical Device Correction" letter dated October 8, 2010, which described the loss of monitoring issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Customers were instructed to remove the units from use. Local service representatives will contact firms to schedule installation of replacement power supply boards.GE Healthcare can be contacted at 1-800-558-7044, Option 2, Technical Support, about this correction. |  | 137 | 137 | Worldwide Distribution: USA in the states of IL, NV, TX, and VA, and the countries of Australia, Brazil, China, Japan, South Korea, and Taiwan. | Terminated | May 29, 2012 | 596 | 137
Z-0361-2011 | 56940 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ENDOPATHÂż ETS-FLEX45 Articulating Endoscopic Linear Cutters, Model # ATS45. Manfuactured by Ethicon Endo-Survery, LLC; 475 Calle C; Guaynabo PR 00969. |  |  | January 13, 2009 | 2009 | Ethicon Endo-Surgery Inc | The mechanism that connects the articulation joint to the device shaft may cause the jaws of the cutters to remain closed and clamped down on tissue after the device is fired. If the jaws of the device remain clamped down on the tissue, there is a risk the surgery will need to be modified to remove the device. This may lead to an increase in procedure time, tissue manipulation, and a change in t | Other | Ethicon Endo-Surgery issued "Urgent Device Recall Event 1952" notification letters dated January 21, 2009 to their customers. They instructed their custormers to immediately discontinue use of the product. They are to complete and fax back the Business Reply Form. Affected product is to be returned with a copy of the completed Business Reply Form to Stericycle. Replacement product should be received within 30-60 days after return.Customers can contact Stericycle concerning this recall at 1-866-861-6921, reference Event 1952. |  | 4113 | 4,113 UNITS | Worldwide Distribution: USA only, in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI. IA. ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | May 10, 2012 | 1213 | 4113
Z-0362-2011 | 56872 | No | MMH | Automated blood grouping and antibody test system. | Hematology | Hematology | 510(k) | SoftBank Versions 23.1.2.x, 23.2.0.x, and 25.0.0.x with SoftReports. |  |  | June 22, 2010 | 2010 | SCC Soft Computer | SCC Soft Computer is issuing a Correction for SoftBank II Versions: 23.1.2.x, 23.2.0.x, and 25.0.0.x software applications. Functionality issue - For clients generating patient history reports using Patient&gt;Reports&gt;Batch Reports&gt;History Report option in SoftBank with SoftReports activated for patient reports, when the report is generated for more than one patient ALL of the results for the qualif | Software design | SCC Soft Computer notified Customers on June 22, 2010, via the proprietary communication software entitled "Risk-to-Health Notification Task Verbiage" dated June 11, 2010. The notification described the product, problem and action to be taken by customer. The customers were instructed to acknowledge receipt of the notification; to grant permission to load the required hot fix, and to make the manual correction. For the manual correction to the datapath for patient report layouts, your SoftBank Support Specialist will Coordinate those changes and your validation of affected environments.If you have any questions, contact (727) 789-0100. |  | 30 | 30 | Worldwide distribution: U.S. and country of Canada. | Terminated | September 05, 2012 | 806 | 30
Z-0381-2011 | 57052 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sorin OptiPack, Perfusion Tubing Set, Sterile EO, REF 084500900, Manufactured by Sorin Group USA, Inc. 14401 W. 65th Way Arvada, CO 80004 |  |  | October 07, 2010 | 2010 | Sorin Group USA, Inc. | Defective product packaging may compromise product sterility. | Packaging process control | Sorin Group sent an "URGENT: FIELD SAFETY NOTICE" letter dated October 7, 2010, to all customers. The letter described the product, problem and action to be taken by customer. The customers were instructed to check all affected inventory to exam for holes. Any damaged products were to be returned to Sorin. If no damage was found, product could be used. The customers were also instructed to contact Sorin Cardiopulmonary Customer Service at 1-800-650-2623 to return product and receive replacement; they were to share this information with all personnel who may use this product; and to complete and return the attached Customer Response Form via the self-addressed prepaid envelope provided or fax to (303) 467-6502 as soon as possible.For further information, contact 1-800-650-2623. |  | 86 | 86 | Nationwide distribution: CA, FL, IL, IN, KS, MD, MO, NE, NJ, NM, TN, TX, and VA. | Terminated | December 17, 2010 | 71 | 3021
Z-0366-2011 | 57001 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | BD Phoenix" ID/AST Panels, Item # PMIC/ID-107, Catalog number 448607, and labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, MD 21152 USA (800) 638-8663*** |  |  | May 14, 2010 | 2010 | Becton Dickinson &amp; Co. | The diagnostic reagent packaging may not be completely sealed, which may cause substrate degradation, and produce incorrect test results, including false positive and false negative results. | Packaging process control | BD Diagnostic Systems issued an "Urgent Product Recall" letter dated May 2010 to direct accounts, identifying the affected product, potential impact, and actions to be taken by the customer. Accounts were requested to respond by returning the enclosed reply form indicating the number of replacements needed for the affected product to be discarded.For assistance iwth replacements, contact BD Customer Service at 800 675-0908. For other inquiries, contact BD Technical Services at 1 800-638-8663. |  | 1675 | 1675 units | Nationwide Distribution: To direct accounts in CA, MD, NJ, PA. and WI. | Terminated | November 29, 2010 | 199 | 1675
Z-0395-2011 | 55999 | No | IXQ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | CIVCO Type-S Extension, REF MTIL6605, Reusable non-sterile extension for use with Universal Couchtop. The firm name on the label is CIVCO, Orange City, IA. |  |  | June 21, 2010 | 2010 | Med Tec Inc | Under certain conditions, the extensions can become detached from the main body of the couch. | Device Design | The recalling firm, CIVCO, issued an 'Urgent Field Safety Notice" dated June 15, 2010, via overnight mail on June 21, 2010, regarding the first issue for the three extension models. The letter described the product, problem, and actions to be taken by the customers. The customers were instructed to load the patients on the couch from the bottom of the couch and not from the superior end with the extension and do not let patients crawl to the top of the couch with their full upper body weight at the end of the couch; notify others within their organization that may have the extension, and complete and return the enclosed response form via fax: Attention: CIVCO MTL6600 Rectangular Extension at 877-271-2310. The letter informs the customer that they will be receiving a kit to replace the locking tab so that the extension will be less likely to come loose. Once the kit is received, the customers are ask to return the replaced part and sign a confirmation form for their records.CIVCO sent a second "Urgent Field Safety Notice" dated July 20, 2010 via overnight mail on July 21, 2010, and by phone on July 22, 2010, regarding the second issue which covered only MTIL6600 regarding the weight being dropped onto the corner of the rectangular extension. The letter described the product, problem and action to be taken by the customers. The customers were instructed to accept a one-piece Universal Couch Top replacement; sign the attached form stating the they will stop using the rectangular extension and return it to CIVCO or sign the form stating they have been advised of the recall and they will exercise the necessary precautions to prevent the event from occurring by always using the thumb screw lock and never allowing the patient to crawl on the couch top or extension, and to notify others within their organization that have the extension. The customer was informed that their current system would be replaced and sent back to the recalling firm. Please contac |  | 3 | 3 extensions | Worldwide distribution: USA including states of: CA, FL, IA OH, and WI; and countires of: United Kingdom, Australia, and Switzerland. | Terminated | April 17, 2012 | 666 | 30
Z-0365-2011 | 56867 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab GUI Version 4.0.1.0 - 4.0.1.16, 4.0.2.0 - 4.0.2.10, 4.0.3.0 - 4.0.3.13, and 4.0.4.0 - 4.0.4.5. |  |  | April 01, 2009 | 2009 | SCC Soft Computer | Clients using Patient Maintenance to perform Moves of Stays may send the wrong billing number during an ADT update to SoftWeb, SoftPath or SoftMedia. The wrong billing number can come from a different patient in Order Entry. Only Moves of Stays are affected. | Software design | Clients were notified via the proprietary communication software on 06/01/2009. This is a MANDATORY CORRECTION. The Risk-To-Health Notification Task gives the background on the issue, the potential hazard, root cause, corrective action, alternative working solution, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software. |  | 49 | 49 | Nationwide Distribution | Terminated | August 07, 2013 | 1589 | 49
Z-0392-2011 | 56988 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 4100/4100IQ GE Innova 4100 / 4100 IQ. , GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | August 06, 2009 | 2009 | GE Healthcare, LLC | GE Healthcare has become aware of a potential video splitter failure with certain Innova 2000 and 4100 X-ray systems that may impact patient safety.A potential video splitter failure inside the GE equipment cabinet could occur with no advance warning and thereby prevent x-ray image display. Should this occur, the systemcannot be recovered without service intervention.GE Healthcare has obse | Nonconforming Material/Component | GE Healthcare sent an " Urgent Medical Device Correction" letter dated July 9, 2009, to Hospital Administrators / Risk mangers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter identified the product, described the Safety Issue, Affected Product Details, Safety Instructions Product correction and Contact Information.Hospitals were instructed to establish emergency workflow procedures should a component stop operating during a patient examination. A GE Healthcare service representative would contact each customer and correct all affected systems.For questions regarding this recall call (262) 548-2731. |  | 43 | 43 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX UT, VA, WA, WV, and WI, and the countries of UNITED KINGDOM, TUNISIA, TAIWAN, PROVINCE OF CHINA, SPAIN, SOUTH AFRICA, SAUDI ARABIA, RUSSIAN D=FEDERATION, POLAND, NORWAY, NEW ZEALAND, NETHERLANDS, MEXICO, MARTINIQUE, LITHUANIA, LEBANON, REPUBLIC OF KOREA, JORDAN, JAPAN, ITALY, IRELAND, ISLAMIC IRAN, INDONESIA, INDIA, GREECE, GERMANY, FRANCE, FINLAND , EGYPT, CHILE, CANADA, BULGARIA, BELGIUM BELARUS, AUSTRALIA. | Terminated | May 17, 2012 | 1015 | 408
Z-0403-2011 | 56876 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Vortex TR, AngioDynamics, Inc., One Horizon Way, Manchester, GA 31816 |  |  | September 27, 2010 | 2010 | Angiodynamics, INC | Labeling Error; The labeling for the specific products incorrectly identified that the packaged device kit contains a Lifeguard Safety Infusion Set. | Other | AngioDynamics issued a Notification of Medical Device Correction letter dated September 27, 2010 to customers. The notification served to make the customer aware of the mislabeling. Customers were instructed not to return the affected products.Customers can contact AngioDynamics at 1 800 772-6446 about this correction. |  |  | see line one. | Nationwide Distribution : Throughout the US. | Terminated | June 21, 2011 | 267
Z-0410-2011 | 56967 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 4100/4100IQ GE Innova 4100 / 4100 IQ. , GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | August 05, 2009 | 2009 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue regarding Innova systems with GE Fluoro UPS 20KVA that may impact patient safety.GE Fluoro UPS 20KVA (Uninterrupted Power Supply) is used to maintain fluoroscopy availability for approximately 5 minutes should there be a main power failure. Once power is restored, the equipment power will be returned to the main hospital power.UPS fluoro ba | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter on August 5, 2009 to all Hospital Administrators / Risk Mangers, Mangers of Radiology / Cardiology and Radiologists / Cardiologists. The letter was dated August 4, 2009, and described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For questions regarding this recall call (262) 548-2731. |  | 117 | 117 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK ,OR, PA, SC, TN, TX, VT, VA, WA, WV, and WI and the countries of ANGOLA, BELARUS, BULGARIA, CANADA, EGYPT, FRANCE, GERMANY, HUNGARY, IRELAND, MEXICO, ITALY, NIGERIA, NORWAY, POLAND, PORTUGAL, RUSSIAN FEDERATION, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, TUNISIA, TURKEY, and the UNITED KINGDOM | Terminated | January 16, 2014 | 1625 | 578
Z-0401-2011 | 57015 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Millar Disposable Angiographic Catheter (Left Heart), manufactured for Millar Instruments, Inc 6001-A Gulf Freeway, Houston, Texas 77023 |  |  | October 14, 2010 | 2010 | Millar Instruments, Inc | The sterile packaging was found to be compromised when tested after shipping conditions. This may result in infection or transmission of disease. A test for transportation of sterile package was conducted and resulted in several samples failing. | Packaging | Millar issued a Recall Report letter on October 14, 2010 to consignees identifying the affected product, sterility issue, and actions to be taken. Customers are to return all unused catheters to Millar for replacement. Customers can contact Millar about this recall at 832-667-7131. |  | 12 | 12 Units | Worldwide Distribution: USA, including the states of MA, OH, PA, and D.C., and the countries of Canada and Germany. | Terminated | February 15, 2011 | 124 | 12
Z-0400-2011 | 57005 | No | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | i-STAT ACT Kaolin Cartridges |  |  | October 06, 2010 | 2010 | Abbott Point Of Care Inc. | Certain lots of i-STAT Celite ACT and Kaolin ACT cartridges may be difficult to fill or will not fill when attempting to transfer blood to the sample entry port of the cartridge. | Other | Abbott Point of Care, Inc. sent an URGENT RECALL NOTICE dated September 2010 on October 6, 2010 by Federal Express to all US customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of cartridges from the identified lots. If customers did not currently have replacement cartridges then they should continue to use the cartridge lot until their facility receives replacement material.Upon receipt of replacement material, customers were to discontinue use of and return the affected cartridges and switch to the replacement material. All unused cartridges will be credited and will be returned.If you have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020, Option 1 or your Abbott Point of Care representative. |  | 95325 | 95,325 cartridges | Worldwide Distribution | Terminated | August 22, 2012 | 686 | 144350
Z-0398-2011 | 57029 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | SKYLight Gamma Camera Gearbox brake, a component of the SKYLight Gamma Cameras 3/8" SPECT, Model number 4535 600 66661; and 5/8" SPECT, Model number 4535 602 20381, manufactured by Philips Medical Systems (Cleveland), Cleveland OH. |  |  | October 15, 2010 | 2010 | Philips Medical Systems | Detector arm assembly may slide to hardware limit, potentially causing impact. | Component change control | Firm will notify consignees via Urgent - Medical Device Correction, dated Nov 3, 2010, via FedEx. The letter identified the affected product, described the problem, the hazard involved, how to identify affected products, actions to be taken, and actions planned by Philips. Customers are to contact their Field Safety Engineer immediately to schedule an inspection of their system. If customers suspect that the identified problem has occurred with their device, they should immediately stop using the system until Philips has confirmed it as safe to use. If customers need further information concerning this issue, they are to contact their local Philips representative. |  | 24 | 24 units | Worldwide Distribution -- USA and the Netherlands. | Terminated | March 21, 2012 | 523 | 24
Z-0412-2011 | 52108 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips EasyVision MM workstation with software R10.2 and R11.1 |  |  | January 30, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | Review by firm of Philips EasyVision MM workstations found that to assure compliance with 21 CFR 900.12(c)(5) corrected software was needed so that all sites are in compliance. | Other | Philips Medical Systems North America Co. sent an URGENT-Field Safety Notice dated January 30, 2009, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were advised to ascertain the quality of a hard or soft copy explicity before further distribution to make sure that no part of the tissue is covered by image identification information. Philips will issue the Field Change Order (FCO) 83000129 and release a level on EasyVision MM R10.2 L6 and EasyVision MM R11.1 L3 to resolve the issue.Customers could contact Philips Healthcare Call Center at 1-800-722-9377, #2, #3 and Reference FCO 83000129 for any questions. |  | 11 | 11 units in US | Nationwide Distribution including DC, FL, MI, MO, NM, NY, OH, AND TX | Terminated | November 19, 2010 | 658 | 11
Z-0415-2011 | 56978 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Rascal LiteWay 214 manufactured by Giant Manufacturing, Taiwan. |  |  | September 27, 2010 | 2010 | Electric Mobility Corp | Scooter seat bases were found to have insufficient welding which may lead to the seat base becoming loose and causing the seat to wobble. This may result in complete failure of the seat base, causing the user to fall off the vehicle. | Device Design | The Rascal Company (Electric Mobility Corporation) sent Urgent Medical Device Recall Notice Scooter Seat Base letters dated October 7, 2010 and November 11, 2010 letter to their affected retail customers and affected dealers. The letter instructs the reader that the seat base may become loose, causing the seat to wobble and if not replaced the wobble could lead to a complete failure of the seat base and cause an individual to fall off their vehicle and become injured. The letters also instruct the reader to replace seat bases immediately. Respone forms are attached. Seats with new seat bases will be shipped with return shipping labels to return the original base to the factory. Service representatives will replace the seat base for customers.Customers can contact the Rascal Company about this recall at 1-800-662-4548. |  | 172 | 172 units | Worldwide Distribution: Throughout the US and Canada | Terminated | April 03, 2014 | 1284 | 653
Z-0346-2011 | 56874 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | ATM100, LB-AP-M0001 |  |  | March 15, 2010 | 2010 | Micromedics, Inc. | Micromedics, Inc is conducting a product recall on a number of medical device products because of weak seals of the sterile pouches potentially resulting in a non-sterile product which may cause transmission of disease or infection. | Packaging process control | Micromedics issued a Product Recall letter dated March 15, 2010 to customers, describing the product being recalled, the product sterility issue, and actions to be taken. The letter advised customers to contact Customer service to return any affected product to Micromedics for replacement.Customers can contact Micromedics at 1-800-624-5662 or 651-452-1977 ext. 226 concerning this recall. |  | 620 | 620 (800 pouches) | Nationwide Distribution in the states of CA, CT, DE, FL, GA, ID, IL, IN, MD, MA, MI, MN, NJ, NY, NC, PA, TX, UT, and WI.FRANCE, NETHERLANDS, BELGIUM, UNITED KINGDOM, CANADA, ISRAEL, PERU, SPAIN, IRELAND. | Terminated | June 21, 2011 | 463 | 4880
Z-0419-2011 | 57041 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew FAST FIX 360, STRAIGHT NEEDLE |  |  | October 14, 2010 | 2010 | Smith &amp; Nephew, Inc. Endoscopy Division | FasT-Fix 360 Needle Delivery System may not allow the T-2 (anchor) to advance for implantation. | Device Design | Smith &amp; Nephew notified customers by an Urgent - Product Recall 1st Notification - Urgent letter sent via Federal Express October 14, 2010. The letter identified the affected product and explained the problem. The letter asked customers to call for a Return Authorization (RA) Number. Customers are to complete the form and return of the affected product along with a copy of the letter to the address provided being sure to reference the RA number. Questions were to be directed to Cindy Burns at 508-261-3655. Consignees out side the US were notified by e-mail. |  | 53 | 53 units | Worldwide Distribution - USA, Dubai, and Great Britain. | Terminated | May 09, 2012 | 573 | 804
Z-0418-2011 | 57094 | No | LTI |  |  | Gastroenterology/Urology | PMA | LAP-BAND Adjustable Gastric Band System; Product codes: |  |  | September 16, 2010 | 2010 | Allergan | In a LAP-BAND System adjustment, failure to inject the needle perpendicular to the access port septum, as described in Directions for Use, may lead to port damage, resulting in subsequent port leakage. Leakage results in deflation of the LAP-BAND to its widest, open position requiring a procedure to replace the port. The patient may have symptoms of reduced satiety and increased appetite. Stalling | Pending | Allergan sent an URGENT FIELD CORRECTIVE ACTION letter dated September 15, 2010, to all affected customers. The recall communication was initiated on 09/16/2010 when the firm begin forwarding Customer letters with attached current Directions for Use distributed via Federal Express, with tracking verification provided by Federal Express, to the direct consignees.The Customer letters informed consignees of the products affected, description of the Issue, clinical implications and firm recommendations.Customers were instructed to follow the Directions for Use for inserting the needle. Customers were instructed to report any malfunction or adverse event related to the device to Allergan Product Support at 1-800-624-4261 ext. 5972 or 1-805-961-5972 and FDA's Medwatch Program (1-800-FDA-1088 or www.fda.gov/medwatch). For questions, consignees were instructed to contact their Allergan field representative or contact Allergan Product Support at 1-800-624-4261. |  | 152000 | Approx. 152,000 | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, AND WY | Terminated | December 22, 2010 | 97 | 152000
Z-0414-2011 | 57039 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Clinical Information Portfolio (ICIP), Clinical Management System |  |  | October 19, 2010 | 2010 | Philips Healthcare Inc. | Software: Scheduling function of the Philips Intellivue Clinical Information Portfolio Releases D.00 through D.03. Edits to a scheduled even frequency order may result in pending interventions not being generated after the first 24 hours. | Software design | Philips Healthcare issued an Urgent Medical Device Correction letter dated October 15, 2010 to customers. Customers are asked to follow the guidelines in the "Action to be taken by Customer/User" section of the Urgent Medical Device Correction Notice until they receive their software correction. The correction will be in the form of a software upgrade and will be performed by Philips service personnel. Contact Philips Customer Care Center with questions or concerns about this correction at 1-800-722-9377 |  | 200 | 200 units | Worldwide Distribution: USA, including states of CA, DC, FL, IL, IN, KY, MA, MD, ME, NJ, NY, OK, RI, TX, VA, VT, and WV; and the countries of: AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, IRELAND, ITALY, KENYA, LATVIA, MALAYSIA, NETHERLANDS, NORWAY, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, UNITED ARAB EMIRATES, and the UNITED KINGDOM. | Terminated | August 01, 2017 | 2478 | 200
Z-0411-2011 | 57047 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | smith &amp; nephew TRIGEN SURESHOT Distal Targeting System, Trauma Interface, V2.0.2, Trauma Launcher V1.01.1 |  |  | October 15, 2010 | 2010 | Smith &amp; Nephew Inc | Targeting software has been found to provide erroneous targeting when used with the 13 mm TAN/FAN Left nails. | Other | Smith &amp; Nephew Inc. sent an URGENT MEDICAL DEVICE CORRECTION letter by FedEx on October 15, 2010, to all customers. The letter was dated October 4, 2010, and identified the product, the problem, and the action to be taken by the customers. Customers were instructed to perform a software upgrade and complete the form provided and return by fax to (901) 566-7835. For questions regarding this recall call (901) 399-5441. |  | 445 | 445 units | Worldwide Distribution - USA including AL, AR, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, MS, NC, NJ, NM, OH, PA, TN, TX, AND WA and the countries of Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Italy, Japan, Netherlands, Norway, South Africa, Spain, Sweden, Switzerland, Turkey, and the UK. | Terminated | February 17, 2011 | 125 | 445
Z-0423-2011 | 57010 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BARDÂż VentrioÂż Large Oval Hernia Patch 5.4" x 7.0" |  |  | October 15, 2010 | 2010 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: Hernia Patch may be a different size than on the product label | Packaging process control | Davol notified accounts by an Urgent: Medical Device Recall letter on 10/15/10 via overnight delivery. The letter stated that the product maybe a different size than indicated on the product label. It also identified the affected product. Accounts are requested to examine their inventory, remove, and return recalled product. Customers are to notify any of their customers that they may have further distributed the affected product to by including a copy of the recall notice. Customers should contact Davol Customer Service 1-800-556-6275 for instructions on how to return the affected product. Customers are also requested to complete and fax back the attached fax sheet to 1-401-825-8753. If there are any additional questions, customers should contact Davol Customer Service Department at 1-800-556-6275 or C.R. Bard's Medical Service Support Department at 1-800-562-0027. |  | 184 | 184 units | Nationwide Distribution | Terminated | October 27, 2011 | 377 | 406
Z-0426-2011 | 56380 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Signa Ovation, 0.35T MR Systems (MFO 1-4). |  |  | August 23, 2010 | 2010 | GE Healthcare, LLC | If the lateral table adjustment crank handle breaks off, it has sufficient attractive force to the magnet and may result in patient injury. | Nonconforming Material/Component | Consignees were sent on 8/23/10 a GE "Urgent Medical Device Correction" letter dated August 20, 2010. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Managers and Radiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, and Contact Information. Customers were asked to discontinue use of their MR system and to contact their local service representative, if the lateral table adjustment crank handle on the table becomes loose. |  | 202 | 202 | Worldwide Distribution -- USA, including states of AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, LA, KS, KY, LA, MD, MA, MI, MN, MS,MO,MT, NV, NH, NJ, NY, NC, OH, OK, OR, PA, PR, SC, SD, TN, TX, VA, WA, WV, and WI and countries of VENEZUELA, ECUADOR, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, SPAIN, RUSSIAN FEDERATION, REPUBLIC OF KOREA, POLAND, PHILIPPINE,S OMAN, MEXICO, KUWAIT, JAPAN, ITALY, HUNGARY, GERMANY, EGYPT, DENMARK, CHILE, CHINA, and BRASIL. | Terminated | July 13, 2012 | 690 | 378
Z-0432-2011 | 49384 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart FRx Defibrillator, Automated External Defibrillator (AED), catalog number 861304. |  |  | August 07, 2008 | 2008 | Philips Medical Systems | The affected device may not deliver appropriate defibrillation therapy. | Employee error | Philips notified the customer with the affected device by phone on August 8, 2008 . A replacement device was shipped to the customer on August 11, 2008. |  | 1 | 1 unit | Nationwide Distribution: to one account in Arkansas. | Terminated | December 01, 2010 | 846 | 1
Z-0429-2011 | 56017 | No | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | diaDexus PLAC Test ELISA Kit; Manufactured and distributed by diaDexus , Inc., South San Francisco, CA; Model Part Number: 90123 |  |  | March 15, 2010 | 2010 | DiaDexus, Inc | The reason for the recall is that Real Time Stability Test program showed the kit lot failed stability prior to expiration, affecting samples with results&gt; 500 ng/ml. Results up to 500 ng/ml exhibit passing stability results and are not impacted by this issue. | Other | DiaDexus notified all customers by phone, including distributor's customers, beginning on March 15, 2010. Customers were told to re-test samples that were tested from 2/1/2010 through 2/28/2010 which have values above 500 ng/ml with a new lot of product.For questions regarding this recall call (650) 246-6477. |  | 239 | 239 kits | Nationwide Distribution - US including NJ, OH, KS, IN, FL, CA, UT, OK, PA, WA and MI. | Terminated | January 11, 2011 | 302 | 239
Z-0437-2011 | 56215 | No | KYI | Wrist joint carpal trapezium polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | The Ascension PyroHemiSphere (aka NuGrip), Model #30-L NUG-443-30L, Catalog #NUG-443-30L, Lot #10-0731, is a single-use uncemented, one-component prosthesis for the basal thumb joint. |  |  | July 12, 2010 | 2010 | Ascension Orthopedics, Inc | Two different product lots were found to be incorrectly labeled and placed into the incorrect box. | Mixed-up of materials/components | Ascension Orthopedics notified by phone on June 25, 2010 and sent a letter dated July 22, 2010, to all customers. The customers were instructed to return the devices to the firm.Should you have any questions or concerns please feel free to contact (512) 836-5001 ext 1585. |  | 9 | 9 | Worldwide distribution: USA including states of: IL, IN, MD, OH, and TN; and countries of : Puerto Rico and Netherlands. | Terminated | June 30, 2011 | 353 | 19
Z-0431-2011 | 56416 | No | CDZ | Testosterone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Sex Hormone Binding Globulin (SHBG) Reagent for use on Access Immunoassay Systems, Part number(s): A48617 |  |  | July 26, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because the Instruction for Use (IFU) for Access SHBG does not specify the units of measure to be usedwhen calculating the Free Androgen Index (FAI %). Access SHBG (nmol/L) values and Access Testosterone (ng/mL) values are reported in different default units. The testosterone value (when reported in ng/mL) requires conversion to nmol/L before the FAI % can be calculated. | Pending | Beckman Coulter sent an Urgent Product Corrective Action (PCA) letter on the week of July 26, 2010 to the affected customers. The letter was dated July 30, 2010, and identified the product, the problem and provided the appropriate equation to recalculate the FAI% values. The customers were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Customers were also asked to complete and return the enclosed response form within 10 days to confirm that they had received the important notification.The customers were instructed to contact Beckman Coulter Technical Support at 1-800-854-3633 in the United States, or contact their local Beckman Coulter Representative if they had any questions regarding the notification. |  | 1113 | 1,113 Units ( 894 in US &amp; 219 in Canada) | Worldwide Distribution - USA including AL, AR, CA, CO, FL, GA, IL, IN, LA, MA, MI, MN, MO, MS, NC, NE, NV, NY, OH, OK, OR, TN, TX, UT, VA, WA, AND WI and the country of Canada | Terminated | May 04, 2012 | 648 | 1113
Z-0430-2011 | 57098 | Yes | LWP |  |  | Cardiovascular | PMA | Medtronic Dual Chamber Temporary Pacemaker, model 5388. Manufacturer: Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604, USA. |  |  | October 11, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | A subset of Model 5388 Dual-Chamber External Temporary Pulse Generators worldwide may be unable to power up or may power down unexpectedly. The issue presents itself in one of two ways. First, during startup, and prior to initiating patient therapy, the instrument may power down in 1 - 2 seconds. Second, the instrument may power up correctly, but power down at a later time while in use. | Nonconforming Material/Component | Medtronic issued an Important Medical Device Correction letters, dated October 2010 and November 2010 to hospital's biomedical engineering department. The letter described the issue, how it presents, and stated that Medtronic will contact the hospitals to return affected product for servicing. The second letter reiterated that Medtronic would notify customers when FDA approves their design upgrade for returning devices for service. In the interim, hospitals should continue to use normally functioning devices. Cusotmers may contact Medtronic Instruments Technical Service at 800-638-1991 concerning this action. |  | 5595 | 5,595 ( 2,696 US, 2,899 OUS) EXPANDED: 944 (852 US, 92 OUS) | Worldwide Distribution: throughout USA and territories of Guam and Puerto Rico and to the countries of: Algeria, Argentina, Australia, Austria, Barbados, Belarus, Belgium, Bolivia, Bosnia And Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, French Guiana, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Latvia, Lebanon, Luxembourg, Macedonia, Malaysia, Martinique, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Oman, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, and Yemen NO ADDITIONAL COUNTRIES ARE AFFECTED BY THE EXPANSION. | Terminated | May 29, 2012 | 596 | 5595
Z-0427-2011 | 56913 | No | KDN | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Belzer UW Cold Storage Solution, 1000 mL. Sterile and non-pyrogenic. Formulated for flushing and cold storage of kidney, liver and pancreas organs. Manufactured by: Preservation Solutions, Inc., Elkhorn, WI 53121. Manufactured for: BTL Solutions LLC, a subsidiary of Bridge to Life, Ltd., 1225 Laurel Street, Columbia, SC 29201. |  |  | September 24, 2010 | 2010 | Preservation Solutions Inc | Mislabeled volume - 2 liter bags of Belzer UW Cold Storage Solution were labeled as 1 liter. | Labeling mix-ups | Bridge to Life, Ltd, issued a Notification of Product Recall letter dated October 18, 2010 to all direct accounts identifying the affected product and labeling issue. Customers were informed that all product from the affected lot has been replaced by the firm. Customers with questions about this recall can contact Bridge to Life at 803 673-4548.. |  | 55 | 55 units | Nationwide Distribution - in US to the states of CA, MO, and PA. | Terminated | April 02, 2012 | 556 | 55
Z-0439-2011 | 56985 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cholestech LDX Multi-Analyte Controls, Level 1 &amp; 2, 1 vial set, Catalog Number 12-712. |  |  | October 06, 2010 | 2010 | Biosite Inc Dba Innovacon Inc. | The recall was initiated because results were being reported as being out of range. Firm is reassigning ranges for HDL cholesterol in product. | Pending | All customers were notified and given instructions, requirements, and expectations for their role in the Correction activities using the Notice of Correction with attached Returned Verification Forms which were supplied via fax, email or direct mail beginning on October 6, 2010.Customers are to share the information provided in the Notice with their laboratory staff and retain the letter as part of their laboratory Quality System documentation. If any affected product had been further distributed to another laboratory, then a copy of the letter should be provided to them. Customers are to complete and fax the enclosed Verification Form within 10 days to confirm their receipt of the notice.Consignees with questions about the information contained in the notification were instructed to contact:Alere San Diego, Inc.9975 Summers Ridge RoadSan Diego, CA 92121U.S.A.Phone: 877 441 7440FAX: 858 805 8457E-mail: Response.ts@alere.com |  | 3382 | 3,382 kits | Worldwide Distribution -- USA, Australia, Canada, Great Britain, India, Italy, Mexico, Norway, Puerto Rico, South Korea, Spain, Switzerland, and Trinidad &amp; Tobago. | Terminated | August 01, 2012 | 665 | 6529
Z-0433-2011 | 57028 | No | FTY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LeMaitre VascuTape Radiopaque Tape, LeMaitre Stent Guide, 20 strips |  |  | October 22, 2010 | 2010 | Lemaitre Vascular, Inc. | Box label (Top Part). Lot and Expiration Date information were switched. | Labeling Change Control | Lemaitre Vascular notified customers by an Urgent: VascuTape Radiopaque Tape Device Field Safety Notice letter on 10/22/10 via USPS with delivery confirmation. The Notice advised users of the incorrect lot and lot number on the outer box and also provided an image of the box and label for ease in identifying the product in their inventory. The letter states that the use and functionality of the device are not affected by the problem; however, if customers feel uncomfortable continuing to use the affected device, they can contact Customer Service for replacement. The firm recommended discarding the outer box of the affected product. The letter contains a form that is requested to be returned to LeMaitre Vascular as record of notification and reconciliation. Accounts can contact customer service at 800-628-9470. |  | 160 | 160 units | Nationwide Distribution -- AL, CA, CT, IL, ME, MI, MS, NC, NJ, NY, PA, PR, TX, and WA. | Terminated | September 15, 2011 | 328 | 1940
Z-0444-2011 | 55889 | No | OFY | Endosseous dental implant accessories. | Dental | Dental | Enforcement Discretion | Zimmer Dental, Drill stop kit, REF DSKIT, for use with dental implants. |  |  | May 18, 2010 | 2010 | Zimmer Inc. | The instructions for use contain pre-vacuum sterilization instructions that may lead to compromised sterility at the lid-kit interface location. | Labeling design | Zimmer sent out "Urgent Device Correction" letters dated 5/18/2010. The consignees were provided with new instructions for use. The letter informed users that the current pre-vacuum cycle of 3 to 18 minutes is not enough time to assure sterility between the lid-kit interface and could lead to cross contamination of multiple patients. The users were informed that a pre-vacuum time of 30 minutes was required to assure sterility. |  | 158 | 158 | Distribution Nationwide, to France, and Spain. | Terminated | May 18, 2011 | 365 | 158
Z-0448-2011 | 55200 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Estradiol, Part Number: 33540 |  |  | March 04, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed that three lots of Access Estradiol reagent (REF 33540) contains an incorrect raw material which may produce incorrect patient results. | Other | A Product Corrective Action (PCA) letter dated March 17, 2010, with attached PCA Customer Response Form was sent to the affected customers. The letter describes the product, problem and actions to be taken by the customers. The customers were instructed to discontinue use of the three Access Estradiol Reagent lots; discard all remaining inventory of affected product; share this information with laboratory staff and retain the notification as part of their laboratory Quality System documentation, and complete and return the enclosed response form within 10 days via fax to 786-639-4000 or mail: Beckman Coulter, Inc., 1000 Lake Hazeltine Drive, Chaska, MN 55318, Attn: Regulatory Affairs.The following product replacement contacts were provided:In the United States, please contact Client Services at 1-800-526-3821.In Canada, please contact Customer Service at 1-800-463-7828.Outside the United States and Canada, contact your local Beckman Coulter Representative.Questions and assistance could be obtained through Technical Support at 1-800-854-3633 in the United States and Canada. Outside the United States and Canada, contact consignees were told to contact their local Beckman Coulter Representative. |  | 20209 | 20,209 units (5475 in US) | Worldwide distribution: USA and Internationally to: Algeria, Argentina, Armenia, Australia, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Czech Republic, El Salvador, France, Georgia, Germany, Ghana, Greece, India, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Malaysia, Mexico, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Oman, Panama, Philippines, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Serbia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab, United Kingdom, Venezuela, Vietnam, and Zambia | Terminated | May 07, 2012 | 795 | 20209
Z-0446-2011 | 57008 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BARDÂż Mesh Pre-shaped (1.8" x 4.0"/4.5cm x 10cm) |  |  | October 15, 2010 | 2010 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: The labeled BARDÂż Mesh Pre-shaped product code 0112700 (1.8" x 4.0"/4.5cm x 10cm), but the contents are the BARDÂż Mesh Large Pre-shaped product code 0113700 (2.4" x 5.4"/6.0cm x 13.7cm). | Process control | Davol issued an "URGENT MEDICAL DEVICE RECALL" letter on October 15, 2010, via Overnight Delivery to all customers. The letter describes the product, the problem and the action to be taken by the customers. The customers were instructed to immediately examine their inventory and quaratine the product; identify and notify customers they may have further distributed the product to of this recall; contact Davol Customer Service at 1-800-556-6275 for instructions on how to return product, and complete and return the attached EFFECTIVENESS CHECK form via fax to 1-401-825-8753. If you have an questions, please contact Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Service Support Department at 1-800-562-0027. |  | 144 | 144 units | Nationwide distribution | Terminated | October 27, 2011 | 377 | 144
Z-0445-2011 | 57282 | [6 empty] | Lumen Biomedical Xtract Aspiration Catheter system, intravascular catheter, consisting of One Catheter, Two 30 mL Syringes, One Extension Tube with Stopcock, and One 40 um Cell Strainer. Single-user, 0.014" guidewire compatible intravascular extraction and aspiration catheter, Part number 1450-0001 size 4.2F, and part number 1451-003 size 5.8F, distributed by Volcano Corporation. |  |  | November 03, 2010 | 2010 | Volcano Corporation | Potential for product failure. Packaging/Storage issue results in compromised integrity and potential product failure during use, leading to potential patient injury. | Device Design | Firm initiated recall on 11/03/2010, and sent notifications to its consignees via letter. The firm has requested that consignees provide the notification to all personnel within the organization and letters will be sent directly to known customers on 11/04/2010 by Volcano. Customers have been instructed to cease use and contact customer service to arrange returns, and credit will be issued. |  | 4724 | 4724 units | Worldwide distribution: USA, Greece, Egypt, Kuwait, Netherlands, Saudi Arabia, Italy, France, Sweden, Portugal, Great Britain, Austria, Switzerland, Germany, Romania, Poland, and Bulgaria | Terminated | December 07, 2011 | 399 | 4724
Z-0447-2011 | 57048 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Medtronic Navigation O-Arm Imaging System, Door Assembly |  |  | October 15, 2010 | 2010 | Medtronic Navigation, Inc. | Safety Notice: Door Cable Periodic Maintenance Schedule and Cable Wear for the O-ARMÂżÂż Imaging System Breakage of the door cable during opening and closing may occur and cause physical harm or delay of therapy | Equipment maintenance | Medtronic Navigation sent all consignees/customers an "Urgent Field Safety Notice" letter dated October 15, 2010. The letter described the product, problem and action to be taken by the customers. The customers were instructed to review and follow the attached information that describes the process to properly detect and proactively manage cable wear; to ensure that the door is completely open when positioning the patient into or removing them from the gantry and when placing the O-ARM Imaging System Tube Drape (BI-900-00003) onto the unit, and report any visual or audible (cracking noise) wear of the cable to Medtronic Representative ASAP. Note: A Medtronic Representative will be contacting the customers within the coming weeks to schedule a time to inspect the current condition of the cable, explain and train the customers on the recommended workflow and inspection procedure and place a permanent label on their O-ARM Imaging System to reinforce these actions. If you have any questions about this notification, please contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or 720-890-3200. |  | 244 | 244 units | Worldwide distribution: USA including states of: AL, AZ, CA, CO, CT, DC, FL, GA, IA, IL IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, and WI; and countries including: Canada, Austria, Australia, Belgium, Brazil, Libya, South Africa, India, Japan, Singapore South Korea, Haiti, Kuwait, Saudi Arabia, Turkey, UAE, Germany, Netherlands, Great Britain, Italy, Poland, Netherlands, Spain,Sweden, and Switzerland. | Terminated | June 04, 2012 | 598 | 244
Z-0461-2011 | 55604 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The Dimension NTP and LNTP Flex reagent cartridge (N- Terminal Pro- Brain Natriuretic Peptide and N-Terminal Pro brain Natriuretic Peptide (low volume) |  |  | April 12, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Some LOCI cartridges have sub-optimal seal of the lidstock that can cause rupture of the package at high elevations, causing the cartridge to leak. Use of leaking cartridges may result in inaccurate test results. | Other | Siemens issued an "Urgent Field Safety Notice" dated April 2010, informing users of the problem and instructing them to inspect cartridges for leaking fluid. Cartridges with leaking fluid should be reported to the Siemens Technical Solutions Center. The Technical Solutions Center can be contacted at 800 441 9250. |  | 3364 | 3364 | Worldwide Distribution: USA, Europe, Australia, China, Japan, Korea, and Saudi Arabia. | Terminated | March 19, 2013 | 1072 | 14037
Z-0456-2011 | 55178 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT CREA Cartridge, Abbott List Number: 06F10-01, 02; IVD. Manufactured by Abbott Point of Care, Inc., Princeton, NJ. |  |  | March 17, 2010 | 2010 | Abbott Point of Care Inc. | The i-STAT CREA Cartridges may produce inaccurate (depressed) results at or near the end of the labeled expiry date when stored above 25 degrees for 2 weeks. | Other | Abbott Point of Care issued Urgent Product Recall Notification letters dated March 2010 to all direct customers identifying the affected product and the actions to be taken. Customers were instructed to determine if they had cartirdges that had been stored above 25 degrees Celsius, and if so, contact Abbott for replacement or credit. Abbott Technical Support can be contacted at 800-366-8020 option 1 about this recall. |  | 34525 | 34525 cartridges | Worldwide Distribution: Throughout USA and Canada. | Terminated | September 22, 2011 | 554 | 34525
Z-0449-2011 | 56345 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens brand ARTISTE MV, Artiste Digital Linear Accelerators in combination with the Flat Panel, Part No. 8139789, |  |  | January 22, 2010 | 2010 | Siemens Medical Solutions USA, Inc | The robotic lift armature shaft may break and allow the Flat Panel Positioner with the digital flat panel imager to fall into the Flat Panel cradle. If the positioner is over a patient, there is a potential for injury. | Other | Siemens sent an Urgent Medical Device Correction letter to all on January 27, 2010, via phone, Fax, certified mail and/or delivery via Siemens Customer Service Group to all its direct consignees, informing them of the affected products, the problem, and provided instructions on the recall.Consigneees were to be contacted by an ARTISTE Service Representative for an appointment for corrective action. Consignees were instructed to include Customer Safety Advisory Notice in their ARTISTE System Owner Manual.For questions regarding this recall call (925) 246-8407. |  | 106 | 106 | Worldwide Distribution - USA including CA, DE, FL, LA, MA, MI, MO, NC, NE, NJ, OH, PA, TN, UT, WI, AND WV and Egypt, Germany, Japan, Lebanon, Malaysia, New Zealand, Poland, Russia, Spain, Sweden, and Turkey | Terminated | April 12, 2011 | 445 | 106
Z-0457-2011 | 56147 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics V60 Ventilator, manufactured by Respironics California Inc. in Carlsbad, CA. |  |  | July 02, 2010 | 2010 | Respironics California Inc | Respironics California Inc has received customer complaints of an odor or burning smell coming from the V60 Ventilator, and identified a component on the Power Management PCB that exhibited thermal damage. | Device Design | Phililps Healthcare issued an Urgent Medical Device Field Correction letter dated July 2, 2010 concerning the RespironicsV60 Ventilator issue. Customers will be contacted by US Field Service Technicians and international distributors to schedule and perform inspections and replacement of the affected component. The unit may continue to be used pending completion of the field inspection.Customers may contact Respironics US Customer Service about this action at 1 800 345 6443, option #5, then option #1. |  | 504 | 504 | Worldwide Distribution: Throughout the US and Puerto Rico,, and to the countries of: Australia, Austria, Belgium, Canada, Denmark, France, Germany, Hong Kong, India, Italy, Kenya, Kuwait, Libya, Malaysia, Oman, Netherlands, Norway, the United Kingdom, Saudi Arabia, and Spain. | Terminated | December 01, 2010 | 152 | 504
Z-0450-2011 | 56912 | No | GXQ | Dura substitute. | Neurology | Neurology | 510(k) | DuraGen XS Dural Regeneration Matrix: 2 inches x 2 inches. |  |  | September 29, 2010 | 2010 | Integra LifeSciences Corp. | Test results no longer support a 3 year shelf life. | Other | Integra LifeSciences Corporation sent an URGENT Product Recall Notification letter dated September 29, 2010, to all affected customers. Integra NeruoSpecialists were informed of the recall via telephone conference call and recall notification letters sent by e-mail. Letters were also sent to customers by Federal Express with copies of e-mail provided to Interga Neuro Specialists on October 1, 2010.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately quarantine any affected product and contact Integra Customer Service at 1-800-654-2873, select option 5, followed by option 3.Customers were instructed to fill out and return the enclosed "Recall Acknowledgement and Return Form" indicating whether or not they possessed any of the affected product.For any questions related to this recall call 1-800-654-2873 and selection option 1. |  | 118 | 118 | Nationwide Distribution including AK, AL, AZ, CA, CO, DC, DE, FL, GA, HI, IL, IN, LA, MA, MD, ME, MN, MO, MT, NC, ND, NH, NJ, NM, NY, OK, PA, TN, TX, UT, VA, VT, WA, WI, AND WV. | Terminated | April 23, 2012 | 572 | 2766
Z-0428-2011 | 57166 | No | NRY | Percutaneous catheter. | Cardiovascular | Neurology | 510(k) | Penumbra System Reperfusion Catheter 032, catalog number PSC032, lot number lot F15020, manufactured by Penumbra, Alameda, CA. |  |  | October 25, 2010 | 2010 | Penumbra Inc. | Reperfusion catheter may fail at the mid-shaft joint and may break off. If the segment cannot be removed, it may result in blocked blood flow, requiring emergency medical treatment. | Process control | Penumbra issued an Urgent Voluntary Field Removal letter dated October 28, 2010 to consignees, identifying the affected device and actions to be taken by consignees. Customers were instructed to inspect their inventory for the affected product and quarantine it, pending return to Penumbra. Penumbra employees will contact customers to arrange return and replacement of affected units. Customers were also requested to complete and return a product identification form and return form to Penumbra within 5 business days.Customers can contact Penumbra at 1 510 748-3223 concerning this action. |  | 61 | 61 | Worldwide Distribution: 17 consignees in the US, including in the states of CA, FL, HI, ID, IL, KY, MA, MS, NJ, NY, PA, and TX, 1 in Canada and 21 in the EU countries of France, Germany, Hungary, Italy, Norway, Slovenia, Spain, and the United Kingdom. | Terminated | June 01, 2011 | 219 | 61
Z-0455-2011 | 56012 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Inc. Eclipse/VARiS Vision, an Eclipse radiotherapy Treatment Planning System (Eclipse TPS); Eclipse Version 7.3.10 (multiple workstations) VARiS Vision System 7. x; |  |  | May 13, 2010 | 2010 | Varian Medical Systems Oncology Systems | Incorrect treatment plan or record: A software design deficiency has the potential for corrupted treatment plan records or treatment records in cases where multiple users editing a record simultaneously can result in a corrupted (but accessible) treatment plan record. | Software design | On May 13, 2010, the firm issued letters to all its direct consignees, informing them of the affected products and providing instructions on the recall.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** |  | 2466 | 2417 units***12-20-12 AMENDED TO: 49 units*** | Nationwide and Worldwide | Terminated | February 24, 2012 | 652 | 2466
Z-0462-2011 | 57142 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SignaÂż Excite" 1.5 T Magnetic Resonance Systems, manufactured by GE Healthcare, P.O. Box 44 Milwaukee, WI 53201. |  |  | May 11, 2009 | 2009 | GE Healthcare, LLC | SignaÂż MR Excite" 1.5T system containing 11.1_M4_0818.a software, associated with 2D Fast Spin Echo based pulse sequences (2D FSE-XL, 2D FRFSE-XL, 2D FSE-IR, 2D T2FLAIR and 2D T1FLAIR) may experience an image flip along the phase encode direction and incorrect annotation, which may result in incorrect localization of anatomy in oblique axial images with the pulse sequences. | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 21, 2009. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. GE Field Engineers will contact customers to schedule a field modification of the affected products to address this issue.Customers can contact GE Healthcare at 262 521-6100 about this issue. |  | 1464 | 1464 | Worldwide Distribution: Throughout USA except states of AK, VT and WY, and the countries of: Argentina, Australia, Bulgaria, Brazil, China, Canada, Switzerland, Chile, Colombia, Czech Republic, Ecuador, Egypt, Spain, Finland, France, Great Britain, Guatemala, Honduras, Croatia, Hungary, Indonesia, Ireland, India, Iran, Iceland, Italy, Japan, Kenya, South Korea, Kuwait, Lebanon, Luxembourg, Libya, Morocco, Malta, Mexico, Malaysia, Netherlands, Norway, New Zealand, Panama, Pakistan, Poland, Puerto Rico, Portugal, Reunion, Romania, Russian Federation, Saudi Arabia, Sweden, Singapore, Syria, Thailand, Tunisia, Turkey, Taiwan, Venezuela, and South Africa. | Terminated | November 24, 2015 | 2388 | 1464
Z-0481-2011 | 50117 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | T2 Ankle Arthrodesis System Nail and Nut Adapter; |  |  | August 02, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | The thread of the nut which should fix the nail adaptor to the targeting arm may seize on the threat of the nail adaptor, interrupting surgery. | Other | Stryker Orthopedics issued an Important Stock Recovery letter dated August 2, 2006 The customer was instructed to inspect inventory for the affected product and return it to the firm. The customer was asked to complete and return a Product Accountability Form and contact their Stryker representative for replacement product.Stryker can be contacted about this issue at 201 831-5118. |  | 8 | 8 sets | Nationwide Distribution: To one location in New Jersey. | Terminated | November 22, 2013 | 2669 | 8
Z-0482-2011 | 55760 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Neuroform 3 Microdelivery Stent System; Distributed and manufactured by: Boston Scientific Corporation, Fremont, CA |  |  | March 01, 2010 | 2010 | Boston Scientific Corporation | The product may contain manufacturing aids (mandrels) which should not be present in the final product. If the mandrel remains in place, the catheter cannot be loaded over the required guidewire, so cannot be used. | Process control | Boston Scientific issued an Urgent Customer Notice dated March 1, 2010 to all US customers, with personal delivery by sales representative where applicable outside the US, identifying the affected device. Customers were instructed to return the device for replacement, and to return an acknowledgement form.Customers can contact Deborah Springer at Boston Scientific at 800-845-4598 about this recall. |  | 638 | 638 units | Worldwide Distribution: Throughout the US and to the countries of: Austria, Australia, Belgium, Brazil, Canada, China, Columbia, Czech Republic, Germany, Denmark, Ecuador, Spain, France, Great Britain, Greece, Hungary, India, Ireland, Italy, South Korea, Luxemburg, Mexico, the Netherlands, New Zealand, Peru, Russian Federation, Sweden, Slovenia, Taiwan and Ukraine. | Terminated | November 26, 2010 | 270 | 638
Z-0479-2011 | 52100 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Product's marketing brochure is labeled in part: "Philips Easy Diagnost Eleva Conventional R/F system...PHILIPS". manufactured by Philips Medical Systems Gmbh, Hamburg, Germany. |  |  | January 21, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | Patient and image data can be exchanged when a patient is selected in the patient list and the examination tab is pressed with background query running at the same time. There is a risk that the data may then be linked to the incorrect patient, which may adversely affect diagnosis and treatment. | Software design | Philips Medical Systems issued URGENT-Device Correction letters dated January 27, 2009 to consignees describing the issue with the Easy Diagnost Eleva GXR RF systems. The consignees were informed of the potential for incorrect diagnosis or treatment if an incorrect image is used for diagnosis. Consignees should contact the Philips Healthcare Call Center at 800-722-9377, #5, #2 and reference "FC0 70600030" with any communication. |  | 314 | 314 units | Nationside Distribution: USA, including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NE, NH, NJ, NV, NY, OH, OK, OR, PA, PA, RI, SC, SD, TN, TX, TX, UT, VA, VT, WA, WI, and WV. Devices are not distributed to any foreign consignees. | Terminated | December 30, 2010 | 708 | 314
Z-0469-2011 | 52040 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | The product's marketing brochure states: "1.5T coils and accessories...Philips Achieva and Intera 1.5T MR systems". |  |  | May 01, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | Combined use of the Synergy Flex-M / Shoulder Coil 1.5T increases the chance of RF interaction and heating up of the coil. This may result in possible burns of the patient. | Other | Philips issued URGENT-Field Safety Notice letters dated May 1, 2009 and October 14, 2009 to their consignees. The letter informed customers that the firm will replace the "old" Flex M/Shoulder Coils with a new Flex-M coil.Consignees could contact Philips Healthcare Call Center at 800-722-9377, #5, #3, #1 and reference FCO 78100290 for any questions. |  | 389 | 389 units | Nationwide Distribution: in the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VA, WA, WI, WV, WY, and Washington DC. There were no foreign consignees. | Terminated | January 06, 2011 | 615 | 389
Z-0473-2011 | 55682 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart FR2+ automated external defibrillator (AED) model: M3841: LAERDAL brand; text prompts only on screen (NO ECG); no configurable manual charge. |  |  | May 28, 2010 | 2010 | Philips Medical Systems | Failure of the voltage detector, U8 may cause the AED battery to drain more rapidly than normal or make the AED unusable. | Nonconforming Material/Component | Philips Medical Systems issued an Urgent Medical Device Recall letter dated May 28, 2010 identifying the affected device and actions to be taken. Customers were instructed to review their inventory for affected devices. If affected devices are found, the customer should contact Philips to arrange for a replacement device. The customer should also run a battery insertion test. If the device passes the test, the AED may remain in service until a replacement is received. For further information, contact Philips at 1 800 263-3342, option 6. |  | 1478 | 1,478 units | Worldwide Distribution: Throughout the USA, and to the following countries: ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BERMUDA, BOSNIA AND HERZEGOVINA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, FINLAND, FRANCE, GERMANY, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PERU, PORTUGAL, PUERTO RICO, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIETNAM. | Terminated | September 13, 2013 | 1204 | 16573
Z-0478-2011 | 56428 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | The software is ARM version 60.80.101.025 (referred to as Nano 1.0.5). |  |  | August 13, 2010 | 2010 | Sonosite, Inc. | Software ARM version 60.80.101.025 (referred to as Nano1.0.5) on the NanoMaxx ultrasound system results in errors when the "print" command is selected. As a result, the displayed MI or TI values may be incorrect. | Software design | Sonosite issued Urgent Medical Device Correction letters dated August 9, 2010 to consignees. Consignees were advised of the software problem and given directions to verify the software version they are using. They are advised to contact SonoSite at 877-657-8118 to obtain a free software upgrade. |  | 69 | 69 | Worldwide Distribution: USA and to the countries of Australia, Great Britain, India, Israel, the Netherlands, New Zealand, Norway, and Spain. | Terminated | July 21, 2011 | 342 | 69
Z-0480-2011 | 57023 | No | KZJ | Gas-generating device. | Microbiology | Microbiology | 510(K) Exempt | GasPak EZ Large Incubation Container, a component of the GasPak EZ Gas Generating Container System, catalog # 260672, labeled in part ***Becton Dickinson and Company Loveton Circle Sparks, MD 21152 USA*** |  |  | August 13, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic reagent containers may be defective and could cause incorrect test results in patient samples. | Nonconforming Material/Component | BD Diagnostic Systems issued an Urgent Product Recall letter dated August 2010 to consignees. Consignees were advised of the defect and were requested to discard product for replacement. Distributors were additionally requested to provide distribution lists of end users for direct BD notification. Accounts were requested to return a response form.BD Customer Service can be contacted at 800 675-0908, and BD Technical Services can be contacted at 800-638-8663. |  | 496 | 496 units | Worldwide Distribution: USA and the countries of Australila, Belgium, Canada, India, Mexico, New Zealand. Singapore, and South Korea, | Terminated | January 05, 2011 | 145 | 496
Z-0465-2011 | 57117 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Rectal Tubes, 22 French, 20 Long (51 cm), X-Ray Opaque Rubber with Funnel end, Open Tip, One Eye , Non sterile, 12 Each, C.R. Bard, Inc., Covington, GA 30014 |  |  | October 28, 2010 | 2010 | C.R. Bard, Inc., Urological Division | The product is misbranded. The product is labeled as 22 French, however the product is actually 20 French. | Employee error | Bard sent an Urgent Product Recall Notification letter dated October 28, 2010 to the consignee. They were instructed to return any outstanding inventory of the affected product. They were advised to notify their consignees of the recall and to retrieve relevant product from those locations. Affected products received by Bard Medical will be destroyed.BMD can be contacted at 1 770 784-6969 for questions or assistance. |  | 108 | 108 eaches | Nationwide Distribution: In the state of Ohio. | Terminated | March 31, 2011 | 154 | 108
Z-0463-2011 | 57115 | No | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE Stainless Steel Primary Molar Crowns, Made in USA by 3M ESPE Dental Products, St Paul, MN. |  |  | October 06, 2010 | 2010 | 3M Company / Medical Division | Unitek Stainless Steel Permanent Crowns with product Codes 902150, 900213, and 900223, is mislabeled.The crown contained within the package is incorrectly labeled; series UL3 crowns are labeled series UR3 and vice versa. The lid of the package correctly identified the product number and lot. This mislabeling could result in placement on the right side of a crown intended for use on the left, and | Packaging process control | 3M ESPE issued a Device Recall letter dated October 6, 2010 to distributors and dentists. The letter listed the product being recalled including code and lot codes, background information including a picture of the product problem, and the recall procedure. The letter included a template letter to be used by distributors to inform their customers who have received the product. Mislabeled product is to be quarantined and returned to 3M ESPE.3M ESPE can be contacted at 651 733-1110. |  | 9819 | 9819 | Nationwide Distribution: Throughout USA, including the states of AZ, CA, GA, IL, IN, IA, LA, MA, MN, MO, NV,NJ, NY, NC, PA, TN, TX, WA, and WI.Canada and Dominican Republic | Terminated | February 04, 2012 | 486 | 9819
Z-0464-2011 | 57109 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DX-G Digitizer, manufactured by Agfa-Gevart NV, Mortsel, Belgium. |  |  | October 08, 2010 | 2010 | AGFA Corp. | Packaging was switched between a DX-G Digitizer and a DX-M Digitizer, resulting in the wrong product, which is not sold in the US, being shipped to the consignee. | Other | Agfa halted Installation of the incorrect unit, and the unit was sent to Agfa warehouse in Carlstadt, NJ, where it is being held in quarantine for export. |  | 1 | One unit | Nationside Distribution: to NJ. | Terminated | December 30, 2010 | 83 | 1
Z-0468-2011 | 57095 | No | HWX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CSS-059-40-4 Modified 4.0mm Modified Depth Gauge, Lot 0985011 |  |  | October 15, 2010 | 2010 | Orthohelix Surgical Designs Inc | The print for modified CSS depth gauges has inaccurate definition of the etched screw length. The error results in a length reading that is longer (5 mm in the case of the 7.0 and 4mm in the cases of the 5.5 and 4.0) than the actual length of the buried guide wire segment and even longer (10 mm in the case of the 7.0 and 8mm in the cases of the 5.5 and 4.0) than the design intention readings. This | Device Design | OrthoHelix issued an Important Recall Notice letter dated October 15, 2010 to their customers, identifying the affected products. Customers were requested to immediately return the afected devices. OrthoHelix can be contacted at 330 247-1444 concerning this recall. |  |  |  | Nationwide distribution: to the states of GA and MO. | Terminated | September 06, 2012 | 692
Z-0475-2011 | 57090 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic Sprint Quattro Secure, model 6947. Steroid eluting, quadripolar, screw-in, ventricular lead with Right Ventricular (RV) and superior vena cava (SVC) defibrillation coil electrodes. Contents Sterile. Manufactured in: Villalba, Puerto Rico, USA. Medtronic, Inc., Minneapolis, MN 55432 U.S.A. |  |  | October 25, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Over-retraction of the helix during initial implant or subsequent repositioning may result in the inability to extend the helix. This does not impact acute or chronic performance of successfully implanted leads. | Device Design | Medtronic issued an Important Medical Device Correction letter, dated October 2010, to doctors. The letter contains a Performance Note. The letter describes the issue and the performance note provides information on how to mitigate the issue through replacement with a new lead.Customers can contact Medtronic Technical Services at 800 723-4636 |  | 154000 | 154,000 total | Worldwide Distribution: Throughout the US, including Guam, Puerto Rico, and US Virgin Islands, and to the countries of: Algeria, Argentina, Armenia, Australia, Austria, Bahamas, Bangladesh, Barbados, Belgium, Bolivia, Bosnia And Herzegovina, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, El Salvador, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Malta, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Saint Lucia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Viet Nam. | Terminated | February 04, 2012 | 467 | 154000
Z-0477-2011 | 57077 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare IDI Mammo Workstation |  |  | August 28, 2009 | 2009 | GE Healthcare, LLC | The measurement values provided by the GE Workstation may be incorrect when applied to magnification images acquired on non-GE digital mammography systems.If not noticed by the caregiver, this could lead to an overestimate of the size of the breast lesion that may impact patient treatment and safety. The measurement values for GE images are not affected. | Software design | GE Healthcare issued "Urgent Medical Device Correction" letters dated August 28, 2009 or August 31, 2009 to either Seno Advantage Customers (Hospital Administrators/Risk Mangers, Radiology Department Managers), and Image Diagnostic Customers (Hospital Administrators/Risk Managers, Radiology Department Managers). The letters described the Safety Issue, Affected Product Details, safety Instructions, Product correction and Contact Information. GE will linstall new software with the release of Field Modification Instruction 12129.Customers can contact GE Healthcare about this issue at 800-437-1171. |  | 158 | 158 | Worldwide Distribution: USA in the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI,MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WY, and to the countries of UNITED KINGDOM, TURKEY, TAIWAN, SWITZERLAND, SPAIN, SOUTH, KOREA, SAUDI ARABIA , PORTUGAL, NEW ZEALAND, NETHERLANDS, KUWAIT, JAPAN, ITALY, GERMANY, FRANCE , CHINA, CANADA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, VENEZUELA, UNITED ARAB EMIRATES, THAILAND, SWEEDEN, SOUTH AFRICA, SLOVENIA, SINGAPORE, RUSSIA, PHILLIPNES, NORWAY, POLAND, MOROCCO, MEXICO, MARTINIQUE, MALAYSIA, LUXEMBOURG, LEBANON, JORDAN, JAMAICA, ISRAEL, IRELAND, INDIA, ICELAND, HUNGARY, HONG KONG, HONDURAS, GUADOPUO, GREECE, GEORGIA, FINLAND, EGYPT, DOMINICAN REPUBLIC, DENMARK, CZECH REPUBLIC, CROATIA, COLOMBIA, BULGARIA, BOTSWANA, and ALGERIA. | Terminated | June 05, 2012 | 1012 | 2672
Z-0485-2011 | 56816 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart FR2+ AED with ECG Display, Model 3840, Laerdal brand, configurable manual charge in advance mode. |  |  | September 22, 2010 | 2010 | Philips Medical Systems | A flash memory component in the HeartStart FR2+ and Automated External Defibrillator was improperly relabeled by a third party. Therefore the component could not be identified by its source, quality, and date of manufacture. | Nonconforming Material/Component | On 10/5/10, Philips began sending out the Voluntary Medical Device Removal notification letter to their consignees.This letter informs the consignees that Philips Healthcare is conducting a voluntary product removal of a limited number of HeartStartFR2+ automated external defibrillators (AEDs), model numbers M3860A, M3861A, M3849A, and M3841A, manufactured by Philips and shipped between March 2007 and August 2010. The owners of affected AEDs are advised to run a battery insertion test. If it passes, the AED should remain in service until its replacement is received. Philips will be replacing affected units.The consignees can visit the firm's website www.philips.com/FR2AEDAction or contact Philips at 1-800-263-3342 for any information or support concerning the issue. |  | 3244 | 3244 units | Devices were distributed to hospitals, medical facilities, and distributors through out the US and to foreign consignees. Foreign consignees including countries:Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Ecuador, El Salvador, Finland, France, French Guiana, Georgia, Germany, Ghana, Gibraltar, Guatemala, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Japan, Jordan, Kampuchea, Kenya, Korea, Latvia, Lebanon, Libya, Luxembourg, Malaysia, Malta &amp; Gozo, Mexico, Netherland, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestinian Territory, Panama, Peru, Philippines, Portugal, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, United Arab Emirates, United Kingdom, Vietnam, and Zimbabwe. | Terminated | September 05, 2013 | 1079 | 27270
Z-0489-2011 | 56775 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Vent Needle used on the CELL-DYN Sapphire and CELL-DYN 4000 Hematology Analyzers; |  |  | July 23, 2010 | 2010 | Abbott Diagnostics Division, 5440 Patrick Henry Drive, Santa Clara, CA | The Vent Needle, which is built into the Vent Head Assembly on the CELL-DYN Sapphire Instrument and which is used directly on the CELL-DYN 4000 Instrument, may be defective, which may result in an undetected short sample, generating incorrect low results, or which may bend, preventing aspiration of the sample. | Process control | Abbott Laboratories issued a Product Recall letter dated July 23, 2010 to all its direct consignees, informing them of the affected products and providing instructions on the recall. Customers were instructed to discard all affected vent needles and replace them with the needle included with the recall letter. Customers may contact customer support at 1-877-422-2688 concerning this recall. |  |  |  | Worldwide Distribution: Thoroughout the US, and in the countries of Argentina, Canada, Germany, Hong Kong, Japan, and Singapore. | Terminated | September 01, 2011 | 405
Z-0490-2011 | 57002 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx E. coli/P. aeruginosa PNA FISH Culture Identification Kit for the diagnosis of E. coli and/or P. aeruginosa bacteremia |  |  | October 06, 2010 | 2010 | AdvanDx, Inc. | Strains of Acinetobacter radioresistens cross-react with both E. coli/P. aeruginosa PNA FISH and EK/P. aeruginosa PNA FISH to produce a false positive red signal. | Device Design | AdvanDx issued a customer letter dated October 6, 2010 with a PDF file of a new package insert and an acknowledgement form to fill out and return to AdvanDx. Customers were instructed to replace the package insert in any existing inventory with the enclosed package insert.Customers can contact AdvanDx technical support about this action at 1-866-376-0009. |  | 260 | 260 units | Worldwide Distribuition: Throughtout the US, and to the countries of Denmark, Italy, Sweden, and UK. | Terminated | June 06, 2011 | 243 | 260
Z-0493-2011 | 55119 | No | HWN | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Reconstruction System Tenaculum Left Tine Straight 9 inch, Item # 00-1179-021-00, Zimmer Inc., Warsaw, IN. |  |  | February 16, 2010 | 2010 | Zimmer Inc. | The firm determined that the recalled instruments were manufactured using the wrong grade of material. As manufactured, the instruments are more brittle, increasing the potential for fracture. | Nonconforming Material/Component | Zimmer sent an Urgent Device Recall letter dated February 16, 2010, to all affected customers. The letter identified the product, the problem, and the action the customers should take. Customers were instructed to monitor all potentially affected patients through regular office consultations. Ensure Zimmer Reconstruction System Tenaculum System Instrument Users at their facility were informed of the urgent device recall.Locate all instruments and quarantine them immediately.Count all affected product in their territory and record the data on theInventory Return Certification Form included with the letter.Fax a copy of the completed Inventory Return Form to:Zimmer, Inc. at (574) 372-4265.Return the recalled product along with the original completed Inventory Return Certification Form to:Zimmer Distribution CenterAttn: Product Service1777 West Center StreetWarsaw, IN 46580For questions customers were instructed to contact their Zimmer Sales Representative.For questions regarding this recall call 1 - (800) 613-6131. |  | 217 | 217 all products | Worldwide Distribution - USA including AL, FL, HI, IA, IL, KS KY, LA, MI, MO, MS, NC, ND, NJ, OH, OK, OR, PA, TN, TX, UT, VA, WA, AND WI and the countries of Australia, Canada, Central or South America, China, Germany, India, Italy, Saudi Arabia, and UK, | Terminated | October 05, 2011 | 596 | 217
Z-0502-2011 | 56892 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System, Micro-Infusion Manifold; and Interlink Manifold. An Rx sterile device with 3 injection sites, volume per injection site 0.1 mL, priming volume 0.8 mL, length 8.7 cm. Product code 2N3410 and A2N3410 - Japan |  |  | October 21, 2010 | 2010 | Baxter Healthcare Corp. | If the manifolds are connected in series, there may be a risk of the manifolds loosening or disconnecting at the luer connection between them. | Other | Baxter sent an "Urgent Device Correction" letter dated October 20, 2010, via first class mail to the direct accounts, to the attention of the Director of Nursing and the Director of Materials Management, informing them that there may be a risk of the manifolds loosening or disconnecting at the luer connection between them. The loosening or disconnection can result in an interruption of therapy, potentially resulting in patient injury and death. The accounts were requested to place the notification where the inventory of affected product is dispensed, so that users are made aware of the procedures to follow: "If these devices are connected in series, and the luer connector is tightened, do not twist the manifolds. This may lead to loosening or disconnection, which may cause a leak. Additionally, Baxter recommends checking all connections before starting and during infusion therapy." The accounts were requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 752965 | 752,965 units in US and all lots in Japan | Worldwide Distribution -- United States and Japan. | Terminated | October 16, 2011 | 360 | 752965
Z-0491-2011 | 56997 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx EK/P. aeruginosa PNA FISH- Culture Identification Kit for identification of E. coli and/or K. pneumoniae and P. aeruginosa bacteremia (In-Vitro Diagnostic) |  |  | October 08, 2010 | 2010 | AdvanDx, Inc. | Strains of Acinetobacter radioresistens cross-react with both E. coli/P. aeruginosa PNA FISH and EK /P. aeruginosa PNA FISH to produce a false positive red signal | Device Design | AdvanDX sent an e-mail notification along with a PDF file to all customers on October 8, 2010, followed by written notification dated October 6, 2010. The letter identified the product, the problem, and the action the customers should take. Customers were instructed to replace the package insert(s) included with any kits in their inventory with the enclosed package insert(s). Customers were also instructed to fill out the enclosed Acknowledgement of Receipt Form and return it to AdvanDx to confirm that they received the new package insert. For questions regarding this recall contact AdvanDx Technical Support at 1-888-376-0009. |  | 27 | 27 kits | Worldwide Distribution - USA including IL, MI, NY, and PA and the country of Denmark | Terminated | June 06, 2011 | 241 | 27
Z-0492-2011 | 56405 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit Model 8000 only when used with Alaris Pump module or Alaris Syringe module |  |  | August 02, 2010 | 2010 | CareFusion Corporation | Than recall was initiated because CareFusion has identified identified a potential risk associated with the Alaris PC unit formerly known as Medley PC unit (PC Unit) model 8000 only when used with Alaris Pump module (Pump module) or Alaris Syringe module' (Syringe module), This recallnotification details the potential risk and recommended steps for users to take if they encounter a specific PC | Device Design | CareFusion will send by registered return receipt mail the customer notification letter to all customers usingthe affected Alaris PC Unit beginning August 24, 2010. The customer notification letter will be addressed tothe Directors of Nursing, Risk Management, and Biomedical Engineering. Customers will be required toconfirm receipt of the notification by returning the Recall Response Card to CareFusion by postage-paid, selfaddressedmail, fax or email. The recall notice informed the customers of the reason for the notice, the unitsaffected, the potential risk, a step-by-step description of the issue and the required actions for users.The customers were also told that CareFusion is working to release a software update to address this display error and will contact your facility as soon as the software update is available. CareFusion does not require that you return your devices. Clinicians should weigh the risk/benefits to the patients beforecontinuing the use of the device.Customers were also told that any adverse reactions experienced with the use of this product, and/or quality problems shouldalso be reported to the FDA's MedWatch Program. The recall notice provided contact information for CareFusion's Recall Center, Customer Advocacy and Technical Support units. |  | 94146 | 94,146 units | Nationwide in the United States, Canada, Mexico, and the European Union | Terminated | July 25, 2013 | 1088 | 94146
Z-0350-2011 | 56753 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI |  |  | September 01, 2010 | 2010 | Elekta, Inc. | It is possible to do couch corrections twice if both Couch Move Assistant in MOSAIQ and Remote Automatic Table in XVI are configured when in SYNERGISTIQ mode. | Other | Elekta sent an "Important Notice A340" dated September 1, 2010 to all customers. The notice described the product, problem and actions to be taken by the customer. The customers were instructed to make a decision if XVI or MOSAIQ does the corrective table movement; configure and ensure that the system only uses RATM or CMA for the corrective table movement, when in SYNERGISTIQ mode, and complete and return the IMPORTANT NOTICE Confirmation of Receipt form via fax at +44(0)1293 654401. If you have any questions contact your local service representative at +44(0)1293 654200. |  | 23 | 23 units | Nationwide distribution: AZ, CA, DC, IL, KY, MA, MI, MS, NV, NY, OH, OR, PA, VT, and WV. | Terminated | December 14, 2010 | 104 | 23
Z-0336-2011 | 56939 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | ENDOPATHÂż XCEL" with OPTIVIEW" Technology trocar |  |  | September 27, 2010 | 2010 | Ethicon Endo-Surgery Inc | This recall has been initiated because Ethicon Endo-Surgery is aware of an increase in reports of leaking in the ENDOPATHÂż XCEL" with OPTIVIEW" Technology trocar, which potentially can lead to loss of pneumoperitoneum during a surgical procedure. Pneumoperitoneum leaks may lead to conversion from laparoscopic to open surgery. | Device Design | Ethicon Endo-Surgery sent an URGENT: DEVICE RECALL EVENT 2244 letter, dated October 8, 2010. The letter identified the product, the problem, and the action to be taken by the customer. The customers were instructed to stop using the product, examine all inventory to determine if they had the affected product on hand, and remove the affected products from inventory. Ethicon Endo-Surgery requested that each customer complete the Business Reply Form and fax it back to Stericycle at (877) 272-6141 with in 3 business days and also return any affected product to Stericycle.For question regarding this recall call (877) 598-5706 and reference 2244. |  | 671 | 671 lots totalling 2,479,134 instruments | Worldwide distribution: USA, ARGENTINA, AUSTRALIA,BELGIUM,BRAZIL,CANADA,CHILE,CHINA,COLOMBIA,COSTA RICA, DOMINICAN REPUBLIC,ECUADOR, EL SALVADOR, GREECE, GUATEMALA, HONG KONG, ISRAEL, ITALY, JAPAN,JORDAN, KOREA, REPUBLIC OF, LEBANON, NETHERLANDS ANTILLES, NICARAGUA, PANAMASAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICASPAIN, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATESURUGUAY, and VENEZUELA. | Terminated | January 25, 2012 | 485 | 671
Z-0506-2011 | 57102 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Merit MEDALLION Syringe, 10ml, REF: MSSW11-Y, Sterile EO, Single Use, Merit Medical Systems, Inc., South Jordan, Utah 84095. |  |  | October 22, 2010 | 2010 | Merit Medical Systems, Inc. | Packaging for syringes may have holes which would compromise the sterility of the package contents. | Packaging process control | Merit Medical notified product managers of the product recall on October 22, 2010 following a voice mail script. Frequently Asked Questions, Sales Rep Letter, and Recall Notification Forms were then sent.. Sales reps were instructed to contact all of their customers and have them identify and quarantine any unused inventory and either return it to Merit or discard. A response form was to completed by the sales rep and the firm's representative, indicating amount of product quarantined, returned , or destroyed.Also, on 10/22/2010, authorized distributors were contacted by phone, mail, and / or fax. They were also provided a Distributor Letter, FAQs, and Recall Notification Forms instructing them to identify and quarantine any unused inventory and return it to Merit or discard. They were instructed to continue the recall to customers to whom they had shipped the affected Merit product.Additional information is available from Merit at 801-316-4822. |  | 25 | 25 units | Worldwide Distribution: USA including the states of AL, CA, GA, OH, PA, TX, UT, and WA, and to the countires of Japan and UK. | Terminated | January 18, 2011 | 88 | 960
Z-0397-2011 | 56048 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | BD Bioscience CD45 APC-H7 Analyte Specific Reagent, manufactured by BD Biosciences, San Jose, CA. |  |  | November 23, 2009 | 2009 | BD Biosciences, Systems &amp; Reagents | Incorrect expiration date: product with an expiration date of 03/31/2010 was marked with an incorrect expiration date of 06/30/2010. | Labeling mix-ups | A Product Field Action Notification letter, dated November 2009, was sent to customers who purchased the affected device. The letter identified the affected device and explained the reason for recall. Customers were asked to not use the current CD45 APC H-7 lot number 58633 past 2010-03-31 until they receive their new lot of CD45 APC H-7 at no cost to them. Customers were also asked to complete and return the attached fax form. Questions should be directed to BD Customer Support Center at 1-877-232-8995. |  | 87 | 87 units | Nationwide Distribution -- CA, WA, NJ, UT, TX and OH. | Terminated | December 03, 2010 | 375 | 87
Z-0528-2011 | 57178 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer, NexGen Complete Knee Solution LPS Femoral Component, Size G Left, sterile, REF 00-5996-017-01. |  |  | September 15, 2010 | 2010 | Zimmer Inc. | These lots of Zimmer NexGen Complete Knee Solution LPS Femoral Component and Zimmer NexGen Complete Knee Solution LPS Flex Gender Femoral Component exhibit a nonconforming internal CAM radius. The condition is intermittent and not all devices have nonconforming geometry. | Process design | The firm sent URGENT DEVICE CORRECTION AND REMOVAL notices to distributors and sales staff, Risk Managers, and implanting surgeons dated 9/15/2010. Sales staff and risk managers were instructed to return any product on hand. Surgeons were informed of the issue but there was no action to be taken by surgeons. |  | 1151232 | 192 all sizes of 5996 products | US, Germany, Spain, France, UK, Italy, Sweden, South Africa, Czech Republic, Bulgaria, Poland, Romania, and Austria. | Terminated | July 27, 2011 | 315 | 1151232
Z-0537-2011 | 56818 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | XELERIS (1.0, 1.1, 2.0 &amp; 2.1) PROCESSING AND REVIEW WORKSTATION |  |  | October 27, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware of inaccurate user protocol setup associated with the Thyroid Uptake Index application on certain Xeleris workstations(1.0, 1.1, 2.0, 2.1) that may impact patient safety.Inaccurate protocol setup by the user while using Thyroid Uptake Index application has been identifiedon certain Xeleris workstations. Incorrect protocol setup during clinical data acquisition a | Other | GE Healthcare sent an "Urgent Medical Device Correction" letter dated October 27, 2010, to all consignees. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology and Radiologists/Cardiologists and Identifed the product, the problem and the action to be taken by the consignees. Consignees were instructed to replace the original protocol instructions provided with their system with the updated protocol instructions provided with the letter. For any questions consignees were to contact their local GE Healthcare Service Representative or contact the GE Healthcare Call Center:United States or Canada (800) 437-1171Other countries: (262) 896-2890 |  | 6020 | 6020 | All States in continental USA including DC and PR. Viet Nam, Venezuela, Uruguay, United Kingdom, Ukraine, Turkey, Thailand, Taiwan, Switzerland, Sweden, Spain, South Africa, Slovakia, Singapore, Saudi Arabia, Russian Federation, Republic of Korea, Portugal, Poland, Philippines, Panama, Pakistan, Norway, New Zealand, Netherlands, Morocco, Mexico, Martinique, Malaysia, Luxembourg, Kuwait, Kenya, Japan, Jamaica, Italy, Israel, Ireland, Indonesia, India, Hungary, Hong Kong, Guadeloupe, Greece, Germany, France, Finland, Ecuador, Dominican Republic, Cyprus, Colombia, China, Canada, Bulgaria, Brazil, Bosnia &amp; Herzegovina, Belgium, Azerbaijan, Austria, Australia, Argentina, Algeria. | Terminated | August 28, 2012 | 671 | 6020
Z-0538-2011 | 57111 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta VBH Head FIX |  |  | October 14, 2010 | 2010 | Elekta, Inc. | The HeadFIX Baseplate is not screwed down to the adapter to allow to compensate for roll and pitch with the HeadFIX leveling screws of the HeadFIX Baseplate. | Other | Elekta, Inc. sent an Important Notice A343 dated October 14, 2010 to all affected consignees. The Notice identified the product, the problem, and the action to be taken by the consignees. Consignees were instructed constantly monitor the patient before and during the treatment. Check the patient's position after every patient movement (e.g. coughing). Advise all appropriate personnel of the content of the notice place a copy of the Notice In the Important Notice Section of the HeadFix User Manual P10103-108.For questions regarding this recallservice.schwabmuenchen@elekta.com |  | 34 | 34 units | Worldwide Distribution - USA including AZ, CA, CT, FL, GA, IL, KY, MI, MS, NY, NC, OK, OR, RI, TN, TX, VA, WA and the country of Canada | Terminated | December 28, 2010 | 75 | 34
Z-0540-2011 | 56047 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Autoimmune EIA ANA Screening Test Kit, 576 Tests, |  |  | March 01, 2010 | 2010 | Bio-Rad Laboratories Inc | The ANA Screening kit generates a high percentage of false positive results, associated with a low check control reading. | Other | Bio-Rad issued an Urgent Medical Device Recall Notification letter dated March 1, 2010 to customers describing the issue and actions to be taken. Responses will be tracked via response forms. The letter instructed customers that affected products should be identified, quarantined, and discarded, or reworked with shorter expiration dates until newer lots are available.Bio-Rad can be contacted at 510-724-7000 concerning this recall. |  |  |  | Worldwide Distribution: Throughout the US, and to the countries of: Australia, Canada, France, Hong Kong, India, Italy, Japan, New Zealand, Panama, Singapore, South Korea and the UK. | Terminated | January 05, 2011 | 310
Z-0545-2011 | 57069 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura XPER X-ray system. |  |  | May 26, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result, the complete system may shut down and not re-start. | Other | Philips began mailing out an URGENT-Device Correction letter on May 26, 2009. Philips sent out a second URGENT-Device Correction letter dated June 4, 2009, to all consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were advised to contact their local Philips representative immediately if their system completely shut down and would not restart. A mandatory Field Change Order 72200144 was issued for replacement of the Printed Circuit Board (PCB) in the Allura XPER FD system. A mandatory Field Change Order FCO 70800091 was issued for replacement of the Printed Circuit Board (PCB) in the MultiDiagnost Eleva with FD system. The consignees were to be contacted by their Philips Representative for implementation of the corrective action. The firm's Field Engineer would replace the power tray of the systems.If the consignees needed any further information or support concerning this issue, they were to contact the call Center at 800-722-9377, # 5 and reference FCO 72200144 or FCO 70800091 with any questions. |  | 16 | 16 units | Nationwide Distribution including the states of AL, AR, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KY, LA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, OK, OR, PA, RI, TN, TX, VA, and WI. | Terminated | December 09, 2010 | 562 | 100
Z-0543-2011 | 57143 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Edwards Protection Cannulae - Monitoring and Infusion Set, Model MIS004, Sterile EO, Edwards Lifesciences LLC, One Edwards Way, Irvine, CA 92614. |  |  | October 19, 2010 | 2010 | Edwards Lifesciences Llc | Monitoring and Infusion Set connectors may break during manipulation. | Employee error | Edwards Lifesciences issued a Urgent Product Recall letter dated October 13, 2010 to customers identifying the affected devices, requesting the return of unused devices and completion and return of a response form.Further information is available at 800-424-3278. |  | 570 | 570 sets | Nationwide Distribution: in the states of FL, KS, MO, TX, VA, and PR. | Terminated | May 18, 2011 | 211 | 892
Z-0563-2011 | 57241 | No | JOR | Cardiopulmonary bypass pulsatile flow generator. | Cardiovascular | Cardiovascular | PMA | Sarns 8000 Roller pump for Sarns System 8000 Heart Lung Machine console, 220V, Catalog No 16407, Terumo Cardiovascular System Corp Ann Arbor, MI. (not distributed in the US) |  |  | November 02, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The system 8000 roller pumps may fail resulting in a persistent pump stop. | Other | The firm sent "URGENT MEDICAL DEVICE CORRECTION" dated 11/02/2010 to customers. The letter describes the product, problem and corrective actions taken and/or needed to be taken by the customers. The letter states the system responds to a failed inductor by displaying an overspeed message and stopping the roller pump without an audible alert. The pump cannot be restarted. The customers/users are instructed to stop the pump, replace the pump with a backup or if necessary use the handcrank. The letter states that Terumo will replace the roller pump at the next preventative maintenance visit - or sooner- if the customer has a service contract. If the customer does not have a service contract, Terumo will contact the customer within 2 weeks to set up an appointment for the service. Terumo CVS expects all affected inductors to be replaced to be replaced within three months. Call Terumo CVS Customer Service upon completion of the case at 800-521-2818. The customer was also instructed to complete and return the enclosed CUSTOMER RESPONSE FORMPlease note that we are not recommending that you stop using the roller pumps for your Sarns system 8000.For any questions, call toll free 800-262-3304 or phone 734-663-4145. |  | 95 | 95 devices both products | Worldwide distribution: USA including states of CA, FL, NC and NY and countries of United Arab Emirates, China, Ecuador, and Panama. | Terminated | October 19, 2011 | 351 | 95
Z-0561-2011 | 57064 | No | CBL | Carbon dioxide absorbent. | Anesthesiology | Anesthesiology | 510(K) Exempt | Grace SODASORB LF Carbon Dioxide Absorbent, Medical Grade (16%) Soda Lime USP-NF with an ethyl violet indicator which turns from white to violet as absorbent is depleted or upon moisture loss; Manufactured by: W.R. Grace &amp; Co., 6050 West 51st St., Chicago, IL 60638; the product is packaged in pre-paks pre-filled 2.5 lb. canisters, 12 per carton |  |  | October 20, 2010 | 2010 | Darex Container Products Div of W.R. Grace &amp; Co. | Pre-Pak units of SODASORB CO2 absorbent manufactured prior to August 2010 may have missing or illegible expiration dates and/or batch numbers. | Packaging process control | Darex/W.R. Grace sent a Customer Notification to their sole direct account of the recall and the account provided Darex/W.R. Grace with lists of their customers who received the recalled products. Darex then sent undated recall letters to those accounts on 10/20-22/10. The accounts were informed of the product, the problem, and the action that should be taken. The accounts were warned not to use Pre-Pak units with missing or illegible expiration dates and/or batch numbers and were requested to examine their inventories of Pre-Pak units, placing all units that had missing or illegible expiration dates and/or batch numbers on hold for destruction and replacement. The accounts were instructed to call W.R. Grace at 1-800-492-7632 to arrange for replacement. The accounts were also requested to complete and return the reply postcard, indicating how much product was destroyed. |  | 8096 | 8,096 cartons | Nationwide Distribution | Terminated | September 21, 2011 | 336 | 103638
Z-0557-2011 | 57176 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Custom Kit, Sterile EO, Rx Only, Catalog No. K09-09815, Manufacturer: Merit Medical Systems, Inc. South Jordan, Utah 84095, USA, |  |  | October 25, 2010 | 2010 | Merit Medical Systems, Inc. | Defective custom kits may draw in air during use which may compromise patient safety and affect treatment. The valve's fixed luer may be over- or under-tightened, causing leaking at the bond joint. | Process control | Merit issued an Urgent Product Recall Notice letter dated October 25, 2010 to consignees. Copnsignees were instructed to quarantine the product immediately for return and to notify customers of recall activity and instruct them to return product.. All distributors and customers are to complete and return a Recall Notification Form.For further information about this recall, contact Merit at 800 356-3748. |  | 1033 | 1033 units | Nationwide Distribution in the states of CA, IL, IN, and UT. | Terminated | March 23, 2012 | 515 | 6943
Z-0584-2011 | 56250 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active I-PTH ELISA, Part Number: DSL-10-8000 |  |  | April 22, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the stated stability of the DSL-10-8000 I-PTH ELISA and DSL-8000 I-PTH IRMA assays may not be achieved for the lot numbers affected by the recall. The affected I-PTH ELISA and the I-PTH IRMA kit lots may produce falsely elevated patient results. Internal stability testing showed an average 16% elevation for I-PTH ELISA lot 99 | Pending | Beckman Coulter sent an Urgent Product Corrective Action Letter to their consignees dated June 18, 2010. The letter identified the product, the problem, and the action to be taken. Consignees were instructed to:(1) Discontinue use of all materials provided with I-PTH ELISA kit lot number 992751 and I-PTH IRMA lot numbers 090013, 090121, and 090372.(2) Review historical external quality control results to ensure the assay performance of these lots was within laboratory specifications. Retesting may be necessary if their laboratory uses only the Control materials provided with the I-PTH ELISA and IRMA kits.(3) Abnormal elevated patient results that are obtained with these lot numbers and that are inconsistent with the patient's clinical status should be reviewed in the context of other clinical diagnostic tests.(4) I-PTH ELISA kit lot number 992751 and I-PTH IRMA lot numbers 090013, 090121, and 090372 will no longer be shipped. Due to the calibrator instability recently encountered with DSL-10-8000 I-PTH ELISA and DSL-8000 I-PTH IRMA assays, Beckman Coulter has decided to immediately discontinue these two part numbers.Consignees were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that they received this important notification.If they have any unused product the were asked to contact the numbers listed below:(1) In the United States, contact Customer Service at 1-800-526-382.1(2) Outside of the United States, consignees were asked to contact their local Beckman Coulter RepresentativeConsignees were told to contact DSL Technical Support Center at 1-800-231-7970 in the United States, or their local Beckman Coulter Representative with any questions.For questions regarding this recall call (714) 961-3634. |  | 3 | 3 kits | Worldwide Distribution - USA including MD, NY, GA, NJ, CA, MI, TX, PA, IA, and IN. Internationally to Austria, Brazil, China, Colombia, Czech Republic, France, Germany, Hong Kong, India, Japan, Mexico, and the Russian Federation | Terminated | May 07, 2012 | 746 | 199
Z-0574-2011 | 56731 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ZimmerÂż Trabecular Metal Humeral Stem Surgical Technique. Catalog Number: 97-4309-002-00, Rev. 3 |  |  | September 10, 2010 | 2010 | Zimmer Inc. | The results of an investigation into complaints revealed that Trabecular Metal Humeral and/or ZimmerÂż Trabecular Metal Reverse Humeral implants engaging in the canal earlier than expected could lead to incomplete seating of the implant, delay in surgery, humeral fracture, or the need for cement in a press-fit application. | Device Design | Zimmer, Inc. sent an URGENT DEVICE CORRECTION letter dated September 10, 2010 to surgeons, risk managers, and distributors. The letter identified the product, the problem, and the action to be taken. Sales Representatives were instructed to:a. provide the Dear Surgeon (Urgent Device Correction) letter and revisedsurgical technique directly to surgeons and risk managers at facilities usingthe device orb. provide a list of facilities and surgeons using the TM Humeral and TMReverse shoulder systems to the Zimmer Recall Coordinator, andc. destroy previous versions of the surgical technique, andd. inform users of when the new surgical technique, instrument case, distalreamers, and distal pilots will be shipped to them based on communicationfrom Zimmer, Inc.Handling of returned product:HCPs were notified to return all of the distal pilots in the scope back toZimmer which were to be be destroyed once the recall has been completed. They willalso be instructed to either return the reamers back to Zimmer or to remove these reamers from the TM Reverse shoulder systems sets and place into theBigliani/Flatow Complete Shoulder Solution set. New surgical techniques will also be introduced and the user will be asked to remove previous versions of the surgical technique from all points of use. This removal will be verified byZimmer Sales Representatives.For questions regarding this recall call (574) 371-8852. |  |  |  | Worldwide Distribution - USA, Nationwide, Switzerland and Eschbach, Germany, Switzerland, Australia, China, India, Japan, Korea, Malaysia, Singapore, Taiwan, Brazil, Venezuela, El Salvador, Colombia, Canada, and Thailand. | Terminated | September 14, 2012 | 735
Z-0582-2011 | 57157 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | McKesson Medical Surgical 107-101 Urine Analyzer, Catalog Number: 107-101 |  |  | October 01, 2010 | 2010 | Teco Diagnostics | This recall has been initiated due to the CLIA-101 Urine Analyzer not having 510(k) clearance by the FDA for use in measuring urinary glucose or occult blood with the CLIA waived, Inc. CLIA-10 test strips. Teco Diagnostics has been notified bythe FDA that the CLIA-101 requires 510(k) clearance and is not CLIA Waived for use with the CLIA-10 test strips. Without FDA approval or clearance, Teco Di | Other | Teco Diagnostics initiated an "Urgent: Medical Device Recall" letter dated October 14, 2010, with an attached "Recall Return Response Form" to the affected consignees via US Postal on October 15, 2010. The recall notice informed the consignees of the reason for recall, the products affected and the actions required.Consignees were instructed to complete and return the enclosed Recall Return Response Form as soon as possible via mail, fax, or email. Upon receipt of the Recall Return ResponseForm, Teco Diagnostics would work with the consignees/distributors to process any refunds or reimbursements for the returned product and return shipping. If consignees/distributor had questions, they were to call Teco's Technical Support Department at (800) 222-9880 or (714) 463-1111. |  | 2140 | 2,140 units total for all products | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NM, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI, AND WV and the country of Canada. | Terminated | August 03, 2012 | 672 | 2140
Z-0587-2011 | 57089 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Three-Way Large Bore Stopcock with Male Luer Slip; a sterile, nonpyrogenic fluid pathway; Baxter Healthcare Corporation, Deerfield, IL 60015; product code 2C6202. |  |  | October 28, 2010 | 2010 | Baxter Healthcare Corp. | Rotating the handle on the 3-Way Large Bore Stopcock past the stop can cause the stopcock to leak. A leak can lead to under-infusion or interruption of critical life-sustaining therapy. | Device Design | Safety Alert letters dated October 28, 2010 were mailed via first class mail to the direct accounts, to the attention of the Director of Nursing and the Director of Materials Management, on the same date, informing them that rotating the handle on the 3-Way Large Bore Stopcock past the stop can cause the stopcock to leak. A leak can lead to under-infusion or interruption of critical life-sustaining therapy. The accounts were requested to post the enclosed attachment, which illustrates the correct direction to turn the handle of the stopcock, where the stopcocks/sets are dispensed. The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at (847) 270-5457. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. The Baxter Field Corrective Action Gatekeepers were also notified of this action via e-mail on 10/27/10, and instructed to notify their customers and ministries of health of the recall.Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  |  |  | Worldwide distribution: USA and countries including Puerto Rico, Guam and American Samoa, and internationally to Canada, Bermuda, Jamaica, Guatemala, Panama, Bahrain, United Kingdom, Hungary, Lebanon, Pakistan, Saudi Arabia and New Zealand. | Terminated | November 01, 2012 | 735
Z-0569-2011 | 57315 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | MONOJECT Prefill Heparin Lock Flush Syringe |  |  | November 15, 2010 | 2010 | Covidien LP | The heparin sodium USP used to manufacture was produced with crude heparin containing Over-Sulfated Chondroitin Sulfate. | Material/Component Contamination | COVIDEN sent an URGENT PRODUCT RECALL letter dated November 12, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and in-use stock to determine if they had any units of the affected product and quarantine the product immediately. Distributors were instructed to notify their customers immediately of the URGENT PRODUCT RECALL letter. Customers were to fill out the enclosed response form and return it by fax to (203) 822-6009 or email to adfeedback@coviden.com following the instructions on the form. For questions customers were to contact Customer Service at 1-800-962-9888 |  | 1539540 | 1,539,540 | Nationwide Distribution including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, LA, KS, KY, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and Bermuda | Terminated | December 04, 2015 | 1845 | 3510660
Z-0551-2011 | 57045 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Dixtal Handheld Pulse Oximetry Model 512 |  |  | September 27, 2010 | 2010 | Dixtal Medical, Inc. | Unit may overestimate Sp02 levels in a high humidity environment. | Equipment maintenance | Dixtal Medical notified accounts by email on 9/27/10 and followed by hard copy Urgent Product Recall letter, dated September 27, 2010. The letter identified the affected product and stated the reason for the recall. The letter also requested that customers to cease distribution and/or use of the suspect product and to return product for replacement. The attached Confirmation Notification should be completed and returned. Customers are to contact Customer Service: sales@dixtal.com, Phone 1-203-269-1112, Fax 1-203-269-1760, with questions or to arrange to have a replacement product provided. |  | 4 | 4 units | Worldwide Distribution -- USA, including states of AK, KY, NJ, and NY and countries of Australia, Austria, Bolivia, Canada, Finland, Indonesia, Italy, Poland, Sweden, and Switzerland. | Open, Classified |  |  | 39
Z-0589-2011 | 55977 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | GE Responder 2000 Defibrillator/Monitor, models 2025653, 2026109, and 2026114 that were manufactured between July 25, 2008 and October 2, 2009 by GE Healthcare. |  |  | April 28, 2010 | 2010 | Cardiac Science Corporation | Some units may be non-operational during a rescue, due to a defective PCB interconnect on traces connected to the reed relay. This will result in an inability to analyze the patient's heart rhythm or deliver therapy during a rescue. | Nonconforming Material/Component | Cardiac Science issued a Urgent Medical Device Corrective Action letter dated April 28, 2010 requesting GE Healthcare cease distribution of affected devices and to instruct end users of distributed devices to schedule a service call to have the device corrected. Cardiac will send rework kits no later than June 30, 2010. CSC can be contacted at 44 161 926 0011 to arrange return material authorization and replacement of affected devices. |  | 130 | 130 units | Worldwide Distribution: Throughout the US, and in the countries of Germany, Korea, Peru, Singapore, and Spain. | Terminated | December 08, 2011 | 589 | 130
Z-0602-2011 | 54143 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Spinal Concepts polyaxial screw driver, AO, Manufactured by Spinal Concepts, Austin, TX, REF 775-01. |  |  | December 01, 2005 | 2005 | Zimmer Inc. | The devices may fracture, resulting in surgical delay or the need for intervention to remove fragments from the surgical site. | Device Design | The affected instruments were removed and replaced in the U.S. Abbott Spine and Spinal Concepts instrument kits between December 2005 and October 2007. User accounts were not advised of the removal Zimmer has confirmed that none of these instruments are currently available for use in the U.S.For further information about this action, contact Dale Miller of Zimmer at 574 372-4962. |  | 220 | 220 | Worldwide Distribution, including US. | Terminated | June 01, 2011 | 2008 | 273
Z-0588-2011 | 55571 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension LIP Flex Reagent Cartridge, Catalog# DF55A |  |  | March 10, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Open Well instability: After calibration, QC shifts high (out of range) after the Flex Well is hydrated. Shifts of up to 20% have been observed in both QC and patient samplesas soon as two hours post hydration. | Other | Recalling firm issued "Urgent Field Safety Notice", dated March 2010, informing users of the problem and to transition to the revised Lipasae method. |  | 21171 | 21171 | Nationwide Distribution throughout the US. | Terminated | August 04, 2011 | 512 | 21171
Z-0594-2011 | 56832 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FUSO Analyzer 300F |  |  | September 20, 2010 | 2010 | Abbott Point Of Care Inc. | i-STAT Analyzers batteries may overheat due to a circuitry failure, resulting in the analyzer becoming uncomfortably hot in the area of the battery component. | Other | Abbott Point of Care issued Urgent Recall Notice letters dated September 20, 2010 to customers. The letter informs users to of the possibility of the unit becoming hot to touch, to remove the unit if this is the case, not to remove batteries unless instructed to do by APOC and to forward notification of the recall to customers if it has been further distributed.Abbott Point of Care can be contacted at 800 366-8020, option 1, about this action.Recall extended in May 2011 to included distribution dates of 9/20/2010-6/6/2011. |  | 2202 | 2202 | Worldwide Distrbution: Throughout USA and the countries of Argentina, Australia, Austria, Belgium.Bolivia, Bosnia-Herz, Canada. Croatia, Chile, Colombia, Costa Rica, Denmark, Ecuador, El Salvador, Finland, Georgian Republic, Greece, Honduras, France, Germany, Italy, Netherlands, New Zealand, Hong Kong, India, Ireland, Israel, Japan, Kingdom of Saudi Arabia, Korea, Kuwait, Mauritius, Mexico, Nigeria, Norway, Panama, Paraguay, Peoples Republic of China, Peru, Poland, Portugal, Qatar, Russia, Singapore, South Africa, Spain, Sweden, UK, Sultanate of Oman, Switzerland, Taiwan, R.O.C., Thailand, Turkey, UAE, Uruguay, and Yemen. | Terminated | January 28, 2013 | 861 | 1295561
Z-0592-2011 | 56864 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microtek Medical, Inc., Microscope Drape for Zeiss Standard |  |  | August 23, 2010 | 2010 | Microtek Medical, Inc. | This product is being recalled due the product is miss-labeled as being sterilized by EO processing; however, the product was processed using Gamma sterilization. | Error in labeling | On 08/23/10 Microtek Medical, Inc., Jacksonville, Florida notified all of their customers via a telephone call and requested immediate return of the product . The firm followed up with each of the customers by sending out an Urgent: Medical Device Recall letter (written notice), dated September 02,2010, to each on 9/8/10. The letter identified the affected product, stated the reason for the recall, and asked that the product be returned immediately. Customers are to complete and return the enclosed response form. Anyone needing further assistance is asked to contact Elizabeth Jolly, V.P. Legal at 1-888-844-0988. |  | 60 | 5 Cases, 12 drapes per case, 60 Drapes total distrubuted | Nationwide Distribution | Terminated | December 14, 2010 | 113 | 60
Z-0600-2011 | 57096 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Medefil Heparin I.V. Flush Syringe 10 units/mL; a polypropylene luer locking syringe containing Heparin Lock Flush Solution, USP, a sterile non-pyrogenic, isotonic injection; each mL in a Heparin I.V. Flush Syringe contains 10 units of heparin sodium derived from porcine intestinal mucosa and 9 mg of sodium chloride in water for injection; Medefil, Inc. Glendale Heights, IL 60139. The pre-filled syringes were packed as 60 syringes per dispenser box, 16 boxes per shipping carton in the following syringe configurations: a) NDC# 64253-222-21, catalog # MIH-2221: 1 mL fill in 6 mL syringe; b) NDC# 64253-222-23, catalog # MIH-2223: 3 mL fill in 6 mL syringe; c) NDC# 64253-222-35, catalog # MIH-2235: 5 mL fill in 12 mL syringe |  |  | November 01, 2010 | 2010 | Medefil Incorporated | The heparin lock flush solution was manufactured with contaminated Heparin Sodium USP Active Pharmaceutical Ingredient (API). | Material/Component Contamination | Medefil faxed/emailed the direct accounts on 11/1/10, and sent the accounts a hard copy of the Urgent Product Recall letter, dated 11/1/2010, on 11/2/10. The accounts were informed of Medefil's recall of 18 lots of Heparin I.V. Flush Syringes produced from the active pharmaceutical ingredient (API) Heparin Sodium, USP, lot 1035-0780, that has been recalled by the manufacturer, Scientific Protein Laboratories, due to the presence of a contaminant, Oversulfated Chondroitin Sulfate. The accounts were instructed to stop use of the 18 lots of Heparin I.V. Flush Syringes immediately, place them under quarantine, and contact Medefil to arrange for the product to be returned. The accounts were requested to complete and return to Medefil the enclosed response form, indicating the quantity on hand of each recalled lot. Any questions were directed to Pradeep Aggarwal, President at 630-682-4600. |  | 583500 | 583,500 syringes | Nationwide Distribution --- including states of Arizona, California, Florida, Illinois, Ohio, Louisiana, Massachusetts, Michigan, Minnesota, New York, Texas, and Virginia. | Terminated | February 13, 2012 | 469 | 2919180
Z-0590-2011 | 57011 | No | LWE |  |  | General Hospital | Enforcement Discretion | Orthofix Bone Marrow Aspiration Needle, Catalog number 21-5000, 21-5011, Distributed By: Orthofix Inc, 1720 Bray Central Drive, McKinney, TX 75068-8207. Product includes handle, stylet, cannula, probe, probe guide, and universal cap. Cannula is available as an 8 gauge needle or 11 gauge needle. |  |  | October 18, 2010 | 2010 | Orthofix Inc | Outer cartons for the product may be labeled with incorrect needle gauge. | Other | Orthofix issued an Urgent Product Recall letter dated October 14, 2010 to consignees. Customers were instructed to inspect their inventory, and remove and return affected product to the firm. Replacement needles will be provided. Orthofix account representatives can be contacted at 1 888 298 5700, and Orthofix Regulatory Affairs can be contacted at 1 214 937 2061. |  | 274 | 274 | Worldwide Distribution - USA, including the states of AL, AR, CA, CT, FL, LA, MI, MO, NC, NJ, NY, PA, SC, and TX, and the country of Spain. | Terminated | February 28, 2011 | 133 | 274
Z-0591-2011 | 57216 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CHROMOPHARE Surgical Lighting System &amp; CHROMOVIEW Flat Panel Monitor, Mechanical Spring Arms. |  |  | November 16, 2009 | 2009 | Berchtold Corp. | A mechanical failure of a welded joint within the spring Arm may result in the rapid drop of a monitor to the mechanical stop. | Other | Berchtold issued a Product Safety Notification letter on November 16, 2009 identifying the affected device and issue prompting the action. Consignees were instructed to:Exercise caution when adjusting the position of the surgical lights or flat panel monitors. While an event is highly unlikely, the likelihood does increase when the spring arm is being repositioned. Identify a specific contact person for their facility to facilitate the communication and field action process with BERCHTOLD. The following action is being planned: To further identify if their units are part of the affected population, determine the supplier serial number identification and specific location for each spring arm in their facility. Report the facility information, supplier spring arm serial numbers and specific location or spring arms to a BERCHTOLD. Detail instructions on serial number identification and the method of reporting to BERCHTOLD will also be provided in a subsequent communication. Once BERCHTOLD has received the serial numbers of their spring arm assembly, customers will be notified whether their particular units are affected and the further required action.Berchtold Corporation can be contacted at 1 800 968 6097 about this action. |  | 7031 | 7031 arms | Worldwide Distribution: Throught the U.S,.and the countries of Canada and Germany | Terminated | July 06, 2012 | 963 | 7031
Z-0614-2011 | 51242 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Heartstring II Proximal Seal / VASOVIEW 6 Endoscopic Vessel Harvesting System, |  |  | February 05, 2009 | 2009 | CTS, Inc. dba Guidant Cardiac Surgery | Sterility may be compromised by a loss of package integrity. | Packaging | Maquet issued an Urgent Device Removal letter dated February 5, 2009 to customers, explaining the reason for the recall and requesting the product be returned for replacement. Customers were instructed to examine their inventory and discontinue distribution and use of the affected devices. Customers also were requested to complete and return a Field Action Response Form.Maquet can be contacted at 888-880-2874 concerning this recall. |  |  |  | Distribution: Nationwide and Internationally. | Terminated | January 19, 2011 | 713
Z-0608-2011 | 55815 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Quest EIA ANA Screen Bulk Kit, model number 96AN-BPU-QUEST. Manufactured by Bio-Rad Laboratories, Inc., Hercules, CA. For labeling and distribution by Quest International Inc. |  |  | February 26, 2010 | 2010 | Bio-Rad Laboratories Inc | The value for the ANA Cutoff and ANA Positive Controls will decrease overtime resulting in an unusually elevated positivity rate for patient samples tested. | Other | Bio-Rad Laboratories, Inc. sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated February 24, 2010, to their customer. The letter identified the product, the problem, and the action to be taken. The customer was instructed to:Examine inventory and quarantine affected product. Identify all customers that have received the affected product and immediately notify them. Instruct them to discontinue use of the affected product.Complete and return the enclosed MEDICAL DEVICE RECALL RESPONSE FORM.For questions regarding this recall call (510) 724-7000. |  | 453 | 453 kits | Nationwide Distribution including FL | Terminated | January 05, 2011 | 313 | 453
Z-0609-2011 | 55363 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | NexGen Complete Knee Solution Stemmed tibial broach impactor, Catalog # 00-5977-011-00, Zimmer Inc., Warsaw, IN. |  |  | April 16, 2010 | 2010 | Zimmer Inc. | The impactor may disassemble or detach during use. The detached piece may not be observed intraoperatively and may be left in the surgical site. Complications may include pain, infection, and prolongation in surgery. These complications may also require revision surgery. | Device Design | Zimmer issued "Urgent Medical Device Safety Alert" letters dated April 16, 2010 to Risk Managers, user facilities, Health Care Professionals and Zimmer Sales force. The users were informed to inspect the device for loosening. The alert also informed users to inspect the device after each use to verify that the device was intact. Customers were instructed to complete and return a User Facility Documentation Form. Zimmer will repair and replace all affected devices.Customers can contact ZImmer, Inc. at 800 613-6131 concerning this recall. |  | 10666 | 10,666 all products. | Worldwide Distribution: Throughout US including Puerto Rico, and the countries of Argentina, Australia, Belgium, Bolivia, Brazil, Canada, Chile, China, Columbia, Denmark, Ecuador, El Salvador, France, Germany, Guatemala, Hong Kong, India, Italy, Japan, Korea, Mexico, Netherlands, Nicaragua, Panama, Peru, Portugal, Singapore, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine, and United Kingdom. | Terminated | September 12, 2012 | 880 | 10666
Z-0613-2011 | 56386 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio-Rad brand Meridian ANA Screen ELISA Test Kit, 96 Test, |  |  | February 24, 2010 | 2010 | Bio-Rad Laboratories Inc | False positive test results caused by decrease in the OD values of ANA Cutoff Controls and ANA Positive Controls. The increased positivity rates were associated with the Meridian ANA Screening kit. False positive ANA values could result in misdiagnosis or mistreatment . | Other | Bio-Rad issued an Urgent Medical Device Recall NOtification letter dated February 24, 2010 to their distributor, identifying the affected product and actions to be taken. The distributor was instructed to examine inventory and quarantive affected kit lots, and to identify and notify customers with affected lots to discontinue use and discard the product. Replacement kits will be distributed. Customers were asked to complete and return a recall response form.Bio-Rad can be contacted at 510 724-7000 concerning this recall. |  |  |  | Nationwide Distribution: to Meridian Diagnostics in Ohio. | Terminated | January 10, 2011 | 320
Z-0607-2011 | 56752 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems, Part Number: 81600 ; Access 2 Immunoassay Systems, Part Number: 81600N; |  |  | August 13, 2010 | 2010 | Beckman Coulter Inc. | This recall was initiated because Beckman Coulter has received reports of erroneous test results being reportedas a result of assay reagent packs being transferred between systems. Erroneous results will be generated as a result of missing or incorrect assay reagents. In these casesboth false positive and false negative results are possible, depending upon the assay configuration. | Pending | Beckman Coulter, Inc. sent two (2) Product Corrective Action (PCA) letters (Access letter and DxI letter) with attached Customer Response form the week of August 09, 2010 to the affected customers. The letters were dated August 13, 2010, and identified the product, the problem, and the action to be taken by the customers. The letter reminded customers that it is critical that they load reagent packs according to the procedure provided in their Access or Access 2Operators Guide, Instructions for Use, or Help system. A copy of the procedure was included with the letter. Customers were reminded that all operators should be properly trained on these important procedures andunderstand the impact on reported results if these procedures are not followed.Customers were instructed to share this information with their laboratory staff, and retain this notification as part of their laboratory Quality System documentation. Customers who needed assistance or had any questions regarding this notification, were instructed to contact Technical Supportat 1-800-854-3633, option 1 in the United States or Canada. |  | 9571 | 9571 units (instruments) | Nationwide Distribution &amp; Canada | Terminated | July 11, 2012 | 698 | 9571
Z-0611-2011 | 56849 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | RotoProne Therapy System, Model 209500, Device List Number E232632 |  |  | September 09, 2010 | 2010 | KCI USA, Inc. | Product complaints indicating that while attempting to resolve an alarm or error condition, operator was unable to return product to the prone position for patient. | Other | KCI USA, Inc. sent an URGENT - Medical Device Safety Alert letter, dated September 8, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the cusotmer. Customers were instructed to have a practiced back-up plan, such as a plan for the manual proning of patients in the event that clinical circumstances dictate the need for more immediate prone therapy. For questions, customers in the US should call KCI's Technical Support Center at 1-800-275-4524, option 5. For customers outside the US they should contact their local sales representative. Customers should contact their accounts with affected product and advise them of the safety notice. |  | 240 | 240 units | Worldwide Distribution - USA Nationwide, United Kingdom, Germany, and Saudi Arabia | Terminated | March 21, 2011 | 193 | 240
Z-0615-2011 | 57214 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Incorrect Label: Endocare 2.4mm VARIABLE ATTACHED CRYOPROBE, REF: CVA 2400, LOT 2751. Correct Label: Endocare 2.4mm CRYOPROBE, REF: CRYO-44F (Sharp Tip). |  |  | November 04, 2010 | 2010 | Endocare, Division Of Healthtronics | Mislabeled : The label on the outer box and inner pouch of the EndoCare 2.4mm CryoProbe (CRYO-44F) inaccurately reference the CVA2400 (2.4mm Variable Probe aka: V-Probe). | Other | Healthronics issued an Urgent Device Recall letter dated November 2, 2010 to consignees. The letter described the product mislabeling, and requested customers to examine their inventory, quarantine all affected product, and complete and return the verfication section of the letter to the firm. A return authorization number and prepaid return label were provide to facilitate return and replacement.Healthronics can be contacted at 877 367-4873 concerning this recall. |  | 33 | 33 Units | Nationwide Distribution: to the states of FL, MI, PA, TN, and WA. | Terminated | April 14, 2011 | 161 | 33
Z-0612-2011 | 56126 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio-Rad Laboratories Autoimmune EIA ANA Screening kit, 96 test, model 96AN, Distributed and Manufactured by Bio-Rad Laboratories, Inc. |  |  | June 24, 2010 | 2010 | Bio-Rad Laboratories Inc | Product mislabeled-- the ANA cutoff Control and Prostiva controls were mislabeled. Use of the incorrect controls will cause run failure - patient results cannot be reported because controls will not pass.. | Labeling mix-ups | Bio-Rad issued Medical Device Recall Notification letters dated June 24, 2010 to consignees, identifying the affected products and requested that all affected product be disposed of in accordance with local procedures. Response forms are to be completed and returned to the firm. Bio-Rad can be contacted for replacement product and about this recall at 510 741-6114. |  | 37 | 37 units | Nationwide Distribution: including the states of California, Ohio, Massachusetts, and Missouri.Product delivered to 11 consignees only in the US. | Terminated | August 08, 2011 | 410 | 37
Z-0618-2011 | 53182 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Titanium Intraline Anchor, Model Numbers 3910400050, 3910400055, 3910400060, 3910400065, Soft Eyelet Anchor, manufactured by Stryker Endoscopy, Puerto Rico |  |  | August 20, 2009 | 2009 | Stryker Endoscopy | Potential breakage -- product may break while inserting it into the bone during surgery. This may prolong surgery, or, if undetected during surgery, result in retained fragments that may cause metal sensitization or immune response, or necessitate further intervention to retrieve the titanium fragments. | Material/Component Contamination | Stryker issued an Urgent Device Removal notification letter dated August 21, 2009 to consignees. Consignees were asked to acknowledge receipt and distribute the letter, track reponses, and destroy any returned material and document through weekly summary and records.Stryker can be contacted at 408 754-2000 about this recall. |  | 447 | 447 units | Nationwide Distribution: to one distributor in NJ. | Terminated | December 14, 2010 | 481 | 447
Z-0619-2011 | 57304 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Inter-Op Acetabular System Hooded Trial Insert size 32mm/57mm, non-sterile, REF 9363-32-057, Zimmer, Warsaw, IN. |  |  | November 15, 2010 | 2010 | Zimmer Inc. | Mislabeled - Device packages labeled as hooded trial inserts contain standard trial inserts. | Employee error | Zimmer contacted the two affected accounts directly by phone on November 15, 2010 to inform them of the recall. Zimmer will monitor the return of the affected instruments and account for the outstanding instruments until the removal action is complete. Each of the affected accounts will be provided with a new trial insert, manufactured and packaged with the correct component. The accounts will be asked to return the affected units back to Zimmer where the returned product will be destroyed. Customers can contact Zimmer at 574 371-8852 concerning this recall. |  | 3 | 3 | Worldwide Distribution: US in the states of KS and UT, and the country of Canada. | Terminated | May 23, 2011 | 189 | 3
Z-0620-2011 | 57370 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Flat Panel Display is a touchscreen display, 19-inch thin film transistor liquid crystal display (TFT-LCD). The PNs are 010-1619-00 and 010-1655-00, Models 94260-19 and 94256. |  |  | March 20, 2009 | 2009 | Spacelabs Healthcare, Llc | The power supply of the Spacelabs Medical Flat Panel Display may fail, causing the display to go blank. | Other | Spacelabs Healthcare sent an URGENT: MEDICAL DEVICE RECALL Spacelabs Medical Flat Panel Power Supply PNs: 010-1619-00 &amp; 010-1655-00 letter dated March 20, 2009 to US consignees. On March 27, 2009, Spacelabs sent the same recall letter with translation option to their international consignees.The letter stated the firm would be sending a replacement power supply. While waiting for the replacement of the power supply, Spacelabs Medical recommended their consignees brief the staff on the possibility of the display going blank to help prevent any adverse event from happening. The consignees were advised to dispose the old power supply per their premise protocol or local regulations. The consignees from US can contact Spacelabs Medical at 1-800-522-7025, #2 for Technical Support , and consignees outside the United States can call 1-425-657-7200 x5089 for any questions about the recall. |  | 3608 | 2877 units in US and 731 units outside US | Worldwide Distribution: USA in the following states: AK, AL, AR, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, PA, RI, SC, SD, TN, TX, VA, WA, WI, and WY, and including the following countries: AUSTRALIA, AUSTRIA, CANADA, Caribbean, CZECH REPUBLIC, EGYPT, ESTONIA, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, ITALY, JORDAN, KUWAIT, LEBANON, MEXICO, NETHERLANDS, PAKISTAN, PANAMA, POLAND, PUERTO RICO, SAUDI ARABIA, SINGAPORE, SWITZERLAND, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED KINGDOM, and VENEZUELA. | Terminated | June 13, 2011 | 815 | 3608
Z-0617-2011 | 57223 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The PercuNav MP008R-B Button Probe is designed to be used with Philips (formerly Traxtal) PercuNav Image Guided Intervention System. Some of the clinical applications of the Button Probe require sterility and others may not. However, it is possible that patient contamination could occur if a non-sterile Button Probe is used in clinical applications where aseptic technique is required. |  |  | November 10, 2010 | 2010 | Philips Healthcare Inc. | The protective cover over the tip of the MP008R-B Button Probe may fall off, causing the probe tip to damage the packaging and compromise the sterility of the probe. | Other | On 11/10/10, Philips issued URGENT - Medical Device Recall notification letters dated November 10, 2010 to their consignees.Consignees are advised to examine their inventory of MP008R-B Button Probes to confirm the probe tip cover is properly placed over the tip of the probe. They should discard any Button Probes that have been confirmed to have the protective probe tip cover removed from the Button Probe tip. Consignees are also reminded to follow the MP008R-B Button Probe and PercuNav Instructions for Use prior to using the device. Do not use the probe if the unit package is damaged, broken or soiled.The firm's representative will contact the consignees to confirm quantities and Lot Codes of MP008R-B Button Probes that the firm may have examined or discarded. The firm estimates replacement probes will be available in the first quarter of 2011.Consignees can contact Philips Customer Service at 800 722-9377 and reference the FCO 79500189. |  | 25 | 25 units | Nationwide Distribution: in the states of FL, IL, MA, NE, NY, SC, and VA. There are no foreign consignees and US government consignees. | Terminated | August 15, 2011 | 278 | 25
Z-0621-2011 | 57269 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ENCORE (R) Microptic (R) Powder-Free Latex Surgical Gloves, 1 Pair, Size 5Âż, Sterile R, CE 0086, EC REP Ansell (UK) Ltd., Tamworth, Staffordshire, B79 7UL UK, Ansell Lanka (pvt) Ltd., Biyagama Export Processing Zone, Biyagama, Sri Lanka |  |  | September 03, 2010 | 2010 | Ansell Healthcare Products LLC | open package seams on sterile gloves | Other | The firm, Ansell, sent a letter dated September 3, 2010 to all customers. The letter described the product, problem and actions taken. An Ansell representatives will contact the customer to remove the gloves bearing the noted lot number from their facility; Ansell will ensure that the gloves removed will be replaced shortly with products from a different manufacturing batch and the recall was extended to the retail (hospital and clinic) level requesting that the product be returned. If you have any questions, contact US (334) 615-2563, international telephone 01.34.24.52.52 and email: infofrance@eu.ansell.com. |  | 4000 | 4000 units | International distribution only: France and Germany | Terminated | November 28, 2011 | 451 | 4000
Z-0678-2011 | 55861 | No | JEY | Bone plate. | Dental | Dental | 510(k) | 2.8 mm Inion CPS, 2.8 x 18 mm screw, REF SCR-1216, sterile, Manufacturer, INION Oy, Finland. |  |  | April 14, 2010 | 2010 | Stryker Craniomaxillofacial Division | Inadequate Instructions for Use. The stated strength retention is ambiguous and the maximum immersion time of the plates/meshed is not defined. | Other | Stryker issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 14, 2010. The notice informed users of the change in the instructions for use (IFU), and included a copy of the revised IFU. The consignees were informed that the affected product does not need to be returned.Stryker can be contacted about this action at 877-534-2464 (option "Regulatory Affairs", reference Product Field Action #2009-445) |  | 8631 | 8631 all products | Nationwide Distribution: Throughout the US | Terminated | June 06, 2012 | 784 | 8631
Z-0624-2011 | 56659 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Cmax Surgical Tables, Model #'s: Cmax110, Cmax 220 manufactured by Steris Corporation in Montgomery, AL |  |  | August 12, 2010 | 2010 | Steris Corporation | Cmax Surgical Tables may lose calibration if the table is operated only on battery power and if the table is in a low battery charge condition. This could result in the table moving into an unexpected tilt position when the Level button is depressed on the hand control keypad. | Other | Steris issued Urgent Field Correction Notice letters dated August 26, 2010 to each affected consignee notifying them of the voluntary field correction. A Steris Representative will contact customers to arrange for STERIS to upgrade the software in the table.For further information about this corrective action, contact Steris at 1 800 333-8828. |  | 1635 | 1,635 tables | Worldwide Distribution: USA in the following states: AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and WV, and in the countries of Australia, Canada, China, Columbia, Dominican Republic, India, Israel, Mexico, New Zealand, Saudi Arabia, Singapore, South Korea and Taiwan. | Terminated | September 27, 2012 | 777 | 1635
Z-0627-2011 | 56716 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Wingspan Stent System, Model Number M003WE0300150, Manufactured by Boston Scientific, Fremont, CA |  |  | August 25, 2010 | 2010 | Boston Scientific Corporation | Products may not meet radial force specifications. | Process control | Recall initiated August 25, 2010. "Urgent Medical Device Recall - Immediate Action Required" notifications, dated 25 August 2010, were delivered by FedEx where possible, and by personal delivery assigned to sales representatives. The letter identified the affected product and the reason for recall. The letter stated that further distribution or use of any remaining product should cease immediately. Customers were asked to read the enclosed instructions and told that their local Sales Rep could answer any questions that they may have regarding the recall. |  | 10 | 10 units | International Distribution Only -- China, Germany, Italy, the Philippines, Portugal and Turkey. | Terminated | December 17, 2010 | 114 | 10
Z-0629-2011 | 56757 | No | DLJ | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | ARCHITECT iVancomycin Reagent; an in vitro chemiluminescent microparticle immunoassay (CMIA) for the quantitative measurement of vancomycin in human serum or plasma on the ARCHITECT i System with STAT protocol capability; Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064 USA; list 1P30-25 |  |  | September 13, 2010 | 2010 | Abbott Laboratories | The assay may produce falsely elevated vancomycin results. Inaccurate test results may result in inappropriate patient treatment. | Other | Abbott Laboratories issued a Product Correction letters dated September 13 and 15, 2010 to all customers that received ARCHITECT iVancomycin reagents in the last 12 months. In the September 13, 2010 letter, the accounts were informed of the falsely elevated results and were instructed to continue to follow all sample handling instruction in the package insert and to take the following additional actions to mitigate the risk of obtaining an incorrect result:1. After following the primary tube manufacturer's recommended centrifugation instructions transfer serum or plasma for all specimens to a microcentrifuge tube for centrifugation at 150,000-300,000 g-minutes, following centrifugation guidelines in the chart in the letter.2. Prior to testing, transfer the serum or plasma to a sample cup or secondary tube for testing. Care must be taken to transfer only the clarified specimen and not the lipemic material or pellet.The accounts were also requested to complete and fax back to Abbott the enclosed customer reply sheet, acknowledging the receipt and understanding of the recall. On September 15, 2010, Abbott discovered that the calculation for the determination of RCF (xg) given in the September 13, 2010 letter was incorrect. Incorrect Calculation previously provided: Relative Centrifugal Force (RCF) = 0.0000118 x rotational radius in centimeters x RPM2 Correct Calculation within revised letter: Relative Centrifugal Force (RCF) = 0.00001118 x rotational radius in centimeters x RPM2 Abbott stopped distribution of the previous letter and initiated redistribution of the revised letter to all impacted customers, globally, on September 15, 2010. Additionally, all customers which had already received the initial letter were sent a cover letter identifying the error.Customers can contact Abbott at 1-877-422688 about this action. |  | 2411 | 2,411 kits | Worldwide Distribution: Throughout USA, including Puerto Rico, and the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Finland, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Japan, Latvia, Malaysia, Mexico, Netherlands, Oman, Pakistan, Poland, Portugal, Saudi Arabia, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates, United Kingdom and Vietnam. | Terminated | June 22, 2011 | 282 | 2411
Z-0625-2011 | 56957 | Yes | MVK |  |  | Cardiovascular | PMA | Model 4000 Monitor component of the LifeVest Wearable Defibrillator |  |  | September 24, 2010 | 2010 | Zoll Lifecor Corporation | The response button function may become intermittent. This can result in the device not progressing beyond the response button test during startup, which may prevent a conscious patient from delaying a treatment shock. | Other | Zoll Life Vest contacted patients that are expected to be in the device more than 60 days by letter and phone to arrange replacement monitors. Instructions for the return of the recalled monitor accompany the replacement monitor The firm will also contact the prescribing physician via telephone for long-term use patients to inform the physician of this voluntary recall and to assist in contacting any patients the firm was unable to reach by telephone.Zoll Life Vest can be contacted at 1-800-543-3267 about this recall. |  | 494 | 494 units | Nationwide Distribution: Throughout the US. including the states of AL, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, and WV. | Terminated | July 05, 2012 | 650 | 494
Z-0623-2011 | 56679 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Vacuette 2 ml Lithium Heparin, Collection Tube, Manufactured by: Greiner-Bio One North America, Inc., 4238 Capital Drive, Monroe, NC. |  |  | August 25, 2010 | 2010 | Greiner Bio-One North America, Inc. | Lithium heparin solution was contaminated by EDTA K (potassium) solution due to an operator error and a misinterpretation of a test result. Contaminated tubes will produce inaccurate test results. | Other | Greiner Bio-One instructed customers to stop distribution and to complete an attached Product Disposition form and return it via fax to Greiner Bio-One. They were further instructed to destroy any product they may have. Distributors were instructed provide a list of their customers that have received the affected item via Excel file so that they can be notified of the recall. |  | 18 | 18 cases (total 21,600 pieces) | Nationwide Distribution: USA, including the states of CA, FL, IL, IN, KS, KY, ND, NE, NH, OH, SC, SD and WI. | Terminated | May 15, 2012 | 629 | 18
Z-0622-2011 | 56855 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | ThermaCare Menstrual 8 Hour Heat Wraps, 3 count, NDC 0573-3020-02, SKU: 302002, UPC: 3-0573-3020-02-9, Wyeth Consumer Healthcare, P.O. Box 26609, Richmond, VA 23261-6609 |  |  | September 24, 2010 | 2010 | Pfizer Global | The chemical components contained in the heat wrap may leak directly onto the skin, which could cause skin injury, including burns. | Packaging process control | Wyeth Consumer Healthcare issued an Urgent Voluntary Medical Device Recall letter dated September 24, 2010 to direct accounts, and issued a press release. The firm is removing the product in question from store shelves and asking consumers who have purchased the product to record the lot number, throw the product away in its entirety without opening the foil pouch and call 1 800 232-3383 for replacement. Direct accounts were instructed to remove the affected product from sale and return it immediately. If they further distributed the product they were instructed to notify their customers of the recall and request return of the products to them credit. Instructions were provided in the event consumers inquired about the recall. A prepaid post card was enclosed to record the quantity on hand to be returned. |  | 192000 | 192000 cartons (3 wraps per carton) | Nationwide Distribution: Throughout the U.S. and Puerto Rico. | Terminated | September 16, 2011 | 357 | 192000
Z-0616-2011 | 57290 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Fresenius Red Blood Cell Set, Catalog Number 9007601. |  |  | October 18, 2010 | 2010 | Fresenius Kabi, LLC | Potential for patient hematocrit dropped to levels lower than expected post-treatment. | Device Design | On October 21, 2010, Fresenius Kabi, LLC in Redmond, WA sent a "Field Safety Corrective Action" Letter dated October 18, 2010 to their US Consignees/Customers. On October 22, 2010, the firm followed up with their customers by telephone call.The letter informed the consignees/customers that the Red Blood Cell (RBC) exchange sets for use on the AS104 Blood Cell Separation Device are being recalled.The consignees/customers were advised to examine their stock and determine if they have any affected products on hand. They were instructed to discontinue distributing, using, dispensing the affected products, and return the product to Fresenius Kabi, LLC, 8635 154th Avenue, NE, Redmond, WA 98052.The consignees/customers were also instructed to notify their sub-account customers if the products were further distributed and to complete the Product Recall Response Form and fax it back to 425-242-2101. TIf you have any further questions concerning the recall, please contact the Quality Assurance Manager at 800-909-3872. |  | 255 | 255 units. | Nationwide distribution: including states of: AL, AR, AZ, CA, CT, GA, KY, MN, NC, NJ, OH, TN, and TX. | Terminated | October 05, 2011 | 352 | 255
Z-0606-2011 | 57104 | No | LZP | Intraocular fluid. | Ophthalmic | Ophthalmic | PMA | Beaver-Visitec International (Beaver-Visitec) BD 1% OVD Sodium Hyaluronate, 10 mg/ml 1.0 mL. Sterile. Rx |  |  | October 04, 2010 | 2010 | Becton Dickinson and Company | Finger flange on the BD 1% OVD syringe can disengage during cataract surgery and and may lead to injury. | Component design/selection | Beaver Visitec issued an Urgent: Field Safety Notice letter, dated 10/4/10 to direct accounts, identifying the affected product and informing users of the potential for the finger grip to disengage during use. The firm offers solutions as follows:Allow the finger grip to remain attached to the syringe and continue to use the product. Position the open side of the finger grip towards the palm to avoid detachment of the grip from the syringe. Remove the finger grip and continue to use the product. Return the unused product to Beaver-Visitec for reimbursement.Users were requested to complete the Customer Response Form. Questions can be directed to: Inside Sales at 781-906-7952. |  | 72278 | 72278 units | Nationwide Distribution | Terminated | February 16, 2012 | 500 | 72278
Z-0626-2011 | 57112 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 Inc. 4600 Nathan Lane, Plymouth, MN ProtÂżgÂż " GPS" Self-Expanding Biliary Stent System, Sterile EO, 14 mm 60mm 80 cm, REF SERB65-14-60-80, Rx only. |  |  | October 30, 2010 | 2010 | Ev3, Inc. | Mislabeled size: Specific ProtÂżgÂż GPS lots have product labeling that does not match the actual stent sizes contained in the packages. The affected product outer packaging is labeled as a 12 mm x 60 mm ProtÂżgÂż GPS stent but contains a 14 mm x 60 mm stent. This situation can potentially result in vessel damage. Conversely, a lot was labeled as a 14 mm x 60 mm ProtÂżgÂż GPS stent but actually contai | Packaging process control | ev3 issued a "Medical Device Recall" letter dated September 29, 2010 to consignees. The letter described the product, Issue Summary, &amp; Required Action. Consignees are to review inventory and remove all affected devices. ev3 representatives will contact customers to arrange for the return and credit of unused product.ev3 can be reached at 800 716-6700 concerning this recall. |  | 30 | 30 | Nationwide Distribution: including the states of AZ, DE, C, FL, GA, ID, IL, IN,LA, MN, MO, NJ, NY, NC, OH, VA, and WA. | Terminated | March 29, 2012 | 516 | 30
Z-0711-2011 | 57230 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Genzyme Diagnostics Total Serum Cholesterol Kit, IVD |  |  | November 10, 2010 | 2010 | Diagnostic Chemicals, Ltd. | Cholesterol Reagent bottles labels reversed for R1 and R2 | Labeling mix-ups | Genzyme Diagnostics P.E.I. Inc. sent an "IMPORTANT PRODUCT FIELD CORRECTION" letter dated November 10, 2010, via Federal Express to customers. The letter described the label error and requested that the reagents not be interchanged with another reagent lot. The firm ask that the customers complete and return the attached IMPORTANT PRODUCT FIELD CORRECTION confirmation of notification via fax at 902-628-6504. Contact a Genzyme Diagnostics Technical Services Representative at 1-800-565-0265, if you require further information, or if you have any questions. |  | 30 | 30 kits | Worldwide distribution: USA including states of: KY, PA, NV, and WA and country of Canada. | Terminated | September 19, 2011 | 313 | 30
Z-0712-2011 | 57220 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Comprehensive Reverse Shoulder Humeral Tray with Locking Ring, 49 mm standard, sterile, REF 115350, Biomet Orthopedics, Warsaw, IN |  |  | September 13, 2010 | 2010 | Biomet, Inc. | The firm received complaints regarding fracturing of the device. | Pending | The firm sent "URGENT MEDICAL DEVICE RECALL NOTICE" dated 9/13/2010 to their distributors and customers. The letters stated the reason for the recall and what may happen as a result of the fracture of the device. The notices stated that the recalled devices should be immediately located, discontinued from use, and shipped back to Biomet. Customers should follow the instructions on the "FAX Back Response Form." If the affected product was further distributed, hospital personnel must be notified of the recall via the enclosed "Dear Biomet Customer" notice. The firm sent "URGENT MEDICAL DEVICE RECALL NOTICES' to implanting surgeons. These letters informed the surgeons of the recall and instructed them to continue to monitor the patient for fracture of the device.Customers are to confirm receipt of the notice by calling 800-348-9500 ext 3755 or 3983. Questions should be directed to 574-371-3755 or 574-372-3983. |  | 45 | 45 both products | Nationwide Distribution -- NY, VA, MD, AL, CO, SD, TX, and OK. | Terminated | July 27, 2011 | 317 | 45
Z-0628-2011 | 57281 | [6 empty] | ASP Automatic Endoscope Reprocessor (AER), Product Code: 20300 &amp; 20301 |  |  | February 19, 2010 | 2010 | Advanced Sterilization Products | The recall was initiated because Advanced Sterilization Products (ASP) received three customerreports of elevated temperature of the high-level disinfectant solution used in the ASP AER as a result of apossible temperature monitoring system malfunction. This recall communication advised consignees that if high-leveldisinfectants are exposed to elevated temperatures, fumes may occur and resul | Pending | Advanced Sterilization Products sent an URGENT: PRODUCT CORRECTION letter dated February 19, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. The recall communication was initiated with ASP forwarding a "Urgent: Product Correction" letter with attached customer fax-back response form to the affected customers. Customers were also instructed to complete and sign the enclosed customer fax-back response form indicating that they have disconnected the heater on the ASP AER according to the instructions provided and fax the form to 1-888-345-1435.If customers have any questions or need assistance, they were instructed to call ASP Professional Services at 888-783-7723, option 3 then option 1. |  | 1881 | 1,881 units (1514 in US) | Worldwide Distribution -- United States, Argentina Brazil, Canada, Chile, Colombia, Costa Rica, Ecuador,Hungary, Mexico, Panama, Guatemala, Honduras, Poland, Puerto Rico, Virgin Islands, Singapore, Taiwan,Venezuela, Croatia, and Bulgaria. | Terminated | July 30, 2013 | 1257 | 1881
Z-0714-2011 | 57236 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System, medical charged-particle radiation therapy system, manufactured by Accuray Inc., Sunnyvale, CA. |  |  | October 28, 2010 | 2010 | Accuray Inc | If users modify the dose calculation box to cover just the contoured anatomy and not the entire CT volume, dosage may not be properly calculated. | Other | Accuray Inc., sent an "URGENT DEVICE CORRECTION" letter dated October 21, 2010, to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to increase the size of the dose box to include CT volume, not just the contoured anatomy; recalculate dose at high resolution, and examine the treatment plan to ensure that applicable organs at risk are accounted for and all portions of the patient receive the desired level of doses. If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1-877-668-8667 (USA), +1-408-716-4700 (non-USA) or email: customersupport@accuray.com. |  | 219 | 219 units | USA and countries of: Athens, Belgium, Canada, China, Czech Republic, France, Germany, India, Italy, Japan, Korea, Malaysia, Russia, Saudi Arabia, Spain, Switzerland, Taiwan, Thailand, The Netherland, Turkey, Ukraine, United Kingdom, Vietnam, Venezuela. | Terminated | April 01, 2011 | 155 | 219
Z-0007-2011 | 54154 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | HiSpeed X/i Computed Tomography operator consoles with the Model Number 2153675 and 2247802. |  |  | May 01, 2009 | 2009 | GE Medical Systems, LLC | It was discovered that the operator console did not have a label affixed in accordance with 21CFR 1010.2 indicating certification to applicable performance standards. | Radiation Control for Health and Safety Act | GE Healthcare will bring customer systems into compliance with 21 CFR by having a GE Healthcare field service representative, who will perform a field corrective action, visit each customer site. This will be carried out via a Field Modification Instruction planned to be released in October 2009 and planned to be completed by February 2010. Additionally, all forward production operator consoles are manufactured so as to comply with 21 CFR and to include a manufacturing check for appropriate labeling. |  | 45 | 45 units (13 Nationwide) | Worldwide Distribution -- USA, including states of ME, FL, IA, AL, CT, CA, TX, KS, KY, OH, and WV and countries of Japan, Russian Federation, India, Argentina, Brazil, Italy, Poland, Puerto Rico, Romania, and Saudi Arabia. | Terminated | May 16, 2012 | 1111 | 45
Z-0717-2011 | 57311 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Luminos dRF |  |  | October 25, 2010 | 2010 | Siemens Medical Solutions USA, Inc | patient leg can be pinched | Other | SIEMENS sent a Customer Safety Advisory Notice dated October 25, 2010, to affected customers via Update Instructions AX072/10/S. The notice described the issue and the actions to be taken by customers. The customers were instructed to retain the annex together with their operating instructions and bring the hazard to the attention of their employees; inform all users and other persons of this notice; if they sold the device, forward this notice and attachment to the new owners of the device and provide SIEMENS with details of the new owners.If you have any questions, please contact +1-888-826-9702. |  | 24 | 24 units | Nationwide distribution: CA, CO, FL, GA, IL, IN, KY, MI, MN, MO, NC, NJ, and NY. | Terminated | September 29, 2012 | 705 | 24
Z-0716-2011 | 57242 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | SDS Kerr Revolution Formula 2 Flowable Light Cure Composite. Net Wt. 4 g, Contents 4 Syringes (1g) each. This Package Contains 4 - 12 g Syringes A1, 20 - Dispensing tips, 29493. |  |  | October 15, 2010 | 2010 | Patterson Dental Supply, Inc. | Patterson Dental is conducting a recall because they have been advised by the manufacturer of Revolution Formula 2 (Kerr Corporation) that some syringes and tips marked with the brand name are not manufactured and distributed by Kerr and, thus, not authentic product. | Other | Consignees were sent a Patterson Dental "Urgent Device Recall" letter dated October 15, 2010. The letter was addressed to "Valued Customer". The letter described the product and the problem. It also advised the consignees to quarantine and return the product with the specific distribution dates. Consignees were to complete and return the enclosed Recall Return Form along with the affected product. Questions should be directed to Nancy Strand at 651-686-1966 or nancy.strand@pattersondental.com. |  | 3204 | 3,204 packages | Nationwide Distribution -- Including DC except VT. | Terminated | September 11, 2012 | 697 | 3204
Z-0715-2011 | 57277 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Polaris Spectra System Control Unit, a component of the StealthStation S7 System. Medtronic Navigation, Louisville, CO 80027. |  |  | November 05, 2010 | 2010 | Medtronic Navigation, Inc | Electronic surgical navigation aid may lose tracking ability or stop working. | Component design/selection | Customers were notified by a Product Correction Notification letter on 11/05/2010. The letter advised customers of the affected product, the potential problem, and the description of the correction. Customers were told that a Medtronic rep will visit them to install a replacement control unit. Further information is available at 800-595-9709 or 720-890-3200. |  | 48 | 48 devices | Worldwide Distribution -- US, including states of AL, AZ, CO, FL, GA, KY, MD, MN, NE, PA, TX, WI, and MD and countries of Australia, Belgium, Canada, Czech Republic, Italy, Netherlands, Poland, Portugal, Saudi Arabia, Singapore, South Korea, Spain, Turkey, and UAE. | Terminated | May 03, 2011 | 179 | 48
Z-0718-2011 | 57312 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobilett XP Hybrid/Digital/CR |  |  | October 04, 2010 | 2010 | Siemens Medical Solutions USA, Inc | battery fuse | Other | SIEMENS released a field update dated October 25, 2010, via SP011/10/S to the representatives. The update described issue and corrective action plan. The representatives are to visit each account to upgrade the systems with a new battery fuse and a new holder.If you have any questions regarding this report, please contact Quality Engineer by phone 610-448-3237 or email: meredith.adams@siemens.com and Director, Product complaints &amp; Field Support by phone 610-448-4634 or email: josephine.mcbride@siemens.com. |  | 170 | 170 units | Nationwide distribution: AZ, CA, DC, MD, PA, SC and VA. | Terminated | September 29, 2012 | 726 | 170
Z-0794-2011 | 57351 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Accessory Cart for The ACHIEVA 3.0T and Panorama 1T Magnetic Resonance Imaging unit. |  |  | May 12, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | The Accessory Cart contained magnetic material and is attracted to the magnet when bringing it close to the system. | Nonconforming Material/Component | On 05/12/2009, Philips began sending out the URGENT-Field Safety Notice MR Accessory Cart to their consignees. The letter identified the affected product and stated that the drawer rails were made from magnetic material and were not according specification. The hazards involved were also provided along with how to identify the affected product. The firm issued the FCO 78100296 and the firm's Field Service Engineer replaced the magnetic rails with the non-magnetic rails at the consignee's site.The consignees can contact contact the Philips Healthcare Call Center at 800-722-9377 #5 (then follow prompts) and reference FCO 78100296 with any communication. |  | 3 | 3 units | Nationwide Distribution -- IL, MI, and MD. | Terminated | December 28, 2010 | 595 | 3
Z-0797-2011 | 57219 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vital, Vitrea fX and Vitrea Enterprise Suite fX |  |  | October 14, 2010 | 2010 | Vital Images, Inc. | Vital Images Inc., reported that a software defect exists with Vitrea fX 2.1, 3.0 and 3.1 CT Brain Analysis whereby the annotated numeric ROI values on a batched perfusion map are incorrect when created using the auto batching function (from the right click menu) with the auto ROI templates enabled. | Software design | Vital Images, Inc sent a Vital "Urgent Software Update Alert" letter dated 14 October 2010 to all Consignees/Customers. The letter described the product and problem. They advised, until new software is installed, that the customers can avoid obtaining incorrect values by tuning off ROI templates prior to selecting the batch option from the right click menu in the CT Brain Analysis application; and recommended that the customers re-examine any ROI perfusion values that were previously generated using automatic collage batching with ROI templates.Please direct any questions or comments to Customer Support at support@vitalimages.com or at 1-800-208-3005. Any help needed to install this software update can also be directed to Customer Support. |  | 182 | 182 | Worldwide distribution: USA including states of: AL, AZ, AR, CA, CO, FL, GA, IA, KY, LA, MD, MA, MI, MN, MO, MT, NY, NJ, ND, OH, PA, PR, TN, TX, WA, WI, and WY, and countries of: CANADA, NETHERLANDS, TAIWAN, CHINA, SINGAPORE, UNITED KINGDOM, BRAZIL, GERMANY, RUSSIA, THAILAND, FRANCE, REPUBLIC OF KOREA, SPAIN, PAKISTAN,IRELAND, JAPAN, AUSTRALIA, INDIA, CZECH REPUBLIC, HONG KONG, EGYPT, ITALY, AUSTRIA, SLOVAK IA, BULGARIA, LEBANON, ALGERA, TAIWAN, SOUTH ARICA, BELGIUM, MALAYSIA, DENMARK, and JORDAN.CANADA, NETHERLANDS, TAIWAN, CHINA, SINGAPORE, UNITED KINGDOM, BRAZIL, GERMANY, RUSSIA, THAILAND, FRANCE, REPUBLIC OF KOREA, SPAIN, PAKISTAN,IRELAND, JAPAN, AUSTRALIA, INDIA, CZECH REPUBLIC, HONG KONG, EGYPT, ITALY, AUSTRIA, SLOVAK IA, BULGARIA, LEBANON, ALGERA, TAIWAN, SOUTH ARICA, BELGIUM, MALAYSIA, DENMARK, JORDAN. | Terminated | May 15, 2012 | 579 | 182
Z-0795-2011 | 57349 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Advanced Bionics HiRes 90K Cochlear Implant device (unimplanted units), Model Numbers: CI-1400-01 |  |  | November 23, 2010 | 2010 | Advanced Bionics Corporation | The recall action is being taken in response to two confirmed instances where the product experienced a malfunction requiring explantation. These recipients experienced severe pain, overly loud sounds and/or shocking sensations at 8-10 days after initial activation of their device. | Pending | Advanced Bionics (AB) Personal contacted all centers with known surgeries about the recall by telephone on November 22, 2010, and sent Blastfax to affected surgeons and clinicians.AB forwarding a Press Release on November 23, 2010 to the Business Wire News Trak. The firm also faxed Health Care Provider letters to the affected Health Care Providers and placed a Patient notification letter on the AB website on November 23, 2010. Health Care Provider letters were FedEx'ed November 24, 2010 to the affected Health Care Providers. Patient Letters were sent via FedEx'ed beginning Monday November 29th and 30th to approx. 10,500 patients. A Recall Acknowledgement Form was included for each recipient to sign as proof that the notification has been read.The press release &amp; recall letter informed the patients/consignees of the reason for recall, the product affected and the potential hazard involved. The Health Care Provider letter also requested providers with recipient under their care who reports pain, overly loud sounds, or sudden shock sensations while the implant is powered to follow these steps:1. Advise them to remove the external equipment immediately,2. Schedule an appointment to complete standard troubleshooting steps listed in the AB Sound Processor User Guide,3. If standard troubleshooting does not resolve the issue, please contact Advanced Bionics for additional troubleshooting steps that may assist in determining the presence of recall problem, 4. If your evaluation of the patient's condition suggests that the implant exhibits this problem, device explantation and replacement is advised. In addition, the Health Care Provider letter informed providers that if AB records show that you have any unregistered implants in their possession, a list of serial numbers has been provided for devices that must be returned to Advanced Bionics. If no unregistered implants have been identified, no list will be enclosed. The Patient letter also |  | 332 | 332 units | Nationwide distribution: AL, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT and VA. | Terminated | March 23, 2012 | 486 | 414
Z-0719-2011 | 57291 | No | EXS | Urine flow or volume measuring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BardÂż CritiCoreÂż Monitor, Software Version 2.1, Catalog number 000002N, C.R. Bard, Inc., Covington, GA 30014. |  |  | November 17, 2010 | 2010 | C.R. Bard, Inc., Urological Division | Error can occur in the reading of the Present Interval Output (PIO) reading when the device is subjected to bumping or forceful replacement of the fluid container. | Other | Consignees were notified by an Urgent: Product Advisory Letter and Medica Device Recall letter on/about November 18, 2010. The letter identified the affected product, stated the reason for recall, the interim precautions, and actions to be taken. The letter requested acknowledgement of the recall and to return the entire affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. Affected product received by Bard Medical will be reworked with the software upgrade and place back into distribution. Consignees will either receive a software upgrade on site by a third party or replacement product with the software upgrade. For international distribution, the hospital will be informed of this issue through Bard affiliate and Notified Body Bard Europe and managed according to the requirement. If there are questions or assistance is needed, customers can contact BD at 1-770-784-6969. |  | 234 | 234 units | Worldwide Distribution -- U.S., Netherlands, and Japan. | Terminated | February 02, 2012 | 442 | 234
Z-0507-2011 | 57376 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Aimshot HeatSeeker Model HS3500, contains Class IIIb infared laser illuminator. |  |  | August 17, 2010 | 2010 | Amstech Incorporated | Product contains infra red laser, which can pose a potential hazard to users who unknowingly shine the beam into people's eyes. | Radiation Control for Health and Safety Act | Firm sent a letter dated August 5, 2010 to customers stating that a recall was being conducted on the AimSHOT HeatSeeker and specified model. The letter stated the reason for the recall and that the firm has redesigned the HeatSeeker and removed the infra red laser feature. The letter also asks customers to check their inventory and return the affected products to the firm. The firm states that they will send customers their new model and pay for the shipping. Questions should be directed towards 770-840-8039. |  |  | All units | Nationwide Distribution | Terminated | November 09, 2011 | 449
Z-0802-2011 | 57209 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | klinoPORT, AC powered Operating Table |  |  | October 26, 2010 | 2010 | Kreuzer Gmbh &amp; Co. KG | The motorized adjustable pendant arm may suddenly descend to its lowest position. | Other | Consignees were notified by a Warning Notice on/about 10/26/2010. Medical facilities were instructed to adjust the pendant arm to the lowest height position and disable the motor and to place the accompanying warning label on the pendant. They will, therefore, not be able to adjust the height of the pendant until corrective action is developed and completed. TRUMPF Medical Systems, Inc. will receive corrective action repair kits by 11/12/2010. |  | 15 | 15 units | Nationwide Distribution -- CT, MA, and MO. | Terminated | March 20, 2012 | 511 | 15
Z-0800-2011 | 57210 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Turon Shoulder System Instrumentation, Turon Trial Pegged Glenoid, DJO Surgical |  |  | November 05, 2010 | 2010 | Encore Medical, Lp | Complaints of fracturing were reported for affected trial products. | Device Design | Encore Medical notified consignees by letter on 11/5/2010. |  | 85 | 85 units | TX, FL, SC, GA | Terminated | April 05, 2011 | 151 | 578
Z-0790-2011 | 56965 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Arrow Ultra 8 IAB (Intra-Aortic Balloon) Catheter, 8 Fr, 30cc. |  |  | October 11, 2010 | 2010 | Arrow International, Inc., Division of Teleflex Medical Inc. | SuperArrow Flex IAB catheter becomes stuck in the sheath, unable to move the IAB catheter forward or backward, causing a delay in therapy, bleeding or arterial injury. | Device Design | Arrow international initiated a voluntary recall on 12/21/10 by letter requesting users to discontinue use and return all unused 5800 Series IAB with SuperFlex Introducers to Arrow. This field action supercedes the Safety Alert IAB 'Stuck in Sheath' letter issued on 10/8/2010. Questions should be directed towards their local sales representative or the IABP support Line at 1-866-396-2111 or 1-617-389-6400. |  | 5289 | 5289 units | Worldwide Distribution -- USA, Canada, Australia, Belgium, Czechoslovakia, Germany, Russia, Spain, Finland, France, Great Britain, Hungary, Italy, Netherlands, Poland, Saudi Arabia, Switzerland, Brazil, Australia, Chile, Colombia, Ecuador, India, Indonesia, Japan, Korea, Malaysia, Mexico, New Zealand, Panama, Peru, Philippines, Taiwan, Thailand, and Venezuela. | Terminated | February 05, 2015 | 1578 | 69523
Z-0789-2011 | 57335 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3 NanoCross 0.14 OTW PTA Dilatation Catheter, sterilized using ethylene oxide gas. REF #'s: AB14W040040150, AB14W030040150, AB14W020040150, AB14W025080150, AB14W020080150, AB14W020120150, AB14W030080150, AB14W035120150, |  |  | November 12, 2010 | 2010 | Ev3, Inc. | ev3 has determined that specific lots of the NanoCross .014 OTW PTA Dilatation catheters may crack or break during use. Cracking or breaking of the catheter shaft can result in the inability to inflate or deflate the balloon, and can result in material separation and potential embolization. | Process control | A Medical Device Recall letter addressed to Risk Managers, dated 11/10/2010, was sent to hospitals beginning 11/12/2010. The letter identified the affected devices, described the issue and requested that affected devices be located and removed from use. ev3 sales representatives will contact customers to arrange return of all unused devices. A credit or replacement will be provided for any unused devices. Questions regarding this issue may be directed towards Customer Service at 800-716-6700. |  | 4962 | 4,962 (3,858 US; 1,104 OUS) | Worldwide Distribution -- USA, AUSTRALIA, BELGIUM, BRAZIL, BULGARIA, CANADA, CHILE, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, KUWAIT, LEBANON, MALTA, NETHERLANDS, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SYRIA, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM, | Terminated | May 17, 2011 | 186 | 4962
Z-0818-2011 | 56894 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Depuy Spine Mountaineer OCT Spinal System Screw Caddy Set, , Non Sterile |  |  | October 02, 2010 | 2010 | DePuy Spine, Inc. | The scale which serves as a secondary means of identifying screw length is incorrectly etched on the screw caddy | Labeling design | Depuy Spine issued an "URGENT-VOLUNTARY PRODUCT RECALL NOTIFICATION" via e-mail on October 2, 2010 to customers. The notification describes the product, problem and action to be taken by customers. The customers were instructed to please distribute this information to appropriate personnel at their facility and return any caddies that they may have to their local distributor for DePuy Spine products. Depuy will return devices to customer promptly (within two business days), after reworking the devices by removing the scale altogether. Once the product is fully corrected, Depuy will provide the customers with a new and updated unit with the corrected scale markings. If you should have any questions or concerns regarding this action, please contact DePuy Spine Customers Service at 800-227-6633. |  | 121 | 121 units | Nationwide distribution: USA including states of: AR, AZ, CA, FL, HI, IA, IL, KY, LA, MA, MI, MN, MO, NC, NY OH, PA, TN, TX, VA, and WA. | Terminated | April 15, 2011 | 195 | 441
Z-0820-2011 | 55976 | No | FAF | Testicular prosthesis. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | Coloplast, Saline Filled Testicular Prosthesis, Sterile R, Manufacturer Coloplast A/S 3050 Denmark. |  |  | April 30, 2010 | 2010 | Coloplast Manufacturing US, LLC | On April 22, 2010, a decision was made by Coloplast to conduct a field correction based on a report that the Saline Filled Testicular Prosthesis is considered mislabeled. This is solely due to a difference in expiration dates between the Terumo Butterfly Needle included in the kit and the testicular implant. The needle has a 3yr expiration date &amp; the Saline Testicularhas a 5yr expiration date. | Labeling False and Misleading | The recall will extend to field personnel who have control of all SFTP Systems. The product is ordered by the customer (physician and/or hospital) directly from a field sales representative who ships the device to them or drops it off in person.Coloplast sent all Field Personnel an email on April 30, 2010, with a copy of a letter dated April 30, 2010 and a follow-up letter dated May 1, 2010. The letters described the problem and the product involved in the recall. The letters included a list of Corrections Required. This included the field personnel receiving a stock of sterile Butterfly Needles based on their home consignment inventory to be provided for use with all Saline Testicular Implants; Copies of customer letters to be provided with each replacement needle; Changes in all Saline Testicular orders, and Field personnel to administer the details of the correction and make the direct customer contacts. For questions, contact the Customer Service/Field Inventory Representative at 612-337-7859 or email: usbam@coloplast.com |  | 1018 | 1018 | Worldwide distribution: USA including states of AL, AK, AZ, AR, CO, CT, CA, FL, GA, HI, IL, IN, IA, LA, MD, MA, MI,MN, MS, MO, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT and WA; and countries of Australia, Canada, Hong Kong and Latin America (Chile, Columbia, Ecuador and Venezuela). | Terminated | December 17, 2011 | 596 | 1018
Z-0816-2011 | 56972 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD Vacutainer Push Button collection Set; |  |  | October 05, 2010 | 2010 | Becton Dickinson &amp; Company | There is the possibility for certain lots of the BD Vacutainer Push Button Blood Collection sets to have separation of the cannula from the hub. | Process control | BD Diagnostics sent "RECALL NOTIFICATION" separate letters October 1 and 5, 2010 and follow up letter dated October 8, 2010, as well as a letter dated December 1, 2010 to all distributors and customers commencing October 5, 2010 by federal express. The letters describe the product, the problem and the action to be taken by the customers. The customers were instructed to immediately examine their inventory for specific catalog and lot numbers, and return product from the lots following the instructions on the enclosed packing slip (BD will send replacement upon receipt of returned product). If you have any questions on returning and replacing product, please contact BD logistics at 1-866-380-0046. If you have any further detail questions regarding this recall, please contact your BD Sales Consultant or BD Technical Services via email at www.BD.com/Vacutainer/Contact or via telephone at 1-800-631-0174. |  | 4223200 | 4,223,200 | Nationwide Distribution | Terminated | June 06, 2012 | 610 | 4223200
Z-0808-2011 | 56846 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Model # 720088-01: AMS AdVance Male Sling System, US; Model # 720088-02: AMS AdVance Male Sling System, INTL; Model # 72404224: Male TO Sling System, INTL. |  |  | September 07, 2010 | 2010 | American Medical Systems, Inc. | American Medical Systems (AMS) is recalling all lots of the AdVance male Sling System with an expiration date between July 2012 and August 2013. Internal testing has indicated the potential for degradation of the bioabsorbable tensioning sutures when the product is stored for longer than 15 months after the date of manufacture. | Package design/selection | Consignees were sent an AMS "Urgent Product Recall" letter dated September 7, 2010. The letter described the problem and the product involved. They requested consignees to return any existing stock product and to return the Product Recall Acknowledgement form enclosed. Customers may request a credit or a full refund or also, arrange for replenishment of AdVance product. Questions regarding the recall should be directed towards AMS Customer Service at 1-800-328-3881. |  | 7069 | 7,069 | Worldwide Distribution -- All states in the Continental USA including DC and Puerto Rico and countries of ARGENTINA, AUSTRIA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, SWITZERLAND, CHILE, CHINA, COLOMBIA, COSTA RICA, GERMANY, ESTONIA, SPAIN, FINLAND, FRANCE, UNITED KINGDOM, GREECE, HONG KONG, CROATIA, IRELAND, ISRAEL, INDIA, ITALY, REPUBLIC OF KOREA, LEBANON, MOROCCO, MEXICO, MALAYSIA, NETHERLANDS, NORWAY, NEW ZEALAND, PANAMA, POLAND, PORTUGAL, RUSSIAN FEDERATION SAUDI ARABIA, SWEDEN SINGAPORE, SLOVENIA, THAILAND, TURKEY, TAIWAN, VENEZUELA, and SOUTH AFRICA. | Terminated | June 28, 2012 | 660 | 7069
Z-0813-2011 | 57348 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Blease Sirius Anaesthesia System is a anaesthesia Gas Machine. The system has the capability to mount two anaesthetic vapourisers on the back bar via a demountable system. The delivery of anaesthetic agent is achieved by the rotation of a control knob which actuates a feature designed "shot bold" mechanism. The purpose of which is to prevent two vapourisers being turned on simultaneously.. |  |  | June 25, 2009 | 2009 | Spacelabs Healthcare, Llc | Users may experience difficulty turning the control knob to an active vapour delivery setting, or may not be able to turn on the Left Hand mounted vapouriser on the Blease Sirius Anaesthesia Systems. | Other | Spacelabs Healthcare issued an URGENT MEDICAL DEVICE CORRECTION letter dated June 25, 2009 to the two consignees in the US. Spacelabs Healthcare UK notified consignees outside the US.The consignees were advised to brief their staff about the issue and test the Anesthesia Systems in advance of a procedure. If there is an issue with turning on a vaporizer in the left hand position, simply move the vaporizer to the right hand position.The firm contacted their consignees for a schedule to replace the flowmeter mouldings. The US consignees can contact Spacelabs Healthcare Representatives at 425-657-7200 ext. 5089 about this recall. Foreign consignees can contact Spoacelabs at 44 1494 784422. |  | 41 | 5 units to US consignees and 36 units Internationally. | Worldwide Distribution: US in the states of LA and SD, and internationally by Spacelabs Healthcare, Ltd in the countries of China, Czech Republic, Hong Kong, India, Poland, Saudi Arabia, Singapore, and the United Kingdom. | Terminated | December 28, 2010 | 551 | 41
Z-0804-2011 | 57274 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | ACL 9000 Coagulation Analyzer, PN 00000990011 |  |  | November 11, 2010 | 2010 | Instrumentation Laboratory Co. | Sporadically lowered APTT results were observed with the PT/APTT Test Groups. | Device Design | On 11/11/10, An Urgent Product Notification informing customers to discontinue use of all IL locked PT/APTT-based Test Groups was mailed to U.S. and Canadian customers, instructing them to run in Single Test or Profile Mode until further notice. The letter identified the affected product, the issue, risk and stated a list of mandatory actions. Customers are to immediately acknowledge the receipt of the notification by signing and returning it via email or fax. A Worldwide Regulatory Notice was sent on November 8, 2010 to international affiliates (non-U.S. or Canada), notifying them of the issue, providing a risk summary and the support documentation for the notification. Secondary international dealers were notified Nov 8 and a fax response of notification of the dealers will be tracked by International Laboratories. The customer notification of the secondary dealers will be handled per the local regulatory requirements of the countries where they are located. |  | 87 | 87 distributed in USA &amp; Canada | Worldwide Distribution -- USA, Canada, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Sri Lanka, Sweden, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Turkmenistan, United Arabian Emirates, Uruguay, Venezuela, Vietnam, Austria, Belgium, China, France, Germany, Holland, Italy, Japan, Spain (and Portugal), and UK. | Terminated | August 23, 2013 | 1016 | 87
Z-0815-2011 | 57292 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CuttingEdge Advantage Rasp Handle, Straight Rasp Handle, &amp; Specialty Head Center Accolade Broach Handle; Stryker Howmedica Osteonics, Non-sterile Instruments; Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; |  |  | November 01, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The CuttingEdge Advantage Rasp Handle, Straight Rasp Handle, and Specialty Head Center Accolade Broach Handle may fracture in two pieces along the width of the body under the striking plate. This may extend surgical time under anesthesia, cause trauma to the surgeon or patient, or contaminate the surgical site. | Device Design | Stryker issued a Product Correction Bulletin dated November 2010 to their consignees, alerting surgeons of the need to inspect the instruments before and after each surgical use, and, if a defect is detected, to contact Stryker for return of the product. Stryker can be contacted at 201 831-5288 concerning this issue.In response to their revised strategy, Stryker issued notification letters and product recall acknowledgement forms to their branches and agencies on 2/17/2011 for 2/18/2011 delivery via Fed Ex and to their hospital risk managers on 2/18/2011 for 2/21/2011 delivery via Fed ex. |  | 459 | 459 units | Worldwide Distribution: Throughout the US, the countries of Canada and Japan, and the regions of Asia Pacific, Europe, Middle East, Africa, and Latin America. | Terminated | March 26, 2014 | 1241 | 459
Z-0814-2011 | 57337 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ThomoTherapy Hi-Art System |  |  | November 08, 2010 | 2010 | TomoTherapy Incorporated | TomoTherapy Inc. is sending this Field Safety Notice to make you aware of an anomaly which may affect the performance of the Hi-Art System and make it so the user cannot generate a Completion Procedure. | Software design | Consignees were sent a TomoTherapy "Urgent Field Safety Notice" dated November 8, 2010. The letter addressed to TomoTherapy Customers described the Issue, Product Affected, Recommended Action and Resolution. Customers are to contact the TomoTherapy Customer Interaction Center for further instruction and assistance in generating a completion procedure, if they experience the issue. Questions should be directed towards the TomoTherapy Customer Interaction Center via email or telephone. |  | 40 | 40 (21US; 19 OUS) | Worldwide Distribution -- USA, including the states of AZ, CO, IL, KS, MI, MO, NV, NY, ND, OH, OK PA, UT, and VA and countries of MALAYSIA, SWITZERLAND, KOREA, ITALY, SPAIN, CANDA, TAIWAN, JAPAN, BELGIUM | Terminated | February 04, 2012 | 453 | 40
Z-0811-2011 | 57317 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Expect Endoscopic Ultrasound Aspiration Needle, 22GA (0.72 mm) sterile, UPN M00550011, 5 pack, 428 Technology Drive East, Menomonie, WI. |  |  | October 20, 2010 | 2010 | Boston Scientific Corporation | The firm has received a high number of complaints related to difficulty advancing and/or retracting the needle. | Other | "Urgent Medical Device Recall - Immediate Action Required" notification letters were sent via Federal Express Priority mail on October 20, 2010. The letters identified the affected product and stated the reason for the recall. The notification instructs the account to immediately discontinue use of, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. |  | 567 | 567 all products. | Worldwide Distribution -- US and Canada. | Terminated | August 03, 2011 | 287 | 567
Z-0831-2011 | 51260 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Manual Wheelchairs with Transit System Option, |  |  | February 02, 2009 | 2009 | Sunrise Medical Inc | The stabilizer bar was inadvertently omitted during the manufacturing process. | Process change control | An Urgent: Medical Device Correction letter, dated February 2009, with certified return receipt was sent to customers in the United States; Canadian consignees were notified via e-mail. The letter identified the affected product and the reason for the recall. The letters asked customers to contact their Quickie dealer to make arrangements to have their device wheelchair corrected at no charge and to also verify their wheelchair serial number on the form. Questions should be directed towards their dealer or Sunrise Medical at 800-456-8168. |  | 23 | 23 | Worldwide Distribution -- United States and Canada. | Terminated | January 03, 2011 | 700 | 23
Z-0830-2011 | 53183 | No | FAT | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Stryker Infravision Esophageal Kit, Model Numbers 220180548, 220180540, 220180560, manufactured by Stryker Endoscopy San Jose. |  |  | August 20, 2009 | 2009 | Stryker Endoscopy | The silicone sheath component of the product detaches from the catheter during removal of the e-kit from the body. | Device Design | Stryker issued Urgent Product Recall letters dated August 20, 2009 to all its direct consignees, informing them of the affected products and providing instructions on the recall. Customers were requested to return the affected product to Styrker. |  | 2985 | 2985 units | Worldwide Distribution: USA in the states of: Alabama, Arkansas, California, Colorado, Connecticut, Delaware, Florida, Georgia, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, New Hampshire, Nebraska, New Jersey, New Mexico, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Virginia, Vermont, Washington, Wisconsin, West Virginia, and Wyoming, and the countries/region of Australia, Canada, Scandinavia, Italy, and Switzerland. | Terminated | January 03, 2011 | 501 | 2985
Z-0825-2011 | 54147 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Stryker brand End Cutter Shaver Blade F-Series, 4.0 mm; Model Number -375-747-000, manufactured by Stryker Puerto Rico |  |  | November 12, 2009 | 2009 | Stryker Endoscopy | Potential breakage of the internal shaft of the End Cutter Shaver Blade; F-Series, 4.0mm. if breakage occurs, the device becomes non-functional and the defect could be observed through the camera/console system. | Device Design | Stryker issued Urgent Product Recall letters dated November 12, 2009 to all its direct consignees, informing them of the affected products and providing instructions on the recall. Customers were requested to return the product to Stryker.Stryker can be contacted concerning this recall at 408 754-2124. |  | 295 | 295 units | Nationwide Distribution includes: Arizona, California, Colorado, Denver, Florida, Hawaii, Illinois, Kansas, Kentucky, Massachusetts, Michigan, New York, Ohio, and South Carolina. | Terminated | January 03, 2011 | 417 | 295
Z-0823-2011 | 56245 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | Steris Reliance EPS Endoscope Processing System , manufactured by Corporation STERIS Canada, Quebec, Canada. |  |  | July 02, 2010 | 2010 | Steris Corporation | Reliance EPS had internal water leaks and interrupted cycles when air leaks triggered false Fault 38 alarm that DG/HLD cup has not been replaced. | Other | Steris issued an "Urgent Field Correction Notice" dated July 2, 2010 to each affected US consignee, notifying them of the issue and affected device. Steris Field Service Representatives and Dealers will contact each affected customer to schedule a service visit to install system upgrades. In the event of an interrupted cycle prior to the upgrade, consignees were instructed that the load should be handled in accordance with the Reliance EPS Operators Manual. Steris can be contacted at 440-392-7043. |  | 936 | 936 | Worldwide distribution: USA, including states of AK, AL, AZ, CA, CO, DE, FL, ID, IL, KS, KY, MD, MI, MN, MO, NC, NV, NY, OH, PA, SC, VA, WA, WI, and WY, and countries of Algeria, Australia, Belarus, Canada, France, Germany, Greece, India, Italy, Japan, Kuwait, Lebanon, New Zealand, Russian Federation, Saudia Arabia, Spain, Turkey, United Arab Emirates and United Kingdom. | Terminated | September 27, 2012 | 818 | 936
Z-0821-2011 | 56774 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Trilogy System with RapidArc, Model Numbers: |  |  | July 20, 2010 | 2010 | Varian Medical Systems Inc. | The system may deliver an inaccurate dose that does not match the treatment plan. As a result of a system configuration problem, the VMAT (Volumetric Modulated Arc Therapy) plans may be delivered as Dynamic Arc therapy plans. If this occurs, the VMAT plan will not be recognized. This issue affects Trilogy machines with the Rapid Arc Treatment delivery option. | Software design | Varian issued Urgent Medical Device Correction letters dated July 20, 2010 to its direct consignees, informing them of the affected product and providing instructions on the actions to be taken. These instruction include: if the unit does not display "RapidArc" or "VMAT" under Treatment Type during dose delivery, discontinue use of the device and call Varian service.Varian Service Help Desk can be contacted at 1-888-827-4265. |  |  |  | Worldwide Distributiuon: Throughout the US, and the countries of Australia, Canada, China, Japan, and New Zealand, and in the regions of North and Latin America, Europe, and South East Asia. | Terminated | July 18, 2011 | 363
Z-0822-2011 | 56170 | [6 empty] | BD Beaver Slit Blade 3.0mm Straight |  |  | June 24, 2010 | 2010 | Bd Medical Systems | Blades may contain defects at the blade tip or cutting edge. | Process control | BD Medical notified customers via letter dated 6/24/10 to discontinue use of product and remove from inventory. Customers were instructed to complete the customer response form and fax to BD Medical. BD Medical was to contact customers upon receipt of the customer response form to arrange for replacement.Unused product was to be returned to:BD MedicalOphthalmic SystemsAttn: WT-12582411 Waverly Oaks RoadBuilding 2, Suite 229Watham, Massachusetts 02452-8402For questions regarding this recall call 1-781-906-7952. |  | 524 | 524 pieces | Nationwide Distribution | Terminated | June 29, 2012 | 736 | 524
Z-0829-2011 | 57273 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Vanguard Knee Instruments offset reamer bushing - angled, REF 32-488519, Biomet Bridgend, UK. |  |  | September 10, 2010 | 2010 | Biomet, Inc. | The firm has initiated the action following a report that during surgery the Offset Reamer Brushign would not mate with the prepartion mask. After a 5 minute delay in surgery the procedure was completed using another attachment that was on hand. There was no injury to the patient. | Other | The firm, BIOMET, sent a "FIELD SAFETY NOTICES" dated September 10, 2010, addressed to the head of orthopaedic departments, trauma departments, operating departments, sterile service departments, procurement, supplies and risk management. The notice described the product, problem and action to be taken by the customers. The letters instructed the customers to locate and immediately quarantine the recalled device; to complete and return the attached "FAX BACK RESPONSE FORM" via fax to +44 (0) 1656 645454, and the recalled device for credit via mail to Biomet UK Ltd, Waterton Industrial Estate, Bridgend, CF31 3XA. The customers were also instructed to inform all in their organization who are using or ordering the products and send a copy to the organization to which potentially affected products may have been transferred.If you have any questions regarding this communication, please contact the Quality Manager at +44 (0) 1656 655221 or email: david.broderick@biomet.com. |  | 16 | 16 units | International distribution only: Netherlands, Germany, France, Spain, Italy and Denmark | Terminated | January 19, 2011 | 131 | 16
Z-0828-2011 | 57286 | No | HWB | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CLP Instrumentation Removal Hook, distributed by DJO Surgical, manufactured by Encore Medical, Austin, TX. |  |  | November 10, 2010 | 2010 | Encore Medical, Lp | Revision A and B CLP Extraction Hooks may fracture at the distal end, resulting in surgical delay. | Other | DJO issued an Urgent Recall Notice letter dated November 10, 2010 to consignees identifying the affected devices and requesting return of all affected devices for replacement. Customers are also requested to complete and return an enclosed Recall Response Form.Customers can contact David Harris of DJO at 512 834-6330 with questions about this recall. |  | 27 | 27 Units | Nationwide Distribution to the states of AZ, CA, FL, GA, ID, KS, MO, NY, TX, and WI. | Terminated | April 25, 2011 | 166 | 27
Z-0827-2011 | 57289 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC 2000 ultrasound system. |  |  | September 29, 2010 | 2010 | Siemens Medical Solutions USA, Inc. | There is the possibility of a short occurring in the transducer which could cause focal heating on the probe face. | Other | SIEMENS issued a filed correction in the form of a "Customer Safety Advisory Notification" initiated on September 29, 2010 and mailed to customers by Federal Express on the same day. The notification described the product, problem and actions to be taken by the customers. The customers were instructed to examine the 4Z1c transducer prior to use (Note: if transducer does not show a blister on its front face or does not feel unusually when it is safe for use); to immediately discontinue use and contact their local Siemens representative to get issue resolved, if the transducers developed a hot spot that exhibits a visible blister on the patient contact surface and feels hot to the touch during operation, and to pass this notice on to all those within their organization who need to be aware.If you have any questions, please contact (650) 969-9112 or (800) 422-8766. |  | 196 | 196 total units, transducers and SC2000 | Worldwide Distribution: USA and countries including: Australia, Belgium, Brazil, China, Czech Republic, Finland, France, Germany, India, Ireland, Italy, Japan, Kenya, Malaysia, Netherlands, Norway, Portugal, South Korea, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom. | Terminated | October 04, 2011 | 370 | 2392
Z-0824-2011 | 57341 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural Nail System Cephalomedullary Lag Screw Inserter Long, Catalog 00-2490-003-50, Zimmer, Warsaw, IN. |  |  | October 29, 2010 | 2010 | Zimmer Inc. | The firm has determined that two lots of this instrument were inappropriately manufactured by the supplier. The manufacturing process created a step in the lumen which impedes the Lag Screw Retaining Shaft from passing through the lumen without the use of excessive force. | Nonconforming Material/Component | The firm sent Correction and Removal Notification Immediate Response Required letters via email to consignees. The letter identified the affected product and the reason for the recall. The letters also discussed the clinical implications and required actions. Customers are to complete the inventory certification form and user facility documentation form. Each of the affected accounts will be provided with a new instrument, manufactured with the correct lumen diameter and will secure and return the affected instruments per the attached notification, along with the completed inventory certification form. If any further information is needed, customers are to contact Aimee Wood at 574-372-4463. |  | 34 | 34 | Worldwide Distribution -- USA, including states of MI, NJ, MS, OR, KS, TN, A, UT, and CA and countries of Belgium and Germany, Spain, Italy, Greece, France, UK, and Switzerland. | Terminated | September 16, 2011 | 322 | 34
Z-0836-2011 | 56024 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems FastPlan software, Versions 5.5 and 5.5.1, Model Number HZ3, manufactured by Varian Medical Systems, Palo Alto, CA |  |  | March 23, 2010 | 2010 | Varian Medical Systems Oncology Systems | Due to software Issue, an error occurs when performing cone planning using MRI, which may result in the patient receiving a higher radiation dose than expected, potentially to healthy tissue. | Device Design | Varian sent Urgent Medical Device Correction letters on March 23, 2010, by Federal Express to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to run test cases or otherwise determine whether the MR images that they receive have the 'Series Date' parameter populated.Customers should contact their MR system manufacturer to get a confirmation that the 'Series Date' field is populated.Customers were instructed to confirm by visual inspection that the fusion results are loaded appropriately in Cone Planning before starting to create or edit a plan.For questions regarding this recall contact your local Varian Customer Support Distirct or Regional Manager.Phone: USA and Canada - 1-888-VARIANS (888-827-4265) Europe -+41 41 749 8844E-mail: North America: support-america@varian.com Australia/New Zealand: support-anz@varian.com Europe: support-emea@varian.com South East Asia: seasia.apps.helpdesk@varian.com China/Asia: china.apps.helpdesk@varian.com Japan: Japan.Apps.Helpdesk@varian.com Latin America: soporte.al@varian.comInternet: Oncology Systems customer site - www.myvarian.com Varian Medical System public site - www.varian.com |  | 218 | 218 units | Worldwide distribution - USA nationwide and the countries of Australia, Brazil, Canada, Chile, Colombia, Georgia, Germany, Hong Kong, India, Ireland, Japan, Macau, Malaysia, Russia, Singapore, Spain, Sri Lanka, and Taiwan | Terminated | February 24, 2012 | 703 | 218
Z-0834-2011 | 56763 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GreenLight MoXy Fiber Optic, part number 0010-2400, for use with the GreenLight XPS Laser System, manufactured by American Medical Systems Inc., San Jose, CA |  |  | July 14, 2010 | 2010 | AMS Innovative Center - San Jose | Product not intended for release was distributed, which may result in old components, not built to specifications, or undergo final testing | Labeling mix-ups | American Medical Systems sent oral notifications that the product used for patient case evaluations was from the 2 year shelf life study. Products have been returned, as the product was with AMS field personnel.For questions regarding this recall call 847-855-6270. |  | 77 | 77 units | Worldwide Distribution, USA including TX, OH, MO, FL, MA, CA, IL, NY, VA, and the country of Australia. | Terminated | January 03, 2011 | 173 | 77
Z-0833-2011 | 57196 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Care E Vac 3 Portable Suction; Battery/AC Portable Aspirator; Ohio Medical Corporation, 1111 Lakeside Drive, Gurnee, IL 60031; Model 758000 |  |  | November 08, 2010 | 2010 | Ohio Medical Corporation | The D-14 diode on the circuit board may fail, resulting in a damaged, overheated circuit board. Also a screw used to attach one of the feet of the unit may damage the battery inside the device if the screw is over-tightened. | Device Design | Ohio Medical sent Urgent Medical Device Correction Notification letters dated November 4, 2010, to all affected customers on November 8, 2010. The letters identified the product, the problems, and the action to be taken by the customer.Customers were instructed to take the affected product out of service, remove diode 14, and replace the foot screw contained therein. Customers were instructed to visit the micro-sites established for these repair procedures for step-by-step instructions on performing the repairs. Accounts with questions after viewing the video demonstration were directed to the Ohio Medical technical assistance line at 224-430-1522. Customers were asked to complete the attached Customer Acknowledgment Form and return the form to the Quality Assurance Department at (fax) 1.847.855.6304, or e-mail to:Complaints@OhioMedical.com, or mail to:Ohio Medical CorproationQuality Assurance1111 Lakeside DriveGurnee, Illinois 60031.For questions regarding this recall call 847-855-6270. |  | 4183 | 4,183 units | Worldwide Distribution - USA nationwide, Puerto Rico, and the coutries of Canada, Israel, Jordan, Kenya, Kuwait, Saudi Arabia, Taiwan, Thailand and Vietnam. | Terminated | March 16, 2012 | 494 | 4183
Z-0835-2011 | 57110 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Mammography Keypad |  |  | October 22, 2010 | 2010 | AGFA Corp. | Product did not receive UL Certification | Other | AGFA sent an Urgent Safety Notice letter dated October 22, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to unplug existing keypad(s)  including USB cable. Plug in loaner Keypad(s). Print out the Fed Ex shipping label that they received via the e-mail notification. Use that label to ship their existing keypad(s) and their corresponding USB cables to the manufacturer. The manufacturer would perform the upgrades and return the keypad(s) along with the USB cables.Upon receipt of the original keypad(s) customers would also receive another e-mail notification with a shipping label so that they may return the loaner keypad(s) to AGFA HealthCare. Customers were to distribute this information to all those who need to be aware of it, and complete the feedback form as soon as possible and return it to AGFA HealthCare.For questions regarding this recall call 838-342-3989.. |  | 26 | 26 units | Worldlwide Distribution - USA including FL, GA, KY, ME, MD, MS, OH, UT and the country of Canada | Terminated | February 01, 2011 | 102 | 26
Z-0832-2011 | 57254 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Urethral Warmer System Model number FPRCH3001 |  |  | November 11, 2010 | 2010 | Galil Medical | This voluntary recall has been initiated because our supplier (Gaymar Industries, Inc.) notified us that these products contain power cords that may crack or fail at or inside the plug due to a prong design problem. Affected plugs can be identified by a round ground pin and black plastic bridges. Cords with plugs that are not affected have U-shaped ground pins and no plastic bridges. Although Gali | Other | Gaymar sent an URGENT: MEDICAL DEVICE RECALL letter dated September 30, 2010, to all consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to: Customers were instructed to:" Inspect units in their possession to determine if they contain affected power cord plugs or if there were any signs of excessive wear or damage. " Immediately discontinue use and quarantine the affected product. " Contact Customer Service Department at 800-828-7341 to arrange for replacement. " Complete the attached form for each unit and follow the instructions given to return the form by fax or e-mail. " If customers further distributed the product to another person or entity, they were to identify their customers and notify them at once of this product recall by including a copy of the recall notification letter and the attached form. For questions regarding this recall call 800-828-7341 or 716-662-2551. |  | 4 | 4 | Worldwide Distribution - USA, including CA, GA, TX and the country of Canada. | Terminated | December 10, 2011 | 394 | 4
Z-0541-2011 | 52039 | No | KLC | Dental chair and accessories. | Dental | Dental | 510(k) | A-dec Decade Plus 1221 Dental Chair |  |  | April 01, 2009 | 2009 | A-Dec Inc | The design of the Decade Plus 1221 chair backrest creates a possible pinch-point between the chair backrest and the underside of the armrests. | Device Design | A-dec sent a letter dated March 23, 2009 to consignees/customers (doctors and network dealers). The consignees were informed that the design of the Decade Plus 1221 chair backrest creates a possible pinch-point between the chair backrest and the underside of the armrests. The potential for this occurs when returning the chair to the upright position. The consignees were advised to reprogram the Position 0/Entry-Exit preset to stop the back before it reaches a full upright position. This increases the gap between the armrest and the back slightly. To reprogram a chair, the firm included excerpts from a chair Instructions for Use (IfU) guide that detail the steps to reprogram a chair. Consignees can contact the firm's Customer Service at 503-538-9471, 1-800-547-1883, or via fax at 503-538-0276.The firm will replace the chair back assembly for all affected dental chairs. |  | 2723 | 2723 units | Worldwide distribution: US and Canada. | Terminated | December 16, 2010 | 624 | 2723