## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities
Z-0790-2011 | 56965 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Arrow Ultra 8 IAB (Intra-Aortic Balloon) Catheter, 8 Fr, 30cc. |  |  | October 11, 2010 | 2010 | Arrow International, Inc., Division of Teleflex Medical Inc. | SuperArrow Flex IAB catheter becomes stuck in the sheath, unable to move the IAB catheter forward or backward, causing a delay in therapy, bleeding or arterial injury. | Device Design | Arrow international initiated a voluntary recall on 12/21/10 by letter requesting users to discontinue use and return all unused 5800 Series IAB with SuperFlex Introducers to Arrow. This field action supercedes the Safety Alert IAB 'Stuck in Sheath' letter issued on 10/8/2010. Questions should be directed towards their local sales representative or the IABP support Line at 1-866-396-2111 or 1-617-389-6400. |  | 5289 | 5289 units | Worldwide Distribution -- USA, Canada, Australia, Belgium, Czechoslovakia, Germany, Russia, Spain, Finland, France, Great Britain, Hungary, Italy, Netherlands, Poland, Saudi Arabia, Switzerland, Brazil, Australia, Chile, Colombia, Ecuador, India, Indonesia, Japan, Korea, Malaysia, Mexico, New Zealand, Panama, Peru, Philippines, Taiwan, Thailand, and Venezuela. | Terminated | February 05, 2015 | 1578 | 69523
Z-0845-2011 | 57192 | No | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Micrus Endovascular Microcoil Delivery System; |  |  | October 21, 2010 | 2010 | Micrus Endovascular Corp., 821 Fox Lane, San Jose, CA 95131 | The firm has received reports of difficulty connecting to the detachment control cable. Loss of the electrical connection will result in failure to detach the coil for deployment inside an aneurysm. | Other | Micrus Endovascular Corp sent Urgent: Medical Device Recall Letters dated October 21, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine inventory (including consignment) and quarantine the product to prevent it from being used or shipped. If the product was further distributed, customers were to notify their customers by sending a copy of the recall letter and were instructed to work with all customers to ensure that all product was returned and that pre-use protocol was followed.The response form was to be completed and returned indicating the number of devices that had been used (or shipped if distributor) and the number that would be returned. Customers were to call the Micrus Customer Service Department at 1-888-550-4129 from 8:00 am to 5:00 Pacific time for arrange for a Return Authorization number and replacement. Recall letters also stressed the importance of always following the pre-use protocol and inspecting the device immediately prior to use. Additionally, connector functionality should be verified and instructions were provided.Questions or concerns should be directed to Micrus Customer Service at 1-888-4120.Consignees who did not respond to the initial recall notification after 15 days were notified again by letter. A third letter was sent after another 15 days if there was still no response. |  | 6755 | 6755 | Worldwide Distribution - USA including AK, AL, AZ, CA, CO, DC, DE, FL, GA, ID, IL, IN, KY, LA, MA, MD, MI, MN, NC, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV and the countries of Australia, Austria, Belgium, Canada, Chile, Czech Republic, Denmark, France, Germany, Hungary, Ireland, Italy, Japan, Netherlands, Poland, Republic of Korea, Russia, Singapore, Slovakia, Spain, Sweden, Switzerland, Turkey, and the United Kingdom | Terminated | January 11, 2011 | 82 | 6755
Z-0846-2011 | 54638 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The device is the NeuViz Dual series Computed Tomography (CT) Scanner System. The device is identified as NeuViz Dual, PN: 989605651321. |  |  | February 10, 2010 | 2010 | Philips And Neusoft Medical Systems Co., Ltd. | Potential for the CT patient table top to become detached from the carrier due to missing washers. | Other | Neusoft Medical Systems Co., China sent an URGENT DEVICE CORRECTION letter dated Feb. 10, 2010 to their distributor in NC. On 02/12/10, The distributor sent the same letter to their consignees.The firm informed this distributor that there was report of the four screws that secure the patient table top to the carrier that did not have the proper washers in place which caused the table top to detach. If the issue would reoccur, there is a potential for the table top to fall and cause patient injury.Customers were instructed to check the patient table top fixing. If it is found that the the assembly is improper customers were instructed to top using the product and contact local Service personnel for repair. Any customer can contact Service Support Department, Neusoft Medical Systems Co, LTD at nms-service@neusoft.com or Service Support Department at helpdesk@neusoft.com for questions about this recall. |  | 26 | 26 devices | Nationwide Distribution including NC | Terminated | April 15, 2011 | 429 | 26
Z-0849-2011 | 56829 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Video Cart Multi-Specialty product labeling is as follows: |  |  | February 23, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Endoscopy Division, dba. ConMed Linvatec, Largo, Florida is recalling five products: (1.) Video Cart Premium 120 VAC (VP6501), (2.) Video Cart Multi-Specialty 120 VAC (VP6502), (3.) Video Cart Articulating Arm 120V (VP6504), (4.) Video Cart (VP8500), and (5.) Power Cord Assembly Self Retracting (VP6547). These products are being recalled due to there is a possibility that the Vid | Component design/selection | ConMed Linvatec sent URGENT MEDICAL DEVICE CORRECTION NOTIFICATION letters to their customers by FedEx. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to review their plugs for damage and determine if they were affected. Distributors were asked to contact affected customers and schedule delivery of their replacement plugs within 30 days. If the account no longer has the product in their possession, or if they have transferred the product, they were asked to respond as such via a Reply form.Customers were asked to document delivery of replacement plugs on the Sales Rep Reply Form and mail back or return by fax to 1-888-370-3062.For questions regarding this recall call the Customer Service Department at 1-888-292-0100 or e-mail Custserv1@Invatec.com. |  | 173 | 173 carts | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, AND WY and the countries of Canada, Chile, Colombia, Denmark, Korea, Mexico, Singapore, South Africa, United Arab Emirates, and Venezuela. | Terminated | March 22, 2012 | 758 | 1761
Z-0775-2011 | 56400 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | HiSpeed FX/I (2200997 with Gantry mo. #2249696) |  |  | February 01, 2010 | 2010 | GE Healthcare, LLC | A defect with the hardware and software displaying the incorrect dose data. | Radiation Control for Health and Safety Act | Important Electronic Product Radiation Warning letter was sent to customers identifying the affected product, explaining the defect and related hazards, giving instructions, and the product correction. Customers are to decrease the CTDI value by a factor of 30% when using 5, 7, 10 mm slice thickness, and by a factor of 20% when using 3 mm slice thickness for more accurate estimate of the CTDI. GE Healthcare will remedy the defect free of charge and provide an updated Operator Manual and configuration update to correct the CTDI display to all affected systems. Questions should be directed to the contact numbers provided in the letter. |  | 102 | 102 total | Nationwide Distribution | Terminated | May 21, 2014 | 1570 | 102
Z-0838-2011 | 57271 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioScore, Inc. AngioSculpt PTA Scoring Balloon Catheter; |  |  | November 16, 2010 | 2010 | AngioScore Inc. | Reports of distal bond failure which has resulted in peeling of the bond and/or detachment of the distal end of the scoring element. May result in vessel injury or embolism. | Device Design | Medical Device Recall notifications were sent to all domestic consignees on Nov 16, 2010 by overnight mail. Foreign distributors, other than Japan, were sent recall letters by e-mail or international overnight delivery on November 16, 2010. The Japanese distributor received notification and notified their customers on November 10, 2010. |  | 17682 | 17682 total units, all sizes and lots | Worldwide Distribution: USA, China, Finland, Italy, Japan, Norway, Saudi Arabia, Spain, Sweden, Switzerland, UAE, United Kingdom | Terminated | October 04, 2011 | 322 | 17682
Z-0859-2011 | 57179 | No | HGC | Obstetric-gynecologic general manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Boston Scientific brand Matrix2 360 Soft SR Coils. |  |  | October 19, 2010 | 2010 | Boston Scientific Corporation, Neurovascular, 47900 Bayside Parkway, Fremont, CA 94538 | The product appears to have premature degradation of the PGLA coating on the Matrix Coil. | Process control | US consignees were notified by letter sent by Federal Express on October 19, 2010. Customers out of the United States were notified by Federal Express if possible and by personal delivery by locally assigned sales representatives as needed. |  | 77364 | 77,364 total units | Worldwide distribution. USA, Argentina, Austria, Australia, Belgium, Bahrain, Brazil, Canada, Switzerland, Chile, China, Colombia, Serbia, Czech Republic, Germany, Denmark, Spain, Finland, France, Great Britain, Greece, Hong Kong, Hungary, Ireland, India, Israel, Italy, Jordan, Japan, South Korea, Lithuania, Latvia, Macau, Mexico, Malaysia, Netherlands, Norway, New Zealand, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovakia, Thailand, Taiwan, Turkey, Ukraine, Vietnam, and Yemen. | Terminated | January 27, 2012 | 465 | 77364
Z-0852-2011 | 57352 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Medical, Universa Soft Ureteral Stent and positioner, 6 Fr x 24 cm, REF USH-624-R, sterile, Cook Urological, Inc., Spencer IN. |  |  | October 18, 2010 | 2010 | Vance Products Incorporated | The firm received a customer complaint that a box labeled as Universa Firm Stents but the box actually contained Universa Soft Stents. The inner product label is correct but the outer box is wrong. | Packaging | The firm sent URGENT: RECALL NOTICES dated 10/18/2010 to its direct accounts. The notice identified the affected product and instructed firms to pull the recalled product from stock and isolate it. Customers are to complete the form and e-mail or fax the form back to Cook. Product should be returned to the recalling firm using the Return Goods Authorization number. Cook Urological will credit customer accounts upon receipt of the returned product. |  | 25 | 25 units | Worldwide Distribution -- USA, including states of WI, TX, NY, SC, AZ, TN, FL, AR, CA, AL, and WA and country of Canada. | Terminated | July 27, 2011 | 282 | 25
Z-0864-2011 | 55541 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Pioneer Plus Catheter PPlus 120 dual lumen device, inserted through a 6F introducer sheath and placed percutaneously into a peripheral vessel. The device tracks to its intended site in the vasculature over a standard length, commercially available 0.014" (0.36 mm) Rapid Exchange (RX) tracking guide wire. It utilizes an extendable, hollow Nitinol guide tip (needle) to facilitate the redirection and placement of a 0.014 (0.36 mm) Over the Wire (OTW) guide wire into peripheral vessels. Specification Developer: Medtronic, Minneapolis, MN. Manufactured by: Accellent, Laconia NH |  |  | April 30, 2010 | 2010 | Medtronic CardioVascular | Needles may fail to retract and thus cause injury. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Safety Information letter to all affected customers. The letter identified the product, the problem,and the action to be taken by the customer. Sales Representatives were to contact each account and complete an acknowledgement form indicating the Physicians had been notified. For questions regarding this recall customers were to contact their local representative. |  | 1080 | 1,080 units | Worldwide Distribution - USA, Germany, Netherlands, Slovakia, and Spain | Terminated | January 06, 2011 | 251 | 1080
Z-0865-2011 | 57457 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Groshong 3 Fr. Single-Lumen PICC Intermediate Tray with Safety Introducer, Reorder Number 7717300, Bard Access Systems, Inc., Salt Lake City, UT 84116. |  |  | December 10, 2010 | 2010 | Bard Access Systems | Prepackaged sterile latex gloves beyond their expiration date were packaged in surgical kits. | Employee error | BARD Access Systems sent an URGENT: GROSHON* 3 FR SINGLE-LUMEN PICC INTERMEDIATE TRAY WITH SAFETY INTRODUCER RECALL letter dated December 10, 2010. The letter identified the product, the problem,and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine the product subject to recall, discontinue use, and cease further distribution of any the affected product. If the product was further distributed, those customers were to be notified immediately with a copy of the recall notification letter. Any remaining product was to be returned to Bard. For information regarding this recall call 800-290-1689. |  | 195 | 195 units | Nationwide Distribution including AL, AZ, CA, GA, IA, MN, MO, NE, NJ, NY, OH, OR, SD, and WI. | Terminated | May 04, 2011 | 145 | 195
Z-0882-2011 | 57416 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Air bubble detect module, catalog 802110, Terumo Cardiovascular System Corp, Ann Arbor, MI. |  |  | December 02, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The firm identified that a component within the System 1 ABD and Occluder Modules that may malfunction. The Air Bubble Detector Module (ABD) is used with an air sensor to detect gross air bubbles in the perfusion circuit. The Occluder module is used with the Occluder head to provide a computer controlled tube clamping mechanism to regulate the flow in the venous line.The firm identified a fai | Component design/selection | Terumo Cardiovascular Systems Corporation issued an URGENT MEDICAL DEVICE CORRECTION letter dated 12/2/2010 to the Director of Cardiac Operating Room Services or Operating Room Nurse Manager. The letter identified the product, the problem, and the action to be taken. The letter stated the firm will replace all modules in the affected population, and that a representative would contact the consignee regarding the replacement schedule.Customers were to review the Medical Device Correction.Assure that all user were aware of the notice.Confirm receipt of the communication by faxing the attached Customer Response Form to the fax number indicated on the form. For questions regarding this recall call 1-800-521-2818. |  | 508 | 508 devices | Worldwide Distribution - US, Austrailia, Belgium, Canada, Chile, Columbia, Costa Rica, Dominican Republic, Guatamala, Hong Kong, Indonesia, Japan, Korea, Malaysia, Mexico, Phillipines, Saudi Arabia, Singapore, Taiwan, Thailand, United Arab Emirates, and Vietnam | Terminated | April 17, 2012 | 502 | 729
Z-0885-2011 | 54685 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Guider Softip TM XF Guide Catheter |  |  | December 18, 2009 | 2009 | Boston Scientific Corporation | Complaint about product quality: degradation of polymer portion of the distal section of the catheter. | Other | Boston Scientific sent an Urgent Medical Device Recall and Correction letter dated December 18, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product and fill out the Customer Instructions and REPLY VERIFICATION TRACKING FORM indicating the status of the affected product and send a copy of the form by fax to 510-624-1600 or e-mail to Joe.Nesci@beci.com. |  | 94296 | 94,296 units | Worldwide. US Domestic including US Federal Government, and foreign countries. | Terminated | December 19, 2011 | 731 | 94296
Z-0867-2011 | 54651 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Stryker Digital Capture Pro 2 device (240-050-800, CD version) and (240-050-810, DVD version) designed for digital documentation of medical procedures. Manufactured by Stryker Endoscopy, San Jose, CA. |  |  | December 11, 2009 | 2009 | Stryker Endoscopy | Device will not able to assign the correct time/date stamp for saved filed after December 31, 2009. On Jan 1, 2010, the device system date for all SDC Pro 2 units will re-start from 1 Oct 2001 and begin incrementing from that point onwards. The system date will be incorrect and users will not be able to correct this issue. Problem will not affect how pictures or videos are captured using SDC Pro | Software design | Stryker Endocoscopy sent an URGENT: Device Correction letter dated December 11, 2009, to the OR Supervisor/Risk Manager and all Stryker Sales Representatives. The letter identified the product, the problem, and the action to be taken by the Sales Representatives. Sales Representatives were to identify affected accounts. Go to customer accounts and install the upgrade for each affected unit. Upon completion of the upgrade, the sales rep was to sign the acknowledgment form indicating that all of their affected accounts received the appropriate software fix and mail to Stryker Endoscopy or fax both sides of the postcard to 408-754-8378 or scan in both sides of the postcard and email to sdcpro2recall@stryker.com.For questions regarding this recall call 408-754-2738 or 408-754-2000. |  | 3070 | 3070 units | Worldwide Distribution - USA including AL, AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL IN, IA, KS, KY, LA, MA, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, NY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and the countries of Africa, Australia, Brazil, Cnanda, China, France, Germany, Hong Kong, Iberia, Italy, Japan, Korea, Latin America, Mexico, Netherlands, Poland, Puerto Rico, Scandinavia, Singapore, Switzerland, Thailand, and the UK | Terminated | January 08, 2011 | 393 | 3070
Z-0874-2011 | 56388 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific brand Synchro2 0.14" Guide Wire, 300 cm, Standard; |  |  | June 07, 2010 | 2010 | Boston Scientific Corporation | The Guidewire has the potential to flake off polytetrafluroethylene (PTFE) particulate, which has the potential to embolize and result in infarction when introduced into a patient's circulatory system. | Other | Boston Scientific Corporation sent an Urgent Medical Device Recall letter dated June 7, 2010, by Federal Express to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use of and segregate the affected product. Complete and return the Reply Verification Tracking Form even if they did not have any affected product. Contact the Customer Service Call Center to obtain a Returned Goods Authorization (RGA) Number at (800) 811-3211, Option #1. E-mail or Fax the completed and signed Reply Verification Tracking Form to email Bob.Ware@bsci.com or the Field Action Fax (510) 624-2475.PackageShip and return the affected product along with the original Reply Verification Tracking Form by Federal Express to:Boston Scientific CorporationUS Distribution CenterBoston Scientific Marina BayCustomer Fulfillment Center500 Commander Shea Blvd.Quincy, Massachusetts 02171RGA: ____________For questions regarding this recall call (510) 624-2563. |  |  |  | Worldwide Distribution - USA including CA, CO, IN, MD, MI, OR, TX, and VA and the countries of Australia, Austria, Belgium, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Georgia, Germany, Great Britian, Greece, Honduras, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Puerto Rico, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Korea, Sweden, Switzerland, Thailand, Tunisia, Turkey, and Ukraine. | Terminated | January 10, 2011 | 217
Z-0884-2011 | 56986 | No | DSA | Patient transducer and electrode cable (including connector). | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 System Interface Cable, 34-pin to 10-pin, Red, 10' - Part Number: D-1286-18; Catalog Number CR3410CT |  |  | October 05, 2010 | 2010 | Biosense Webster, Inc. | When connecting this cable to the mapping system, loss of Body Surface ECG signals may occur. This may result in inaccurate locations appearing for catheters displayed 3D on the CARTO monitor. | Device Design | Biosense Webster issued an Urgent Field Safety Notice letter dated October 5, 2010. The letter identified the product, the problem,and the action to be taken by the customer. Customers were instructed to return the affected cable, accompanied by the attached Return Authorization form.For questions regarding this recall call Biosense Webster at 909 839-8874. |  | 1692 | 1692 | Worldwide Distribution: USA, including the states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, KS, KY, MA, MD, MI, MN, NC, NE, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, WA, WI, and WV, and the countries of Belgium, China, Israel, Japan, Singapore, and Thailand. | Terminated | May 07, 2012 | 580 | 1692
Z-0866-2011 | 57418 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Merit One Step Centesis Drainage Catheter |  |  | November 17, 2010 | 2010 | Merit Medical Systems, Inc. | The catheter's valved hub may fail resulting in an inability to drain fluid. | Device Design | Merit Medical Systems, Inc. sent a Product Safety Notice letter dated November 18, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to provide a copy of the Safety Notice form to their applicable customers. Advise them to immediately stop using the affected product and to discard it.Customers were asked to complete and sign the Safety Notice forms and fax them to 1-801-208-3378 and send the original Safety Notice Form with both signatures via Federal Express 2nd day air to:Customer ServiceMerit Medical Systems, Inc.1600 West Merit ParkwaySouth Jordan, UT 84095-2416entirety and faxed them For questions regarding this recall call 801-316-4929 or Customer Service at 801-208-4365. |  | 20863 | 20,863 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, AND WV and the countries of Canada, Finland, Germany, Netherlands, Singapore, Sweden, Switzerland, and the UK | Terminated | March 23, 2012 | 492 | 20863
Z-0886-2011 | 56347 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Varian brand C-series Clinac, Trilogy and Novalis Tx, Software Versions 6.X and 7.X, Model Numbers: H14, H27, H29, HCX, Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | June 10, 2010 | 2010 | Varian Medical Systems Oncology Systems | The Auto goto or Auto setup functions in C-Series version 7 software ignore the couch angle and exceed motion zone (unless a tolerance is defined by the user); it has the potential for collision with the patient on the couch. | Other | Varian sent a Customer Technical Bulletin letter to all affected users on May 27, 2010, and was also posted to Varian's customer support site: http://www.MyVarian.com. The letter identified the product, the problem, and provided instructions on the recall. Customers were instructed to return the proof of notification (via completed "Proof of Notification Form or appropriate Postal system certified receipt of delivery to Varian. For questions regarding this recall call 650-483-3153. |  |  |  | Worldwide Distribution - USA including CA, TX, IL, GA, NJ, and NV and the countries of Australia, Belgium, Brazil, Canada, China, Denmark, France, Germany, Hong Kong, India, Italy, Japan, Middle East, Netherlands, Russia, Spain, Switzerland, and the UK | Terminated | December 23, 2011 | 561
Z-0888-2011 | 51217 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Floorstand. An accessory of the Optical Guidance Platform, Model number HZ5 for use with a charged particle accelerator (such as the Clinac 600C). |  |  | January 26, 2009 | 2009 | Varian Medical Systems Oncology Systems | Potential patient impact -- limit switches for Floorstand may be inadequate for some models and allow for potential impact. | Device Design | Varian sent an urgent device correction notification letter 01/26/2009 to all affected users, with a description of the problem and user corrective action steps via certified mail. The letter will also be distributed to Varian Sales, Marketing and Service organizations. |  | 3 | 3 units | Product was distributed to 3 consignees in Florida and Wisconsin. | Terminated | February 24, 2012 | 1124 | 3
Z-0891-2011 | 54146 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console, versions: 8.0.12, 8.0.23 SP2, 8.1 SP2, 8.1.1.3, 8.1.1.4, 8.1.2, 8.1.3, 8.1.4, 8.3, 8.6, 8.6.1, 8.6.2, 8.8.0; Model Number: H51; Distributed and manufactured by: Varian Medical Systems Inc., Palo Alto, CA |  |  | October 16, 2009 | 2009 | Varian Medical Systems Oncology Systems | Software anomaly may result in a design deficiency causing Plan Edit Settings to malfunction under specific circumstances, where the user selects to use a Cone Beam CT set up field. | Device Design | The firm, Varian Medical Systems, sent a Product Notification Letter/Field Safety Letter entitled "URGENT MEDICAL DEVICE CORRECTION" dated October 16, 2009, to all affected customers. The letter described the product, problem, actions to be taken by the customers and Varian. The customers were instructed to use caution when creating CBCT imaging field and limit use of the 'select all' button when acquiring treatment parameters; review departmental policies and procedures regarding the acquisition of any gantry angle at the 4D ITC, and advise the appropriate personnel of this letter. The firm, Varian, recommended that only vertical, longitudinal and lateral couch parameters are acquired when creating a CBCT field at the 4D ITC; notified all affected customers of this document, and will correct this issue in a future release of the 4D ITC software. If owners of $D ITC, version 8.0.12 and higher, have any questions regarding this product notification letter, please contact Varian Medical Systems Help Desk at 1-888-VARIANS (888-827-4265) for USA and Canada, +41 41 749 8844 or support-emea@varian.com for Europe;Emails for North American: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com and Latin American: soporte.al@varian.com, and Internets for Oncology Systems- www.myvarian.com and Varian Medical Systems- www.varian.com. |  | 3629 | 3629 | Worldwide distribution: USA and countries including: Australia, Asia, China, Japan, Latin American, New Zealand , North America, South East Asia. | Terminated | July 20, 2011 | 642 | 3629
Z-0890-2011 | 56326 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Treatment Planning System, a subsystem of the CyberKnife Robotic Radiosurgery System radiation therapy device, manufactured by Accuray Inc., Sunnyvale, CA. |  |  | July 06, 2010 | 2010 | Accuray Inc | Latches responsible for securing the cover to the device may come lose if not properly secured and the cover may unexpectedly drop off. A design change is planned. | Device Design | Accuray sent an "URGENT DEVICE CORRECTION" letter dated July 6, 2010, to all customers. The letter described the product, problem and actions to be taken by customers. The customers were instructed to assure the LINAC cover is secure following reattachment and adhere to and follow the instructions included in the letter. Note: A new design is being developed, and will be distributed to user base once available.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1-877-668-8667 (USA) or +1-408-716-4700 (non-USA) or customersupport@accuray.com. |  | 207 | 207 units | Worldwide distribution: US including: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI and WV, and countries including: Japan, China, Vietnam, Taiwan/China, South Korea, Hong Kong/China, Thailand and Malaysia. | Terminated | October 04, 2011 | 455 | 207
Z-0887-2011 | 57338 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | TiTLE 2 Polyaxial Spinal System, Counter Torque Assembly Tool, Part ET1132-01, Zimmer, Warsaw, IN. |  |  | October 21, 2010 | 2010 | Zimmer Inc. | There is a chance that the ball plunger could fall out during surgery and fall into the surgical site. There could be a delay in surgery to remove the ball plunger. If the problem was undetected, the ball plunger could remain in the patient causing pain and potentially necessitating subsequent surgery for removal. | Other | Zimmer sent an URGENT MEDICAL DEVICE RECALL LETTER dated October 21, 2010, to all affected customers. The letter identified the product, the problem, the action to be taken by the customer.Customers were instructed to:Inspect their inventory, stop using the device, and place it in quarantine. Contact Zimmer Customer Service at 1-800-774-6368 to obtain a return authorization number. Complete the Inventory Return Certification form and fax a copy to (952) 837-6958, or scan and email to qualitycompliance@zimmer.com.Return the recalled product along with the original completed Return Certification Form to:Zimmer Spine, Inc.Attn: RA Receiving5301 Riata Park Ct.Austin, Texas 78727For questions regarding this recall call Zimmer Spine Customer Service at 1-866-774-6368. . |  | 68 | 68 | Nationwide Distribution including AL, AZ, CA, CO, FL, MI, MN, NC, PA, OH, OK, TN, TX, WA, and WI. | Terminated | October 12, 2011 | 356 | 68
Z-0843-2011 | 57293 | Yes | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Medical CombiSet BVM Hemodialysis Blood Tubing Set |  |  | November 19, 2010 | 2010 | Fresenius Medical Care Holdings, Inc. | Hemodialysis Blood tubing Sets may develop kinking of the arterial line. | Device Design | Fresenius Medical issued recall notification dated 11/19/10, "URGENT RECALL FMCNA CombisetÂż BVM Hemodialysis Blood Tubing Sets (Part Numbers 03-2695-9 and 03-27957)", by certified mail with return receipt. Customers have been instructed to check their stock immediately to determine if they have any affected lots on hand. If customers have the affected lots, they have been instructed to discontinue use immediately and place the product in a secure segregated area for return to Fresenius Medical Care North America. Customers have been instructed to contact Fresenius Medical Care Customer Service Team Representative for instructions on how to return the recalled product. |  | 190080 | 190,080 units (7920 cases x 24 units) | Distribution Nationwide and Canada | Terminated | November 03, 2012 | 715 | 190090
Z-0837-2011 | 57099 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook 400ES Safety Infusion System, Model Number 621-400ES. For use with B. Braun Medical Inc. Horizon Pump IV Sets |  |  | August 24, 2010 | 2010 | B Braun Medical, Inc | Possible pump stop with audible alarm, but no visual error indication that the pump is not delivering the medication. Under certain conditions in a network environment, Management Processor may become non-responsive, ceasing normal operation. | Component design/selection | B. Braun Medical Inc. notified consignee by letter on August 24, 2010. |  | 2107 | 2,107 Pumps (582 Pumps to Winchester) | Virginia and New York | Terminated | May 29, 2014 | 1374 | 2107
Z-0889-2011 | 57414 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | TandemHeart Transseptal Cannulas-EF set (THTC-EF) - Dilator assembly |  |  | December 02, 2010 | 2010 | Cardiac Assist, Inc | Component (Dilator of TandemHeart Transseptal Cannula Set-EF) does not meet specifications | Other | CardiacAssist, Inc. sent a Voluntary Recall letter dated November 12, 2010, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignees. Consignees were instructed to examine inventory for the affected product and remove from inventory and identify it as a recalled product by applying the enclosed "CAUTION: DO NOT USE" sticker(s) and quarantine to prevent usage.Consignees were to complete the form included with the letter and fax a copy of the form to (412) 963-9739 to coordinate the return and replacement of any affected product. For questions regarding this recall call (412) 963-7770, ext. 271. |  | 164 | 164 | Nationwide distribution including AZ, CA, CO, CT, DC, FL, GA, IL, LA, MA, MD, MI, MN, NC, NH, NJ, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, and WI. | Terminated | December 01, 2011 | 364 | 164
Z-0853-2011 | 53047 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | RoboCouch Patient Support System, model number 025007, a subsystem of the Cyberknife Robotic Radiosurgery System, manufactured by Accuray Incorporated, Sunnyvale, CA. |  |  | July 21, 2009 | 2009 | Accuray Inc | Mechanical Failure-- Patient Support System's table top may unexpectedly descend several inches. | Device Design | Consignees were notified verbally about the issue when an inspection was scheduled. |  | 26 | 26 | Worldwide Distribution -- USA, including states of CA, FL, VA, AL, CT, MA, NC, TN, PA, RI, and NE and countries of Turkey, France, UK, Saudi Arabia, and Switzerland. | Terminated | January 12, 2011 | 540 | 26
Z-0892-2011 | 56350 | [6 empty] | Varian brand C-Series Ximatron Radiotherapy Treatment Stimulator, |  |  | January 19, 2010 | 2010 | Varian Medical Systems Oncology Systems | There is a potential for failure of a lead screw drive system or image intensifier lead screw drive assembly, if normal maintenance inspections are not followed correctly. | Other | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter via certified mail receipt on January 19, 2010, to all direct consignees. The letter identified the product, the problem, and the action the consignee should take.Consignees were instructed to have maintenance personnel perform the recommended measurements as described in the Varian Customer Technical Bulletin.Ensure that all measurements made as part of the procedure are recorded in the 'wear table' that is part of the procedure.Previous wear measurements should be compared to new measurements and any large increase in wear should be investigated immediately.Pay special attention to the "Recommendations based on wear measurements" table.Ensure that new wear measurements comply with the operating parameters described in the table. For any questions consignees are to contact Varian Medical Systems Customer Support Services at one of the locations described in the CTB.For questions regarding this recall call (650) 424-5731. |  |  |  | Worldwide Distribution including USA, Australia, New Zealand, Europe, Southeast Asia, China, Japan, Latin America, and Canada. | Terminated | February 24, 2012 | 766
Z-0894-2011 | 57463 | No | OFM | Impedance phlebograph. | Cardiovascular | Cardiovascular | Enforcement Discretion | Blood Center Kit w/Chloraprep One-Step, Sterile/Disposable/Single Patient Use, Reorder No. Q921CL, Manufactured for : Blood Center of Wisconsin, Milwaukee, WI. |  |  | December 06, 2010 | 2010 | Continental Medical Labs,Inc | Continental Medical labs, Inc, is recalling various lots of Blood Center Kit and Port VAD Kits because Covidien reported that testing could not pick up an Over-Sulfated Chondroitin Sulfate (OSCS) contaminant in the purified heparin sodium USP. | Material/Component Contamination | Covidien sent an URGENT PRODUCT RECALL letter dated November 12, 2010, to their distributors. Continental Medical Labs,Inc telephoned on 12/06/2010 and later sent a letter dated December 7, 2010, with an enclosed copy of the recall notice. The letters were addressed to McCormick Medical Distribution, Inc and to Medico Mart, Inc. The letter described the problem and product affected. Advised consignees to inform all their customers to return the product as well as the distribution information as soon as possible so CML could notify Covidien for replacements.For questions regarding this recall call (262) 723-6715, ext. 313. |  | 900 | 30 cases, 30 kits per case | Nationwide Distribution including WA and WI | Terminated | July 05, 2012 | 577 | 900
Z-0899-2011 | 51401 | No | GBS | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Concentric 7F Balloon Guide Catheter (percutaneous catheter per 21 CFR 870.1250), REF 90072, manufactured by Concentric Medical, Inc., Mountain View, CA |  |  | February 27, 2009 | 2009 | Concentric Medical Inc | Product malfunction-- hub may leak at connection to the shaft due to insufficient adhesive curing. | Nonconforming Material/Component | Concentric Medical products sent an URGENT VOLUNTARY MEDICAL DEVICE REMOVAL letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to examine their inventory and identify any affected product.Quarantine the product and call Concentric Medical Customer Service at 877-471-0076 (USA toll-free) or 650-938-2100 to receive instructions for returning the devices.For questions regarding this recall call 877-471-0076. |  | 36 | 36 units affected. | Worldwide Distribution - USA including AL, FL, GA, IL, MN, NY, OH, OR, PA, TX, VA, WA, and WV and one consignee in Turkey and one in Switzerland | Terminated | January 13, 2011 | 685 | 36
Z-0897-2011 | 51274 | No | INQ | Powered table. | Physical Medicine | Physical Medicine | 510(K) Exempt | Precedence SPECT/CT System, Precedence 16 3/8, Upper Patient Pallet, A single photon nuclear medicine and image from an X-ray computed tomography system, |  |  | January 26, 2009 | 2009 | Phillips Nuclear Medicine | A problem has been identified with some upper patient pallets on Precedence SPECT/CT System, which may reduce its load bearing capacity and may result in its cracking. If this were to happens, could pose a risk to patient. | Process change control | The firm, Philips Healthcare, sent an "Urgent - Field Safety Notice" dated January 26, 2009 to all consignees/customers. The notice described the product, problem and action to be taken by customers and by Philips. The customers were instructed to reference their system's System Operation and System information Manuals regarding loading and unloading of patient; monitoring the patient during a study ,and careful monitoring during unsupported (cantilevered) table translation. Philips has ordered replacement parts and they will have a Philip Healthcare Field Service Engineer contact and visit the customers site within two months to replace the affected upper patient pallet.If you need any further information or support concerning this issue, please contact your local Philips representative at: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: and Oncology or your local Philips Healthcare office. |  | 2 | 2 units | International distribution only: Mexico and Saudi Arabia. | Terminated | January 13, 2011 | 717 | 2
Z-0896-2011 | 54718 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Acuson Antares Ultrasound System |  |  | September 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Software safety guard measures on the device malfunctioned to prevent the temperature from going over its limits when the device is used. | Software design | Field notification in the form of a Customer Safety Advisory Notice was initiated on 09/30/2009 by the firm, and the notice was mailed to all affected customersSiemens sent a CUSTOMER SAFETY ADVISORY NOTIFICATION letter on September 30, 2009. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconnect and reconnect the V5Ms transducer between uses until their local Customer Service Engineer installs the software update to their system. For any questions customers were to contact their local service support person.For questions regarding this recall call 650-694-5398. |  | 452 | 452 | Worldwide Distribution - United States (nationwide) and the countries of Australia, Austria, Azerbaijan, Brazil, Belarus, Canada, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Guatemala, Croatia, India, Indonesia, Italy, Jordan, Japan, Latvia, Malaysia, Mexico, Moldova, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Taiwan, Turkey, Ukraine, United Kingdom, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Terminated | June 18, 2012 | 992 | 452
Z-0898-2011 | 57218 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow-Howes (TM) Multi Lumen Central Venous Catheterization Set with Blue FlexTip(R) Catheter. Catalog #: CV-15703. |  |  | November 09, 2010 | 2010 | Arrow International Inc | One lot of product was inadvertently mislabeled with a product lidstock identifying the product as a triple-lumen catheter with product number CV-15703, but a double-lumen catheter with part number CV-16702 was actually packaged in the kit. | Process control | The recalling firm issued Urgent Medical Recall letters dated 11/9/10 to inform their customers of the problem and the need to return any unused product. Customers are to check their stock for the affected product, contact Arrow Customer Service for a Return Authorization Number, and complete the enclosed Recall Acknowledgement &amp; Stock Status Form. For customers with no affected product in stock, the form should still be completed and also faxed to Arrow Customer Service. Affected product should be returned along with the original completed form to Arrow International, Inc. Arrow will replace the affected product once it has been returned. Additional information or clarification can be obtained from Arrow Customer Service at 1-919-361-3964. |  | 1150 | 1150 units | International Distribution Only -- Argentina and Columbia. | Terminated | September 29, 2012 | 690 | 1150
Z-0895-2011 | 57462 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Disc Space Distractor |  |  | December 07, 2010 | 2010 | Stryker Spine | Stryker Spine has received reports of the Disc Space Distractors failing during surgical procedures. | Other | Stryker Spine sent an URGENT PRODUCT RECALL letter dated December 7, 2010, via Fedex to all affected customers.. The letter identified the product, the problem, and the action they should take. Customers were instructed to examine their inventory and hospital locations to identify the product.Reconcile IBD Disc Space Distractors with the catalog and lot numbers referenced. Complete the Customer Response Form and distribution sheet and fax a copy of the Customer Response Form to Regulatory Compliance at 201-760-8370.Return the affected product using the enclosed pre-paid mailing label to Regulatory Compliance,Stryker Spine 6 Pearl CourtAllendale, New Jersey 07401For questions regarding this recall call 201-760-8206. |  | 21 | 21 units | Nationwide Distribution including AZ, FL, MA, NY, PA, and TX | Terminated | November 14, 2012 | 708 | 21
Z-0900-2011 | 57436 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology ECG Management |  |  | June 16, 2010 | 2010 | McKesson Provider Technologies - Medical Imaging Group | When importing ECG files from Mortara ECG carts using DICOM interface, the Horizon Cardiology ECG Management System does not display the amplitudes of the waveform properly. | Device Design | McKesson Medical sent an advisory notice letter dated June 16, 2010, to all affected customers. Advisory notice ADV01646 was distributed to 13 affected clinical sites from August 30, 2010 to September 28, 2010. The letters identified the product that is affected, the circumstances under which patient safety issue may occur, the risk, and the immediate actions that need to be taken along with a temporary fix. Customers were instructed to use the ECG cart printout to verify results of procedures acquired by Mortara ECG carts. Users should be aware of the problem when referrring to past ECG procedures and when performing a visual comparison with ECG procedures acquired by Mortara ECG carts.Customers should disable the Glasgow Serial Comparison for ECG procedures.McKesson Support would contact the customer to coordinate the deployment of the software update on all the affected applications and Horizon Cardiology client applications. For any questions regarding this recall, customers should call the Customer Support department at 1-866-777-0202 and speak to their Support Manager. |  | 16 | 16 units | Nationwide Distribution including AZ, CA, FL, GA, MA, PA, UT, and WI. | Terminated | August 17, 2012 | 793 | 16
Z-0904-2011 | 52716 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | VISU 200 on S8 Floor Stand, Manufactured by Carl Zeiss Surgical GmbH, Germany |  |  | May 18, 2009 | 2009 | Carl Zeiss Meditec Inc | Potential Retina Damage-- An error in the user manual may result in misapplied filter and injury during treatment. | Labeling mix-ups | Carl Zeiss Meditec Inc sent a SAFETY ALERT letter dated May 18, 2009, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were reminded to use the correct retinal protection filter (GG475) when using the OPMI VISU 200 system during ophthalmic surgery.Verify the position of the retinal protection filter and inform their staff of the customizable feature. Verify that the correct filter was in place. Customers were asked to complete the acknowledgement form indicating receipt of the alert. Sign and return the form to:Carl Zeiss Meditec in Dublin, California via the enclosed self-addressed, pre-paid Federal Express envelope.For any questions customers should call (925) 557-4422. |  | 2540 | 2540 units | Worldwide Distribution including USA (nationwide) and the UK | Terminated | February 09, 2011 | 632 | 2540
Z-0901-2011 | 57377 | No | LSZ |  |  | Anesthesiology | PMA | One Cat. Life Pulse High Frequency Ventilator, Catalog Number: 203A, Manufactured by: Bunnell Incorporated Salt Lake City, Utah, Rx Only. |  |  | October 25, 2010 | 2010 | Bunnell, Inc. | Product software coding error may compromise patient safety. | Software design | Customers were notified by an "Urgent: Medical Device Recall or Correction" letter starting 10/22/10. The letter identified the product and described the issues. Customers were given the option of either returning their ventilator(s) for rework or to receive installation/re-test instructions with an upgrade kit to rework themselves. The letter states that all Life Pulse ventilators in clinical use must be upgraded and there is no charge for the software upgrade. |  | 660 | 660 total | Worldwide Distribution -- USA, Australia, Canada, Czech Republic, Korea, Malaysia, Trinidad &amp; Tobago, and Uruguay. | Terminated | April 17, 2012 | 540 | 660
Z-0903-2011 | 57021 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PROCEED Surgical Mesh; Product Code PCDG1; :Lot BMG618; Ethicon, Inc. a Johnson &amp; Johnson Company. |  |  | October 18, 2010 | 2010 | Ethicon, Inc. | There is a potential for delamination in one lot of Proceed surgical mesh. | Nonconforming Material/Component | Ethicon sent Urgent Voluntary Product Recall letters via Fed Ex on October 18, 2010. The letters described the issue with the affected product and instructed customers not to use it. They are to immediately remove the recalled product from their inventory and return it according to the instructions provided. A reply card is attached and must be filled out and returned to Stericycle whether or not the customer has product in their possession. Separate correspondence has been sent to end customers stating that if they have purchased the recalled product through a distributor, to return the product directly to Stericycle. Replacement product will be provided at no additional charge for all recalled product. |  | 87 | 87 products | Worldwide Distribution | Open, Classified |  |  | 87
Z-0909-2011 | 51550 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife System Robotic Radiosurgery System, (Robocouch Patient positioning System) Synchrony Respiratory Tracking System, a Radiation therapy device, Model Numbers: 026430, 028186, Product is manufactured and distributed by Accuray, Inc., Sunnyvale, CA |  |  | March 16, 2009 | 2009 | Accuray Inc | The product has the potential for the extension coupling of the Synchrony extension mounting hardware to become loose over time. The assembly weighs approximately 17 pounds; an unexpected contact with users or patients may result in serious injury if the boom descends unexpectedly. | Device Design | Accuracy Inc., sent an URGENT ADVISORY NOTIFICATION letter, dated March 16, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their system to see if it was affected. Customers were instructed to contact Accuracy Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 or e-mail customersupport@accuracy.com to schedule service. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Turkey, Russia, France, Italy, the Netherlands, the UK, Spain, Greece, Switzerland, Germany, India, Taiwan, China, Hong Kong, Korea, Thailand, Malaysia and Japan | Terminated | January 19, 2011 | 674
Z-0910-2011 | 53778 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | C-Series Clinac, Trilogy and Novalis Tx used for SRS treatments, Model Numbers H14, H18, H27, H29, HCX, manufactured by Varian Medical Systems, Palo Alto, CA. |  |  | October 08, 2009 | 2009 | Varian Medical Systems Oncology Systems | Product may deliver radiation treatment to areas larger than intended,to healthy tissue. | Nonconforming Material/Component | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter dated October 8, 2009, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Varian recommended to consignees that test runs be performed prior to actual patient treatment to ensure that the entire system is behaving as intended and that appropriate quality assurance procedures should be followed prior tothe delivery of any treatment plan involving radiation therapy. Operators should monitor the patient and the treatment console continuously during treatment. For questions call USA and Canada - 1-888-827-4265, Europe +41 41 749 8844. |  | 5208 | 5208 units | Worldwide Distribution | Terminated | February 24, 2012 | 869 | 5208
Z-0914-2011 | 56624 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Instruments, SmartPump Tourniquet System, Report Printer Kit, REF 5920-012-000, Rx only, Stryker Instruments, Kalamazoo, MI (USA), 49001, (269) 323-7700, (800) 253-3210. |  |  | April 13, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | This recall is related to the Electri-cord recall. There is a risk when excessive force is applied to the plugs, a fracture can develop on the prongs inside the molded section of the plug. Over time the fracture may lead to arcing inside of the molded section of the plug and a build up of heat which can cause charring, electric shock or fire. To this date there have been zero complaint for this | Component design/selection | Stryker Instruments sent an Urgent - Medical Device Recall Notification letter dated April 13, 2010, to all affected customers. The letter identified the product, the problem,and the action to be taken by the customer. Stryker notified customers again on May 24, 2010, and again on September 14, 2010, via email, fax, or telephone.Customers were instructed to locate the units and visually inspect each affected cord. Remove the affected cord and dispose of it using their healthcare facilities disposal procedure. If the cord did not show signs of damage, the unit could remain in service until a replacement cord was available. Customers were to return the enclosed business reply form to confirm receipt of the notification letter and identify how many affected units were currently in their inventory. Stryker instruments would send out replacement cords.If the customer loaned or sold any of the units listed in the notification, the customer was to forward a copy of the notice to the new users and advise Stryker of the new user's name and their new location. If products were disposed of and were no longer in use, the customer was to advise Stryker of their obsolescence by providing them with their serial numbers. For questions regarding this recall call 800-800-4236, ext 3808 or ext 3584 Monday -Friday 8am-5pm (EST). |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Portugal, Canada, China, Taiwan, France, Germany, Greece, Spain, India, Italy, Korea, Switzerland, Singapore, and England. | Terminated | May 16, 2011 | 398
Z-0906-2011 | 56871 | No | FGS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | InfraVision IR Illuminator Console, model number 220-180-521, manufactured by Stryker Endoscopy San Jose, for use with the Stryker InfraVision system. |  |  | September 29, 2010 | 2010 | Stryker Endoscopy | Product may not adequately visualize IR light emitted from U- and E- kits due to a manufacturing error. | Nonconforming Material/Component | The firm, Stryker Endoscopy, sent an "Urgent: Device Correction" letter dated September 28, 2010, t o all consignees/customers via a Fed Ex Mailer that included notification letter, acknowledgment receipts and updated instructions for use. International Stryker sites received product advisory notification. Stryker reps at those sites are responsible for notifying customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to discard any old IFUs from the InfraVision Console and ensure that they only use InfraVision compatible scopes; read the newly updated IFU provided with the letter; and complete and sign the enclosed "acknowledgement of receipt" form and fax to (408) 754-8378 or scan in or email to infravision@stryker.com. Note: Not returning the "acknowledgement of receipt" card will result in additional notices.Please contact your local Stryker sales representative to obtain information on scopes that are specific to the InfraVision system.If you have any questions, please contact us at 1-800-624-4422 or email at infravision@stryker.com. |  | 861 | 861 units. | Worldwide distribution: USA and countries including: Australia, Africa, Brazil, Canada, China, France, Greece, India, Italy, Korea, Latin America Mexico, New Zealand, Portugal, Romania, Singapore, Sweden, Switzerland, Taiwan, Thailand, and United Kingdom. | Terminated | February 09, 2011 | 133 | 861
Z-0907-2011 | 56773 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE and ONCOR Digital Linear Accelerators configured with the 160 leaf Multi-Leaf Collimator (MLC) and MEVATRON and PRIMUS Digital Linear Accelerators configured with the High Performance Defining Collimator (HPD); Product is manufactured by Kemnath, Germany-Siemens AG OCS and Siemens AG, Erlangen, Germany-Siemens AG OCS and distributed by Siemens Medical Solutions, Oncology Care Systems, Concord, CA; |  |  | August 13, 2010 | 2010 | Siemens Medical Solutions, Oncology Care Systems, 4040 Nelson Ave., Concord, CA 94510 | The potential safety issue with the 160MLC or the HPD collimator of mechanical noises and/or sticking of the collimator when rotating the collimator. The potential for mistreatment to a patient should the collimator not be aligned correctly if the bearings are damaged. | Pending | Urgent: Medical Device Correction Safety Advisory Notice letters were sent by certified mail commencing August 3, 2010. The letter described the issue and when it occurs, the measures users could take, and how Siemens would address the issue. It is recommended that customers include a check of the field alignment at 90 and 270 degrees gantry angle in the daily check prior to start of patient treatments. A copy of the Safety Advisory Notice should be included in the Owner's Manual and questions should be directed towards the local Siemens Service Engineer or Application Specialist. |  | 1167 | 1167 | Worldwide Distribution. | Terminated | January 31, 2012 | 536 | 1167
Z-0911-2011 | 56772 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Eclipse Client software versions with build numbers from 7.3.10 through 8.6.15 (i.e. excluding 8.6.17); Model Number: H48; Radiation Treatment Planning System |  |  | August 17, 0010 | 10 | Varian Medical Systems, Palo Alto, CA | Possible incorrect dose calculation: An anomaly in certain Eclipse versions where, under certain specific sequences of user actions, the dose distribution may fail to be re-calculated even though changes have been made which affect dose. | Software design | Varian Medical Systems Urgent Medical Device Correction letters were sent to all affected users on August 6, 2010. Letters were also sent to Varian Sales, Marketing and Service organizations to inform them of the recall. |  | 7019 | 7019 | Worldwide Distribution. | Terminated | July 21, 2011 | 730823 | 7019
Z-0919-2011 | 54229 | No | JFC | Intermittent pressure measurement system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Product labeling of the X3 Display reads in part: |  |  | July 20, 2009 | 2009 | Xsensor Technology Corporation | The plastic enclosures for the X3 Display and X3 PRO Pressure Mapping System has a UL94 HB flammability rating while the medical equipment specification UL 60601-1 requires V-2 rated plastic for electronic enclosures. | Component change control | On July 6, 2009, Xsensor sent an "URGENT: MEDICAL DEVICE RECALL" notifications to their consignees/customers.The consignees/customers were informed that the incorrect material of the plastic enclosures was used. The incorrect plastic had a UL94 HB flammability rating while the medical equipment specification UL 60601-1 required V-2 rated plastic for electronic enclosures. The firm provided prepaid shipping box to consignees/customers with instructions for returning the components. The components were refurbished and returned to consignees for free of charge.For further information please contact XSENSORs Customer Support at 1-866-927-5222. |  | 66 | 33 units of X3 Display and 2 units of X3 Pro | Devices were distributed to one distributor, medical facilities, and rehab centers through out the US, including following states:AL, AR, AZ, CA, CO, FL, GA, IA, IL, KS, ME, MI, MN, MO, NC, NY, OH, OK, OR, PA, SC, TN, TX, VA, VT, and WI. | Terminated | January 31, 2011 | 560 | 66
Z-0920-2011 | 57535 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ultra Surgical Gown, Large with Polypropylene Fabric and Towel, Reference number 95111-35, Single use only, Sterile, Made in Honduras. |  |  | December 23, 2010 | 2010 | Kimberly-Clark Corporation | Product was not properly sterilized. | Pending | Kimberly-Clark contacted all ten (10) Distributors with a Distributor Recall Letter (sent FedEx Priority overnight) on 12/23/2010, as well as via telephone. This letter instructs the Distributor to extend the Recall to end-user level by notifying each of their customers of the Recall via the provided Customer Recall Letter. Both the Distributor Letter and Customer Letter include Response Sheets that allows for both Distributor and Customer (end-user) to indicate the quantity of affected product they have in their inventory or indicate whether they do not have any affected product. The Distributor is instructed to immediately cease further shipment of the product and to physically quarantine the product, and fax the completed Response Sheet to our firm'sRecall Coordinator within 5 business days of receipt of the letter. The Customer is instructed to immediately cease further use of the gowns and physically quarantine the product, and fax the completed Customer Product Response Sheet to Kimberly Clark's Recall Coordinator within 5 business days of receipt of the letter. Upon receipt of the Response Sheets, the firm will provide instructions for the return of the impacted gowns for delivery to a quarantine location at our warehouse. |  | 200 | 200 cases (32 units per case) | Nationwide Distribution -- CA, CT, LA, NJ, NM, RI, CO and TX. | Terminated | July 27, 2011 | 216 | 240
Z-0917-2011 | 57068 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva with Digital Image (DI) Rel.2 system. |  |  | March 06, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | Images from two different patients may be combined in one patient folder in the Digital Image (DI) or may be combined as one (printed) image. It is possible that an image of a previous patient registered in the DI will be overwritten with an image of a later patient. | Other | On March 6, 2009, Philips sent an "URGENT-Device Correction Easy Diagnost Eleva, MultiDiagnost Eleva, and OmniDiagnost Eleva with DI Rel. 2 systems notification" to their consignees/customers.The consignees/customers are advised that before diagnosing images, always check the stored (Digital Image, Viewform or PACS) or printed images to ensure that these represent the data and images captured from the examination of the respective patient. Do not use images if they do not represent the patient that was examined. Consignees will be contacted by Philips to schedule a software upgrade of their units.The consignees can contact Philips Healthcare Call Center at 800-722-9377, #5, #2 and reference "FCO 70800070" with any questions about this recall. |  | 43 | 43 units | Nationwide distribution: USA including states of: AZ, CA, CO, CT, DE, FL, GA, HI, ID, LA, MA, ME, MI, MN, MO, MS, NC, NJ, NV, NY, OH, OR, PA, SC, TN, TX, VA, WA, WI, and WY. | Terminated | June 16, 2011 | 832 | 95
Z-0923-2011 | 50074 | No | MGA | Alpha-globulin immunological test system. | Immunology | Immunology | 510(K) Exempt | K-ASSAY Alpha-1 Microglobulin Reagent Kit. Kit contains: 1 vial Reagent 1, Buffer 50 mL, and 1 vial reagent 2, Antibody Reagent 50 mL. The product is packaged in a kit box along with with the package insert. |  |  | September 17, 2008 | 2008 | Kamiya Biomedical Company, LLC | Standardization of the reagent performance data did not match the standardization used to assign the calibrator values. | Employee error | Kamiya Biomedical Company, LLC sent an urgent medical Device Recall letter dated September 17, 2008, and the corrected package insert via e-mail. . The letter identified the product, the problem, and the action the customer should take. Customers were instructed to replace their previous package inserts with the corrected version and asked to sign and FAX the letter back once that had been completed. As one customer is also a distributor, the recalling firm told the this distributor's representative to advise their single customer of this correction and replace the package insert, and return the signed letter back to the firm. On September 22, 2008, both customers were contacted by phone and reminded to fax the signed letters once completed.For questions regarding this recall call 1-800-222-0342. |  | 5 | 5 kits | Kits were distributed to one laboratory in MN and to one distributor in PA who further distributed product to one customer. | Terminated | January 20, 2011 | 855 | 5
Z-0922-2011 | 56356 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Stryker brand Cart with Isolation Transformer, |  |  | April 16, 2010 | 2010 | Stryker Endoscopy | The electrical cord power plug design may allow the prongs to develop a fracture inside or outside the plug, if the plug is misused. | Device Design | Stryker Endoscopy sent an URGENT: Device Recall letter dated April 16, 2010, to Stryker Sales reps and all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Locate the affected units.Visually inspect each cord for damage.If the cord has bent or cracked prongs the unit should be taken out of service immediately. The customer was instructed to call a Regulatory Affairs Recall Representative at (408) 754-2738 for further instruction. If the cord did not show signs of damage, it could remain in service.An acknowledgement form was included to document confirmation of receipt of the letter. Upon completion of the power cord inspection, customers were to sign the acknowledgement form indicating that all of their units were inspected. Customers were asked to mail the form to Stryker Endoscopy or fax to (408) 754-8378 or scan the card and email to PowerCord@stryker.com.For questions regarding this recall call the Stryker RA/QA Department at (408) 754-2738 or (800) 435-0220, Reference RA 2010-034. |  | 7371 | 7371 | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, MA, ME, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, MY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and the countries of Canada, Chile, China, India, Korea, Mexico, Singapore, Taiwan and Latin America. | Terminated | January 19, 2011 | 278 | 7371
Z-0926-2011 | 57573 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DPOC-20, Direct Peroral Cholangioscopy Balloon, G54951, 20 mm/17 mL, Rx Only, Cook Medical, Winston-Salem, NC 27105 |  |  | December 23, 2010 | 2010 | Cook Endoscopy | There is a risk of gastrointestinal perforation or air embolism due to balloon deflation difficulty, resulting in poor patient outcomes. | Other | The firm, Cook Medical, sent an "URGENT: PRODUCT RECALL" letter dated January 7, 2011, to all customers. The letter described the product, problem and the actions to be taken by the customers. The customers were instructed to quarantine any affected product and return to Cook Medical. If they are distributors, they were instructed to contact their customers, direct them to quarantine any affected product and coordinate the return of the affected product to Cook Medical. The customers were also instructed to complete and return the bottom of the letter and send to Cook Medical as well as the PRODUCT RECALL EFFECTIVENESS SURVEY via fax at 812-339-7316 or scanned and emailed to CustomerRelationsNA@cookmedical.com. Note: All non-responding customers will receive a total of three (3) recall notifications ( two (2) follow up requests for a response in addition to the initial recall notification). Should you have any questions, please contact the Customer Relations Department at Cook Medical for more information. Please use CustomerRelationsNA@cookmedical.com or call toll free at 1-800-457-4500, press 4 then enter extension 2146. |  | 1 | see line 1. | Worldwide distribution: USA including states of: AL, CA, FL, IL, IN, LA, MA, NC, NY, OH, PA, SC, TN, TX, WA, and WI; and countries of: Belgium, Canada, Germany, Hong Kong, and the Netherlands.. | Terminated | August 04, 2011 | 224 | 1
Z-0928-2011 | 56931 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Scorpio Fixed Femoral A/P Sizing Guide; non-sterile instruments; Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; A subsidiary of Stryker Corporation |  |  | September 27, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that there is a potential for the shoulder screw on the Scorpio Sizing Guide (the "Sizing Guide") to disassociate from the main body of the instrument. | Device Design | Stryker Orthopaedics sent an 'Urgent Update Product Recall' letter dated December 30, 2010, along with Product Acknowledgement forms to all customers (US and international). The letter described the product, problem and action to be taken by the customers. The customers were asked to complete and return the PRODUCT RECALL ACKNOWLEDGMENT FORM within five (5) days via fax at 201-831-6069; contact the hospitals in their territory that have the affected product to arrange return of the product; and return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100 or the Associate Product Manager, Knee Marketing Team, at (201) 831-5028. |  | 6 | six units | Worldwide distribution: USA and countries including: Australia, Brazil, China, Columbia, France, Greece, India, Italy, Korea, Poland, Singapore, South Africa, and United Kingdom. | Terminated | July 31, 2013 | 1038 | 243
Z-0924-2011 | 56859 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mizuho ProFx 6850 Operating Room Table, with orthopedic accessories, AC Powered, manufactured by Mizuho OSI, Union City, CA. |  |  | March 30, 2010 | 2010 | Mizuho Orthopedic Systems Inc | Device component spar gear box locks up and may result in failure of push pull tube, which may result in release of breaking and impact. | Device Design | All customers will be contacted to arrange for removal and shipping arrangements. Verification will be done by the Repair and Verification instructions for the leg spars, via registered letters.*******UPDATED INFO---The firm sent another notification letter---dated 12/22/10---which states in part: "****The upgrade involves the correction of your 6850 PROfxÂż spars (6850-2100) at your facility by a Qualified Technician. The spars shall be upgraded to remove the 6850-2123, Drive Bracket, Spar Push Rod and replace it with the 6850-2170, Drive Bracket, Spar Push Rod. Upon completion of the upgrade, the 6850-2123, Drive Bracket, Spar Push rod shall be returned to Mizuho OSI for processing. ******* |  | 17 | 17 units | Worldwide Distribution -- USA &amp; Germany. | Terminated | May 17, 2011 | 413 | 17
Z-0929-2011 | 56403 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT CG8+ Cartridge; |  |  | August 03, 2010 | 2010 | Abbott Point of Care Inc. | Two cartridge lots were identified that may generate falsely elevated PCO2 results. | Other | Abbott Point of Care Inc. sent an Urgent Recall Notice letter dated July 2010 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the cartridges from the identified lots and follow the instructions on the enclosed reply card to receive replacement cartridges. If cartridges were distributed from the lot numbers to another institution, a copy of the recall letter should be provided to them.For any questions regarding this recall contact Abbott-Point of Care Technical Support at 800-366-8020, Option 1. |  | 55350 | 55350 total cartridges | Worldwide Distribution - USA including AL, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV, and WY and the countries of Oman and Turkey. | Terminated | December 05, 2011 | 489 | 55350
Z-0933-2011 | 57334 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter FloGard 6201 Volumetric Infusion Pump. Baxter Healthcare Corporation, Medication Delivery Division. Product code 2M8063 |  |  | November 17, 2010 | 2010 | Baxter Healthcare Corp. | The FloGard Infusion Pump was released from service without having the Current Leakage test performed. This could lead to the patient/user to receive an electrical shock. | Process control | The firm, Baxter, telephoned the one customer on November 17, 2010, informing them that electrical safety testing was not performed on their FloGard pump serial number 11091330FA during the last service. Since the test was not conducted on the pump, the device may not function as designed and could lead to a hazard of electrical shock. Baxter requested that they locate the pump and remove it from use as soon as possible. Baxter advised the customer that a Baxter Field Service Engineer would be sent to their facility to perform electrical safety testing or alternatively, the customer may choose to send the pump to Baxter for testing.If you have any questions regarding this information, please contact Vice President, Quality, GIS at 847-270-5507. |  | 1 | 1 pump | Nationwide distribution: WA only | Terminated | March 16, 2011 | 119 | 1
Z-0936-2011 | 57572 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Super Arrow-Flex(R) PSI Set |  |  | December 20, 2010 | 2010 | Arrow International Inc | Packaging may have been compromised, and sterility of product cannot be guaranteed. | Other | The firm, Teleflex Medical thru its subsidiary Arrow International, Inc., sent an "Urgent Field Safety Notice" letter dated December 20, 2010, with Field Acknowledgement form to all customers. The letter described the product, problem and action to be taken by the customers. The letter instructed the customers to check their stock for affected product, cease use and distribution of product and quarantine all products from affected lot immediately; contact their local Teleflex Medical Customer Service/Sales Rep (or Distributor) for a return authorization number prior to filling out the FIELD SAFETY CORRECTIVE ACTION ACKNOWLEDGEMENT AND STOCK STATUS FORM; and complete and return the form immediately by email. If you do not have the product from the affected lot, indicate it on the form and fax it to your local Teleflex Medical Customer Service/Sales Rep (or Distributor). The Teleflex Medical Customer Service/Sales Rep (or Distributor) will coordinate the product return with the customer.If you have any questions, please contact the Regulatory Affairs Project Manager at 610-378-0131 ext 3318 or email: Jody.Cadd@teleflexmedical.com. |  | 320 | 320 units | International distribution only: countries include: Austria, Belgium, Czech Republic, Germany, Greece, Italy, and Luxembourg. | Terminated | September 25, 2014 | 1375 | 320
Z-0931-2011 | 57515 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Single Channel Volumetric Infusion Pump. Baxter Healthcare Corporation, Medication Delivery Division. Product Code: 2M8161 |  |  | December 01, 2010 | 2010 | Baxter Healthcare Corp. | While the infusion pump was being serviced at the Baxter Service Center, it was inadvertently routed to the packaging area before the final inspection. | Employee error | Baxter telephoned the one customer on December 1, 2010, informing them that that an outgoing Quality Control Evaluation was not performed on their Colleague pump serial number 15030799CC during the last service. Since this was not conducted on this pump, the device may not function as designed and could lead to an interruption of therapy. Baxter requested that they locate the pump and remove it from use as soon as possible. Baxter will then send a shipping container for retrieval, reinspection and return of the device back to the customer at no cost.For questions regarding this recall call (800) 437-5176. |  | 1 | 1 pump | Texas | Terminated | May 13, 2011 | 163 | 1
Z-0932-2011 | 57527 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console (4DITC); |  |  | December 14, 2010 | 2010 | Varian Medical Systems, Inc. Oncology Systems | Multiple isocenter plans will load on the 4DITC and are available for imaging and treatment, but the 4DITC software does not recognize that mulitplex isocenter is present. | Software design | Urgent Medical Device Correction/Field Safety notification letters were distributed to all affected users, and to Varian Sales, Marketing and Service Organizations on December 14, 2010 by certified mail. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letters described the issue and recommended user actions along with Varian's actions. Users are to use caution when acquiring actual couch positions for any plan. In addition, each plan should be reviewed for quality assurance purposes in accordance to set departmental or organizational policies. The firm updated the procedure and enclosed a copy with the letter. It was also posted to the customer support site. If further clarification is needed, customers are to contact their local Varian Customer Support District or Regional Manager. |  | 4188 | 4188 | Worldwide Distribution. | Terminated | July 22, 2011 | 220 | 4188
Z-0934-2011 | 57612 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Lubricating Jelly, Sterile, Non-staining, Made in the USA, packaged into the following size containers 5 gram tube, 2 ounce and 4 ounce tubes, 4 ounce bottle, 3 and 5 gram packets |  |  | December 23, 2010 | 2010 | H &amp; P Industries, Inc. dba Triad Group | Triad Group is recalling all lots of Sterile Lubricating Jelly manufactured by Triad Group. This recall has been initiated due to concerns expressed by the Food and Drug Administration regarding the validation of the gamma radiation sterilization cycles for these products. | Process control | The firm, Triad Group, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 22, 2010, to all customers via US Postal Service certified. The letter described the issue, identified affected product as that which contained a lot number beginning with the digits 7, 8, 9, or 0, and the action to be taken by the customers. The customers were instructed to immediately examine their inventory and quarantine the product subject to recall; notify their customers at once if the recalled product was further distributed, and complete and return the enclosed Recall Acknowledgement form via fax at 262-538-2947 or mail along with the recalled product to: Recall Coordinator, Triad Group, 700 West North Shore Drive, Hartland, WI 53029.If you have any questions, please call Triad Group Customer Service Monday thru Friday, between the hours of 8:30 a.m. and 4:00 p.m. Central Time at 262-538-2900 ext 2761. |  | 453787 | 453,787 cases | Worldwide distribution: USA including states of: CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MA, MI, MO, MS, MY, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI and Puerto Rico; and countries including: Australia, Belgium, Canada, China, Lebanon, New Zealand, Philippines, Northern Marianna Islands, Saudi Arabia, Singapore, Taiwan, and United Arab Emirates. | Terminated | June 11, 2012 | 536 | 453787
Z-0935-2011 | 57516 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX and CXE Infusion Pumps, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2M8161 and 2M9161. |  |  | December 01, 2010 | 2010 | Baxter Healthcare Corp. | The Air In Line (AIL) test was not properly performed by the service technician at the Service Center. | Employee error | Baxter telephoned the two customers on December 1, 2010, informing them that that a post repair test was not performed correctly on their Colleague pumps (serial numbers 12070599CC and 17086323CP) during the last service. As a result, the device Air Sensor may not function as designed and could lead to undetected air in the IV line. Baxter wants to ensure that the devices are properly tested, safe and effective. Baxter requested that they locate the pumps and remove them from use as soon as possible. Baxter will then send a shipping container for retrieval, reinspection and return of the device back to the customer at no cost.*** Baxter expanded the recall to include twelve additional pumps on June 22, 2011 by telephone calls to the affected accounts. The recall strategy remains the same as the December 1, 2010 instructions. **** |  | 14 | 14 pumps | Nationwide Distribution -- Arizona, Indiana, *** California, Iowa, New Jersey, Ohio, Virginia and West Virginia***. | Terminated | January 20, 2012 | 415 | 14
Z-0938-2011 | 50791 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Crosser 14S Catheter, Catalog/REF CRU14S, Lot number 5053 manufactured by Flowcardia Inc, Sunnyvale, CA |  |  | January 09, 2009 | 2009 | Flowcardia Inc. | Mislabeled expiration date: Some units were mislabeled with a three year shelf life, but the product is validated only for a two year shelf life. | Employee error | The firm, FlowCardia, Inc., sent an "URGENT! Medical Device Recall" letter dated January 9, 2009 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to return the product to FlowCardia, at the firms expense, for replacement and Confirm that this product has been used by completing and returning the attached fax document via fax to 408-616-0053. The customers were ask to please make an attempt to contact the customer service representatives, Operations specialist at 888-456-7286 or Vice President, Clinical, Quality &amp; Regulatory Affairs at 408-617-0352 x302, within the next 30 days. Note: FlowCardia Sales, Clinical and Customer support will send notification via certified mail or traceable courier and track response.If you have any questions or need any additional information email (dmichaels@flowcardia.com) or telephone at 408-617-0352 x302. Alternatively you may contact customer service at 888-456-7286. |  | 35 | 35 units | Nationwide distribution: CA, NJ, LA, TX, GA, MI and CO. | Terminated | January 25, 2011 | 746 | 35
Z-0937-2011 | 56804 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | Welch Allyn PIC 50 Portable Intensive Care System. Model 971081 series. |  |  | September 17, 2010 | 2010 | Welch Allyn Protocol, Inc | Potential likelihood that the Portable Intensive Care (PIC) 50 Defibrillator's screen display may fail. The loss of the ability to view the patient's Electrocardiogram (ECG) on the screen could cause the user to delay or fail to administer defibrillating shocks needed to resuscitate a patient. | Other | On 9/17/10, Welch Allyn began sending out the Important Product Alert letter to all their direct consignees. On 9/24/2010, the firm sent out the letters to their distributors.The Product Alert letter informs the consignees of a low potential likelihood that the Portable Intensive Care (PIC) 50 Defibrillator's screen display may fail.The consignees are advised to follow the instruction in the PIC 50 user's manual requiring routine shift equipment checks daily or at each shift change to minimize the likelihood of the malfunction from occurring. The checks ensure that the PIC 50 and its component (including the screen display) are fully operation and undamaged.The letter further states "our analysis indicates that routinely doing such checks will likely lead to discovery of this malfunction as it occurs, and thus timely maintenance and repair".The distributors are advised to communicate this Product Alert to their customers if the devices were further distributed. For customers in US and Canada, distributors are advised to provide the electronic Excel file with customer's name, address, serial number and part number to WA2221@Stericycle.com. For customers in other countries, they are advised to provide information to PIC50Product Alert@welchallyn.comConsignees can call Welch Allyn at 800-289-2500 for any questions during normal business hours. |  | 3606 | 2002 units to US consignees and 1604 units to foreign consignees | Worldwide Distribution -- US, including states: AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and countries ofAustralia, Austria, Belgium, Canada, China, Columbia, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Jordan, Kuwait, Kyrgyzstan, Latvia, Lebanon, Liechtenstein, Lithuania, Malaysia, Mexico, Oman, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Switzerland, Syrian Arab Republic, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam, and Yemen. | Terminated | July 25, 2011 | 311 | 3606
Z-0949-2011 | 57470 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | The Stryker InterPulse system is a single patient use, sterile, disposable, pulsed lavage device. The system consists of an irrigation handpiece, irrigation tubing, irrigation tips, suction tips, suction tubing, and splash shields. |  |  | October 04, 2010 | 2010 | Stryker Instruments Division of Stryker Corporation | There is potential for a breach in sterility in the blister packaging. This can lead to the use of non-sterile product which, if contaminated, could lead to a patient infection. | Packaging process control | Stryker sent a Medical Device Recall Notification letter dated October 8, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to destroy and dispose of all affected product and to complete the business reply form and return to Stryker to receive replacement interpulse products at no charge.For questions regarding this recall call (269) 323-7700. |  | 954 | 954 | Worldwide Distribution - USA (including all 50 states) and the countries of Brazil, Canada, France, Germany, Japan, Netherlands, South Africa, South Korea, Spain, Sweden, Switzerland and the UK. | Terminated | June 06, 2011 | 245 | 3187
Z-0955-2011 | 57380 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Newdeal (Newdeal Compression Plate) Model #UNI-CP 339900 Sets, Lot Numbers: E71H, E7V1, E86R, EELE, EGTC |  |  | November 18, 2010 | 2010 | Integra LifeSciences Corp. | Integra LifeSciences Corporation has decided to initiate a voluntary recall of the Uni-Comp Kits due to inaccurate information relative to the recommended sterilization parameters described in the products Instructions for Use (IFU). The sterilization parameters noted in the Uni-Comp IFU were incorrect. It is noted that the sterilization recommendations in the UNI-Comp IFU in which the Exposure | Other | The firm, Integra, contacted the Sales Representatives on November 8, 2010 via telephone conference calls. Integra also sent a "URGENT: PRODUCT RECALL NOTIFICATION" letter dated November 9, 2010 all customers via FedEx overnight delivery. The letter described the product, problem and the action to be taken by the customers and the firm. The customers were instructed to immediately examine their inventory, QUARANTINE THE PRODUCT, and complete and return the RECALL ACKNOWLEDGMENT AND RETURN FORM via mail: Integra LifeSciences Corporation, Attn: Sandy Yates, 311 Enterprise Drive, Plainsboro, NJ 08536, fax to (609)-275-5363 or Scan and email to sandra.yates@integralife.com. The firm will have a Integra LifeSciences Corporation Sales Representative contact the customer to assist them in retrieving and returning their Uni-CP Compression Plate System to Integra. Integra will also ship replacement Uni-CP Compression Plate (Uni-CPII and Large Uni-Clip) Systems back to the customers as soon as possible at no charge.Should you have any additional questions, contact the Director of Marketing, Hindfoot-Extremity Reconstruction or your regional managers at 513-533-7929, 609-275-0500 or toll free at 800-654-2873. |  | 67 | Sixty seven (67) Kits | Nationwide distribution: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA and WI. | Terminated | November 18, 2011 | 365 | 67
Z-0966-2011 | 57708 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Urological, Inc. Koyle Diaper Stent 6.0 Fr, Sterile, GPN G17667, REF KDSCF-060050. |  |  | December 13, 2010 | 2010 | Vance Products Inc. dba Cook Urological Inc. &amp; Cook OB/GYN | The firm received a complaint that 15 units of the Koyle Diaper Stent were received in packages that had not been sealed. Each stent should be sealed in a poly-Tyvek pouch. | Packaging process control | The firm sent an URGENT RECALL NOTICE to one consignee on 12/13/2010 (the other consignee returned all their product as a complaint). The letter informed the consignee of the recall and asked the customer to return the recalled product. |  | 25 | 25 stents | CA, CT | Terminated | January 31, 2011 | 49 | 25
Z-0954-2011 | 57498 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien, LigaSure, Blunt Tip Laparoscopic Sealer/Divider, 5 mm-37 cm, REF LF1537, Sterile EO, Covidien, Valleylab, a division of Tyco Healthcare Group LP, 5920 Longbow Drive, Boulder, CO 80301 |  |  | December 08, 2010 | 2010 | Covidien LP | During surgical procedures, the device may energize when clamping without pushing the activation button, which could result in unintended energy delivery to tissue. | Device Design | The firm, Covidien, sent a "URGENT: PRODUCT RECALL" letter dated December 6, 2010, to the customers. The letter described the product, problem and the action to be taken by the customers. The customers were instructed to remove product from inventory, immediately quarantine the product and return any remaining stock of the lots listed to: Covidien Energy-based devices, Attention: R82, Building 4/Receiving, 5920 Longbow Drive, Boulder, CO 80301. The customers were also instructed to complete and return the Customer Response Form immediately via mail (in the envelope provided) or fax to (303) 516-8307.To return any identified product for credit or replacement, contact Customer Service at 1-800-722-8772 Option 1.Should you have any medical questions concerning this recall, please contact your local sales representative or the Covidien Clinical information Hotline at (800) 255-8522 (option 1). |  | 25009 | 25,009 units | Worldwide distribution: USA and countries including: Australia, Canada, France, Germany, Japan, Puerto Rico, Singapore, and Uruguay. | Terminated | March 27, 2012 | 475 | 25009
Z-0939-2011 | 57630 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Linemaster Wireless Footswitch, VASER System, REF Vaser 2.0, Sound Surgical Technologies LLC, 357 South McCaslin Blvd., Suite 100, Louisville, CO 80027 |  |  | December 21, 2010 | 2010 | Sound Surgical Technologies, LLC | Surgical equipment, activated by a wireless foot switch, may remain powered on when switch is no longer depressed. | Process control | Customers were notified by e-mail with an Advisory Notice on 12/21/2010. The Notice identified the affected product and described the issue. The Notice also explained what the potential hazard involved along with a mechanism to eliminate the potential hazard. Sound Surgical would arrange a service call to implement the correction at the firm's expense. If customers experience this potential failure, they are to contact Sound Surgical Technologies at 1-888-677-8765 or 1-888-471-4777. |  | 27 | 27 units | Nationwide Distribution -- AL, AR, AZ, CA, CO, DC, FL, GA, IA, IL, IN, MA, MI, MO, NJ, NY, PA, and TX . | Terminated | February 08, 2011 | 49 | 27
Z-0967-2011 | 57188 | [6 empty] | StaXxÂż XD Expandable Cartridge Convex Expandable Implant 25mm (L) x 9mm (W) x 7mm (H) |  |  | October 22, 2010 | 2010 | Spine Wave Inc | Package Insert Incorrect | Packaging process control | Spine Wave notified each affected customer on October 22,2010, via telephone of the product, the issue and action they should take. Customers were asked to segregate the product from other products and immediately return to Spine Wave.For questions regarding this recall call 203-944-9494. |  | 12 | 12 units | Nationwide Distribution including Delaware, Massachusetts, Maryland, and New York | Terminated | March 15, 2011 | 144 | 12
Z-0963-2011 | 57609 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | SRD S5 Mast Roller Pump System 85, S5 Component, Part No.: 50-80-70, Made in Germany, Distributed by: Sorin Group USA, Inc., Arvada, CO 80004. |  |  | November 12, 2010 | 2010 | Sorin Group USA, Inc. | Touch screen may become unresponsive, inhibiting user input. | Device Design | All affected US customers and distributors were notified by certified mail on Nov 12, 2010 via a Field Correction letter. They were told that they would be contacted by Sorin to arrange an appointment to inspect and replace affected touch screens on site. They were told that they could safely continue using their S5 systems in accordance with the Operator's Manual and previously supplied instructions until the inspection and necessary replacement had been completed. A customer response form was also included for completion and return to the firm. Customers can contact the firm if they have any questions. |  | 16 | 16 units | Worldwide Distribution. | Terminated | July 13, 2012 | 609 | 6671
Z-0968-2011 | 57231 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sternal Saw II Sterile Blades for the Sarns Sternal Saw System, Catalog number 5589, Terumo Cardiovascular System Corp, Ann Arbor, MI. |  |  | October 20, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The firm discovered that some outer box labels contain incorrect information regarding the method of sterilization. The label incorrectly indicates the product was ethylene oxide sterlized but the product actually was sterilzed with gamma radiation. | Labeling design | Terumo Cardiovascular Systems Corporation firm sent out an 'URGENT MEDICAL DEVICE CORRECTION: SAFETY ADVISORY" letter dated October 20, 2010, to the Director of Cardiac Operating Room Services or Operating Room Nurse Manager.. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the Safety Advisory, assure that all users were aware of the notice, and confirm receipt of the Safety Advisory by faxing the attached Customer Response Form to the fax number on the form. For questions regarding this recall call 1-800-521-2818. |  | 181902250 | 42,650 units, 4265 boxes of 10 each | US, Indonesia, UAE, Hong Kong, Japan, Singapore, Malaysia, Taiwan, Belgium, Uraguay, Venezuela, Columbia, Mexico, Guam, Vietnam and Thailand. | Terminated | December 14, 2011 | 420 | 181902250
Z-0986-2011 | 54350 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | KING LT -D Laryngeal Tube, Size 4, KLTD204, King Systems Corporation 15011 Herriman Boulevard, Noblesville, IN 46060. |  |  | December 15, 2009 | 2009 | King Systems Corp. | King Systems was issued a warning letter from the Center for Devices and Radiological Health, Office of Compliance on October 28, 2009. The warning letter asserts that Kings Systems KLT(S)-D airway exceeds cleared indications in that it is promoted for "airway management in patients over 4 feet in height, for controlled or spontaneous ventilation, difficult and emergent airway cases, and ambulat | Labeling False and Misleading | On October 29, 2009, the firm, Kings System, sent a "Urgent - Product Correction" letter to the customers. The letter described the product, problem and actions to be taken. The customers were instructed to remove the promotional material from the King System website which includes literature, and Instructions for Use. The firm instructed all Sales force and distributors to cease dissemination of promotional materials, literature and communications.Note: The firm revised indications, instructions for use, product labels and labeling so it is aligned with the 510Ks. The customers were also instructed to replace current instructions for use with revised instructions for use and destroy the current IFU; return Kings LT(S)-D oropharyngeal airway product inventory to King Systems; and complete and return the attached Instructions for Use Destruction Signature Form via fax to King Systems at 317-776-5197.For further information, please call 317-776-6823 ext 257 or email: LTSD@kingsystem.com. |  |  |  | Worldwide distribution: USA and countries including: Canada, Costa Rica, and Panama. | Terminated | August 16, 2012 | 975
Z-0973-2011 | 53198 | No | LPB |  |  | Cardiovascular | PMA | RF Ablation System Foot Switch, used with the EPT-1000 RF Generator, Foot Switch Model Number UPN M0048400, and in kits with UPNs M0044810 and M0044811. |  |  | July 14, 2009 | 2009 | Boston Scientific Corp | Product malfunction causes inadvertent delivery of energy. | Nonconforming Material/Component | The firm, Boston Scientific, sent an "URGENT: Medical Device Recall" letter dated January 14, 2009, with attached recall instructions to all customers. The letter described the product, problem, and the action to be taken by the customers. The customers were instructed to complete and return the Reply Verification Tracking Form and Obtain an RGA via fax to 408-957-6204; and return the recalled product using BSC Federal Express Number 1-809-4713-2 via second day delivery to Boston Scientific Corporation, Distribution Center, 500 Commander Shea Blvd., Quincy, MA 02171. Boston Scientific Corporation sales representatives will contact each customer.If you have any questions, please contact 408-935-4952 or email: wareb@bsci.com. |  | 1044 | 1044 units | Worldwide distribution | Terminated | January 27, 2011 | 562 | 1044
Z-0987-2011 | 57501 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Physiological Monitoring, Patient Monitor |  |  | December 17, 2010 | 2010 | Philips Healthcare Inc. | Some PAS-201 active speaker assemblies in use with the Philips IntelliVue Information Center (PIIC) have experienced intermittent audio or loss of audio. Speaker failures may delay recognition of an alarm condition, if operators do not recognize that there is no alarm sound from the monitor and the monitor is not connected to another audible alarm annunciation system. | Device Design | Philips notified all affected customers with an Urgent Medical Device Correction notification letter dated December 17, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to observe their system for any of the described symptoms. If they experience intermittent audio or there is no sound from their PAS-210 speaker while it is not connected to another audible alarm annunciation system, customers were instructed to remove the speaker from use and contact their local Philips service representative. Customers were advised not to rely exclusively on the audible alarm system for patient monitoring, the most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment.ACTIONS PLANNED BY PHILIPSPhilips will replace affected speaker assemblies free of charge.A Philips Healthcare representative would contact the customer when the replacement was available. Field Service Engineers will remove and replace defective speakers to correct the problem. The defective speakers will be returned to vendor (Beach Wire) to be discarded. This correction will be done free of charge for customers. |  | 1071 | 1071 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, MS, MT, MS, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA WI, WV, WY, and Canada | Terminated | April 18, 2017 | 2314 | 1071
Z-0970-2011 | 57703 | [6 empty] | RLSense, Restless Legs Sensor, Proto6, Rev-J Firmware |  |  | September 15, 2010 | 2010 | Sleep Modus | Device had no 510(k) clearance prior to patient use. | Other | Sleep Modus, Inc. notified consignees by e-mail on 9/15/2010 and 10/7/2010. The e-mail identified the product, the problem, and the action to be taken. Customers were instructed to let them know if they experienced any of the issues identified in the e-mail. For questions call 214-405-4055. |  | 110 | 110 devices (2 per kit) | Nationwide Distribution including Texas | Terminated | January 31, 2011 | 138 | 110
Z-0972-2011 | 57051 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Bio Med Devices Disposable Pediatric Breathing Circuit, One Set |  |  | October 19, 2010 | 2010 | Biomed Devices | Breathing circuit may disconnect between the adapter adjacent tubing cuff | Component change control | The firm,Bio Med Devices, sent a "URGENT MEDICAL DEVICE RECALL" letter dated October 18, 2010 via fax to the customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to immediately examine their stock for the affected lots, if found, discontinue distributing and using the product and promptly return via UPS #025106 to Bio-Med Devices 61 Soundview Rd, Guilford, CT 06437, Attention: 4040LC Recall Return. The customers were also instructed to immediately contact their accounts, if the product was further distributed , advise them of the recall situation and have their outstanding product returned, and to complete and return the RETURN FAX-IMMEDIATE RESPONSE REQUIRED form via fax to 203-458-0440. Note: Bio-Med informed the customers that they will credit the return product; however, they only plan to replace the product in an emergency situation, because they have very few of the 3/8" x 15mm adapters left and are looking to discontinue the 4040LC due to the discontinuance of this product by the manufacturer.If you have any questions, please contact (203) 458-0202 x 217. |  | 15 | 15 cases (20/cs) | Nationwide distribution: OH, WI. | Terminated | October 23, 2014 | 1465 | 16
Z-0991-2011 | 57522 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | EndoPlege Sinus Catheter Kit, REF EP, 9 Fr x 18.9", Contains Phthalates, Sterile EO, Sterile R, Edwards Lifesciences LLC, Irvine, CA 92614. |  |  | December 03, 2010 | 2010 | Edwards Lifesciences Llc | Defective product may result in the silicone valve becoming dislodged from the introducer which can result in injury. | Device Design | The firm, Edwards Lifesciences, sent an "URGENT PRODUCT RECALL - ACTION REQUIRED" letter dated December 3, 2010 to customers. The letter describes the product, problem and actions to be taken by the customers. The customers were instructed to check their inventory and return the unused devices; complete and return the INTROCSC Introducer Recall Response form to the attention of Customer Service, Edwards LifeSciences, One Edwards Way, Irvine, CA at 949-250-3489. If you have any questions, please contact Edwards Customer Service at 800-424-3278 from 8:00am - 4:30pm Pacific Time. |  | 3427 | 3,427 kits | Nationwide Distribution | Terminated | October 27, 2011 | 328 | 3517
Z-0941-2011 | 57540 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Prelude Short Sheath Introducer, 7F-SMT, 4 cm, REF No: PSS-7F-4-038MT, Sterile EO, Merit Medical, South Jordan, Utah. |  |  | December 17, 2010 | 2010 | Merit Medical Systems, Inc. | Introducer tips may detach and embolize during use, or the tip material may elicit a thrombogenic response. | Component change control | Merit Medical Systems, Inc., issued an "Urgent Product Recall Notice" dated December 17, 2010, and contacted all customers (sales reps and authorized distributors) via email, phone and/or fax. The notice described the product, problem and action to be taken by the customers (Sales reps/Distributors). The customers were instructed to contact all their customers to identify, quarantine and return all unused inventory and extend the recall to any of their customers who may have received the affected products. The customers were instructed to complete and return the Manufacturer RECALL NOTIFICATION and Product Retrieval Form via fax to 1-801-208-3378 and send original to: the Manager, Customer Service, Merit Medical Systems, Inc., 1600 West Merit Parkway, South Jordan, UT 84095-2416.Questions should be directed to the Senior Product Manager at 801-208-4344 or the Manager, Customer Service at 801-208-4365. |  | 378 | 378 units | Nationwide Distribution | Terminated | July 11, 2011 | 206 | 378
Z-1001-2011 | 56354 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand C Series ClinicÂżs (Includes Trilogy, and Novalis Tx), Model Numbers: H14, H26, H27, H29, Distributed by and/or Manufactured by: Varian Medical Systems Inc., Palo Alto, CA |  |  | December 01, 2009 | 2009 | Varian Medical Systems Oncology Systems | A circuit design anomaly whereby a faulty balance potentiometer can lead to an undetected electron beam asymmetry. | Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 1, 2009 via Federal Express to consignees/customers. The letter described the product, problem and actions to be taken by the customers and the firm. The customers were instructed to perform a verification of field symmetry on electron beam energies to determine the threshold and report any field symmetry issues to their local Varian service office for attention. The firm, Varian, is developing new functionality to detect the symptoms of this component failure condition under all conditions of electron beam operations and asserting a beam interlock if it is detected. Note: Varian indicated until the solution can be developed, they will conduct an audit of installed Clinacs to identify those the suspect "blue" electron beam balance potentiometers installed, and will execute a program of prophylactic replacement of these potentiometers.If you require further clarification, please contact your local Varian Customer Support District or Regional Manager at phone: USA and Canada-1-888-VARIANS (888-827-4265); Europe- +41 41 749-8844 or email support-emea@varian.com; North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: seasia.apps.helpdes@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America: soporte.al@varian.com; Internet: Oncology Systems - www.myvarian.com; Varian Medical Systems -www.varian.com |  | 4411 | 4411 | USA and countries including: Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Belorussia, Bolivia, Brazil, Canada, Chile,, China, Columbia, Costa Rica, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France,French Polynesia, Germany, Guadeloupe, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Kazakhstan, Kuwait, Latvia, Lebanan, Luxembourg, Maca, Macedonia, Malaysia, Martinique, Mexico, Moldova, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan,, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea,, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Uruguay.. | Terminated | March 14, 2012 | 834 | 4411
Z-0988-2011 | 56963 | No | GDX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Southmedic/Personna Safety Scalpel with Personna Plus Blades, Reorder No. 73-0610, Size #10, General Surgical Use, Sterile, unless package is damaged or opened., Southmedic Inc., 50 Alliance Blvd., Barrie, ON., Canada L4M 5K3 www.southmedic.com, Made in Canada. |  |  | September 03, 2010 | 2010 | Southmedic, Inc. | Scalpels labeled as sterile have the potential to be unsealed and therefore not sterile. | Packaging process control | On 9/3/2010, the firm notified consignees about the recall by facsimile letter. The letter instructed consignees to return scalpels for credit or replacement. The firm enlarged the scope of the recall on 11/15/2010. An additional 927 cases were recalled. This includes 11 additional lots. The firm sent out URGENT MEDICAL DEVICE RECALL letters dated 11/15/2010 to its direct accounts. The letter identified that affected product and the reason for the recall. The consignees were instructed to check their inventory, complete and return the attached Voluntary Recall Notification Confirmation Form, return the affected product to the manufacturer. Product return question should be directed towards Marc Pauze at 800-463-7146 ext 299 and regulatory questions should be directed towards Tish Anger at 800-463-7146 ext 342. |  | 1330 | 1330 all products | Worldwide Distribution -- USA, including states of TX, MO, NJ, LA, IL, NY, KY, PA, VA, WA, CA, MA, MD, IN, NV, TN, NC, AZ, IA, WI, UT, NE, GA, FL, MN, ME, OH, SD, and KS and countries of Australia and Canada. | Terminated | May 18, 2011 | 257 | 1330
Z-1005-2011 | 53780 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | 10032746 Acuson Antares PE Product Version 5.0, manufactured by Siemens Medical Solutions, Mountain View, CA. |  |  | October 02, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | ECG waveform does not accurately align with 2D, M-Mode or Spectral Images. | Nonconforming Material/Component | The firm's Customer Safety Advisory notice was mailed to all affected customers. The letter describes when the malfunction occurs and the potential risk. It also recommended that users collect corresponding evidence with spectral waveforms that would confirm or deny findings based on ECG, in order to avoid potential risks for misdiagnosis. The firm planned to exchange the faulty hardware component in the system. Customers would be contacted to schedule for the hardware exchange. Questions regarding this issue should be directed to the local service support person. |  | 67 | 67 units total | Worldwide Distribution | Terminated | March 29, 2011 | 543 | 67
Z-1008-2011 | 54535 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Treatment Planning System, a subsystem of the CyberKnife Robotic Radiosurgery System, with MultiPlan Treatment Planning System Software version 3.5 |  |  | December 29, 2009 | 2009 | Accuray Inc | If electron density values are left empty, the calculation of radiation dose in a patient will be modeled as air-like density material rather than correct density. A plan may be created and saved, thus creating the risk of mistreatment. | Employee error | The firm, Accuray, sent an "URGENT DEVICE CORRECTION" letter dated December 29, 2009, to all customers. The letter described the product, problem, and action to be taken by the customers. The customers were instructed to verify correctness of the CT number to relative electron density and Mass Density Calibration curves. Note: A patch is being developed to address this issue. Accuray Customer Support will contact the customers regarding further resolution.If you have any questions or concerns regarding this issue, please contact Accuray Customer Support at +1-877-668-8667 (USA) or +1-408-716-4700 (non USA) or by email: customersupport@accuray.com. |  | 71 | 71 units are subject to this action | Worldwide Distribution: USA and countries including: Canada, Turkey, India, Italy, France, the Netherlands, the UK, India, Spain, Greece, Saudi Arabia, Switzerland, Germany and China. | Terminated | February 11, 2011 | 409 | 71
Z-1002-2011 | 56357 | No | FAT | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Stryker Endoscopy, Infravision Esophageal Kit, |  |  | May 11, 2010 | 2010 | Stryker Endoscopy | The silicone sheath component of the product detaches from the catheter upon removal of the e-kit from the body. | Device Design | On 5/11/10, Stryker Endoscopy initiated the recall notification letters explaining the reason for the recall and requesting the product be returned. |  | 4125 | 4125 | Nationwide and 3 foreign consignees (Canada, Singapore and Switzerland). | Terminated | February 01, 2011 | 266 | 4125
Z-1003-2011 | 56349 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand C-Series Clinacs and Acuity, Exact Couch Top, Model Numbers: |  |  | April 19, 2010 | 2010 | Varian Medical Systems Oncology Systems | The product has been shipped since 2005 with the wrong label regarding max weight to be used on accessory rails. | Labeling mix-ups | Varian issued via customer technical bulletin on April 19, 2010, to Varian Sales, Marketing and Service organization. Notification. The bulletin identified the product, the problem, and the action to be taken by the customer. Customers were asked to fill out the attached form and send it to their local Varial Medical Systems Field Service Office.For questions regarding this recall call 650-483-3153. |  | 13 | 13 | Worldwide Distribution - Japan only | Terminated | December 23, 2011 | 613 | 13
Z-0940-2011 | 56868 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Maquet Heartstring II Proximal Seal System (HS-1045) manufactured by Maquet Cardiovascular, Wayne, NJ |  |  | May 05, 2010 | 2010 | Maquet Cardiovascular, LLC | Deployment tube may detach during use due to insufficient adhesive, which may result in product failure to operate. | Process design | The firm, MAQUET Getinge Group, sent an "URGENT DEVICE REMOVAL IMMEDIATE ACTION REQUIRED" letter dated May 5, 2010, to all customers via Federal Express. The letter described the product, problem and actions to be taken by the customer. The customers were instructed to immediately examine their stocks to determine if they have any of the affected product, if found, discontinue dispensing (distributing) the lot and complete the Field Action Response Form; immediately contact accounts if they acted as a distributor, and acknowledge receipt of this notification whether or not they have any affected products by completing the Field Action Response Form via fax to (U.S.): 1-888-899-2874.If you have any additional questions, please contact your local MAQUET Cardiovascular sales representative or our Customer Service at 1-888-880-2874. |  | 1630 | 1630 | Worldwide distribution: USA and countries including: Japan, Mexico, Germany, Switzerland, Israel | Terminated | June 21, 2011 | 412 | 1630
Z-1004-2011 | 56759 | No | NQQ | Ultrasonic pulsed echo imaging system. | Radiology | General &amp; Plastic Surgery | 510(k) | Optovue iVue optical coherence tomography system, computer controlled opthalmic imaging system, manufactured by Optovue, Fremont, CA |  |  | July 20, 2010 | 2010 | Optovue Inc. | Use of the CAM attachment may result in a positioning anomaly and mistreatment in the eyes of the patient. | Software design (manufacturing process) | Optovue Inc. sent a Field Modification/Correction Notice dated July 20, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were informed that they could continue to use the iVue until the correction has been made. Customers were instructed to review any existing iVue patient data to ensure that this situation and an improper designation of eye assignment has not occurred. Upon completion of the updates, customers were to sign the attached form indicating they acknowledge that their iVue has been updated and they understand that they must review existing data. For any questions regarding this recall call 1-866-344-8948 or by mail at 45531 Northport Loop W, Fremont, CA 94538. |  | 22 | 22 devices in the US. | Nationwide | Terminated | February 08, 2011 | 203 | 22
Z-1009-2011 | 57576 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AutoSuture SURGIPRO Polyproylene Mesh Clear 3" x 5" Sterile |  |  | December 15, 2010 | 2010 | Covidien LP | Sterility compromised due to breach in sterile barrier | Packaging process control | Covidien notified customers via letter dated November 15, 2010. The letter identified the product, the problem, and action to be taken by the customer. Customers were instructed to examine inventory and return the affected lot to Field Returns Department, 195 McDermott Road, North Haven, CT 06473. For questions regarding this recall call 203-492-8091. |  | 132 | 132 units | Worldwide Distribution - USA including CA, CT, FL, HI, IL, IN, KS, MA, MD, MI, MN, NC, NJ, NY, OH, PA, SC, TX, and WI and the countries of Canada and Italy | Terminated | January 10, 2013 | 757 | 132
Z-1010-2011 | 57595 | No | ONG | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Fraxel re:storeÂż Dual Laser System----a laser surgical instrument for use in general and plastic surgery and in dermatology. |  |  | December 15, 2010 | 2010 | Solta Medical, Inc. | The product has the potential to cause inadvertent laser firing which may lead to patient/operator injury if the problem recurs. | Software design | The firm, Solta Medical, issued three "DEVICE CORRECTION" letters, one dated December 15, 2010, the second one date December 17, 2010 and the third one dated December 20, 2010, to all affected customers. The letters described the product, problem and action to be taken by the customers. The customers were instructed to follow the instructions carefully, sign and date the bottom of the instruction form upon completion of the update and promptly return the signed form and USB stick in the enclosed, return-addressed envelope (Product Support, Solta Medical, 25881 Industrial Blvd., Hayward, CA 94545) Note: Solta Medical sales personnel will assist with the software update process so, if your local representative has already updated your system, you do not need to do anything further.If you have any questions, please call Solta Product Support at (510) 259-7291. |  | 352 | 352 units | Worldwide Distribution: USA and countries including: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Egypt, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Kuwait, Mexico, Netherlands, New Zealand, Pakistan, Philippines, Poland, Russia, Singapore, South Africa, Spain, Thailand, Turkey, Ukraine, United Kingdom, and Vietnam | Terminated | November 14, 2011 | 334 | 352
Z-1017-2011 | 57661 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Port Dressing Tray, Reorder DT12885, Sterile, Centurion Medical Products Corp. Williamston, MI. |  |  | December 06, 2010 | 2010 | Centurion Medical Products | The firm is conducting a recall of kits that contain the MONOJECT PREFILL ADVANCED flush syringes containing heparin. The syringes are under recall by their manufacturer Coviden because of testing that showed "that the crude heparin used to produce a specific batch of purified heparin sodium USP contained low level of Over-Sulfated Chondroitin Sulfate (OSCS)". | Other | The firm, Covidien, sent an "URGENT PRODUCT RECALL" letter dated November 12, 2010, to affected customers via Certified Mail, Return Receipt Requested. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory and in-use stock to determine if they have any units of the affected product, if found, immediately remove from inventory and quarantine; if they are a distributor they were ask to notify their customers of this letter; and complete and return the PRODUCT RECALL FORM whether or not they have the affected product via fax to: 203-822-6009 or scan and email to: sdfeedback@covidien.com. Note: The returned product will be collected and counted prior to destruction or return to the vendor.Should you have any medical questions concerning this recall, please call 508-261-6393.For Customer Service questions regarding the recall, please use one of the following methods: phone 1-800-962-9888 option 1 (8am to 6:30pm Eastern); email: sdfeedback@covidien.com or fax: 302-822-6009. |  | 2250 | 2250 kits. | Nationwide Distribution: USA including states of: OH, WI, IL, MO and TX. | Terminated | February 16, 2011 | 72 | 8580
Z-1012-2011 | 57259 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Arrow WH-02220 - Epidural Catheterization Kit with FlexTip Plus(R) Catheter |  |  | November 11, 2010 | 2010 | Arrow International Inc | Sterility may be compromised. | Employee error | The firm, Arrow International, issued an "URGENT MEDICAL DEVICE RECALL" letter dated November 11, 2010 to their consignees/customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to check their inventory for the affected products, cease use and distribution, quarantine all affected product immediately, return any affected product freight collect along with the completed form via mail to: Arrow International, Inc., 312 Commerce Place, Asheboro, NC 27203, Attn: Asheboro Quality Assurance Manager, and complete and return the FIELD SAFETY CORRECTIVE ACTION ACKNOWLEDGEMENT AND STOCK STATUS FORM via fax to: 800-343-2935.If you require additional information or clarification regarding this matter, please contact your sales representative, or Arrow's Customer Service Department at 866-396-2111. |  | 2320 | 2320 | Nationwide Distribution: USA states including IL and LA. | Terminated | July 05, 2012 | 602 | 2320
Z-1011-2011 | 57275 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL90 FLEX Blood Analyzer Part Number: 393-090. |  |  | November 01, 2010 | 2010 | Radiometer America Inc | Due to a software error, a patient mix up can occur on the ABL90. If a sample is pre-registered and the sample "fails" the next result will inherit the patient data from the previous result. | Software design | On 11/16/10 the firm sent Medical Device Field Correction letters to their customers. The letter states that all systems will be upgraded and until each system is upgraded, they instructed their customer to "always query for patient ID prior to introducing the sample." They also instruct each customer to complete the "Fax Return Form" acknowledging receipt of the letter. If customers have questions regarding the letter or the upgrade, they should contact Technical Support at 1-800-736-0600 Opt 4. |  | 432 | 432 units | Worldwide Distribution -- USA, including states of MI, NJ, &amp; PA and countries Australia, Austria, Belgium, Denmark, Finland, France, Germany, Indonesia, Italy, Japan, Kuwait, Malaysia, Netherlands, Norway, Poland, Russia, Spain, Sweden, Switzerland, United Arab Emirates &amp; United Kingdom. | Terminated | May 09, 2012 | 555 | 432
Z-1027-2011 | 52903 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER component of MOSAIQ versions 1.30 (all builds), 1.40 (all builds), 1.50 (all builds), 1.60 (all builds), and 2.0 (all builds). Manufactured by IMPAC Medical Systems, Inc. Sunnyvale, CA. |  |  | July 17, 2009 | 2009 | Impac Medical Systems Inc | Potential for improper treatment. A software error may result in potential major underdosing to targeted areas from jaw blockage of the MLC aperture (jaw intrusion), and dosing to non-targeted areas from interstitial leaf leakage (jaw extrusion). | Software design | Impac Medical Systems Inc sent an Important Safety Notice letter to all affected consginees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Impact Software support as soon as possible if they have the affected configurations identified in the letter. Customers were asked to complete the form included with the letter and and fax to:European customers: +44 1293 654401 American &amp; rest of world customers 702-992-5002 or e-mail to support@impac.com Customers were asked to distribute the Safety Notice to any and all users of IMPAC software at their organization who are potentially affected by this issue. For questions e-mail support@impac.com or callEuropean customers: 44 1293 654320American &amp; rest of world customers: 800-488-4672. |  |  |  | Worldwide Distribution - USA including AL, CA, CT, FL, GA, IA, IL, KY, MI, MN, MS, ND, NJ, OH, OR, PA,TN, TX, VA, WA, WI, and WV and the country of Canada. | Terminated | November 25, 2011 | 861
Z-1367-2011 | 53012 | Yes | NIQ |  |  | Cardiovascular | PMA | Endeavor Resolute Zotarolimus-Eluting Coronary Stent System with Rapid Exchange Delivery System, UPN 00613994400505, model #IDERES25024X. Manufacturer: Medtronic, Inc. Minneapolis, MN. Manufactured in: Medtronic Ireland, Galway, Ireland. Product is packaged inside a sterile, inner pouch, labeled in part "Use immediately". Inner pouch is supposed to be within a foil pouch containing desiccant, and moisture absorber. Foil pouch label reads in part "RX Endeavor Resolute Zotarolimus-Eluting Coronary Stent System...2.5 mm X 24 mm...REF IDERES25024X...Contents: One (1) Endeavor Resolute zotarolimus-eluting stent mounted on a rapid exchange stent delivery system...Guide Catheter/Minimum Inner Diameter &gt;=5F/0.056" (1.42 mm)...Maximum Guidewire Diameter 0.014" (0.36 mm)...Maximum Stent I.D. 3.00 mm (0.12"...Nominal Pressure 9 atm (912 kPa)...Rated Burst Pressure 15 atm (1520 kPa)...MR Conditional...Sterilized using ethylene oxide Consult instructions for use...Manufacturer: Medtronic, Inc. 710 Medtronic Pkwy NE Minneapolis, MN 55432 USA...Manufactured In: Medtronic Ireland Parkmore Business Park West, Galway, Ireland Tel:+353-91-708000 Fax:+353-91-757524 Investigational Devices / Returned Goods: Tel: (800) 556-4247 / Fax: (877) 518-2469...BioLinx Medtronic Polymer System". Foil pouch contents and Instructions for Use are placed within an outer box, which is labeled identically to the foil pouch. |  |  | August 06, 2009 | 2009 | Medtronic Vascular Galway Limited | The product was removed from investigational sites due to a packaging issue, in that the device units were packed without the foil pouch, oxygen absorber and Instructions For Use. | Packaging | The firm, Medtronic, sent their initial "Urgent Medical Device Recall Notice" to customers with affected product via e-mail on August 7, 2009. An additional "Urgent Medical Device Recall Notice" was sent to all customers on August 12-13, 2009 via FedEX and UPS. The notice described the product, problem and action to be taken by the customers. The customers were instructed to remove the affected units from their inventory and isolate the units in a secure location to prevent additional use. A Medtronic CardioVascular Clinical Customer Service (CCS) representative will contact their site to arrange for return of the units. This will be followed by a shipment of replacement units to their sites. If there were any questions regarding the return process, customers were to contact CCS at 800-556-4247.If you have any further questions or concerns regarding this notice, please contact the Sr. Clinical Research Manager at 707-543-2284 or ann.depaoli@medtronic.com. |  | 49 | 49 units | Nationwide distribution: AL, AR, CA, FL, IN, MD, MI, NY, OH, PA, RI, TX, VA, and VT. | Terminated | February 17, 2011 | 560 | 49
Z-1022-2011 | 53014 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PEAK brand PlasmaBlade EXT, Tissue Dissection Device, |  |  | May 19, 2009 | 2009 | Peak Surgical | A customer reported that a patient suffered a burn while using the device. | Other | Peak Surgical notified afftected customers by phone on May 19, 2009. They were notified of the product, the problem, and the action to be taken. Peak Surgical sales representatives removed PlasmaBlade EXT devices from consignees affected by the recall. |  | 493 | 493 units | Worldwide Distribution - United States including CA, CO, FL, IL, MD, MN, NV, NY, OH, OK, TN, TX, UT, and WAand the countries of Turkey and Ireland | Terminated | April 21, 2011 | 702 | 493
Z-1021-2011 | 56923 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vital Connect 4.0 &amp; 4.1, radiological image processing software |  |  | August 23, 2010 | 2010 | Vital Images, Inc. | There is a potential alignment error of fused PET/CT images in the Vitrea Core component of Vitrea Enterprise Suite 1.2 and 1.3. The error occurs during panning of the fused PET/CT images. Because the PET data has a larger voxel size than the CT data, it is possible to pan the CT image in more frequent, smaller steps than the PET image, so that the fused PET image does not move until the CT data | Software design | The firm, Vital Images, Inc,., sent a "Urgent Software Alert" letter dated August 23, 2010 to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to not use the PET/CT fusion functionality until the software updates are provided and to re-examine any past results. If you have any questions or comments, contact Customer Support at support@vitalimages.com or at 1-800-208-3005. |  | 178 | 178 | Worldwide Distribution: USA including states of: AK, AZ, CA, CO, CT, FL, GA, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, NH, NJ, NM, NY, NC, ND, OH, PA, SC, TN, TX VA, WV, WI, and WY; and countries of: AUSTRALIA, HONG KONG, UNITED KINGDOM, NETHERLANDS, FRANCE, BELGIUM, CANADA, CHINA, AUSTRALIA, SWITZERLAND, and ITALY.AUSTRALIA, HONG KONG, UNITED KINGDOM, NETHERLANDS, FRANCE, BELGIUM, CANADA, CHINA, AUSTRALIA, SWITZERLAND, ITALY. | Terminated | November 06, 2012 | 806 | 178
Z-1023-2011 | 54658 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation therapy device. Medical charged particle radiation therapy system with software version. Model Numbers: Multi Plan Treatment Planning System, Software version 3.0, 3.1, 3.5.0, and 3.5.1. |  |  | January 15, 2010 | 2010 | Accuray Inc | When a user modified the align center of an existing plan during the treatment planning process the effective depths of each beam in the plan were not updated by the system. These depths were incorrect and led to an incorrect dose calculation. Software patch /update to address this issue. Implementation at affected sites will initiate, once the patch is available. | Software design | The firm, Accuray, sent an "URGENT DEVICE CORRECTION" letter dated January 18, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to load the plan and recalculate the dose to reflect the current state of beam, compare the reloaded plan and the recalculated plan to verify maximum dose and prescription values, and re-prescribe and save to overwrite to existing plan if necessary. The firm is developing a patch/software update to address the issue and a Accuray Customer Support person will contact the customer regarding further resolutions.If you have any questions or concerns regarding this issue, contact Accuray Customer Support at +1-877-668-8667(USA) or +1-408-716-4700 (non USA) or email: customersupport@accuray.com. |  | 99 | 99 units | Worldwide Distribution: USA and countries including: Canada, Turkey, India, Russia, Italy, France, Belgium, Netherlands the UK, Spain, Athens, Saudi Arabia, Switzerland, the Ukraine, Germany, Korea and Hong Kong. | Terminated | February 11, 2011 | 392 | 99
Z-1025-2011 | 57555 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb, Renu Fresh Multi-purpose Solution, 2 fl oz (60 mL), Travel Kit, Manufactured by: Bausch &amp; Lomb Place, Rochester, NY 14609 |  |  | December 22, 2010 | 2010 | Bausch &amp; Lomb Inc | There is a potential for some packages to not meet sterility requirements. | Other | On December 8, 2010, Target was visited by FDA/PHI-DO and was informed on the violation. The Target customers that received lot GH0078 were requested to return all affected product directly to the Bausch &amp; Lomb Greenville, SC Distribution Center. Bausch and Lomb sent an Urgent Voluntary Medical Device Recall to all first consignees who were shipped the affected lot GH0066 of Renu Fresh Multipurpose Solution on/about December 22, 2010. Since Target customers had already been notified previously regarding lot GH0078, this lot number was not included in the subsequent recall notification. All consignees were requested to review their inventory, and return any of the affected product to Stericycle Pharmaceutical Services, 2670 Executive Drive, Indianapolis, IN 46241. Stericycle is conducting this recall under the control of Bausch &amp; Lomb.For questions abut the recall process contact Stericycle, Inc. at 1-888-345-8316.For questions regarding this recall call 1-800-828-6974. |  | 292992 | 146,496/ 2 oz Bottles | Worldwide Distribution - USA including AL, CA, CT, FL, GA, HI, IA, IN, LA, MA, MI, NE, NJ, NY, OH, PA, RI, SC, TN, TX, UT, VA, and WI and the countries Taiwan and Brazil | Terminated | September 04, 2012 | 622 | 360288
Z-1024-2011 | 57464 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Laryngoscope Handle, Reusable, Rx, Product Code 4558GSP or 4559LED, Vital Signs Colorado, Inc., 11039 E. Lansing Circle, Englewood, CO 80112. |  |  | December 20, 2010 | 2010 | Vital Signs Colorado Inc. | Battery which powers the laryngeal scope light may dislodge resulting in loss of illumination. | Device Design | Healthcare Systems sent an URGENT MEDICAL DEVICE CORRECTION letter dated December 20, 2010, to all affected consignees advising them of the possible problem and giving them interim safety instructions until new battery cartridges were available. Customers were instructed to complete and return the Confirmation Form which was included with the letter.For information regarding this recall call 973-956-5300. |  | 2202 | 2202 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, FL, GA, HI, IA, IL, IN, KY, MA, MN, MO, NH, NJ, NY, OH, OK, SC, TN, TX, WA, and WV and the countries of Australia, Belgium, Canada, Dominican Republic, Ecuador, Germany, Hong Kong, Latvia, New Zealand, Taiwan, Turkey, and the UK | Terminated | July 01, 2011 | 193 | 2202
Z-1020-2011 | 56392 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | i-STAT Level 2 Control Value Assignment Sheets |  |  | August 03, 2010 | 2010 | Abbott Point of Care Inc. | Three i-STAT Level 2 Control Value Assignment Sheets and the i-STAT CLEW A20 Software System Manual Update CD contain the incorrect range for pCO2, BUN and Glucose for the i-STAT EC8+, 6+, EC4+, E3 and G cartridge types. | Other | Abbott Point of Care Inc., sent an URGENT RECALL NOTICE to all affected customers on 8/3/2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to replace all copies of the value sheets with the ones provided in the letter. ASLO, if they have a disc with the Value sheets on it, it is to be replaced with the updated disc and the original discarded. If the control materials, discs or value sheets have been forwarded to another facility, a copy of the letter was to be sent to them. The Business Reply card is to be completed and returned to Abbott. Questions or comments are to be directed to Abbott Point of Care Technical Support at 800-366-8020, Option 1. |  | 8293 | 8293 Value sheets worldwide | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of Argentina, Austria, Belgium, Bolivia, Brazil, Chile, China, Columbia, El Salvador, France, Georgia Republic, Germany, Hong Kong, Israel, Italy, Kuwait, Mexico, Netherlands, New Zealand, Panama, Paraquay, Peru, Poland, South Africa, Sweden, Switzerland, Taiwan, Turkey, UAE, UK, Uruquay, and Yemen. | Terminated | September 28, 2011 | 421 | 8293
Z-1030-2011 | 57570 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Precise Treatment Table |  |  | March 10, 2008 | 2008 | Elekta, Inc. | Possible error in the table longitudinal position. | Other | Elekta sent an updated IMPORTANT NOTICE letter dated December 24, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Elekta recommended that all users, and Eleckta field service engineers do regular QA checks to compare the movement of the treatment table in a given axis with the digital readout. Elekta Service representatives would attach the screws with Loctite and calibrate the axis on these sites. For questions regarding this recall call 770-670-2548. |  | 597 | 597 units | Nationwide Distribution and Puerto Rico | Terminated | August 09, 2012 | 1613 | 597
Z-1031-2011 | 57571 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Presice Cryosurgical System, model FPRCH2047, and the following system is not approved in the US: CryoHit-P Cryoblation System, model FPRCH2058. The system is a software controlled cryoblation system designed to destroy targeted tissue using a minimally invasive percutaneous approach. Freezing and thawing of the tissue is achieved through the use of the Joule-Thompson effect, using a high-pressure gas distributed through closed-end needles placed into the patient;s target organ. |  |  | December 20, 2010 | 2010 | Galil Medical, Inc. | This voluntary correction has been initiated due to the discovery of a damaged optocoupler on one of our Presice Systems during periodic preventive maintenance. The optocoupler sends a signal across the locking mechanism in the Presice and Seednet MRI Systems, enabling the software to recognize if the channel on the system manifold is locked or unlocked. If an optocoupler is damaged or malfunction | Device Design | Galil Medical sent an "Urgent Medical Device Correction" letter dated December, 30, 2010, to all affected consignees. The letter described the product, the problem involved, and the action to be taken by the consignee. Consignees were instructed to make sure the locking mechanism on the Precise and Seednet MRI Systems were in the "locked" position when a needle is inserted into the manifold. Gently pull on the black plug to ensure needle is secure in locking mechanism. Apply the enclosed alert sticker on the bottom of the distribution panel and the bottom of the system monitor of the Precise System and on the side of the mobile unit itnerface panel and the top of the system monitor of the Seednet MRI System. Consignees were asked to take a picture of the alert sticker after it has been applied to the appropriate location on their Presice or Seednet MRI System(s) and e-mail to yoni.kadosh@galilmedical.com. Consignees were instructed to complete and return the enclosed form as soon as possible.For questions regarding this recall call Customer Service at (877) 639-2796.problem involved. |  | 63 | 63 | Worldwide Distribution - USA including CA, CT, FL, GA, MA, MD, MN, MO, NC, NY, OH, PA, SC, TN, TX, VA, and WA and the countries of CANADA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, ISRAEL, ITALY, KOREA, NETHERLANDS, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, CHINA, and JAPAN.CANADA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, ISRAEL, ITALY, KOREA, NETHERLANDS, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, CHINA, JAPAN. | Terminated | March 27, 2012 | 463 | 85
Z-1033-2011 | 57650 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, CADDÂż-Solis Ambulatory Infusion Pump, Model Number: 2100 and 2110, Reorder Number 21-2101-00, 21-2101-0100-50, 21-2101-0100-51, 21-2101-0200-02, 21-2101-0200-03, 21-2111-0100-51, 21-2111-0200-02, and 21-2102-51. |  |  | January 13, 2011 | 2011 | Smiths Medical ASD, Inc. | Reports of the Cassette Lock mechanism on certain CADDÂż-Solis Pumps becoming stuck in the locked position, making it nearly impossible to remove the medication cassette reservoirs (or administration set). The purpose of the Cassette Lock is to secure the medication cassette reservoir (or administration set) to the pump. | Nonconforming Material/Component | The firm, Smiths Medical, sent an "URGENT MEDICAL DEVICE CORRECTION NOTICE" dated January 13, 2011 to all consignees/customers. The notice described the product, problem, and action to be taken by the customers. The customers were instructed to have their Pumps undergo a correction which consist of removal and reattachment of the Cassette Latch handle using a new adhesive process. The customers were also instructed to remove the Pump from service if unable to remove cassette (or administration set); to inspect their inventory and identify Pumps; prepare a plan for correction of these Pumps; complete and return the attached CONFIRMATION FORM via fax to 1-800-237-8033 or by email to CADDDSolis_CassetteLock@smiths-medical.com, and if they have distributed the affected product to other persons or facilities, promptly forward this notice to the recipients. If you should have any questions regarding this information, please contact Smith Medical at 1-866-216-8806. |  | 2903 | 2903 (2539 US; 364 OUS) | Worldwide Distribution: USA including states of: AK, AZ, C, CO, DC, FL, GA, ID IL, MA, MI, MN, MO, MT, NE, NH, NM, NY, NC, OK, OR, PA, RI, TN, TX, UT, VT, VA, WA, WI, and WY; and countries including: Australia, Austria, Belgium, Canada, France, Germany, Great Britain, Ireland, Italy, Luxembourg, New Zealand, Qatar, Saudi Arabia, Switzerland and The Netherlands. | Terminated | April 17, 2012 | 460 | 2903
Z-1038-2011 | 57721 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Irrigation Kit, Large, Vinyl, REF: 6114, Apogee Medical, LLC, 90 Weathers Street, Bournesville, NC 27596. |  |  | January 18, 2011 | 2011 | Apogee Medical | This recall has been initiated due to concerns expressed by the Food and Drug Administration regarding the validation of the gamma radiation sterilization cycles for these products of Sterile Lubricating Jelly manufactured by Triad Group. | Other | The firm, TRIAD Group, sent a "Urgent Medical Device Recall" letter dated January 18, 2010 to all consignees/customers. The letter described the product, problem and the action to be taken by the customers. The customers were instructed to review there inventory, cease distribution of the affected lots and to complete and return the enclosed Product Recall Response Form via fax to 1-919-570-9611 or e-mail to Apogee Medical. If they distributed the product, they were instructed to notify their customers of the recall; to include a copy of the recall notification, and to instruct them to return the recalled product to the wholesaler who would then return the product to Apogee Medical. Non-responders to the recall notice will be contacted on a weekly basis.Should you have any questions regarding this notification, please call the Regulatory Affairs/Quality Assurance Manager at (919) 435-5409. |  | 1 | see line 1. | Nationwide Distribution: USA including states of: AL, AZ, CA, FL, GA, IL, IN, KS, LA, MA, MO, NC, NJ, NY, OH, OK, PA, TN, TX, VA and WA. | Terminated | February 28, 2012 | 406 | 4454545
Z-1028-2011 | 57504 | No | BZQ | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | ThermiSense model 5700-A5 |  |  | December 21, 2010 | 2010 | Salter Laboratories, Division of Regulatory Affairs | The incorrect length and connector were utilized in lot 110810 of the ThermiSense Alice 5, making the product unusable as it will not interface with the sleep labs equipment. | Employee error | Salter Laboratories, Inc. sent an URGENT PRODUCT RECALL LETTER dated December 20, 2010, to all affected customers. The letter identfied the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for this product and lot number. Contact Salter to arrange for the return of the product. Complete the attached reply form, confirming receipt of the letter, and fax it back to Salter Labs at 1-800-628-4690.For questions regarding this recall call 661-854-3166. |  | 41 | 41 units | Nationwide Distribution including CA, CO, DE, ID, IN, KY, MD, MO, and OH, | Terminated | June 06, 2011 | 167 | 41
Z-0993-2011 | 57507 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Arstasls One Access System; Models: AAD100, AAD101, Part Numbers: FG-02279, FG-03010; The system consists of the following components: (1) Device Sheath/Anchor, (2) Shaft |  |  | October 19, 2010 | 2010 | Arstasis, Inc | Device has the potential to fracture and sheath separation during use. | Device Design | The firm, Arstasis, initiated a market withdrawal and sent notification via letters dated October 19, 2010, to customers informing the customers that Arstasis is implementing a product improvement and will replace all existing inventory that the customer has on their shelf with the improved product. The customers were instructed to work closely with their local Arstasis territory manager to ensure that their product replacement is handled satisfactorily, and complete and return the PRODUCT REPLACEMENT and PRODUCT RETURN forms to Arstasis Customer Service via fax at 1-877-688-8417 or email to customerservice@arstasis.com.If you have any questions or concerns, contact Customer Service toll free at 1-877-594-4545. |  | 330 | 330 units | Worldwide Distribution | Terminated | February 16, 2011 | 120 | 330
Z-1029-2011 | 55087 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER for Mosaiq interfaced to Electa SL75-5 machines. |  |  | March 10, 2010 | 2010 | Impac Medical Systems Inc | Computer interface is defected. Computer based verification system for recording patient treatment on a radiotherapy treatment machine fails to record treated field. | Software design | A User Notice was sent by the firm to its consignees. It identified the affected product and the reason for the recall. It also discussed the clinical impact and interim mitigations. The firm states in the letter that a fix for the problem and been developed and is available. Customers are to contact IMPAC Support for the patch to the SL75-5 interface. If customers have questions, concerns, or requests, they can email support@impac.com or one of the numbers listed in the letter for the applicable country. The letter should be distributed to all users affected by the issue at the customer's organization. |  | 40 | 40 | Worldwide Distribution -- US, Canada, Germany, and France. | Terminated | April 06, 2011 | 392 | 40
Z-1042-2011 | 57549 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Vertex Max Drill Tube, Catalog Number 9733140, a component of the Navigated Vertex Max Drill Guide Kit, Catalog Number 9733370. Medtronic, Inc., Louisville, CO. |  |  | December 15, 2010 | 2010 | Medtronic Navigation, Inc | Incorrectly manufactured drill tube could cause incorrect screw placement during spinal surgery. | Component design/selection | Consignees were notified by a Product Correction Notification letter on 12/15/2010 and given the affected product and a description of the problem. The letter stated that a Medtronic representative would contact customers to visit their facility and replace the affected guide with one that has been correctly manufactured. Additional information is available at 800-595-9709. |  | 15 | 15 devices | Worldwide Distribution -- USA, including states of GA, IL, IN, MN, MO, OH, OR, PA, and TX and countries of Denmark, Finland, Singapore, Spain, and UK. | Terminated | May 04, 2011 | 140 | 15
Z-1019-2011 | 57558 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | PCR Eleva. The marketing brochure is labeled in parts: "Philips Medical Systems is part of Royal Philips Electronics... www.medical.philips.com... Philips Medical Systems Nederland B.V". |  |  | December 23, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | There is a risk of selecting the wrong patient, when patient is selected in the patient list and the examination tab pressed with background query are running at the same time. | Software design | On 1/30/2009, Philips began sending out URGENT-Device Correction PCR Eleva notifications to their consignees.The consignees were informed of the hazards as when a patient is selected in patient list and examination tab pressed with background query running at the same time, there would be a risk that the wrong patient is selected.The firm issued the Field Corrective Order (FCO) 73200029 to their representative to correct the problem. The corrective action would be a software upgrade that resolve the issue.All consignees can contact Philips Call Center at 800-722-9377, #5, #2 and reference "FCO 73200029" with any questions. |  | 686 | 686 units | Nationwide Distribution -- USA, including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WI, WV, and WY and to LANDSTUHL ARMY BASE, APO in United Arab Emirates. | Terminated | August 25, 2011 | 975 | 686
Z-1046-2011 | 56448 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Omnifit M/S PSL Porous Coated Acetabular Shell 40MM; |  |  | August 02, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics determined that specific lots of the Omnifit Acetabular Shells may have a screw hole wall thickness that falls above the design specification. | Device Design | The firm, Stryker Orthopaedics, sent an "URGENT PRODUCT RECALL" letter dated August 2, 2010, via Federal Express to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory to identify the affected product; retrieve all affected product and return it their branch or agency warehouse for reconciliation; reconcile all product(s) on the attached Product Recall Acknowledgement Form and fax to (201) 831-6069, and return the affected product via Inter-Org Transfer and ship to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100. |  |  |  | Worldwide distribution: USA including states of: AL, AR, CA, FL, IN, ME, MI, MS, NC, NJ, OK, OR, WA , and WI; and countries including: China, Hong Kong and Korea. | Terminated | November 09, 2012 | 830
Z-0943-2011 | 57193 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Discovery ST (16 Slice) System; Discovery ST (16 Slice) Systems with technical Reference Manual numbers 2351785-xxx and 2390935-xxx distributed from November 2001 to August 2010. |  |  | October 27, 2010 | 2010 | GE Healthcare, LLC | Although the system x-ray output is accurate for the technique factors given, and the dose display is also accurate, certain CTDI adjustment factors used for the manual calculation from the Technical Reference Manual (TRM) are incorrect. CTDI100 Aperture Adjustment Factors for 4x0.625 mode contain incorrect values. This issue may lead to inaccurate CTDI manual calculation. The results of the diffe | Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter will be provided together with the TRM addendum informing the customers of the issue and related hazards. The letter also identified the affected product and gave instructions, which include referring to the TRM addendum, along with the product correction. This correction will be executed via a Field Modification Instruction IFMI 254281 released in November, 2010 with an estimated completion date of May, 2011. Questions or concerns should be directed to the call numbers listed or a local GE Healthcare Service Representative for countries not listed. |  | 1450 | 1450 units total in US | Worldwide Distribution | Terminated | July 13, 2012 | 625 | 1450
Z-1064-2011 | 53054 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Linear Medical Accelerator ARTISTE MV systems, when equipped with a syngo RT Therapist, Therapist Express Basic, syngo RT Therapist Assist or Therapist Express Assist workspace with software version NET VA10A for RTT 4.1.122 or NET V10B for RTT 4.1.125, manufactured by Siemens Medical Solutions, Concord, CA. |  |  | July 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Potential for Mistreatment-- when treatment plans intended for non-ARTISTE machinese are loaded, if the plans contain unspecified beam parameters, attributes or unknown block codes, the therapy system may remove those parameters. | Software design | The firm,Siemens, sent two Customer Advisory letters (TH004/09/S and TH008/09/S) contained in the Update Instructions to all consignees/customers with active devices. The letters describe the product, problem and action to be taken by the customers. The customers were instructed to follow the updated instructions. Siemens service representative will sign a document noted as CORRECTED upon completion of the software installation.If you have any questions, call 1-925-602-8157. |  | 151 | 151 units | Worldwide distribution: USA and countries including: Germany, France, Poland, Norway, Australia, Malaysia, Spain, New Zealand, the UK, India, Republic of Korea, Canada, Ireland, Italy, Egypt, Belgium, Philippines, Saudi Arabia, Netherlands, Hungary, Norway and South Africa. | Terminated | April 12, 2011 | 621 | 151
Z-1063-2011 | 57531 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Inrad sample master aspiration biopsy needle with hiliter sharp bevel needle, 22 ga x 2 3/8'' (6 cm), model 41222, Inrad, Inc. Kentwood, MI. |  |  | December 20, 2010 | 2010 | Inrad, Inc | This lot of product has the wrong expiration date on the carton label. The pouches contained the correct expiration date. | Employee error | The firm, Inrad Inc., sent an "URGENT FIELD CORRECTION" notice dated December 20, 2010 to its customers. The notice described the product, problem and actions to be taken by the customers. The customers were instructed to immediately examine their inventory for the affected product. If the recalled product was in stock, the customers were told to removed the affected product from the carton and dispose of carton; to disregard carton label expiration date and reference the expiration date appearing on the pouch label only, and to complete, sign and return the attached "Carton Expiration Date" INFORMATION FORM via fax to 616-301-779 or mail to Inrad Inc., ATTN: Director RA/QA, 4375 Donker Court SE, Kentwood, MI 49512.If you have any additional questions about this action, please call 800-558-4647 extension 102. |  | 250 | 250 | Nationwide Distribution: USA including states of: FL, GA, KY, NH, NJ, NM, PA , SC, WA, and WI. | Terminated | February 09, 2011 | 51 | 250
Z-1045-2011 | 57610 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian High Energy Linear Accelerator: |  |  | December 10, 2010 | 2010 | Varian Medical Systems, Inc. Oncology Systems | An anomaly was identified whereby the bolts used to fasten the counterweight to the gantry may not, in some cases, be fully tightened to the required torque specification. | Process control | An Urgent Medical Device Correction letter was sent to all affected users with a description of the problem and user corrective action steps. Letters were sent on 12/10/2010. Users are to immediately investigate any unusual noises coming from the affected product and report them to whoever is responsible for servicing the device. For customers that do have a service contract with Varian, Varian service will perform an inspection of the counterweight bolts as part of the regular Periodic Maintenance Inspection procedure. Customers are to advice the appropriate personnel of the contents of the letter. For further clarification, customers can contact their local Varian Customer Support Services Manager. |  | 1882 | 1882 | Worldwide Distribution. | Terminated | August 02, 2011 | 235 | 1882
Z-0998-2011 | 57524 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Diabetes Care, Precision Point of Care Blood Glucose Test Strips; |  |  | December 22, 2010 | 2010 | Abbott Diabetes Care, Inc. | Certain lots of Precision Xtra, Precision Xceed Pro, MediSense Optium, Optium, OptiumEZ and ReliOn Ultima Blood Glucose Test Strips may generate erroneously low blood glucose results. | Process control | On 12/22/10, Abbott Diabetes Care notified customers via Press Release and letters explaining the reason for the recall and requesting the products be returned. |  | 359000000 | 359 total lots, 1,000,000 strips per lot. | Products were released for distribution in US and Puerto Rico. | Terminated | October 25, 2011 | 307 | 359000000
Z-1065-2011 | 57333 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Bond Polymer Refine Detection Kit; an in vitro diagnostic immunohistochemistry reagent kit; Catalog No. DS9800; |  |  | November 16, 2010 | 2010 | Leica Microsystems, Inc. | Users of the Bond Polymer Refine Detection System, DS9800, may experience instances of "weak staining" due to the DAB Part B component, which may give a false negative result. | Other | Leica Microsystems sent Medical Device Field Correction letters dated 11/16/10 to the affected users of the Bond Polymer Refine Detection System, DS9800 via first class mail on the same date, informing them that they may experience instances of "weak staining" due to the DAB Part B component in the affected lots, which may give a false negative result. The users were requested to test their DS9800 System prior to use to rule out that the DAB Part B components have been affected, using the enclosed testing protocol. If their system is affected, the users were instructed not to use it and contact Leica Technical Assistance Center for a replacement. The accounts were asked to complete the attached Acknowledgement Form, acknowledging the receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. |  | 653 | 653 units | Worldwide Distribution -- USA and Canada. | Terminated | October 17, 2011 | 335 | 653
Z-1067-2011 | 57676 | [6 empty] | Fujinon Endoscope Series; |  |  | January 07, 2011 | 2011 | Fujinon Inc | The FICE feature installed in video endoscopic processors is being disabled since this feature should have been reviewed under the premarket notification/510k process. | PMA | Fujinon sent an Urgent Medical Device Corrective Action Letter, dated 1/7/2011, via overnight mail to all consignees. The letter advises users that the FICE feature should no longer be used and that a Fujinon Field Service Engineer will be coming to physically disable the FICE feature on the processor. Instructions labeled as Attachment B are provided to assist in ceasing use of the FICE function. Attachment A is provided for frequently asked questions. Customers are to complete the Response/Acknowledgement Form and either fax it or email back to the firm. If there are additional questions, customers can call 800-385-4666 or email products.fice.@fujinon.com. |  | 783 | 783 units | Nationwide Distribution | Terminated | June 04, 2012 | 514 | 783
Z-1066-2011 | 57321 | No | CKA | Alanine amino transferase (ALT/SGPT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abaxis brand Piccolo Comprehensive Metabolic Panel; Used with the Piccolo Blood Chemistry Analyzer or Piccolo xpress Chemistry Analyzer; Catalog/Part Number: 400-0028; |  |  | November 19, 2010 | 2010 | Abaxis, Incorporated | Abaxis has determined that during the manufacture of product lot 0396BC5 may report inaccurate values of Alanine Aminotransferase (ALT) when used in the Piccolo Blood Chemistry Analyzer or Piccolo xpress Chemistry Analyzer. | Other | The firm, Abaxis Inc., sent an "URGENT PRODUCT RECALL NOTICE" dated November 19, 2010, to all customers. The notice described the product, problem and action to be taken by the customers. The customers were instructed to check their inventory of the lot for boxes #0665 or higher; if found, to discontinue use; contact Abaxis for immediate replacement; complete and return the enclosed PRODUCT RECALL RESPONSE form to Abaxis Technical Support via fax at 1-877-349-2087, and ensure that their laboratory staff is informed of this recall notice.If you have any questions regarding this Product Recall Notice, please contact Abaxis Medical Technical Support at 800-822-2947, or at medtechsupport@abaxis.com. |  | 797 | 797 units | Worldwide distribution: USA and countries of: Kowloon and Germany. | Terminated | August 23, 2011 | 277 | 797
Z-1080-2011 | 56391 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Centurion Medical Products Insertion/PPE Kit CV12590, Sterile. |  |  | June 02, 2010 | 2010 | Centurion Medical Products | The kits contain heparin lock flush solution USP code 504505 that is under recall by APP Pharmaceuticals. APP Pharmaceuticals is recalling the heparin because of incomplete documentation associated with test results. | Nonconforming Material/Component | Centurion Medical Products sent out Urgent Product Recall Notices on 6/3/2010. A second recall notice was sent to nonresponding consignees on 6/21/2010. The letter instructed the consignees to quarantine the recalled product and contact Centurion so that the recalled product can be retrieved. |  | 20 | 20 | Nationwide distribution. | Terminated | February 09, 2011 | 252 | 4080
Z-1094-2011 | 57202 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald, CELL-DYN Emerald CPU Board Hematology System; |  |  | October 15, 2010 | 2010 | Abbott Laboratories | There is the potential for the instrument to not power up due to the flash memory issue. This will result in an error message displayed is "No Memory Available" and has the potential for delay in generating patient results. | Other | The firm, Abbott Laboratories, sent a "Product Correction" letter dated October 15, 2010 to all customers. The letter described the product, problem and action to be taken by the customer. The customers were instructed to install a new printer driver in the software location using the provided instructions and to complete and return the Customer Reply Form via fax 1-800-777-0051 or email to QAGCO@abbott.com even if they no longer have the instrument.If you have any questions regarding this information, U.S. customers should call Customer Support at 1-877-4ABBOTT; customers outside the U.S., please contact your local hematology customer support representative. |  |  |  | Worldwide distribution: USA including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, O, OR, PA, RI, SC, TN, TX, UT, VA, WA, and WI; and countries including: Argentina, Australia, Brazil, Canada, Dominican Republic, Germany, Singapore, South Korea, St. Eustalius &amp; St. Martin, Trinidad and Tobago. | Terminated | June 07, 2011 | 235
Z-1061-2011 | 57651 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMedimplantable drug pump |  |  | January 14, 2011 | 2011 | Medtronic, Inc. - Neuromodulation | This letter provides important reminders concerning the potential for a pocket fill during a SynchroMedÂż II or SynchroMed EL implantable drug pump refill procedure, and important patient management recommendations that will be added to our product labeling. A pocket fill is the inadvertent injection of all or some of the prescribed drug into the patients subcutaneous tissue, which includes the pum | Other | The firm, Medtronic, sent an "Urgent Medical Device Correction Cover " letter dated January 2011 to all customers. Enclosed with each letter, an "Urgent Medical Device Correction Important Clinical Information about Pocket Fills" letter dated January 2011 as well as a Clinician Refill Reference Card, Refill Appointment Cards, Physician Reply Card and a Self Addressed Stamped Envelope. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to always check the needle placement within the pump septum throughout the procedure to help perform a successful refill; be aware that the pump refill difficulty may increase due to a variety of factors (listed); monitor the patient closely for signs and symptoms of overdose according to the provided instructions. Note: Medtronic is in the process of supplementing the labeling related to improper injection with the information included in this communication.Patients with questions are encouraged to talk with their physician or contact Medtronic Patient Services at 1-800-510-6735, Monday - Friday, 8 a.m. to 5 p.m. CDT. Physicians with medical questions related to this issue or Medtronic therapies should contact Medtronic at 1-800-328-0810, Monday - Friday, 8 a.m. to 5 p.m. CDT. UPDATE - June 2013 - Medtronic Field Representative began notifying Healthcare Professionals June 3,2013,via visit to deliver and discuss the Pump Refill Procedure Safety Update letter. The letter informs clinicians that SynchroMed Implantable Infusion System labeling and the Clinician Refill Reference Card have been updated. The labeling and reference card were updated to provide important reminders concerning the potential for a pocket fill during a SynchroMed II or SynchroMed EL implantable drug pump refill procedure. The Medtronic Rep will complete a Physician Notification Confirmation form documenting the delivery of the information. A copy of the recall notification letter a |  | 91369 | 91,369 total for all devices. | Worldwide distribution: USA including DC and PR and countries including: AUSTRALIA, NEW ZEALAND, HONG KONG, TAIWAN, CHINA, SINGAPORE, MALAYSIA, INDONESIA, SOUTH KOREA, THAILAND, ARUBA, AUSTRIA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, FAROE ISLANDS, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HUNGARY, ICELAND, INDIA, IRELAND, ISRAEL, ITALY, JAMAICA, JORDAN, KUWAIT, LEBANON, LITHUANIA, LUXEMBOURG, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NETHERLANDS ANTI LES, NORWAY, PANAMA, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, SAN MARINO, SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VATICAN CITY. | Terminated | February 04, 2015 | 1482 | 91369
Z-1222-2011 | 56860 | No | FZT | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet HEARTSTRING Aortic Cutter 4.3 mm, AC-3043, Lot numbers 9081071 and 9082071, Manufactured by Maquet Cardiovascular, Wayne, New Jersey. |  |  | June 04, 2010 | 2010 | Maquet Cardiovascular | Pouch integrity may be compromised, resulting in loss of sterility. | Process control | Urgent Device Removal letters were sent out on June 10, 2020 by Federal Express. The letter identified the affected product, described the reason for the recall, and asked for all affected product to be returned. The firm asked customers to examine their inventory and to discontinue distributing the affected lots. In addition, customers are to complete and return the Field Action Response form. If the affected product was further distributed, those customers also need to be notified of the recall and return their affected product to their distributor. All affected returned products may be exchanged with other unaffected HEARTSTRING Aortic Cutter 4.3 MM devices. Customers are to contact their local MAQUET Cardiovascular sales representative or Customer Service at 1-888-880-2874 to arrange for an exchange and if they have any questions. |  | 395 | 395 units | Worldwide Distribution -- USA, including states of IA, KY, IL, WI, TX and NC and countries of Japan, Germany and Hong Kong. | Terminated | June 21, 2011 | 382 | 395
Z-1221-2011 | 57733 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom TotalCare Bariatric Bed, P1840, Hill-Rom, Batesville, IN. |  |  | February 03, 2011 | 2011 | Hill-Rom, Inc. | The firm received one report that a patient was weighed on the bed at 53 Kg when the patient actually weighed 73 Kg. The patient expired the next day and it is inconclusive as to whether the scale was a contributing factor. The firm investigated and found that the software for the scale system did not clear the last patient weight from the history of the scale system. | Software design | The firm, Hill-Rom, sent a "URGENT MEDICAL DEVICE CORRECTION" letter dated January 19, 2011, along with a software upgrade (USB stick/flash drive) with instructions, to all affected customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to requests that facility service personnel immediately upgrade their software; complete and return the Customer Response Form via fax to 812-931-2410 or mail to: Hill-Rom Company, 1069 State Route 46 E, Batesville, IN 47006, Attn: MOD Coordinator J12 and forward a copy of this notice to any other facility personnel. Note: The customer can request that Hill-Rom technicians complete the software upgrade.If you have any questions concerning this Urgent Medical Device Correction, please contact Hill-Rom Technical Support at 800-445-3720. |  | 40 | 40 beds. | Worldwide distribution: USA and countries including: AUSTRALIA, BRAZIL, CANADA, CHINA, EGYPT, GERMANY, GREECE, ITALY, JAPAN, KUWAIT, MALAYSIA, MEXICO, NETHERLANDS, OMAN, POLAND, QATAR, SAUDI ARABIA, SPAIN, SWEDEN, THAILAND, UNITED KINGDOM, VENEZUELA, and YEMEN. | Terminated | September 26, 2012 | 601 | 945
Z-1224-2011 | 57597 | No | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Pump in Style Advanced (PNSA)Breast Pump; a daily use, double pump, electric powered breast pump; Manufactured exclusively by and for Medela, Inc., P.O. Box 660, 1101 Corporate Drive, McHenry, IL 60051-0660; the affected breast pumps were sold in the following configurations: |  |  | January 10, 2011 | 2011 | Medela Inc | The performance of the pumps may be compromised due to the potentiometer used on the circuit board. Affected pumps may not turn on, may turn off on its own, may reduce or lose suction during the pumping process, or may make a clicking noise. | Vendor change control | This firm, Medela, sent an "Important Device Recall" letter dated January 11, 2011 to all of direct accounts via UPS or US Mail. The accounts were informed that a component on the control board, after a period of usage, may cause the motor to slow down or stop, resulting one or more of the following symptoms: the pump will not turn on; the pump turns off on its own; the pump reduces or loses suction during the pumping process or the pump is making a clicking noise. If the users have a pump that exhibits any of these symptoms, they were requested to visit Medela's website at www.medelareturns.com to expedite shipment of a replacement unit. Questions about the recall process were directed to Medela Customer Service at 1-800-435-8316 or e-mail at medelareturns@medela.com.Distributor/retailers were requested to post the recall notice within their establishment visible to their customer base, and to notify their customers of the recall. Information on the recall was posted on Medela's website www.medelareturns.com on January 11, 2011. |  | 11154 | 11,154 units | Nationwide, including Puerto Rico | Terminated | May 21, 2012 | 497 | 11154
Z-1093-2011 | 57405 | No | KQG | Coagulation instrument. | Hematology | Hematology | 510(k) | Hemochron Whole Blood Coagulation Systems Activated Clotting Time (ACT); Package label states -International Technidyne Corporation Response Celite ACT Tube HRFTCA510; Lot number and expiration date are listed. Each box contains 95 ACT tubes tests. |  |  | November 30, 2010 | 2010 | International Technidyne Corp. | Some Activated Clotting Time (ACT) Tubes may contain a component that does not meet its specification, making it possible for a failure to detect a tube or the presence of a clot. | Process control | ITC notified their distributors and end user customers with an Urgent Medical Device Recall letter dated November 29, 2010 through UPS overnight mail. The letter identified the affected product and explained the reason for the recall. Distributors and customers should check their inventory for any affected product. Distributors should immediately discontinue shipment of the affected product and place the product on hold. Then regardless of whether they have any remaining inventory, they should complete the enclosed Distributor Account Tracking Form. The firm will contact them to arrange for any affected product to be returned and will credit them for any unused, returned product. Customers are to stop using the affected product and remove it from their inventory. They should also complete the attached Customer Account Tracking Form and return it to the firm. The firm will contact them to arrange for the return of any affected inventory and will credit the customers for any unused, returned product. If customers believe that they have experienced a problem with the product, they should take any additional actions regarding their records or patients. Questions should be directed to ITC Nexus Dx Technical Support at 800-631-5945 (US) or 732-548-5700 (International) or e-mail techsupport@itcmed.com. |  | 265 | 265 boxes (95 ACT Tubes per box) | Worldwide Distribution | Terminated | August 22, 2012 | 631 | 265
Z-1095-2011 | 57660 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Femoral Distal Fixation Peg; Catalog Number 5575-X-00; Manufactured by Stryker Howmedica Osteonics Corp., Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430 USA A Subsidiary of Stryker Corp. |  |  | December 22, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that the labels on the Triathlon Distal Femoral Fixation Pegs contain an incorrect expiration date of January 2015 (2015-1) when the actual expiration date should read June 2014 (2014-6). | Incorrect or no expiration date | The firm, Stryker, sent a "UPDATED (Corrected Lot Code) URGENT PRODUCT RECALL" letter dated January 6, 2011, via Fed Ex ,with return receipt to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to fax the attached Product Recall Acknowledgement form to (201) 831-6069; contact the hospitals in their territory that have the affected product to arrange return of the product, and return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100 or the Associate Product Manager, Knee Marketing Team at (201) 831-5028. |  | 15 | 15 units | Nationwide distribution. | Terminated | October 03, 2012 | 651 | 15
Z-1219-2011 | 57730 | No | MIE |  |  | Cardiovascular | Contact ODE | Pulmonary Valve &amp; Conduit SG |  |  | January 07, 2011 | 2011 | CryoLife, Inc. | A pre-processing culture performed on a companion allograft (Aortic Valve &amp; Conduct SG) tested positive for Staphylococcus aureus. | Other | The firm, Cryolife, sent an "Urgent-Human Tissue Recall" letter dated January 7, 2011 to the consignee/customer. The letter describe the product, problem and action to be taken by the customer. The customer was instructed to determine if any additional patient surveillance or intervention was necessary, and to complete and return the enclosed postcard affirming that the notification was received.If you have any questions, please contact Cryolife's Medical Director at 1-678-290-4403. |  | 1 | One | Nationwide distribution: USA state of: MI | Terminated | November 01, 2011 | 298 | 1
Z-1236-2011 | 57505 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker 5190 TPS Series Burs. Powered simple cranial drills, burrs, trephines, and their accessories. |  |  | November 30, 2010 | 2010 | Stryker Instruments Division of Stryker Corporation | Product was packaged with an incorrect packaging material and therefore has a potential to be non-sterile which, if contaminated, could lead to a patient infection. | Packaging | The firm, Stryker, sent a "MEDICAL DEVICE RECALL NOTIFICATION" letter dated December 2, 2010 to all customers. The letter described the product, problem and the action to be taken by the customers. The customer were asked to complete and fax in the business reply form indicating they have received the notification and to indicate whether or not they have affected product on hand, and send back any affected product to Regulatory Affairs at Stryker Instruments. Stryker will then send a replacement product to the customer.Upon receipt of the completed business reply form, the customer will be contacted to arrange return of any affected product they have in their inventory. Once the affected product is received, a replacement 5190 TPS Series Bur will be sent to the customer.For questions regarding this recall, please contact Stryker Instruments at 1-800-800-4236 ext. 3584 or kathleen.kelly.smith@stryker.com. |  |  |  | Worldwide distribution: USA including states of: CA, FL, LA, MI, MT, NE, NH, OH, PA, SD, TN, TX, VT, and WA; and countries including: Sweden and Australia. | Terminated | April 13, 2011 | 134
Z-1231-2011 | 57550 | No | GYK | Shunt system implantation instrument. | Neurology | Neurology | 510(K) Exempt | Integra Peritoneal Introducer Sheaths, 61 cm;Integra NeuroSciences/REF 901-124; Manufacturer Integra NeuroSciences Implants S.A., 2905 Route Des Dolines, 06921 SOPHIA ANTIPOLIS Cedex, France; Distributed by Integra NeuroSciences, 311 Enterprise Drive, Plainsboro, NJ 08536, USA |  |  | December 22, 2010 | 2010 | Integra LifeSciences Corp. | Peritoneal Introducer Sheaths were incorrectly labeled as 61 cm instead of 46 cm. | Labeling mix-ups | Integra Life Sciences sent an Urgent Product Recall letter dated December 22, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and quarantine any of the affected product. Contact Integra Customer Service at (800) 654-2873, select option 2 to arrange for product returns. Customers were asked to complete the attached "Recall Acknowledgement and Return Form" indicating the current status of quantities of affected product in their inventory. For questions call (609) 936-2237. |  | 95 | 95 units | Nationwide Distribution including CO, FL, GA, MO, NC, NJ, PA, TX, UT, and WI. | Terminated | May 07, 2012 | 502 | 95
Z-1228-2011 | 57679 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Thomas Medical Products, Acuity Bleed back Control Valve, Acuity Rotating Hemostasis Valve, Acuity Universal Cutter, and Contamination Sleeve for use during vascular interventional procedures. Contamination Sleeve part no. CON-006-02; Acuity Bleedback Control Valve part no. CSP-001-00; Acuity Rotating Hemostasis Valve part no. CSP-002-01; Acuity Universal Cutter part no. CSP-003-00; Acuity Universal Cutter part no. CSP-003-01. |  |  | December 16, 2010 | 2010 | Thomas Medical Products Inc | Compromised sterility/weak or open seals. | Packaging process control | GE Healthcare Thomas medical products Recall notification dated 16 December 2010 (GEHC Ref # 35003) was sent by letter via UPS Next Day Air and/or USPS certified mail. |  | 28679 | 28,679 for all products | Distribution Nationwide USA and Germany. | Terminated | August 08, 2012 | 601 | 28679
Z-1230-2011 | 57735 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Bipolar cord for use with the VirtuoSaph Endoscopic Vein Harvesting System, MCBICORD1, Terumo Cardiovascular System, Ann Arbor, MI. |  |  | January 10, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The Bipolar cord product specification indicates the product has a service life of 50 uses. However the IFU does not specify an end of service life. The firm has recieved 46 complaints on the Bipolar cord. The firm investigated and found that the units involved in the complaints were often over 12 months old. The firm believes the units involved in the complaints were beyond their product life | Other | Terumo Cardiovascular Systems Corporation sent an URGENT MEDICAL DEVICE CORRECTION letter dated January 10, 2011, to Operating room managers. The letter informed the customers of the Bipolar cord product life of 50 uses. The letter states that the IFU for the bipolar cord is being updated to include the information regarding the limit of 50 uses and sterlization cycles. The letter recommends that user keep a sterile back up bi-polar cord in a convienent place during endoscopic vein harvesting procedures. The firm provided a tracking sheet in each package to help users track the number of sterlization cycles for the cord. Customers were instructed to review the Medical Device Correction notice, assure that all users were aware of the notice, and confirm receipt of the communicaiton by faxing the attached Customer Response Form to the fax number indicated on the form. For questions regarding this recall call Terumo CVS Customer Service at 1-800-521-2818. |  | 2782 | 2782 cords | Worldwide Distribution - United States including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, KS, KY, LA, MA, MD, ME, MI, , NY, MO, MS, MT, NC, NE, NH, NJ, NM, OH, OK, OR, PA, SC, TN, TX, VA, WI, AND WV and the countries of Chile, Hong Kong, Malaysia, United Arab Emirates, Belgium, Japan, Korea, The Phillipines, Singapore, Taiwan, Thailand, Germany, and United Kingdom. | Terminated | July 13, 2011 | 184 | 2782
Z-1223-2011 | 57607 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell GammaPlan, Model 5.34. |  |  | October 01, 2005 | 2005 | Elekta, Inc. | Investigation found that if a user accidentally selects the wrong image/tube position images can be displayed flipped. | Other | Product bulletin Leksell GammaPlan - "Leksell GammaPlan version 5.34" (dated October 2005) was sent to affected Customers as notification of the problem. Field Change Order FCO-DCO-05-0006 (dated November 22, 2005) and Technical Notes Leksell GammaPlan Leksell SurgiPlan Important Notice - "Potential issue using flipped angiographic images" (dated October 2005) was sent to Elekta Service Engineers as notification of the problem. The letters described the issue and the FCO asked the Service Engineers to confirm their receipt of the documentation. The Technical Notes also discussed image visualization requirements and listed recommendations. |  | 272 | 272 units | Worldwide Distribution | Terminated | March 02, 2011 | 1978 | 272
Z-1229-2011 | 56058 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AMS brand GreenLight PV Laser System, Model-P/N: 0010-9230; |  |  | February 08, 2010 | 2010 | AMS Innovative Center - San Jose | A failure mode whereby the fiber attachment to the GreenLight PV laser system could inadvertently be broken. If fiber were to bend sharply at point of fiber connection, it may break and cause a burn. | Other | AMS Innovative Center - San Jose issued Important Product Update Information letters via express mail on February 8, 2010, and September 15, 2010 to all its direct consignees. The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to attach the Fiber Guard to their GreenLight PV Laser System, and to complete the Response Form in the enclosed, self-addressed, postage paid envelope to AMS. For any questions regarding this recall call the Customer Care Department at 1-800-356-7600 or send an e-mail to CustomerCare@ammd.com. |  | 242000 | ~242,000 units | Worldwide Distribution including US | Terminated | February 14, 2011 | 371 | 242000
Z-1258-2011 | 51399 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle Blood Glucose Monitoring System Test Strips and Freestyle Lite Blood Glucose Monitoring System Test Strips, for use with FreeStyle and FreeStyle Lite Test Strips respectively, 100 count, manufactured by Abbott Diabetes Care, Alameda, CA. |  |  | March 13, 2009 | 2009 | Abbott Diabetes Care, Inc. | Product dispositioned for destruction was potentially diverted to end users. | Other | The firm, Abbott Diabetes Care (ADC), will be posting notification on its Abbott Diabetes Care website. Also, the firm issued an "IMPORTANT: MEDICAL DEVICE NOTIFICATION" letter dated March 13, 2009, to customers ( registered users of Free Style, Free Style Flash, FreeStyle Freedom or Free Style Navigator). For foreign consignees/customers, the firm has issued communication worldwide. The letter described the product, problem and action to be taken by the customers. The customers were instructed to please check the lot number on the test strip carton or vial for the affected product number; if found, discontinue use and call the ADC Customer Care line at: 1-800-310-1500 and the ADC representative will provide information on how to return the test strips and arrange to have replacement test strips sent.If you have any questions, please call ADC Customer Care line at 1-800-310-1500. |  | 1860 | 1860 units | The firm, ADC, quarantined the eight lots and disposed them to scrap. These eight lots were not released from the firm ADC's control; therefore, the firm was unable to determine the number of consignees. | Terminated | February 15, 2011 | 704 | 1860
Z-1259-2011 | 54746 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio-Rad Variant II Turbo Link Hemoglobin Testing System, including Hemoglobin A, Program Reorder Pack 1600 Tests, Program Buffer A, Buffer B, Cartridge Set, Analytical Cartridge, Guard Cartridge, Calibrator, Accessory Box, Instruction Manual and Quick Guide |  |  | February 19, 2010 | 2010 | Bio-Rad Laboratories Inc | The results from the blood device had the occurrence of some ramping baselines on their chromatograms. The ramping baseline can affect quantitation of the blood results. | Nonconforming Material/Component | The firm, Bio-Rad Laboratories, Inc., sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated February 5, 2010, to all customers and letter dated February 16, 2010 to subsidiaries via Fed-Ex or by fax. The letter described the product, problem and action to be taken by the customers. The customers were instructed to review there chromatograms form the lots listed and compare them to what documented; follow the noted instructions; destroy reorder packs from the affected lots if a ramping baseline issue is observed, and complete and return the attached CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM via fax to (510) 741-3954, attention: Bio-Rad CSD Regulatory Affairs Department. If you have any questions, please contact your regional Bio-Rad Office or (510)724-7000. |  | 27 | 27 | Worldwide distribution: USA including states of: OH, MN, SC and VA; and country including: Japan. | Terminated | March 03, 2011 | 377 | 27
Z-1249-2011 | 56182 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Radiotherapy delivery system (C-Series Clinac, Trilogy and Novalis Tx), model numbers H14, H17, H18, H23, H26, H29, HCX |  |  | June 23, 2010 | 2010 | Varian Medical Systems Oncology Systems | Position errors may result in misdelivery of radiation. | Device Design | Varian sent Urgent Medical Device Correction letters on June 23, 2010 by Federal Express to all customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***Field Safety notification including customer distribution lists, was e-mailed to Varian international Customer Support Managers. Managers will translate to local language if necessary, notify sites by appropriate methods. |  | 5741 | 5741 units | Worldwide distribution. | Terminated | May 03, 2011 | 314 | 5741
Z-1248-2011 | 57485 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex Dental Systems, Gendex eXpert DC Intraoral X-Ray System; a dental x-ray system; Model 110-0205G1, catalog number EXPERTDC75NA |  |  | December 17, 2010 | 2010 | Gendex Dental Systems | The intraoral x-ray systems were installed improperly with incorrect screws which allow the unit to pull away from the wall when the arm is fully extended. This condition may harm the patient. | Component design/selection | Gendex sent a Urgent Medical Device Field Correction letter dated December 17, 2010 to the consignee, informing them that the x-ray units were installed with improper sized lag bolts and requesting that they contact Gendex Customer Service at 1-800-323-8029 to complete the field correction. |  | 5 | 5 units | California | Terminated | February 15, 2011 | 60 | 5
Z-1246-2011 | 57587 | No | DZD | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | ClassOne Orthodontics Buccal SS Tubes Label ID Reference Number 351-327 (NF DB UR 6 Convertible .018-14T 14DO HK-Each); |  |  | November 01, 2010 | 2010 | Ortho Organizers Inc | Products were mis-labeled as nickel free. | Pending | An Urgent: Product Recall Notice letter, dated November 1, 2010 was sent to consignees at the distribution level who received recalled products. Consignees were instructed to do the following:1) Carry out a physical count of current inventory and record this data on the business reply card provided with this notice. Return the enclosed Reply Card with the amount remaining inventory to Ortho Organizers, Inc., Carlsbad, CA 92008 within 10 days of receipt of this notification.2) Immediately cease distribution/use of this product and remove from your shelves.3) Notify all of your retail customers with the Urgent: Product Recall Notice (Retail Level) letter who purchased the affected item numbers from you that the affected items are stainless steel and contain nickel.4) Contact your sales representative to return any unused items. Return any unused product within 30 days. Please clearly mark the return carton Recall Material Enclosed. Ortho Organizers shall accept any product covered by this recall and credit your account for all product and shipping costs. |  | 1300 | 1300 | Worldwide Distribution -- USA, including states of CA, CO, IL, LA, MI, MN, PA &amp; TX and countries of CA (Canada), BE (Belgium), NO (Norway), and CO (Columbia). | Terminated | April 27, 2011 | 177 | 1300
Z-1255-2011 | 57626 | No | BYL | Nonpowered oxygen tent. | Anesthesiology | Anesthesiology | 510(K) Exempt | Face Tent |  |  | January 07, 2011 | 2011 | Convatec Inc. | Face tents and Oxygen masks may have a sticky glue-like substance on the inside or outside of the mask consisting of dioctyl phthalate, epoxidized soybean oil and glycerol monoleate. | Nonconforming Material/Component | Unomedical, Inc issued recall letters with attached return response forms dated January 7, 2011 via UPS. The letters request the customer stop use and quarantine product. The recall response form is to be completed and returned to Unomedical.The letters/return response form were also sent out in Spanish.Another recall letter was sent dated February 16, 2011 to capture additional lots that werent included in the initial letter. |  | 1579100 | 1579100 pcs. | Worldwide distribution, including USA, Puerto Rico, Canada, Mexico, Colombia, El Salvador, Paraguay, UK, Austria, Belgium, Cyprus, Czech Republic, Denmark, Estonia, Faroe Islands, Finland, France, Germany, Hungary, Iceland, Ireland, Italy, Netherlands, Norway, Poland, Spain, Sweden and Switzerland, Israel, Jordan, Lebanon, Libya, Morocco, Saudi Arabia, UnitedArab Emirates, Russia, Australia, New Zealand,Japan and Vietnam. | Terminated | February 27, 2013 | 782 | 2692300
Z-1253-2011 | 57592 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | FlexicairÂż II (Model number 13000) Low Airloss Therapy Units |  |  | December 06, 2010 | 2010 | Hill-Rom Manufacturing, Inc. | Due to product age and parts availability the units are being withdrawn to reduce a potential health risk to patients when bed rails are not used. | Other | Hill-Rom Manufacturing, Inc. sent an Urgent Medical Device Withdrawal letter dated December 15, 2010, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. The FLEXICAR MC3 AND FLEXICAIR 13000 were no longer available for rental as of November 15, 2010. Customers in Michigan were required to switch to another product by December 31, 2010. All other customers were required to switch to another product by March 4, 2011. Customers advised to continue the use of the product until arrangements could be made to replace the product. were notified by letter on/about 12/06/2010. The letter indicated the units will no longer be available for rental as of 11/15/2010. Patients in Michigan currently using the products will be required to switch to another product by 12/31/2010. All other patients using the products will be required to switch to another product by 03/04/2010. They were advised that a Hill-Rom representative will be contacted them to discuss details of the product withdrawal and possible alternative surfaces for their patients.For questions regarding this recall call 800-445-3720. |  | 1 | 1 Flexicair II unit | Nationwide Distribution including CA, FL, GA, IL, MI , MS, NJ, NY, OH and PA | Terminated | November 01, 2011 | 330 | 46
Z-1069-2011 | 57580 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | FlexTray Procedure Delivery System, Product Codes: KBW71SXL- Bariatric Surgery Trays, TBW13SXL- Bariatric Surgery Trays, XBB14B - Bariatric Surgery Trays, and XBB65S - Bariatric Surgery Trays. |  |  | March 27, 2009 | 2009 | Ethicon Endo-Surgery Inc | The firm initiated this second subsequent recall because the affected devices can remain closed and clamped on tissue during surgical procedures due to the fact that the mechanism which connects the articulation joint to the device shaft is insufficient. | Other | Ethicon Endo-Surgery issued "Urgent Expanded Device Recall Event 1981" notification letters dated April 7, 2009, to their customers. They provided a list of all Affected Product Codes of the ETS FLEX45 Articulating Endoscopic Linear Cutters (including the Product Code: ATS45 from the 1st recall) and Affected Flex Tray Codes subject to recall. The notification document reminds the Operating Room and Materials Management customers of the problem which necessitated the previous recall (Z-0361-2011). It goes on to instruct their customers to immediately discontinue use of the affected products. They are to complete and fax back the Business Reply Form, Affected product is to be returned with a copy of the completed Business Reply Form to Stericycle, Replacement product should be received within 30-60 days after return, Customers can contact Stericycle concerning this recall at 1-888-365-4680, reference Event 1981. |  | 135 | 135 kits | Nationwide Distribution -- USA including the following states, AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MO, MS, MT, NC, NE, NH, NJ, MN, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, and Puerto Rico. | Terminated | February 15, 2013 | 1421 | 29538
Z-1257-2011 | 57499 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba TSX-301A AQUILION ONE (MPDCT0297EA) - Whole Body X-Ray CT Scanner |  |  | September 07, 2010 | 2010 | Toshiba American Medical Systems Inc | Incorrect Dose Calculation. The software that performs the pre-scan dose calculation uses an incorrect parameter for an exam Plan using a tube voltage of 80kV or 100kV. This results in a pre-scan expected dose that is approximately 10% lower than what the displayed value should be when the scanner executes the exam Plan. | Software design | Toshiba America Medical Systems (TAMS) issued notification to consignees 9/7/2010, via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. |  | 396 | 44 units distributed between 09/2007 - 07/2010 | Nationwide distribution. | Terminated | April 20, 2012 | 591 | 396
Z-1251-2011 | 57604 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Elekta Leksell Stereotactic System, used for localization (spatial reference) for cranial surgery using X-ray or CT and MRI Image data. |  |  | October 02, 2009 | 2009 | Elekta, Inc. | New protocols for MR sequences may result in higher RF energies deposited during MR scanning, generating heat in uninsulated fixation posts. | Labeling design | Elekta sent Important Notice Affected Product: Fixation Posts, A331, dated June 22, 2009. The Important Notice will be added to the Customer User manual upon receipt. |  | 379 | 379 units | Nationwide | Terminated | February 17, 2011 | 503 | 379
Z-1247-2011 | 57665 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Mantis Tap Sleeve,Mfg by: Stryker Spine, 2 Pearl Court, Allendale, NJ 07401-1677. packaged in the following products: |  |  | January 07, 2011 | 2011 | Stryker Spine | Stryker Spine has received reports that after sterilization but prior to surgery, cracks were noticed at the tantalum marker / PEI tube interface. This can likely be attributed to differences in the rate of thermal expansion between the markers and the sleeve during the sterilization cycle. | Other | Stryker Spine sent an Urgent Product Recall letter dated January 07, 2011, to Stryker Spine sales branches and agencies where affected product was distributed. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and hospital locations to identify the product.Reconcile any Mantis Tap Sleeves with the catalog and lot numbers referenced in the letter. Complete the Customer Response Form with distribution sheet provided to assist with product reconcilation. Customers were asked to fax a copy of the Customer Response Form to Regulatory Compliance at 201-760-8370.Customers were to retrieve and return the affected product using the enclosed pre-paid mailing label to Regulatory Compliance, Stryker Spine 6 Pearl Court, Allendale, New Jersey 07401. For questions call 201-760-8180. |  | 25 | 25 units | Nationwide Distribution including MA, MN, and PA. | Terminated | March 08, 2015 | 1521 | 25
Z-1252-2011 | 57731 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI |  |  | January 05, 2011 | 2011 | Elekta, Inc. | Flexmap calibration did not cover the full 360 degrees of gantry rotation (0 to 355 degrees) and that this caused an isocenter displacement with 2D PlanarView images required at 180 to -180. | Other | The firm, Elekta Limited, sent an "Important Notice A344" dated January 5, 2011 to all consignees/customers. The notice described the product, problem and action to be taken by the customers. The customers were instructed to follow any instructions or recommendations covered in the notice. The notice provides suggestions on how to check and prevent this situation from occurring; it instructs the User to file the document in the Important Notice section of the appropriate User Manual; to complete a FlexMap Check each time they acquire a new FlexMap and amend their local procedures to include the FlexMap Check and complete and return the Important Notice Confirmation of Receipt via Fax to: Field Support Administration, Elekta Limited, Linac House, Fleming Way Crawley, West Sussex, RH10 9RR UK +44 (0)1293 654401.If you have any questions, contact Elekta Limited at +44 (0)1293 654200 or www.elekta.com. |  | 223 | 223 units | Nationwide distribution. | Terminated | April 07, 2011 | 92 | 223
Z-1250-2011 | 57159 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Xeleris Processing and Review Workstation 1.0, 1.1, 2.0, 3.0 |  |  | January 12, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of safety issues relevant for the Xeleris Workstation 1.0, 1.1, 2.0, 2.1, 3.0:1. Measurements taken on the CT Images on the CT system are different from the Xeleris Workstation2. Inaccurate PET SUV values may be calculated when data is retrieved from PACS after being processed in the Xeleris.3. Wrong marking of Early and Late result series on display in Bra | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 12, 2011 to all consignees/customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to please ensure that all potential users in their facility are made aware of this safety notification and the recommended actions. A GE Healthcare Service representative will perform the required software update on each affected system.If you have any questions or concerns regarding this notification, contact your local GE Healthcare Service Representative. You may also contact GE Healthcare call center in US or Canada at 800-437-1171 and other countries at 262-896-2890. |  | 6046 | 6046 | Worldwide distribution: USA including DC, PR, and Guam; and countries including: AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA, AZERBAIJAN, BELGIUM, BOSNIA &amp; HERZEGOVIA, BULGARIA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KENYA, KUWAIT, LUXEMBOURG, LEBANON, MALY ASIA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PHILIPPINES, PORTUGAL POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SPAIN, SOUTH AFRICA, SOUTH KOREA, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, and VIET NAM. | Terminated | October 16, 2012 | 643 | 6046
Z-1245-2011 | 57379 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory software; GE Healthcare IT, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. |  |  | December 03, 2010 | 2010 | Ge Healthcare It | The GE Centricity Laboratory System Software has safety issues with the "Telephone List" function, the "Anatomic Pathology" application, and the Instrument Interface IO module that may impact patient safety. | Software design | GE Healthcare, sent "Urgent Medical Device Correction" letters dated December 12, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the product, the problem, and the action to be taken by the customer. The letters provided the users with specific safety instructions to follow until the software was updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. |  | 26 | 26 units | Worldwide Distribution -- USA, including AR, AZ, CA, CO, DC, DE, FL, GA, HI, IA, IL, KY, MA, MI, MN, MO, NC, NH, NJ, NY, OR, TN, VA, and WI, , and internationally to the countries of Australia, England, Hong Kong, India, Malaysia, New Zealand, Qatar, Scotland and Singapore. | Terminated | July 15, 2011 | 224 | 26
Z-1334-2011 | 50094 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | S T-Bone Plate; |  |  | May 26, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | A package labeled as a S T-Bone Plate, Catalog Number 57-05260, Lot Code G4M00F1576, may actually contain a M T-Compression Plate, Catalog Number 57-10250, Lot Code G4N00F1584 and vice versa. | Process design | Stryker Orthopedics sent an Important Market Withdrawal letter dated May 26, 2006, by Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to examine their inventory, including their hospital consignment location, to identify the product.If the affected product remained in its original packaging, customers were instructed to retrieve all of the affected product lots and return it to their branch or agency warehouse for reconcillation.Reconcile all product on the attached Product Accountability Form and fax a copy of the form to 201-831-6069 within 5 days of receipt of the notice.Customers were asked to return the affected product via Inter-Org Transfer and ship to Product Surveillance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, USA 07430-2104. Customers were also instructed to contact their Stryker Orthopaedics Customer Service representataive to re-order the product that they were returning.For questions regarding this recall call (201) 831-5825. |  | 21 | 21 | Nationwide Distribution | Terminated | December 08, 2014 | 3118 | 21
Z-1320-2011 | 57295 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | WalkMed Infusion Triton 90" 20 drop IV Administration Set with Needlefree Y-site |  |  | November 16, 2010 | 2010 | Churchill Medical Systems, Inc. | Routine testing conducted to challenge the success of sterilization was not fully completed for sterile products. | Vendor change control | The firm, Churchill Medical/Vygon, sent an "URGENT MEDICAL DEVICE RECALL" letter dated November 16, 2010, via overnight mail to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to check their stock for affected product; immediately cease use and distribution and quaratine the affected product; contact Vygon's Customer Service Department at 1-800-473-5414 for a return Merchandise Authorization number; complete and return the Recall Acknowledgement and Inventory Return Form via fax to 215-390-2019 and mail return product with form to: Vygon, 103A Park Drive, Montgomeryville, PA 18936, and if the customer is a distributor, they were to provide their customers with a copy of this notification and attached form. Note: An additional follow up communication will be provided to the customers after reconciliation of the returned inventory and all inventory will be credited or replaced.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by email at customerservice@vygonus.com. |  | 12150 | 12150 units | Nationwide Distribution: USA including states of: CO, IA, OH, MA, MD, MO, NY, FL, SC, FL, and VA. | Terminated | May 01, 2012 | 532 | 80232
Z-1341-2011 | 57454 | No | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | Celite ACT Cartridge |  |  | December 03, 2010 | 2010 | Abbott Point Of Care Inc. | Celite ACT cartridges have increased numbers of values outside the accepted allowable error. | Device Design | On behalf of Abbott, Genco Pharmaceutical Services sent an "URGENT RECALL NOTICE" dated November 2010, on December 3, 2010 via Fedex to all consignees/customers. Abbott sent out notification packages to foreign consignees/customers on December 3, 2010. The notice described the product, problem and action to be taken by customers. The firm recommended that customers use cartridges within 12 hours of removal from refrigerated storage based on the instructions provided with the product. The customers were instructed to complete and return the Business Reply cards to acknowledge receipt of the communication and provide the letter to any customers that have recieved the product. If you have any additional questions, please contact Abbott Point of Care Technical Support at 800-366-8020, option 1 or your Abbott Point of Care Representative.On February 27, 2012 Abbott sent an Important Product Information letter via Fed Ex to inform customers that i-Stat Celite ACT cartridges manufactured from the beginning of 2012 can now be stored for two weeks at room temperature as indicated on the cartridge box. Foreign consignees were notified on 2/28/2012. Users are instructed to complete business reply cards to acknowledge receipt of the communication and provide the letter to any customers that have received the product. Any additional questions can be answered by contacting 8003668020 option 1. |  | 1164275 | 1164275 cartridges | Worldwide distribution. | Terminated | July 16, 2013 | 956 | 1164275
Z-1261-2011 | 57318 | No | KTW | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 02.122.550, 50mm Spiral Blade, 14mm, Right - Synthes Spiral Blade Component of the Synthes Modular Blade Plate System |  |  | October 27, 2010 | 2010 | Synthes USA (HQ), Inc. | Potential for out of specification star grind teeth. | Other | Synthes initiated their recall on October 18, 2010, notifying affected hospital accounts via consultant communication. Affected Synthes Trauma Sales Consultants were notified via email (receipt requested). A copy of the Urgent Medical Device Recall letter dated October 20, 2010, was mailed via UPS second day air on October 20, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify any affected devices and call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. A replacement order would be taken at that time.Customers were also instructed to complete the Verification Section at the end of the letter by checking the appropriate box indicating the affected product had been located or indicating in the appropriate box that no affected product was located. Also, customers were to indicate the number of devices found and note the Return Authorization Number. Return the Verification Form with the recalled product to:Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.For questions regarding this recall call 1-800-620-7025, x 6883. |  | 19 | 19 | Nationwide Distribution including AL, CA, GA, IA, IL, IN, MA, MI, MS, NC, NJ, NV, OH, PA, SC, SD, TN, and WV. | Terminated | April 16, 2015 | 1632 | 561
Z-1340-2011 | 57792 | No | HMR | Ophthalmic surgical marker. | Ophthalmic | Ophthalmic | 510(K) Exempt | Walcott Rx Products, Surgical Division, Corneal Marking Pad, REF RX40490, Sterile. For use in ophthalmic surgery. |  |  | December 23, 2010 | 2010 | Aspen Surgical Products, Inc. | Fragments from the pad may disengage and stick to the instrument. The condition could cause fragments to come into contact with the intact cornea prior to surgery. | Component design/selection | Aspen Surgical sent URGENT MEDICAL DEVICE RECALL NOTICES dated 12/23/2010 by Certified Mail. The letters instructed the users to examine their inventory and place the recalled products in quarantine. Consignees were instructed to return the product and complete the response form. Distributors were asked to notify their customers or requesting that Aspen notify the customers.The firm does not intend to continue manufacturing or distributing the product. |  | 77850 | 77850 all products | Worldwide distribution including USA Canada, England, Sweden, Saudi Arabia, England, Hong Kong, Belgium and Israel. | Terminated | October 07, 2011 | 288 | 77850
Z-1362-2011 | 57544 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ConMed System 2450 Electrosurgical Generator, REF 60-2450-100, 100 V, ConMed Electrosurgery, Centennial, CO. |  |  | December 27, 2010 | 2010 | ConMed Electrosurgery | Electrosurgical generator may allow possibility of electrical shock and/or burns to user or patient. | Device Design | ConMed Electrosurgery sent an Urgent Device Recall letter dted December 27, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify anyone who may have received the affected product.. They were to review their inventory and return a Business Reply Form via fax to 303-699-9854 and to contact ConMed Electrosurgery Technical Services at 800-552-0138 to schedule a time to have their equipment upgraded. |  | 119 | 119 units | Worldwide Distribution, including USA (nationwide) and the countries of Australia, Belgium, Bolivia, Canada, Chile, China, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, Honduras, Hong Kong, Hungary, India, Iraq, Israel, Italy, Jamaica, Japan, Jordan, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, Uruguay, Venezuela, Vietnam. | Terminated | August 01, 2011 | 217 | 2323
Z-1346-2011 | 57287 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | ARTERIAL LINE KIT/Product SKU (Code) AK-04510 |  |  | October 29, 2010 | 2010 | Arrow International Inc | Sterility may be compromised. | Other | Arrow International, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated October 29, 2010, to all affected customers. The letter identified the products, the problem, and the action to be taken by the customer. Customers were instructed to quarantine the product.If they had no affected product in stock, check the appropriate box on the Recall Acknowledgement &amp; Stock Status Form and fax the form to Arrow's Customer Service fax number 800-343-2935.If they had affected stock, immediately quarantine the product and contact Arrow Customer Service at 800-523-8446 for a Return Authorization Number. If customers wished to return product for replacement, they were to consult their sales representative or Arrow Customer Service for information on product availability, substitutability, and lead times. Complete the enclosed Recall Acknowledgement &amp; Stock Status Form and immediately fax the compoleted form to Arrow, Fax number 800-343-2935, Attn: Customer Service.Return any affected product freight collect, along with the original completed Recall Acknowledgement &amp; Stock Status Form to:Arrow International, Inc.312 Commerce PlaceAheboro, NC 27203Attn: Asheboro Quality Assurance Manager.For any questions regarding this recall call 800-523-8446. |  | 361640 | 361,640 | Worldwide Distribution -- United States - all 50 States, and the countries of Canada, South America, China, South Africa, Australia, Puerto Rico, Philippines, Singapore, South Korea, New Zealand, Mexico, Japan, India, Indonesia, Bahamas, Taiwan, Thailand, Finland, France, Cyprus, Czech Republic, Belgium, Germany, Ireland, Italy, Latvia, Netherlands, Poland, Slovakia, Spain, Bahamas, Costa Rico, Dominican Republic, Ecuador, Guatamala and the United Kingdom. | Terminated | September 29, 2012 | 701 | 963380
Z-1342-2011 | 56387 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio Rad brand VARIANT II TURBO Hemoglobin A1c Program Reorder Pack, Model Number/Catalog No.: 270-2417, which contains the update kit CD-ROM, Model/Catalog No.: 270-2423, Distributed and Manufactured by: Bio-Rad Laboratories, Inc., Hercules, CA. |  |  | March 24, 2010 | 2010 | Bio-Rad Laboratories Inc | The firm received customer complaints associated with the Update Kit CD-ROM not being able to generate a valid calibration after being uploaded, resulting in patient results not being generated. | Other | Medical Device Correction letter dated March 24, 2010 was sent to customers. The letter identified the affected product and explained the reason for recall. Customers were asked to destroy Lot No. AA00190 and replace it with the Lot No. AA00838 CD-ROM that is enclosed with the letter. Customers are to complete and fax back the Medical Device Correction Response Form at the number provided. |  | 48 | 48 units | Worldwide Distribution -- USA and Italy. | Terminated | March 04, 2011 | 345 | 48
Z-1370-2011 | 57850 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | MitraXs" Self Adjusting Atrial Retractor, Sterile EO, Models (REF): A0030/1; A0031/1; A0032/1; A0033/1, Cardio Life Research s.a., Green Hill, Hoge Wei, 16, B-1930 Zaventem, Belgium. |  |  | January 26, 2011 | 2011 | St. Jude Medical Cardiovascular Division | St. Jude Medical Cardiovascular Division is conducting a voluntary recall of all batches and model numbers of the MitraXs Self Adjusting Atrial Retractor. St. Jude Medical has been distributing this product since their acquisition of CLR in late 2008.They have determined that some distributed batches have very small burrs on the polycarbonate material at the distal edge. As these burrs may | Process control | The firm, St. Jude Medical, sent an "Urgent Medical Device Recall Notice" dated January 26, 201,1 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to discontinue use of the product (MitraXs Self Adjusting Atrial Retractor batches) and remove from inventory. St. Jude began production of a new batch of units that will bear the labeling of St. Jude Medical. Shipment of these units started on January 10, 2011. Note: A St. Jude Medical Sales Representative will contact the customer to facilitate removal and reconciliation of the affected product.If you have any questions or concerns, please contact your St. Jude Medical Sales Representative or 651-490-4470 and/or 612-310-7936. |  | 450 | 450 | Worldwide distribution: USA including states of: NJ, OH, MI, and MN; and countries including: BELGIUM, FRANCE, GERMANY, SWITZERLAND, and UNITED KINGDOM. | Terminated | May 31, 2012 | 491 | 450
Z-1374-2011 | 56394 | No | LZS |  |  | Ophthalmic | PMA | AMO brand WaveScan WaveFront System, OPM software versions 3.67 and 3.67, |  |  | July 13, 2010 | 2010 | AMO Manufacturing USA, LLC | The failure to comply with Minimum System Requirements, increased system errors could result, and a miscreated treatment table (incorrect LASIK refractive treatment) could result. | Software design | Abbott Medical Optics, Inc. sent an Urgent Medical Device Field Correction letter to all affected customers starting July 15, 2010. The letter identified the product, the problem, and the action to be taken by the customer.Customers were asked to take the following actions:Surgeons were instructed to complete the facsimile attachment to acknowledge receipt of the notification letter and fax to 714-247-4530.For those using the affected product a 3.68 OPM disc must be requested by e-mailing a request to OPM@pacful.com.For any questions please contact AMO Technical Service Representatiave at 1-877-266-4543. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Great Britian, Greece, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Malaysia, Mexico, Netherland, New Zealand, Peru, Philippines, Portugal, Russia, Russia Federation, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Taiwan, ThailandChina, India, South Korea, Singapore, Taiwan, Thailand, Turkey, United Arab Emirates, Venezuela, and Vietnam. | Terminated | June 22, 2012 | 710
Z-1373-2011 | 57502 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Orthoralix 8500 DDE Panoramic Dental X-Ray System; Extraoral x-ray unit with timer; Gendex Dental Systems, Des Plaines, IL 60018; part number 110-0234G1 (domestic) and 110-0234G3 (international) |  |  | December 21, 2010 | 2010 | Gendex Dental Systems | The Automatic Exposure Control (AEC) feature located on the control panel of the Orthoralix 8500 Panoramic Dental X-Ray System was functional, but was subsequently disabled in units manufactured starting 5/17/10. North American units manufactured between 12/4/09 and 5/16/10 had functioning AEC buttons, but their manuals stated they were disabled. Units intended for International distribution bet | Labeling Change Control | Gendex sent a recall letter to their accounts on December 21, 2010, stating that the AEC button on the control panel was a non-functional feature that becomes illuminated upon activation and that Gendex would send them a kit to correct this issue and inactivate the button. A second letter was sent to the accounts on January 31, 2011, explaining that a corrective kit was no longer necessary, letting them know if the AEC feature was active or inactive for their system, and providing them with the appropriate user manual addendum for their system.Customers were asked to complete and return the required Acknowledgement Form by fax to 215-997-5665 or e-mail: regulatoryaffairs@gendex.com.For questions regarding this recall call 800-323-8029. |  | 574 | 574 systems | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, GA, CO, CT, FL, GA, ID, IL, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OK, PA, TN, TX, VA, WA, and WI and the countries of Australia, Canada, Germany, New Zealand, Taiwan, and Vietnam. | Terminated | November 09, 2011 | 323 | 574
Z-1260-2011 | 57567 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon SecureStrap 5 mm Absorbable Strap Fixation Device Product Code STRAP25. |  |  | December 29, 2010 | 2010 | Ethicon, Inc. | The Ethicon SecureStrap 5 mm Absorbable Strap Fixation Device may have the potential for the sterile barrier to be compromised due to a defect in the packaging. | Pending | Ethicon sent Urgent Voluntary Product Recall letters dated, December 28, 2010, to their customers on December 29, 2010 via UPS. The letter identified the affected product and described the reason for recall. Customers are to discontinue use and sampling immediately. They are to return the product using the enclosed prepaid shipping label along with the enclosed business reply card regardless of whether they have any affected product or not. Customers should contact 1-877-551-7147 or speak to their manager if they have questions regarding this recall. |  | 4110 | 4,110 | Worldwide Distribution -- USA and Canada. | Terminated | November 07, 2012 | 679 | 4110
Z-1372-2011 | 57736 | No | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medstorm Adult Radiolucent Zoll Medical (Part # 16381, HeartSync, Ann Arbor, MI. |  |  | December 17, 2010 | 2010 | Heart Sync, Inc | The seal on the electrode pouch may be open, exposing the electrode to the outside environment. The package does have a caution statement that states, " Do not open package until immediately prior to using electrodes. and Electrodes may dry out when removed from packaging and exposed to air. | Packaging | Heart Sync, Inc sent an URGENT MEDICAL DEVICE RECALL letter dated December 23, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Examine their inventory and quarantine product that was subject to the recall that was not sealed properly. Contact any customers if they further distributed the product and notify them at once of this product recall. Contact the firm if they had any unused product. The firm would issue a UPS Call Tag to have the product returned to them and the firm would send replacement product as soon as they were notified.Complete and return the Recall Return Response Form to the firm in the attached prepaid envelope.For any questions regarding this recall call 734-213-5530. |  | 300 | 300 | Worldwide Distribution - USA including AK, AZ, CA, FL, IL, KS, LA, MA, ND, NM, OH, TX, WA, WI and WV; and Canada | Terminated | July 13, 2011 | 208 | 300
Z-1378-2011 | 57666 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | InfusO.R. Infusion Pump, an Rx syringe infusion pump; Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A.; product codes 2L3100, 2L3100R, 2L3100U, 6461500 and 6461500R. These product codes includes refurbished units and InfusO.R. pumps previously sold under the Bard label. |  |  | January 10, 2011 | 2011 | Baxter Healthcare Corp. | Should a battery with the minimum contact diameter (.51 inches) be placed on the battery spring (.53 inches), an electrical circuit will not be completed and the InfusOR will not operate. The reason being that the InfusOR (2L3100) Battery Spring has a flat contact that measures .53 inches Âż.01 inches on the inner diameter of the spring coil. The IEC 60086-2 and ANSI 08 standards specify a minimum | Other | Safety Alert letters dated January 11, 2011 were mailed via first class mail to the InfusOR direct accounts, to the attention of the Director of Biomedical Engineering, the Director of Anesthesia and the Director of Materials Management, on the same date, informing them that over-the-counter Energizer C size alkaline batteries should not be used with the InfusO.R. Pumps (Product codes 2L3100, 2L3100R, 2L3100R, 2L3100U, 2L3100N, 6461500 and 6461500R). That particular battery incorporates a feature that may inhibit electrical contact and affect the ability of the pump to power-up. The accounts were requested to check their inventory of InfusO.R. Pumps for the presence of over-the-counter Energizer C size alkaline batteries. The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST.*** Baxter expanded the scope of their action to include the Mini-Infuser System and sent Urgent Device Correction letters dated 3/4/11 via First Class mail to all InfusOR and Mini-Infuser customers who had purchased either of the affected product lines. The letters informed the accounts that the Mini-Infusers use the same size battery and incorporate the same battery contact interface as the InfusOR pump, and that the same feature that inhibits electrical contact in over-the counter Energizer C size batteries may also be incorporated by other battery manufacturers. The accounts were requested to verify that the InfusOR and Mini-Infuser will power-on before use; if the device does not power on, replace the batteries with a different brand of size C batteries and verify that the device will power-on before use; and to place a copy of the letter with each operator's manual for the InfusOR and Mini-Infuser System for future refer |  | 39725 | 39,725 pumps | Nationwide Distribution | Terminated | August 27, 2012 | 595 | 39725
Z-1377-2011 | 57270 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Optima MR450w 1.5T MR systems |  |  | November 18, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware that during certain service procedures, there is a potential for electric shock associated with the gradient cooling manifold hoses of your MR750, MR450, and MR450w MR systems that may impact service or maintenance personnel's safety.The gradient cooling manifolds used on the affected systems may contain electrically conductive hoses, which become live when the gra | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 16, 2011, to all affected customers. The letter was addressed to Hospital Administrator / Risk Manger, Radiology Department Manager and Maintenance / Biomedical Engineering Department. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product correction and Contact Information.Customers were instructed to follow appropriate lockout/tagout (LOTO) procedures.Customers were instructed to contact GE Healthcare engineering at 843-687-7213 or 843-697-3009 for instructions on how to proceed. For questions concerning the notiifcaton customers were instructed to contact their local Support Engineer. |  | 68 | 68 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, LA, MD, MA, MI, MO, NV, NJ, NY, NC, OR, PA, SC, TN, TX, WA, WI. and the countries of AUSTRALIA, BRAZIL, CANADA, EHYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, ITALY, JAPAN, KUWAIT, QATAR, REPUBLIC OF KOREA, SWITZERLAND, TAIWAN, and the UNITED KINGDOM.UNITED KINGDOM, TAIWAN, SWITZERLAND, REPUBLIC OF KOREA, QATAR, KUWAIT, JAPAN, ITALY, INDIA, GREECE, GERMANY, FRANCE, FINLAND, EGYPT, CANADA, BRAZIL, AUSTRALIA. | Terminated | February 04, 2012 | 443 | 104
Z-1390-2011 | 56670 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The GMD UNIVERSAL URINARY INCONTINENCE SLING (Code 1010) is supplied sterile and is intended for single patient use only. |  |  | February 14, 2011 | 2011 | Generic Medical Devices | Protective sheath sticks to itself and sleeve breakage has resulted in complaints and MDRs. | Component design/selection | The firm, GMD, issued an "URGENT MEDICAL DEVICE RECALL" letter dated February 14, 2011 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed not to use any GMD Model 1010 Universal Slings remaining in their inventory; return any Model 1010 slings in their possession to the Distributors or GMD; notify their customers to whom they shipped Model 1010s of this voluntary recall and request them to return any Model 1010 slings in their possession to them; and return all Model 1010 inventories to GMD using the attached Recalled Model 1010 Return Form. Note: Full credit in the amount of the returned products purchase price will be issued for all sterile, unexpired product returned to GMD.Contact the GMD Product Performance Senior Manager, at phone: (253) 853 3512 or email: 1010recall@gmd-us.com, if you have any questions/issues related to this recall. |  | 1711 | approximately 1711 | Worldwide distribution: USA and countries including: Australia, Argentina, Brazil, Canada, South Africa, and Scandinavia. | Terminated | September 24, 2012 | 588 | 1711
Z-1392-2011 | 53691 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | V5M TEE Transducer, when used with an ACUSON S2000 ultrasound system that has cardiac options installed and is at software version VA10, VA10A, VA10B, VA10C, VA15, VA15A, VA15B, VA16 or VA16A, manufactured by Siemens Medical Solutions, Mountain View, CA. |  |  | August 28, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | If the transducer is not disconnected from the system after an automatic shutdown due to thermal limits being exceeded, then the next time it is used, the thermal safety guards do not take effect. Normally, the thermal limit before automatic image freeze (and therefore power shut off) occurs at 43 degrees Centigrade. | Software design | Siemens sent a Customer Safety Advisory Notification letter on August 28, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that their local Customer Service Engineer would install a software update as soon as it becomes available. Customers were instructed to pass the Customer Safety Advisory Notification letter to all who need to be aware of these issues until the corrective action is completed. For questions regarding this recall call 650-969-9112 or 800-422-8766. |  | 97 | 97 units | Worldwide Distribution - USA (nationwide) and internationally to Canada, China, Czech Republic, Denmark, Germany, Great Britian, Hong Kong, Indonesia, Italy, Japan, Norway, Poland, Portugal, Russian Federation, Singapore, South Korea, Taiwan, and Turkmenistan. | Terminated | March 29, 2011 | 578 | 97
Z-1395-2011 | 57715 | No | GTY | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | SP Brand Strep A Cassette; Part Number: B1077-26; Lot #: STA0010037 |  |  | January 12, 2011 | 2011 | Biosite Inc Dba Innovacon Inc. | The recall was initiated after the Investigation of customer-reported complaints, testing of market retention and returned product, It was determined that this lot of product hasdemonstrated an elevated Incidence of non-specific binding, potentially causing false positive results. A falsepositive test result could lead to the administration of antibiotics that are unnecessary. This exposes the | Pending | The firm, Alere San Diego, sent a "URGENT MEDICAL DEVICE RECALL" letter dated January 17, 2011. The letter described the product, problem and action to be taken by customers. The customers were instructed to: (1) Immediately discontinue use of the affected lot (if applicable) and inventory remaining full or partial kits at their site (if applicable);(2) Destroy any remaining inventory of this lot and document destruction per their site requirements (if applicable);(3) Complete and FAX the enclosed Verification Form within 10 days to confirm their receipt of this notice and to receive replacement of their remaining inventory of this lot (if applicable). The customers were ask to please return the form even if they have no remaining inventory or do not have any involved product in order to confirm receipt of this notification; and (4) to please ensure that all users of the device have received a copy of this Urgent Medical Device Recall.Customers with questions about the information contained in the notice were instructed to contact Technical Services at 888-246.7483, option 2 or by e-mail attechservices@alere.com. |  | 19877880 | 814 kits/24,420 devices | Nationwide distribution | Terminated | June 26, 2012 | 531 | 19877880
Z-1383-2011 | 57790 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | AUTOMIX High Speed Compounder System, product 2M8077; automated nutrition compounders; Baxter Healthcare Corporation, Clintec Nutrition Division, Deerfield, IL 60015 |  |  | January 27, 2011 | 2011 | Baxter Healthcare Corp. | The AUTOMIX compounder may not respond correctly to the key pressed by the operator using the keypad. The incorrect keypad response issue can occur both in manual data entry mode (STD Mode) and when entering order numbers to download from host software (MAN ID Mode). It is possible that an incorrect volume and/or an incorrect solution could be admixed into the final bag. | Other | Baxter Healthcare Corp. sent Urgent Device Correction letters dated January 27, 2011, via first class mail to all affected customers. informing personnel that use Automix compounders to exercise care when entering data, and verify the accuracy of the keypad response. The accounts were requested to convey this information to all users of the Automix compounders, emphasizing the need to verify data accuracy each time the keypad is used. The accounts were also requested to complete the attached customer reply form listing the model and serial number of each of the Automix compounders at the facility and confirming their receipt of the letter, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to their local Baxter representative, and technical questions were directed to Baxter's Global Technical Services at 1-800-626-2667.After determining the root cause of the incorrect keypad responses, Baxter issued a follow-up correction letter on 4/27/11 outlining the recommended work practices to reduce fluid penetration into the Automix compounder keypad/control module. |  | 270 | 270 units | Worldwide Distribution - USA - Nationwide including Puerto Rico, and internationally to Brazil, Canada, Colombia, Costa Rica, Czech Republic, El Salvador, Egypt, Germany, Guatemala, Honduras, Luxembourg, Mexico, Panama, Saudi Arabia, Singapore, Switzerland, Taiwan, Trinidad, Turkey, United Arab Emirates and United Kingdom. | Terminated | February 16, 2012 | 385 | 5160
Z-1393-2011 | 56367 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Tyvek Self-Seal Pouch with STERRAD Chemical Indicator 12332. |  |  | June 23, 2010 | 2010 | Advanced Sterilization Products | The recall was initiated after Advanced Sterilization Products (ASP) discovered a non-conforming lot. The misaligned perforations may have caused the seal to be incomplete on some of the pouches, creating a potential breach for pouches being used to store sterilized equipment on shelves. | Process control | Advanced Sterilization Products (ASP) released Urgent: Product Recall notifications to its consignees identifying the affected product and explaining the reason for recall. The letter also discussed the impact of using this affected product and return instructions. The firm asked customers to examine their inventory, discontinue use of Tyvek Self-Seal Pouch with STERRAD Chemical Indicator 12332 and to return all unused affected products to Stericycle. Customers are to complete the Business Reply Card and return it to Stericycle. |  | 470 | 470 | Worldwide Distribution -- USA, Argentina, Australia, Canada, Chile, Czech Republic, Hong Kong, India, Japan, Panama, Peru and Puerto Rico. | Terminated | March 14, 2012 | 630 | 470
Z-1394-2011 | 57082 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet Cardiopulmonary AG, MAQUET RPM 20: Roller Pump for Heart Lung Machine HL20. |  |  | September 28, 2010 | 2010 | Maquet Cardiovascular Us Sales, Llc | A wrong ball bearing was installed in the pump heads of product. This may make the pump head wear out and cease function. | Component change control | Maquet hand delivered a letter to the consignee on 10/4/2010. |  | 4 | 4 units | One hospital in New York City | Terminated | August 11, 2011 | 317 | 4
Z-1388-2011 | 57245 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-ROM Distal Femoral Cutting block Guide SZ X- SM/SM, REF 216311001, non-sterile, Depuy International Leeds, England. |  |  | October 19, 2010 | 2010 | Depuy Orthopaedics, Inc. | The firm was notified that the Hinge Block Femoral Box Guide cut being off alignment when using the S-ROM Distal Femoral Cutting Block Guide SZ SML/X-SML pin holes to locate. The firm then found that the same issue was found int the MED block size. | Device Design | The firm, DePuy, sent an "URGENT INFORMATION - DEVICE CORRECTION" email and letter dated October 19, 2010 to all customers. The notices described the product, problem and actions to be taken by the customers. The notices stated that: Depuy is issuing a device correction; however, the product may still be used; the devices do not need to be returned; the distributors were to provide a letter to clinicians that were using the device; the surgical technique will be revised in October 2010; the product will undergo a design change, and the recalled product will be traded out in December 2010 and January 2011.If you have any additional questions, please contact DePuy's Scientific Information office at 1-888-554-2482.On 3/4/2011, The firm sent out a revised URGENT DEVICE RECALL asking that all affected cutting guides be taken out of service immediately. Devices should be returned to DePuy and new cutting guides will be sent out within 2 weeks. |  | 229 | 229 | Worldwide distribution: USA including states of: AZ, AR, CA, CO, FL, GA, IA, ID,IL, IN, KY, KS, LA, MA, MD, MI, MN, NC, NV, PA, TN, TX, VA, WA, and WI; and countries including: Canada and Ireland. Australia, Austria, Canada, Denmark, Ireland, Israel Italy, New Zealand, Russia, Sweden, Switzerland, Spain, and UK. | Terminated | December 03, 2012 | 776 | 448
Z-1387-2011 | 57700 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Masimo Hemoglobin |  |  | December 23, 2010 | 2010 | Masimo Corporation | Masimo Corporation is initiating a voluntary recall on the Pronto-7 rainbow 4D Resuable Sensor because when performing an SpHb measurement, the sensor may incorrectly read actual finger temperature. As a result, Pronto-7 Check Pulse CO-Oximeter may provide an incorrect SpHb measurement. | Process design | Masimo sent an Urgent Voluntary Medical Device Recall letter dated December 21, 2010, to all affected customers via FexEx on December 23, 2010. The Urgent Voluntary Recall Notice provides instructions on the actions that the customers and field personnel must take to correct the issue.Customers were instructed to choose any of the three options below. Option 1 would ensure accurate results are achieved when using the Pronto-7 rainbow 4D Reusable Sensor with the Pronto-7 Spot Check Pulse CO-Oximeter. (1) Option 1: Prior to performing an SpHb measurement, wait at least seven (7) minutes with the patient's finger inserted inside the Pronto-7 rainbow 4D Reusable Sensor. The seven minute wait will help ensure that the actual finger temperature and the sensor temperature are close to equalized.Masimo is working to quickly develop a revision to the Pronto-7 rainbow 4D Reusable Sensor that will address this issue and plan to have it available in early 2011. We will contact you when the new revision of the sensor becomes available to replace the current sensor;(2) Option 2: Return only the Pronto-7 rainbow 4D Reusable Sensor and request a loaner Pronto device with rainbow DCI sensor (part numbers 9167 and 2646) to be utilized until the new version of the sensor is available for use with the Pronto-7. Please let us know if you wish to try the Pronto or see the product brochure for Pronto. We will contact you when the new revision of the sensor becomes available;(3) Option 3: In addition, the Pronto-7 rainbow 4D Reusable Sensor and the Pronto-7 Spot Check Pulse CO-Oximeter may be returned for a refund.Customers were also instructed to complete the Tracking/Verification Form provided with the recall notice and fax it to Masimo at: 1-949-297-7499. |  | 497 | 497 | Worldwide Distribution - Nationwide in the US and worldwide to Argentina, Austria, Bahrain, Belgium, Benin, Canada, Chile, China, Costa Rica, Croatia, Czech Republic, Egypt, Finland, France, Germany, Ghana, Greece, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebaron, Libyan Arab, Jamahiriya, Malaysia, Malta, Mauritius, Morocco, Netherlands, New Zealand, Norway, Oman, Philippines, Poland, Portugal, Reunion, Romania, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United kingdom | Terminated | March 26, 2013 | 824 | 497
Z-1391-2011 | 57805 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics cardioPAT Cardiovascular - autotransfusion System |  |  | January 25, 2011 | 2011 | Haemonetics Corporation | Update Instructions for Use for the CardioPAT&reg; Manual for Leak Detection. User may mis-interpret the air-leak rate displayed by the cardioPAT device when making a determination about chest-drain removal, which could result in the patient experiencing a pneumothorax. | Labeling design | Haemonetics issued a Field Correction Notice on 1/25/11 . with the Air Leak Detection addendum and index be inserted into the appropriate document, applied to all cardioPAT@ devices on site. The serial number for each cardioPAT device on site should be recorded on the attached acknowledgementand the completed acknowledgement should be returned to your Haemonetics Sales Consultant. A Haemonetics Sales Consultant will train on the content of these items and insert them into the Operator Manuals |  | 309 | 309 units | Worldwide distribution, including USA, Australia, Belgium, Canada, Switzerland, China, Germany, Finland, France, United Kingdom, Hong Kong, Israel, Italy, Japan, Netherlands, Russian Federation, and Sweden. | Terminated | January 14, 2013 | 720 | 309
Z-1401-2011 | 57781 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Verio |  |  | January 10, 2011 | 2011 | Siemens Medical Solutions USA, Inc | On some systems the gradient cable connections did not meet firm's specifications after installation | Other | Siemens Medical Solutions USA, Inc sent customers a letter on January 10, 2011 notifying them of an Update Instruction and informing them that recalling firm would preventatively check the gradient cable connections of all systems. For questions regarding this recall call 610-219-6300. |  | 92 | 92 | Nationwide Distribution including AK, AZ, CA, CT, FL, HI, IL, IN, KY, MA, MD, MI, MN, MO, MT, NC, NJ, NM, NV, OH, OK, OR, PA, RI, TN, TX, VA, WI, and WV | Terminated | September 25, 2012 | 624 | 92
Z-1421-2011 | 57775 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM SIRESKOP SD, SIRESKOP SX, AXIOM LUMINOS |  |  | January 14, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm became aware of a potential malfunction and possible hazard to patients in the event the tilt buttons on the OPTI Grip become stuck under the housing | Other | Siemens Medical Solutions USA, Inc issued a Customer Safety Advisory Notice via Update Instruction XP005/10/S to affected customers. This letter informs customers of the potential malfunction and of possible danger for patients, users or other persons. A future Update Instruction will be released for the replacement of the tilt buttons on all OPTI Grip handles. Customers were instructed to forward the safety notice to any new owners of this device and to provide Siemens Medical with names of the new customers. For any questions regarding this recall call 610-448-3237. |  | 1323 | 1323 | Nationwide Distribution including AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Terminated | July 16, 2013 | 914 | 1323
Z-1164-2011 | 57663 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Spectrum Central Venous Tray, Minocycline/Rifampin Impregnated Double Lumen Polyurethane, 5.0 Fr/ 15cm, C-UDLMY-501J-LSC-ABRM-HC-IHI-FST, G48067, Sterile, Cook Medical, Cook Incorporated, Bloomington, IN. |  |  | October 04, 2010 | 2010 | Cook, Inc. | Cook, Inc. is recalling Single, Double, Triple, and Five-Lumen Central Venous Catheter Trays AND Single and Double Lumen PICC Peripherally Inserted Central Venous Catheter Trays. These kits and trays contain 0.9% Sodium Chloride Injection USP 5 mL fill in 6 mL flush syringes that have been recalled by Excelsior Medical. During a routine test for syringe sterility, a possible integrity failure wa | Nonconforming Material/Component | Cook, Inc. sent URGENT DEVICE RECALLS letters to its customers dated 10/4/2010. The letters asked customers to examine their inventory and notify Cook for product return. The letter included a response form. |  | 51 | 51, 473 total trays | Worldwide distribution: USA, Chile, Colombia, Panama, Barbados, Canada, Guatemala, and Saudi Arabia. | Terminated | June 21, 2011 | 260 | 51
Z-1400-2011 | 57818 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Integra Radionics HRAIM Intubation Head Ring Assembly |  |  | January 12, 2011 | 2011 | Integra LifeSciences Corp. | Overall length of the intubation hoop in the HRAIM Intubation Head Ring Assembly is too long and will not allow a device to to attach | Manufacturing material removal | Integra sent a Medical Device Field Corrective Action letter dated January 12, 2011, to all affected customers via Federal Express. Integra Sales Specialists were provided with the replacement part and written instructions on how to replace the affected intubation hoops from customer inventory identified in the traceability report. Upon completion of the field correction each sales specialist was instructed to return the affected hoop along with a completed copy the Recall Acknowledgement and Return Form. If additional information is required contact me at (609) 936-2485 or james.tillman@integraJife.com. |  | 8 | 8 units | Worldwide Distribution - USA, Canada, and Australia | Terminated | April 04, 2011 | 82 | 8
Z-1396-2011 | 57861 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Comfort Sling Plus Model 300, Part number 35300305, Size Medium, LIKO, Sweden. Liko a Hill-Rom Company. |  |  | February 09, 2011 | 2011 | Hill-Rom, Inc. | Slings were improperly assembled and the loops that are used to support the sling when hung from a lift were too long, potentially putting the patient being lifted at risk of falling out. | Pending | The firm sent an URGENT MEDICAL DEVICE RECALL notice dated 02/09/2011 to Facility Risk Manager/Facility Administrators. The letter Identified the affected product, explained the recall situation, and asked firms to check their facility for the affected serial numbers. Consignees were asked to return the recalled products, and a replacement sling would be sent to them. Customers can contact Hill-Rom Technical Support at 800-445-3720, if they have any questions concerning this matter. |  | 10 | 10 units | Worldwide Distribution -- US, Switzerland, Germany, Denmark, Finland, France, UK, Italy, Netherlands, Sweden, and Canada. | Terminated | October 07, 2011 | 240 | 10
Z-1399-2011 | 57431 | No | MKT |  |  | Microbiology | PMA | Abbott RealTime HBV Assay, List 2N40, in vitro diagnostic; Abbott Molecular Inc., Des Plaines, IL 60018 USA |  |  | December 20, 2010 | 2010 | Abbott Molecular | Internal testing identified the potential of under filled and over filled enzyme reagent vials which may result in obtaining less than the total number of tests per kit as described on the package insert. | Other | Abbott Molecular sent a Field Advisory Notice letter to all RealTime HIV-1, CT/NG and HBV assay customers on December 20, 2010, informing them that internal testing identified the potential of under filled and over filled enzyme vials which may result in obtaining less than the total number of tests per kit as described in the package insert. The customers were told that specimens that meet acceptance criteria, per the product labeling are not affected by the situation and are considered valid results, and that the issue has no impact on results for those runs that generate valid results. Any questions were directed to Abbott Technical Support at 1-800-553-7042, Option 2. |  |  |  | Nationwide Distribution | Terminated | November 19, 2012 | 700
Z-1432-2011 | 57084 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | 9) Signa Excite 1.5 T (Excite II) |  |  | November 09, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue associated with the gradient cables being inadvertently swapped while servicing the gradients on some GE Healthcare MR Systems that may impact patient safety.If the gradient cables are inadvertently swapped while servicing the gradients, and the Service Engineerdoes not perform the geometry check as required in the service procedure, images m | Nonconforming Material/Component | On November 9, 2010, the firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated November 5, 2010 to all Consignees/Customers. The letter described the product, problem and action to be taken by the firm. GE instructed the customers to contact their local GEHC support Engineer, if they have any questions about the image orientation, or the procedures for geometry checks on thier system. Additionally, a GE Healthcare representative will install a device, which will reduce the potiential for the gradient cables to be swapped.Contact your local GEHC Support Engineer if you have any questions concerning this notification or phone: 262-548-2731. |  | 1175 | 1175 | Worldwide distribution: USA including DC and PR; and countries including: VIET NAM, VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, UKRAINE, TURKEY, TUNISIA, FORMER YUGOSLAV REPUBLIC OF MACEDONIA, THAILAND, TAIWAN, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, REUNION, REPUBLIC OF KOREA, PORTUGAL POLAND, PHILIPPINES, PERU,PARAGUAY, PANAMA, PAKISTAN, NORWAY, NEW ZEALAND, NETHERLANDS, MOROCCO, MEXICO, MAURITIUS, MARTINIQUE,MALTA, MALAYSIA, LUXEMBOURG,LITHUANIA, LIBYAN ARAB JAMAHIRIYA, LEBANON, LATVIA, KUWAIT, KENYA, KAZAKHSTAN, JORDAN, JAPAN, JAMAICA, ITALY, ISRAEL, ISLAMIC REPUBLIC OF IRAN, IRELAND, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, HONDURAS, GUYANA, GUATEMALA, GUADELOUPE, GREECE, GERMANY, GEORGIA, FRENCH POLYNESIA, FRANCE, FINLAND, EL SALVADOR, EGYPT, ECUADOR, DOMINICAN REPUBLIC, DENMARK, CECH REPUBLIC, CYPRUS, CROATIA, COLOMBIA, CHINA, CHILE, CANDA, BULGARIA, BRAZIL, BOSNIA AND HERZEGOVINA, BOLIVIA, BELGIUM, BELARUS, AZERBAIJAN, AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA. | Terminated | July 03, 2012 | 602 | 8937
Z-1434-2011 | 57821 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Cardiovascular Review Station |  |  | January 27, 2011 | 2011 | AGFA Corp. | Echocardiography volume measurements performed within CRS were not reflected correctly in associated Results Management (RM) reports. | Other | The firm, AGFA Healthcare, sent a "Safety Notice" dated January 27, 2011 to all customers. The notice described the product, problem and action to be taken by firm and customer. AGFA informed the customers that a software upgrade was made between December 2008 and February 2010 and instructed the customers to distribute this information to all individuals within their facility and to complete and return the attached URGENT SAFETY NOTICE FEEDBACK FORM via fax to 864-421-1664 or email to debbie.norris@agfa.com. Note: Service Bulletin DD+DIS100.09E, was issued March 2009 and an upgrade from Software Versions 2.12.05 through 2.14.03.SU1 to Software Version CRS 2.14.03.SU2 was required for the impacted sites. Most identified consignees/customers required the software upgrade, but some consignees/customers had already been upgraded prior to the event and did not require additional upgrades.If you have any questions about this matter, please contact the Post Market/QARA North America via email at debbie.norris@agfa.com or call at 864-421-1754. |  | 107 | 107 | Worldwide distribution: USA and country of: Canada | Terminated | March 29, 2011 | 61 | 107
Z-1422-2011 | 57704 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens brand ARTISTE MV with Artiste Digital Linear Accelerators in combination with the Flat Panel Positioner; Part No. 8139789; |  |  | December 03, 2010 | 2010 | Siemens Medical Solutions USA, Inc., Concord, CA | The firm become aware of a potential issue with ARTISTE system that may result in patient injury should the positioner belt break, and the cam disc (secondary safety mechanism) is not installed and if the flat panel is over the patient. | Other | The firm, Siemens, issued a "Customer Information" letter on December 3, 2010 via mail to all its direct consignees/customers. The letter described the product, problem and actions taken by the firm to the customer. Siemens arranged for an inspection of their customers systems to make sure the safety mechanism is fully functional. Siemens instructed the customers to include the Customer Information letter in their ARTISTE System Owner manual chapter 'Safety Advisory Letters' .If you any questions, please call 925-246-8407. |  | 20 | 20 units | Worldwide distribution. | Terminated | June 30, 2011 | 209 | 20
Z-1423-2011 | 57629 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | HomeFill Oxygen Compressor IOH200 Homefill, The serial numbers are as follows: 10JF034081, 10JF034082, 10JF034083, 10JF034084, 10JF034085, 10JF034086, 10JF034087, 10JF034088, 10JF034089, 10JF034090, 10JF034091, 10JF034092, 10JF034093, 10JF034094, 10JF034095, 10JF034096, 10JF034097, 10JF034098, 10JF034099, 10JF034100, 10JF034101, 10JF034102, 10JF034103, 10JF034104, 10JF034105, 10JF034106, 10JF034107, 10JF034108, 10JF034109, 10JF034111, 10JF034112, 10JF034113, 10JF034114, 10JF034115, 10JF034116, 10JF034117, 10JF034119, 10JF034120, 10JF034121, 10JF034122, 10JF034123, 10JF034124, 10JF034125, 10JF034126, 10JF034127, 10JF034128, 10JF034129, 10JF034130, 10JF034131, 10JF034132, 10JF034133, 10JF034134, 10JF034135, 10JF034136, 10JF034177, 10JF034178, 10JF034179, 10JF034180, 10JF034181, 10JF034182, 10JF034183, &amp; 10JF034184. |  |  | January 11, 2011 | 2011 | Invacare Corporation | Invacare Corporation decided to recall the products because risk analysis reveals that inadequate clearance between a motor ground clip and the motor windings may result in an increased risk for shock to service personnel. It also may result in contact with the motor winding, resulting in a shorted winding; this condition has the potential to result in motor failure/shutdown or to cause the motor | Other | The firm, INVACARE Corp., sent an "URGENT PRODUCT RECALL: IOH 200 Homefill Recall" letter dated January 10, 2010 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to "immediately examine their inventory and quarantine product subject to recall. In addition, if they may have further distributed this product, please identify their customers and notify them at once of this product recall." The customers were also instructed to complete and return an enclosed response form. Note: All product is to be return to the recalling firm.If you have any questions, contact the Manager, Regulatory Affairs at (440) 329-6356. |  | 62 | 62 Compressors | Nationwide distribution: USA including states of: FL, GA, KY, LA, MO, NC, NE, NJ and SC. | Terminated | November 21, 2011 | 314 | 62
Z-1409-2011 | 57487 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stryker Navigation System - iNtellect Cranial Planning Software - Enable, Part number 6000-652-300. It is part of the product series of the Stryker Navigation System. The system comprises software for pre-operative surgical planning and supports computer assisted surgery based on a wireless optical tracking localization device for the use in navigated cranial surgery. It includes automatic multi-modality correlation. The system provides an easy-to-use work flow concept for image importation, FESS, Craniotomy, Biopsy, etc. Additional, automatic segmentation of brain surface, tumor, vessels, etc. are supported. |  |  | November 10, 2010 | 2010 | Stryker Instruments Division of Stryker Corporation | When users imports more than one image set and changes the viewing direction for both images, the orientation of every other image series might be stored incorrectly, as a mirror image. This may lead to insufficient or inaccurate correlation results or incorrectly shown orientation of every second image series during navigation. | Software design | Urgent: Medical Device Recall Notification letters were sent on 12/3/2010, and software upgrades were to be performed in the field by the end of 2010. Specific Customer Instructions: 1) Locate the units listed in the notification.2) Distribute this notification to all departments affected.3) Return the enclosed business reply form to confirm receipt of this notification and identify how many affected units are currently in your inventory. 4) Once the new software is available, Stryker will dispatch a service tech in order to update the software. Note: Software upgrades for all applications will be available by the end of the year. The upgrades will be scheduled as we receive your business reply form.Note: If you currently have a Time License or Pay Per Use card, a new one will be issued to you at the time of the update.5) If you have loaned or sold any of the units listed in this notification, please, forward a copy of this notice to the new users and advise Stryker of the new users name and their new location.6) If you have disposed of any of the affected units and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers.Questions regarding this recall should be directed to Stryker Field Action Team at 800-800-4236, ext 3808, 3584, or 4354 between the hours of 8 a.m. - 5 p.m. Monday - Friday. |  | 1155 | 1,155 total | Worldwide Distribution -- USA, Canada, Mexico, Australia, Argentina, Brazil, United Kingdom, Japan, Poland, S. Korea, Hong Kong, Taiwan, India, and New Zealand. | Terminated | June 01, 2011 | 203 | 1155
Z-2809-2011 | 57829 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ***EXPANDED RECALL*** Curlin Infusion Administration Set, REF 340-4165, Sterile EO, Zevex, Inc., 4314 Zevex Park Lane, Salt Lake City, UT 84123. |  |  | January 24, 2011 | 2011 | Zevex International, Inc. | Defect in tubing may trigger false Air-In-Line warning. | Process control | The firm, MOOG, sent an "Expansion of Moog/Curlin Ambulatory Pump Administration Set Recall" letter dated May 2, 2011 to its customers. This letter informed the customers that this recall has been expanded to included specific lot numbers of Curlin model ambulatory infusion administration sets. The customers were instructed to follow the outlined process in the letter: remove suspect product code from inventory, contact MOOG Customer Service at 1-800-970-2337, prompt #7, contact your distributor to arrange replacement product, request a shipping return label as per the distributor's process and ship back to their location and follow the temporary solution outlined in the letter. Moog Customer Service will provide a call tag for convenient product return.If you have any questions contact MOOG Customer Service at 1-800-970-2337. The firm, MOOG, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 24, 2011 to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to remove suspect product code from inventory; contact their distributor to arrange replacement product, if purchased from distributor and request shipping return label per distributors process and ship back to their location or contact MOOG Customer Service at 800-970-2337, prompt #7, if purchased directly from MOOG and MOOG Customer Service will provide a call tag for convenient product return; follow the temporary solution provided until replacements are administered; and complete and return the MEDICAL DEVICE CORRECTION REPLY FORM in the self addressed envelope provided.If you have any questions, contact MOOG Customer Service at 800-970-2337. |  | 8080 | 8,080 units | Worldwide Distribution: USA including states of: AR, CA, FL, IA, KS, MS, NE, NY, OH, OK, TN, TX, and UT; and countries including: Australia, Croatia, France, Germany, Kuwait, New Zealand, Qatar, Saudi Arabia, South Africa, United Arab Emirates, and UK. | Terminated | September 27, 2011 | 246 | 3024740
Z-1435-2011 | 57820 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Results Manager/Results Manager Administration Tool |  |  | April 15, 2009 | 2009 | AGFA Corp. | User could modify a measurement's units and that change would not be reflected in echocardiography formulas used within the Adult echocardiography RM. | Other | The firm, Agfa Healthcare, sent individual letters dating from April 15, 2009 to August 26, 2009, to the customers. The letters described the product, problem and actions to be taken. Agfa provided the affected customer with a complete list of their impacted patients and the measurements. Agfa also scheduled unit corrections with their customers and received acknowledgement of the letters at the time of scheduling the unit corrections. Service Bulletin DIS026.09E was issued January 2009 to inform users that Agfa will no longer allow customers to request unit changes for measurements that have formulas. Note: Technical Support ran diagnostic scripts for each customer and documented in GSC cases for each consignee/customer.If you have any questions, contact Agfa Healthcare at (949) 830-4959. |  | 6 | 6 devices | Nationwide distribution: USA including states of: IL, MT, NY, OR and SC. | Terminated | October 13, 2011 | 911 | 6
Z-0893-2011 | 54768 | No | RFZ |  |  |  |  | Hemochron Signature Elite Whole Blood Microcoagulation (IVD) System. |  |  | March 31, 2010 | 2010 | International Technidyne Corp. | There is a mismatch between the class of the laser barcode scanner and the labeling of the Hemochron Signature Elite: some of the devices were inadvertently equipped with class II laser barcode scanner instead of the intended class I barcode scanner. | Labeling mix-ups | Urgent - Medical Device Recall - Hemochron Signature Elite letters, dated February 24, 2010, by Federal Express to all direct customers. The letter informed customers of the issue and stated that the problem does not negatively affect the performance of the product and to continue using the product as intended. In addition, customers are to notify the appropriate personnel about the notice, complete and return the attached form, and an ITC representative will contact customers to discuss arrangements to re-label their devices. Customers can contact ITC Technical Support at 800-579-2255 (US) or 732-548-5700 ext 4700 (International) if they have any questions regarding this issue and notice. Customers can also email sigelitescan@itcmed.com with questions. |  | 1640 | 1640 units (1303 US, 337 foreign) | Worldwide Distribution -- United Arab Emirates, Belgium, Canada, Switzerland, Germany, Denmark, France, United Kingdom, Ireland, Iceland, Italy, Kuwait, Liechtenstein, Netherlands, Oman, Qatar, Saudi Arabia, Sweden, and Singapore. | Terminated | September 25, 2012 | 909 | 1640
Z-1457-2011 | 57402 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | QuickVue One-Step hCG Urine, Quidel Catalog Number: 20109, Lot Number: 706842 |  |  | November 23, 2010 | 2010 | Quidel Corporation | The recall was initiated after Quidel Corporation determined that a small number of QuickVue One-Step hCG Urine kits have been incorrectly packaged with QuickVue H. pylori gll cassettes inside the kit boxes. These foil pouches and testcassettes are clearly labeled as H. pylori. Nevertheless, there is a risk that customers may overlook this labeling. | Pending | Quidel Corporation initiated an "Urgent Medical Device Recall" letter on November 23, 2010, with an attached "Distributor Recall Notification Form" to the affected distributor consignees notifying them of this issue and requesting return of all kits from the affected lot. Quidel is further requesting that distributor consignees notify end-users of the issue and request return of kit boxes containing the H. pylori pouched cassettes. The recall notices sent to the distributors were sent via US Postal, certified, with return receipt.The consignees were instructed to:(1) Locate kits with the affected lot number and immediately remove the kits from your existinginventory and place them into a quarantine location, (2) Contact Quidel Corporation for instructions on how to return the kits and obtain replacement product.Quidel Technical Support(800) 874-1517, Monday-Friday, 7am - 5pm PST,(3) Fax the enclosed Inventory Form to Quidel at (858) 431-3522. Indicate the total number of kits currently on hand with lot number 706842 or if you do not have any kits of this lot numberremaining in inventory. |  | 90600 | 3624 kits = 90,600 tests total | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, GA, HI, IL, KS, KY, MA, MD, MN, NC, NJ, NM, NV, OK, OR, PA, TN, TX, UT, VA, and WA and Canada. | Terminated | June 26, 2012 | 581 | 90600
Z-1458-2011 | 57561 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | The Panorama HFO system has the Traceable Item Identification (12NC): |  |  | February 05, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | In the ExamCard MobiFlex the Patient Position is "Prone" which must be "Supine". The result is mislabeling of left and right side of the patient. If this is not detected by the operator there is a risk of wrong treatment or the need of additional X-RAY during treatment. | Other | Philips began sendng out the URGENT-Device Correction Notice letter dated February 5, 2010, ro all affected consignees on February 10, 2010. The consignees were advised to change the patient position from "Prone" immediately to "Supine" and store the corrected ExamCard in the hospital database if they use the ExamCard, MobiFlex dual Injection.Consignees were instructed to call Philips Healthcare Call Center at 800-722-9377 #5 (then follow prompts) and also reference "FCO 78100318" with any communication. |  | 82 | 82 units | Nationwide Distribution including the following states: AL, CA, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, NE, NH, NJ, NY, OH, OR, SC, TN, TX, VA, VT, WI, and WV.AL, CA, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, NE, NH, NJ, NY, OH, OR, SC, TN, TX, VA, VT, WI, and WV. There was no foreign consignees. | Terminated | August 10, 2011 | 551 | 82
Z-1472-2011 | 57514 | No | NRS | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | Veridex, Cell Search Circulating Tumor Cell Control Kit. Product code 7900003 |  |  | November 09, 2010 | 2010 | Veridex, LLC | There is a discrepancy between the assigned range of values on the high control label and the range listed on the Cell control Assigned Value Label for the Cell Search Circulating Tumor Cell Control Kit. | Process control | Veridex notified all affected consignees on November 11, 2010 by Federal Express overnight letter and prodivded with the informaiton regarding the label discrepancy. Foreign consignees were also notified on November 11, 2010 of the issue and instructed to notify their affected consignees. Consignees are asked to return the confirmation of receipt form by 11/29/2010. |  | 150 | 122 US, 28 International | Worldwide distribution, including: USA, Canada, France, UK, Germany, Italy, China, Spain &amp; Portugal. | Terminated | March 28, 2014 | 1235 | 150
Z-1459-2011 | 57560 | No | MKT |  |  | Microbiology | PMA | COBAS AmpliPrep/COBAS TaqMan HBV v2.0 US-IVD; |  |  | December 28, 2010 | 2010 | Roche Molecular Systems, Inc. | There is an issue regarding liquid level detection of the CS1 Magnetic Glass Particles (MGP) Reagent Cassette of the COBAS AmpliPrep /COBAS TaqMan HBV v2.0 Test on the COBAS AmpliPrep instrument. | Employee error | Roche sent Urgent Medical Device Correction Letters to their consignees on December 28, 2010 via Fed Ex. The letters advise the customers to verify the correct configuration of the CSI reagent cassette, remind customers to follow the package insert instructions for including all required controls for each batch and advise that review of previously generated results may be warranted. A UMDC Faxback form was attached for the consignee to complete.For questions regarding this recall call 908-253-7569. |  | 22212 | 93 kits to US; 22,119 kits to foreign | Worldwide Distribution - USA including CA, CT, IN, MI, MN, NC, NE, NY, OR, TX, WA, and WI and the countries of Australia, Austria, Azerbaijan, Belgium, Brazil, Canada, China, Croatia, Czech Republic, Egypt, Finland, France, Germany, Ghana, Greece, Hong Kong, India, Israel, Italy, Japan, Kuwait, Luxembourg, Malaysia, Mauritania, Mongolia, Netherlands, New Zealand, Norway, Poland, Romania, Saudi Arabia, Singapore, Slovakia, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, and the United Kingdom | Terminated | August 27, 2012 | 608 | 22212
Z-0286-2011 | 56922 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | BrightSpeed Excel/Edge/Elite Select Computed Tomography X-Ray System. |  |  | May 01, 2010 | 2010 | GE Healthcare, LLC | Failure to properly document the CTDI in the Technical Reference or User Manual. | Radiation Control for Health and Safety Act | Planned action(s) to repair defect or to bring product into compliance: GE Healthcare will bring the Technical Reference Manual into compliance by correcting the noncompliance in forward production and the installed base units. This will be accomplished by delivering an "Important Electronic Product Radiation Warning" letter and a TRM addendum which contains the corrected information to customers. This will be carried out via a Field Modification Instruction IFMI 229251 planned to be released in October, 2010 with an estimated completion date of April, 2011. The letter identifies the affected product and describes the defect and related hazards. It also asks customers to refer to the TRM addendum for the scan time accuracy, the helical scan exposure time, and the CTDI adjustment factors for 4 and 8 slices. Contact information is also provided. |  | 103 | 103 units | Nationwide Distribution | Terminated | December 05, 2013 | 1314 | 103
Z-1495-2011 | 53690 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PhD System EIA/IFAversion 2.1A software, model number 426-0246, a component of the Bio-Rad PhD System. Manufactured by Bio-Rad Laboratories, France, Distributed by Bio-Rad Laboratories, Hercules, CA. |  |  | September 29, 2009 | 2009 | Bio-Rad Laboratories Inc | Incorrect patient results are being generated from out-of-range samples tested withc ompetitive enzyme immunoassays on the PhD system with software version 2.1A. | Software change control | BIO-RAD sent an URGENT: MEDICAL DEVICE CORRECTION letter dated September 29, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their customer records to identify all PhD System customers that are using version 2.1A of the EIA/IFA Software and Methods CD-ROM.Customers who are not running competitive immunoassays on the PhD System may continue using the 2.1A software.Customers who run competitive immunoassays must revert to software version 2.0B.Customers were instructed to translate if necessary and add local contact information to the customer letter and distribute to all PhD System customers that are using 2.1A of the EIA/IFA Software and Methods CD-ROM. A Customer Medical Device Correction Response Form must be collected from each customer to ensure that they have received this important communication. Once forms have been collected from all applicable customers, they should be faxed to Bio-Rad CSD Regulatory Affairs Department along with the completed Subsidiary Medical Device Correction Response Form.For questions regarding this recall call 510-741-3954. |  | 262 | 262 units | Worldwide Distribution - USA including AL, CA, CO, CT, DC, FL, IL, MA, MD, MN, MO, MS, MT, NC, NY, OH, PA, SD, TN, TX, UT, and VT and the countries of Australia, New Zealand, Singapore, China, France, Sweden, and Canada | Terminated | March 11, 2011 | 528 | 262
Z-1479-2011 | 56385 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio Rad brand Autoimmune EIA ANA Screening Kit, 96 Tests, |  |  | March 01, 2010 | 2010 | Bio-Rad Laboratories Inc | The firm received customer complaints of increased positivity rates associated with the ANA Screening kit due to a decline in stability. | Nonconforming Material/Component | An Urgent: Medical Device Recall Notification, dated March 1, 2010, was provided to the firm's consignee, identifying the affected product and stating the reason for the recall. The letter also asked the customer to examine their inventory and quarantine the affected kits. If the product had been further distributed, those customers must be identified, notified, and asked to discontinue use of the affected product. Customers were to discard all affected product and complete and return the enclosed Medical Device Recall Response Form. The letter also stated that replacement kits would be available in mid-March. |  | 2109 | 2,109 bulk packs | Nationwide Distribution -- New York, US. | Terminated | March 11, 2011 | 375 | 2109
Z-1477-2011 | 57427 | No | LTK | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur CA 125 II Assay packaged in 100 and 500 test kits as follows: |  |  | December 14, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc | Positive bias (40-55%) with patient results greater than the upper limit of normal and less than the upper limit of normal | Process control | Siemens Healthcare Diagnostics on 12 /14/10 issued Urgent Device Recall notice via Federal Express to accounts. This notice informs the customer of the recall and the actions that need to be taken. The Urgent Device Recall notice directs customers to immediately discontinue the use of ADVIA Centaur CA 125 II assay kits that contain CA 125 II reagent lot 144. It also recommends that patient results obtained with this reagent lot be reviewed. A confirmation fax-back form is included to ensure that all customers worldwide have been notified and have complied with the mandatory action. Technical questions, please contact the Technical Solutions Center at 1-877-229-3711 |  | 2960 | 2960 units | Worldwide distribution, including USA, Canada, Argentina, Austria, Belgium, China, Denmark, France, Germany, Great Britain, Greece, India, Italy, Japan, Korea, Lavitia, Lithuania, Malyasia, Mexico, Netherlands, Norway, Peru, Philippines, Poland, Portugal, Romania, Saudi Arabia, Switzerland, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Czech Republic, Taiwan, Turkey, Uganda, and Venezuela | Terminated | April 12, 2013 | 850 | 2960
Z-1478-2011 | 55293 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | DiaSorin ETI-ANA Screen Kit, 96 Test, Model Number N0093, manufactured by Bio-Rad Laboratories. |  |  | March 01, 2010 | 2010 | Bio-Rad Laboratories Inc | Inaccurate results-- Consumer complaints of increased positivity rates with the ANA screening kit. | Nonconforming Material/Component | Recall was initiated on 3/1/2010. DiaSorin S.R.I. (the sole consignee) was alerted via email using the firm's Medical Device Recall Notification. Consignee informed of the recall and instructed to dispose of the affected product according to local laws. |  | 999 | 999 kits | Foreign Only:One distributor, DiaSorin S.R.I. in Italy.No wholesalers.No retailers.No domestic distribution, No Federal Government sales/distribution centers. | Terminated | March 11, 2011 | 375 | 999
Z-1473-2011 | 57083 | [6 empty] | GE ADVANTAGE SIM MD versions 5.x, 6.x, 7.x up to and including version 7.7.0 |  |  | November 24, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware that inconsistent DRR (Digitally Reconstructed Radiographic) visualization may occur when the blended DRR feature or the Isocenter Move mode feature is used in combination with specified other functions of the AdvantageSim application on Advantage Workstation. These inconsistencies may affect patient safety. Other AdvantageSim functions are not affected. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 23, 2010. The letter was addressed to Hospital Administrators / Risk mangers, Oncology Department Managers / Radiation Oncologists, &amp; Radiation Physicists / Dosimetrists / Therapists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For any questions or concerns regarding this recall, please call one of thefollowing phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. |  | 1531 | 1531 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MI, MA, MI, MN, MS,MO, NT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, POLAND, PAKISTAN, NOWAY, NIGERIA, NETHERLANDS, MEXICO, MALAYSIA, LITHUANIA, LATVIA, JAPAN, ITALY, ISRAEL, IRELAND, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, EGYPT, ECUADOR, DOMINICA REPUBLIC, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BRAZIL BELGIUM, AUSTRIA, AUSTRALIA. | Terminated | August 03, 2012 | 618 | 1531
Z-1476-2011 | 57692 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PrecisePLAN Treatment Planning System |  |  | January 28, 2008 | 2008 | Elekta, Inc. | PrecisePLAN generates digitally Reconstructed Radiographs (DRR;s) with a shift in the superior direction. | Other | Elekta sent an Important Notice A308 letter dated January 28, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Elekta representative directly if they had any questions. . |  | 277 | 277 units | Worldwide Distribution - USA and the countries of Bulgaria, Switzerland, Czech Republic, Germany, Spain, France, Hungary, Ireland, Italy, Liechtenstein, Netherlands, Poland, Portugal, Sweden, Slovak Republic, Australia, China, Hong Kong, and Japan, | Terminated | April 21, 2011 | 1179 | 277
Z-1474-2011 | 57419 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Aesculap MicroSpeed Uni Control Box, Part #GD670 |  |  | November 22, 2010 | 2010 | Aesculap Inc | Potential safety issue in that the control box can misinterpret the information sent from the low speed motor if the control magnet inside the motor changes its posthion. | Software design | AESCULAP sent an IMPORTANT CORRECTION NOTICE letter dated November 22, 2010, to all affected consignees. The letter identified the product, the problem, and the action the consignee should take. Consignees were instructed to not use the motors for thread cutting and removal of screws.If any unusual reactions of any cutting tools were observed consignees should check to ensure that the motor is not running in the opposite direction.If any error message is displayed, consigness were instructed to not use the motor.Consignees were also instructed to check both directions of running before use and do not use a motor which produces an error message in one direction. Aesculap would be providing all customers with a software update to the MicroSpeed Uni control box either by the unit being returned to Aesculap for update or by firm representative doing the upgrade on site at their facility.For questions regarding this recall call 800-258-1946 x5254 or 800-258-1946 x5267. |  | 91 | 91 units | Worldwide Distribution - USA including CA, CO, DE, FL, GA, IN, LA, MO, NC, NJ, OK, PA, TN, TX, and VA and the country of Canada. | Terminated | July 05, 2012 | 591 | 91
Z-1475-2011 | 57100 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Vu a-POD Intervertebral Body Fusion Device. Packaged as surgical kits containing all implants and instruments required for a procedure. Also packaged as individual instruments and implants in poly bags distributed as replenishment for surgical kits. Theken Spine, LLC, Akron, OH 44306. |  |  | October 19, 2010 | 2010 | Integra LifeSciences Corp. | The IFU provided with the device contains incorrect information in the Indications for Use Section, stating the following: The Theken Spine Vu a"POD Intervertebral Body Fusion Device is intended for use as a stand alone implant when used with the SpinPlate provided in the set. | Employee error | On 10/19/2010 the firm sent "URGENT : FIELD CORRECTION ACTION Letters to their customers. The letters identified the affected product and explained the reason for the recall. The letter also stated that the issue is not believed to have a health impact on patients. Physicians are asked to continue to exercise their clinical judgment with any periodic patient follow up. Attached to the letters were IFUs containing the correct information. Customers are to complete and return the enclosed Recall Acknowledgement Form per the instructions provided. Questions can be directed to Suzanne B. Wojcik, Sr. Dir Regulatory Affairs and Quality Assurance, at 330-835-8528. |  | 2163 | 2,163 pieces | Nationwide Distribution -- AL, AZ, CA, CT, ID, KS, MO, NH, NV, NY, OH, OK, TX &amp; UT. | Terminated | September 19, 2012 | 701 | 2163
Z-1497-2011 | 57884 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | High Viscosity Dermabond ProPen Topical Skin Adhesive 2-Octo Cyanoacrylate, Sterile, Rx only, Ethicon, Inc. |  |  | February 02, 2011 | 2011 | Ethicon, Inc. | Some units of Dermabond Topical Skin Adhesive and Dermabond ProPen Topical Skin Adhesive contain discolored product. | Nonconforming Material/Component | Ethicon sent their recall letters with business reply cards attached on February 2, 2011 via UPS overnight mail. Customers are informed of the issue and asked to return the attached business reply card. |  | 725722 | 725,722 each (total) | Worldwide distribution, including USA, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Czech Repub, Chile, China, Denmark, Dominican Repub, Egypt, Finland, France, French Guiana, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, ireland, Israel, italy, Japan, Korea, Lithuania, Luxembourg, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Palestine, Puerto Rico, Reunion, Russia, Senegal, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Ukraine, and Venezuela. | Terminated | November 07, 2012 | 644 | 1451494
Z-1494-2011 | 54620 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Male Driver 4.0mm (catalog # MDr140L) and Male Driver 3.2mm (catalog # MDr132L), an instrument used for the insertion of the Fassier-Duval Telescopic IM System intramedullary nail, for bone fixation. |  |  | January 29, 2010 | 2010 | Pega Medical Inc. | Possible breakage. Structural weakness has been detected on the Male Driver 3.2mm and 4.0mm sizes. In certain cases, this weakness could lead to a premature failure of the instrument if it is submitted to large torsional loads higher than the normal forces required to introduce the male component. | Process control | On 01/29/2010, Pega Medical began contacting their consignees by e-mail and telephone.Pega Medical followed up with their consignees by sending out the replacement components along with the explanation letter. The subject on the letter is Notice for 3.2 mm and 4.0 mm Male Drivers of the Fassier-Duval IM Telescopic System.The consignees were informed that there could be a structural weakness in some of the Male Driver instruments. This problem has been detected on the 3.2mm and 4.0mm sizes. In certain cases, this weakness could lead to a premature failure of the instrument if it is submitted to large torsional loads higher than the normal forces required to introduce the male component. Consignees can contact Pega Medical at 450-688-5144 x 242. |  | 12 | 4 units of MDr132L (3.2 mm) and 3 units of MDr140 (4.0 mm) distributed in US. | Components were distributed to four hospitals in the US in the state of GA, OK, TN, and UT. | Terminated | March 26, 2011 | 421 | 12
Z-1493-2011 | 57827 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Neotrode Wire Attached, Catalog Number: 1731-003; Neotrode II Wire Attached, Catalog Number: 1741-003; Softrace Neonatal Wire Attached, Catalog Number: 2321-003; Softrace Pediatric Wire Attached, Catalog Number: 2331-003; Softrace Limb Band, Catalog Number: 2310-003; TRU-LINK ECG Electrode (Neonatal Pre-Wired), Catalog Number: 685-0037-00; for use with electrocardiographic monitoring equipment to detect neonatal/pediatric heart action voltages. |  |  | September 23, 2010 | 2010 | ConMed Corporation | Poor signal or loss of signal. Certain recalled devices have a loose wire attachment that may lead to intermittent or poor signal quality. | Nonconforming Material/Component | ConMed Corporation distributed an Urgent Device Recall notification letter on Tru-Link ECG Electrode (Neonatal Pre-wired) dated September 23, 2010. ConMed Corp. has advised to cease the use of the devices immediately and conducted a recall to the user level. Direct consignees were asked to contact sub-accounts if recalled devices were further distributed and notified them of such recall. Recalled devices were to return to ConMed Corp. at 525 French Road, Utica, NY 13502 with completed Attachment II to indicate the amount returned. Customers with no recalled devices were asked to fill out Attachment II and faxed to 315-624-3225. Questions can be directed to Ms. Patricia Cotter, ConMed Recall Coordinator at 315-624-3237 or fax to 315-624-3225 or email to neorecall@conmed.com. |  | 9009000 | 9,009,000 Pediatric/Neonatal ECG Electrodes with Attached Wire | Worldwide distribution, including USA, Dubai, Ecuador, Egypt, Germany, Holland, Indonesia, Israel, Italy, Jamaica, Kenya, Korea, Lithuania, Malaysia, Mexico, New Zealand, Norway, Qatar, Saudi Arabia, Singapore, South Africa, Sweden, Switzerland, US Virgin Islands, Canada and Belgium. | Terminated | October 28, 2014 | 1496 | 9009000
Z-1379-2011 | 57566 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | XenMatrix Surgical Graft Sterile, Porcine Dermal Matrix |  |  | January 07, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Davol cannot guarantee that all Xenomatic product released from July to October 2010 meets FDA guidelines for endotoxin limits; and therefore, the grafts are potentially contaminated with bacteria. | Process control | Davol issued recall notification to accounts by letter dated 1/06/11 requesting recalled units be returned. contact Davol Customer Service at 1-800-556-6275 for instructions on how to return product to Davol. A fax sheet attached to the letter to be completed and fax to 1-401-825-8753. |  | 1456 | 1456 units | Nationwide | Terminated | October 31, 2013 | 1028 | 5227
Z-1518-2011 | 57552 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | LICOX CMP Tissue Oxygen Monitor; Catalog numbers REF AC31. Manufacturer GMS Gesellschaft fur medizinische, Sondentechnik mbH, Dorfstrasse 2, 24247 Mielkendorf, Germany; Distributed by Integra NeuroSciences, 311 Enterprise Drive, Plainsboro, New Jersey 08536 USA. |  |  | December 17, 2010 | 2010 | Integra LifeSciences Corp. | Certain LICOX CMP Tissue Oxygen Monitors could allow for underestimation of temperature values. | Nonconforming Material/Component | Integra Life Sciences notified their Integra Neurospecialist via email and their customers via Fed Ex overnight mail on December 17, 2010 with an Urgent: Product Recall letter. The letter identified the affected product and the reason for the recall. The letter states that an Integra Neurospecialist will contact customers to arrange for the return of their Licox monitor to the Service Center for recalibration. A loaner will be provided until the recalibration is complete. The attached "Recall Acknowledgement and Return Form" must be completed and returned to Integra. If there are any additional questions, customers should contact Gary Rabinovich at 609-936-6864. |  | 26 | 26 units | Worldwide Distribution | Terminated | November 05, 2012 | 689 | 26
Z-1522-2011 | 56660 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, ApexPro Telemetry System |  |  | January 28, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential for loss of patient monitoring when ApexPro Telemetry System is used in conjunction with a mirrored view CARESCAPE CIC Pro in combination monitoring mode (combo mode). The safety issue includes the potential for a delay in annunciation of audible alarms when used in combo mode, and may result in thedelay of treatment of a cardiac event. | Software design | GE sent an "Urgent Medical Device Correction" letter dated December 29, 2010. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Affected Product Details, Product Correction, Safety Instruction and Contact Information.Customers were instructed to contact Technical Support or their local Service Representative with any questions regarding this medical device.For any questions regarding this recall call 262-573-4122. |  | 2346 | 2346 | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO,MT,NE, NJ, NM, NY, NC, OH, OK, OR, PA, PR, SC, TN, TX, UT, VT, VA, WA, WV, WI, WY. and the UNITED KINGDOM, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH KOREA, RUSSIA, NORWAY, NEW ZEALAND, NETHERLANDS, MALAYSIA, JAPAN, ITALY, ISRAEL, IRELAND, GERMANY, FRANCE, FINLAND, DENMARK, CUBA, CHINA, CANADA, AUSTRIA, AUSTRALIA, POLAND. | Terminated | May 29, 2012 | 487 | 2346
Z-1521-2011 | 57847 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vari Lase WireFiber Laser Fiber, RED 7137, Rx ONLY, Sterile EO, Vascular Solutions Inc., 6464 Sycamore Court, Minneapolis, MN 55369 |  |  | January 10, 2011 | 2011 | Vascular Solutions, Inc. | Investigations of a recent Device Experience Report has made us aware of a potential problem with our Vari-Lase WireFiber, Model 7137, manufactured with the following lot numbers: 5500898 and 550896.Recently it was reported that the endovenous laser fiber cage and guidewire tip of the WireFiber became separated. It is possible that separation could occur during patient procedure and has the p | Nonconforming Material/Component | Vascular Solutions sent an "Urgent Medical Device Recall" letter dated January 10, 2011. The letter described the issue and the product affected. Advised consignees to immediately check their inventory , place affected product in a secure location, and to complete and return the Customer Inventory Form.For questions regarding this recall call 763-656-4210. |  | 41 | 41 | Nationwide Distribution - USA including CT, MA, MN, NY, VA | Terminated | May 03, 2012 | 479 | 41
Z-1498-2011 | 57303 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | SSH-880CV Aplio Artida; System, Diagnostic Ultrasound |  |  | September 28, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on the SSH-880CV Aplio Artida Diagnostic Ultrasound System due to software issues. | Software design | Toshiba America Medical Systems (TAMS) sent an URGENT MEDICAL DEVICE CORRECTION letter to consignees via USPS return receipt mail. The letter contained a return reply form that is to be faxed to TAMS at 877-349-3054 as proof of notification.Customers were instructed to share the information with all users and reviewing radiologists as well as clinical engineering or biomedical group at their facility.For questions regarding this recall call 800-421-1968 or 800-521-1968. |  | 40 | 40 units | Nationwide Distribution to CA, GA, IA, IL KY, LA, MA, MN, MT, OH, OR, PA, TX, WI, and Puerto Rico. | Terminated | May 02, 2012 | 582 | 40
Z-1520-2011 | 57841 | [6 empty] | VCareÂż Vaginal-Cervical Ahluwalia's Retractor-Elevator devices. The VCare IFU is packaged one per carton of eight VCare Pouched Devices. The Instructions for Use is an 8.5" x 12" 12-page booklet and is packed loose in the carton. Only the carton will have to be opened to replace the old IFU; there is no need to open the sterile packaging of the individual VCare pouches. |  |  | February 02, 2011 | 2011 | ConMed Corporation | Product manufactured after 4/11/2008, but before 8/01/2009, contains the earlier Instructions For Use, IFU (PN 17416, Rev. D) which does not contain the revisions on the addition of Warning Notes 10 and 11 and the expansion of Warning Note 6 to the VCareÂż Instructions for Use. The reason for the field correction is to provide a copy of the revised IFU to all customers who may have unexpired VcareÂż | Device Design | ConMed Corporation sent a communication to provide its direct consignees with important supplemental information concerning the Instructions for Use (IFU) for the VCAREÂż Vaginal-Cervical Ahluwalia's Retractor-Elevator uterine manipulators. The direct consignees were instructed to place a copy of the enclosed Instructions for Use, Part No. 14071 (Rev. A), with any inventory in your possession with an expiration date prior to August 2011 (2011-08) and replace any other Instructions for Use in thier possession. A copy of the enclosed Instructions for Use is already packaged with all devices bearing later expiration dates. Any inventory in their consignees possession with an expiration date prior to January 2011 (2011-01) is expired and should not be used. If the devices were further distributed to other facilities or distributors, or their consignees are a distributor, they were instructed to forward this information as appropriate to any known users of these devices.For questions regarding this recall call 315-624-3225. --- The IFU Warning changes are as follows: A new warning statement # 6 asks the user to verify the presence of all the components of the device upon removal of the VCAREÂż device from the patient. Although the prior IFU contained warnings about verifying that all components of the device were properly removed following use during surgery, the new statement expands the earlier, briefer statement. 6. Upon removing VCARE, the surgeon should visually inspect the VCARE device, and the patient, to make sure that the entire VCARE device was properly removed and that no components or fragments of these components were retained in the patient. There are 5 parts/components to the VCARE Cervical Elevator Retractor. These are: 1) the balloon; 2) the forward "cervical" cup; 3) the back or vaginal cup; 4) the locking assembly with thumb screw; 5) the metal shaft and handle with balloon inflation valve. --- A new warning statement # 10 further addresses |  | 31028 | 31,028 VCareÂż Vaginal-Cervical Ahluwalia's Retractor-Elevator devices | Worldwide Distribution - USA including NY, TX, CA, IA, AZ, PA, IN, and KS and the countries of Austria, Jordan, Turkey, China, Portugal, Chile, Belgium, Ecuador, Brazil, Romania, Japan, Malaysia, India Switzerland, Colombia, Netherlands, Spain, Saudi Arabia, France, and Canada. | Terminated | August 30, 2012 | 575 | 31028
Z-1501-2011 | 57702 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical's LAP CHOLE KIT 5 GK; Model Number: GK105 |  |  | September 10, 2010 | 2010 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall of specific lot numbers of the Direct Drive Laparoscopic Clip Applier, model number CA090, due to a potential defect in the jaws, possibly resulting in jaw breakage. | Process design | Applied Medical sent an URGENT: MEDICAL DEVICE RECALL letter dated September 10, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to examine their inventory and complete the attached Voluntary Recall Notification Confirmation Form and fax it to 949-713-8832. Customers were asked to return all affected product.Secondary consignees were notified by Applied Medical on September 16, 2010. For any questions regarding this recall call 949-713-8041. |  | 2843 | 2,843 units for all products | Worldwide Distribution - USA including AL, AR, AZ, CA, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OR, PA, SC, TN, TX, VT, WA, and WI and the countries of Australia, Europe, Japan, and Kuwait. | Terminated | March 09, 2011 | 180 | 2843
Z-1562-2011 | 54551 | No | JNB | Phenylalanine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Quantase Neonatal Elution Buffer, a component of Quantase Neonatal Accessory Kits (#1, 2 and 5), Model number 532-5035, included in 532-5036, 532-5037 and 532-5040, Manufactured by Bio-Rad Laboratories, Hercules, CA |  |  | January 26, 2010 | 2010 | Bio-Rad Laboratories Inc | Pressure contained within the Elution Buffer vial may expel the rubber stopper, causing the potential for injury. | Other | Direct accounts were notified by phone, followed by fax (domestic). Urgent Device Recall notification letters were sent out on January 26, 2010 to subsidiary along with a Bio-Rad customer form to be send to each customer. The letters identified the affected product and explained the reason for the recall. Customers were asked to examine their inventory and quarantine the affected lots listed. In addition, they are to complete and return the enclosed response form according to the instructions provided. Customers are contact their regional Bio-Rad office for Elution Buffer replacement. |  | 714 | 714 | Worldwide Distribution -- USA, including states of MI and MN and Bio-Rad facilities in Italy, France, Czech Republic, Canada, Argentina, Panama, United Kingdom, Hong Kong, China, Canada, Singapore, India, South Korea, Brazil, and Sweden. | Terminated | March 17, 2011 | 415 | 714
Z-1523-2011 | 56344 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens brand ARTISTE System Digital Linear Accelerator with the syngo RT Therapist version 4.1 Connected to MOSAIQ, Part No. 8168754, Distributed by Siemens Medical Solutions Oncology Care System, Concord CA |  |  | June 29, 2010 | 2010 | Siemens Medical Solutions USA, Inc | A software problem that has the potential for data loss during the transfer of treatment records to the OIS, which may not be recorded in MOSAIQ, and subsequently may lead to mistreatment. | Other | Siemens sent Urgent Medical Device Correction, Safety Advisory Notices to all affected customers on June 29, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to include the Safety Advisory Notice in their System Owner Manual, Chapter, "Safety Advisory Letter.For questions customers were instructed to contact their Siemens Service Engineer or Application Specialist.For questions regarding this recall call 925-602-8157. |  |  |  | Worldwide Distribution - USA including NC,OH, and the countries of Belgium, Sweden, Germany, Lebanon, Japan, Australia, and Iran. | Terminated | March 09, 2012 | 619
Z-1559-2011 | 57500 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Wall Mounting Plate, an accessory to the Gendex eXpert DC (REF 110-0205G1) and 765DC (REF 110-0155G1)Intraoral X-Ray Systems; Gendex Dental Systems, Des Plaines, IL USA |  |  | December 20, 2010 | 2010 | Gendex Dental Systems | The X-Ray unit may break loose from the mounting plate and fall off the wall due to defects in the mounting plate. | Nonconforming Material/Component | The firm, Gendex, sent an "URGENT MEDICAL DEVICE FIELD CORRECTION" letter dated December 20, 2010 to their customers. A revised notification letter dated January 20, 2011, was sent to the customers providing further clarification on instruction to correct the problem, as well as a template for end user notification. The letters described the product, problem and actions to be taken. The customers were instructed to contact their customers, inform them of the issue and replace the wall-mounting plate. The customers were also instructed to complete and return the attached acknowledgement form via fax to 215-997-5665 Attn: Regulatory Affairs or email to: regulatory.affairs@gendex.com and contact Gendex Customer Service at 1-800-323-8029.If you have any questions or concerns, contact the Regulatory Affairs Associate at 847-364-3290 or regulatory.affairs@gendex.com. |  | 34 | 34 wall plates | Worldwide distribution: USA and countries including: Australia, Canada and Germany. | Terminated | August 14, 2012 | 603 | 34
Z-1539-2011 | 57605 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cylindrical Bullet Nose Reamer 14.0 mm x 250 mm, REF 475816, Biomet Orthopedics 66 East Bell Drive Warsaw IN. |  |  | December 23, 2010 | 2010 | Biomet, Inc. | The firm was notified by their supplier that during manufacture of this instrument, the 12 degree cutting edge was not created. Testing found that if the cutting edge is missing, the reamer will not perform as intended. | Other | The firm, BIOMET, sent an" URGENT MEDICAL DEVICE RECALL NOTICE" dated December 23, 2010 to its customers. The notice described the product, problem and actions to be taken by the customer. The customers were instructed to immediately locate, discontinue use of product and return the product to Biomet. Biomet distributors were told to notify hospital personnel of the recall if they had distributed the product to hospital customers. The Customers/Distributors were instructed to confirm receipt of the notice by calling 800-348-9500 extension 3755 and 3983 and to complete and return the Response Form included with the recall notice via fax to 574-372-1683.Questions related to this notice should be directed to 574-371-3755 or 574-372-3983, Monday through Friday, 8am to 5pm. |  | 51 | 51 all products | Worldwide distribution: USA including states of: FL, OH, PA, IN, CA, and NJ; and countries of Argentina, Canada, Mexico, Netherlands and Spain. | Terminated | October 01, 2012 | 648 | 51
Z-1548-2011 | 57563 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK12 Defibrillator/Monitor. |  |  | October 01, 2010 | 2010 | Physio Control, Inc. | L1FEPAK 12 Operating Instructions that describethe process for disconnecting the device from AC (wall) power are not always being followed. When theinstructions are not followed and the Power ON button is pushed within 2 seconds after disconnectingthe device from AC power, the service indicator will illuminate and the device may fail to charge the defibrillator and/or initiate pacing. | Device Design | Physio Control, Inc. began mailing the URGENT-MEDICAL DEVICE INFORMATION letter to their consignees on October 1, 2010. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were advised to follow the current LIFEPAK 12 Operating Instructions-Section 7.5-AC Power Adapter Operation and wait at least 2 seconds to push the Power ON button after disconnecting the device from AC power. If the instructions are not followed and the Power ON button is pushed within 2 seconds after disconnecting the device from AC power and the service indicator illuminates, it's important to be aware that turning the device off, then back on again, will reset the device and restore full functionality.The complete copy of the current LIFEPAK 12 Operating Instructions can be downloaded from the company's website at www.physio-control.com/products/ defibrillators, Select LIFEPAK 12, Select Documents.The consignees can call the firm's Technical Support Team at 1-800-442-1142, option 5, between 6:00 A.M. to 4:00 P.M. (Pacific), Monday - Friday. |  | 39 | 39, 180 units (17168 units internationally and 22012 units in US) | Worldwide Distribution - USA and the countries of ARGENTINA, AUSTRALIA, BAHRAIN, BERMUDA, BOLIVIA, BRAZIL, BRUNEI, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, EGYPT, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, LEBANON, MALAYSIA, MALTA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PUERTO RICO, QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SRI LANKA, TAIWAN, THAILAND, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UAE, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN, and ZIMBABWE.Devices were distributed to the following countries:ARGENTINA, AUSTRALIA, BAHRAIN, BERMUDA, BOLIVIA, BRAZIL, BRUNEI, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, EGYPT, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, LEBANON, MALAYSIA, MALTA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PUERTO RICO, QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SRI LANKA, TAIWAN, THAILAND, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UAE, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN, and ZIMBABWE. | Terminated | April 05, 2012 | 552 | 39
Z-1556-2011 | 57693 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PrecisePLAN Treatment Planning System |  |  | February 05, 2008 | 2008 | Elekta, Inc. | When using the 'Beam Edit" function to modify a bean in an unlocked Fraction Group, there are two conditions that may cause the user to unitentionally modify a locked fraction group. | Other | The firm, Elekta, sent an "Important Notice A312" titled "Unintentional Modification of a locked fraction group in PrecisePLAN" dated February 5, 2008 to all affected customers. The notice described the product, problem and action to be taken. The customers were instructed to use care when using the "Make Opposed" function and when changing the Fraction Group of a beam. The customers were also advised to review the final plan and any locked fraction groups to make sure the planned beam arrangement is what was intended. Note: A Field Change Order 540 073 "All PrecisePLAN 2.15 users" was issued on September 30, 2008 to all customers.If you have any questions, call 770-670-2548 or +44(0)1293 654400. |  | 282 | 282 units | Worldwide distribution: USA and country of: Mexico. | Terminated | April 07, 2011 | 1157 | 282
Z-1549-2011 | 57589 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Precedence Imaging System; 16 Slice 5/8, 16 slice 3/8; |  |  | December 28, 2010 | 2010 | Philips Medical Systems | Philips Healthcare Nuclear Medicine determined that the detector arm assembly could be compromised. | Other | Urgent - Medical Device Correction notifications were sent to all domestic consignees on December 28, 2010 via Federal Express. The letter identified the affected products, the reason for the recall, the hazard involved, and actions to be taken by the customer/user. Customers have been instructed to immediately stop using the system and contact a Philips service representative for an immediate inspection if one of the listed reasons apply. Customers were also instructed to follow the instructions in the system's Instructions For Use to avoid a collision or force on the detector. The firm's service will contact customers to arrange a time for Philips to conduct an inspection and to conduct the appropriate field safety correction. The letter should be placed in customers' Instructions For use. Notices to international consignees will be sent as soon as translation is complete. If customers need further information or support concerning this issue, they should contact their local Philips representative. |  | 540 | 540 units, all varieties | Worldwide Distribution | Terminated | April 02, 2012 | 461 | 540
Z-1555-2011 | 57688 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Desktop Pro 4.2, 5.0, 5.1, 6.0, 6.2 |  |  | August 23, 2007 | 2007 | Elekta, Inc. | Under certain conditions it has been found that movement of the collimators can be induced during multi segment beams such as IMRT and Omniwedge which will not be detected by the Desktop Pro if high values are sent within the iCom tolerance table. | Other | The firm, Elekta, sent an "Important Notice A302" dated August 23, 2007 titled "Update to iCom Tables for Desktop Pro when used with Third Party R&V; Systems",to all affected customers. On September 28, 2007, Elekta issued a second Important Notice A304 which reiterated the same information as was in Important Notice A302. The notices described the product, problem and actions to be taken. This notice provides a workaround on how to avoid the issue. In order to avoid this issues documented, Elekta recommended that the customers set the tolerance for collimators in the iCom tolerance table to be 3.0 degrees-no less than 3.0 and to use the Linac Record, if configured, as a means of checking the final collimator position. Elekta noted that the issue will be resolved in a future release of Desktop Pro.If you any questions related to this issue, please contact your local Elekta, third party R&V; system representative or Elekta Limited at +44 (0)1293 654200. |  | 229 | 229 units | Worldwide distribution: USA including Puerto Rico and countries including: Canada and Mexico. | Terminated | April 07, 2011 | 1323 | 229
Z-1525-2011 | 57854 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | TI ANCHOR 2.5MM W / 2-0 SUTURE, G-FORCE (R) SUTURE ANCHOR, Tissue Anchor, REF 86TA-2520, Rx ONLY, Sterile EO, Wright, Arlington, TN USA |  |  | January 11, 2011 | 2011 | Wright Medical Technology Inc | Driver may break during use. | Other | Wright Medical Technology Inc initiated their recall of the product by telephone to their distributors on January 11, 2011. The firm followed with letters sent on February 1, 2011, to distributors, hospital administrators, and surgeons. Distributors were asked to immediately complete the attached response form, confirming receipt of the notice, and return a copy to Wright Medical by fax to 901-867-4788. For questions they were instructed to call 800-874-5630.For questions regarding this recall call 901-867-4788. |  | 45 | 45 units | Nationwide Distribution including FL, NC, OH, IL, IN | Terminated | July 11, 2011 | 181 | 55
Z-1526-2011 | 57851 | [6 empty] | Fresenius 2008T Hemodialysis Machine |  |  | January 19, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Cursor on the dialysis screen may disappear and not respond to user input | Device Design | Fresenius Medical Care contacted customers by mailing a Customer Notification letter "Urgent correction FMNCA Model 2008T Hemodialysis Machine" via Certified-Mail, Return Receipt requested on January 19, 2011. The Technical Services department began contacting consignees by telephone on January 13, 2011, to schedule an appointment for installing the electronic board on the 2008T machines at the customers' facilities. For questions call Technical services 1-800-227-2572. |  | 1268 | 1268 units | Worldwide Distribution - USA and Canada | Terminated | July 08, 2013 | 901 | 1268
Z-1557-2011 | 58008 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | AquaRefresh, Multi-Purpose Solution, 2 fl.oz/60mL, Made in the Netherlands, Distributed by Marietta Vision, 397 N. Sessions Street, Marietta, GA 30060. |  |  | February 21, 2011 | 2011 | Marietta Vision | Product is misbranded; Product labeled as, Made in the Netherlands, was actually manufactured in Mexico. | Other | John Patterson, President of Marietta Vision called each optical retailer on February 21, 2011, and told them of the misbranding issue in regards to country of origin. In addition, an Urgent Product Recall letter was sent to each account on February 23, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their AquaRefresh inventory for lot number KD005, and discontinue distribution of the samples immediately and to return any outstanding stock to Marietta Vision for exchange. Marietta Vision would return the misbranded product back to the Mexican factory at their expense to have the product packaging corrected. If they elect not to do so, Marietta Vision would have the product disposed of.Customers were instructed to fax the enclosed form back to Marietta Vision at 678-797-1099.For questions regarding this recall call 866-300-6257. |  | 1000 | 500/2 fl.oz. Bottles | Nationwide Distribution including AR, AZ, CA and TX. | Terminated | November 03, 2011 | 255 | 1000
Z-1561-2011 | 57974 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Pulmonary Hemi-Artery SG |  |  | February 17, 2011 | 2011 | CryoLife, Inc. | Staphylococcus aureus (not methicillin resistant) identified in companion tissue. | Other | CryoLife Inc. issued a letter to the physician on February 17, 2011, advising him of the results of the testing performed on a companion cardiac allograft. The letter did not recommend explant of the allograft or special antimicrobial treatment of the patient unless medically necessary.The physician was asked to complete and return the enclosed, stamped, self-addressed postcard affirming the receipt of the notification.For any questions the physician was advised to call 770-419-3355. |  | 1 | One | Nationwide Distribution including AZ &amp; NV | Terminated | April 24, 2012 | 432 | 1
Z-1554-2011 | 57988 | No | OIG | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Medline Industries, Inc. "Cedar" brand Vinyl Economy Exam Glove, Powder-Free, Large Size, Catalog Number: MSV403P, Lot Number: JCU04958569 |  |  | February 03, 2011 | 2011 | Medline Industries Inc | The recall was initiated because the affected lot, lot number JCU004958569, of gloves did not meet the requirements for the water-leak test. | Release of Material/Component prior to receiving test results | MEDLINE INDUSTRIES, INC. sent a RECALL NOTICE IMMEDIATE ACTION REQUIRED letter by US Mail on February 10, 2011, to all affected cusomters. The letter identified the product, the problem, and the action to be taken by the customer. The customers were instructed to:1. Immediately quarantine ALL inventory of the affected lot for return to Medline. 2. Complete the attached recall verification form. Please complete this form EVEN IF you have none of the affected lot in stock.3. Fax or email the completed form to Medline. The email address is Recalls@Medline.com or faxed the form to 866-767-1290. Customers with questions regarding the recall were instructed to contact Medline at 1-866-359-1704 or via e-mail at Recalls@Medline.com. |  | 696 | 696 cases (696,000 units) | Nationwide Distribution including CA, HI, and NV. | Terminated | March 19, 2012 | 410 | 2796
Z-1558-2011 | 57696 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Polaris Spectra Position Sensor Unit, a component of the StealthStation S7 System, Medtronic Navigation, Louisville, CO 80027. |  |  | December 22, 2010 | 2010 | Medtronic Navigation, Inc | Position Sensor Unit may generate tracking errors or may stop tracking immediately. | Component design/selection | The firm, Medtronic, sent a "Product Correction Notification" letter dated December 22, 2010, to all US Consignees/Customers on December 23, 2010 and foreign Consignees/Customers on December 27, 2010. The letter described the product, problem and actions to be taken. The customer were instructed to discontinue use of the impacted navigation system(s) until a replacement camera (PSU) is installed. A Medtronic Navigation Representative will visit each facility to remove the navigation camera and install a replacement. If you have any questions, contact your Medtronic Navigation Representative or Medtronic Navigation Technical Support services at 1-800-595-9709 or +1-720-890-3200. |  | 11 | 11 (US), 8 (Intnl) | Worldwide distribution: USA including states of: CA, CO, GA, SC, TN, and WI; and countries including: Australia, Canada, Chile, Russia, Saudi Arabia, and South Africa. | Terminated | March 11, 2011 | 79 | 11
Z-0905-2011 | 58077 | No | IZO | Diagnostic x-ray high voltage generator. | Radiology | Radiology | 510(K) Exempt | BuckyDiagnost general purpose x-ray system with Optimus Generator. |  |  | January 17, 2011 | 2011 | Philips Healthcare Inc. | Affected systems were determined not to prevent run-on of a selected technique factor if the technique selector and the exposure control were activated simultaneously. | Radiation Control for Health and Safety Act | A Customer Information BuckyDiagnost Diagnostic X-Ray System letter, dated January 6, 2011, was sent to customers. The letter identifies the affected product, the reason for the recall, actions that are to be taken by customers, and actions that are planned by the firm. Customers are advised to make sure that the desired exposure values are displayed and that they have released the technique button BEFORE they start the exposure by pressing the exposure switch whenever manually changing the exposure parameters and in order to avoid the problem. Philips will issue a Field Change Order (FCO) to resolve the issue by updating the system software free of charge. Philips will contact customers to implement the FCO. Further information or support concerning this issue can be obtained by contacting your local Philips representative. The Field Service Engineers began software implementation the week of March 21, 2011. |  | 1810 | 1810 units in the US | Worldwide Distribution | Completed |  |  | 1810
Z-1568-2011 | 56120 | No | HQG | Prescription spectacle lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | Polycarbonate prescription single vision lenses for use in sport goggles. |  |  | June 30, 2010 | 2010 | Luxottica Optical Manufacturing | The polycarbonate prescription lenses may not meet the frame manufacturer's minimum lens thickness specification of 2.0 millimeters. LNAD made the decision to remove these sport goggles from the market because of the potential risk that the lenses could break or disengage from the frame in the event of an impact. | Other | Luxottica Optical Manufacturing sent notification letters on June 30, 2010, to their customers informing them of the problem. Customers were instructed to complete the enclosed Sport Goggle Return Form, and to indicate whether they would like a replacement pair of goggles or a refund.Place the completed Sport Goggle Return Form and their sport goggles (without the sport goggle case) in the envelope provided and return it to the firm via the United States Postal Service. Customers were instructed not to return to the retail store where purchased. If customers were unable to return their sport goggles because they were lost or discarded, customers should complete the appropriate information on the Sport Goggle Return Form, being sure to indicate whether they would like a replacement pair of goggles or a refund, and then sign the form. Place the completed Sport Goggle Return Form in the envelope provided and return it to the firm via the United States Postal Service. For any questions, call 1-877-488-2407. They instructed their customers to return the product for replacement or refund. |  | 2916 | 2,916 individual lenses (1,458 pair of sport goggles) | Worldwide Distribution - USA including the following states: AK, AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI &amp; WV; and the country of Canada. | Terminated | September 19, 2012 | 812 | 2916
Z-1592-2011 | 56994 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Straumann coDiagnostiX, dental implant planning system. |  |  | September 30, 2010 | 2010 | Straumann Manufacturing Inc. | Software: Surgical Error may occur when the implant length is changed after positioning the surgical sleeve. | Software design | Straumann notified accounts of this correction by letter datedSeptember 30, 2010 and delivered via Federal Express Delivery Service. The upgrade is available on-line by logging into the coDiagnostiX system and enabling the automatic upgrade to occur. |  | 20 | 20 systems | Nationwide | Terminated | September 19, 2011 | 354 | 20
Z-1594-2011 | 57915 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | ARCOS Modular Revision System Torque Limiting T-Handle 55 inch- pounds, REF 31-301850, Biomet Orthopedics Warsaw, IN. |  |  | January 28, 2011 | 2011 | Biomet, Inc. | The firm received two complaints of the tabs on the Torque Limiting T-Handle fracturing during surgery. The firm performed internal testing and found that the tabs may fracture when used in the "OFF" or "locked" setting for hand-reaming the femoral canal. | Device Design | Biomet, Inc. made telephone calls to distributor accounts on January 28, 2011, with specific instructions that the Arcos Torque Limiting T-Handle was not to be used in the "OFF" position. Distributors were required to use the Arcos Torque Limiting T-Handle on the "55" setting for insertion of locking screws and troch bolt. Biomet, Inc. sent an URGENT MEDICAL DEVICE RECALL NOTICE letter dated February 21, 2011, to their distributors. The letter stated that the firm was currently manufacturing replacement handles which would be distributed to the field shortly. Consignees were advised to immediately discontinue use of the T-Handle in the "OFF" position.Send back the current T-Handle as soon as they receive a replacement.Carefully follow the instructions on the enclosed "FAX Back Response Form."Fax a copy of the Response Form to 574-372-1683 prior to return of product.Use priority carrier for shipment.If the product has been further distributed, consignees must notify hospital personnel of this action, via the enclosed "Dear Bioment Customer" notice. This letter must be give to hospital personnel responsible for receiving recall notices. Consignees are charged with the location and return of these products. Consignees were instructed to confirm receipt of this notice by calling, 800-348-9500, ext. 3755 or 3983. For any questions call 574-371-3755 or 574-372-3983. |  | 40 | 40 units | Worldwide Distribution - USA including the states of CA, NY, FL, VA, UT, IL, MI, KY, OH, MA, PA, SD, MO, TX, NV, Puerto Rico, and the countries of Australia, New Zealand, Canada. | Terminated | July 27, 2011 | 180 | 40
Z-1572-2011 | 57732 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | January 11, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The potential for injury to a patient should the treatment data be incorrectly uploaded to patient records. | Software design (manufacturing process) | Sent an Urgent Medical Device Correction Safety Advisory Notice to customers on 1/11/11. Letters will be delivered by Siemens Customer Service Group or sent by certified mail. Letter identified the affected product and stated there is a potential risk due to a problem identified by the manufacturer. Customers are asked to refer to the attached Safety Notice from Elekta IMPAC Software. A Siemens Service Engineer will contact customers site to arrange for a patient matching criteria audit to determine if you have a problem. Customers should include the Safety Advisory Notice including the Safety Notice from Elekta IMPAC Software in their product related documentation until further notice. If customers have any questions, they should contact their responsible SIEMENS Service Engineer or Application Specialist. |  | 10 | 10 units | International Only Distribution -- Germany and Japan | Terminated | April 19, 2012 | 464 | 10
Z-1593-2011 | 57519 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Hemashield Finesse Ultra-Thin Knitted Cardiovascular Patch |  |  | December 07, 2010 | 2010 | Maquet Cardiovascular, LLC | One unit of Hemashield Finesse Ultra-thin Knitted Cardiovascular Patch which did not meet the internal release criteria for water permeability was distributed. | Other | The firm, Maquet, contacted their sole consignee/customer via telephone informing the customer of the recall and instructing the customer to put the unit to the side. On December 8, 2010 a sales representative visited the site to ensure the unit had been put to the side and to ensure the unit was shipped back to Maquet. If you have any questions, call the Manager, Regulator Affairs at 973-709-7487. |  | 1 | one unit | Nationwide distribution: SC only. | Terminated | September 12, 2012 | 645 | 1
Z-1591-2011 | 57432 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Anchor Tissue Retrieval System; ETO sterilized, single use disposable device; port size 15 mm, 1550 mL capacity; 3 units per box; Manufacturer: Anchor Products Co., 52 Official Rd, Addison, IL 60101 U.S.A.; Ref TRS175SB |  |  | December 09, 2010 | 2010 | Anchor Products Company, Inc. | The metal component within the unit can break and cause a malfunction. Use of the device may represent a potential health hazard if the metal component is not removed after breakage. | Device Design | The firm, Anchor, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated December 8, 2010 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory and return any remaining units to their distributor for replacement; and complete and return the attached ANCHOR RECALL RETURN RESPONSE FORM via fax to (314) 961-4535 or mail to: Progressive Medical, Inc., Attn: Marketing Manager, 11085 Gravois Industrial Court, St. Louis, MO 63128, even if they have no inventory. A representative of Anchor Products then visited each of the direct accounts to help them implement their subrecall of the devices from their customers.For additional information please contact: Vice President, Anchor Products Company at 1-800-543-9124 or 1-630-543-9124; and the distributor at 1-800-969-6331 or 1-314-961-4535. |  | 2256 | 2,256 units | Worldwide distribution: USA including state of: Missouri; and countries including: Australia, Belgium, Canada, France, Greece, Ireland, Japan, Spain and the United Kingdom. | Terminated | January 20, 2012 | 407 | 91041
Z-1580-2011 | 57396 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | PolyAx Locking Plate System, POLYAX FEMORAL PLATE - RIGHT, 15 hole, Sterile, REF 8141-30-115, Depuy Orthopaedics, Inc., Warsaw, IN. |  |  | November 01, 2010 | 2010 | Depuy Orthopaedics, Inc. | The firm determined that the 8.0 mm hole threads in the PolyAx Distal Femoral Locking Plate were incorrectly manufactured, which prevents the mating screw from locking in place. The screw should lock with the plate, instead the screw passes through the plate. | Other | Urgent Information - Recall Notice letters, dated November 1, 2010 were sent to customers. The letters described the reason for recall and identified the product affected along with the affected units. The letters also provided clinical guidance and stated that the problem had been corrected and that DePuy was in the process of replenishing its inventory. Hospitals were to determine if any of the recalled components they received were still on hand, and return them immediately to their DePuy Representative or directly to DePuy for credit. |  | 246 | 246 all products | Nationwide Distribution -- ME, MA, FL, VA, CO, TX, WY, AZ, MO, WI, GA, PA, AL, NY, OH, NJ, NC, and CA. | Terminated | June 27, 2012 | 604 | 246
Z-1569-2011 | 57314 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | MONOJECT Blood Collection Tubes containing heparin |  |  | November 15, 2010 | 2010 | Covidien LP | The heparin sodium USP used to manufacture blood collection tubes was produced with crude heparin containing Over-Sulfated Chondroitin Sulfate. | Material/Component Contamination | Covidien sent an URGENT PRODUCT RECALL letter dated November 12, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop using the product, and return it to the firm. Distributors were instructed to notify their customers. All customers were required to respond to the notice regardless of their remaining inventory. Responses will be tracked for follow-up and effectiveness checks. For Customers Service questions regarding this recall call 1-800-962-9888 option 1. |  | 18100 | 18,100 | Worldwide Distribution - USA including AL, AZ, CA, CO, DE, FL, GA, IA, ID, IL, KS, KY, MA, MI, MN, MO, MS, NC, ND, NE, NJ, NM, MV, MY, OH, OK, PA, SC, SD, TN, TX, WA, and WI and the country of Canada | Terminated | December 04, 2015 | 1845 | 176100
Z-1589-2011 | 57965 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Biomet INNERVUE DISPOSABLE SCOPE KIT 1.2MM WITH CANNULA SET, PRODUCT NUMBER 922141, STERILE |  |  | February 10, 2011 | 2011 | Biomet, Inc. | The cannula could not be connected to the scope and scopes disengage during use. The outer hub of Innervue scopes may become detached from the inner hub of the scopes during assembly of the scope to the Innervue hand piece. This condition may prevent the Innervue scope from being securely attached to the Innervue hand piece. | Process design | On 2/10/2011, Biomet issued an "URGENT MEDICAL DEVICE NOTICE" to their Direct Accounts instructing customers to identify and discontinue use of the recalled products and return them to Biomet. Response forms were provided. |  | 264 | 264 | Distribution USA nationwide and to the Netherlands | Terminated | August 27, 2013 | 929 | 607
Z-1566-2011 | 57969 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 100/120V System 1 Base, REF 801763, Terumo Cardiovascular Systems Corp., Ann Arbor, MI. |  |  | February 14, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (CVS) has received 19 reports of a circuit breaker for the Terumo Advanced Perfusion System 1 tripping and causing the system to switch to battery back-up power. In two cases, the incident was caused when the circuit breaker came into incidental contact with the base of an IV pole; in others there was no known cause. Terumo CVS' investigation has determined that | Device Design | Terumo sent an URGENT MEDICAL DEVICE CORRECTION letter dated February 14, 2011, addressed to Chief of Perfusion, Department of Cardiovascular Surgery or Director of Operationd Room Services. The letter explains the problem with the circuit breakers and hazards of the problem. The letter states that a service technician will visit the facility to install a guard. The service will be performed at the next scheduled preventative maintenance visit for those consignees with a service contract. For customers without a service contract, Terumo CVS will contact the consignees to schedule an appointment for the correction.Customers were instructed to review the Medical Device Correction notice.Assure that all users are aware of the notice.Confirm receipt of the communication by faxing the attached Customer Response Form to the fax number indicated on the form. For questions or concerns customers were instructed to contact Terumo CVS Customer Service at 1-800-521-2818. |  | 728 | 728 devices. | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IN, KY, LA, MA, MD, ME, MI, MO, MT, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and WV and the countries ofCANADA, COLUMBIA, JAPAN, RUSSIA, EGYPT, TURKEY, PAKISTAN, KUWAIT, COSTA RICA, TAIWAN, SAUDI ARABIA, INDIA, SOUTH KOREA, JORDAN, THE PHILLIPINES, INDONESIA, HONDURAS, DOMINICAN REPUBLIC, REPUBLIC OF GEORGIA, HONG KONG, UNITED ARAB EMIRATES, SINGAPORE, AUSTRALIA, MALAYSIA, TAIWAN, GERMANY, BELGIUM, MEXICO, THAILAND, VEITNAM, SOUTH AFRICA | Terminated | March 20, 2012 | 400 | 1113
Z-1577-2011 | 57885 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Automatic Cutting Needle, Soft Tissue Biopsy Needle, NAC-1825M, 18ga (1.3mm) x 25cm (10), Remington Medical, Inc. |  |  | February 07, 2011 | 2011 | Remington Medical Inc. | There is a potential for the needle hubs to break during use. | Device Design | Remington Medical consignees were notified via facsimile, email and certified mail starting 02/07/2011. Distributors are requested to communicate this recall to their customers. If a customer does not respond to the recall notice within one week of the recall notification, the customer will be contacted by telephone. If this is unsuccessful, the customer will be sent a second Recall Notice letter and Customer Inventory Reconcile Fax by certified mail by 28 February 2011. If there is no response at that point, RMI will consider the customer's inventory depleted. All return product for this recall will follow RMIs Returns Goods Authorization procedure (SOP7-200). All product returns and customer responses will be documented. Upon completion of the recall, the status will be updated and submitted for the FDA's recall closure process. |  | 8161 | 8161 units | U.S. Nationwide | Terminated | August 12, 2011 | 186 | 46878
Z-1565-2011 | 57944 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | VersaCare P500 Microclimate Management Surface Upgrade Kits P3218A and P3219A Hill-rom Batesville, IN. |  |  | February 21, 2011 | 2011 | Hill-Rom, Inc. | Hill-Rom has become aware of a potential issue with some TotalCareÂż and VersaCareÂż beds equipped with the integrated P500 Microclimate Management(MCM) therapy surface, and with the VersaCareÂż P500 MCM surface upgrade. Because of an unauthorized vendor part modification, it is possible that the MCM blower may not perform as originally designed. Given this modification, the MCMblower may experi | Other | Hill-Rom sent an "IMPORTANT MEDICAL DEVICE CORRECTION" notice dated February 21, 2011, to Facility Risk Manager/Facility Administrators. The letter states that because of an unauthorized vendor part modification, it is possible that the MCM blower may not perform as originally designed. Given this modification, the MCM blower may experience the following: delayed start, slow running, or failure to operate.This issue does not pose a hazard to patients as long as normal patient monitoring is in effect in relation to managing skin temperature and moisture. Customers were instructed to continue to use the product as they normally would. The mattress system would provide an alarm in the event the MCM blower does not operate properly. A Hill-Rom representative would be contacting their facility to make arrangements for correcting the affected MCM blowers. If customers had any questions concerning this request or the procedure as outlined, they were instructed to contact Hill-Rom Technical Support at 800-445-3720.For questions regarding this recall call 812-934-7777. |  | 740 | 740 units | Worldwide Distribution - USA including AL, CA, CO, CT, DC, DE, FL, HI, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NH, NJ, NY, OH, OK, PA, SC, TN, TX, VA, WA, and WI, and the countries of ARUBA, AUSTRIA, CANADA, FRANCE, GERMANY, LEBANON, PORTUGAL and SAUDI ARABIA | Terminated | August 06, 2012 | 532 | 2465
Z-1605-2011 | 57611 | No | IXL | Spot-film device. | Radiology | Radiology | 510(K) Exempt | Philips BV Pulsera, Rel. 2.3, 718095. |  |  | December 14, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | The fixing strap of the belt holding the CBX Laser to the Image Intensifier may be detached during a procedure when used with the BV Libra and the BV Pulsera/Endura Systems. | Device Design | On 01/06/11, Philips sent an URGENT Device Correction Notice, dated 01/05/2011, to their consignees. The consignees were informed that the fixing strap of the Laser Aiming Device Image Intensifier may inadvertently detach when used with BV Libra, BV Endura Rel. 2, and BV Pulsera Rel 2.3 Mobile X-Ray systems. Specifically the draw latch of the fixing strap holding the CBX Laser to the Image Intensified can be accidentally released during a procedure when used with the above Mobile X-Ray systems. While waiting for Philips to implement their corrective action, the consignees are advised to fix the closed latch body of the draw latch to the metal band using a Tie Wrap. Without securing the latch with a Tie Wrap, the device should not be used.The consignees can contact Philips Call Center at 1-800-722-9377, select 5, and reference FCO 71800035 with any questions about this recall. |  | 214 | 214 units | Nationwide Distribution -- AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and Washington, DC. | Terminated | February 28, 2012 | 441 | 227
Z-1606-2011 | 56427 | No | OAW |  |  | Immunology | HDE - Humanitarian Device Exemption | MESOMARK Humanitarian Device. Authorized by Federal Law for use as an aid in monitoring of patients diagnosed with mesothelioma. Each kit contains 96 tests. |  |  | July 29, 2010 | 2010 | Fujirebio Diagnostics, Inc. | presence of material attached to the bottom of some wells in places visible to the naked eye, which caused an overestimation of the sample. | Other | Notification dated 7/7/10 was sent to the customer via e-mail and FedEx Priority Overnight.FUJIREBIO Diagnostics, Inc. sent an URGENT PRODUCT CORRECTION NOTIIFICATION letter dated July 7, 2010, to all affected customers via e-mail and FedEx Priority Overnight. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to immediately discontinue using MESOMARK kit. Lot 57K04510 and discard any remaining inventory.Complete and return the attached Confirmation of Receipt form upon receipt of this notification, no later than July 30, 2010.Notify your ordering physician or health care provider as appropriate.For questions regarding this recall call 877-861-7246 or 610-240-3800. |  | 42 | 42 kits | Nationwide Distribution including UT | Terminated | July 05, 2012 | 707 | 42
Z-1596-2011 | 57892 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Power cord use on Symbiq One-Channel infuser and |  |  | October 27, 2010 | 2010 | Hospira Inc | Complaints of broken, bent or missing prongs, charring, sparks, visible smoke, burnt smell have been reported on Symbiq AC Power cords. | Component design/selection | Hospira Recall notification letters were sent by Federal Express on November 1, 2010. |  | 32376 | 32,376 power cords | Nationwide distribution and to Australia, Canada, and Malaysia. | Terminated | November 07, 2013 | 1107 | 32376
Z-1598-2011 | 57697 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Plum XL Micro/Macro Infusion Pump, Single Channel, with a 0.1-999 mL/hr flow rate range; list number 11846 |  |  | February 14, 2011 | 2011 | Hospira Inc. | Hospira has received reports of the Plum XL, XLM and XLD Infusion Pumps with no audible alarm at the low audio level setting. | Labeling design | Hospira sent Urgent Device Recall letters dated February 14, 2011 to their customers on the same date, informing them that Hospira had received complaints of the Plum XL, XLM and XLD infusion pumps in which the audible alarm failed. If the audible alarm fails and the user does not notice the visual alert, the user may not be aware of the change in pump status such as air-in-line or occlusion. This may result in a delay or interruption of therapy which may result in serious injury and/or death. Clinicians were instructed to weigh the risk associated with continued use of the device versus removal from service. If they elect to continue the devices, they were provided with instructions to perform alarm tests prior to each clinical use (during their pump cleaning process). If the alarm is not audible, they were instructed to discontinue use of the pump and contact Hospira Global Product Safety and Complaints at 1-800-441-4100. They were also instructed to ensure that all personnel in their facility are following Hospira's Technical Service Manual for routine decontamination of the pumps and cleaning of the buzzers during preventative maintenance, including inspection of the buzzers for leg lifting. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance.The accounts were also requested to complete the attached reply form and return it to Hospira via fax at 1-866-382-4228. |  | 23385 | 23,385 pumps | Worldwide distribution, including United States, the U.S. Virgin Islands, Puerto Rico, Argentina, Australia, Bahamas, Bahrain, Barbados, Belgium, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Ecuador, EI Salvador, Finland, France, Greece, Guatemala, Hong Kong, Hungary, Ireland, Israel, Italy, Jordan, Korea, Kuwait, Luxembourg, Mexico, New Zealand, Oman, Panama, Pakistan, Peru, Philippines, Poland, Portugal, Qatar, San Marino, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Taiwan, Turkey, United Kingdom, Uruguay, United Arab Emirates and Venezuela. | Open, Classified |  |  | 205689
Z-1607-2011 | 57759 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Latex Free NeoGuard Transducer Cover, Ref 676-117, Sterile (2.6 x 30cm) cover, packaged 24/box. The firm name on the label is CIVCO, Kalona, IA. |  |  | January 21, 2011 | 2011 | Civco Medical Instruments Inc | Non-sterile transducer covers were incorrectly labeled as sterile and contained incorrect temperature storage information. | Labeling design | The recalling firm notified their customers via phone and e-mail letter entitled Urgent - Medical Device Correction, dated 1/21/11, explaining the reason for recall. The letter stated that the firm will replace the affected product and requested that customers dispose of the product from the affected lots. A response form was enclosed to report final disposition of the product and should be returned to the firm. If there are any questions, customers can call their customer service representative at 800-445-6741 between 8 am and 5 pm CST. |  | 31 | 31/24-cover boxes | Nationwide Distribution -- WA, CA, and MO. | Terminated | July 09, 2013 | 900 | 31
Z-1471-2011 | 57767 | No | MCY | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Dressing Change Tray, DT5880A, Sterile, Centurion Medical Products Corp., Williamston, MI. |  |  | January 14, 2011 | 2011 | Centurion Medical Products Corporation | The firm is conducting a subrecall of kits containing Triad Groups alcohol prep pads. Triad initiated a recall of prep pads due to a concern about potential contamination of the products with an objectional orgainism, Bacillus cereus. | Nonconforming Material/Component | The firm, Centurion Medical Products, sent an "URGENT PRODUCT RECALL NOTICE" dated January 12, 2011 to affected customers via Certified Mail, Return Receipt Requested. The notice described the product, problem and action to be taken. The customers were instructed to identify and remove all inventory of the recalled kits in their possession; forward a copy of this notice to any customers that this product was further distributed to, and complete and return the attached Accountability Record form via fax to 517-546-3356 or email to: lcarpenter@centurionmp.com. The notice states that a Centurion representative will contact them regarding retrieval of the product from the facility.Should your require additional information, please contact the Director of Quality Assurance &amp; Regulatory Affairs at (517) 546-5400 ext. 1135. |  | 68056 | 68056 all kits | Nationwide distribution: USA including state of: CT, ID, NM, MA, MO, MT, VA and SC. | Terminated | June 21, 2011 | 158 | 68056
Z-1595-2011 | 57819 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | AMPLATZER Sizing Balloon II, REF 9-SB-034, Sterile EO, AGA Medical Corporation, 5050 Nathan Lane North, Plymouth, MN 55442. |  |  | December 17, 2010 | 2010 | AGA Medical Corporation | AGA Medical is initiating a voluntary recall of a specific lot of mislabeled AMPLATZER Sizing Balloon II. AGA Medical received 2 reports from customers who received sizing balloons that were labeled as 34 mm; however, the package contained 24 mm devices. The proper sizing (24 mm) was printed on the device hub visible through the product package. | Employee error | AGA Medical Corporation sent an Urgent Medical Device Voluntary Recall letter, dated December 17, 2010, via e-mail to all affected customers. A hard copy was sent via Fed Ex to customers on December 20, 2010. The letter identified the product, described the issue, and the action that should be taken by the customer. Customers were requested to cease distribution and use of the devices, check inventory for unused devices, and complete and return a response form to facilitate return of the affected devices to AGA.For questions regarding this recall call 1-763-531-3175. |  | 54 | 54 | Worldwide Distribution - USA including AZ, CA, KY, PA, SC, TX and OUS to include: Argentina, Germany, Greece, Netherlands, Peru, Poland, Portugal, Slovakia, United Kingdom | Terminated | March 29, 2012 | 468 | 54
Z-1609-2011 | 57846 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Defibrillator Monitor (M3535A/M3536A) with software (SW) versions F.01.00 or F.01.01 and with all of the following hardware options EtCO2, IP and Temp. |  |  | February 07, 2011 | 2011 | Philips Healthcare Inc. | Software: EtCO2 and Ventilation Rate values may be labeled incorrectly. | Software design | Philips Healthcare representatives issued an Urgent Medical Device letter on 2/7/11. Customers are asked to follow the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notification. In this section, customers are given the following instructions: Philips software (SW) version F.01.03 (for M3535A and M3536A) and R.01.00 (for M3536A, English only) corrects this problem.If you have affected product a Philips Service Engineer will contact you to install software version F.01.03 or R.01.00. Until your software is updated your HeartStart MRx monitor/defibrillator can remain in service. If you are using the device in one of the use modes listed under problem description, make sure that operators are aware that even though the EtCO2 and ventilation rate values can display interchangeably (unexpectedly or randomly switch) under either the ventilation rate label or the EtCO2 label, the values themselves can be identified by their color (for example if the ventilation rate label is white, the ventilation value is displayed as white). Customers are to contact their local Philips representative at 1-800-722-9377, if they need further information or support regarding this issue. |  | 2382 | 2382 units | Worldwide -- USA, Canada, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA,FRANCE, GERMANY, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, KENYA, LATVIA, MALAYSIA, MEXICO, NEW ZEALAND, NORWAY, PHILIPPINES, REUNION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | March 16, 2017 | 2229 | 2382
Z-1610-2011 | 57936 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Natura Durahesive ConvaTec Moldable Technology Post-Operative/Surgical System (57mm), 2 1/4 inch |  |  | February 01, 2011 | 2011 | Convatec Inc. | Natura Durahesive ConvaTec Moldable Technology Post Operative/Surgical System (57mm) units may contain pouches with 45 mm Natura flanges, therefore the pouch, low pressure adapter and barrier would not fit together. | Employee error | Convatec sent a Medical Device Recall letter dated February 1, 2011, via UPS with return response forms attached to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop distributing and quarantine the recalled lot (0100814).Perform a count of recalled product currently in inventory. Complete the enclosed Recall Response Form and return it to the address on the response form. Return the attached Recall Response Form even if no recalled product is in inventory. Ensure that their account number is correctly identified on the attached Recall Response Form. To return the recalled product call the toll free number 1-800-582-6514 to receive an RGA number and instructions on how and where to return the recalled product. If the product has been distributed to wholesalers then customers should forward the letter to them and ask that they follow the Required Actions and return the Recall Response Form to the address listed on the form. Product must be labeled for shipment to: UPS Supply Chain Solutions1920 Outer LoopDock Door #123Louisville, KY 40219For questions regarding this recall call 908-904-2149. |  | 77 | 77 market units | Nationwide Distribution including AZ, IL, KS, MA, MI, NC, OK, OR, PA, SC, and TX. | Terminated | January 11, 2012 | 344 | 77
Z-1611-2011 | 58045 | [6 empty] | MMP 200 Articulating Booms, Catalog Number 0682-000-604SX, Product Number 0682-001-060C. Used for operating room equipment. |  |  | January 28, 2011 | 2011 | Stryker Communications Corp | Utilization of incorrect screws resulted in inadequate thread engagement reducing the load bearing capability of the product. | Process control | Stryker notified consignees verbally on January 28, 2011 and by letter dated February 11, 2011. |  | 2 | 2 | MA, CO | Terminated | March 30, 2011 | 61 | 2
Z-1608-2011 | 57211 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | ORTHOFIX Firebird Spinal Fixation System, Pedicle Screw Spinal System, Reduction Case, Part #68-0111, Lot #16518-LC18. |  |  | October 22, 2010 | 2010 | Orthofix Inc | There is a possibility that certain Anti-Splay Caps may not mate properly during assembly of the construct to the Screw Body with Multi-Axial Reduction Screw Assemblies. | Other | The firm, Orthofix, sent an "URGENT: PRODUCT RECALL" letter dated October 22, 2010 to all customers. The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to: 1) identify and remove the affected product 2) complete the enclosed Tracking and Verification Form via fax to: Orthofix Inc. at 214-937-2764, and 3) contact their Account Services Representative at 1-888-298-5700 and obtain a return authorization number to return product to Orthofix, Inc.The letter also states that Orthofix Inc., will provide replacement Anti-Splay Caps via priority delivery.If you have any questions regarding the removal and return of this product to Orthofix Inc., please contact your Account Service Representative at 1-888-298-5700. Should you have any other questions or concerns, please contact the Director of Regulatory Affairs at 214-937-2061 or by email: DarlaChew@Orthofix.com. |  | 166 | 166 Units | Nationwide distribution: USA including states of: CA, CO, FL, LA , MD, MI, MO, NJ, NY, OH, TN, VA, and WA. | Terminated | April 25, 2011 | 185 | 166
Z-1619-2011 | 56746 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GLU Reagent (GLU cartridge assay) used on Synchron LX and UniCel DxC Synchron Clinical Systems, Part Number 442640; |  |  | August 10, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Cartridge Glucose (GLU) on UniCel DxC and Synchron LX instruments may give falsely high results for moderately hemolysed samples.The Impact:(1) Neonatal samples drawn by lancets are more likely to be hemolysed to a degree sufficient tocause interference with glucose estimation and are a special at-risk population,(2) Other hemolysed samples may be similar | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 10, 2010 with an enclosed fax-back Customer Response Form to the customers. The letter provided the customers with an explanation of the problem identified, the impact and an action to be taken. The customers were instructed to run samples using another method/technology and to complete and return the enclosed FAX BACK RESPONSE FORM via fax to (714) 961-4232 or mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs, 250 S. Kraemer Blvd., PO Box 8000, Brea, CA 92822.Firms with questions regarding this Product Corrective Action letter, were instructed to call our Customer Support Center (Hotline) at (800) 854-3633 in the United States and Canada http://www.beckmancoulter.com/customersupport/support/default.asp. |  | 81871 | 81,871 units | Worldwide distribution: USA and country: Canada. | Terminated | October 17, 2013 | 1164 | 81871
Z-1613-2011 | 57831 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | STA-Flex SUBTALAR SCREW, 8 mm x 14 mm, REF 982362, Sterile, Biomet Sports Medicine, 56 East Bell Drive, Warsaw, IN. |  |  | January 21, 2011 | 2011 | Biomet, Inc. | The firm has initiated this recall following internal fatigue testing that shows possible initiation of fatigue cracking in implants under load. | Pending | The firm sent URGENT MEDICAL DEVICE RECALL NOTICES dated 1/21/2011 to distributors, implanting physicians, and medical facilities. The letter to the distributors and medical facilities states the recalled product should be immediately located and returned to Biomet. In addition, the "FAX Back Response Form" should be completed and returned to the number provided. If product has been further distributed, distributors are to contact hospital personnel via the enclosed "Dear Biomet Customer" notice. The letter to the implanting physicians informs the physician of the recall and lists the patients that were implanted with the device. The physician letter quotes from the surgical technique "the implant should also be removed if any postoperative complications occur". The letter to the implanting physicians does not recommend any action be taken with the devices that are implanted. The notice just makes the surgeons aware of the recall. Distributors and Physicians are to confirm receipt of the letter by contacting 800-348-9500, ext 3755 or 3756. Also, questions related to this matter should be directed to 574-371-3755 or 574-371-3756. |  | 254 | 254 all products | Nationwide Distribution. | Terminated | November 10, 2011 | 293 | 254
Z-1621-2011 | 56637 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT G3+ Cartridge: PCO2 |  |  | September 25, 2006 | 2006 | Abbott Point Of Care Inc. | Certain lots of G3+ cartridges may be difficult to fill. | Other | The firm, Abbott Point of Care, sent a "Important notice" letter dated September 25, 2006 to all customers. The letter described the product, problem and actions to be taken. The customers were to destroy and discard the i-STAT G3+ Cartridges and complete and return the customer reply form. Abbott sent replacements cartridges to their customers.If you or any of the healthcare providers you serve have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020, Option 1. |  | 12800 | 12,675 in US, 125 ROW | Worldwide distribution: USA and country of : Canada. | Terminated | February 27, 2012 | 1981 | 12800
Z-1492-2011 | 58015 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ACCU-CHEK FlexLink Plus, subcutaneous infusion set. Part Number 05511089001. |  |  | February 24, 2011 | 2011 | Roche Insulin Delivery Systems Inc. | There is the potential to kink or bend the cannula when inserting the ACCU-CHEK FlexLink Plus infusion set. This can lead to under delivery and elevation of blood glucose levels. | Device Design | Roche Insulin Delivery Systems sent urgent Product Recall letters, dated 2/21/2011. The letter identified the affected product and described the issue regarding kinked or bent cannulas. The firm requests that customers discontinue use of the ACCU-CHEK FlexLink Plus infusion set and return the unused stock. Customers are to contact their health care provider to determine if changes to their pump therapy are needed and how to temporarily continue insulin pump therapy without the ACCU-CHEK FlexLink Plus infusion set. The attached Customer Reply Card should be completed and return via mail or fax per the instructions on the card. The letter also lists some additional information pertaining to signs and symptoms. Questions regarding this recall should be directed to ACCU-CHEK Customer Care at 1-800-688-4578, 24 hours a day, 7 days a week. |  | 1454918000000 | 1,454,918 million worldwide | Worldwide Distribution -- US and Canada. | Terminated | July 27, 2011 | 153 | 1454918000000
Z-1438-2011 | 55432 | No | LQF | Mycobacterium tuberculosis immunofluorescent reagents. | Microbiology | Microbiology | 510(k) | Mycobacterium Tuberculosis Complex Culture Identification Test, 2860. |  |  | December 30, 2008 | 2008 | Gen-Probe Inc | This recall has been initiated following the identification of certain assay tubes that do not contain probe reagent. The affected kits may contain a combination of tubes with correct probe amount, partial probe amount and some probe tubes that are empty. Possible false negative results. | Process control | Gen-Probe issued a Recall Notice to customers recommending that they discontinue the use of the identified batch (lot) numbers listed. For customers that have used these batches and obtained negative results, they recommend that the referring physician be notified about the possibility of a false negative result.With the Recall Notice, consignees received a Customer Response Form to complete. Gen-Probe requested that they segregate the affected batches listed on the Customer Response Form, then reconcile, and destroy all partial and/or full kits left in inventory, being sure to document the amount destroyed on the Customer Response Form. Gen-Probe should be updated via phone, fax, and email information. Customers should then sign and return the completed form to Gen-Probe via fax to 858-410-8250, or scanned and emailed to technicalsupport@genprobe.com.The Technical Support team should be contacted for replacement kits. Technical Support was also available for additional information at 888-484-4747 or technicalsupport@gen-probe.com. |  | 1293 | 1293 kits | Worldwide Distribution -- USA, including Puerto Rico and countries of Australia, Guatemala, Malaysia, Bahamas, Hong Kong, Mali, Singapore, Bermuda, India, Mexico, South Africa, Canada, Israel, Netherlands, South Korea, Finland, Kuwait, Panama, Thailand, France, Latvia, and Philippines. | Terminated | June 08, 2011 | 890 | 2659
Z-1625-2011 | 54384 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Medtronic brand KyphX HV-R Bone Cement with Kyphon mixer, a polymethyl bone cement kit consisting of a quantity of powder (20 g) and a volume of liquid (10 ml/9 g) mixed in a Kyphon mixer for up to two minutes. Model Number: CO1B; Manufactured by Elmdown LTD, London, England, for Medtronic Spine LLC, Sunnyvale, CA. |  |  | November 10, 2009 | 2009 | Medtronic Spine LLC, formerly Kyphon Inc | Product may not reach proper viscosity (dough state) and inconsistency or lack of homogeneity during cement preparation and during delivery period. | Other | The firm, Medtronic Spine LLC, sent an "URGENT DEVICE RECALL NOTICE" dated November 9, 2009 to the customers. The notice described the product, problem and actions to be taken. The customers were instructed to return the affected lots to Medtronic Spine; to not use any of the affected product that they may have received (Medtronic will replace any affected product in their inventory), the customer should contact their Medtronic Spine sales representative or customer service at [866-959-7466 or for geography] for assistance with replacement product, and complete and return the URGENT DEVICE RECALL NOTICE: Action Required form to Medtronic Spine LLC via fax at (408) 543-6192 or mail to: Medtronic Spine LLC, Attn: Recall Coordinator, 1221 Crossman Avenue, Sunnyvale, CA 94089. If you have any questions about this notification, please do not hesitate to contact Sr. Director, Regulatory Affairs via email: pamela.n.segale@medtronic.com or phone at (408) 548-5235. |  |  |  | USA including states of: AL, AR, AZ, CA, CO, DC, FL, GA, IA, IN, KY, MA, MD, MO, MT, NC, NM, NV, NY, OH, PA, SC, TN, TX, UT, VA, WI, and WV; and countries including: Argentina, Austria, Belgium, Brazil, Canada, Chile, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Ivory Coast, Jordan, Kuwait, Lebanon, Luxembourg, Lybia, Malaysia, Malta, Mexico, Morocco, Netherlands, New Caledonia, Norway, Panama, Philippines, Poland, Portugal, Puerto Rico, Reunion, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Terminated | March 11, 2011 | 486
Z-1627-2011 | 57506 | No | DHX | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | CEA ELISA Kit. |  |  | December 22, 2010 | 2010 | BioCheck Inc | Product labels indicate that the products are for IVD use and they do not have an approved 510(k). | PMA | The firm, BioCheck, sent an "URGENT: DEVICE RECALL" letter dated December 22, 2010 to all customers. The letter described the product, problem and actions to be taken by customers. The customer were instructed to examine their inventory for the affected lots, if found, discontinue distributing the lots and promptly return via Fedex to BioCheck at 323 Vintage Park Drive, Foster City, CA 94404, Attn: RETURNED GOODS; immediately contact their accounts, if they further distributed the product, and inform them of the Recall, and complete and return the response form via email at: info@biocheck.com or fax to (650-573-1969) immediately.If you have any questions, please call (650) -573-1968 anytime during office hours (8:30am to 5:30pm Pacific time). |  | 25000 | estimate 25,000 kits, all varieties | Nationwide distribution. | Terminated | March 25, 2011 | 93 | 25000
Z-1623-2011 | 54468 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN 4000 Vent Head Spring Assembly, a component of the CELL-DYN 4000 Automated Hematology Analyzer, List Number 07H95-01, manufactured by Abbott Diagnostics Division, Santa Clara, CA |  |  | December 29, 2009 | 2009 | Abbott Laboratories | A component used in assembly is out of specification, resulting in failure to retract, eject specimens, exposure below the safety zone, piercing of specimen tubes, or error message to display. | Other | The firm, Abbott Laboratories, sent an "Urgent Field Safety Notice Product Recall" letter dated December 28, 2009 to the Customers via Federal Express. The letter described the product, problem and action to be taken. The customers were instructed to destroy any CELL-DYN 4000 Vent Head Spring Assemblies that they have on hand and have not installed; (No action is necessary for assemblies that are currently installed on the CELL-DYN 4000 Analyzer); complete and return the Customer Reply form (even if they do not have the product) via fax to 1-800-777-0051 or email: QAGCO@abbott.comProduce Recall; indicate the number of Vent Head Spring Assemblies destroyed. the number identified will be replaced with Vent Head Spring Assemblies that meet specifications, and to please keep this communication with their CELL-DYN 4000 System Operator's Manual. If you have any questions regarding this communication, U.S. customers should call Customer Support at 1-877-4ABBOTT (1-877-422-2688). Customers outside the U.S., please contact your local hematology customer support representative. |  | 157 | 157 assemblies | Worldwide distribution: USA including states of: CA, CO, FL, GA, KY, NC, NJ, NY, OR, VA, and WV; and countries including: Argentina, Australia, Canada, Germany and Japan. | Terminated | March 11, 2011 | 437 | 157
Z-1647-2011 | 56044 | No | LRM | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio Rad brand Autoimmune EIA Anti-dsDNA Test Kit, 96 Tests, |  |  | March 02, 2010 | 2010 | Bio-Rad Laboratories Inc | The firm received customer complaints of increased positivity rates associated with the Anti-dsDNA test kits. | Nonconforming Material/Component | The firm, Bio-Rad, sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated March 2, 2010 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory and quarantine the affected kit lots; discard all affected product in accordance with all local, state and federal disposal procedures; complete and return the enclosed MEDICAL DEVICE RECALL RESPONSE FORM via fax to: (510) 741-3954, attn: Bio-Rad CSD Regulatory Affairs Department, and to please re-evaluate all positive patient samples and retest if necessary. If you have any questions, contact Bio-Rad CSD Regulatory Affairs Department at 510-724-7000. |  |  |  | Worldwide distribution: USA and countries including: Canada, Carribean, Chile, France, Germany, Hong Kong, India, Italy, Japan, New Zealand, Panama, Singapore, South Korea and United Kingdom. | Terminated | March 15, 2011 | 378
Z-1651-2011 | 57534 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MagNA Pure LC 2.0 Instrument, catalog number 05197686001, Roche Diagnostics, Indianapolis, IN |  |  | October 26, 2010 | 2010 | Roche Diagnostics Operations, Inc. | If the bar code is printed from the Batch Results Screen, a wrong batch ID may be printed, while the correct batch ID is still assigned to the batch in the data repository of the MPLC 2.0 software. This may result in sample mis identification. If printed from the ordering screens, the bar code is printed with the correct batch ID. | Software design | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 04, 2010 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to print batch ID barcode labels from the ordering or stage set up screens only; print Cooling Block barcode labels from the Post Elution Edit screen only; complete and return the attached UMDC Faxback Form for MagNA Pure LC 2.0 Instrument [UMDC 10-228] via fax to 1-877-551-7152 and file this correction notice for future reference. The firm plans to issue a software update in first quarter 2011. If you have questions about the information contained in this UMDC, for Clinical customers, please contact Roche Molecular Diagnostics Technical Service at 1-800-526-1247; for Research customers, please contact Roche Applied Sciences Technical Service at 1-800-262-4911. |  | 119 | 119 devices | Worldwide distribution: AL, AZ, CA, CO, DC, FL, GA, HI, IL, KY, LA, MA, MD, ME, MN, MO, MT, NC, ND NE, NH, NJ, NM, NY, OH, OK, OR, SC, SD, TN, TX, UT, VA, VT, WI and Puerto Rico; and country of Europe. | Terminated | August 30, 2011 | 308 | 119
Z-1644-2011 | 57724 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; the pumps were sold under the following configurations: |  |  | February 14, 2011 | 2011 | Hospira Inc. | Hospira has received customer reports of the Plum A+ infusion pumps with no audible alarm conditions, which have been associated with failure of the piezoelectric assembly due to component quality issues. | Nonconforming Material/Component | The firm, Hospira, sent "URGENT DEVICE RECALL" letters dated February 14, 2011 to their customers on the same date. The letter described the product, problem and action to be taken by the customers. The Customer/Clinicians were instructed to weigh the risk associated with continued use of the device versus removal from service. If they elect to continue the devices, they were provided with instructions to perform alarm tests prior to each clinical use (during their pump cleaning process). If the alarm is not audible, they were instructed to discontinue use of the pump and contact Hospira Global Product Safety and Complaints at 1-800-441-4100. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance. The accounts were also requested to complete the attached Reply Form and return it to Hospira via fax at 1-888-216-7330. |  | 144477 | 144,477 pumps | Worldwide distribution: United States including the U.S. Virgin Islands and Puerto Rico, and countries including: Argentina, Australia, Bahamas, Barbados, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Denmark, Dominican Republic, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Uruguay, United Arab Emirates and Vietnam. | Open, Classified |  |  | 167595
Z-1649-2011 | 57914 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-ART System, Model # H-0000-0003 |  |  | January 26, 2011 | 2011 | TomoTherapy Incorporated | It was determined that the Treatment Planning Station (TPS) can potentially under dose. | Software design | Consignees were sent a TomoTherapy "Field Safety Notice Medical Device Correction" letter dated January 26, 2011. The letter was addressed to "TomoTherapy Customer". The letter described the Planning Issue, Related DQA Consideration, Graphical Presentation of Results, Product Affected, Recommended Action and Resolution. The letter suggests to use the information presented in order to avoid creating plans with an increased potential for discrepancies in planned vs. delivered dose. Specifically include measurements for small, off-axis targets that could be susceptible to the effects of this issue. DQA dose measurements should be performed at each unique full fraction dose that will be used for patient treatment, without using DQA scaling.If Customers have any questions they are to contact TomoTherapy Interaction Center by email or phone. |  | 326 | 326 | Worldwide Distribution -- US, including States of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, MT, NE, NV, NJ, NM, NY ,NC, ND,OH, OK OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, and WY and countries of SPAIN, GERMANY, FRANCE, INDIA, KOREA, SWEDEN, ITALY, MALAYSIA, SWITZERLAND, CANADA, UNITED KINGDOM, SINGAPORE, TURKEY, TAIWAN, AUSTRALIA, JAPAN, BELGIUM, CHINA, MEXICO, NETHERLANDS, POLAND, SAUDI ARABIA, and UNITED ARAB EMIRATES. | Terminated | March 18, 2014 | 1147 | 326
Z-1658-2011 | 54136 | No | LKV |  |  | Clinical Chemistry | PMA | Rapid fFN for the TLi IQ system and specimen collection kit in the following configurations: 0797Specimen Collection Kit (8 pack); manufactured by Hologic, Sunnyvale, CA. |  |  | December 15, 2009 | 2009 | Cytyc Prenatal Products/Hologic Inc. | Product name change requires conversion to a different type of PMA and implies that the product could be used to assess whether a patient's pregnancy is full term, or to predict imminent delivery in women who are at full term gestation. | PMA | The firm, Hologic, sent a "Important-Product Recall/Correction" letter dated December 18, 2009 to its customers. The letter described the product, problem and the actions to be taken. The customers were instructed to 1) replace the directional inserts for the specimen collection kit and health care providers' brochure with provided instructions, 2) over-label specimen collection kit boxes with provided stickers and individual kits, 3) permanently destroy any marketing and educational materials with the "full term" name and 4) complete and return the CUSTOMER RECALL RESPONSE FORM via fax at 408-744-1905 or scan and email to recallfullterm@hologic.com.If you have any questions about this notification, please contact the Hotline (800-442-9892, Option 7# or 508-263-8510). |  |  |  | Worldwide distribution. | Terminated | March 16, 2011 | 456
Z-1681-2011 | 56332 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SCORPIO CR TIBIAL TRIAL, non sterile, made in the USA |  |  | July 13, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that there is the potential for the Scorpio T72 tibial Insert Trials to be oversized by a maximum of .024" (0.6mm). | Other | Stryker Howmedica Osteonics Corp.sent Urgent Product Correction letters Correction letters dated July 23, 2010, via Federal Express to all affected customers. were sent to all Stryker Branches on July 23, 2010 via Federal Express. Asia Pacific, Canada, Europe, Middle East, Africa, Japan and Latin America sites were sent letters on July 26, 2010 by Federal Express. The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the attached Product Correction Acknowledgement Form by fax as soon as possible to 201-831-6069.For any questions customers were to call 201-972-2100 or 201-831-5028. |  | 159095 | 159,095 units, all types. | Worldwide Distribution - USA including OH, TX, MS, and CO and the countries of Australia, Hong Kong, Taiwan, India, Korea, Japan, Chile, Brazil, Columbia, Venezuela, , Panama, Santo Domingo, Mexico, Argentina, Costa Rica, Guatemala, Peru,Sweden, Germany, Netherlands, France, Spain, Italy, Romania, South Africa, United Kingdom. and Canada. | Open, Classified |  |  | 159095
Z-1665-2011 | 56407 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301A: AQUILION ONE System |  |  | July 22, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) has found that communication may be cut off in the Aquilion ONE/ Premium Systems, owing to the failure of a part used in the console. | Software design | The firm, Toshiba America Medical Systems (TAMS), sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated July 22, 2010, to its customers. The letter described the product, problem and actions to be taken. The letter informed the customers that a new relay will be installed to prevent the occurrence of the problem within the next 90 days and that their Toshiba Service Representative will contact them to install the part to their Aquilion ONE/Premium System. The customers were instructed to contact their Toshiba Service Representative, if the problem occurs before the corrective action has been taken; to share the information with all users and review radiologists as well as clinical engineering or Biomedical group at their facility, and to complete and return the attached CUSTOMER REPLY FORM via fax to 877-349-3054 or email to raffairs@tams.com.If you have any questions regarding this letter, please feel free to contact the Director, Regulator Affairs at (800) 421-1968 or your local Toshiba Representative at (800) 521-1968. |  | 43 | 43 units | Nationwide distribution: USA including states of: AL, AR, AZ, CA, CO, FL, GA, IA, KY, LA, MA, MD, MI, MN, MT, ND, NJ, NV, NY, OH, PA, PR, TX, VA, WI, and WY. | Terminated | April 05, 2011 | 257 | 43
Z-1675-2011 | 56415 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Toshiba DUA-450F Kalare Radiographic Tilting Table. |  |  | July 15, 2010 | 2010 | Toshiba American Medical Systems Inc | Top table lateral movement on their DUA-450F Kalare Table cannot be performed, or the table top lateral movement is performed at an extremely low speed. As a result, positioning of the patient cannot be performed by tabletop movement. | Software design | Toshiba America Medical Systems (TAMS) issued Urgent Medical Device Correction notification dated July 8, 2010, to consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. |  | 238 | 238 systems | US Nationwide | Terminated | April 20, 2012 | 645 | 238
Z-1674-2011 | 57199 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | BLACKSTONE Medical Inc. Firebird Spinal Fixation System 6.5mm x 30mm Bone Screw, Self Tapping, Cannulated. BLACKSTONE Medical Inc. Firebird Spinal Fixation System 6.5mm x 30mm Multi-Axial Screw, Self Tapping, Cannulated. |  |  | October 22, 2010 | 2010 | Orthofix Inc | There is a possibility that certain Multi-Axial Screws have spherical heads measuring less than the allowable tolerance. | Other | Orthofix sent an URGENT: PRODUCT RECALL letter dated October 22, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their accounts immediately and advise them of the recall situation and have them return any affected inventory.For any questions customers were instructed to contact their Account Services Representative at 1-888-298-5700.For questions regarding this recall call 214-937-2061. |  | 27 | 27 (19 of 77-7630) (8 of 77-8630) | Nationwide Distribution including AZ, NY, TX, MS, LA, FL, NC, and AR. | Terminated | May 31, 2011 | 221 | 27
Z-1679-2011 | 57513 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Acetaminophen Test System for use on the Roche/Hitachi, part number 03255379190, Roche Diagnostics Corporation, Indianapolis, IN. |  |  | December 07, 2010 | 2010 | Roche Diagnostics Operations, Inc. | The current product labeling for Acetaminophen does not provide specific interference information for bilirubin, hemoglobin or lipemia. Additional testing has been performed to quantify the impact of these interferences at different concentrations and different levels of acetaminophen. Low serum concentrations of acetaminophen were included in the study. | Other | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 07, 2010 to its customers. The letter described the product, problem, and action to be taken. The customers were instructed to be aware that endogenous interferents (i.e., bilirubin, hemolysis, and/or lipemia) may produce a falsely elevated value for samples containing concentrations less than 50 ug/ml of acetaminophen; complete and return the attached UMDC Faxback Form for Acetaminophen Assay via fax to 1-817-868-5343, and file this letter for future reference. The firm ask the customers to please consult with the physician or pathologist at their facility to determine specific clinical implications for their patients. The firm is also updating the product labeling to include information regarding the interference of icteric, hemolytic or lipemic samples.Please contact Roche Diagnostics Technical Support at 1-800-428-2336 if you have any questions about the information contained in this Urgent Medical Device Correction (UMDC). |  | 8668 | 8,668 units | Nationwide distribution. | Terminated | March 13, 2012 | 462 | 8668
Z-1669-2011 | 57690 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Vertical Brake Hubs of the Computed Tomography X-Ray Systems under the following Brands and Model #s: Brilliance 6, M/Ns 3001-3708, 30001-30131; Brilliance 10, M/Ns 4001-4067, 40001-40035 ; Brilliance 16, M/Ns 5001-6180, 50002-50139 ; Brilliance 16P, M/Ns 6007-6163, 60001-60029 ; Brilliance 40, M/Ns 9009-9272, 29001-29153, 90028-92034; Brilliance 64, M/Ns 9501-10300, 95003-95697; Brilliance Big Bore, M/Ns 7001-7520; Brilliance iCT, M/Ns 100003-100169 &amp; Brilliance iCT SP, M/Ns 200004-200034. |  |  | December 10, 2010 | 2010 | Philips Medical Systems (Cleveland) Inc | Philips Medical Systems has decided to recall the patient support vertical brake hub in all its Brilliance systems due to the collapse in the vertical direction. The short key connecting the vertical brake hub to the motor shaft can disengage due to the lack of thread locker on set screws. | Other | Philips Medical Systems (Cleveland) Inc sent Field Safety Notice to all affected customers on December 10, 2010. The Safety Notices identified the product, the problem, and the action to be taken by the customer. Customers were instructed to monitor their units and if the vertical brake hub malfunctioned to immediately contact the recalling firm. The letter also stated that a Field Service Engineer would visit to perform a brake hub rework on the units. For further information or support customers were to contact their local Philips representative or local Philips Healthcare office for North America and Canada contact the Customer Care Solutions Center 1-800-722-9377, option 5: Enter Site ID or follow the prompt(s).For questions regarding this recall call 440-483-7672. |  | 5700 | 5700 units/vertical brake hubs | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO,CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT,NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY, and the countries of Afghanistan; Algeria; Angola; Argentina; Australia; Austria; Azerbaijan; Bahamas; Bahrain; Bangladesh; Belgium; Bolivia; Bosnia &amp; Herzegovina; Brazil; Brunei Darussalam; Bulgaria; Cambodia; Canada; Chile; China; Columbia; Costa Rica; Cote d'Ivoire; Croatia; Cyprus; Czech Republic; Denmark; Dominican Republic; Ecuador; Egypt; El Salvador; England; Equatorial Guinea; Finland; France; French Polynesia; Germany; Greece; Hong Kong; Hungary; Iceland; India; Indonesia; Iran; Iraq; Ireland; Israel; Italy; Jamaica; Japan; Jordan; Kazakhstan; Kenya; Kuwait; Latvia; Lebanon; Libyan Arab Jamahiriya; Lithuania; Luxembourg; Madagascar; Malaysia; Malta; Martinique; Mexico; Monaco; Mongolia; Morocco; Mozambique; Namibia; Nepal; Netherlands; Netherlands; Antilles; New Zealand; Nicaragua; Nigeria; Norway; Oman; Pakistan; Panama; Paraguay; Peru; Philippines; Poland; Portugal; Romania; Russia; Saudi Arabia; Serbia &amp; Montenegro; Singapore; Slovakia; Slovenia; South Africa; South Korea; Spain; Sweden; Switzerland; Syrian Arab Republic; Taiwan; Tanzania; Thailand; Tunisia; Turkey; Ukraine; Unidentified KM; United Arab Emirates; United Kingdom; Uzbekistan; Venezuela; Viet am &amp; Yemen. | Terminated | August 21, 2013 | 985 | 5700
Z-1667-2011 | 57880 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Innovate Esmark Bandages, a component of various Presource Custom Sterile Packs/Presource PBDS Modules; Packaged by Cardinal Health Medical Products and services, McGaw Park, IL 60085 U.S.A.; There is no label on the Esmark bandage in the Presource pack/module, Some are placed in Tyvek pouches and some are just loose in the pack, but neither have any labeling on them. |  |  | January 28, 2011 | 2011 | Cardinal Health | Esmark bandages placed in some Presource kits/modules may not have been adequately sterilized. | Nonconforming Material/Component | The firm, Cardinal Health, sent an "URGENT: PRODUCT RECALL" letter dated January 28, 2011 to all affected consignees/customers via UPS next day air. The letter described the product, problem and action to be taken. The customers were instructed to quarantine and return the packs/modules to Cardinal Health for destruction; contact the appropriate Customer Service group listed in the letter for return instructions, and complete and return the enclosed Acknowledgment Form via fax to 847-689-9101, regardless of whether or not they have inventory. Should you have any questions, relating to this recall, please feel free to contact Cardinal Health Quality Systems, 800-292-9332. |  | 4.395 | 4.395 modules | Nationwide distribution. | Terminated | September 04, 2012 | 585 | 4.395
Z-1677-2011 | 57926 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | OSTEOMATRIX bone graft substitute. OMG10. |  |  | January 31, 2011 | 2011 | Biostructures, LLC | The OsteoMatrix Synthetic Bone Graft Composite Instructions for Use document, P4G10A/Rev. 01 had language that was not consistent with the cleared indications. | Other | An email was sent on 06/15/2010 to the firm's distributors that had products with lot #0809D110 or lot #0307I208 in their possesion. The email was in regards to the distributors' OMG10 granule inventory. The distributors need to return all the recalled granule inventory to be swapped out for new inventory. |  | 248 | 248 units | Products were distributed to 6 consignees that were located in Michigan, Ohio, Maryland, California, Texas, and Utah. | Terminated | March 17, 2011 | 45 | 248
Z-1654-2011 | 57777 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Allura Xper Systems with Velara Generator, x-ray system. FCO72200184. |  |  | October 07, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | A wire in the connector of the Generator Grid Switch Supervisor box might short-circuit between 12V power and ground, and it will cause a System breakdown. | Nonconforming Material/Component | On 10/15/2010, Philips began sending out the URGENT-Device Correction Notice Failure Causing the Generator to Stop Working notification letter to their consignees.The consignees were informed that a short circuit may occur due to a faulty wire connection in the Supervisor Box Connector in the Generator, and it will cause the generator stops working. When the system fails during a critical interventional case, this may cause risk to patients (no Fluoro available).The systems involved in this recall are Allura Xper with the Velara Generator, Allura Integris with OMCP Generator, MultiDiagnost Eleva, and OmniDiagnost Eleva Imaging system.Consignees can contact Philips Healthcare Call Center at 800-722-9377, #5 and Reference FCO 70800114, 70800115, 72200184, and 72200185. |  | 1556 | 1556 units distributed in US (1452-3107=1556) | Devices were distributed through out the US. | Terminated | February 23, 2012 | 504 | 2739
Z-1676-2011 | 57932 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Remote Cardiology Review Station, a cardiovascular information system. Model Number CV7.8 |  |  | February 10, 2011 | 2011 | AGFA Corp. | A discrepancy in the validation testing which resulted in the product not performing as intended. | Device Design | AGFA sent an "URGENT SAFETY NOTICE" letter dated 2/10/11 via FED-EX to the customer. The letter included an acknowledgment which was to be FAX-Back indicating that the information was received and understood. Agfa has discussed with the customer how the de-installation of their CRS Remote product - Software Version 7.8_HL would be performed. |  | 1 | One | TX | Terminated | September 15, 2011 | 217 | 1
Z-1664-2011 | 57967 | No | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | LUCAS 2 Chest Compression System. |  |  | February 08, 2011 | 2011 | Jolife AB | LUCAS 2 Chest Compression System may have malfunctions resulting from the failure of a hood probe on the device that connects to the battery. The device may stop in fail safe mode and the LEDs will start flickering or the device will switch off. | Other | Medical Device Correction letter, dated February 2011, was sent to customers and states that Jolife AB with the servicing support assistant of Physio-Control is conducting a voluntary upgrade of the LUCAS 2 Chest Compression System with the Serial number range of 30090016 through 30101932.Consignees were advised to keep LUCAS 2 in service and continue to use it in accordance with the Instruction for Use, Section 3.8-Safety Precautions. WARNING-MALFUNCTION. If there are interruptions, or the compressions are not sufficient or something unusual occurs during operation: Push ON/OFF for 1 second to stop LUCAS and remove the device. Start Manual chest compression.Physio-Control service representative will contact the consignees to schedule a service call to replace the affected battery connector board in the hood of the device. The new connector board will include a new probe design in which current cannot affect the spring.Consignees can contact Physio-Control Technical Support at 1-800-442-1142, #5 between 6 am-4 pm (Pacific Time), Monday-Friday or visit the website www.physio-control-notices.com/LUCAS2. If customers no longer have the affected device or the device location has changed they should contact Physio-Control Technical Support. |  | 1843 | 1843 units | Worldwide Distribution -- United States, Australia, Austria, Belgium, Canada, China, Croatia, Czech Republic, Denmark, France, Germany, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, United Arab Emirates, and United Kingdom. | Terminated | April 24, 2012 | 441 | 1843
Z-1673-2011 | 57747 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Ovation Hip Stem, REF 201-0104, Offset Stem Inserter Neutral, Non-Sterile, Ortho Development, Draper, Utah 84020. |  |  | January 11, 2011 | 2011 | Ortho Development Corporation | Component of total hip arthroplasty may break during surgical procedure. | Process design | Ortho Development Corporation notified Field representatives of the issue and conditions that could lead to failure by e-mail on December 14, 2010. Redesigned product is being produced for replacement.For questions regarding this recall call 801-619-3450. |  | 33 | 33 units | Nationwide Distribution including AZ, CA, CO, FL, MI, NV, OR, and UT. | Terminated | November 04, 2011 | 297 | 80
Z-1666-2011 | 58053 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Labeling on box reads: MAQUET BEQ-HG 0284 3/8" X 3/8" Connectors, Lot 70041098 but the product packaged inside was MAQUET BEQ-HG 1113 3/16" x 3/16" Connectors, Lot 70041096. The HG 0284 and HG 1113 connectors are packaged sterile as single components in a container holding 20 units. |  |  | February 23, 2011 | 2011 | Maquet Cardiovascular, LLC | Tubing connectors incorrectly packaged and labeled. | Other | The firm, Maquet, sent a "PRODUCT RECALL URGENT MEDICAL DEVICE CORRECTIVE ACTION" letter dated February 22, 2011 via Federal Express to its customer. The letter described the product, problem and the action to be taken. The customer was instructed to examine their inventory immediately to identify and remove the affected connectors and complete and return the attached RECALL- Device Corrective Action Response Form via fax to Director of Regulatory Affairs at (973) 709-7016 or email: whitney.torning@maquet.com. If you have any additional questions, please contact your local MAQUET Cardiovascular Sales Representative or customer service at 1-800-777-4222. |  | 11360 | 40 BEQ-HG 0284 3/8" x 3/8" Lot 70041098 Connectors | Nationwide distribution: USA state of CA. | Terminated | September 13, 2012 | 568 | 11360
Z-1701-2011 | 51955 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott CELL-DYN 3700 Hematology Analyzer CS (Closed Sampler), , and CELL-DYN 3700 SL (Sample Loader), Manufactured by Abbott Diagnostics Division, Santa Clara, Ca |  |  | March 27, 2009 | 2009 | Abbott Laboratories | FIRE HAZARD-- Some power supply modules may have incorrect fuses installed, which may result in fires and smoking. | Component design/selection | Abbott Laboratories sent Product Correction letters, dated March 22, 2010, to its customers. The letter listed the additional products that are now included in the recall and gave an explanation of the reason for recall. The letter also discussed the safety impact and provided a list of necessary actions for customers. Customers are to check that the ampere rating of the fuse matches the voltage of operation that is currently being utilized with their CELL-DYN 3000 Analyzer. If the fuse is correct, then the steps provided should be followed. If the fuse is incorrect, then customers are to replace the fuse with the correct fuse and follow the instructions provided. Customers are to contact their local Customer Support Representative in the event that they do not have their replacement fuse that was provided with the Accessory Kit. Customers can refer to the Product Information Letter dated April 20, 2010 for additional labeling to assist with the correct fuse replacement. The Customer Reply form should be completed and the letter should be kept with the Operator's Manual. Questions should be directed to Customer Support at 1-877-4ABBOTT.The firm, Abbott Laboratories, sent a "Product Correction" letter dated March 27, 2009 to its customers. The letter included a Customer Reply form and CELL-DYN accessory kits to inform new customers of the necessity of replacing fuses correctly. The letter described the product, problem and actions to be taken. The customers were instructed to check that the ampere rating of the fuse matches the voltage of operation that is currently being utilized with their CELL-DYN 3700 System and follow the instructions provided, and complete and return the Customer Reply form via fax to 1-800-777-0051 or e-mail QAGCO@abbott.com, even if they no longer have the instrument. The firm intends to service items in the field. In April 2010, additional labeling was sent to customers providing additional information when replacing fuses. |  | 2961 | 2961 units | Worldwide Distribution -- USA, including states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY; and countries of Argentina, Australia, Bahamas, Barbados, Brazil, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Curacao, Dominica, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Haiti, Honduras, Hong Kong, Jamaica, Japan, Mexico, New Zealand, Panama, Paraguay, Peru, Puerto Rico, Singapore, South Korea, Taiwan, Thailand Trinidad and Tobago, Turks &amp; Caicos, Uruguay, and Venezuela. | Terminated | December 21, 2011 | 999 | 6881494
Z-1698-2011 | 54665 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | The Waste Line Assembly and the Waste Outlet Tubing and the Waste Bottle Cable are components that are used on one or more CELL-DYN Systems. |  |  | November 30, 2009 | 2009 | Abbott Laboratories | The use of the device assemblies has been extended to a 6 month useful life. | Other | The firm, Abbott Laboratories, sent a "Product Correction" letter dated November 30, 2009 to all consignees/customers. The letter described the product, problem and actions to be taken. The letter stated that Abbott recommended that the customers change their Waste Line Assembly and Waste Outlet Tubing every six months and included a tag in the letter to record installation and replacement dates. Additionally, the customers were instructed to keep this communication with their CELL-DYN System Operator's Manual and complete and return the Customer Reply form via fax to 1-800-777-0051 or email QAGCO@abbott.com even if they no longer have the instrument.If you have any questions regarding this communication, U.S. customers should call Customer Support at 1-877-4ABBOTT (1-877-422-2688). Customers outside the U.S. please contact your local hematology customer support representative. |  | 874 | 874 | Worldwide distribution: AL, AR, AZ, CA. CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO MS, MT , NC, NE, NJ, NM, NY, OH, OK, PA, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries including: Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Dominican Republic, Germany, Guatemala, Hong Kong, Japan, Mexico, New Zealand, Panama, Peru, Puerto Rico, Singapore, South Korea, Thailand, Trinidad and Tobago, and Uruguay. | Terminated | March 18, 2011 | 473 | 874
Z-1668-2011 | 56968 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald Diluent Reagent, 10 mL, List Number: 09H48-02, |  |  | September 30, 2010 | 2010 | Abbott Laboratories | The product bar code label, when scanned, indicates that the volume is 10 mL rather than the correct volume of 10,000 mL. Analytical results are not impacted, although, there is a potential for delay in generating results. | Error in labeling | Product Recall notification letters were sent on 9/30/2010. The letter identified the affected product, an explanation of the problem, patient impact, and necessary actions. Consignees are to contact their local hematology customer support representative to arrange for replacement of the CELL-DYN Emerald Diluent Lot. Once the replacement lot has been received, customers should destroy the affected lot according to their laboratory procedures. A copy of the letter should be retained for their records and the Customer Reply Form should be completed and returned via fax or e-mail. US customers should contact Customer Support at 1-877-4ABBOTT, if they have questions. |  | 322 | 322 units | Worldwide Distribution -- USA, including states of PA, NY, CA, NC, IL, TX, KS, UT, VA, FL, AR, TN, MN, NE, MI, and DE and the country of Germany. | Terminated | April 07, 2011 | 189 | 322
Z-1691-2011 | 57734 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Sabina EE Lift, Liko, Sweeden. |  |  | January 27, 2011 | 2011 | Hill-Rom, Inc. | The firm has received reports of injuries related to the Sabina Sit-in-Stand lifts. Of six injury reports, two were injuries in patints during use and four injuries were in non-patients when they tripped over the lift device and fell on the unit. | Device Design | The firm, Liko, issued two "URGENT FIELD SAFETY NOTICES" one dated January 27, 2011 and a follow-up notice dated February 9, 2011 to its consignees/customers. The notices described the product, problem and actions to be taken. The customers were instructed to keep the device in a separate unoccupied area when not in use; use extra care and support when lifting patients with the device; not allow unauthorized persons, especially children, to play around or operate the lift; examine the sling hooks daily to make sure they are secure, and to forward a copy of this notice to any other facility personnel they deem appropriate. The notices state the firm is investigating a design change to mitigate the risk of injuries with the device.If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. |  | 1009 | 1,009 all products | Worldwide distribution: USA including: AL, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK ,OR, PA, RI, SC, SD, TN ,TX, UT, VA, WA, WI, and WV; and country of: Canada. | Terminated | September 19, 2014 | 1331 | 1009
Z-1699-2011 | 58075 | No | EPP | External nasal splint. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | An external nasal splint is a rigid or partially rigid device intended for the use externally for immobilization of part of the nose. |  |  | February 21, 2011 | 2011 | Medtronic Xomed, Inc. | Medtronic Xomed, Inc. Jacksonville, FL issued a recall on Feb 21, 2011 for all lots of the alcohol pad packaged with the Medtronic's External Nasal Splints REF 1528116 (sm), 1528126 (med), 1528136 (lg) and the Thermasplints REF 1529100 (sm), 1529110 (med), and 1529120 (lg) distributed from Jan 2008 through Aug 2010. The alcohol supplier, Cardinal Health, issued a recall on the alcohol pads on Jan | Material/Component Contamination | Medtronic Xomed, Inc. sent an URGENT PRODUCT SAFETY ADVISORY (PSA) letter dated February 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately discontinue use and dispose of the alcohol prep pad co-packaged in their External Nasal Splint or Thermasplint kits. If there has been further distribution of this item, the Urgent Product Safety Advisory letter should be forwarded to the appropriate recipient. For any questions regarding this recall call 904-279-7532. |  | 26305 | 26,305 units (5,261 boxes) | Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Singapore, Canada, Belgium, Croatia, Cyprus, Denmark, France, Germany, Greece, India, Ireland, Israel, Italy, Latvia, Martinique, Netherlands, Norway, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Switzerland, United Arab Emirates, United Kingdom, Hong Kong, Taiway, Korea, Puerto Rico, and Costa Rica. | Terminated | April 03, 2012 | 407 | 34690
Z-1044-2011 | 57763 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE DS FULL FIELD MAMMOGRAPHY SYSTEM |  |  | July 20, 2010 | 2010 | GE Healthcare, LLC | It was noticed that 3 units received in Japan were missing the rating plate on the X-ray tube cover. There is a potential that other X-ray tube covers are missing the rating plate. | Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter was issued to customers identifying the affected product and describing the defect along with the related hazards. The letter states that there are no actions that are required to be taken on the part of the customer. A GE Healthcare Service Representative will apply the missing rating as necessary and free of charge. Questions or concerns regarding the letter should be directed to 1-800-437-1171 in the US. |  | 112 | 112 total | Worldwide Distribution | Terminated | May 16, 2012 | 666 | 112
Z-1709-2011 | 57970 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | KING LTS-D, SIZE 4, QUANTITY 5 EACH, |  |  | January 14, 2011 | 2011 | King Systems Corp. | King Systems has been notified by Triad Group, our supplier of sterile lubricating jelly, that a recall has been initiated for this product because the product may be "inadequately sterilized" and use of this product may result in "patient infection". Because this particular sterile lubricating jelly is packaged in the KING LT-D Kit and KING LTS-D Kit, King Systems is sending a Notice of the Reca | Process control | The firm, King Systems, sent an "URGENT PRODUCT RECALL" letter dated January 14, 2011 to its customers. King Systems sent a follow-up "URGENT PRODUCT RECALL" letter dated March 3, 2011 which "SUPERSEDES" the January 14, 2011 recall letter with attached instructions to its customers. The letters described the product, problem and actions to be taken by the customers. The customers were instructed to notify all their purchasers of this recall; immediately cease sales and use of the affected products; identify all inventory on hand or returned to them by their customers; REMOVE Triad Group lubricant from King LT-D and LTS-D kits according to King instructions; quarantine the Triad lubricant, re-seal the LT-D and LTS-D kits and destroy Triad lubricant or return to Triad according to King instructions, and complete, sign and return Recall Acknowledgement Form via fax to: 317-776-6827 or email to: triadluberecall@kingsystems.com. King Systems will provide replacement lubricant packs to all customers. Customers with questions regarding product return, replacement, hospital notification or concerning the information requested above, please contact Customer Service at 1-800-642-5464. If you have any regulatory questions about the recall, please contact Director of Quality Assurance and Regulatory Affairs at 1-800-642-5464. |  | 577426 | 577,426 total units | Worldwide distribution: USA including: CA,CO, FL, ID, IN, KS, NC, NH, NY, OH, PA, PR, SC, TN, and TX; and Country of: Canada. | Terminated | February 20, 2013 | 768 | 577426
Z-1690-2011 | 57900 | No | JDK | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DePuy, a Johnson &amp; Johnson company, CEMENT RESTRICTOR, Size 5, Sterile, DePuy Orthopaedics. This device is intended to restrict bone cement migration into the distal medullary canal during implantation of a femoral hip prosthesis |  |  | January 27, 2011 | 2011 | Depuy Orthopaedics, Inc. | The firm became aware of a manufacturing error resulting in a potentially incomplete seal of the outer pouch. Although the sterility of the product is not compromised, the sterility of the inner pouch containing the product cannot be assured. Lack of sterility of the inner pouch could potentially cause contamination of the sterile field, including the implant. Non-sterile product and/or contaminat | Employee error | DePuy Orthopaedics sent an URGENT INFORMATION  RECALL NOTICE dated January 27, 2011, to DePuy Area Directors, FMDs, Distributors, and Office Managers. The notice identified the product, the problem, and the action to be taken. All territories were required to ensure that all recalled products that remain in customer accounts and field inventory are returned immediately to Warsaw to the attention of Returns. Hospitals that purchased and potentially implanted the now recalled product should be provided appropriate notification of the recall and should complete and return a signed Reconciliation Form(s) within seven business days. If the hospitals have any inventory on hand that is subject to the recall, distributors should remove it or help the hospitals return it for credit. All Reconciliation Forms should be faxed to 574-372-7567.For questions regarding this recall call 574-372-7333. |  | 20 | 20 | Nationwide distribution including the states of MA, PA, FL, LA, MN, WA, WI, CO, GA, NY, NC and SC and the countries of Canada, Colombia, Ireland and Argentina. | Terminated | April 23, 2013 | 817 | 140
Z-1704-2011 | 55617 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Stiletto Electrosurgical Probe (E-Probe); Model number: 0270-070-550, 0270-070-551, 0270-070-552, 0270-070-553, 0270-070-554, 0270-070-555; Manufactured by Stryker Endoscopy Puerto Rico. |  |  | February 25, 2010 | 2010 | Stryker Endoscopy | The heat shrink, PTFE E-Probe Tip were out of specification (OOS) in the longitudinal change (+5% -20%). | Device Design | All Sales reps were notified of the recall on February 25, 2010. Urgent Device Removal advisory notices were also sent to all consignees.Device Removal letters identified the specific part numbers and lot numbers involved. Letters also list the potential risk to patients as a result of the issue. Customers are asked to inspect their inventory for part and lot numbers and quarantine any affected product. If recalled product is found, customers are to contact Stryker Customer Service at 800-624-4422 and request an RMS to send back recalled units. New units will be provided at no charge. The acknowledgement form is to be completed and returned. Questions or concerns should be directed to 408-754-2000 or e-mail at eprobe@stryker.com. |  | 31014 | 31, 014 units | Worldwide Distribution -- US, Australia, Canada, Germany, Greece, Iberia, Korea, Switzerland, and Newbury. | Terminated | March 18, 2011 | 386 | 31014
Z-1685-2011 | 56399 | No | CGF | Estrone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Estrone Radioimmunoassay (RIA) Kit, DSL-8700. |  |  | April 22, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter has confirmed that eight Estrone RIA kit (Ref DSL 8700) lots are inaccurately calibrated to the firm's internal reference calibrator. Lots 992637, 100115A, 100205A, 100205RA, and 100226A have now expired and should not be in active use any more. | Other | Beckman Coulter released an Urgent: Product Corrective Action letter dated July 9, 2010 identifying the affected product and informing consignees of the reason for recall, asking them to take the following actions:*For lots 992637, 100115A, 100205A, 100205RA, and 100226A, review historical external quality control results relevant to the patients' results obtained with these lots to ensure the assay performance was within laboratory specifications.*For lots 100319A, 100409A, 100409RA, review patents' results obtained with the results of external quality control and interpret them in the light of additional applicable tests, patient's clinical status and clinical history of patient.Beckman Coulter has finalized the process of re-adjusting Estrone calibrators' values assignment to the new internal reference calibrator starting with lot 100403A.If consignees need products for additional tests they were told to follow replacement instructions below:*In the United States, please contact Customer Service at 1-800-526-3821*Outside the United States, contact your local Beckman Coulter Representative.Consignees were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. They were to also complete and return the enclosed response form within 10 days so the firm can be assured that consignees received the notification.Consignees were instructed to forward any other questions regarding the notification to the DSL Technical Support Center at 1-800-231-7970 in the United States, or contact their local Beckman Coulter Representative. |  | 917 | 917 units | Worldwide Distribution -- US, including states of CA, FL, GA, IL, IN, KS, MI, NC, NY, OH, OR, TX, VA, WA, and WI and country of Canada. | Terminated | May 04, 2012 | 743 | 917
Z-1733-2011 | 55146 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald Hematology Analyzer System; List Numbers: 09H39-01; The product is manufactured in France for Abbott Diagnostics Division of Abbott Laboratories, Abbott Park, IL |  |  | March 23, 2010 | 2010 | Abbott Laboratories | Abbott has determined that gain settings for the non-standard specimen types may not match the gain settings for the standard specimen type. There is the potential for impact to product performance. | Software design (manufacturing process) | The firm, Abbott, sent two "Product Correction" letters one dated March 23,2010 and a follow-up letter dated July 1, 2010, via express mail to all its consignees/customers. The letters described the product, problem and actions to be taken. The customers were instructed to perform the following steps, noted in the letter, to print and verify their gain setting. Abbott informed the customers that they would receive a mandatory software update and were instructed to install the CELL-DYN Emerald according to the attached instructions, and to complete and return the Customer Reply form via fax to 1-800-777-0051 or email to QAGCO@abbott.com.Should you have any questions or would like to discuss in more detail, contact Director Quality Assurance and Regulatory Affairs Abbott Hematology at 408-567-3403. |  | 891 | 891 units | Worldwide distribution: USA including: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, and WV; and countries including: Australia, Brazil, Canada, Columbia, Dominican Republic, Germany, Puerto Rico, Singapore, South Korea, St. Eustalius &amp; St. Martin, and Turks &amp; Caicos. | Terminated | March 18, 2011 | 360 | 891
Z-1727-2011 | 56962 | No | OJZ |  |  | Orthopedic | HDE - Humanitarian Device Exemption | Stryker Biotech OP-1 Putty (OP-1, BMP-7, bone morphogenetic protein) |  |  | October 25, 2010 | 2010 | Stryker Biotech | A change to the Preparation for Use sectionwas identified where 2.5cc was replaced with 2cc-3cc volume and was not approved | Labeling Change Control | The firm, Stryker Biotech, sent a "Medical Device Correction" letter on October 15, 2010 via return receipt letter to all customers. The letter contains labeling with the approved Preparation for Use Section of 2.5 cc volume and included the revised copy of the Package Insert. Additionally, the letter described the product, problem and actions to be taken. The customers were instructed to follow the new instructions provided and to keep the new instructions in mind when using the OP-1 Putty.If you have any questions concerning this communication, please contact 508-416-5326 or Stryker Biotech's Pharmacovigilance department at PVG@stryker.com. |  | 2861 | 2861 units | Nationwide distribution. | Terminated | March 20, 2012 | 512 | 2861
Z-1722-2011 | 55875 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 20 defibrillator/monitor and LIFEPAK 20e. Product part numbers: 3202487-xxx, 3202488-000, U3202487-000, 3202488-015, and U3202488-000. |  |  | May 27, 2010 | 2010 | Physio Control, Inc. | A component failure on the AC power supply assembly results in "NO AC POWER" operation.A failure of AC power can result in a delay of defibrillation therapy if backup battery power is depleted. | Device Design | PhysioControl, Inc. began sending an Urgent Medical Device Correction letter via certified receipt to all affected customers describing the potential issue. Customers were advised to keep the defibrillators in service and to follow the daily Operator's Checklist. All power supplies will be updated.For questions customers were instructed to contact Technical Support at 1-800-442-1142.For questions regarding this recall call 425-867-4000. |  | 53550 | 53550 units | Worldwide distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MI, MD, MI, NC, NY, OH, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY and the coutnries of ANGUILLA, ARGENTINA, AUSTRALIA, BAHAMAS, BAHRAIN, BARBADOS, BOLIVIA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, CYPRUS, DOMINICAN REPUBLIC, EGYPT, EL SALVADOR, GERMANY, GUAM, GUATEMALA, HONG KONG, INDIA, IRAN, ISRAEL, JAMAICA, JORDAN, KUWAIT, LEBANON, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PUERTO RICO, SAUDI ARABIA, SOUTH AFRICA, SYRIA, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIRGIN ISLANDS, and YEMEN. | Terminated | September 29, 2014 | 1586 | 53550
Z-1732-2011 | 57615 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus(R) CS Acute Care Troponin I TestPak, Troponin I Assay |  |  | December 23, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Low frequency of non-repeatable falsely elevated CTNI results without an associated error message. | Other | The firm, SIEMENS, sent an "Urgent Field Safety Notice" letter dated December 2010 to all Stratus(R) CS Instrument customers. The letter described the product, problem and actions to be taken. The customers were instructed to repeat testing of samples with CTNI results above 0.07 ng/mL; discuss content of this letter with their laboratory director regarding the need to review test results reported from CTNI TestPak lots 230200002 or greater; complete and return the FIELD CORRECTION EFFECTIVENESS CHECK form via fax to Technical Solutions Center at 302-631-8467, and forward this notification to anyone to whom they may have distributed this product.If you have any questions, please contact the Siemens Technical Solutions Center at 800-405-6473. |  | 11635 | 11635 | Worldwide distribution: AL, AK, AZ, AR, CA, CO, DC, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY.Product was distributed to foreign accounts in Algeria, Armenia, Australia, Austria, Bahrain, Bangladesh,Belgium, Bosnia &amp; Herzegovina, Canada, Canary Islands, Czech Republic, Cyprus, Croatia, Denmark, Finland, France, Germany, Great Britain, Greece, Guadeloupe, Indonesia, India, Israel, Ireland, Italy, Japan, Libya, Malaysia, Morocco, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Arab Emirates. | Terminated | March 19, 2013 | 817 | 11635
Z-1731-2011 | 57757 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter MAXM Hematology Analyzer and MAXM Hematology Analyzer with Autoloader. |  |  | December 14, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed the Coulter MAXM, MAXM AL, HmX and HmX AL Analyzers omit the tilde (~) character when thetilde is used as part of the Sample 10 within a barcode label scanned by the primary mode barcode reader. There is a potential for Specimen or Patient misidentification to occur. | Other | The recall communication was initiated on 12/14/2010 with Beckman Coulter forwarding an Urgent: Product Corrective Action (PCA) letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the COULTER HmX Hematology Analyzer, COULTER HmX Hematology Analyzer with Autoloader, COULTER MAXM Analyzer, COULTER MAXM Hematology Analyzer with Autoloader. The letter provides the customers with an explanation of the problem identified and an action to be taken.Customers were instructed not to use the tilde (~) character in Specimen or Patient identifiers. In addition, customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with any technical questions regarding this Product Corrective Action were instructed to call Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 4288 | 4288 units total | Worldwide Distribution -- US, Algeria, Andorra, Antigua and Barbuda, Australia, Austria, Bahrain, Bangladesh, Bermuda, Bosnia and Herzegovina, Botswana, Brunei Darussalam, Bulgaria, Burkina Faso, Burundi, Canada, Cayman Islands, China, Congo, Cote d'Ivoire, Croatia, Czech Republic, Djibouti, Egypt, Finland, France, French Guiana, French Polynesia, Germany, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Iraq, Italy, Japan, Jordan, Republic of Korea, Kuwait, Latvia, Lebanon, Libyan Arab, Jamahiriya,Lithuania, Macao, Malawi, Mexico, Mongolia, Morocco, Namibia, Netherlands, New Caledonia, New Zealand, Nigeria, Oman, Pakistan, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia,Serbia, Singapore, Slovakia, Slovenia, South Africa, Sri Lanka, Sudan, Swaziland, Switzerland, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, U.S. Virgin Islands, Yemen, and Zambia. | Terminated | September 02, 2014 | 1358 | 4288
Z-1724-2011 | 57674 | No | JEQ | Powered toothbrush. | Dental | Dental | 510(K) Exempt | W Vital Health Pulsating Clean Rechargeable Toothbrush; Distributed by: Walgreen Co., 200 Wilmot Rd., Deerfield, IL 60015-4616, Made in China; |  |  | January 14, 2011 | 2011 | Brushpoint Innovations, Inc. | The power toothbrush may make a loud noise and the battery cap becomes separated from the main body of the toothbrush. The battery cap can become a projectile that could cause potential injury if it comes in contact with the body. | Device Design | The firm, BrushPoint Innovations, sent a email dated January 7, 2011 notifying their customers of the recall. The email described the product, problem and actions to be taken. The customers were instructed to pull and quarantine the toothbrushes due to potential product quality concerns, and hold their stocks until January 14, 2011, call in the number of toothbrushes quarantined, and then return them to the Walgreen distribution centers. A stop sale was placed on the affected item numbers in their cash register scanners.If you have any questions, call 905-833-5122 or 905-944-1000. |  | 8100 | 8,100 units | Nationwide distribution. | Terminated | December 16, 2012 | 702 | 8100
Z-1041-2011 | 57784 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE ESSENTIAL FULL FIELD MAMMOGRAPHY SYSTEM (SENO ESSENTIAL) |  |  | January 19, 2011 | 2011 | GE Healthcare, LLC | During installation it was discovered that the duplicate of the 21 CFR 1010 Certification label for the X-ray Generator Cabinet was missing from the installation kits that allow installation of the Senographe Essential in a van (mobile unit). | Radiation Control for Health and Safety Act | The firm plans to repair the defect or to bring product into compliance. The field action is being implemented to inspect the affected units and if found to be missing the certification label, a certification label will be applied. This will be carried out via a Field Modification Instruction planned to be released in March, 2011 with an estimated completion date of July, 2011. |  | 14 | 14 | Nationwide Distribution | Terminated | May 16, 2012 | 483 | 14
Z-1726-2011 | 57778 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PhD System EIA/IFA Software; Model/Catalog Number: 426-0246; |  |  | January 20, 2011 | 2011 | Bio-Rad Laboratories, Inc. | Firm's Correction involves advising customers who have implemented a Laboratory Information System (LIS) that they need to ensure quantitative results have passed all validation criteria prior to reporting patient results. When a patient result fails a Validation Rule, the qualitative result will indicate " INVALID". | Other | The firm, Bio-Rad, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 20, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the PhD Worklist Results to verify that all validations have passed and that the quantitative results are valid; to translate (if necessary) the customer letter, add local contacts information, distribute to all active PhD System customers in their region and provide tracking documentation to CSD-RA for all customers who received the customer letter.If you have any questions, contact the Bio-Rad CSD Regulatory Affairs Department at (510) 741-4618. |  | 650 | 650 units | Worldwide distribution: USA and countries including:Canada, China, Australia, New Zealand, Czech Republic, France, Hong Kong, India, Italy, Korea, Singapore, Sweden, Taiwan and United Kingdom. | Terminated | August 17, 2011 | 209 | 650
Z-1737-2011 | 56639 | No | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | diaDexus PLAC Test ELISA Kit, Enzyme Immunoassay for Quantitative Determination of Lp-PLA2 in human plasma or serum, part number 90123, manufactured by diaDexus Inc, South San Francisco. |  |  | August 20, 2010 | 2010 | DiaDexus, Inc | Product fails stability tests prior to expiration date. | Nonconforming Material/Component | A phone script has been prepared to contact customers. Customers were informed of the affected product and its reason for the recall. If customers didn't use the product after May 11, 2010, the action is considered complete. If the product was used after May 11, 2010, a follow up letter will be sent to the customer and they will be asked to re-test any samples beyond May 11, 2010 with an alternate lot of reagent. |  | 482 | 482 units | Worldwide Distribution -- US, including states of UT, OH, CA, NJ, NC, FL, OK, MA, NY, TX, NJ, WA, and MI and countries of Germany and Israel. | Terminated | March 21, 2011 | 213 | 482
Z-1736-2011 | 57443 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Total BhCG (BhCG) Assay Part Number 33500 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter has confirmed that non-reproducible falsely elevated test results may occur when using the Access Total BhCG assay. These elevated results are often attributed to pre-analytical factors and are particularly noted at the low end of the analytical measuring range.The role of preanalytical factors in laboratory testing has been described in a variety of published literature. | Pending | Beckman Coulter sent an Urgent Product Corrective Action letter dated November 11, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:." Evaluate current laboratory protocols to ensure the necessary procedures are in place for identifying and mitigating falsely elevated test results." Follow the sample handling instructions provided in the Access Total ÂżhCG Instruction for Use (IFU)." Review the enclosed pre-analytical sample handling information bulletins:1. The Role of Preanalytical Factors in Chemistry and Immunoassay Testing (P/N A34880A).2. Sample Handling Makes a Difference (P/N A28226A)." At the discretion of the laboratory director, any total ÂżhCG result that is questioned due to lack of clinical correlation should be repeated.NOTE: Due to the clinical utility of the Access Total ÂżhCG assay as a biomarker of pregnancy, Beckman Coulter feels it is unnecessary to require a review of all previously obtained Access Total ÂżhCG test results. The Access Total ÂżhCG assay is intended as a marker for pregnancy at the time of testing. A discrepant or questionable result would have been brought to the attention of the laboratory. The pregnancy status of a patient at this point, post testing, would be known to the patient and the physician." Please share this information with your laboratory staff and your phlebotomy director and/or phlebotomy staff and retain this notification as part of your laboratory Quality System documentation." Complete and return the enclosed response form within ten days so they can be assured that this important notification has been received.If you have any questions regarding this notification, please contact the Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call1-800-854-3633 in the United States and Canada. Outside the United States and Canada |  | 158919 | 158,919 units | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Angola, Antigua and Barbuda, Argentina, Armenia, Australia, Bangladesh, Belgium, Brazil, Bulgaria, Chile, China, Colombia, Congo, Croatia, Czech Republic, Djibouti, Egypt, Ethiopia, Finland, France, French Polynesia, Georgia, Germany, Ghana, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Moldova, Monaco, Morocco, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Zambia, Zimbabwe. | Terminated | September 02, 2014 | 1413 | 158919
Z-1738-2011 | 57809 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | TEG 5000 Thrombelastograph Hemostasis Analyzer System; a multipurpose system for in vitro coagulation studies; Haemonetics, Haemoscope Division, Niles, IL 60714; list number 07-022 |  |  | January 28, 2011 | 2011 | Haemoscope Division of Haemonetics Corp | Haemonetics has received a report of smoke from the unit with a power supply failure. | Other | Haemoscope sent letters to all of their U.S. accounts on January 24, 2011, via certified mail. The letters listed guidelines regarding proper use of the TEG 5000 System Power Supply to eliminate any potential risks such as electrical shorting or power supply damage. The information, which was printed in the letter, will also be available in an updated TEG User Manual soon to be released. A Caution label with installation instructions was included with the letter, to be applied to the back of the TEG unit. The accounts were requested to complete the enclosed acknowledgement form and fax it back to Haemoscope at 847-588-0455. Any questions were directed to 1-800-438-2834. |  | 1155 | 1,155 units | Nationwide Distribution including CA, FL, MD, PA, OH, SC, TN, TX, and VA | Terminated | March 12, 2012 | 409 | 1155
Z-1734-2011 | 57232 | microtome, cryostat | IDP | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Thermo Scientific HM 550 Cold Disinfection Cryostat, Microm International GmbH, Waldorf, Germany |  |  | September 23, 2010 | 2010 | Thermo Fisher Scientific | Through internal quality review, firm determined the need to conduct a more comprehensive system validation to support the claim of complete bacterial, fungal and viral disinfection efficacy. | Device Design | ThermoFisher sent Advisory Notice letters beginning on September 28, 2010, to all affected customers. The letters stated that the firm needed to conduct more comprehensive system validation data on the Cold-Disinfection sytem option available on the HM550-D cryostat to support the claim of complete bacterial, fungal and viral disinfection efficacy. The letter recommended that the Cold Disinfection option on the HM550-D not be used as the sole disinfection method on the HM550-D until the validation work is completed. The firm provided addtional safety protocols pertaining to the cryostat disinfection. Customers were instructed to complete and fax the Confirmation of Receipt Form to Technical Support, COLD-D Advisory Notice at (269) 372-2436. |  | 100 | 100 units | Worldwide Distribution - USA including the states of LA, OH, FL, MD, PA, TX, MO, WI, KS, AL, CA, VA, MS, IL, NE, MA, and the countries of Australia, Belgium, China, Czech Republic, Denmark, Estonia, Finnland, France, Germany, Hong Kong, India, Italy, Japan, Netherlands, Norway, Poland, Portugal, Russia, Spain, Sweden, Taiwan, and the UK. | Terminated | April 02, 2012 | 557 | 100
Z-1739-2011 | 57366 | No | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | Dimension Vista(R) Gentamicin (GENT) Flex(R) reagent cartridge, K4012 |  |  | November 30, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | False elevation of Dimension Vista(R) Gentamicin results for patients and QC when prior reagent aspirated by the reagent probe is any of the following plasma protein methods: ASL, B2MIC, B2MU, CRP,CYSC, HCYS, hcCRP, IGE, IGG3, IGG4, RF, and STFR resulting in inappropriate patient treatment. | Other | Siemens Healthcare Diagnostics, Inc. sent an Urgent Field Safety Notice letter dated November, 2010, to all affected customers. Customers were requested to run GENT on a Dimension Vista(R) system that does not run the listed plasma protein methods. Instructions were provided for disabling methods on the instrument to prevent inadvertently testing Gent along with plasma protein methods.Customers were instructed to :Forward the notification to anyone to whom they may have distributed the product. Complete and return the Field Correction Effectiveness Check Form to Siemens Healthcare Diagnostics at (302) 631-8467.For further information contact Siemens Technical Solutions Center at 1-800-441-9250 or their local Siemens representative. |  |  |  | Nationwide Distribution including AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | March 19, 2013 | 840
Z-1743-2011 | 56397 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | SYNCHRON System(s) Chemistry RF Rheumatoid Factor Part Number 475136 |  |  | June 21, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter testing indicated that Rheumatoid Factor (RF) reagent lot M907325 will not remain stable to the labeled expiration date. | Other | The firm Beckman Coulter, sent a "URGENT: PRODUCT CORRECTIVE ACTION" letter dated June 25, 2010 to the consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to take the following actions:*Discontinue use and discard appropriately Rheumatoid Factor (RF) reagent lot M907325*Contact their Customer Service Representative at (800) 526-3821 in the United States or (800) 463-7828 in Canada to discuss reagent replacement options.*Complete and return the FAX BACK RESPONSE FORM via fax to (714) 961-4234 or mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs, 250 S. Kraemer Blvd., Brea, California 92821, Mail Code E2.SE.08Consignees/customers were asked to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If consignees/customers forwarded the affected products listed above to another laboratory, then they need to also provide a copy of the Product Corrective Action letter to the laboratory.If they have any questions regarding this Product Corrective Action letter, consignees were told to call Beckman Coulter Customer Support Center (Hotline) at (800) 854-633 in the United States and Canada. |  | 58 | 58 units | Worldwide distribution: USA including states of: AZ, CA, CO, IA, PA, NH, NV, NY, and WA; and countries of: Belgium, China, Czech Republic, France, Germany, Hong Kong, India, Italy, Japan, Kuwait, Lebanon, Lithuania, Mexico, Netherlands, Russian Federation, South Africa, Spain, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | July 02, 2012 | 742 | 58
Z-1368-2011 | 57106 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura XPER FD 10 and Field Extension for XPER Cardiovascular System and Allura XPER FD 20 and Allura CV20. |  |  | November 04, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Adjustment of the micro-switches in the footswitch used on the Philips cardiovascular X-ray systems was not implemented. If it is programmed for exposure, it is possible that after pressing the pedal on the footswitch it does not release itself. The patient may be exposed to X-ray longer than intended. | Pending | On 11/5/10, Philips began sending the Electronic Product Radiation Warning notification letter, dated November 4, 2010, to their consignees.The consignees were informed that the middle pedal of the foot switch of the Allura FD10, FD20, CV20 systems, and the Allura Field Extension series does not not always become inactive (sticks) when pedal is released. The function of the middle pedal on the foot switch is freely programmable by the customer. If it is programmed for exposure and if the pedal does not release by itself, the patient will be exposed to X-ray longer than intended.The consignees were advised to check if their footswitch is possibly affected. Footswitch identification 452270000141 is noted. If the pedal does not release by itself press the red emergency stop button on the system.Consignees can call Philips Healthcare Call Center at 800-722-9377, #5 and Reference FCO 72200189. |  | 4 | 4 units | Worldwide Distribution -- US, including states of CA, IL, MA, NC, and OC. | Terminated | July 21, 2011 | 259 | 11
Z-1622-2011 | 57947 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Optiflux F250NRe Hemodialyzer, Single Use, Catalog Number: 0500325E, Advanced Fresenius Polysulfone. |  |  | February 16, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | There is an increased risk of an internal blood leak. | Process control | FMCNA notified consignees on 2/16/11 by telephone and follow-up with an Urgent FMCNA Optiflux F250NRe Recall letter via Certified Mail, return Receipt requested. The firm states in the letter that there is an increased risk of an internal blood leak affecting a small percentage of the two lots of dialyzers. They are asking customers to examine their inventory, to discontinue use, and segregate the affected product in a secure place. Customers are to contact their FMCNA Customer Service Team at 1-800-323-5188 for instructions on how to return product. Used dialyzers should not be returned. If there are any questions, Customer Service should be contacted. |  | 38364 | 38,364 units | Worldwide Distribution -- USA and Canada. | Terminated | February 12, 2013 | 727 | 38364
Z-1643-2011 | 57916 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | AFFIXUS Lag Screw Drill, REF 2112-01-303, Non-Sterile, DePuy Orthopaedics, Inc., Warsaw, IN; DePuy International Leeds, England. |  |  | January 12, 2011 | 2011 | Depuy Orthopaedics, Inc. | The firm recieved a complaint that the cannula for the drill is not the same size at both ends. The drill is grabbing the guidewire and advancing and not allowing to finish reaming, resulting in the possibility of the guide wire progressing through and penetrating the acetabulum. | Other | On 1/12/2011 DePuy contacted their distributors by telephone and directed the distributors to cancel all pending surgeries and discontinue use of the sets until the investigation is completed. On 1/14/2011 DePuy again contacted all distributors by telephone and directed all distributors to return the instruments to DePuy for inspection as part of the investigation. |  | 296 | 296 units | Worldwide Distribution -- US, including the states of AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MS, NC, NV, NY, OH, PA, TN, TX, VA, WA, WI; and country of Switzerland. | Terminated | June 27, 2012 | 532 | 296
Z-1744-2011 | 57937 | No | LOM |  |  | Microbiology | PMA | ETI-AB-AUK PLUS kits,REF P001931, DiaSorin S.p.A., 13040 Saluggia (VC), Italy, Distributed by: DiaSorin Inc., 1951 Northwestern Avenue, P.O. Box 285, Stillwater, MN 55082-0285 |  |  | August 19, 2010 | 2010 | Diasorin Inc. | Through internal testing, DiaSorin has become aware that the kit lots identified (ETI-AB-AUK PLUS kits, catalog number P001931, lot numbers 0650270A and/or 0650270A/1) may not consistently exhibit stability to defined expiration dating, resulting in decreased binding and lower absorbance. Internal testing has confirmed that when defined assay validation criteria pass, analytical performance is co | Other | The firm, DiaSorin, sent a "Customer Notification Letter" dated 19 August, 2010 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to destroy all remaining inventory and complete and return the receipt of notification and destruction of product form, found in the bottom portion of the letter, via fax to: (651) 351-5669, ATTN: TECHNICAL SERVICES or mail to: ATTN: TECHNICAL SERVICES, 1951 Northwestern Ave., Stillwater, Minnesota 55082-0285.If you have any questions or concerns, please contact DiaSorin Inc., Technical Services at 1-800-328-1482 or 651-439-9710. |  | 710 | 710 | Worldwide distribution: USA including states of: AL, AR, CA, CT, DE, DC, FL, GA, IL, IA, KY, LA, MD, MA, MI, MN, MO, NH, NJ, NM, NY, NC, ND, OH, OK OR, PA, TX, VT, VA, WA,and WI; and country of: SOUTH KOREA. | Terminated | August 16, 2012 | 728 | 710
Z-1723-2011 | 57960 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI R4.5 |  |  | October 25, 2010 | 2010 | Elekta, Inc. | If another patient is selected on the imaging system (XVI) release 4.5) during the transmission of end of treatment data from Desktop Pro R7.01 to R&V; system, the XVI information is prioritized and the end of treatment data is never received by the R&V; system. | Other | Important Notice A341 titled "Mandatory Upgrade Desktop Pro R7.01 Service Pack and Cancelled Notices" dated January 26, 2011, is in distribution to all affected customers. Field Change Order 20000501011, "Service Pack 2 upgrade to Desktop Pro R7.01 Service Pack 1", dated October 25, 2010 will be utilized by Elekta Service personnel to upgrade the Desktop Pro to include Service Pack 2. The Important Notice letter instructed customers to file the letter in the Important Notice section of the appropriate User Manual. Before doing so, users are to ensure that Desktop Pro R7.01 SP2 is installed on their treatment control system. If it is not installed, customers are to contact their local Elekta representative. Once the installation of Service Pack 2 has taken place, the notices listed in Table 1 of the letter will no longer be applicable and should be removed from the User Manual. Customers should direct their questions to their local Elekta representative. |  | 13 | 13 units | Class 2 Recall -- Nationwide Distribution -- Including CA, MI, MO, ND, OH, OR, RI, VT, and WI. | Terminated | May 24, 2011 | 211 | 13
Z-1745-2011 | 52904 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Siemens Medical Solutions USA, Inc., ACUSON AcuNav ultrasound catheters when using SwiftLink catheter connectors; |  |  | March 20, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Screen Image Quality. An assembly issue on SwiftLink connectors may cause reduced/degraded image quality when using ACUSON AcuNav ultrasound catheter transducers. | Process design | Siemens Field correction was initiated on 3/20/2009. A device correction letter was distributed to all affected users, with a description of the problem and its corrective actions. Siemens service personnel will make a site visit. |  | 21 | 21 units | Distribution in US and Internationally. | Terminated | April 01, 2011 | 742 | 21
Z-1746-2011 | 57824 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Organogenesis Apligraf Interactive burn and wound dressing, supplied as a living, bi-layered skin substitute |  |  | January 30, 2011 | 2011 | Organogenesis, Inc. | Product sterility compromised | Material/Component Contamination | Organogenesis notified accounts on Sunday, January 30, 2011, via fax and phone. The Company will collect written confirmation from the affected customers that they have been notified of the recall. Sales representatives for the affected customers have been directed to visit the customers as soon as possible Monday morning to assist in the recall and prevent additional applications of the product. Apligraf Medical Affairs personnel will provide clinical support to treating physicians, and update them as new information on microbial contamination and antibiotic susceptibility becomes available. Units not applied have been requested to be returned to Organogenesis |  | 50 | 50 units | nationwide. | Terminated | July 28, 2011 | 179 | 50
Z-1748-2011 | 58137 | No | NEZ | Ingestible telemetric gastrointestinal capsule imaging system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Given Imaging Limited, PillCam Express Delivery Device, For use with PillCam SB, endoscopic imaging device. Part numbers FGS-0367 (3-Pak) or FGS-0366 (1-Pak). |  |  | March 08, 2011 | 2011 | Given Imaging Ltd. | The capsule holder can break away from the catheter during endoscopic procedures, necessitating retrieval. | Device Design | Given Imaging, Inc. consignees were notified via FedEx letter on/about 03/08/2011. They instructed to return all affected product. A Customer Response form with return instructions was included to be completed and returned to Given Imaging. |  | 389 | 389 | U.S. Nationwide, Canada, Germany Israel and Australia. | Terminated | December 02, 2011 | 269 | 389
Z-1747-2011 | 58138 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Organogenesis Apligraf - Interactive burn and wound dressing, supplied as a living, bilayered skin substitute |  |  | February 21, 2011 | 2011 | Organogenesis, Inc. | Product sterility may be compromised | Packaging process control | Organogenesis contacted the accounts by telephone and fax on 2/24/11 to advise of the contamination and identification fo the organism for appropriate patient follow-up. |  | 7 | 7 units | Nationwide. | Terminated | July 28, 2011 | 157 | 7
Z-1796-2011 | 57177 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | ASR 300 Size 46, DEPUY ASR 300 ACETABULAR CUP SYSTEM. |  |  | August 23, 2010 | 2010 | Depuy Orthopaedics, Inc. | As part of our ongoing post-market surveillance of all products, DePuy is continually evaluating data from a variety of sources including national joint replacement registries, published literature, company sponsored clinical trials, internal complaints data and unpublished clinical research reports.DePuy Orthopaedics issued a Field Safety Notice in March 2010 after receiving new data from the | Device Design | DePuy Orthopaedics issued a Field Safety Notice in March 2010 to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to share the Field Safety Notice with their organization and any organization where the ASR products may have been transferred. Customer were instructed not to implant the ASR devices. Patients who received the ASR System should be informed of this recall and instructed to return for a follow up visit.Patients with radiographic changes indicative of product failure should be addressed according to normal procedures.All other patients should be followed according to the April 22, 2010, and May 25, 2010 UK Medicines and Healthcare products Regulatory Agency (MHRA) Device Alerts. Follow up for all patients implanted with ASR acetabular cups at least annually for five years postoperatively. Beyond five years, follow up in accordance with locally agreed protocols.For patients who are symptomatic or implanted with a cup angle greater than 45 degrees, particulary where a small component has been implanted.:" Consider measuring cobalt and chromium ion levels in whle blood and/or performing cross sectional imaging including MRI or ultrasound scans." If metal ion levels in whole blood are elevated above 7 parts per billion (ppb) for either metal ion, a second test should be performed three months after the first in order to identify patients who require closer surveillance, which may include cross sectional imaging" If MRI or ultrasound scan reveals soft tissue reactions, fluid collections or tissue masses, then revision surgery should be considered.DePuy representatives were to assist with returns of any remaining inventory.For questions regarding this recall call 574-372-7333. |  | 270 | 270 | Worldwide | Terminated | August 29, 2013 | 1102 | 120984
Z-1818-2011 | 57946 | No | JNL | Thyroid autoantibody immunological test system. | Immunology | Immunology | 510(k) | Access Thyroglobulin Antibody II, Part Number: A32898 |  |  | January 26, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed customer reports of reagent pack to pack variability for the Access Thyroglobulin Antibody II (TgAb) reagent lot numbers identified in the recall. Erroneous results for patient and quality control samples can be generated. The ThyroglobulinAntibody II (TgAbII) reagent lots may produce:(1) Erroneously high TgAb results for both quality co | Other | Beckman Coulter sent a Product Corrective Action (PCA) letter on January 26, 2011 via US Postal Service to the affected customers. The letter explained the problem identified and the action to be taken.Customers were instructed to:(1) Discontinue use of Access Thyroglobulin Antibody II reagent lot numbers identified in the tableabove;(2) If customers obtained questionable thyroglobulin antibody results, or received feedback from clinicians on reported thyroglobulin antibody results using the reagent lot numbers identified above, rerun or request a redraw of the patient sample;(3) Share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded any of the affected product listed above to another laboratory, they should provide a copy of the letter to them;(4) Complete and return the enclosed response form within 10 days so the firm can be assured that you have received this important notification and are taking appropriate action.For questions regarding the recall communication, customers were told to contact Beckman Coulter Customer SupportCenter:(1) Via our website, http://www.beckmancoulter.com/customersupport/support(2) Via phone, call 1-800-854-3633 in the United States and Canada(3) Outside the United States and Canada please contact your local Beckman Coulter representative. |  | 3658 | 3658 units (1800 in US) | Worldwide Distribution - USA including AR, CA, FL, GA, HI, IA, IL, KY, MA, MI, MN, MO, MS, NC, NJ, NM, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI and the countries of Argentina, Armenia, Australia, Belgium, Brazil, Bulgaria, Canada, China, Czech Republic, France, French Polynesia, Georgia, Germany, Greece, Hungary, Israel, Italy, Kazakhstan, Kuwait, Lebanon, Morocco, Netherlands, New Zealand, Oman, Poland, Puerto Rico, Romania, Russian Federation, Serbia, South Africa, Spain, Switzerland, Taiwan, Tunisia, Turkey, Ukraine, United Arab Emirates, Yemen | Terminated | May 04, 2012 | 464 | 3658
Z-1819-2011 | 57962 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Extended Brilliance Workstation (EBW) a component of the Computed Tomography X-Ray System Model Numbers EBW (CT) 728260 &amp; EBW (Portal) 728269. |  |  | October 03, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | During an FDA inspection on 10/1/2010 the firm was notified by FDA that their field action they initiated in Oct 2009, constituted a recall. The firm determined that the functional analysis feature in the Comprehensive Cardiac Analysis (CCA) was displaying incorrect values on the polar fields. The calculated results for the myocardial wall thickening was incorrect in that the numbers displayed f | Other | Philips sent an URGENT - Field Safety Notice letter dated September 29, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to be aware that the numbers displayed for "Wall Thickening" actually represent a reduction in wall thickness from ES to ED.Philips will be providing a sotware update to resolve the issue. For further information or support customers were instructed to contact their local Philips representative:Customer Care Service Center in the U.S. at 1-800-722-9377, option 5:Diagnostic Imaging option 1: CT or their local Philips representative. |  | 2068 | 2,068 units | Worldwide Distribution - USA including the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX , UT, VA, VT, WA, WI, WV &amp; WY and the following countries: ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BANGLADESH, BELGIUM, BRAZIL, BRUSSEL, CANADA, CHILE, CHINA, COLUMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, EDUADOR, EGYPT, FRANCE, GERMANY, GREECE, GUADELOUPE, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, LATVIA, LEBANON, MALAYSIA, MALTA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PHILIPPINES, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB, MIRATES, UNITED KINGDOM, VENEZUELA &amp; YEMEN. | Terminated | September 21, 2012 | 1084 | 2068
Z-1820-2011 | 57564 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 S/N 9089, 9551, 9548. |  |  | December 20, 2010 | 2010 | Philips Medical Systems (Cleveland) Inc | During the version 2.6 software upgrade, the system was not recalibrated with the infant phantom for the infant protocol. | Other | Philips sent an URGENT - Medical Device Correction letter dated January 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers may consider reviewing any infant images that may have been affected during the Field Test to confirm reading accuracy. For further information or support concerning this issue, customers could contact their local Philips representative or local Philips Healthcare office for North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts) |  | 3 | 3 | Worldwide Distribution - USA including CO, IN, and TN and the countries of Norway, Israel, and Germany | Terminated | February 22, 2013 | 795 | 3
Z-1725-2011 | 58123 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Millar Instruments, Inc. Millar Sensors Systems Solutions Mikro-Tip Angiographic Catheters Sterile and Nonpyrogenic Model SPC-454D and SPC454F. |  |  | March 08, 2011 | 2011 | Millar Instruments, Inc | Possibility of epoxy debris particles inside the catheter lumen. | Process design | The firm, Millar, notified their consignees/customers by email and sent hardcopy of email on March 8, 2011. The notification described the product, problem and actions to be taken. The customers were instructed to return the catheter(s) to Millar or report back to Millar if the catheter(s) has been consumed.If you have any questions, please do not hesitate to contact Director, RA/QA at 832-667-7000. |  | 1080 | 1080 units | MN, NY, Washington DC, NC, MA, MN, PA, CA, MI, Australia, Germany, Canada, Japan, United Kingdom, and Norway. | Terminated | June 27, 2011 | 111 | 1080
Z-1824-2011 | 57455 | No | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcium Reagent, Calcium Test System for cobas c 111, Part number 04718933190, Roche Diagnostics Corporation, Indianapolis, IN. |  |  | October 11, 2010 | 2010 | Roche Diagnostics Operations, Inc. | Internal investigations have confirmed imprecision and erroneous patient results (up to 35% discrepant results) for specific lots of the Calcium assay.Internal investigations revealed the presence of precipitate in the R2 of the affected calcium reagent lots. The precipitate observed was classified as o-cresolphthalein complexone (o-CPC), a reactive component of the calcium R2.The root cau | Other | The firm, Roche, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated October 11, 2010 to all consignees/customers via UPS(signature required) . The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of the recalled product; discard from inventory according to their site's local regulations; complete and return the attached Faxback form via fax to 1-877-598-5705 even if they do not have the affected product, and file this letter for future reference. The firm added lot 63020401 to the recall in 3/2011. The firm is mailing an URGENT MEDICAL DEVICE REMOVAL notice to customers that purchased that lot of product.Please contact Roche Diagnostics Technical Support at 1-800-428-2336 if you have questions about the information contained in this letter. |  | 14528 | 14,528 units | Nationwide distribution. | Terminated | September 13, 2011 | 337 | 27933
Z-1821-2011 | 58017 | No | JCY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ***REF 24k100***QTY 1***24k Arthroscopy Outflow/Suction Tubing Set***MADE IN USA***Rx Only***CONMED LINVATEC, CONMED LINVATEC 11311 CONCEPT BLVD, LARGO, GL 33773-4900*** |  |  | September 10, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | On 09/13/2010 ConMed Linvatec, Largo, FL initiated a recall on 24k100 Arthroscopy Outflow/Suction Tube Set used with the ConMed Linvatec 24k Pump. The 24k100 Arthroscopy Outflow/Suction Tubing Set Lots 22747-000, 22748-000, and 2749-000 may have been incorrectly assembled and could potentially allow waste to enter the sterile field. | Other | Linvatec Corp. dba ConMed Linvatec sent an Urgent Medical Device Recall Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their facility's inventory for the affected product. If they had any of the product they were instructed not to use and segregate the product and return for credit.Return the affected products and the completed REPLY FORM to their local ConMed Linvatec office.If the product was transferred to another facility, customers were instructed to complete the REPLY FORM documenting the transfer of the product and mail or fax the REPLY FORM to their local ConMed Linvatec office at 727-319-5701.If the customers do not have the affected product they should complete the REPLY FORM and mail in the enclosed envelope or fax to 727-319-5701. |  | 4598 | 4,598 | Worldwide Distribution - USA (nationwide) and internationally to Saudi Arabia, Italy, France, Canada, Portugal, South Africa, Switzerland, Denmark, Germany, New Zealand, Australia, and Spain. | Terminated | January 31, 2012 | 508 | 4598
Z-1822-2011 | 58002 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ossur Total Knee, Model Number: 2100 |  |  | February 11, 2011 | 2011 | Ossur Engineering, Inc | The recall was initiated based on a finding by Ossur that some units of the Total Knee 2100 produced from October 19th, 2010 until January 21st 2011, contained back links that may compromise knee function prematurely. | Process design | Ossur initiated a Safety Notice (Voluntary Product Recall) letter on February 11, 2011, via US Postal Service Registered Mail. The recall letter informed the affected customers of the reason for the recall and the product affected.Ossur requested their distributor and providers take the following action: (1) Please identify the users of each of the Total Knee 2100 model devices listed above; (2) Also, please check your inventory and isolate any recalled devices you have on hand but have not fitted on users. Customers were also informed that a representative from Ossur customer service will be contacting them to coordinate replacement of the affected devices.The letters were followed by phone calls placed by a trained Ossur Customer Service Representative to each customer after US distributors provided Ossur with customer contact information.For questions regarding this recall call 517-629-8890. |  | 142 | 142 units | Worldwide Distribution - USA including CA, CO, FL, LA, MD, MI, MO, NJ, NY, OH, TN, VA,, and WA and the countries of Australia, Canada, Czech Republic, Germany, France, United Kingdom, Ireland, Italy, Japan, The Netherlands, Russian Federation, Sweden and Turkey. | Terminated | February 29, 2012 | 383 | 142
Z-1834-2011 | 54706 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Coherence Dosimetrist, version 2.2. |  |  | November 06, 2009 | 2009 | Siemens Medical Solutions USA, Inc | A problem in the software has the potential to be associated with the dose; when a template is created the user has the possibility to predefine organs. For the Target the user defines the max. dose and the min dose, for the OAR (Organs at Risk) the user only defines the max. dose. The firms engineering team determined that the software defect was repeatable. | Software design | Affected consignees will be notified by the Customer Service Engineer to schedule a software update beginning November 2009. |  | 60 | 60 units | Worldwide distribution: USA including: CA, KY, MI, NJ and TN; and countries of: Canada and Mexico. | Terminated | April 06, 2011 | 516 | 60
Z-1832-2011 | 54842 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson S2000 Ultrasound System with software and imaging. |  |  | March 01, 2010 | 2010 | Siemens Medical Solutions USA, Inc. | "Any 2D measurement performed in Exam Review or offline, on a zoomed Retrospective clip capture may be erroneous. This applies to 2D, 2D/PW, 2D/CW and 2D/M-mode." | Software design | Siemens Medical Solutions USA, Inc. sent a Consumer Safety Advisory Notification letter on February 16, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the Customer Safety Notice to all those within their organization who need to be aware of the issue until the corrective action was completed. For questions regarding this recall call 650-694-5580. |  | 1903 | 1903 | Worldwide Distribution - USA (nationwide) and the countries of United Arab Emirates, Australia, Belgium, Bolivia, Brazil, Bahamas, Canada, Switzerland, Chile, China, Colombia, Czech Republic, Germany, Denmark, Ecuador, Eguypt, Spain, Finald, France, Great Britian, Germany, Hong Kong, Croatia, Indonesia, Ireland, Israel, India, Ireland, Italy, Jordan, Japan, Kenya, South Korea, Kuwait, Lithuania, Mexico, Malaysia, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Sweden, Slovenia, Slovak Republic, Thailand, Turkmenistan, Tunisia, Turkey, Taiwan, Ukraine, Venezeula, Yemen, and South Africa.. | Terminated | June 18, 2012 | 840 | 1903
Z-1837-2011 | 56840 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 ultrasound systems (all versions) with cardiovascular options installed; |  |  | August 17, 2010 | 2010 | Siemens Medical Solutions USA, Inc., Mountain View, CA | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. | Software change control | Siemens Medical Solutions USA, Inc., sent a Customer Safety Advisory Notification letter on August 17, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the notice on to all those within their organization who need to aware of the issue until the corrective action was completed. For questions customers were to contact their local service support person. For questions regarding this recall call 650-694-5993. |  |  |  | Worldwide Distribution - USA (nationwide)and the countries of Antares distribution includes: United Arab Emirates-AE, Austria-AT, Australia-AU, Azerbaijan-AZ, Brazil-BR, Belarus-BY, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Egypt-EG, Spain-ES, France-FR, Great Britain-GB, Guatemala-GT, Croatia-HR, Hungary-HU, Indonesia-ID, Ireland-IE, India-IN, Italy-IT, Jordan-JO, Japan-JP, South Korea-KR, Moldova-MD, Mexico-MX, Malaysia-MY, Netherlands NL, Norway-NO, New Zealand-NZ, Poland-PL, Portugal-PT, Romania-RO, Serbia-RS, Russian Federation-RU, Saudi Arabia-SA, Sweden-SE, Singapore-SG, Slovak Republic-SK, Suriname-SR, Turkey-TR, Taiwan-TW, Ukraine-UA, United States-US, Uruguay-UY, Uzbekistan-UZ, Venezuela-VE, Vietnam-VN, South Africa-ZA S2000 distribution includes: Austria-AT, Australia-AU, Brazil-BR, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Spain-ES, Finland-FI, France-FR, Great Britain-GB, Hong Kong-HK, Indonesia-ID, India-IN, Italy-IT, Japan-JP, South Korea-KR, Kazakhstan-KZ, Lebanon-LB, Mexico-MX, Norway-NO, Philippines-PH, Poland-PL, Portugal-PT, Romania-RO, Russian Federation-RU, Sweden-SE, Slovak Republic-SK, Turkmenistan-TM, Turkey-TR, Taiwan-TW, United States-US X300 distribution includes: AE (United Arab Emirates) , AL (Albania) , AN (Netherlands Antilles), AO (Angola), AR (Argentina), AT (Austria), AU (Australia), AZ (Azerbaijan), BA (Bosnia-Herzegovina), BD (Bangladesh), BE (Belgium), BG (Bulgaria), BN (Brunei), BO (Bolivia), BR (Brazil), BY (Belarus), CA (Canada), CH (Switzerland), CL (Chile), CN (China), CO (Colombia), CV (Cape Verde), CZ (Czech Republic), DE (Germany), DK (Denmark), DZ (Algeria), EC (Ecuador), EG (Egypt), ES (Spain), ET (Ethiopia), FJ (Fiji), FR (France), GB (United Kingdom), GE (Georgia), GR (Greece), HK (Hong Kong), HN (Honduras), HR (Croatia), HU (Hungary), ID (Indonesia), IE (Ireland), IN (India), IR (Iran), IT (Italy), JO (Jordan), JP (Japan), KE (Kenya), KR (Korea, South), KW (Kuwait), KZ (Kazakhstan), LT (Lithuania), LV (Latvia), LY (Libya), MD (Moldova), ME (Montenegro), MK (Macedonia), MM (Myanmar), MU (Mauritius), MX (Mexico), MY (Malaysia), NG (Nigeria), NI (Nicaragua), NL (Netherlands), NO (Norway), NP (Nepal), NZ (New Zealand), PE (PerÂż), PH (Philippines), PK (Pakistan), PL (Poland), PS (Palestine), PT (Portugal), QA (Qatar), RS (Serbia), RO (Romania), RU (Russia), RW (Rwanda), SA (Saudi Arabia), SC (Seychelles), SE (Sweden), SG (Singapore), SI (Slovenia), SK (Slovakia), SV (El Salvador), SY (Syria), TH (Thailand), TM (Turkmenistan), TN (Tunisia), TR (Turkey), TT (Trinidad/Tobago), TW (Taiwan, ROC), UA (Ukraine), US (United States of America), UY (Uruguay), UZ (Uzbekistan), VE (Venezuela), VN (Viet Nam), YE (Yemen), ZA and (South Africa) | Terminated | June 18, 2012 | 671
Z-1831-2011 | 57887 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The enGenTM Laboratory Automation System consists of 3 main components including Instrument Management Subsystem (Instrument Manager), Automated Track Subsystem (TCAutomation) and Configuration File developed by Ortho-Clinical Diagnostics (OCD). The Instrument Manager is a Windows-based software product that operates on a Pentium-class (or higher) PC processor. The instrument management subsystem is capable of connectivity with clinical analyzers, a Laboratory Information Systems (LIS), barcode label printers and other third party devices. The TCAutomation is a pre-analytical and post-analytical sample management system. The configuration file enables operational support of the site specific layout and instrument selection. |  |  | January 27, 2011 | 2011 | Ortho-Clinical Diagnostics | A software anomaly that can occur with analyzers connected to the Bypass modules listed below on an enGen" Laboratory Automation System using TCAutomation" Software Version 3.1.1 and below. When this anomaly occurs, it is possible for a sample other than the intended sample to be aspirated, potentially leading to test results being assigned to the wrong patient without error notification. | Software design | The firm, Ortho Clinical Diagnostics (OCD), sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated January 27, 2011, by Federal Express overnight mail to all customers. Foreign affiliate consignees/customers were notified by email informing them of the issue on the same day. The letter described the problem, product and actions to be taken. The firm also included temporary procedures to help prevent the occurrence of the software anomaly. The customers were instructed to follow the enclosed Procedure to STOP the Automation System and/or Bypass Module; complete and return the attached Confirmation of Receipt - Important Response Required form by February 4, 2011, via fax to: Ortho Clinical Diagnostics, 1-888-557-3759 or 1-585-453-4110, and display this communication near the enGen Laboratory Automation System to ensure compliance by all operators until updated TCAutomation Laboratory Automation System software is installed. A Technical Bulletin will be sent to customers in the near future to be stored with user documentation. Ortho will implement a TCAutomation Laboratory Automation System software change in the near future to resolve this anomaly.If you have any questions regarding this notification, please call Customer Technical Services at 1-800-421-3311 (Options 1, 2). |  | 63 | 63 units | Worldwide distribution: USA and countries including: Canada, Chile, England, France, Germany, Italy, Mexico, Singapore and Spain. | Terminated | May 09, 2017 | 2294 | 63
Z-1839-2011 | 58104 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | 10 ft. Interface Cable, 10 Pin - Shielded Tip Pins to 12 Pin - Blue, Manufacturing Part Number 01287-09, Catalog Number D128709. |  |  | February 15, 2011 | 2011 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster has confirmed that there is an inaccessibility of electronic version of the instruction for use (e-IFU) for the "10 ft. Interface Cable, 10 Pin-Shielded Tip Pins to 12 Pin-Blue", catalog number D128709, through the J&J; Gateway (ie, e-IFU website). | Other | Biosense Webster, Inc. sent an Urgent Field Safety Notice letter dated March 16, 2011, to all affected customers (via express mail). The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the enclosed Acknowledgement Form to confirm that all customers who received the product have been notified accordingly of this concern. If customers have any questions or require additional information, they are to contact their BWI representative and return the Acknowledgement form to:Biosense Webster, Inc.l5715 Arrow HighwayIrwindale, CA 91706Attn: Recall Coordinator - Virgil EllsworthPhone: (909) 839-8500Fax Number: (909) 839-7207.For questions regarding this recall call 909-839-7207. |  | 606 | 606 units total | Nationwide Distribution - including GA, KY, NV, WA, PA, and UT | Terminated | May 07, 2012 | 447 | 606
Z-1833-2011 | 58025 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | REF 32-485050 VANGUARD KNEE INSTRUMENTATION MICROPLASTY SLIDEX ANTERIOR/POSTERIOR, QTY: 1 |  |  | February 10, 2011 | 2011 | Biomet, Inc. | Biomet has initiated this action due to the scale on the stylus of the Vanguard MP Slldex AP Sizer being incorrect for this assembly resulting in the scale being out of position. | Process change control | The firm, Biomet, sent a "FIELD SAFETY NOTICE" dated February 11, 2011 to all customers (heads of Orthopaedic Departments/Operating Departments /Sterile Services Departments/Procurement /Supplies/ Risk Management). The notice described the product, problem and actions to be taken. The customers were instructed to immediately locate and discontinue use of the product; return to Biomet or their local Biomet distributor; ensure that the operating staff are made aware of this issue, and complete and return the FAX BACK RESPONSE FORM to Biomet UK Ltd or their local Biomet distributor via fax at +44 (0) 1656 645454. If you have any questions regarding this communication, please contact the Regulatory &amp; Compliance Manager at +44 (0) 1793 645317 or email: edward.spearpoint@biometeurope.com. |  | 17 | 17 | Foreign only; UK, Malta, The Netherlands, Italy, Turkey, Switzerland. | Terminated | May 16, 2011 | 95 | 17
Z-1835-2011 | 57857 | [6 empty] | Triathlon PKR Peg Drill, Catalog Number: 5650-4-512 (Size 1-2 Peg Drill); Catalog Number: 5650-4-536 (Size 3-6 Peg Drill); |  |  | February 09, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has identified a trend of complaints involving reports of the peg drill cold welding to the drill guide, causing the Peg Drill to cease functioning. If cold welding of these components occurs, the following potential patient and/or surgeon or surgical staff harms may results. | Component change control | Stryker sent Urgent Product Recall letters on February 9, 2011, and February 10, 2011, via Fed Ex. to all affected customers. The letters identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete and fax back the attached Product Recall Acknowledgement Form as soon as possible to 201-831-6069.For questions regarding this recall call 201-972-2100. |  | 13990 | 5650-4-512- 754 units; 5650-4-536- 880 units | Worldwide Distribution - USA , Sweden, Germany; France, Portugal, UK, Canada, Australia, and New Zealand, | Terminated | November 18, 2013 | 1013 | 13990
Z-1830-2011 | 57899 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Interventional Pain 6" 15 G (1.85 mm) Percutaneous Discectomy Probe, REF 407-266, Sterile, Stryker Instruments, Kalamazoo, MI. |  |  | February 03, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | The firm received five complaints that the stylet did not fit down the cannula. The investigation revealed that the incorrect probe assembly was packaged for two DeKompressor part numbers with the lot number of 10214012. | Other | Stryker Instruments Division of Stryker Corporation sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated February 3, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their inventory for any recalled product and quarantine any affected product. If the product was further distributed, customers were asked to forward the letter and the completed Reply Form to all affected locations. Customers were to be contacted to coordinate return of the affected product. Upon receipt of the product, a replacement would be issued to their account. For questions regarding this recall call 1-800-253-3210. |  | 44 | 44 units both products | Worldwide Distribution - USA including the states of PA, NY, KS, NJ, TX, MI, OK, and FL, and the countries of Italy and Switzerland. | Terminated | June 01, 2011 | 118 | 44
Z-1846-2011 | 55905 | No | GKA | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | The VARIANT Sickle Cell Short Program ROM Card; Distributed by and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CA |  |  | April 22, 2010 | 2010 | Bio-Rad Laboratories Inc | The presence of parameters for Version 40400-S on the labeled VARIANT Sickle Cell Short ROM Card instead of Version 80300-S. | Other | The firm, Bio-Rad, sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated April 22, 2010, to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory and quarantine the product; complete and return the MEDICAL DEVICE RECALL RETURN RESPONSE FORM via fax to: (510) 741-3954, Attn: Bio-Rad CSD Regulatory Affairs, and contact there regional Bio-Rad office for ROM Card replacement. If you have any questions, contact Bio-Rad CSD Regulatory Affairs Department at 510-724-7000. |  | 10 | 10 units | Worldwide distribution: USA including: MD, NJ and UT; and countries of: Italy and Singapore. | Terminated | April 07, 2011 | 350 | 10
Z-1852-2011 | 58028 | container, specimen, non-sterile | NNI | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | Large Port Specimen Sock, used to convert the pour/accessory port of a Cardinal Suction canister into a specimen retainer. The sock is used as an accessory in a suction collection system. Its placed in a suction canister to remove solids from the liquid aspirated during surgery. The solids may consist of patient tissue that are then sent to pathology or retained for other purposes. There is no direct contact between this device and the patient. |  |  | February 18, 2011 | 2011 | Cardinal Health Inc | In one lot of product (lot# 1091745), the label for the small version of the sock (catalog number 65652-123) was improperly used for the larger version of the product (65652-124) | Labeling mix-ups | On February 18, 2011 Cardinal Health sent a consignee letter with the following instruction: 1) Quarantine either Catalog Number with the inticated Lot Number.2) Contact the appropiate Customer Service group for return instructions:Hospital - 800-964-5227Federal Government - 800-444-1166Distributirs - 800-635-6021All other customers - 888-444-54403) Return the enclosed acknowledge form via fax to 847-689-9101, even if you do not have affected product, as Cardinal Health is required to confirm receipt of this notification from you, and to prevent further notices.4) Notify any customer to whom you may have distributed product affected by this recall. |  | 3360 | 3,360 each | Worldwide Distribution - USA including AL, AR, AZ. CA. CO,, CT, DC, DE, FL, GA, HI, IA, ID,IL,IM,KS, KY, LA, MA, MD, ME. MI MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, &amp; Puerto Rico - and the country of Canada. | Terminated | October 04, 2011 | 228 | 3360
Z-1850-2011 | 58105 | No | NRU | Dental chair and accessories. | Dental | Dental | 510(K) Exempt | Pelton &amp; Crane Spirit 3000 dental chair (Model #SP30) |  |  | March 04, 2011 | 2011 | Pelton &amp; Crane Company | The Link Arm Assembly that connects some rear mounted options to the Spirit 3000 or Spirit 1800 dental chairs may potentially break and fall off the dental chair. | Process control | Pelton &amp; Crane sent an "Urgent Medical Device Recall" letter dated 3/04/2011. They were instructed to: Locate the affected devices this recall is in regards to. Schedule service with the end user for Pelton &amp; Crane representative to replace the affected components this recall is in regards to. International accounts were instructed to order link arm assemblies from Pelton &amp; Crane and replace the affected components. All consignees were to complete the Recall Acknowledgement/Return Form and return the form to Pelton &amp; Crane. |  | 9 | 9 units | Distribution to USA, Canada, and the Netherlands. | Terminated | August 16, 2011 | 165 | 15
Z-1847-2011 | 58066 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | BD Phoenix" Panel NMIC/ID-123, Catalog # 448723, packaged 25 panels/carton, and labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, MD 21152 USA (800) 638-8663*** |  |  | February 28, 2011 | 2011 | Becton Dickinson &amp; Co. | In vitro diagnostic test kit is defective and may cause false positive or false negative results. | Mixed-up of materials/components | BD Diagnostics Systems sent an Urgent Product Recall letter dated February 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinsue use of the affected product and discard any remaining packages. BD would issue replacements for the discarded material. Customers were instructed to complete the attached form whether or not they had any inventory remaining so that the firm may acknowledge receipt of this notification. Complete and fax the form to BD Regulatory Compliance to 410-316-4258. For further assistance contact BD Customer Service at 1-800-675-0908. For other inquiries, contact BD Technical Services at 1-800-638-8663. |  | 377 | 377 cartons | Worldwide Distribution - USA including AZ, CA, FL, GA, MA, MD, NC, NY, PA, TX, WA, WI, and WY and the countries of The product was distributed to medical facilities nationwide and to foreign consignees in Argentina, Brazil, Chile, and Mexico. | Terminated | January 30, 2013 | 702 | 377
Z-1817-2011 | 57004 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | 115V Blanketrol III Model 233 Hyper/Hypothermia System, Part number: 86107; and 115V CoolBlue Model 25-01 Hyper/Hypothermia System, Part number: 86000. |  |  | October 18, 2010 | 2010 | Cincinnati Sub-Zero Products Inc | The firm initiated this recall due to changes that have been made to the Blanketrol III Model 233 device's Operation and Operation/Technical Service Manuals, in order to stay in compliance with labeling regulations. | Other | On 10/18/2010 &amp; 11/5/2010 the firm sent URGENT Medial Device Field Action letters, dated 10/18/2010, to their customers. The letters identified the affected product and stated the reason for recall. Customers were instructed to immediately collect and discard all affected product Operation and Operation/Technical Service manuals currently in their possession. They were to replace all the discarded manuals with the revised versions that were provided with the letter. The enclosed response form should be completed and returned by fax to 513-772-9119, scanned and e-mailed to B3manual@cszinc.com. If the product had been further distributed, then those customers should also be identified and notified of the field action via a copy of the letter. Customers should direct questions to the field action coordinator, David Niehaus at 513-773-8810 ext. 3234. |  | 1385 | 763 - 115V Blanketrol III Model 233; &amp; 133 - 115V CoolBlue Model 25-01 | Worldwide Distribution -- US, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, &amp; WI and countries of ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, CHILE, CROATIA, ECUADOR, EL SALVADOR, GERMANY, HONG KONG, INDONESIA, ISRAEL, ITALY, KOREA, MALAYSIA, MEXICO D.F., NEW ZEALAND, PEOPLE'S REPUBLIC OF CHINA, PERU, REPUBLIC OF SIGNAPORE, SPAIN, TAHITI, THAILAND, THE NETHERLANDS, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY &amp; VENEZUELA. | Terminated | March 08, 2013 | 872 | 1385
Z-1845-2011 | 57613 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CellQuest Pro SW, version 4.0.2. and later; |  |  | September 07, 2007 | 2007 | BD Biosciences, Systems &amp; Reagents | Results and gates could be mismatched if you utilize the undo function to undo a gate that was deleted. | Software design | Important Product Information letters, dated September 2007, were sent the week of 9/14/2001 via Federal Express to all affected customers. The letters identified the affected product along with the reason for recall. The letter states that "as long as customers do not use the Undo function to undo a gate deletion, the software can be used with full functionality." The letter also states that BD will change the issue in the next software version. If customers believe that they may have analyzed sample files by utilizing the Undo function, then it is recommended that the sample files be re-analyzed. If customers need more information they should contact the Customer Support Center at 877-232-8995. |  | 4552 | 4552 | Worldwide Distribution | Terminated | March 31, 2011 | 1301 | 4552
Z-1849-2011 | 57602 | No | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | HEMOCUE AB Hb 201 DM ANALYZER, in vitro diagnostic device for quantitative hemoglobin determination. |  |  | October 07, 2010 | 2010 | Hemo Cue, Inc. | Incorrect configurations settings in software. As a result, certain error codes do not trigger appropriately. These error codes are: E01, E02 and E05 which are all related to the optronic performance of the analyzer. Possible incorrect results. | Process control | Hemocue AB sent Customer Notification "Notice of Medical Device Correction on HemoCue Glucose 201 DM and HemoCue Hb 201 DM Analyzers" dated October 7, 2010, to customers notifying them of incorrect settings on error codes. The corrective action consists of an update of the configuration settings on the the affected HemoCue Glucose 201 DM and Hb 201 DM analyzers that are in the field. The affected units in the field will be corrected in accordance with a rework procedure that has been established by HemoCue AB. |  | 79 | 79 | Worldwide distribution, including USA, and Glucose 201 DM analyzers to Norway and the Netherlands; Hb 201 DM analzyers to the Netherlands, France, Ireland, Great Britain, Poland and Germany. | Terminated | June 22, 2011 | 258 | 102
Z-1853-2011 | 54844 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson Antares Ultrasound Imaging System with software versions 4.0 and 5.0. |  |  | March 01, 2010 | 2010 | Siemens Medical Solutions USA, Inc. | Image captured from the device may be erroneous. | Software design | Acuson Customer Safety Advisory notification letters were sent March 1, 2010 by certified mail. |  | 3539 | 3539 | Worldwide distribution to USA, Andorra, United Arab Emirates, Netherlands Antilles, Angola, Austria, Australia, Azerbaijan, Bosnia, Bangladesh, Belgium, Bahrain, Brunei Darussalam, Brazil, Belarus, Canada, Switzerland, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, Faroe Islands, France, Great Britain, Georgia, Greece, Guatemala, Hong Kong, Croatia, Hungary, Indonesia, Ireland, Israel, India, Iran, Italy, Jordan, Japan, South Korea, Kuwait, Lebanon, Sri Lanka, Lithuania, Latvia, Morocco, Moldova, Mexico, Nicaragua, Netherlands, Norway, New Zealand, Oman, Peru, Philippines, Pakistan, Poland, Portugal, Qatar, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovak Republic, Suriname, Thailand, Turkey, Trinidad and Tobago, Taiwan, Ukraine, UA, Uruguay, Venezuela, Vietnam, Yemen, and South Africa. | Terminated | June 18, 2012 | 840 | 3539
Z-1854-2011 | 57773 | No | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | HGBPROKIT which contains one meter and one container of 25 test strips. HGB TEST 100; HGB PRO TEST STRIPS -100, includes four containers of 25 test strips per container; International Technidyne Corporation |  |  | January 21, 2011 | 2011 | ITC-Nexus Dx, Inc. | ITC has determined that results obtained using certain lots of Hgb Pro Professional Hemoglobin Test Strips may differ from results obtained using a reference hemoglobin testing system. | Other | ITC-Nexus Dx, Inc. sent an Urgent Medical Device Recall letter with return response forms to end users and distributors dated January 21, 2011 via UPS overnight. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory to determine if they had an of the affected lots.If they have full boxes or loose containers of the affected lots customers were instructed to stop using the product and remove it from inventory.Customers were requested to fill out the return response form and send it back to ITC via fax, email or mail.For questions customers were instructed to call ITC Technical Support at 800-631-5945 (US) or 732-548-5700 (International)), Extension 4707 or e-mail techsupport@itcmed.com. |  | 50563 | 2500 containers + 50,563 addtnl containers =53063 containers | Worldwide Distribution - USA (nationwide) and the countries of India, Oman, South Africa, and Singapore | Terminated | November 02, 2012 | 651 | 50563
Z-1859-2011 | 57987 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Stanley Bed-Check 90-Day LT Bed Sensormat, With 5.5 Ft. Cord, Model 74001, 5 mats per inner box, 5 inner boxes per overpack box (total 25 units). The firm name on the box is Stanley, 1-800-523-7956. The mat contains only the phone number. |  |  | November 25, 2009 | 2009 | Stanley Security Solutions, Inc. | The cord could come loose from the mat potentially causing a short to the internal wiring. | Process change control | The firm, Stanley Healthcare Solutions, issued an "Important Technical Bulletin" dated November 25, 2009 to all customers via e-mail. In addition, the firm issued a "Compliance Notice" dated August 30, 2010 via certified mail because the firm had not collected all of the units with the defect. The bulletin described the product, problem and actions to be taken. The customers were instructed to check the cords on the mats to make sure they weren't pulled out, if the cords were pulled out, they were to return the mats for replacement; contact Customer Service at 1-800-824-2996 ext: 3440 to obtain a return authorization number and take immediate action to return the units to the firms location in Lincoln, Nebraska. The Compliance Notice informed the customers who had not sent in their units to promptly read the notice and take the necessary steps to respond appropriately for their facility. Additionally, the customers were instructed to complete and return the Bed Check Sensormat Return Form and page 2 of the "Compliance Notice" via at 402-475-4281 or mail to: Compliance Engineer, Stanley Security Solutions, Inc., 1550 N. 20th Circle, Lincoln, NE 68503. Follow-up phone calls to nonresponders were made using a phone script. A "Second Notice Compliance Notice" dated October 20, 2010 was issued via certified mail to customers asking for immediate response to the previous communication. In addition, a third notice dated October 28, 2010 which was also flagged as "Second Notice" was again used asking for a response from the customers.If you have any questions regarding the notice or any action that may need to be taken by you, please contact Quality Manager at 1-800-824-2996 extension 3419 or Compliance Officers at 402-742-9414 or 402-742-9425. |  | 3500 | 3,500 mats | Worldwide distribution: USA and the country of Canada. | Terminated | March 31, 2011 | 491 | 209183
Z-1862-2011 | 57852 | No | HTJ | Depth gauge for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Xcelerate Specialty Patella Milling System Per File K2399; Catalog Number I-K2399PR00, I-K2399PR10; Stryker Howmedica Osteonics, Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430 |  |  | January 27, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that the 10 mm on the bone remaining depth gage is incorrect thus allowing for the reamer to leave as little as 8.9mm (in a nominal condition) patella bone. | Device Design | Stryker sent an Urgent Product Recall letter dated January 27, 2011, via FedEX for January 28, 2011, delivery with return receipt. The letter identified the product, the problem, and the action to be taken by the customer. Customers were insructed to fax back the attached Product Recall Acknowledgement Form within 5 days to 201-831-6069. Customers were also instructed to contact the hospitals in their territory that have the affected product to arrange return of the product. Product should be returned to the attention of: Regulatory ComplianceStryker Orthopaedics325 Corporate DriveMahwah, New Jersey, 07430. The orange Product Remediation sticker should be attached to the return, indicating the Product Remediation # RA 2009-296. Mark the outer box with words "Product Recall."For questions regarding this recall call (201) 972-2100. |  | 52 | 52 units | Nationwide Distribution | Terminated | October 15, 2012 | 627 | 52
Z-1866-2011 | 58136 | No | MGP |  |  | General &amp; Plastic Surgery | 510(k) | HydroHydrofilm is for use as a post-operative and primary dressing to protect against secondary infection and to cover non-exudating wounds and for use as a secondary dressing for fixation of catheters or cannulae.film, 10cm x 12.5 cm, Paul Hartmann AG, 89522 Heidenhiem, Germany |  |  | February 11, 2011 | 2011 | Hartmann USA, Inc | Misbranded; the expiry date on the primary packing incorrectly indicates 2010-10, instead of 2015-10. Consignees were notified by letter on/about 02/11/2011. They were advised to examine their inventory and not to use any of the affected product and to return it to Paul Hartmann. A response form was included to be completed and returned to Paul Hartmann. They were further instructed to pass the no | Other | Hartmann USA, Inc. sent letter to the Consignees on 2/11/11. They were advised to examine their inventory and not to use any of the affected product and to return it to Hartmann USA, Inc. A response form was included to be completed and returned to Hartmann USA, Inc. They were further instructed to pass the notice on to their consignees if distributed. For questions, call (803) 325-7600. |  | 1250 | 100 boxes (10 x 12.5 cm per box) | Nationwide -- Including CA, GA, IL, MO, MS, TX, and VA. | Terminated | August 12, 2011 | 182 | 1250
Z-1875-2011 | 58051 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | CVS Pharmacy Plastic Sesame Street Adhesive Bandages; a sterile, latex free, assorted sizes with multiple designs adhesive bandages; 20 bandages per box, 24 boxes per case; Distributed by: CVS Pharmacy, Inc., One CVS Drive, Woonsocket, RI 02895, Made in Egypt; REF CVS 826159, UPC 0 50428 95118 7 |  |  | February 21, 2011 | 2011 | Pharmaplast S.A.E. | The bandages are labeled as sterile, but failed sterility testing. | Process control | Pharmaplast S.A.E. called the direct account on 2/17/11 and sent a Urgent Medical device recall letter dated 2/21/11 to the direct account via overnight mail, requesting that they examine their inventory for the non-sterile adhesive bandages and quarantine them for return to Pharmaplast via sea freight. The account was instructed to contact their customers concerning the recall and return their stocks of the affected lot, which was done by telephone calls on 2/17/11.For questions regarding this Recall please call 00203-4500264 |  | 5336 | 5,336 cases | Nationwide Distribution - Including Illinois. | Terminated | August 25, 2011 | 185 | 5336
Z-1865-2011 | 57890 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | Axxent Flexishield Mini |  |  | February 03, 2011 | 2011 | Xoft Inc. | Xoft Axxent Flexishield Mini product may may shed particles identified as tungsten, and may look like suspicious calcifications on follow up scans if the product was used during Intraoperative radiation therapy. | Device Design | Xoft, Inc. sent an Urgent Field Removal Action letter dated February 3, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using and return all units and lots of Flexishield Mini Catalog Number: F5300 in their inventory. Customers were instructed to fill out the attached form and fax to 866-222-3404, then call Customers Service toll free directly at 877-963-8327 to make arrangements to return the product and receive a credit if applicable. Customers were asked to complete and return the form even if they did not have any units to return.For questions regarding this recall call 408-419-2300. |  | 165 | 165 units | Nationwide Distribution including AL, CA, CO, FL, IL, MI, MO, MS, NJ, NY, OK, PA, TN, and WI. | Terminated | June 08, 2011 | 125 | 165
Z-1863-2011 | 57742 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cardiac Marker Control Level 1, List Number 136601-06F12-02; Cardiac Marker Control Level 2, List number 136602/06F13-02; and Cardiac Marker Control Level 3, List number 136603/06F14-02. |  |  | December 17, 2010 | 2010 | Abbott Point Of Care Inc. | Presence of CO2 in the head space of the vial can result in the generation of high results in Cardiac Marker Control Level 1, 2, and 3 which is used along with the Cardiac Marker Control Calibration Verification Set. | Device Design | Urgent Recall Notice letters with business reply cards were sent Fed ex on December 17, 2010 to all accounts. International notifications were sent on December 21, 2010. The letter identified the affected product and lot numbers along with a background explanation of the issue. The firm recommends that customers remove the cap from the vial; squeeze the vial 10 times, prior to thawing; replace cap on the vial; and then follow the remaining steps for use as outlines in section 14-12 of the i-STAT 1 System Manual. Customers are to complete and return the Business Reply card to acknowledge receipt of the letter. If the affected product had been further distributed then a copy of the letter should be provided to those customers as well. If additional information is needed, customers should contact Abbott Point of Care Technical Support at 800-366-8020, Option 1 or their Abbott Point of Care representative. |  | 2360 | 2360 units total | Worldwide Distribution | Terminated | October 16, 2012 | 669 | 2360
Z-1620-2011 | 51208 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BD FACS Sample Prep Assistant (SPA and SPA II) Instrument, |  |  | January 03, 2009 | 2009 | BD Biosciences, Systems &amp; Reagents | To prevent a needle stick while replacing the probe, the firm has initiated a Field Correction, and issued an improved Users Guide procedure (User Bulletin, Part Number 645050). | Nonconforming Material/Component | The firm, BD Biosciences, sent an information letter dated December 2008 with an enclosed New Probe Replacement Service/User Bulletin, to the customers. The letter described the product, problem and action to be taken. The customers were instructed to review the attached User Bulletin (Part Number 645050); ensure that their laboratory follows the updated procedure for removing and replacing probe; place the User Bulletin in their BD FACS SPA or SPA II User's Guide; always follow the instructions in User Bulletin to properly remove and replace probe, and complete and return the Effectivity Form via fax to 408-577-1409. Should you have any comments or questions, please contact BD Biosciences at 1-877-232-8995 in the US. Outside the US, contact your regional BD office or field service engineer for assistance. |  | 70 | 70 | Worldwide distribution: USA and countries including: Hong Kong, Australia, Belgium, Canada, El Salvador, Hong Kong, Japan, and Mexico. | Terminated | March 10, 2011 | 796 | 70
Z-1650-2011 | 57404 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | BIO-RAD BioPlex 2200 Syphilis IgG Control Set. Lot # 46636 |  |  | November 23, 2010 | 2010 | Bio-Rad Laboratories | An upward shift in the control values can result in a control value that falls outside of the current manufacturing ranges. | Other | The firm, Bio-Rad, sent a "Voluntary Field Correction: QC Ranges for Syphilis IgG Control Set Lot#46636" letter to Domestic customers via FedEx on November 24, 2010 and to the International customers via e-mail on November 23, 2010. The letter described the product, problem and actions to be taken by customers. The Customers were advised to temporarily establish their own means and acceptable ranges for Control Lot 46636 for the r15 and r47 markers. The firm will provide adjusted recommended means and acceptable ranges. Additionally, the customers were to complete and return the Customer Response Form for USA and/or the Subsidiary Response Form (Outside US) via fax to (510) 741-4846, Attn: Bio-Rad BPD-Product Support.If you have any questions about this notification, contact your local Bio-Rad Technical Support. |  | 202 | 202 | Worldwide distribution: USA and countries including: Canada, Czech Republic, France, Germany, and Italy. | Terminated | August 04, 2011 | 254 | 202
Z-1519-2011 | 57107 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Pacific Hemostasis Âż Thromboplastin-DS, catalog #100354, packaged in 4 mL lypholyzed vials, 10 bottles/carton. Vials are labeled, in part ***Fisher Diagnostics a division of Fisher Scientific Company, LLC Middletown, VA 22645-1905 USA*** |  |  | October 06, 2010 | 2010 | Fisher Diagnostics, A Company of Fisher Scientific LLC | An in vitro diagnostic reagent used in coagulation testing may cause inaccurate clotting times in quality control and patient plasma samples. | Nonconforming Material/Component | The firm, ThermoFisher Scientific, sent two "Urgent Medical Device Removal" letters one via Federal Express on October 20, 2010 and certified mail for domestic customers on October 21, 2010. The letters described the product, problem and action to be taken by the customers. The customers were instructed to indicate the quantities of all affected product currently in stock at their facility; complete and return the attached Response Form/table to Fisher Diagnostics via fax to (540) 869-8129 or by mail to the attention of Technical Support; dispose of any remaining inventory of the affected lot numbers; provide a copy of this letter to their customers, and file this letter for future reference.If you have any questions regarding this notification, please contact our Technical Support staff at 800-528-0494, Option 2, or email techsupport.diagnostics.mtn@themofisher.com. |  | 1190 | 1190 kits | Worldwide distribution: USA including states of: Alaska, Florida, Kentucky, Ohio, New Jersey, New York, and Texas; and countries including: Argentina, Dubai, Italy, Mexico, Philippines, Romania, and Turkey. | Terminated | November 02, 2011 | 392 | 1190
Z-1740-2011 | 57705 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | UC Tibial Insert, Size 4, 13mm, REF 163-3413, Sterile R, Ortho Development, Draper, Utah 84020. |  |  | January 07, 2011 | 2011 | Ortho Development Corporation | Components of a replacement knee system could be mislabeled. | Error in labeling | The firm, Ortho Development, sent a "Field Notification Memo" date January 7, 2011 to its customers. The memo described the product, problem and action to be taken. The customers were instructed to promptly return any remaining parts of the listed product; locate all the parts in their inventory; identify any of the listed parts that may have been used in surgery; verify the quantity they have on hand and the quantity used in surgery and reconcile them with the quantity sent; if discrepancy exists, notify customer service resolution; prepare items for return and contact customer service to obtain RMA number. Once reconciled, arrangements would be made for sending new inventory and the return of unused parts.If you have any questions regarding this issue, please notify Regulatory Affairs Manager at (801) 619-3450. |  | 25 | 25 units | Worldwide distribution: USA including states of: AZ, CA, FL, MI, TX, and UT; and country of: Japan. | Terminated | December 27, 2011 | 354 | 72
Z-1728-2011 | 57814 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hc2 System Software Suite 4.0 version 3.0, available as component in Qiagen user guide, catalog #5050-00162, and as a component within a Qiagen PC, available as catalog # 5050-11140A. |  |  | January 12, 2011 | 2011 | QIAGEN Gaithersburg, Inc. | The software component of a diagnostic medical device may be defective and could cause incorrect patient values. | Software design | The firm, Qiagen, sent a "Notice of Product Field Action" letter dated January 12, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue the process of running multiple assay protocols on a single microplate to eliminate the potential for inaccurate patient result reporting, and to complete and return the acknowledgment section via fax to QIAGEN 661-702-3854 or scan the document and email to techservice-na@qiagen.com. The firm informed users that they will provide further detailed instructions for a workaround and software fix. A QIAGEN representative will contact the customers who indicated that they run multiple assays on a plate and are currently using the HC2 System Software Version 3.0 (Suite 4.0), to assist in the identification of those patients that may have been affected by the software defect.If you have any questions or concerns, please contact Qiagen Technical Services at 800-344-3631, Option 2. |  | 163 | 163 | The recalled product was sold to medical laboratories nationwide and in Canada. | Terminated | November 14, 2011 | 306 | 163
Z-1729-2011 | 57631 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension(R) Creatinine Flex(R) Reagent Cartridges |  |  | December 27, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed the potential for falsely elevated creatinine results with specified CREA Flex(R) reagent cartridge lots when processing tests from a well that has been punctured. This issue may affect creatinine results for quality control (QC) and patient samples. | Material/Component Contamination | The firm, SIEMENS, sent an "Urgent Field Safety Notice" dated December 2010 to all customers. The letter described the product ( Dimension(R) CREA Flex(R) Reagent Cartridges DF33A lots BB1286, CA1260 and FB1251), problem and actions to be taken. The customers were instructed to please discontinue use and discard any remaining inventory of the product; complete and return the fax form to the Customer Service Support for replacement of the product; if they must continue use of the product until they receive a replacement, open a new well set every 8 hours to minimize any impact to CREA test results, and complete and return the FIELD CORRECTION EFFECTIVENESS CHECK form via fax to the Technical Solution Center at 302-631-8467.If you have any questions, contact the Technical Solution Center at 800-441-9250. |  | 20177 | 20,177 | Worldwide distribution: United States and the following foreign countries: Austria, Belgium, Bosnia and Herzegovina, Canary Islands, Croatia, Czech Republic, Egypt, Estonia, France, Germany, Great Britain, Greece, Hungary, Italy, Latvia, Lithuania, Kuwait, Macedonia, Morocco, Norway, Poland, Portugal, Qatar, Romania, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland, Turkey and the United Arab Emirates. | Terminated | December 21, 2013 | 1090 | 20177
Z-1877-2011 | 57445 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CYTO-STAT tetraCHROME CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5, Part Number: 6607073. |  |  | September 16, 2010 | 2010 | Beckman Coulter Inc. | A labeling error resulted in the specimen age stability and the prepared sample stability being stated incorrectly in the instructions for use. There is no impact to the product safety or functionality. | Pending | An Urgent: Product Corrective Action letter dated, September 27, 2010, was sent to customers. The letter identified the affected product and stated the reasons for the recall. Customers were also asked to take the following actions: ACTION/RESOLUTIONFor issue 1: The specimen stability claims should be as follows:- 48 hours for both % positives and absolute counts for the CD45/4/8/3 tube- 24 hours for both % positives and absolute counts for the CD45/56/19/3 tubeFor Issue 2: Results produced are not impacted when samples are prepared and run within the post collection time frames provided in the product instructions for use:- Preparation within 72 hours of collection for the CD45/4/8/3 tube- Preparation within 24 hours of collection for CD45/56/19/3 tubeBeckman Coulter requested consignees complete and return the enclosed response form within 10 days so they may maintain their records. The firm also asked that they share the information with laboratory staff, and retain this notification as part of their Quality System documentation.Questions regarding the Product Corrective Action were directed to (800) 526-7694 in the United States or (800) 463-7828 in Canada. If outside the US or Canada, consignees were instructed to contact their local Beckman Coulter Representative. |  | 26822 | 26822 units | Worldwide Distribution -- United States, Albania, Australia, Bahrain, Bangladesh, Barbados, Belarus, Belgium, Bosnia and Herzegovina, Bulgaria, Burundi, Cameroon, Chile, China, Colombia, Croatia, Czech Republic, Denmark, El Salvador, Finland, France, French Guiana, French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Republic of Kuwait, Libyan Arab Jamahiriya, Macao, Malaysia, Mexico, Monaco, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan Panama, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sudan, Sweden, Switzerland, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | June 14, 2012 | 637 | 105810
Z-1883-2011 | 55554 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems enhanced AccuTnl Reagent Kit, Part Number: A78803 |  |  | December 10, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed that under certain circumstances, the following assays may produce elevated values at the low end when run in conjunction with the enhanced Access Accu Tnl assay (P/N A 78033):(1) Access Total BhCG (P/N 33500);(2) Access HYPERsensitive hTSH/Fast hTSH (P/N 33820);(3) Access hFSH (P/N 33520).The issue can occur only when the assays are | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated December 16, 2009 to all customers. The letter described the product, problem and actions to be taken. Beckman Coulter instructed the customers to do the following mitigation to reduce the risk of encountering an erroneous result:Systems with multiple reagent pipettors (UniCel Dxl 800 and 600; UniCel DxC 880i, 860i, 680i, and 660i):1. Designate enhanced Access AccuTnl to a different reagent pipettor (s) than Access Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH.2. Discard any open Access Total BhCG, hTSH/Fast hTSH and hFSH reagent packs3. Calibrate Access Total BhCG, hTSH and hFSH using fresh reagent packs4. Run QC following acceptable calibration curves. Proceed with operation following acceptable QC performanceSystems without multiple reagent pipettors (Access, Access 2; SYNCHRON LXi 725, UniCel DxC 600i):1. Discard any open Access Total BhCG, hTSH/Fast hTSH and hFSH reagent packs.2. Calibrate Access Total BhCG, hTSH/Fast hTSH and hFSH using fresh reagent packs3. Run QC following acceptable calibration curves.Ensure that controls and patient samples for Access Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH are not run immediately following enhanced Access AccuTnl controls or patient samples.For example, during routine QC runs, ensure that these assays are requested prior to enhanced Access AccuTnl.4. Proceed with operation following acceptable QC performance.NOTE: An enhanced AccuTnl sample loaded with a system priority of STAT will interrupt routine Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH sample aspiration. Avoid loading enhanced AccuTnl sample aspiration. Avoid loading enhanced AccuTnl samples with a system priority of STAT during aspiration of Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH.Diagnostic test results should be interpreted in light of the clinical presentation of the patient, including: symptoms, clinical history, data from other |  | 22032 | 22,032 kits | Worldwide distribution: USA and country of Canada. | Terminated | May 07, 2012 | 879 | 22032
Z-1884-2011 | 52820 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | TEE Transducer V5M, when used with the Acuson S2000 Ultrasound System, software versions: VA15, VA15A or VA15B; Model Numbers: 10041461 (Acuson S2000) and 8263703 (V5M Transducer); |  |  | April 07, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Device function failure--When performing a TEE exam with the transducer, CW (Continuous Wave) mode is not functional. Due to work flow, this may not be apparent until the transducer is inserted into the patient's esophagus. | Software design | Siemens sent a Customer Safety Advisory Notification letter on April 7, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to pass the notice on to all those within their organization who need to be aware of the issue until the corrective action is completed.For questions regarding this recall call 650-969-9112. |  | 61 | 61 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Switzerland, China, Czech Republic, Germany, Denmark, Great Britain, Hong Kong, Indonesia, India, Italy, South Korea, Norway, Turkmenistan, and Taiwan | Terminated | April 05, 2011 | 728 | 61
Z-1879-2011 | 57810 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King System , King LTS-D Kit Size 5, Sterile, REF KLTSD405, Manufacturer: King Systems Noblesville, IN. |  |  | January 24, 2011 | 2011 | King Systems Corp. | The firms contract manufacturer reported that the products were below the specification for hardness. This could contribute to a higher incidence of problems during the insertion and final positioning of the airway in the patient. In rare instances the softness of the tube could cause the tube to fold over. | Other | King Systems Corp. sent an URGENT DEVICE RECALL NOTICE letter dated January 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. The letters instructed users to cease all sales of the recalled product, to notify the consignees for any product that was further distributed, and to return all recalled product to King Systems. The notice did include a response form.For questions call 317-776-6823, ext 257. |  | 1430 | 143 cases of 10 each case | Worldwide Distribution - USA and Canada | Terminated | June 29, 2012 | 522 | 2120
Z-1882-2011 | 58044 | No | LLC |  |  | General Hospital | 510(k) | iSite PACS with iSite ViewForum Applications, R7.4V6 L1 L2, L3, |  |  | February 23, 2011 | 2011 | Philips Medical Systems | A software anomaly that occurs when clinicians use Volume Vision to view images in 3D and uses the measurement tool to measure regions of interest. This could lead to an incorrect interpretation of the measurements which could cause an incorrect diagnosis and incorrect treatment plan. | Software design | Philips sent an Urgent Field Safety Notice letter dated February 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to assure that all colleagues were aware of the problem.When using the Volume Vision software, check the correctness of measurements (graphics and numerical values) prior to saving a captured image containing measurements.Save captured images containing measurements immediately after creating them and before altering the measurements in the original image view.For any subsequent measurements, create a new capture and save it immediately.For further information or support concerning this issues, customers should contact their local Philips representative at +1-877-328-2808. |  |  |  | Worldwide Distribution. | Terminated | August 18, 2011 | 176
Z-1880-2011 | 58013 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | POUCH/CARTON LABEL: |  |  | November 30, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | On Nov 30, 2010 ConMed Linvatec, Largo, FL recalled their Micro Sagittal Blade, Fine Lot number 180209, Product Number 5023-138. Some units of Micro Sagittal Blade, Fine from Lot 180209 may have compromised seals in the sterile packaging. | Process control | Linvatec Corp. dba ConMed sent an Urgent Medical Device Recall Notification letter dated November 30, 2010, by FedEx, email, Fax, etc) to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the reply form within 30 days of receipt. Immediately check their facility's inventory for the affected product and lot number. If they had any of the product with lot number 180209. Do Not Use. Segregate and return for credit. If the product was transferred to another facility, customers were to complete the REPLY FORM documenting the transfer of the product and fax the form to 727-319-5701.For questions regarding this recall call 727-319-5701 or 800-535-8536. |  | 415 | 415 | Worldwide Distribution - USA including CT, FL, GA, HI, IN, MO, MS, MT, NJ, NY, OH, PA, SC, and TX and the countries of Canada, and South Africa. | Terminated | November 09, 2011 | 344 | 415
Z-1896-2011 | 58011 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Rusch Irrigation Tray w/Piston Catalogue Number 67890 |  |  | February 24, 2011 | 2011 | Teleflex Medical | The irrigation trays contain alcohol prep pads which are included in the scope of a recall initiated by Triad Group, due to a potential bacterial contamination. | Other | the firm,Teleflex Medical, sent two "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letters one dated February 24, 2011 and the other dated March 30, 2011 to all consignees and customers. The letters described the product, problem and actions to be taken. The consignees were instructed to communicate the recall to any of their customers who have received product included within the scope of the field corrective action via the customer letter. The customers were informed in the letter of concerns about potential bacterial contamination of the alcohol prep pads in the affected trays. If contaminated alcohol prep pads are used, there is a potential for infection to occur. The customers were instructed "that upon opening the affected trays for use, they should discard the prep pads as medical waste". The customers were also instructed to immediately discontinue use and quarantine any products with the product numbers and lot numbers listed; return the product, and complete and return the Recall Acknowledgement Form, even if they no affected product, via fax to 1-866-804-988, Attn: Customer Service (a customer service representative contact the customer with a Return Goods Authorization number and instructions for the return of the product to Teleflex Medical). If you have any questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 95647 | 95,647 eaches (total for both Catalogue Numbers) | Worldwide distribution: USA including states of: AL, AZ, CA, FL, GA, IL, MS, MI, MO, MS, NC, NJ, NH, NY, OH, PA, TX, UT, VA, Puerto Rica and the country of: Canada. | Terminated | April 17, 2012 | 418 | 95648
Z-1893-2011 | 58009 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ***REF VP6528***Qty 1***Articulating Arm, Secondary Flat Panel***RX Only***See Instructions for Use***CONMED LINVATEC, CONMED LINVATEC, 11311 CONCEPT BLVD, LARGO, FL 33773-4908 USA*** |  |  | September 17, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | If the VHM Arm is left in the locked position, the gas spring can lose pressure and releasing the lock can result in a sudden downward movement. | Device Design | An Urgent Service Bulletin Notification was mailed to each ConMed Linvatec customer who received the affected product. The letter identified the affected product and stated the reason for notification. The letter asks customers to review the GCX Corporation Service Bulletins and updated Installation/User Manual. ConMed Linvatec recommends that customers periodically check the VHM Arm to see that it is functioning properly. In addition "Sudden Movement" stickers, provided by GCX, were also included to all consignees as part of the mailing. If customers in the US have questions or need technical support, they can contact ConMed Linvatec Endoscopy Customer Service at 888-292-0100 or e-mail Customerserv1@linvatec.com. |  | 373 | 373 | Worldwide Distribution -- Australia, Canada, Morocco, Saudi Arabia, Taiwan, Province of China, and throughout the U.S. | Terminated | May 17, 2011 | 242 | 373
Z-1894-2011 | 58109 | [6 empty] | BD Beaver XStar Safety 2.5mm Crescent Knife; |  |  | February 28, 2011 | 2011 | Beaver-Visitec International Inc. | Loose blades in the handle of knives due to uncured epoxy. | Nonconforming Material/Component | Beaver Visitec International, Inc.sent an Urgent Recall Notification letter dated February 25, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all unused product and to complete a Customer Response Form (to be returned via fax to 866-906-4304.For questions regarding this recall call 781-906-7917. |  | 1680 | 1680 units (1130 domestic, 550 foreign) | Worldwide Distribution - USA including AZ, MA, MI, MO, NC, NJ, NY, OH, OR, PA, RI, TN, TX, VA and the country of China | Terminated | June 14, 2012 | 472 | 1680
Z-1886-2011 | 58107 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Harmony LA/LED Monitor Arm Adapter. Model Number: LB61 |  |  | February 18, 2011 | 2011 | Steris Corporation | The arm adapter does not contain labelling specifying the applicable weight and moment limitations of the arm adapter | Other | Steris Corporation sent an Urgent Field Correction Notice letter dated February 18, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A Steris representative would contact customers to arrange for replacement of the product. For questions regarding the firms visit to their facility, customers were instructed to contact Steris Field Service Dispatch at 1-800-333-8828.For any questions regarding this matter call 440-392-7705. |  | 14 | 14 Units | Worldwide Distribution - USA including CO, IL, IN, MO, NC, NV, NY and OH and the countrries of Canada, China, and Australia. | Terminated | March 13, 2013 | 754 | 14
Z-1868-2011 | 57939 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Curlin 6000 IOD Ambulatory Infusion Pump |  |  | February 28, 2011 | 2011 | Zevex International, Inc. | Infusion pumps may exhibit false error codes which render the device inoperable until a recovery process is performed. | Software design | The firm, MOOG, sent an "URGENT DEVICE RECALL NOTIFICATION" letter dated February 28, 2011 to its customers. The letter described the product, problem and actions to be taken. The Customers were instructed to do the following:1) Identify the affected Pumps:a) If they have an affected pump, fill out and return the enclosed Business Reply Card within five (5) working days.b) If their pump is not affected by this recall, the pump can be used for its intended purposes.2) Removal of affected Product:a) remove affected products from service and use non-affected pumpb) If you do not have an alternate pump available for use, MOOG recommended that they weigh the risks and benefits for delivering medications using the affected pump versus the risk of not using any pump for each patientWARNING: Affected pumps should not be used to deliver critical medications where a delay in therapy could result in serious adverse health effects, injury or death; and affected pumps should not be used with the neonatal population. 3)Should they Have an Affected pump which exhibits Error Code 45 or is used in therapeutic situations specified in the warning, please contact MOOG Customer Service at (800) 970-2337, Monday thru Friday, 7:00am to 5:00pm MST to arrange for a loaner pump and the return of your affected pump.If you have any additional questions regarding this recall, please contact MOOG's Customers Service Department at (800) 970-2337. |  | 3407 | 3,407 units | Worldwide distribution: USA and countries including: Australia, Canada, Costa Rica, Croatia, Dubai, France, Germany, Ireland, Italy, Japan, Jordan, Kuwait, Malaysia, New Zealand, Qatar, Saudi Arabia, Singapore, South Africa, Taiwan, Turkey, UAE, United Kingdom. | Terminated | August 06, 2012 | 525 | 18306
Z-1900-2011 | 51497 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Siemens brand ACUSON S2000 Ultrasound System with software versions: |  |  | January 14, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | The product has a software problem in which previous patient measurement data gets associated with another patient's image. | Software change control | The firm, SIEMENS, issued a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that the first image in Exam Review is from the current patient; ensure that the Heart Cycle measurement tool is not selected on the Cale menu when entering Exam Review; be aware of this issue as described in the letter and perform four (4) or less measurements of the same measurement label. Should you have any questions, please call (650) 694-5398. |  |  |  | Worldwide distribution. | Terminated | April 06, 2011 | 812
Z-1828-2011 | 52193 | No | RCE |  |  |  |  | Phoenix x/ray nanotom series. |  |  | March 16, 2009 | 2009 | GE Inspection Technologies, LP | Some of the systems do not have a safety interlock on the service access panel. Although the systems have two independent means for indicating when x-rays are being generated, some systems may have one indicator that is not discernible. Also, some systems lack proper placement of a caution label and others do not have this label affixed. | Radiation Control for Health and Safety Act | The firm has completed all of the following necessary actions and no further actions have been indicated:A Field Service Memo was issued to customers. It described the affected product and described the issue. GE Sensing and Inspection provided and installed retrofit kits, and the systems in question were tested by qualified Field Service Engineers at no charge to customers. GE asked customers to post the notification letter on the affected equipment until retrofit was completed. Questions were to be directed to the GE Service Center at 800-766-6062. |  | 118 | 118 units total | Nationwide Distribution | Terminated | April 17, 2017 | 2954 | 118
Z-1898-2011 | 54684 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Device brochure, labeled in part : BIO -RAD***L70242709 VARIANT TM II TURBO***Hemoglobin A 1C***Cartridge Set, Cartridge Set, Analytical Cartridge, Guard Cartridge Insert***UNITED STATES, Bio-Rad Laboratories, Inc. Hercules, CA 94547***" |  |  | December 08, 2009 | 2009 | Bio-Rad Laboratories Inc | Wrong lot number was found on the Hemoglobin Cartridge Insert . The insert did not match the lot number to the device cartridge or the device cartridge label or on the Update Kit CD. | Other | The firm, Bio-Rad, sent a "MEDICAL DEVICE CORRECTION" letter dated November 29, 2009 to all customers by Federal Express. The letter described the product, problem and action to be taken. The customers were instructed to do the following:1) If their Lot No. LV90170B VARIANT II TURBO Hemoglobin A(10) Cartridge Set includes an insert referencing Lot No. 80080, please destroy the insert and use the new insert enclosed with the letter.2) If their Lot No. LV90170B VARIANT II TURBO Hemoglobin A(10) Cartridge Set includes an insert referencing Lot No. 90170, they may continue using the insert.3) Complete and return the attached CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM via fax to: Bio-Rad CSD Regulatory Affairs Department at (510) 741-3954.If you have any questions, please call 510-724-7000 or contact your regional Bio-Rad office. |  | 51 | 51 kits (Lot # 70292242; 45 kits( Lot # 70292318) | Worldwide distribution: USA including states of: CA, CT, GA, IL, IN, MA, ME, MS, NH, NV, OH, PA, TX, VA and WV; and countries including: France, Hong Kong, Italy, Korea and New Zealand. | Terminated | June 29, 2011 | 568 | 51
Z-1891-2011 | 57972 | No | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Universal F2 Anesthesia Breathing Circuit, REF 5D440-61X-17, Manufacturer: King Systems, Nobleville, IN. |  |  | January 24, 2011 | 2011 | King Systems Corp. | The firm is conducting a subrecall of facetent masks because the masks could have a sticky substance on the inside or outside of the mask that may manifest itself 6-8 months after the manufacture date of the product. The sticky substance could result in inconvenience or in some remote cases cause patient harm. | Other | The firm issued an URGENT PRODUCT RECALL letter dated 1/24/2011. The letter identified the affected product and stated the reason for recall. The letter instructed consignees to cease all sales, to immediately examine their inventory, and to quarantine the product. The letter included a response form that is to be completed and returned. Distributors were instructed to inform their customers of the recall. King Systems Customer Service Department should be contacted at 1-800-642-5464 in order to receive a Return Authorization Number. The letter stated that consignees should return the recalled products to King Systems and the firm will replace the products or provide an acceptable substitute. Customers should contact Customer Service if they have any questions regarding the recall and the letter. |  | 12480 | 416 cases of 30 per case 12480 units | Nationwide Distribution -- NM, GA, TX, FL, IL, VA and LA. | Terminated | August 14, 2012 | 568 | 82905
Z-1899-2011 | 57855 | No | KSD | Heat-sealing device. | Hematology | Hematology | 510(K) Exempt | 2380 MINI Hand-Held Tube Sealer Power Source |  |  | January 31, 2011 | 2011 | Haemonetics Corporation | Battery pack may have exposed wire becoming hot | Equipment maintenance | The firm, Haemonetics, issued a "Field Notification: 2380 Sealer battery pack" letter dated January 28, 2011 to its customers. An international English version of this letter (CL101020-IE) has been released to foreign Distribution countries. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Record the manufacturing date of battery, equipment model, serial number, and photo of the affected battery.2) Report the defect through the Haemonetics PIR system by emailing: sebra2380@haemonetics.com with all of the information listed including the appropriate contact information for their company.3) Recycle or dispose of the nickel-cadmium (NiCad) battery pack in accordance with the local, state, federal, or country specific regulations. Under no circumstances should the battery pack be incinerated as this can cause explosion and personal injury. The operation manual will be updated to provide additional information about the proper installation and removal of the battery. This updated manual will be distributed with new devices upon release.Should you have any questions or concerns regarding this matter, please contact the firm at sebra2380@haemonetics.com. |  | 50031 | 50,031 units | Worldwide distribution: USA and countries including: Australia, Austria, AEU, Belgium, Bangladesh, Canada, China, France, Germany, Hong Kong, India, Indonesia, Italy, Japan, Korea, Malaysia, Mexico, New Zealand, Pakistan, Philippines, Singapore, Sri Lanka, Taiwan and Turkey. | Terminated | March 06, 2012 | 400 | 50031
Z-1903-2011 | 58241 | No | GAB | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | REF 906764, CURVTEK EYE NEEDLE, 12 MM LARGE, STAINLESS STEEL GAMMA STERILIZED, PKG/3, BIOMET SPORTS MEDICINE, 56 EAST BELL DRIVE, P.O. BOX 587, WARSAW, IN 46581 USA, STERILE |  |  | March 10, 2011 | 2011 | Biomet, Inc. | The firm initiated a recall due to the discovery that the devices were not sterilized, as required, prior to distribution. | Process control | The firm, BIOMET, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated March 10, 2011 to all customers. The letter described the product, problem and actions to be taken. The letter noted that -Use of a non-sterile Curvtek Eye Needle may lead to patient infection, possibly requiring a revision procedure. The customers were instructed to immediately locate, discontinue use of the product , remove the product from circulation and return the product to Biomet; carefully follow the instructions on the enclosed FAX BACK RESPONSE FORM; complete and return the response form via fax to 574-372-1683 prior to return of product; use priority carrier for shipment, and please confirm receipt of this notice by calling 800-348-9500, extension 3755 or 3756. The firm also informed physicians that had previously utilized the device to monitor patients for infection and provide treatment as appropriate. Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8am to 5pm. |  | 35 | 35 | Worldwide distribution: USA including states of: IN and KY; and countries of: Canada and The Netherlands. | Terminated | July 27, 2011 | 139 | 35
Z-1901-2011 | 58208 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Prismaflex System, Catalog Number 107493, |  |  | March 14, 2011 | 2011 | Gambro Renal Products, Incorporated | Prismaflex touch screen may become unresponsive during Continuous Renal Replacement Therapy procedure. | Software change control | The firm, GAMBRO Renal Products, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated March 14, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were provided with response procedures to use if the frozen screen problem occurs while using the product. The customers were instructed to closely monitor the return line for potential air and/or clots in the circuit; review the instructions provided in the attached copy of Chapter 5: Troubleshooting, "Softkeys won't work." of the Prismaflex Operators Manual in conjunction with the attached Troubleshooting Checklist; be aware if performing Manual Termination with Blood Return, air detection is not provided, and complete and return the Customer Reply Form via fax to: Regulatory Affairs at 1-866-447-4471 or 303-222-6916, or by email to: GRPRegulatory.Affairs@us.gambro.com. If you have any questions or concerns, please contact for Clinical questions: ICON 800-554-1312 (available 24/7) or your local Gambro Representative; for Technical questions: Gambro Customer Support 800-525-2623, option 2 (Available 6:00am to 5:30pm MT, Monday thru Friday). |  | 1300 | 1300 units | Nationwide distribution. | Terminated | September 01, 2011 | 171 | 1300
Z-1904-2011 | 58099 | No | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | Zoll Auto Pulse Resuscitation System; Zoll Circulation Inc., Sunnyvale, CA 94085 |  |  | January 28, 2008 | 2008 | Zoll Circulation, Inc. | Complaints of batteries discharging prematurely and that during use the Auto Pulse Units would cease operating (providing compressions). | Other | The firm, Zoll Circulation Inc., initiated a Corrective and Preventive Action Report on January 28, 2009. According to the firm, the customers were contacted by telephone calls and were provided with a Battery Maintenance program. For further questions you may call (408) 541-2140 ext.341. |  |  | not provided | Nationwide Distribution -- including, VA, NY, OH, IN, MO, UT, FL, NC, ME, WA, SD, NM, PA, MS, ID, TX, GA, CA, CO, IL, AR, WV, TN, OR, and KS. | Terminated | April 28, 2011 | 1186
Z-1905-2011 | 58112 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA versions 8.0 and 8.1 |  |  | February 18, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly was identified in the ARIA version 8.0 and 8.1, when an image sequence template is added to a field in RT chart and a CIAO (Completely Irradiated Aperture Opening) is calculated, the resulting CIAO display at 4DITC may be incorrect. | Other | Varian Medical Systems sent an Urgent Medical Device Correction letter dated February 18, 2011, to all affected customers via Federal Express.. The letter identified the product, the problem, and the action to be taken by the customer. Varian has made design changes to address the issue in ARIA Radiation Oncology versions 8.2 and later. A corrective action will be undertaken to upgrade all affected items.For questions regarding this recall call US and Canada 1-888-VARIAN5 (888-827-4265.Europe +41 41 749 8844 |  | 645 | 645 | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Australia, Azerbaijan, Bahrain, Bangladesh, Barius, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Guadeloupe, HOng Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Lebanon, Lithania, Malaysia, Maurtama, Morocco, Netherlands, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Yemen. | Terminated | June 19, 2012 | 487 | 645
Z-1915-2011 | 56977 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Penner Manufacturing Transfer Electric Bariatric with Scale, Patient Transfer/Lift System, Cascade Transfer, Models 383500-1 (right entry with scale). The firm name on the label is Penner Manufacturing Inc., Aurora, NE. |  |  | January 19, 2009 | 2009 | Penner Mfg Inc | Incomplete weld on the pillar mounting bracket | Employee error | The firm, Penner, sent an undated letter flagged "Important Safety Inspection: Action Required" via certified mail on January 19, 2009 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect the Penner Patient Transfer for any missing Welds of the Pillar Mounting Bracket; inspect the areas indicated in the enclosed photos to confirm the bracket is welded on all edges to the main frame of the base; complete the inspection, and complete and return the attached WELD INSPECTION RESULTS OF PENNER PATIENT TRANSFER form via fax to 402 694-5844 or mail to Penner Manufacturing Inc, 102 Grant Street, Aurora, NE 68818, Attn: Quality Assurance Manager. Penner stated in the letter, if welds are not complete (no weld), discontinue use and contact the Quality Assurance Manager at Penner Manufacturing immediately at 402-694-5003.If you have any questions, contact the Quality Assurance Manager at 402-694-5003. |  | 20 | 20 lifts | Worldwide distribution: USA and countries including: Canada and Mexico. | Terminated | April 08, 2011 | 809 | 2290475
Z-1929-2011 | 57288 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE |  |  | March 08, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of potential safety issues associated with the CARESCAPE" Monitor B850 when Patient Data Module (PDM), Tram Acquisition Module (Tram), Patient Side Module (E-PSM or E-PSMP), E- MASIMO or E-(P)RE(S)TN modules or Bed-to-bed alarming feature is used.1. When CARESCAPE Monitor B850 is used with the above listed modules, some SpO2 alarms may be delayed or missed.2. Whe | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 25, 2011 to its consignees/customers. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letter described the recall product the problem, Safety Issues, Affected Product details, Product Correction, Safety Instructions and Contact Information. The customers were instructed to follow the safety instructions provided in the letter. GE Healthcare is working on a software correction that will be provided to their customers once it is available.If you have any questions regarding this medical device correction or the identification or affected items please contact Technical Support at 1-800-558-7044, or your local Service Representative. |  | 3817 | 3817 | Worldwide distribution: USA including states of:AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA,WA, WI, and WV; and countries including: JERUSALEM, UNITED KINGDOM, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, QATAR, PORTUGAL, POLAND, NORWAY, NEW ZEALAND, NETHERLAND, MALTA, MALAYSIA, LEBANON, KUWAIT, ITALY, ISRAEL, INDIA, GERMANY, FRANCE, FINLAND, ESTONIA, DENMARK, CZECH REPUBLIC, CHILE , CANADA, BELGIUM, BAHRAIN, AUSTRIA, AUSTRALIA. | Terminated | June 05, 2012 | 455 | 3817
Z-1925-2011 | 57866 | No | NDR | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Inhibin A ELISA kit, Part Number: DSL-10-28100-4 |  |  | December 21, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because the lots of Inhibin A identified above may contain microbial contamination in the conjugate diluent bottlesincluded in the kit. | Pending | Beckman Coulter sent an Urgent Product Corrective Action letter dated December 21, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product and discard all remaining inventory. In addition customers were requested to share the information with their laboratory staff and retain the notification as part of their laboratoryQuality System documentation. Customers were instructed to complete and return the enclosed response form within 10 days soBeckman could be assured that customers received this important notification.For product replacement:1. In the United States, please contact Customer Service at 1-800-526-3821, option 1.2. In Canada, please contact Customer Service at 1-800-463-7828.3. Outside of the United States and Canada, contact your local Beckman Coulter representative.Customers who needed assistance or had any questions regarding this notification, were instructed to contact Technical Support at 1-800-854-3633 in the United States and Canada. For those customers outside the United States and Canada, instructions were given to contact their local Beckman Coulter representative. |  | 740 | 740 total units for both products | Worldwide Distribution - USA including NY, IN, PA, NM, OH, TX, AZ, NC, TN, MI and CT and the countries Denmark, Germany, Spain, Turkey, United Arab Emirates, United Kingdom. | Terminated | August 31, 2012 | 619 | 740
Z-1907-2011 | 58006 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | 8V3c transducer used with the ACUSON SC2000 ultrasound system. 8V3c transducer model number 10433823; ACUSON SC2000 model number 10433816. |  |  | February 04, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | When using the 8V3c transducer in color mode, the derated Ispta 3 (Spatial Peak-Temporal Average Intensity) may exceed the limit stated in the instructions for use. | Software design | A Customer Safety Advisory Notification was initiated on 2/4/2011 and sent via certified mail. The letter identified the affected product and described when the problem occurs along with the associated potential risk. The firm recommends that users not use the 8V3c until the software issue is fixed. The firm is working to release a software update and customers' local Customer Service Engineer will install new software on their system when it is available. Customers should direct their questions to their local service support person. The information in the letter should be shared with all personnel within the organization who need to be aware. |  | 386 | 386 units | Worldwide Distribution. | Terminated | October 04, 2011 | 242 | 386
Z-1909-2011 | 58093 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | SPIDER Surgical Platform |  |  | February 25, 2011 | 2011 | Transenterix, Inc | Plastic shards may break off and fall into the abdominal space which may require retrieval during surgery. | Other | TransEnterix sent an Urgent Device Recall letter dated February 25, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the lot numbers affected, complete the attached inventory reconciliation form and return the form, along with any affected product toTransEnterix using the enclosed shipping label. Customers were asked to return the inventory form even if they had none of the lot numbers affected in their inventory.Their sales representative and customer service will work with them to replace the product in full at not cost to them.For questions regarding this recall contact TransEnterix Customer Care at 1-888-879-4111. |  | 305 | 305 units | Nationwide Distribution including AZ, FL, NC, NY, OH, OR, SC, TX, VA and WA | Terminated | November 07, 2011 | 255 | 397
Z-1928-2011 | 58014 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-G5 Ventilator with software version V2.00x and Neonatal Option or V2.01x and Neonatal software. Hamilton-G5: catalog number 159001; |  |  | February 16, 2011 | 2011 | Hamilton Medical Inc | The adaptive volume/pressure regulation can be automatically and permanently reduced to PEEP + 3 mbar when a closed suctioning maneuver is performed on the patient and may be undetected by the operator. | Software design | A Medical device Safety Alert and Corrective Action along with a cover letter, was sent to all locations on February 16, 2011. The letter identified the affected product along with the reason for recall. The letter states that operators deactivate the "high tidal volume" alarm as described in appendix K.4 of the operator's manual until the ventilator is upgraded with a corrected software. Distributors will upgrade the affected devices with the corrected software as soon as possible. Questions can be directed to 800-426-6331, ext 215. |  | 97 | 97 units | Nationwide Distribution -- MA, FL, GA, NC, MI, NC, MO, TX, and WV. | Terminated | March 21, 2012 | 399 | 97
Z-1930-2011 | 58275 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650 |  |  | March 23, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue associated with the CARESCAPE Monitor B650 &amp; B850 Cardiac Output measurement using Patient Data Module (PDM), and activation of ECG Leads Off alarm in specific cases.1. When the CARESCAPE Monitor is used to measure Cardiac Output values with the Patient Data Module (PDM) or TRAM Module, Cardiac Output values may be calculated incorrectl | Software design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated March 23, 2011 to its consignees/customers. The letters were addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letters described the Safety Issues, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. The customers were instructed to follow the safety instructions documented in the letter. GE Healthcare is working on a software correction that will be provided to their customers once it is available and will contact the customer to schedule the software update.If you have any questions regarding this medical device correction or the identification of affected items, please contact Technical Support at 1-800-558-7044, or your local Service Representative. |  | 1443 | 1443 | AL, AZ, AR, CA, CO, CT, FL,GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA,PR, SC, TN,TX, UT, VA, WA, WV, WI.UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, SWITZERLAND, SWEDEN SLOVENIA, SLOVAKIA, SPAIN, SINGAPORE, SAUDI ARABIA, QATAR, PORTUGAL POLAND, NORWAY, NORTH KOREA, NEW ZEALAND, NETHERLANDS, MALTA, MALAYSIA, LITHUANIA, LEBANON, KUWAIT, ITALY, ISRAEL, IRELAND, INDIA, ICELAND, HONG KONG, GERMANY, FRANCE, FINLAND, DENMARK, ESTONIA, CZECH REPUBLIC, BELGIUM BAHRAIN, COLOMBIA, CHILE CANADA, BRAZIL, AUSTRIA, AUSTRALIA. | Terminated | February 04, 2012 | 318 | 8345
Z-1927-2011 | 58143 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Inflow/Outflow Cannula; |  |  | March 07, 2011 | 2011 | Stryker Endoscopy | Insufficient cleaning instruction were included for the Inflow/outflow Cannula in the user manual. | Other | The firm, Stryker, issued an "URGENT: DEVICE CORRECTION" memo dated March 8, 2011 to all customers. The memo described the product, problem and actions to be taken. The customers were instructed to discard any old Instructions for Use (1000-400-816 rev. A or B) for their Inflow/Outflow Cannual; download the new IFU (1000-400-816 rev. C) from the URL link: http://stryker.com/InflowOutflowCannual; ensure they follow the new cleaning instructions provided in the IFU and complete and return the enclosed "acknowledgement of receipt" form and fax to (408) 754-8378 or scan in and email to cannula@stryker.com. Stryker stated in the memo that "It is important to send the self addressed confirmation forms back to Stryker RAQA". If you have any questions regarding this letter, please contact us by phone at 1-800-624-4422 or via email at cannula@stryker.com. |  | 2618 | 2618 devices | Worldwide distribution: USA including states of: AL, ALASKA, AK, AR, CA, CO, CT, DC, DE, FL, GA, ID IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, NY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY; and countries of: Australia, Canada, China, France, Germany, India, Japan, Korea, Latin America, Mexico, Netherlands, Poland, Switzerland and United Kingdom. | Terminated | October 04, 2011 | 211 | 2618
Z-1924-2011 | 57141 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Quest MPS and MPS 2 Consoles, Quest Medical, Inc., An Atrion company. Single software controlled device that incorporates a pump, a heater/cooler, temperature monitoring, pressure monitoring, a heat exchanger, an arrest agent pump and an additive pump. The MPS console is used with the MPS sterile disposables with a built in heat exchanger. The device is used with on-pump or off-pump cardiac surgery. |  |  | July 29, 2010 | 2010 | Quest Medical, Inc | User update regarding installation of external heater-Cooler units with MPS and MPS2 consoles. | Other | Quest Medical Incorporated sent a "User Update" letter in July 2010 to consignees providing information concerning use of external heater-cooler untis with their device - Myocardial Protection System ( MPS) consoles. If you are installing an external heater-water cooler unit or have questions about this user update, please contact the Field Service department at Quest Medical, Inc. The number is 800.627.0026 or 1 972.390.9800. |  | 771 | 771 consoles | Nationwide, Canada, Saudi Arabia, Kuwait, Japan, and Italy | Terminated | April 26, 2011 | 271 | 771
Z-1932-2011 | 54483 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program, 1600 and 2000 tests Cartridge Sets. Catalog Numbers: 270-2417 and 270-2415; Distributed and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CA |  |  | December 01, 2009 | 2009 | Bio-Rad Laboratories Inc | Some customers have reported the occurrence of some ramping baselines on the chromatograms. The ramping baseline can affect hemoglobin test results. | Software change control | Bio Rad Laboratories, Inc. sent an Urgent Medical Device Correction letter dated December 1, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediatey examine their records to identify all customers who received one of the affected lots.Collect a Customer Medical Device Correction Response Form from every customer.Once forms were received from all applicable customers, complete the following Medical Device Correction Response Form and fax it to Bio-Rad CSD Regulatory Affairs at 510-741-3954 or pdf to CSD Recall Coordinator (excluding USSD).For questions regarding this recall call 510-724-7000. |  | 1099 | 1099 | Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Taiwan, Singapore, India, China, S. Korea, Hong Kong, France, Sweden, United Kingdom, Italy, Hungary, Germany, Brazil, Mexico and Colombia. | Terminated | August 30, 2011 | 637 | 1099
Z-1943-2011 | 57189 | No | DTY | Cardiopulmonary bypass in-line blood gas sensor. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Systems Corporations, CDI 101 Hematocrit/Oxygen Saturation Monitoring System. In-line blood-gas sensor used in cardiopulmonary bypass. |  |  | November 01, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | Inaccurate result. Possibility for displayed values for potassium, CO2, hematocrit, O2 saturation, pH, hemoglobin and other measured blood parameters to be inaccurate. | Device Design | On 11/1/2010, Terumo sent "URGENT MEDICAL DEVICE CORRECTION" notices to Chief Perfusionist, Department of Cardiovascular Surgery or Director of Operating Room Services. The letter informed the consignees about the incidents of the CDI 100 and 101 displaying inaccurate values. The letter states that when a displayed value is significantly different from expectations based on the clinical situation, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operators manual. |  | 468 | 468 monitors | US, India, Colombia, Taiwan, Venezuela, Vietnam, Puerto Rico, Malaysia, Sri Lanka, Brazil, Indonesia, Pakistan, Chile, Germany, Hong Kong, Japan, AE, Singapore, Austria, Belgium and Thailand. | Terminated | April 17, 2012 | 533 | 1500
Z-1941-2011 | 58098 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Transmitter Sensor Cables, Part Numbers: 1002008, 1004587, 1007914-NAV, GE Healthcare Surgery, Salt Lake City, UT 84116. |  |  | February 24, 2011 | 2011 | GE OEC Medical Systems, Inc | Sensor cables may suffer material degradation when exposed to certain sterilization procedures. | Component design/selection | GE OEC Medical Systems, Inc sent an Urgent Recall Notice dated March 1, 2011, by certified letter on February 24, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised if their Transmitter Sensor Cables or Receiver Sensor Cables have been exposed to any STERRAD sterilizer other than STERRAD 1005, discontinue use of these sterilizers for this product. The Transmitter Sensor Cables should be inspected for signs of material degradation.Further information is available at 800-874-7378. |  | 708 | 708 cables | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NY, OH, OR, PA, RI, SC, SC, TN, TX, UT, VA, VT, WA, WI and WV and the countries of Austria, Canada, China, Denmark, England, France, Germany, Hong Kong, India, Ireland, Italy, Japan, San Salvador, Slovakia, Spain, and Switzerland. | Terminated | September 09, 2011 | 197 | 708
Z-1933-2011 | 57953 | No | FNK | Hydraulic adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Rolko Adjustable (retractable) Strap and Handle, type TGN-R100GR; a component of a Lifting Pole, an accessory to long term care hospital beds such as ArjoHuntleigh's Minuet 2 Community bed and Enterprise series hospital beds; Distributed by ArjoHuntleigh, a Branch of Arjo Limited Med AB, Wednesbury, United Kingdom |  |  | February 24, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | Part of the housing of the strap retracting mechanism built into the handle may break while the patient is using the handle and strap. | Other | ArjoHuntleigh sent an Urgent Field Safety Notice to all affected end users via UPS Ground on February 24, 2011. The notice informed the accounts of the potential for the strap retracting mechanism to break during use. The accounts were requested to check their inventory for the presence of adjustable straps and handles procured from ArjoHuntleigh, note if they have a red button on the retractor housing and the affected production dates as explained in the photographs in Figures 1 and 2 in the letter. Any straps and handles found as identified in Figures 1 and 2 should be removed from use and returned to ArjoHuntleigh. Any straps and handles with a yellow button design as shown in Figure 3 are not affected and can continue being used. The accounts were requested to complete and fax or mail the enclosed Customer Response Form to Arjo at 630-576-5012, indicating the number of strap and handles checked and how many were removed from use. Questions were directed to the ArjoHuntleigh Regulatory Department at 1-630-785-4885. |  | 90 | 90 units | Nationwide Distribution including California, Georgia, Hawaii, Michigan, South Carolina and Texas | Terminated | June 14, 2011 | 110 | 90
Z-1940-2011 | 57952 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC 660i, 680i, 860i, 880i |  |  | January 26, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated due to Beckman Coulter receiving a report of results that were associated with incorrect patientidentifications when two sample racks advanced into the sample presentation area in one push. The firstrack of the two was not scanned. This is referred to as a "pushed pair." If a pushed pair event occurs, results may be associated with the incorrect patient identification. | Pending | Beckman Coulter sent an Urgent: Field Safety Notice Product Correction letter dated January 26, 2011, via US Postal Service to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:1. To prevent a pushed pair, be sure your sample racks are free of debris prior to use. Do not applyadditional labels to the sample racks. Extra labels can obscure the edges of a sample rack, orobstruct the gap between two racks. Be sure sample labels do not extend beyond the edges of asample rack;2. If a pushed pair occurs on the UniCel DxI system, the system generates a yellow Event Log message- Clear the presentation tray and press ROUTINE button. This Event Log message is associatedwith a variety of common Sample Presentation Unit (SPU) errors. For that reason, you must reviewthe Event Log details to confirm the pushed pair event has not occurred;3. Detach the Appendix included with this notification and store it near your UniCel DxI system. Use theAppendix as a reference if you suspect a pushed pair event has occurred on your system.Beckman Coulter also requested their customers share the recall information with their laboratory staff, and retain pages 1 and 2 of this notification as part of your laboratory Quality System documentation.For questions regarding ththe recall notification, customers were instructed to contact Beckman Coulter Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support, or call1-800-854-3633 in the United States or Canada. Outside the United States and Canada, contact theirlocal Beckman Coulter Representative. |  | 2639 | 2639 units (892 in US) | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Angola, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Republic of Korea, Kuwait, Liechtenstein, Malaysia, Mayotte, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Philippines, Poland, Puerto Rico, Romania, Russian Federation, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Terminated | June 04, 2012 | 495 | 2639
Z-1945-2011 | 58163 | No | CEC | Free thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Free T4 Calibrator, Part Number: 33885, Lot Numbers: 014769 |  |  | November 11, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because the stability claim for Access Free T4 calibrator lot 014769 is not met. The affected T4 calibrator lot can produce quality control and patient results that are erroneously low (by greater than 2 standard deviations). An erroneously low patient result may lead to misdiagnosis and/or delayed treatment. | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" (PCA) letter dated November 29, 2010 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: ACTION/RESOLUTION:" Discontinue use of Access Free T4 calibrator lot 014769." Review historical quality control performance and determine whether any reported patient results were obtained while using a faulty calibration. If results were affected, you will see a shift downward in control recovery that produces out-of-range low QC results." If patient results were obtained with Access Free T4 calibrator lot number 014769, it is possible they are erroneously low. If any patient results were reported in the period between the time the erroneous calibration with the affected lot was performed and when a different calibrator lot was used, the laboratory should consider advising physicians to whom results were reported so they can evaluate whether diagnosis or treatment decisions should be revisited." No further orders of Access Free T4 Calibrator lot number 014769 will be shipped. Alternate lots will be provided for current and future replenishment orders or requests.Consignees/customers were asked to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation; and complete and return the enclosed RESPONSE FORM via fax to Beckman Coulter at 786-639-4000. For product replacement:" In the United States, please contact Client Services at 1-800-526-3821, option 1." In Canada, please contact Customer Service at 1-800-463-7828." Outside of the United States and Canada, contact local Beckman Coulter representative.Questions and answers regarding this issue were enclosed for consignee/customer reference. Additional questions, were directed to the Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customer |  | 459 | 459 units | Worldwide distribution: USA and countries including: Argentina, Belgium, Bosnia and Herzegovina, Canada, China, Colombia, Czech Republic, French Polynesia, Hong Kong, Ireland, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Oman, Panama, Philippines, Russian Federation, South Africa, Switzerland, Taiwan, Thailand, Tunisia, United Kingdom, Venezuela, and Zambia. | Terminated | July 11, 2012 | 608 | 459
Z-1961-2011 | 52928 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Auxiliary 2 MHz Continuous Wave (CW) transducer, model number 8251046, when used with ACUSON Sequoia ultrasound systems at SW revision 12.X. Manufactured by Siemens Medical Solutions USA, Inc., Ultrasound Business Unit, Mountain View, CA. |  |  | March 21, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Overheating. A software bug may cause the transducer to overheat when used with certain transducer models. The transducer can become noticeably hot within a few minutes, which may cause patient discomfort or burns. | Software design | The firm, SIEMENS, sent a "Customer Safety Advisory Notification" to its customers on 3/21/2009. The notification described the product, problem and actions to be taken. The customers were instructed to disconnect any transducers that may be in an active transducer port and chose another imaging transducer prior to use of the Aux CW transducer. If they do not have another transducer to select other than those listed in the notification, they should close the study in progress and start a new study before using the Aux CW transducer-this procedure should be followed for each new patient and /or exam type. In addition, the customers were instructed to pass this notice on to all those within their organization who need to be aware of this issue until the corrective action is completed. SIEMENS will release a software update that will correct the issue. The customers local service engineer will install the updates on their system as soon as it becomes available.Should you have any questions regarding this notification, you may contact your local service support person and/or call (650) 969-9112. |  | 2591 | 2591 units | Worldwide distribution. | Terminated | March 09, 2012 | 1084 | 2591
Z-1841-2011 | 58076 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Ellex Ultra Q LQP3106-U ophthalmic laser |  |  | February 15, 2011 | 2011 | Ellex Medical Pty Ltd | Unintended laser emission while the user was moving the joystick for targeting. | Process control | Ellex notified all distributors that have received the affected devices per Global Customer Service Bulletin GSB-11 -01a, dated 25 February 2011. |  | 233 | 233 units | worldwide distribution | Terminated | October 02, 2012 | 595 | 2335
Z-1958-2011 | 58122 | No | NRU | Dental chair and accessories. | Dental | Dental | 510(K) Exempt | Dental chairs distributed under several brand names such as DCI Equipment; Marus MaxStar; Marus NuStar; and Pelton &amp; Crane. |  |  | March 01, 2011 | 2011 | Dental Equipment LLC | Link arm assembly that mounts the rear mounted accessories to the dental chair may potentially break and fall off the dental chair. | Process control | On 03/08/11, Dental Equipment LLC. started sending out the Medical Device Recall notification letter via USPS 1st class mail to US consignees and Federal Express mail to foreign consignees.The firm also sent a recall notification letter to consignees who only purchased dental accessories such as Rear Mounted Cuspidor with Coverset, Arm Articulating Rear Mounted Cuspidor - Marus Gray, Maxstar Orbit Cuspidor utility W/ Porcelain bowl, and zinc plated bolts.The letter identified the affected dental chair as Marus brand with the following model numbers: DC1490, DC1690, DC1700 and DC1702.The firm's corrective action including the replacement of the affected link arm assembly with a new link arm assembly. The firm's field representative will contact each consignee for an appointment to replace the affected link arm assembly.Consignees can contact the Dental Equipment LLC at 503-537-3617 for any questions. |  | 66 | 55 chairs and 11 replacement bolts distributed to US and foreign consignees outside the US. | Worldwide distribution to USA, Australia, Canada, Ireland, South Africa, and Vietnam. | Terminated | September 21, 2011 | 204 | 66
Z-1959-2011 | 58110 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Aristos FX and FX Plus System |  |  | February 17, 2011 | 2011 | Siemens Medical Solutions USA, Inc | During product monitoring, firm discovered a possible hazard to patients when using the detector alpha rotation of all Axiom Aristos FX and FX Plus systems with serial numbers up to and including 1216. When using the detector alpha rotation, the detector may descend unexpectedly during movement or in vertical or table mode. If this occurs in vertical modes, the possibility exists that the detect | Other | The firm, SIEMENS, sent a "Safety Advisory Notice" dated February 15, 2011 to the affected customers via Update Instructions XP016/10/S on February 17, 2011. The notice described the product, problem and actions to be taken. The customers were instructed to immediately instruct their personnel accordingly; ensure that the safety advisory is placed in the systems's instructions for use; if they have sold this device/equipment and it is no longer in their possession, forward this safety notice to the new owner of the device/equipment; inform SIEMENS about the new owner of the device/equipment, and complete and return the Acknowledgement of receipt form. SIEMENS is preparing a modification of AXIOM Aristos FX and FX Plus that will resolve this potential malfunction. The field modification will be available from February 2011 and is expected to be carried out May 31, 2011. If you have any questions, call 09131-18-9908 or +1-888-826-9702. |  | 101 | 101 | Nationwide Distribution: USA including states of: AZ, AK, CA, CO, FL, GA. IL, IN, IA, KS, KY, ME, MI, MN, MS, MO, NE, NV, NJ, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, and WI. | Terminated | July 16, 2013 | 880 | 101
Z-1957-2011 | 58072 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Varis Vision - RT Chart version 7.3.10. For radiation therapy. |  |  | February 07, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue with the Varis Vision software where treatment can occur without the dynamic Multileaf Collimator component and may deliver a higher than expected dose of radiation. | Software design | Varian Medical Systems Urgent Medical Device Correction letters were sent to all affected customers on February 7, 2011. |  | 421 | 421 | Worldwide Distribution. | Terminated | December 23, 2011 | 319 | 421
Z-1956-2011 | 57990 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | Coulter GEN's |  |  | January 31, 2011 | 2011 | Beckman Coulter Inc. | This recall was initiated because Beckman Coulter identified a problem with some date fields that are transmitted to the Laboratory Information System (LIS). Starting Jan 1, 2011, the date fields will transmit as a blank field to the LIS if the year is greaterthan 2010. The following dates are affected: Run date, Control expiration date, Date of birth(For individuals born in 2011 or later), a | Software design | The firm, Beckman Coulter, sent an "URGENT: PRODUCTION CORRECTION" letter dated January 31, 2011 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Coulter LH 750 Analyzer, Coulter LH 780 Analyzer, Coulter LH 500 Series System, and Coulter GEN S System. The letter describes the product, problem and actions to be taken. The customers were instructed to work with their US vendor to assess how this affects their laboratory. The actions they choose are dependent on the configuration of their LIS.ACTION FOR ISSUE 1:Use Sample Drawn Date and Date of Birth information available at your LIS in lieu of the analyzer transmitted fields BIRTH and UDATE.Use the current date when the results are received at the LIS in lieu of the transmitted Run Date [DATE]. Print results from the Workstation to keep a copy of the analysis date on the analyzer, if necessary.Use the control expiration date from the analyzer display or printout.ACTION FOR ISSUE 2:If you use the LIS 1G1 transmission format, use the General Information group field names in place of the Retic Parameter group field names to identify each reticulocyte run (see Table in Issue 2 section).Check with your LIS vendor to determine if it is possible to switch to the LH 750 Workstation transmission format. Format requirements are LIS driver specific.For patients with Dates of Birth in 2011 for both Issue 1 and Issue 2, results should be reviewed to determine if decision rules or age-related limits sets were triggered inappropriately.RESOLUTIONThese issues will be corrected in a future software release for the LH Series products. The GENÂżÂżS product has been discontinued; no software revisions will be forthcoming for the GENÂżÂżS product line. Customers were also instructed to complete and return the PRODUCT CORRECTION RESPONSE FORM via fax to 786-639-7500/7501/7502/7503/7504; mail to Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 310B06, 11 |  | 9419 | 9419 (total units distributed in U.S. &amp; foreign countries) | Worldwide distribution: USA (nationwide) and countries including: Algeria, Andorra, Angola, Antigua and Barbuda, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bermuda, Bulgaria, Burundi, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Ethiopia, France, Germany, Greece, Guadeloupe, Guam, Guatemala, Hong Kong, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macao, Malaysia, Mayotte, Mexico, Monaco, Morrocco, Myanmar, Namibia, Netherlands, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and Virgin Islands (US). | Terminated | September 25, 2012 | 603 | 9419
Z-1934-2011 | 57450 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid Waste Bottle |  |  | October 05, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter confirmed Access waste bottles that ruptured and expelled liquid waste, exposing the operator to the system liquid waste. Over time, waste bottles can degrade, become brittle and potentially crack. | Pending | The recall communication was initiated on October 5, 2010 with Beckman Coulter forwarding an Urgent: Product Corrective Action (PCA) letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Access immunoassay Systems, UniCel DxC 600i Access, and Immunoassay Systems, Synchron LxI 725 Access Clinical Systems. The letter explains the problem identified, the impact, and actions required to be taken by customers. The letter instructs customers to inspect the waste bottle for signs of aging and deterioration. New bottles are opaque white in color.Signs of deterioration include:o Any color change to the bottle (yellow, orange or pink)o Crackingo Crazing (spider web cracking)Any defect that is not removed by routine cleaning is a sign of deterioration. It is important to routinely inspect the area nearest the handle and along all seams of the bottle. If upon inspection consignees detect any issues with their waste bottles, they were asked to contact Beckman Coulter Customer Service using the enclosed Replacement Order Form to obtain a new waste bottle, and keep this information should they require a replacement bottle in the future.Consignees were also asked to share this information with laboratory staff and retain the notification as part of laboratory Quality System documentation. They should also complete and return the enclosed Product Corrective Action Response Form within 10 days so Beckman Coulter can be certain they received this important notification.Beckman Coulter offers a service kit which may be used to replace the waste bottle and direct liquid waste to a floor drain. Beckman Coulter requires a consultation with consignees technical support representative to determine if their system can be fitted with a floor drain kit. If assistance is needed or consignees have any questions regarding this notification, they were asked to contact Technical Support at 1-800-854-3633 i |  | 7700 | 7700 units total (3178 in US) | Worldwide Distribution -- United States, Afghanistan, Algeria, Angola, Antigua and Barbuda, Australia, Austria, Azerbaijan, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, China, Colombia, Congo, Costa Rica, Croatia, Czech Republic, Djibouti, Dominican Republic, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, French Polynesia, Georgia, Germany, Ghana, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Mexico, Morocco, Myanmar, Namibia, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestinian Territory, Panama, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam, Yemen, Zambia, and Zimbabwe. | Terminated | July 25, 2012 | 659 | 7700
Z-1952-2011 | 57830 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ArthroCare Opus SpeedLock Knotless Fixation Device; Catalog Number: OM-7500. |  |  | September 30, 2010 | 2010 | ArthroCare Corporation | During surgical procedure the product has the potential for separation of the plug guide. Separation of the plug guide may compromise the ability of the inserter to fully drive the implant completely into the bone, which may also compromise the ability to tension the suture properly. | Process control | On 9/30/2010, Arthrocare notified customers via letters explaining the reason for the recall and requesting the product be returned. |  | 170 | 170 units | Product was released for distribution in US. | Terminated | August 01, 2011 | 305 | 170
Z-1950-2011 | 57815 | No | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medstorm Adult Radiolucent Leads-Out Electrode - Physio-Control, Part 16242. |  |  | December 05, 2008 | 2008 | Heart Sync, Inc | The "leads out" connector had bent pins which would not allow connection to the Medtronic Physio Control Defibrillator cable during a product demonstration. | Equipment maintenance | The firm sent a recall notice to their distributors on 12/3/2008 asking them to quarantine product and return it to HeartSync. The letter stated that customers' current inventory would be exchanged for new product. Customers were to call 1-800-828-4681 to provide the necessary information. |  | 21000 | 2100 (210 cases of 10 each) | Nationwide Distribution | Terminated | April 13, 2011 | 859 | 21500
Z-1947-2011 | 57385 | No | HBX | Cranioplasty material forming instrument. | Neurology | Neurology | 510(K) Exempt | DePuy Mitek FMS Irrigation Intermediary Tube Without One-way Valve Arthroscopy tubing |  |  | December 03, 2010 | 2010 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Product sterility may be compromised due to breach in sterile packaging | Packaging process control | Depuy Mitek sent an Urgent Voluntary Product Recall letter dated December 2, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to isolate all inventory of the affected product and call 1-800-382-4682, PROMPT 3, for return good authorization and replacement product. Customers were instructed to return all identified affected product to:DePuy MitekATTN: Recall Coordinator50 Scotland BoulevardBridgewater, MA 02324Record their assigned return material authorization (RMA) number on the outer shipping box.Use the Prepaid Federal Express shipping label for the return of the product.Include a photo copy of the completed business reply form.For questions or concerns call 1-800-382-4682, PROMPT 3. |  | 7344 | 7344 units (372 boxes x 24 pieces) | Worldwide Distribution - USA (nationwide) and the coutnries of United Arab Emirates, Australia, Belgium, Switzerland, Czech, Denmark, France, UK, Italy, Netherlands, Poland, and Slovakia. | Terminated | April 02, 2013 | 851 | 149720
Z-1967-2011 | 58304 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ExactaMix EVA Container, 3000mL, REF: 741, Sterile R, Baxa Corporation, Englewood, CO 80112. |  |  | March 08, 2011 | 2011 | Baxa Corporation | TPN bags may leak during filling operation, due to bag manufacturing error. | Nonconforming Material/Component | The firm, BAXA, sent a "ExactaMix Bag Product Notice" dated March 8, 2011 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to do the following:If they do not intend to use the manual addition port, verify that the cap is firmly seated onto the port by gently tugging on the cap. If they do intend to use the manual addition port, they should snap off the additive port cap prior to compounding the bag.If they encounter a faulty additive port, contact Baxa Technical Support to receive credit and/or replacement on defective product only.To return affected bags, follow the instructions provided in the notice if the bags were purchased directly from Baxa or purchased from wholesaler, distributor or other 3rd party vendor. Additional copies of the return form can be retrieved by going to www.baxa.com/ebug OR by calling Baxa Technical Support at 800-678-2292.If you have any questions or concerns, contact the Sr. Vice President, Quality and Regulatory at (303) 400-2838 or Baxa Technical Support at 800-678-2292. |  | 336650 | 336,650 bags | Worldwide distribution: USA (nationwide) and countries including: Argentina, Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, Hong Kong, Ireland, Japan, Malaysia, Mexico, New Zealand, Panama, Puerto Rico, Singapore, South Korea, Taiwan, UK, Uruguay, Venezuela. | Terminated | March 23, 2012 | 381 | 6561200
Z-1960-2011 | 58290 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | STERIS SYSTEM 1E Liquid Chemical Sterilant Processing System |  |  | March 14, 2011 | 2011 | Steris Corporation Hopkins Facility | The printout, display and Operator Manual for the STERIS System 1E Liquid Chemical Sterilant Processing System are subject to recall/correction by the firm because in the event of an incomplete or aborted cycle, the language of the warning messages used in the printout, display, and Operator Manual omits the qualifying term: *liquid chemically* (not **sterilized), when alerting the user to inco | Other | The firm, STERIS, Issued an 'Urgent Field Correction Notice' dated March 14, 2011 addressed to their Hospital Administrator, or Manager, and Risk Manager customers on March 18, 2011. The notification letter describes the product subject to field correction and the problem that exists with the language omission present in the printout display and Operator Manual messages which warn of an incomplete or aborted processing cycle. The notification goes on to state that there is no health hazard involved in this error, as the users are still made aware that the processing cycle has not been successfully completed. The notification goes on to remind the user/customers that in case of any incomplete, cancelled or aborted cycle, the users are instructed to process devices in the cycle following the directions contained in the Operator Manual. Lastly, the notification letter informs the user/customers that the company will be installing new software on the recalled systems in order to revise the language on the cycle printout and the display screen warning of a disrupted cycle. It also states that a STERIS Service Representative will be contacting each facility in order to schedule installation of the new software. The notification also provides the customers with a 1-800 telephone number for STERIS Field Service Dispatch, to be used if the customers have questions or want to request additional information. It also provides a local telephone no. 1-440-392-7455, for obtaining additional information from the firm's Director, Div. of Low Temperature Sterilization at: 1-440-392-7455. |  | 106 | 106 units | Nationwide distribution: USA including states of :CA, DE, FL, GA, IA, ID, IL, KS, MA, MO, MT, NE, NJ, NM, NY, OH, OR, SD, TX, VA, WA, WI, and WV. | Terminated | March 13, 2013 | 730 | 106
Z-1949-2011 | 58169 | No | KLI | Total thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Total T4 Reagent Kit, Part Number: 33800, |  |  | March 10, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the concentration of total T4 in samples from pregnant women is erroneously low (20%)when measured using the Access Total T4 assay. Erroneously low results which are actually high (outside of the normal reference range)could mask hyperthyroidism. Failure to diagnose or treat hyperthyroid disease associatedwith pregnancy could result in serious, permanent harm | Other | Beckman Coulter sent an Urgent Product Correction letter dated March 20, 2011, to all affected consignees. The letter identified the product, the problem and the action to be taken by the consignee. Consignees were issued the following instructions:(1) If you suspect pregnant patient results were evaluated at your laboratory using only Total T4 to report thyroid status, consider reviewing Total T4 results reported in the last 9 months. Consult with your Laboratory Medical Director and requesting physician toresolve any concerns you may have regarding previously reported patient results during this time period.(2) Do not use the Access Total T4 assay as the only marker for evaluating pregnant patients for thyroid disorders.(3) Please share this information with your laboratory staff and retain this notification as part ofyour laboratory Quality System documentation. If you have forwarded Access Total T4 reagent kits (Part Number 33800) to another laboratory, please provide a copy of this letter to them.(4) Please complete and return the enclosed Response Form within 10 days so we are assured you have received this important communication.For questions regarding this notice, consignees were instructed to contact Beckman Coulter Customer Support Center.(1) Via our website, http://www.beckmancoulter.com/customersupport/support(2) Via phone, call 1-800-854-3633 in the United States and Canada(3) Outside the United States and Canada please contact your local Beckman Coulter |  | 151740 | 151,740 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina Armenia, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Djibouti, Egypt, El Salvador, Ethiopia, France, Germany, Greece, Guatemala, Hong Kong, India, Israel, Italy, Japan, Jordan, Kazakhstan, Republic of Korea, Kuwait, Lebanon, Libyan Arab Jamahiriya, Liechtenstein, Malaysia, Mexico, Republic of Moldova, Morocco, Myanmar, Netherlands, New Zealand, Nigeria, Pakistan, Panama, Philippines, Poland, Puerto Rico, Reunion, Romania, Russian Federation, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam, Yemen, and Zambia | Terminated | August 10, 2012 | 519 | 151740
Z-1906-2011 | 58175 | No | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Penumbra brand Penumbra Coil System, Penumbra Coil 400, Product is manufactured and distributed by Penumbra Inc., Alameda, CA |  |  | March 04, 2011 | 2011 | Penumbra Inc. | The diameter of a component of the device is out of tolerance specification; the Delivery Pusher pull wire which secures the coil in place can prolapse and allow premature detachment of the coil. | Process change control | Penumbra sent an Urgent Voluntary Field Removal Notice dated March 3, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to immediately inspect all stock of Penumbra Coil 400 and identify product from the Lots.Remove product from these Lots from inventory and place it in quarantine for return to Penumbra Inc. (Do not use product from these Lots).Complete and return the attached product identification/ return form with in 5 business days via fax to +1 510-217-6414.Penumbra employees will contact customer directly to arrange return and replacement of product.For questions call Regulatory Affairs at 510-748-3224 or e-mail RAQA@prenumbrainc.com. |  | 2082 | 2082 | Worldwide Distribution - USA (nationwide) and the countries in the European Union (Denmark, Germany, Italy, Netherlands, Slovakia, Slovenia, Spain, United Kingdom and Switzerland) and South Africa. | Terminated | June 06, 2011 | 94 | 2082
Z-1970-2011 | 53195 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | LANTIS oncology Information System, an accessory to the Linear Accelerator, as part of the COHERENCE RT Therapist, but also cleared under the ONCOR Avante-Garde and Expression, as well as the syngo RT Therapist system, as part of the ARTISTE And PRIMUS system. Manufactured by IMPAC Medical Inc., Sunnyvale, CA, device distributed by Siemens Medical Solutions Oncology Care System, Concord, CA |  |  | August 21, 2009 | 2009 | Siemens Medical Solutions USA, Inc | When editing the note type of a previously entered note, then reverting to the original type, data may be lost. | Software design | Siemens sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to configure their system to lock all new notes automatically.To automatically lock new notes, open eChart Navigator, right mouse click in the notes section, select "configuration" and select "Automatically Lock New Notes."If the customer chooses not to configure their system to automatically lock new notes, follow at least one of the suggestions listed below:If the customer must edit a note type, use the append feature and document that they are making an edit to the note.Lock all new notes that are entered via eChart Notes. To lock a note manually, after creating the note, right mouse click and select "Lock,"Access and enter patient notes via eChart Notes until a LANTIS upgrade has been performed at your site.When entering patient notes via eChart Navigator, make sure that you are entering the notes in the correct section.\When entering patient notes via eChart Notes, make sure that the correct note type has been selected.Customers are to keep the Customer Safety Advisory Notice together with their LANTIS 8.30 User Manual where it should remain until the measures defined above are performed. For questions call 1-888-826-9702. |  | 30 | 30 affected units | Worldwide Distribution - USA including DE, WI, TX, TN, NJ, PA, NE, MA, and AL and the countries of South Africa, Australia, Ireland, Germany, Brazil, South Korea, India, and Japan. | Terminated | April 13, 2011 | 600 | 30
Z-1969-2011 | 56807 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio Rad brand D-10 Hemoglobin A1c Program Reorder Pack, which contains Update Kit D-10 A1c Program Diskette; |  |  | September 08, 2010 | 2010 | Bio-Rad Laboratories, Inc., Hercules, CA | The firm discovered that the Update Kit Floppy Diskette included in the reorder pack was mislabeled with an incorrect Lot Number for the Calibrator/Diluent Set as: S01819; The correct Lot Number for the Calibrator/Diluent Set is: S01704. | Other | The Bio-Rad Laboratories, Inc, sent a Medical Device Correction letter dated 9/8/10, to all affected customers via Federal Express addressing this mislabeling issue, a new Update Kit Floppy Diskette, Lot No. BA02210 has been produced, and Bio-Rad requests that you please destroy the Lot No. AA01815. If you have any questions please contact your regional Bio-Rad office. You may also call (510) 724-7000. |  |  |  | Nationwide Distribution including : PA, OH, MO, CA, TX, FL, MA, MD, KY, OH, MI, NJ, and WA. | Terminated | June 29, 2011 | 294
Z-1976-2011 | 57714 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | SoftPerm (synergicon A) Daily Wear Contact Lenses, Rx Only, Sterile. Contents: One sterile contact lens, rigid center/soft hydrophilic skirt (25% water), in normal saline solution buffered with sodium bicarbonate. Ciba Vision Group, Duluth GA 30097 USA |  |  | February 01, 2011 | 2011 | CIBA Vision Puerto Rico, Inc. | Out of specification values were obtained for sterility testing. The pH of the saline solution may be higher than the spec (up to pH 8.6 compared to a spec limit of pH 7.8). The osmolarity of the saline solution may be lower than specification (263 mOsm/kg compared to a spec limit of 270) | Other | The firm, CIBA Vision, sent two letters of notification to its customers. First letter "IMPORTANT PRODUCT INFORMATION" dated June 2010 and the second "URGENT: Voluntary Product Recall" dated February 1, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to do the following: 1) complete the requested information on the attached SoftPerm (synergicon A) Lenses Reply Form and drop the form in the mail within 10 business days and 2) If they have any of the recalled SoftPerm (synergicon A) lenses (shipped to you after March 1, 2009) still in their possession, return as follows: Package the product and affix the prepaid UPS return label and ship within 10 business days or mail to CIBA VISION Corporation, c/o Stericycle Pharmaceutical Services, 2670 Executive Drive, Ste. A, Indianapolis, IN 46241, Atth: SoftPerm Recall. CIBA ask customers to please not include other products, lots, kits, packs, or competitor's products with their shipment. In addition, CIBA will issue a credit to the customers account upon receipt of the returned product. Please feel free to contact Stericycle at 877-542-5928 or CIBA VISION Technical Consultation Department at 1-800-241-7468 should you have any questions or concerns. |  | 28124 | 28, 124 | Worldwide distribution: USA (nationwide) including PR; and countries including: Australia, Austria, Argentina, Bahrain, Belgium, Canada, Colombia, Czech Republic, Denmark, Dubai (United Arab Emirates), Finland, France, Great Britain, Greece, Hungary, Iceland, Israel, Italy, Jordan, Kuwait, Lebanon, Luxembourg, Netherlands, Oman, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Turkey, and Venezuela. | Terminated | November 30, 2011 | 302 | 28124
Z-1968-2011 | 58054 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Lube Jelly, 1 oz. water-base sterile lubricating jelly, a component of various Presource Custom Sterile Convenience Packs/Modules; Packaged by Cardinal Health Medical Products and Services, McGaw Park, IL 60085 U.S.A.; The affected lots of modules actually contain a 1 oz. tube of Non-Sterile Kendall Vaseline White Petroleum Jelly. |  |  | February 28, 2011 | 2011 | Cardinal Health | The Presource kits were assembled with non-sterile petroleum base lubricant instead of the labeled sterile water base lubricating jelly. | Employee error | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated February 28, 2011 to all affected consignees via UPS next day air, informing them that their Presource kit/modules listed in the attached spreadsheet contain a non-sterile petroleum-based lubricating jelly instead of the labeled water-based sterile lubricating jelly. Cardinal instructed the accounts to quarantine the pack/modules with the indicated work order/lot number on the enclosed spreadsheet and contact the appropriate Customer Service group listed in the letter for return instructions. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules being returned. If the product was further distributed, then those customers should also be notified. |  | 4236 | 4,236 modules | Nationwide Distribution | Terminated | April 19, 2012 | 416 | 4236
Z-1973-2011 | 58035 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Rapid Response H.U.T. (Head Up Tilt) Table. Model 1017. Powered electrophysiology tilt table. Table can tilt -15 degrees to 90 degrees (Trendelenburg). Medical Positioning Incorporated.800-593-3246. |  |  | February 01, 2010 | 2010 | Medical Positioning Inc | The tilt tables work intermittently and some have become stuck in an inclined or reclined position during procedures. | Other | An undated Urgent - Voluntary Medical Device Product Correction letter was issued on 2/1/2010 by certified mail with return receipt requested. The letter identified the model and serial numbers of the affected product. The consignees were advised to stop using the table immediately and remove it from service until further notice. Customers were to contact Mr. Joe Hurtig at 1-800-593-3246 or jhurtig@medicalpositioning.com to confirm receipt of the notice and to provide the firm with the correct contact information. Individualized correction packages including replacement parts and instructions were sent to each consignee. The instructions consisted of (1) Rapid Response HUT Motor Replacement Instructions; (2) Control Box Replacement Instructions; and (3) ILD External Power Supply Removal Instructions. |  | 58 | 58 total | Worldwide Distribution -- USA, including Puerto Rico, and Canada. | Terminated | June 15, 2011 | 499 | 58
Z-1972-2011 | 58022 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC Clinical Systems with Serial Number less than 3412 Including: UniCel DxC 600 Part #: A10405, |  |  | November 16, 2010 | 2010 | Beckman Coulter Inc. | The hose clamps in the DxC (with serial numbers less than 3412, labeled behind the lower right side instrument door) hydropneumatic system are sharp and pose a possible risk of injury. An operator could be cut by the hose clamps, which could in turn result in increased exposure to any waste products (biohazards) that might be present on the assembly. | Pending | The recall communication was initiated with Beckman Coulter forwarding a Product Corrective Action Letter (PCA) dated November 16, 2010 to the affected customers. The recall notice explained the reason for recall and asked that customers take the following actions:(1) Customers with a service contract will have hose clamp covers installed by their Field Service Engineer during the next scheduled Preventive Maintenance (PM). A kit can be ordered through their Customer Service Representative if consignees prefer not to wait for the next Preventive Maintenance visit and(2) Customers without a service contract can obtain a kit by contacting their local Customer Service Representative.Customers were also informed that Kit (PN A68986) containing hose clamp covers and installation instructions is available at no charge from their Customer Service Representatives. One kit is required for each DxC system.The following telephone numbers were given for customers to obtain the kits: (800) 526-3821 (United States) and (800) 463-7828 (Canada).Consignees were asked to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If any of the affected product was forwarded on to another laboratory they were asked to provide a copy of this letter to them.Questions are to be directed to Beckman Coulter's Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. |  | 1535 | 1,535 units | Worldwide Distribution--United States and countries of Angola, Australia, Belgium, Bulgaria, China, Columbia, Czech Republic, Denmark, Eritrea, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Mexico, Netherlands, Reunion, Serbia, South Korea, Spain, Slovakia, Swaziland, Sweden, Switzerland, Tunisia, Turkey, Taiwan, United Arab Emirates, and the United Kingdom. | Terminated | June 28, 2012 | 590 | 1535
Z-1977-2011 | 58016 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | There is a possibility that a specimen could be diluted if a user requests Diluent Dispense while specimens are being automatically processed in cassette presentation. The dilution may cause erroneous results.IMPACT:Erroneous results from a diluted specimen could impact clinical decision making in cases such as blood transfusion, chemotherapy or antibiotic administration- particularly for popu | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" (PCA) letter with PCA Response Form (via US Postal Service for US customers) dated October 20, 2010, to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter describes the product, problem and the following actions to be taken. The customers were instructed to remove all cassettes from the mix station prior to initiating the Diluent Dispense function to prevent the likelihood of a specimen being diluted; do not use pre-diluted samples for RETIC counts; verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and the Body Fluid results are properly validated at the LIS. This will be corrected in a future version; use the lip on the reagent drawer as shown in the picture above; avoid touching the drawer slides while replacing containers; clean all spills promptly to reduce the risk of exposure to blood borne pathogens; complete and return enclosed PRODUCT CORRECTIVE ACTION RESPONSE FORM via fax to 786-639-7500/7501/7502/7503/7503/7504; mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 31-B06, 11800 S.W. 147th Avenue, PO Box 169015, Miami, FL 33116-9015 or email to: Regulatory.Notifications@beckmancoulter.com, and share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If you have any questions regarding this PCA, contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative. |  | 422 | 422 units in total (175 in US) | Worldwide distribution: USA (nationwide) and countries including: Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | September 02, 2014 | 1413 | 422
Z-1981-2011 | 57776 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | SmartBlock PainPump, Part numbers: P49524-S, P49534-S, P49624-S, and P49634-S. |  |  | March 17, 2010 | 2010 | Hsmg, Inc. D/b/a Smartinfuser Usa | Elastomeric tube in product's compression unit has the potential to burst following filling. | Other | Firm notified consignees by e-mail and telephone on 3/17/2010. The correspondence identified the affected product and lot numbers; and asked that customers specify which parts, lot numbers and in what quantities they still had in their possession. |  | 368 | 368 units | Nationwide Distribution | Terminated | April 25, 2011 | 404 | 1851
Z-1990-2011 | 58225 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Posterior Referencing Instrumentation for the NEXGENÂż COMPLETE KNEE SOLUTION MODULAR T-HANDLE, REF 00-5901-061-00, NONSTERILE, Zimmer UK, Ltd.,, Zimmer Warsaw,IN |  |  | March 14, 2011 | 2011 | Zimmer Inc. | The firm is initiating a recall due to the potential for breakdown of the AlTiN PVD (black) coating during surgery and the coating could fall into the surgical site. The firm has received 22 reports of the black coating coming off the instruments after cleaning and sterilization cycles and 5 resulted in a delay of surgery. | Device Design | The firm, Zimmer, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 14, 2011 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:Instructions for Handling of Affected Product:1) Stop using the devices and to quarantine immediately2) Carry out a physical count of all affect product in their territory and record data on Inventory Identification Form included with this letter-Note: The instruments must be returned even if the instrument is not exhibiting the black powder3)Ensure they have provided acknowledgement via "email return receipt" upon receiving this notification4) Fax copy of completed form to (574) 372-42655) Return the recalled product along with the completed Inventory Identification Form and the Sterilization Certification Form to: Zimmer Distribution Center, Attn: Product Service, 1777 West Center Street, Warsaw, IN 46580A new or refurbished instrument will be provided as replacement for their returned product once the inventory is available.Instructions for Health Care Professional Notification1) Provide a list of all hospital that currently use or have used the affected product listed in notice by utilizing the User Facility Documentation Form and the Health Care Professional Documentation Form2) Fax a copy of the completed form(s) to: Zimmer, Inc. at (574) 372-4265Zimmer will provide a similar notification via FedEx to these hospitals making them aware of this recall and provide a "Dear Doctor" letter for their surgeons.Important: If you have distributed these affected lots further, please provide a copy of this letter to these customers when you implement these recall instructions.For shipping assistance, questions or assistance in notifying your accounts about the recall, please contact Zimmer, Inc., at 1-800-613-6131.ADDENDUM INFOThe information listed above was updated to include the additional 5 |  | 552 | 552 | Worldwide distribution: USA (nationwide) and countries including: Australia, Canada, Germany, Japan, New Delphi and Singapore. Addendum: Finland, Switzerland, UK, Denmark, Spain, Portugal, France, Netherlands, Italy, South Africa, Norway, Spain, and Sweden. | Terminated | March 19, 2013 | 736 | 9462
Z-1980-2011 | 58288 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Floor Cabinet Part Number: 723335 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | Sharp edges on drawer slides within the DxH800 Floor Stand (Cabinet) may expose operators to physical injury while replacing reagents and waste containers.IMPACT:There is potential for exposure to blood borne pathogens such as HIV, HBV, and HCV. | Other | Beckman Coulter sent an Urgent Product Corrective Action letter dated October 20, 2010, with a PCA Response Form (via US PostalService) to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. This letter informed the customers about the problems identified and provided an action/resolution to this issue. Beckman Coulter asked consignees to complete and return the enclosed response form within 10 days and share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees local Beckman Coulter representative |  | 422 | 422 units in total (175 in US) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom | Terminated | April 02, 2012 | 530 | 422
Z-1983-2011 | 58048 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxi Sky 1000/BHM Medical V10 Bariatric Ceiling Lift System; a patient transfer lift; ArjoHuntleigh, Addison, IL 60101; Model numbers LF20010, LF20011, LF20013, LF20014, LF20015, LF21015, LF21019, 9120001, 9120004, 9120005, 9120008, 9120009 and 9121009 |  |  | March 21, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | The battery-powered control board in the ceiling lift motor may overheat, resulting in smoke and possibly flames and the lift becoming inoperable. | Device Design | ArjoHuntleigh sent an Urgent Field Safety Notice dated March 4, 2011, to all affected customers via UPS Ground on March 21, 2011. The notice informed the accounts of the potential for smoke and flames coming out of the Maxi Sky 1000/V10 bariatric ceiling lifts. The accounts were requested to immediately cease use of the affected lifts, powering them down by activating the emergency stop, and tag the lifts to prevent further use. If there is a clinical need to continue use of the lift before it is retrofitted or replaced, specific instructions were given to reduce the risk of a device malfunction. The accounts were requested to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 1246 | 1,246 units | Nationwide Distribution | Terminated | February 16, 2012 | 332 | 1246
Z-1984-2011 | 55910 | No | GKA | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT nbs Sickle Program Resin Update CD-ROM; Model Number:250-3020, included in 250-3000; Distributed and Manufactured by Bio-Rad Laboratories, Inc. The VARIANT nbs Sickle Cell Program is intended as a qualitative screen for the presence of hemoglobins F, A, S, D, C and E in eluates of neonatal blood collected on filter paper by high-performance liquid chromatography(HPLC). The VARIANT nbs Sickle Cell Program is intended for Professional use only. |  |  | April 14, 2010 | 2010 | Bio-Rad Laboratories Inc | The affected resin Update CD-ROMs cause all user-defined settings in the Setup/Test/Data Setup sub screen and Set/Test/Pattern Setting sub screen to be overwritten when the Update Kit procedure is performed. | Software design (manufacturing process) | Bio-Rad Laboratories sent out the Urgent Medical Device Correction notice on 4/14/2011, by Federal Express. An Update was sent on May 13th, advising customers of a minor discrepancy in the parameters included on the affected Resin Update CD-ROMs. Customers were instructed to examine your inventory for the affected reorder pack lots. Please complete and return the Medical Device Correction Response FormQuestions regarding this recall please call (510) 724-6114. |  | 658 | 658 units | Worldwide Distribution -- USA , Brazil, Canada, Caribbean, France, Italy, and United Kingdom. | Terminated | August 08, 2011 | 481 | 658
Z-1979-2011 | 58173 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER AC-T 5diff Autoloader Hematology Analyzer, Part Number: 175356, Serial |  |  | February 23, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter identified a problem with occasional unexpectedly high results for RBC, HCT, PLT and MPV without instrument generated messages, for the first run of a capped patient sample. Upon repeat analysis the results are correct. Patients could be affected as described below: (1) False high RBC and HCT may result in a delay of treatment,or the delivery of | Pending | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTION" letter dated February 23, 2011 with an enclosed customer response form to all customers who have the affected instruments via US Postal Service (in US). The letter described the product, problem and actions to be taken. The customers were instructed not to analyze specimens using the Autoloader. Instead, analyze their samples uncapped (open vial) using the Manual (Stat) Mode. The end of this notice includes a Quick Reference for RunningManual Samples. For more detailed instructions, please refer to your instrument's Instructions for Use (IFU), part number 624026, Section 8.4 Running Manual (Stat) Samples. An electronic version of the IFU can be found in the instrument's On-Line Help screen as well as the Beckman Coulter web site. For inconsistent data during patient's follow-up, please consider unusual/unanticipated high or elevated platelet levels as possible falsely elevated results due to this failure. An appropriate action for the need to look-back at previous results is at the discretion of the Laboratory Director, given the fact that laboratories have a variety of measures that are employed to ensure the validity of the results. Some of these provisions include delta checks, laboratory protocols for validating results and interaction with physicians.Additionally, the customers were also requested to complete and return the enclosed PRODUCT CORRECTION RESPONSE FORM within ten (10) days via fax to 786-639-7500/7501/7502/7503/7504; mail to Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 31-B06, 11800 S.W. 147th Avenue, PO Box 169015, Miami, FL 33116-9015, or email to: Regulatory.notifications@beckmancoulter.com.For any questions concerning this notice, contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact your local Beckman Coulter representative. |  | 1211 | 1,211 units | Worldwide distribution: USA (nationwide) and countries of: Angola Mexico Australia, Azerbaijan, Bangladesh, Bermuda, Botswana, Bulgaria, China, Cote d'Ivoire, Finland, France, Germany, Ghana, Hong Kong, India, Italy, Mexico, Netherlands, New Zealand, Nigeria, Palestinian, Puerto Rico, Romania, Russian Federation,Saudi Arabia, Slovenia, South Africa, Spain, Switzerland,Taiwan, Tanzania, and Turkey. | Terminated | July 16, 2012 | 509 | 1211
Z-2000-2011 | 51504 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens LANTIS 6.1 Oncology Information System with MEDOne Module, Linear Medical Accelerators; Product Numbers: 4503178, 5835991, 5851055; Product is manufactured by IMPAC Medical, Inc., Sunnyvale, CA. |  |  | March 06, 2009 | 2009 | Siemens Medical Solutions USA, Inc | The product has a software problem; it does not calculate the Creatinine Clearance correctly for patients less than two years old. | Software change control | On 3/6/2009, a Safety Advisory Letter was distributed to all affected users, with a description of the problem and identifying the affected product. Customers were instructed not to use the LANTIS to calculate Creatine Clearance for any patient under two years of age. Customers should direct their questions with regards to the letter to their local Siemens Service organization. A Siemens Representative will update the systems. |  | 17 | 17 | Worldwide Distribution | Terminated | April 15, 2011 | 770 | 17
Z-1998-2011 | 56636 | No | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | Cyclosporine Sample Pretreatment Reagent, COBAS INTEGRA, IVD, Roche Diagnostics GmbH, Assembled for and Distributed by: Roche Diagnostics Corporation, Indianapolis, IN. |  |  | August 13, 2010 | 2010 | Roche Diagnostics Operations, Inc. | The cyclosporine sample pretreatment reagent lot was assigned an incorrect expiration date. | Other | Roche Diagnostics Corporation sent an Urgent Medical Device Removal letter dated 8/13/2011, by UPS ( signature required ) to the consignees. The letter identified the product, the problem, and the action to be taken by the consignee.Consignees were instructed to discontinue use of affected product upon receipt of replacement product. Consignees were instructed to dispose of the product at their facilities. Therefore, no product returns are anticipated. Questions concerning this recall please call (317) 521-3911. |  | 22 | 22 kits | Nationwide Distribution -- including, IN, NJ, NY, TX, CA, PA, OH, and FL. | Terminated | May 04, 2011 | 264 | 22
Z-1999-2011 | 58202 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar SP Infusion Suite Database, Version 1.0; an Rx PC software application; Hospira, Inc., Lake Forest, IL 60045; list 13092-04-01 |  |  | March 23, 2011 | 2011 | Hospira Inc. | The software enables programming of the GemStar device for a Bolus Lower Lockout Limit outside of the hard limits entered into the GemStar SP Infusion Suite Database PC application. | Software design | The firm, Hospira, sent "URGENT DEVICE FIELD CORRECTION" letters dated March 23, 2011 to their customers on the same date. The letter described the product, problem and preventative action to be taken by the customers until the software can be upgraded. The customers were requested to contact Hospira Global Product Safety and Complaints at 1-800-441-4100 to report adverse events, and Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance. The accounts were also requested to complete the attached Reply Form and return it to Hospira via fax at 1-262-577-6921.For additional information and technical assistance, contact Hospira Advanced knowledge Center at 1-800-241-4002, option 4 (available 24 hours a day/7 days per week). |  | 139 | 139 units | Worldwide distribution: USA (nationwide) and countries including: Canada, Chile, Colombia and Singapore | Open, Classified |  |  | 139
Z-1874-2011 | 58134 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix-i, DFP-8000D |  |  | December 21, 2010 | 2010 | Toshiba American Medical Systems Inc | Under some conditions during high doses procedures the cumulative air kerma display may reach its maximum capability and fail to display the accrued value. | Radiation Control for Health and Safety Act | An Important: Electronic Product Radiation Warning notification was sent to all consignees. The letters described the reason for the recall and identified the affected product. The system software will be modified to allow the shift of the decimal to one position to the right allowing for a display of 99999.9 mGy. This will allow for the retention of the resolution until such time as the system exceeds the 9999.99 mGy range. A Toshiba Service person will contact customers when the software is available. Installed systems with the current version of software ( version 4.23) will begin to receive the corrected software April 2011. Previous versions of the software will begin to receive corrected software June 2011. The software will be provided free of charge. Questions related to the letter should be directed to the Regulatory Affairs Department at 800-421-1968. |  | 220 | 220 units Nationwide | Nationwide Distribution | Terminated | March 18, 2013 | 818 | 220
Z-2002-2011 | 58125 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sarns Flexible Weighted Sucker in K9 Pack (custom pack), catalog number 63635-01, Sterile and Non-pyrogenic. Sterilized by ethylene oxide. Terumo Cardiovascular Systems, Ashland, MA . |  |  | February 25, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has received two reports of the Terumo Weighted Flexible Sucker becoming ensnared in cardiac tissue after it was inserted into a ventricle through a valve. The firm has determined that the products's instructions for use are inadequate; the product should not be used to drain a cardiac chamber or inserted through a valve. However this is not stated in the IFU. The firm is conducting | Device Design | SEE UPDATE BELOW IN ALL CAPS.The firm, Terumo Cardiovascular Systems (Terumo CVS), sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were informed that the product's Indications For Use (IFU) are inadequate; the product should not be used to drain a cardiac chamber or inserted through a valve. However, this is not stated in the IFU. In addition, the customers were instructed to read and understand the changes and clarifications to the product's IFU; review the Medical Device Correction notice; assure that all users are aware of this notice, and confirm receipt of this communication by faxing the attached Customer Response Form via fax to: Terumo CVS Customer Service at 800-292-6551 (US only).If you have any questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818. Customers Service hours: Monday-Friday, 8AM-6PM.ON 8/30/2011 THE FIRM ISSUE A LETTER ENTITLED "URGENT MEDICAL DEVICE RECALL" THAT STATED "AFTER FURTHER CONSIDERATION, TERUMO CVS DECIDED TO DISCONTINUE SALES OF THE WEIGHTED FLEXIBLE SUCKER AND IS ASKING USERS TO RETURN ALL UNUSED PRODUCT" TO THE FIRM. |  | 11690 | 11,690 | Worldwide distribution: USA (Nationwide) and countries including: BRAZIL, BELGIUM, BRAZIL, DOMINICAN REPUBLIC, UNITED ARAB EMIRATES, and MALAYSIA. | Terminated | December 10, 2012 | 654 | 11690
Z-2022-2011 | 58041 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT CHEM 8+ Cartridge, List Number 03P91-25; 510 K053110 |  |  | February 15, 2011 | 2011 | Abbott Point Of Care Inc. | i-STAT CHEM 8+ cartridge may generate false results for sodium, chloride and hematocrit. | Other | The firm, Abbott Point of Care (APOC), sent an "Urgent Recall Notice" letter dated February 2011 with business reply cards attached via Fed Ex overnight to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of cartridges from the box numbers of identified lots; discontinue use as a precaution, if their cartridges are no longer traceable to a specific box; return all unused cartridges as instructed on the enclosed Business Reply Card; if they've provided another institution with the cartridges, provide the institution with a copy of this letter, and complete and return the business reply card to APOC.If you have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-336-8020, Option 1 or your Abbott Point of Care representative. |  | 6575 | 6575 cartirdges | Nationwide distribution. | Terminated | July 10, 2012 | 511 | 6575
Z-2013-2011 | 58148 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Lubricating Jelly packaged as follows: |  |  | February 07, 2011 | 2011 | DeRoyal Processing Center | Possible contamination due to sterilization issue. Inspection of contract manufacturer found cGMP deficiencies. | Process control | The firm notified their consignees by a Re: DeRoyal Recall Of Triad Group Lubricating Jelly letter, dated Friday, March 18,2011. The letter discussed the reason for recall and the affected products. The letter asked customers to identify inventory using the attached spreadsheet and quarantine to prevent further use. Customers are to complete the spreadsheet regardless if they have effected product or not. The completed form can be faxed to 865-362-3716 or emailed to recalls@deroyal.com. All affected products should be return to DeRoyal at the address provided. Distributors are to forward a copy of the notice to their end-users. If there are questions or assistance is needed with the recall, Stephanie DeBusk should be contacted at 865-362-1037 or Sharon Cook at 865-362-6157. |  | 323440 | 323,440 units | Nationwide Distribution | Terminated | August 05, 2014 | 1275 | 323440
Z-2021-2011 | 58080 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Avanto, Espree, Trio A Tim System. |  |  | March 03, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm discovered during product monitoring that on a few systems the gradient cable connections did not meet their specifications after installation. The potential exists for a cable connection to become overheated, leading to the emission of smoke and possible material damage. | Other | Siemens Medical Solutions Inc., issued a Customer Safety Advisory Notice dated November 24, 2010, to affected customers via Update Instructions MR029/10S. Affected customers will be notified via hand delivered letter by Siemens Service Engineers at the time of inspection. A future Update instruction will be released for installation of detectors in appropriate locations on systems to switch off the power electronics in the event of a malfunction in the power supply of the gradient coils.Further questions regarding this recall please call (610) 448-3237. |  | 1419 | 1419 | Product was distributed domestically to hospitals and imaging centers in the District of Columbia, Puerto Rico, AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WY(every State except VT). | Terminated | December 09, 2014 | 1377 | 1419
Z-2023-2011 | 55577 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Siemens brand 550 TxT Patient Treatment Table, part number 7346534; Product is manufactured and distributed by Siemens AG, Medical Solutions, Roentgenstrasse 19-21, Kemnath, Germany 95478. |  |  | March 31, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Potential for causing overheating of the power supply during treatment, and potential for unexpected table descent. | Component design/selection | Siemens issued a Customer Information letter dated March 21, 2010. There were two safety issues. There were two occurrences of smoke caused by an overheating of the power supply during patient treatment. Also, there had been a report of an unexpected lowering of the 550TxT table in patient rescue mode. With the released kit that accompanies this letter, Siemens has implemented a minor modification on the 550TxT to include an additional fuse within the power supply circuit to address the first issue described above. They also included an additional modification to address the lowering of the table by the integration of additional relays. As soon as the "move enable" button is released, the table brake will be enabled which is controlled by the implemented relays. No further action is required for the continued operations of the 550TxT table.Further questions regarding this recall please call (925) 246-8407 |  | 326 | 326 units | Worldwide Distribution -- Including USA, states of NJ, WI, NY, and MA. | Terminated | March 01, 2012 | 701 | 326
Z-2025-2011 | 58276 | No | KDC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hospira LifeShield Latex-Free HEMA Y-Type Blood Set, Nonvented, 100 Inch with Clave and Secure Lock; a sterile Rx tubing set for the alternate administration of fluids from a blood bag and I.V. solution container; Made in Costa Rica; Hospira, Inc., Lake Forest, IL 60045 USA; List No. 12450-48 |  |  | March 28, 2011 | 2011 | Hospira Inc. | There is a potential for the air filter at the distal end of the set to become detached. This could result in the contamination of the sterile fluid path inside the tubing. | Other | The firm, Hospira, sent an "URGENT DEVICE RECALL" letter dated March 28, 2011 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to check their inventory and immediately quarantine any affected product; complete and return the Reply Form via fax to 1-866-324-3734, even if you do not have the affected product; inform healthcare professionals in their organization of this recall; if they have distributed the product further, notify their accounts and have them fax the reply form to 1-866-324-3724; return the affected product to Stericycle using labels provided, and call Stericycle at 1-877-877-0481, if they have not received a return label. For medical inquiries, please call Hospira Medical Communications at 1-800-615-0187. Call Hospira Global Product Safety and Complaints at 1-800-441-4100 (8am-5pm CST, M-F) or by e-mail at (ProductComplaintsPP@hospira.com). |  | 47472 | 47,472 sets | Worldwide distribution: USA (nationwide) and countries including: Canada, Korea and Singapore. | Terminated | May 09, 2012 | 408 | 147646
Z-2047-2011 | 55154 | No | CFT | Prolactin (lactogen) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Active Prolactin ELISA, 96 Wells, DSL-10-4500 |  |  | August 03, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Active Prolactin ElLISA kit, lot number 890291 over estimates the expected dose values for patient serum samples due to a shift in the assigned calibrator values.Beckman Coulter recommends that patient results generated with the affected lot be evaluated in the context of other diagnostic tests and the clinical presentatio | Equipment maintenance | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" Letter dated August 03, 2009 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:Action/Resolution:Discontinue use and discard remaining product with lot number 890291 of Active Prolactin reagents. Contact Beckman Coulter at 1-800-231-7970 for replacement product.Newer lots (lot number 991210 or higher) of the product have been realibrated to restore the performance as described in the Instructions for Use (IFU). For method correlation information, they were told to contact Technical support at the phone number listed below.Consignees were asked to share this information with their laboratory staff and retain this letter for their Quality System documentation and asked to complete and return the enclosed response form within 10 days via fax to: 786-639-4000 so that Beckman can be assured that they received the notification.If you have any other questions regarding this notification, please contact DSL Technical Support Center at 1-800-231-7970 in the United States and Canada , or contact their local Beckman Coulter Representative. |  | 30 | 30 Kits | Nationwide distribution: USA including states of: FL, GA and NC. | Terminated | April 27, 2012 | 998 | 30
Z-2045-2011 | 57750 | [6 empty] | IMMAGEÂż Immunochemistry Systems Wash Solution Part Number: 447060 |  |  | August 24, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed an ImmageÂż Systems Wash Solution Lot 71744 was manufactured by a Beckman European facility and not intended to be used in the U.S. However, this lot was inadvertently shipped to the U.S. and delivered to five customers. | Other | The firm, Beckman Coulter, contacted the customers who received the IMMAGE Immunochemistry Systems Wash Solution, lot #71744 by telephone. Beckman sent a follow-up "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 25, 2010, to all affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of ImmageÂż Immunochemistry Systems Wash Solution lot 71744; return the material to Beckman Coulter's Chino warehouse: Beckman Coulter Inc., 15989 Cypress Ave, Chino, CA 91710, Phone: 909-597-3967;package the material the way it arrived and use Beckmans internal Fedex account to ship the material; complete and return an enclosed FAX BACK RESPONSE FORM within 10 days via fax to (714) 961-4234 or mail to: Beckman Coulter, Inc., Attn: Regulator Affairs, 250 S. Kraemer Blvd, Brea, California 92821, Mail code E2.Se.08; share this information with laboratory staff; retain the notification as part of their Quality System documentation, and provide a copy of the letter if they forwarded the affected products.If you have any questions regarding this letter, please call Beckmans Customer Support Center (Hotline) at (800) 854-3633 in the United States and Canada. |  | 9 | 9 units total in US | Nationwide distribution: USA including states of : AL, CA, MI and IN. | Terminated | June 26, 2012 | 672 | 9
Z-2044-2011 | 58174 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTERÂż HmX Hematology Analyzer |  |  | March 02, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has identified a problem with Integrated Circuit Chips used in different boards throughout the Beckman Coulter LH500/HmX/HmX Autoloader Hematology analyzer. | Other | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated March 2, 2011, with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the COULTERÂż LH 500 Series Analyzer, COULTERÂż HmX Hematology Analyzer, and COULTERÂż HmX Hematology Analyzer with Autoloader. The letter described the product, problem and actions to be taken. Customers were informed that the suspect Integrated Circuit Chips have been removed from all systems in inventory and manufacturing.The customers were instructed to share this information with their laboratory staff and retain this notification as part of their laboratory Quality Systemdocumentation. If customers have forwarded any of the affected product(s) listed on the PCA letter to another laboratory, please provide a copy of the letterto them.The customers were instructed to complete and return the enclosed Response Form within 10 days so Beckman are assured they have received this important communication and are taking appropriate action.The customers with any technical questions regarding the Product Corrective Action were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States and Canada or contact their local Beckman Coulter Representative. |  | 288 | 288 units total (158 in US) | Worldwide distribution: USA (nationwide) and countries of: Canada, France, Germany, Italy, Spain, United Kingdom, and Singapore. | Terminated | May 29, 2012 | 454 | 288
Z-2028-2011 | 57340 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Codman Disposable Twist Drill |  |  | November 30, 2010 | 2010 | Codman &amp; Shurtleff, Inc. | Packaging defect may compromise sterility of device | Package design/selection | The firm, Codman, sent an "Important Notice Urgent- Voluntary Product Recall Notification" letter dated November 30, 2010 to its customers via FedX. On January 7, 2011, Codman expanded the recall to four additional product codes.On January 31, 2011, the firm issued a supplemental recall letter and identified one additional product code: 80-11197. The letters described the product, problem and actions to be taken. The customers were instructed to identify the unused affected products in their inventory and follow the recall instructions provided: 1) If you have NO product to return, please advise Codman as soon as possible using the FACSIMILE form attached and fax to 508-828-6118 and 2) if you DO have product to return, please call 1-800-225-0460, Option 4, for instructions for return.Codman is presently working toward supplying replacement products as soon as possible. Please call Customer Service to check on the status of product availability at 1-800-225-0460.Should you have any questions regarding this recall, please call Customer Service at 1-800-225-0460 or your local Codman Representative. |  | 1549 | 1549 units | Worldwide distribution: USA including states of :AL, FL,GA, KY, MA, MI, MN, MS, OR, PA, RI, UT,WI and WV; and countries including: Argentina, Australia, Canada, Bahamas , Belgium, Dubai, Japan, Madrid, Mexico, New Zealand, Saudia Arabia, UAE, Uruguay, and United Kingdom. | Terminated | March 14, 2013 | 835 | 14792
Z-2035-2011 | 57933 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ***REF MCA5-15SB ***1.5 mm Fluted Ball, 5.6 cm. Use with Micro Curved Attachment (MCA).***ANSPACH***Sterile, Sterilized using Irradiation***Rx Only*** Made in the USA***Manufacturer: The Anspach Effort, Inc. 4500 Riverside Drive, Palm Beach Gardens, FL 33410 USA Tel: (800) 327-6887/ +1 561 627 1080***Anspach Europe Ltd. Aston Court, Kingsmead Business Park Frederick Place High Wycombe, HP11 1LA United Kingdom Tel: +441494 616126*** |  |  | September 29, 2010 | 2010 | The Anspach Effort, Inc. | Anspach Effort, Inc. Palm Beach Gardens, FL is recalling their Sterile Bone Cutting Burs that was manufactured with incomplete flutes which could cause degraded cutting performance and in some cases not cut at all. Anspach has not received any complaints related to this problem. | Other | The firm, Anspach, sent an "URGENT: Medical Device Product Removal" letter dated October 6, 2010, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Screen their inventory immediately for any of the part numbers listed in the letter. Do not include any product received after October 7, 2010.2) If they have any of the products listed, contact Anspach Customer Support at (800) 327-6887 to arrange for their immediate return and replacement.3) Complete and return the attached Customer Reply Form indicating how many of each part number is being returned and confirming their receipt of this letter, and fax to 1-800-327-6661. If they have distributed any of the products to other services or facilities, please forward this information as appropriate.Should you have any queries, please do not hesitate to contact Anspach Customer Support at (800) 327-6887. |  | 271 | 271 cutting burs affected) | Worldwide distribution: USA including states of: AZ, CA, FL, GA, IL, IN, LA, MA, and MI; and countries of: Australia, Belgium, England, France, Italy, Japan, South Africa, Spain and Sweden. | Terminated | August 12, 2011 | 317 | 271
Z-2040-2011 | 57888 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Orthoralix 8500 DDE Panoramic Dental X-Ray System; Extraoral x-ray unit with timer; Gendex Dental Systems, Des Plaines, IL 60018; part numbers 110-0199G2, 110-0199G4, 110-0234G1 and 110-0234G3. |  |  | February 04, 2011 | 2011 | Gendex Dental Systems | There is an anomaly in the Orthoralix firmware which can cause a message corruption which will distort the patient image being transferred from the X-Ray system to the PC. This image distortion/corruption (distorted, fragmented, flipped, negative, reversed or had lines across them) is most likely to occur during periods of high network traffic and is significant enough that it is clear and obviou | Other | Gendex sent a recall letter to their accounts on February 4, 2011 via first class mail, stating that an anomaly in the Orthoralix firmware can cause a message corruption which will distort the patient image being transferred from the X-Ray system to the PC. This image distortion/corruption (distorted, fragmented, flipped, negative, reversed or had lines across them) is most likely to occur during periods of high network traffic and is significant enough that it is clear and obvious these images cannot be used for diagnostic purposes. The letter stated that resending the image from the X-Ray system to the PC will correct the issue present in certain images sent during periods of high network traffic. Instructions were provided to download an Appendix to Orthoralix DDE User Manual that explains how to resend the images. If that does not resolve the issue, users were requested to contact Gendex Technical Support at 1-800-323-8029. Customers were asked to complete and return the enclosed Acknowledgement Form by fax to 215-997-5665 or e-mail: regulatoryaffairs@gendex.com. For questions regarding this recall call 800-323-8029. |  | 6835 | Total 6,835 units of all products | Worldwide Distribution--United States, including Puerto Rico, Australia, Canada, Chile, Colombia, Germany, Hong Kong, India, Italy, Korea, Mexico, New Zealand, Peru, Singapore, Taiwan and Vietnam. | Terminated | February 15, 2012 | 376 | 6835
Z-2054-2011 | 58324 | No | MRA |  |  | Orthopedic | PMA | R3 TM CERAMIC ACETABULAR LINER, BIOLOXÂż FORTE, 32 MM O.D. 55 MM O.D., REF 71338950, STERILE R, QTY (1), Smith &amp; Nephew, Inc., Memphis, TN 38116, 2017/06 |  |  | March 11, 2011 | 2011 | Smith &amp; Nephew Inc | During the manufacturing process for several batches of R3 Ceramic Liners, the titanium rings were pressed onto the ceramic component with a higher force than allowed by manufacturing specifications. This has the potential to result in lower than expected strength for the liners. | Employee error | All affected Smith &amp; Nephew Distributors were notified via email and the hospitals were notified by letter delivered by Fed-Ex on 3/11/2011. They were instructed to immediately cease distribution, quarantine affected product for return, if product further distributed, identify customers and immediately notify them of recall.Please contact Mr. David Archer, Group Director, at 701.396.2121, ext. 5373, if you have questions. |  | 325 | 325 units | AK, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, MA, MD, ME, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OR, PA, SC, TN, TX, WA, and WI | Terminated | February 08, 2012 | 334 | 2271
Z-2051-2011 | 58231 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 ultrasound system |  |  | March 15, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | The Firm discovered a software error on the Accuson S2000 ultrasound systems affecting the measurement results in the patient report. | Software design | Siemens issued a Customer Safety Advisory Notification letter to all customers on March 15, 2011, regarding a software error on the ACUSON S2000 ultrasound systems affected measurement . They are in the process of releasing a software update that will correct this issue. Siemens recommends reviewing the patient report at the end of each study to confirm that the report contains only measurements relevant to the exam just performed. No patient injury has been reported. Further questions regarding this recall please call (650) 694-5398. |  | 1533 | 1533 units | Worldwide Distribution including USA, states of AZ, VA, CA, NE, IL, MN, TN, FL, NM, and DC | Terminated | June 18, 2012 | 461 | 1533
Z-2061-2011 | 53053 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X300 Premium Edition Systems with software versions 5.0.00 and 5.0.01 (model numbers 10348531, 10348532 and 10348533), when using the V5Ms transducer (model number 08263703) and the Plus option (10427795) |  |  | July 24, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Product Failure-- When the 64-Channel System V5M is used in the 128-Channel mode, an over current fault occurs which shuts the system down. | Software design | Siemens sent a Customer Safety Advisory Notification letter on July 24, 2009. The letter identified the product, the problem, and the action to be taken by the customer. A software update will be conducted in the field. Until that update occurs, customers were instructed to image without the CW doppler to avoid a system shut down.Customers were instructed to contact their service support person with any questions.For questions regarding this recall call 650-694-5993. |  | 70 | 70 units | Worldwide Distribution - USA including CA, CT, FL, MA, and TX and the countries of United Arab Emirates, Germany, Ecuador, Spain, France, Great Britian, Saudi Arabia, Sweden, Thailand, and South Africa. | Terminated | April 25, 2011 | 640 | 70
Z-2065-2011 | 53209 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson X300 systems at software versions 1.00.00 to 1.0.10 and 2.0.00 to 2.0.02, model numbers 10037409, 10132987, and 10038837, manufactured by Siemens Medical Solutions ultrasound business unit, Mountain View, CA |  |  | July 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Software Error with potential for misinterpretation of data-- The previous patient's last capture clip image can appear in the next patient's study. | Software design | The firm, Siemens, sent a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to its customers on July 30, 2009. The letter described the product, problem and actions to be taken. The customers were instructed they need to update any of the patient registration settings to reset the cine buffer so that no data is transferred between studies, if the system software is not being updated. Siemens informed the customers that a mandatory update software is being installed on their system free of charge.If you have any questions, contact (650) 969-9112. |  | 2312 | 2312 units | Worldwide distribution: USA (nationwide) and countries including: Albania, algeria, Argentina, Australia, Austria, Azerbaijan, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile,China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Hungary, India, Indonesia, Iran, Ireland, Italy, Japan, Kazakhstan, Kenya, Kosovo, Lativia, Lithuania, Luxembourg, New, Zealand, Nicaragua, Norway, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Viet Nam. | Terminated | April 25, 2011 | 634 | 2312
Z-2060-2011 | 58261 | [6 empty] | LuxtecÂż MLX 300 Watt Xenon Light Source, |  |  | March 07, 2011 | 2011 | Integra LifeSciences Corp. | Malfunction of the MLX Control Board may cause loss of manual control and Lamp shut down. | Process change control | Integra notified all customers/distributors via written communication, dated 3/08/2011. On 3/09/11, Integra Sales Representatives and Service Technicians were notified via telephone conference call and provided copies of written communications. Customers were advised to immediately discontinue use of any affected product and to place in a secure area. Note: This defect only applies to MLX units that were repaired between 12/01/2009 and 1/31/2011. This defect does not occur in manufactured units that have not been serviced, nor does it apply to MLX units that were serviced outside of this specific timeframe. Customers with devices already known to require repair and/or servicing will be contacted directly to arrange for the return of their device(s). Additionally, sales specialists were instructed to examine each customer's inventory for those device(s) that may potentially have the jumper-wire and to perform a stand-by test to determine if repair and/or servicing are required. |  | 159 | 159 units | Nationwide, Canada, Australia, Germany, Netherlands, New Zealand, Norway, Singapore and Sweden | Terminated | March 26, 2014 | 1115 | 159
Z-2063-2011 | 58246 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03899; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | March 03, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The optical guidance platform may not be properly enforcing a 24 hour time limit between optical camera recalibrations. | Software change control | The firm, Varian, sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check the Calibration timeout on their optical guidance platform. If this exceeds 24 hour, please set the factory default of 16 hours to ensure that the camera will be calibrated daily. The instructions are located in the document. The customers were also instructed to download the most recent version of the Optical Guidance Platform Instructions for Use at www.MyVarian.com or request a copy from their local Varian Customer Support District or Regional Manager. In addition, the customers were instructed to advise the appropriate personnel working in their radiotherapy department of the content of this letter. Varian is preparing a software patch which will be applied to affected units to correct this issue.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager; USA and Canada -1888-VARIAN5 (888-827-4265) or Europe- +41 41 749 8844. |  | 599 | 599 units | Worldwide distribution. | Terminated | June 19, 2012 | 474 | 599
Z-2062-2011 | 58320 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | S8-3t Micro TEE Transducer is an Ultrasonic Diagnostic. The Part # is 989605379471. |  |  | March 27, 2011 | 2011 | Philips Medical Systems Nederlands | Users of the Philips model S8-3t Micro TEE Transducer have reported that significant image quality degradation may occur after a number of use/cleaning cycles. If users do not inspect the transducer for such damage prior to each use, imaging may not always be available, which could pose a risk for patients. | Other | Philips Healthcare began mailing Urgent Medical Device Correction letter on March 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers are being instructed on how to inspect the distal tip for defects that cause image quality degradation, and to plan for epicardial / transthoracic ultrasound imaging as an adjunct to the S8-3t. Customers are also reminded to carefully follow cleaning described under Handling the S8-3t Transducer, of Philips' iE33 System User Information. Customers can call Philips Customer Service at 800-722-9377 and reference to FCO 79500198 for any questions about this recall. |  | 155 | 155 units | Worldwide Distribution - USA AL, AR, AZ, CA, CO, DC, DE, FL, GA, IA, IL, KY, LA, MA, MD, MI, MN, MS, NC, ND, NH, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI, and the countries of Australia, Austria, Belgium, Chile, Denmark, Finland, France, Germany, India, Israel, Italy, Japan, Korea Rep, Netherlands, Singapore, South Africa, Spain, Switzerland, Taiwan, Thailand, Ukraine, and United Kingdom. | Terminated | February 14, 2012 | 324 | 155
Z-2074-2011 | 57779 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCelÂż DxC 680i |  |  | October 22, 2010 | 2010 | Beckman Coulter Inc. | This recall was initiated because Beckman Coulter has confirmed that for User-Defined Chemistries (UDRs), the polychrome correction is not automatically updated when wavelengths are modified. | Environmental control | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased Synchron LX and UniCelÂżÂż DxC Clinical SystemsOperating Software (all versions). The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:-After modifying any UDR wavelength, they MUST reboot to apply the update of the polychrome correction. -If the modified UDR requires calibration, recalibrate after rebooting. - Complete and return the enclosed RESPONSE FORM within 10 days via fax to (714) 961-4234 or mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs, 250 S. Kraemer Blvd, PO Box 8000, Brea, California, 92282-8000; Mail Code E2.SE.08. In addition, the customers were instructed to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If customers forwarded any of the affected product listed above to another laboratory, please provide a copy of this letter to them.If customers have any questions regarding this Product Corrective Action letter, please call Beckman's Customer Support Center (Hotline) at (800) 854-3633 in the United States and Canada or www.beckmancoulter.com/customersupport/supportrequests. |  | 4879 | 4879 units total (3490 in US) | Worldwide distribution: USA (nationwide) and countries including:Angola, Australia, Belgium, Bulgaria, China, Colombia, Costa Rica, Czech Republic, Denmark, Ecuador, Egypt, Eritrea, Finland, France, French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Macao, Mexico, Namibia, Netherlands, Nez Zealand, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Rwanda, Saudi Arabia, Serbia, Singapore, Slovakia, Spain, South Africa, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | September 02, 2014 | 1411 | 6289
Z-2052-2011 | 57744 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific's LATITUDE Patient Management System: |  |  | January 10, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has determined that data within certain printed reports obtained from the LATITUDE Patient Management System may not be accurate, in that, the Most Recent Presenting Electrogram (MRP-EGM) for some of the selected patients may be either missing or incorrect. | Software design | Consignees were sent a Boston Scientific "LATITUDEÂż Medical Records Information" letter dated January 2011. The letter was addressed to "Dear Doctor". The letter described the problem and the product involved in the recall.The letter also included the Rate of Occurrence, Action and Further Information sections. |  | 25836 | 25,836 | Worldwide Distribution -- USA, including DC and PR, and countries of AUSTRALIA, CANADA, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NORWAY, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | November 15, 2012 | 675 | 25836
Z-2068-2011 | 58042 | No | GKE | Osmotic fragility test. | Hematology | Hematology | 510(K) Exempt | Advanced Instrument Model 2020 Multi-Sample Osmometer |  |  | January 19, 2011 | 2011 | Advanced Instruments, Inc. | A poorly soldered power cord may pose a potential electrical shock hazard. | Employee error | Advanced Instruments Product Service Department contacted each customer on 1/19/11 via an Urgent Product Recall Notice. The letter identified the reason for recall along with the affected products. The letter provided instructions to safely examine the date code of the AC Power Inlet assembly on the affected product. If customers were unable to perform the inspection or had questions concerning the instructions, they were to contact Advanced Instruments Product Service Manager at 781-471-2147 or via email jorges@aicompanies.com. The attached response forms were to be completed and returned as instructed. |  | 3 | 3 units | Worldwide Distribution -- USA, including states of CA, GA IL, NJ, and MA and country of Belgium. | Terminated | December 01, 2011 | 316 | 3
Z-2066-2011 | 58371 | No | HEM | Obstetric-gynecologic ultrasonic imager. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | OB-Link release 2.0.7. distributed by Digisonics, Inc., 3701 Kirby Drive, Houston TX 77098 |  |  | May 12, 2008 | 2008 | Digisonics, Inc | Data/measurements from a OB report sent from ultrasound systems populated the wrong OB-View report | Other | Digisonics visited consignee in April 2008 as follow-up to complaint. |  | 2 | 2 units | Nationwide Distribution including OK | Terminated | May 02, 2011 | 1085 | 2
Z-2059-2011 | 58209 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice Computed Tomography Scanners, equipped with Software Version 1.1.1. Model #E989605657371. |  |  | March 10, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The Philips MX 16-slice CT scanners are subject to recall/field correction due to a problem encountered with the units when using the 'Combine Viewing' Option. The firm determined that inaccurate dimension measurements display when using the Combine Viewing Option. Moreover, if the dimension measurements from merged images created using this option are used, there is a potential for applying inac | Other | Philips Healthcare issued an 'URGENT- Field Safety Notice MX 16-slice' (dated 3/10/2011) concerning a software issue to customers. The notification informed the firm's users/customers of the potential for inaccurate dimension measurements produced by the MX 16-slice CT scanners when using the Combine Function Viewing Option. Moreover, the letter notified customers of the potential for patient misdiagnosis in the event that inaccurate dimension measurements are used by a physician for reference. The Field Safety Notice described the following: 1) How to identify the affected product; 2) What actions should to be taken by the customers with regard to avoiding the issue -discontinue use of the Combine Viewing Option and instead use the MPR Batch for diagnostic reviews until a planned software upgrade is installed by the recalling firm; and 3) The firm's plans to carry out the installation of a field change order kits to correct the issue.The planned software upgrade is expected to be completed on or before 6 months after the new software is released. If you need any further information or support concerning this issue, pleasecontact your local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center(1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 193 | 25 units were distributed domestically &amp; 168 system units were distributed both domestically &amp; internationally | Worldwide Distribution -- United States (AL, AR, CA, FL, GA, IL, IN, MS, NJ, NV, NY, OH, PA, SC and TX) including Puerto Rico and countries of Albania, Argentina, Australia, Belgium Brazil, Chile, Columbia, Dominican Republic, Ecuador, Germany, Greece, India, Indonesia, Italy, japan, Kenya, Lebanon, Malawi, The Netherlands, Philippines, Poland, Romania, Saudi Arabia, South Africa, Spain, Switzerland, Turkey, Uganda, United Arab Emirates and Venezuela | Terminated | March 01, 2013 | 722 | 193
Z-2092-2011 | 58289 | No | FTN | Plastic surgery kit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Argon Medical Devices Arterial Line Kit/1, 18ga x 3", Part Number 498110, Lot numbers 99768048 and 99768520 |  |  | March 14, 2011 | 2011 | Argon Medical Devices, Inc | Packaging for sterile tray has the potential to have a hole which could compromise sterility for the products inside. | Other | Argon Medical Devices sent an Urgent Medical Device Recall letter dated March 14, 2011, by certified mail to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify, quarantine and return any product remaining at their facility utilizing the instruction on the 2011 Tray Recall Response Form. Customers were instructed to complete the response form included with the letter and return by fax or email as indicated as soon as possible. Customers were instructed to pass the notice on to all those who need to be aware within their organization or to any organization where the potentially affected devices have been transferred. For questions regarding this recall call (903) 677-9375. |  | 480 | 480 units | Worldwide Distribution - USA including AL, AZ, CA, CO, DE, GA, HI, IL, IN, LA, MD, MN, ND, NE, NH, NJ, NY, OK, PA, SC, TX, VA, WA, and WI and the countries of Nationwide, Japan, Australia, Canada, Singapore, and the Netherlands | Terminated | April 24, 2012 | 407 | 5780
Z-2050-2011 | 58135 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK 2 Gram Negative Susceptibility Card, AST-N184, REF 411 154, For In Vitro Diagnostic Use, 20 cards/box. The firm names on the label are bioMerieux, Inc., Durham, NC, and bioMerieux sa, Marcy-l'Etoile France. |  |  | April 04, 2011 | 2011 | Biomerieux Inc | The piperacillin/tazobactam drug on the antimicrobial susceptibility cards is providing false susceptible and false resistant results for E. coli, as well as an additional 7 organisms; and the same drug is providing false resistant results for K. pneumoniae. | Nonconforming Material/Component | BioMerieux sent an Urgent Product Correction Notice (EXPANDED RECALL), VITEK 2 Piperacillin/Tazobactam Test, Document #9300786," letter dated March 30, 2011, via Fed EX on April 4, 2011. to all affected customers. The letter describes the issues, how the issues impact the test results, describes the product changes that are going to take place, and provides instructions on required actions to be taken.The product changes will involve the eventual changeover to new cards that will be created which do not contain Piperacillin/Tazobactam (TZP2). The new cards will begin shipping by July 16, 2011, but in the meantime, the cards currently in possession of the customers can be used with specific required actions that need to be followed. Until the TZP redevelopment is completed and a resolution can be implemented, BioMerieux requires their customers to immediately take the following actions: The customers will be required to suppress the TZP2 resistant and susceptible results for Escherichia coli, Morganella morganii, Proteus mirabilis, Proteus vulgaris, Providencia rettgeri, Providencia stuartii, Pseudomonas aeruginosa, and Salmonella enterica and use an alternate method of testing. In addition, the customer is to suppress the TZP2 resistant result only for Klebsiella pneumonia and use an alternate method of testing.The letter provides instructions for configuring the VITEK 2 software to suppress the results from reporting.The instructions contained in the recall notification are to be applied to all current and future product lots that contain TZP2 until further notice.The customer is to complete the enclosed acknowledgement form and return it via FAX to acknowledge receipt of the notification and is requested to file the correction notice with their VITEK 2 Systems permanent records according to their internal policy or procedure.For questions or additional information customers were instructed to call 1-800-682-2666, option 3.In addit |  | 9000 | 3,000 boxes as of 3/31/11 | There was no U.S. distribution as the cards were for export only to Canada and Sweden. | Terminated | June 01, 2012 | 424 | 10479
Z-2103-2011 | 58271 | No | OCJ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ConfidenceÂż Introducer Needle Side Hole 13G x 4-Inch |  |  | March 22, 2011 | 2011 | DePuy Spine, Inc. | Mislabeled: Package labeled (11g x 6 inch) contains a 13g x 4-inch side hole needle | Packaging process control | The firm, DePuy Spine, sent an "URGENT-VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated March 22, 2011 to its consignees/customers ( hospital sites and domestic sales distributors). A product reconciliation form was provided to assist in the retrieval of affected product. The letter described the product, problem and actions to be taken. The customers were instructed to return any needles that they may have to their local Distributor for DePuy Spine products and complete and return the Product Reconciliation Form via fax to 508-828-3762; call Customer Service at 877-379-4871 to receive a return goods authorization number and/or mail the form with the return product to DePuy Spine Inc., 50 Scotland Boulevard, Bridgewater, MA 02324, Attn: CONFIDENCE NEEDLE RECALL COORDINATOR, RGA #.If you have any questions or concerns with regard to this notice, please contact the DePuy Spine Regulatory Compliance Department at 508-828-2790. |  | 67 | 67 units | Worldwide distribution: USA including states of: AZ, IA, LA, TX, NE, NJ ,NY, PA, VA, and WA; and countries of: Canada, Australia, Belgium, France,Hungary, Italy, Malaysia, Solvenia, UK, South Africa, Switzerland, and UAE. | Terminated | February 20, 2013 | 701 | 282
Z-2082-2011 | 58291 | No | LSE | Epstein-Barr virus serological reagents. | Microbiology | Microbiology | 510(k) | BioPlex 2200 EBV IgG Calibrator Set. Catalog #663-1200. |  |  | March 11, 2011 | 2011 | Bio-Rad Laboratories | Bio-Rad Laboratories confirmed a downward shift in control values for EBV NA-1 analyte when using the BioPlex 2200 EBV IgG Calibrator Set Lot #46704. | Other | The firm, Bio-Rad Laboratories, sent an "Voluntary Field Correction: BioPlex 2200 EBV IgG Calibrator Set, Lot# 46704 for EBV NA-1 Analyte" letter dated March 16, 2011 domestic customers via Fed-Ex and and to international subsidiaries via e-mail. The letter described the product, problem and actions to be taken. The consignees/customers were informed that Bio-Rad confirmed downward shift in control values for the EBV NA-1 analyte when using the BioPlex 2200 EBV IgG Calibrator Set, Lot # 46704. If this occurs, the customers were instructed to readjust the positive control mean using the procedure presented in the customer notification "BioPlex 2200 Quality Control Values (L60013402)". If the re-adjusted mean of the EBV NA-1 controls is within the manufacturer's limit, patient results may be reported. If the new assay positive control mean still does not fall within the manufacturer's limit, the customers were instructed to contact their local Bio-Rad Technical Support and do not report the EBV NA-1 values. The customers were also instructed to complete and return the CUSTOMER RESPONSE FORM via fax to 510-741-5775.If the BioPlex 2200 EBV VCA IgG and EBV EA-D IgG positive and negative control means are within the acceptable QC range, and subsequent passing QC runs are obtained, patient results may be reported for these analytes. Bio-Rad will be manufacturing a new lot of the BioPlex 2200 EBV IgG Calibrator Set which will bring the control values within the manufacturer's limit. It may be necessary to readjust the QC ranges after the consignees receive this new calibrator set lot.Bio-Rad will send their consignees/customers the replacement calibrators and calibrator CDs when available. The consignees/customers were also instructed to discard the calibrator set lot# 46704 in accordance with national and/or local regulations.Customers can call Clinical Diagnostics Group at 510-741-4431 or email to BPD-ProductSupport@bio-rad.com for any questions |  | 270 | 270 sets | Worldwide distribution: USA including the following states: AL, AZ, CA, CT, FL, GA, HI, IL, IN, LA, MA, MD, MI, MN, NC, NJ, NY, OH, OR, PA, TN, TX, VA, WA, and WI; and countries of: Australia, Czech Republic, France, Germany, Italy and Republic of Korea. | Terminated | December 22, 2011 | 286 | 270
Z-2067-2011 | 57912 | No | LNR |  |  | Gastroenterology/Urology | PMA | Therakos UVAR XTS Procedural Kits |  |  | February 01, 2011 | 2011 | Therakos Inc | Therakos received an increased rate of complaints of centrifuge bowl leaks for certain lots of Therakos UVAR XTS Procedural Kits (Product Code XT 125). | Process control | Therakos notified all direct Therakos US and ex-US End Users and Distributor consignees via Federal Express courier with an Urgent Product Correction Notification, dated 2/1/2011. The letter identified the reason for recall along with the affected product. Customers were to immediately discontinue using and discard the identified affected lots of the product. Customers were to use alternative lots. In order to receive credit for the discarded product, customers were to complete the attached confirmation of receipt and destruction. If there were any questions, a question and answer section was provided but additional questions could be directed towards Therakos Customer Service Center at 1-800-828-6316, Option 2,1. |  | 4323 | 4,323 kits | Worldwide Distribution | Terminated | March 13, 2014 | 1136 | 4323
Z-2110-2011 | 58407 | No | KHE | Occult blood test. | Hematology | Hematology | 510(k) | ClearviewÂż iFOB, ClearviewÂż - Clearview iFOB Specimen Collection (accessory) |  |  | March 21, 2011 | 2011 | Alere North Americal | Potential for false negatives test results and negative/faint positive controls | Other | Alere sent a Voluntary Product Recall letter dated March 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to indicate zero on the attached Recall Reply Form, sign, and fax to 1-877-877-0147, if they have no remaining product.For any remaining product customers were instructed to discontinue use and indicate the quantity of kits in the attached Recall Reply Form so that the firm can arrange for them to receive credit. Customers were instructed to destroy any affected units at their site. For any further distribution of product, customers will need to be contacted. Physicians and hospitals were advised to retest patients with negative results. Questions relating to this Voluntary Product Recall call 1-877-597-9586. |  | 1800 | 1800 units | Nationwide Distribution | Terminated | March 07, 2013 | 717 | 15543
Z-2140-2011 | 58342 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Interlok 67 MM fixed Cruciate Tibial Plate with locking bar CO-CR-MO/TI 6AL 4V ALLOY, REF 141232, Sterile, BIOMET ORTHOPEDICS, WARSAW. IN I-Beam Tibia Plates are knee joint replacement prostheses intended for application with or without bone cement. |  |  | March 03, 2011 | 2011 | Biomet, Inc. | The firm is recalling the product due to the product's packaging being labeled "I-Beam 75mm. Tibial Plate", but the package actually contains the "Interlok 67mm. CR Tibial Plate". | Packaging process control | Biomet Orthopedics Inc initiated an electronic communication with Direct Accounts on March 3, 2011 and instructed them to quarantine and return all remaining/unused devices to the firm. For questions regarding this recall call 574-372-3983. |  | 8 | 8 | Worldwide Distribution -USA including FL, IN, and NC and the countries of Japan and the Netherlands | Terminated | July 27, 2011 | 146 | 8
Z-2105-2011 | 58255 | No | GZJ | Transcutaneous electrical nerve stimulator for pain relief. | Neurology | Neurology | 510(k) | Inter x 5002 Device, Inter x Professional Sport Device, Inter x 1000 Personal Device, and Inter x Personal Sport packaged with Triad alcohol wipes. The alcohol wipes manufactured by the Triad Group were labeled under brand names Triad, Select, Select Medical Products, PSS Select or PSS World Medical Inc. |  |  | February 07, 2011 | 2011 | Neuro Resource Group Inc | Product package contains alcohol prep pads recalled by Triad Group because the prep pads have the potential to be contaminated with an objectionable organism. | Material/Component Contamination | Triad Group issued a press release on 2/7/2011, and notified consignees by e-mail 3/2/2011. Neuro Resource Group, Inc. ( NRG ) removed the recalled Triad wipes from all kits in house and obtained non recalled wipes from an alternate supplier. For replacement wipes customers may contact NRG customer service at 1-877-314-6500. |  | 57252 | 57,252 wipes | Worldwide Distribution | Terminated | September 23, 2011 | 228 | 57252
Z-2112-2011 | 58273 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL90 FLEX Portable Blood Gas Analyzer System, Part Number: 393-090, 510 (k) #K092686. All units equipped with software versions: 2.4.1680.25 and all previous versions are subject to recall. |  |  | January 12, 2011 | 2011 | Radiometer America Inc | Following the receipt of two (2) reported customer complaints, the firm determined that the current operating software for their ABL 90 Analyzer does not prevent the further use of the sensor cassette when there is the possibility of iron rich plaque build-up in the reference electrode. If the device operator/user ignores the error message and continues to obtain biased measuring results using the | Software design | The firm, Radiometer America, sent two Urgent "Medical Device Field Correction" letters issued on February 4, and on February 24th, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to replace the sensor cassette. In order to prevent this situation from occurring, the firm informed the customers in the recall/correction notification that they intend to upgrade all ABL 90 Flex analyzers to a new software version. The new software will perform a check during a normal system FLUSH in order to allow the analyzer to detect the bias and will notify the customer immediately of this condition and will automatically disable the cassette thus preventing further use. The letter goes on to remind the device customers that until such time as the software upgrade is installed, they should make sure to adhere to the procedures stated in the Reference Manual which call for the replacement of the sensor cassette immediately upon receiving message #1310. The letter further instructs the customer to complete and return the attached FAX RETURN FORM via fax to 440-871-0463. The letter also states that the expected release date for the new software should be sometime in the second half of February 2011. If you have any questions regarding this letter, or the upgrade, please contact Technical Support staff at 1-800-736-0600 Opt 4. |  | 565 | 32 units were distributed in the U. S. and 533 units were distributed world-wide | Worldwide distribution: USA including states of CA, DC, PA, MN, NY, NJ, TX, OH, MO, WI, UT, and MI; and country of Canada. | Terminated | May 09, 2012 | 483 | 565
Z-2118-2011 | 53013 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON Antares 5.0 Base System, model number 10042832. Distributed by Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA. |  |  | May 15, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Shock Hazard. The manufacturer failed to set the system power switch to "ON" during the line level leakage and high pot current draw testing. | Release of Material/Component prior to receiving test results | Siemens Medical Solutions USA, Inc. sent a Customer Safety Advisory Notification letter out on May 15, 2009. Service engineers were to visit each customer to conduct necessary testing.For questions regarding this recall call 650-694-5993. |  | 1 | 1 unit | Worldwide Distribution - USA including CA, IL, MD, NC, OH, and UT and the countries of Australia, China, Denmark, Spain, South Korea, Norway, Poland, Russia and Sweden. | Terminated | April 29, 2011 | 714 | 30
Z-2016-2011 | 58188 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Shear Valve Assembly used on CELL-DYN 3500SL Analyzer and CELL-DYN 3500SL Analyzer (refurb). |  |  | March 15, 2011 | 2011 | Abbott Laboratories | The Nylon Washer has an outer diameter which is undersized resulting in the inability to generate patient results. | Pending | Product Correction letters were sent on March 15, 2011. via Federal Express. The letter identified the affected product, explained the problem, the patient impact and the necessary action that were to be taken. Customers were to review the letter and inform their Laboratory Manager of the pending action.; complete and return the Customer Reply Form; and save a copy of the letter for their records. If there were any questions, customers were to contact Customer Support at 1-877-4ABBOTT or their local hematology customer support representative for customers outside of the US. |  | 174 | 174 instruments total, all varieties | Worldwide Distribution -- AZ, CA, FL, IA, ID, IN, MA, MI, MN, MS, ND, NY, OK, OR, SC, SD, TX, WI, WY and Germany, Singapore, Mexico, China, South Korea, Argentina, Japan, Brazil, Canada, Australia, and Chile. | Terminated | November 22, 2011 | 252 | 174
Z-2122-2011 | 58482 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Flo-Gard 6201, Single Channel Volumetric Infusion Pump; Baxter Healthcare Corporation, Deerfield, IL 60015; Product Code # 2M8063 |  |  | March 24, 2011 | 2011 | Baxter Healthcare Corp. | Pumps submitted for service may not have been properly serviced before being returned to the customer. | Employee error | Baxter visited the direct accounts starting on 3/24/11 informing them that the affected Flo-Gard pumps may have been released from the preventative maintenance with incorrect accuracy testing. The affected pumps were located, taken out of service and placed in quarantine on the same date. A technician not involved with the previous testing visited both accounts on 3/25/11 and performed preventative maintenance service of the three affected devices, and were found to be within specification. |  | 3 | 3 units | Puerto Rico | Terminated | April 29, 2011 | 36 | 3
Z-2121-2011 | 58286 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | RETIC results may be invalid when the instrument user interface allows reflex processing of a pre-diluted specimen in the RETIC test panel. Beckman Coulter has validated pre-diluted specimen processing in CBC panels. RETIC test panel using pre-diluted samples has not been validated.IMPACT:The reticulocyte value is an indicator for assessing red cell production. Erroneous results could impact c | Other | The firm, Beckman Coulter, issued an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010, with PCA Response Form to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Remove all cassettes from the mix station prior to initiating the Diluent Dispense function. Do not initiate Diluent Dispense while specimens are being processed in cassette presentation.2)Do not use pre-diluted samples for RETIC counts. The instrument is validated for pre-diluted CBC's only, as indicated in the Instructions for Use (IFU).3)Verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and that the Body Fluid results are properly validated at the LIS.4) Use the lip on the reagent drawer to move the drawer as shown. Avoid touching the drawer slides while replacing containers. Clean all spills promptly to reduce the risk of exposure to blood borne pathogens.5)Complete and return the enclosed PRODUCT CORRECTIVE RESPONSE FORM within 10 days, share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Beckman Coulter will make corrections in future version software.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative. |  | 422 | 422 units in total (175 in US) | Worldwide distribution: USA (nationwide) and countries including:Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom | Terminated | September 02, 2014 | 1413 | 422
Z-2120-2011 | 58519 | Yes | NVZ |  |  | Cardiovascular | PMA | Medtronic, Kappa 600 |  |  | April 04, 2011 | 2011 | Medtronic Inc. Cardiac Rhythm Disease Management | A rare measurement lock-up condition that may (at a rate of approximately 1 in 18,000 devices) inappropriately trigger ERI/RRT in the Medtronic Dual Chamber Pacemakers: Kappa 600, 700, 800, 900, EnPulse, Adapta, Versa, Sensia. This is caused by a random lock-up of the measurement system hardware that may result in an incorrect battery voltage reading of zero. This issue does NOT impact battery lon | Component design/selection | The firm, Medtronic, sent an "Important: Medical Device Correction" letter dated April 2011 to its customers. The letter described the product, problem, and actions taken. Medtronic informed the customers that this issue does NOT impact battery longevity and does NOT require device explant. Currently, the device can be reset to normal operation by a Medtronic representative. Reset devices are no more likely to experience a recurrence of this issue. Medtronic is planning to release a programmer software update that will allow the clinician to reset the device. In addition, Medtronic has consolidated their product performance data with the vision of creating a primary reference tool for performance updates.If you have any questions, please contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 1989180 | 1,989,180 | Worldwide distribution: USA (nationwide) including Guam, Puerto Rico and DC; and countries including: Algeria, Andorra, Anguilla, Argentina, Armenia, Aruba, Australia, Austria, Bahamas, Bangladesh, Belgium, Bermuda, Bolivia, Brazil, Brunei, Darussalam, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Georgia, Germany, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Mexico, Morocco, Netherlands, Netherlands Antilles, New Zealand, Norway, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Virgin Islands (British). | Terminated | November 27, 2012 | 603 | 1989180
Z-2113-2011 | 58367 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Doctors Review System, OB-Windows/OB-View software release numbers 4.1 through 4.1.3, Digisonics, Inc., 3701 Kirby Drive, Houston, TX 77098 |  |  | February 09, 2007 | 2007 | Digisonics, Inc | Study would on occasion have fetal heart rate and/or cervical length on the report that was not entered nor did it belong to that patient. Software displayed incorrect data for fetal heart rate and cervical length in OB ultrasound reports. | Other | The firm, Digisonics Inc, notified the affected customers of software release via e-mail on January 24, 2007 and follow-up e-mail on February 7, 2007. The e-mails described the product, problem and the actions taken. The firm was able to trap the error and after this report the developer was able to determine the cause and a new executable was sent and installed at the customers site.If you have any question, please call 800-940-3240 or e-mail: service@digison.net. |  | 128 | 128 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | May 02, 2011 | 1543 | 128
Z-2119-2011 | 58230 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The 4D Integrated Treatment Console; Model Number: H51; Reference: 4DITC versions 8.6 and below; File Reference: CP-2007-0288; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | February 28, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly, in that the Multileaf Collimator (MLC) message displayed on the 4DITC on versions 8.6 and below, does NOT interlock beam delivery. | Software design | Varian Medical Systems sent an Urgent Medical Device Correction Letter on February 28, 2011, to all affected customers.Varian has created the procedure described in CTB-GE-579 to inform users on proper interfacing of the demountable BrainLAB m3 on the Varian Clinac. The procedure ( CTB-GE-579) is enclosed and posted to the customer site: http://www.MyVarian.com.This design fix in 4DITC v10.1 or later will be provided to all affected sites.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager.USA and Canada - 1-888-VARIANS (888-827-4265)Europe - + 41 41 749 8844 |  | 17 | 17 | Class 2 Recall - Worldwide Distribution -- USA including states of ; CA, NV, NJ, GA, IL, and TX and countries of Australasia, Belgium, Brazil, China P.R., Denmark, France, Germany, Germany ( BrachyTherapy) Hong Kong, India Pvt. Ltd, Italy, Japan, Middle East, Netherlands, Russia, Spain, Switzerland and United Kingdom. | Terminated | March 30, 2012 | 396 | 17
Z-2114-2011 | 58282 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | For Panels containing CBC: The RBC and HCT are improperly corrected and not flagged when there is a WBC Pump Failure and the WBC ~ 140x103/uL. For Body Fluid Panels:The RBC is improperly corrected and not flagged when there is a WBC Pump Failure. IMPACT:Erroneous RBC and/or HCT could affect the interpretation and management of patients at risk with various types of anemias. | Other | The recall communication was initiated on 10/20/2010, with Beckman Coulter forwarding a Product Corrective Action (PCA) letter with PCA Response Form (via US Postal Service for US customers) to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter informs the customers about the problems identified and provides an action/resolution to this issue. When a WBC Pump Failure occurs, the WBC or TNC results will be flagged with 'R' and the WBC Pump Failure System message. The RBC and/or HCT results associated with these affected analyses also require review. This will be corrected in a future version of software.Beckman Coulter asked consignees to complete and return the enclosed response form within 10 days and share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees local Beckman Coulter representative. |  | 422 | 422 units in total (175 in US) | Nationwide in the US and worldwide to: Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom | Terminated | September 02, 2014 | 1413 | 422
Z-1895-2011 | 57948 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coulter PrepPlus 2 (part # 378600) is a microprocessor-controlled pipetting and diluting system. |  |  | January 18, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated after erroneous, though potentially credible, results due to intermittent reagent level sense failures could be generated. During these failures, the instrument fails to produce a warning that a reagent vial is low or out of reagent. The low or absent levels of reagent used during these failures may lead to insufficient antibody being dispensed into a daughter tube there | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTION" letter dated January 18, 2011 to all their customers. The letter described the product, problem and actions to be taken. The customers were instructed to: 1) Follow PrepPlus 2 Instructions for Use (Part # 4277264-BA) for all reagent and specimen requirements prior to initiating a run on the PrepPlus 2 including:A) Ensuring all dead volume requirements for reagents as indicated in 7.11 and 7.12 of the Sample Processing sections and 12.2 of the Troubleshooting Guide. Appropriate volume of reagent in all vials per the planned number of tests to be run must contain at least the volume listed on the Worklist (plus an additional 10% of the full volume of the container to account for dead volume)".B) Ensuring specimen processing dead volume requirements are met as indicated in section B of Specimen Tube Size to ensure the minimum amount of specimen needed.C) Monitoring for any foaming/bubble accumulation in the reagent vial.2) A review of all clinically relevant results generated from samples prepared on the system be conducted. Abnormally dim or unexpected negative sample staining patterns should be reviewed to ensure correlation with clinical symptomatology and other diagnostic data. If the clinical information Is in accordance with the clinical data released, the validity of the results could be considered accurate and no retrospective action may be needed. Ifinconsistent data is observed, please adhere to good laboratory practices within your institution and consider data re-analysis and re-testing.The customers were also instructed to share this information with their laboratory staff and retain it as part of their laboratory Quality System documentation; if they have forwarded any of the affected product(s) listed above to another laboratory, provide a copy of the letter to them, and complete and return the enclosed PRODUCT CORRECTION RESPONSE FORM via fax to: (786)639-7500/7501/750 |  | 369 | 369 units (140 in the U.S.) | Worldwide distribution: USA and countries including: Australia, Barbados, Belarus, Belgium, Canada, Croatia, France, Germany, Ireland, Italy, Japan, Malawi, Mozambique, Namibia, Poland, Puerto Rico, Russian Federation, Slovenia, South Africa, Spain, Swaziland, Sweden, Switzerland, Ukraine, United Kingdom, Zambia. | Terminated | September 02, 2014 | 1323 | 369
Z-1885-2011 | 57743 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | MDI(R), S.A.L.T. Airway Device, Supraglottic Airway Laryngopharyngeal Tube, Intended for use with 6.5mm -9.00mm ET tubes, Including SALT(TM) airway, ET tube clamp, Latex-free securing strap, Tongue blade, Lubricating jelly, NON STERILE, REF 79-580, Microtek Medical, Inc., Columbus, MS 39702, www.microtekmed.com, Manufactured in the Dominican Republic of U.S. Materials |  |  | January 05, 2011 | 2011 | Microtek Medical Inc | Product contains Triad Lubricating Jelly which was recalled. | Other | The firm, Microtek Medical, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 17, 2011 to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and place any S.A.L.T. Airway Device (79-580) in quarantine; if they have distributed any of the product, they should identify their customers and notify them at once of this recall; complete and return the enclosed URGENT: MEDICAL DEVICE RECALL-RETURN RESPONSE form via fax to: (662)244-3011 or email: tom.bonner@ecolab.com, and return any unused product inventory to Microtek Medical. The firm will provide the customers with a replacement product as soon as product becomes available.If you have any questions in regards to this recall, please feel free to contact the V.P.-Regulator Affairs at 1-888-844-0988 or email: Tom.bonner@ecolab.com. |  | 2230 | 223 cases of 10 units | Nationwide distribution: AZ, CA, CT, FL, IN, MS, NC, NH, PA, TX and WI. | Terminated | June 20, 2011 | 166 | 2230
Z-2046-2011 | 58240 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac High Energy Medical Linear Accelerator, a Trilogy Radiotherapy Delivery System; Model #s: H14, H18, H27, H29, HCX; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | February 24, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The wedge angle labeling on the wedge tray may not correctly match the wedge body. However, the wedge body is labeled with the correct wedge angle. | Labeling mix-ups | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 24, 2011, to all affected users. The letter described the product, problem and actions to be taken. The customers were instructed to IMMEDIATELY inspect all wedges to verify that the wedge angle label on the wedge body matches the wedge angle on the wedge tray; if any wedge is identified with mismatched labeling follow the written instructions provided in the letter; always measure wedge data at installation of new or replacement wedges; verify all treatment accessories of "matched" linear accelerators, and please advise the appropriate personnel working in your Radiotherapy Department of the content of this letter.Varian has enhanced their manufacturing quality control process to further ensure correct labeling of wedges. Also, Varian will replace any wedge that is identified as mislabeled, at the request of the customer. If you require further clarification, please feel free to contact your local Varian Customer Support Service Manager at: 1-888-VARIAN5 (888-827-4265) for USA and Canada; +14 41 749 8844 Europe; North American: support-americas@varian.com; Australia/New Zealand:support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America:soporte.al@varian.com. |  | 5541 | 5541 units | Worldwide distribution. | Terminated | November 05, 2012 | 620 | 5541
Z-2048-2011 | 58331 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Percutaneous Sheath Introducer Kit |  |  | March 25, 2011 | 2011 | Arrow International Inc | Labeling discrepancy - the outer corrugate label as well as a card inside the package state the product contains dry natural rubber - this is not duplicated on the unit packaging. | Other | The firm, Arrow International, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated March 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine any products with the listed lot numbers; complete and return the Immediate Attention Request form via fax to 800-343-2935 indicating how much product they have in stock and whether they need labels. In the event that an alternative approach is needed, the customers are to contact Arrow Customer Service for more information. Once the labels have been received and applied the product can be released from quarantine.An Arrow's Sales Representative will pay a personal visit to their customers and obtain the quantity and the amount of labels they need. If you have any other questions, feel free to contact your local sales representative or Customer Service at 866-396-2111. |  | 840 | 840 | Nationwide distribution: USA including states of: PA and NJ. | Terminated | January 29, 2014 | 1041 | 840
Z-0836-2011 | 56024 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems FastPlan software, Versions 5.5 and 5.5.1, Model Number HZ3, manufactured by Varian Medical Systems, Palo Alto, CA |  |  | March 23, 2010 | 2010 | Varian Medical Systems Oncology Systems | Due to software Issue, an error occurs when performing cone planning using MRI, which may result in the patient receiving a higher radiation dose than expected, potentially to healthy tissue. | Device Design | Varian sent Urgent Medical Device Correction letters on March 23, 2010, by Federal Express to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to run test cases or otherwise determine whether the MR images that they receive have the 'Series Date' parameter populated.Customers should contact their MR system manufacturer to get a confirmation that the 'Series Date' field is populated.Customers were instructed to confirm by visual inspection that the fusion results are loaded appropriately in Cone Planning before starting to create or edit a plan.For questions regarding this recall contact your local Varian Customer Support Distirct or Regional Manager.Phone: USA and Canada - 1-888-VARIANS (888-827-4265) Europe -+41 41 749 8844E-mail: North America: support-america@varian.com Australia/New Zealand: support-anz@varian.com Europe: support-emea@varian.com South East Asia: seasia.apps.helpdesk@varian.com China/Asia: china.apps.helpdesk@varian.com Japan: Japan.Apps.Helpdesk@varian.com Latin America: soporte.al@varian.comInternet: Oncology Systems customer site - www.myvarian.com Varian Medical System public site - www.varian.com |  | 218 | 218 units | Worldwide distribution - USA nationwide and the countries of Australia, Brazil, Canada, Chile, Colombia, Georgia, Germany, Hong Kong, India, Ireland, Japan, Macau, Malaysia, Russia, Singapore, Spain, Sri Lanka, and Taiwan | Terminated | February 24, 2012 | 703 | 218
Z-2127-2011 | 58313 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Cirrus HD-OCT Model 4000 instruments, new and refurbished, distributed with, or upgraded to, version v 1.0 - 4.5, 5.0.0 or 5.1.0 |  |  | March 10, 2011 | 2011 | Carl Zeiss Meditec Inc | A software defect exists in versions 5.0.0 and 5.1.0 where when following a specific workflow, exam data may be saved under the name of the first patient | Software design | Carl Zeiss Meditec, Inc. sent an Field Correction Action letter dated March 9, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letters requested acknowledgement of the letter and a response and the fact that a follow up packet with revised software will be sent to all customers with V 5.0.0 or 5.1.0 software.For questions, please call 1-855-853-9461. |  | 6888 | 3465 US; 3423 OUS | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, LA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SC, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Algeria, Argentina, Australia, Bangladesh, Belgium, Boliva, Brazil, Canada, France, Chili, China, Columbia, Denmark, Ecuador,Germany, Great Britain, Greece, Guatemala, Hong Kong, India, Japan, Lebanon, Malaysia, Mexico, New Zealand, Pakistan, Peru, Singapore,South Korea, Spain, Sri Lanka, Taiwan, Uruguay and Venezuela | Terminated | March 29, 2012 | 385 | 6888
Z-2125-2011 | 58424 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA, RT Chart Application v 8.0 through 8.9 |  |  | March 09, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue was identified with the gantry rotation direction when creating opposing field within the RT Chart function of ARIA, or possibly with third party planning systems and delivering the plan at 4DITC. | Software design | Varian Medical Systems, Inc. Oncology Systems sent Urgent Medical Device Correction Field Safety Notice letters on March 9, 2011, to all affected customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letters identifeid the product, the problem, and the action to be taken by the customer. Customers were instructed to:Manually enter the gantry angel 0.0 when the plan indicates angel 0.0. Manual entry is only allowed to one decimal place, so the range that triggers the issue cannot be created if the keyboard is ued to enter the gantry angel.Users with third party plan sources should verify whether gantry angles with two or more decimal places are supported. If so, customers may be affected by this issue, and should take steps suited to the device.In all cases, the display at the Treatment console is indicative of what direction the gantry will move - users are advised to be especially attentive to these values, particulary when non-zero couch rotation values are used.Users are advised to always run a simulation of the plan, including all plan motions, at the treatment unit prior to first patient treatment. For further clarification customers were insturcted to contact their local Varian Customer Support District Manager. For questions regarding this recall call 650-424-6471. |  | 612 | 612 units | Worldwide Distribution - USA including CA, DC, FL, GA, MD, MN, MS, NY, OH, and TN and the countries of Canada, Australia, Austria, Bahrain, Bagladesh, Belarus, Belgium, Brazil, Columbia, Costa Rica, Cypress, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lithuania, Luxemborg, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Reunion, Romania, Russia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tobago, Tunisia, Turkey, Ukraine, Emirates, and the United Kingdom. | Terminated | March 13, 2012 | 370 | 612
Z-2128-2011 | 58287 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | Body Fluids results may be incorrectly interpreted by the Laboratory Information System (LIS). Body Fluid results ~ 1,000 celis/uL are reported using the US numeric separators to demark thousandths places (comma, ",") and partial units (decimal, "."). An LIS that uses non-US formats could interpret the comma as a partial units separator, which could lead to reporting erroneously low body fluid r | Other | The firm, Beckman Coulter, issued an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010, with PCA Response Form to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Remove all cassettes from the mix station prior to initiating the Diluent Dispense function. Do not initiate Diluent Dispense while specimens are being processed in cassette presentation.2)Do not use pre-diluted samples for RETIC counts. The instrument is validated for pre-diluted CBC's only, as indicated in the Instructions for Use (IFU).3)Verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and that the Body Fluid results are properly validated at the LIS.4) Use the lip on the reagent drawer to move the drawer as shown. Avoid touching the drawer slides while replacing containers. Clean all spills promptly to reduce the risk of exposure to blood borne pathogens.5)Complete and return the enclosed PRODUCT CORRECTIVE RESPONSE FORM within 10 days, share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Beckman Coulter will make corrections in future version software.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative. |  | 422 | 422 units in total (175 in US) | Worldwide distribution: USA (nationwide) and countries including:Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | September 02, 2014 | 1413 | 422
Z-2123-2011 | 58574 | No | LTI |  |  | Gastroenterology/Urology | PMA | Lap-Band Access Port Needle 10-Pack, 20 G x 51 mm (2 in.), Product Code: 8-20302-10, Lot Number: 9L01N |  |  | March 21, 2011 | 2011 | Allergan | The recall was initiated because Allergan has determined that a number of Access Port Needle pouches from one specific lot number may have small gaps in a pouch seal. Consequently the product may no longer be sterile. | Pending | Allegan sent an Urgent Medical Device Recall letter dated March 21, 2011, to all affected customers. Allergan requested that upon receipt of letter, that the customer check their existing inventory of product for affected lot. Customers were requested to complete Business Reply Card, stating quantity of inventory of affected lot, if any, and return to Allergan Medical. Customer with remaining inventory were instructed to contact Allergan Medical Product Support at 1-800-624-4261, x5972 or 1-805-961-5972 for return and immediate replacement. . Customers will be issued credit for returned product.For questions regarding this recall call 805-319-0589. |  | 1888 | 1878 x 10-packs | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, MN, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV and the countries of Australia, Canada, Germany, Great Britian, and the Netherlands | Terminated | April 06, 2012 | 382 | 1888
Z-2124-2011 | 58470 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂżVentralex" Hernia Patch Medium Circle with Strap Diameter 6.4cm /2.5",\ |  |  | April 11, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: BardÂż Medium Ventralex Hernia Patch 6.4cm /2.5" packaged in a pre-printedcarton identified for a BardÂż Ventralex Small Circle 4.3cm /1.7" | Labeling mix-ups | Davol sent an Urgent: Medical Device Recall letter on April 11, 2011, via FedeX Priority Overnight to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and remove and return the recalled lot. Distributors were requested to notify their accounts.For questions contact the Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Services &amp; Support Department at 1-800-562-0027. |  | 146 | 146 units | Nationwide Distribution - USA including AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, MA, MD, MI, MO, MT, NC, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, VA, WA, and WV. | Terminated | October 27, 2011 | 199 | 146
Z-2135-2011 | 58325 | No | HWD | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Low Profile Self-Tapping Bone Screw, 6.5 mm X 25 mm, Qty. 1, Biomet Orthopedics, 66 East Bell Drive, P.O. Box 587, Warsaw, IN 46581 USA. STERILE. |  |  | March 10, 2011 | 2011 | Biomet, Inc. | One lot of 30 mm screws was switched with one lot of 25 mm screws. | Mixed-up of materials/components | Biomet Orthopedics sent an Urgent Medical Device Recall Notice letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and remove the affected lots from their inventory, and return them to Biomet. Complete and enclosed FAX Back Response Form to 574-372-1683 prior to return of product. Distributors were instructed to notify all sub-accounts of the recall.Confirm receipt of notice by calling, 800-348-9500, extension 3755 or 3983.For questions regarding this recall call 574-372-3983, Monday through Friday, 8AM to 5PM. |  | 100 | 100 | Worldwide Distribution - USA including CA, GA, IA, IL, IN, KS, KY, MD, MI, MN, MO, NC, ND, NJ, NY, OK, PA, SD, TX, VA, WA, and WI and the country of Japan. | Terminated | February 17, 2012 | 344 | 100
Z-2134-2011 | 58402 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Patient Monitor, Model 91387. All modules may be affected except Models 90470, 90478, 90496, &amp; 91496. |  |  | March 24, 2011 | 2011 | Spacelabs Healthcare, Llc | The circuit that supplies backup power to the modules in the Medical Patient Monitor, Model 91387, is not working. If there is a power interruption, the three-minute backup may not be operational and customers may lose any changes they have made to the alarm settings. | Other | SPACELABS Healthcare sent an Urgent Medical Device Correction letter dated April 1, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to weigh the benefits versus the risks when deciding whether or not to continue to use their monitors until they can be upgraded. Customers were recommended to brief their staff about the possibility of losing changes to the alarm settings and to restore them after a power interruption. Field service engineers will contact customers to schedule a mutually agreeable time to come to their hospital and upgrade their affected patient monitors at no cost to the hospital. US Consignees can call 1-800-522-7025, Select 2 for Technical Support. Consignees outside of the US can call 1-425-657-7200 x5089 for any questions about this recall. |  | 4470 | 4470 units total (1894 units in US and 2579 units internationally) | Worldwide Distribution - USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, EGYPT, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PAKISTAN, PANAMA, POLAND, PORTUGAL, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, Slovakia, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, Republic OF CHINA, Thailand, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VENEZUELA.AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, EGYPT, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PAKISTAN, PANAMA, POLAND, PORTUGAL, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, Slovakia, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, Republic OF CHINA, Thailand, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VENEZUELA. | Terminated | February 15, 2012 | 328 | 4470
Z-2133-2011 | 58278 | No | OHR | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | Enforcement Discretion | Catheter Insertion Kit (containing, Under pad, PVP Pad, BZK Prep Pad, Sterile Lubricating Jelly, Vinyl Gloves and Poly Zipper Bag). |  |  | March 24, 2011 | 2011 | Apogee Medical, Inc | Kits contain PVP (Providone Iodine) prep pads that are being recalled by the Triad Group due to possible micro contamination, Elizabethkingia meningoseptica. | Other | The firm, Apogee Medical LLC, sent an "Urgent Medical Device Recall" letter dated March 23, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review their inventory; complete and return the enclosed Product Recall Response Form as soon as possible via fax to (919) 570-9611, mail to: Apogee Medical LLC, Youngsville Commerce Center, 90 Weathers Street, Youngsville, MC 27596, or e-mail to: dpeper@apogeemed.com; cease further distribution and return any product that is the subject of this recall to Apogee Medical LLC. All returned product must be labeled with a returned goods authorization label that has been provided with the letter. Apogee will replace returned product upon receipt of the Product Recall Response Form and identified product.In addition, if the customer distributed the product, they are to identify their customer(s) and notify them at once of this product recall. The notification should include a copy of the this recall notification letter and should instruct the customer to return the recalled product so that it can be returned to Apogee.Should you have any questions regarding this notification, please call the Manager Regulatory Affairs/Quality Assurance at (919) 435-5409. |  | 5503430 | 5,503,430 units | Nationwide distribution: USA including states of: AL, AZ, CA, FL, GA, IL, IN, KS, LA, MA, MO, NJ, NY, OH, OK, TN, TX, VA and WA. | Terminated | February 28, 2012 | 341 | 5503430
Z-2129-2011 | 58383 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | APD Disposable Sets and Extension Sets used with Baxter HomeChoice and HomeChoice PRO Automated Peritoneal Dialysis Systems; Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A. |  |  | March 16, 2011 | 2011 | Baxter Healthcare Renal Div | Clinicians are being reminded of the current warning in The HomeChoice APD Systems Patient At-Home Guide, 07-19-61-244 (October 2,2009), which states on page 3-13 "To reduce the risk of bums, electrocution, fire, or injury to persons: Close supervision is necessary when this product is used by, on, or near children or those unable to care for themselves." Tubing entanglement around the neck may r | Other | Baxter sent Safety Alert letters dated March 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to consider the following safeguards for at-risk patients on HomeChoice Therapy:Have a caregiver present while the patient is performing therapy.Move the HomeChoice cycler to the foot of the bed, if possible, to minimize the likelihood of entanglement that may lead to asphyxiation.Install bed railings for patients at risk for seizures or falls from bed.Complete the attached customer reply form and fax it to Baxter at t847-270-5457.For questions regarding this recall call 888-736-2543, option 2. |  |  |  | Worldwide Distribution - USA including AZ, CA, CT, DC, FL, KY, MO, NC, NM, Ny, OH, PA, TN, TX, and VA and the countries of Austria, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Iraq, Ireland, Israel, Italy, South Korea, Kuwait, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Portugal, Qatar, Russia, South Africa, Saudi Arabia, Singapore, Slovania, Spain, Sweden, Switzerland, Thailand, Taiwan, Tunisia, Turkey, United Arab Emirates, United Kingom, and Venezuela. | Terminated | March 08, 2012 | 358
Z-2136-2011 | 58052 | No | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | HealthPort BioAnologics Patient Electrode Sensor Test Kit, Part # SEN-0025-0, 25 kits/carton |  |  | February 08, 2011 | 2011 | Vermed, Inc. | Vermed kits contain a product component (Triad Alcohol Prep Pads) recalled due to potential Bacillus cereus contamination. | Other | Vermed sent consignee a recall notification letter on February 8, 2011, followed by a phone call and email. The consignee was instructed to respond with a Recall Reply Form via fax and to notify any customers to whom the affected product was further distributed. All recalled alcohol prep pads were to be returned to Triad Group. Firm is replacing Triad alcohol pads in kits with pads from another supplier and has removed Triad from its approved vendor list.For questions regarding this recall call 802-722-9976. |  | 1250000 | 2500 kits total; 500 from each lot except 1000 from Lot 203031 | OR | Terminated | July 25, 2011 | 167 | 1250000
Z-2138-2011 | 58449 | No | DTH |  |  | General &amp; Plastic Surgery | PMA | SharpointTM 4-0 Black Monofilament Non-Absorbable Nylon 18"/DS24mm Reverse Cutting Edged Needle. |  |  | March 30, 2011 | 2011 | Surgical Specialties Corp | Mislabeled - a number of units of the product labeled as 4-0 Black Mono Nylon suture product contain a 4-0 Black Braided Silk suture. | Mixed-up of materials/components | Surgical Specialites sent an Urgent Product Recall Notification letter on March 30, 2011, along with an attached Customer Acknowledgment Form to the affected consignees. It was reported that upon opening a box labeled as 4.0 Black Mono Nylon, it was observed that the package acutally contained 4-0 Black Braided Silk suture. Customers were to review their current inventory, segregate the affected lot number and / or contact their customer accounts that have purchased this specific item / lot code requesting a return of the product to customers' facility. To ensure customers have received this notification and to arrange for credit and the return of identified product, the Customer Acknowledgment Form should be completed and returned via fax to our Quality Assurance department at (610) 404-8488. |  | 2628 | 2,628 | Class 3 Recall - Worldwide Distribution-- USA and Australia. | Terminated | August 04, 2011 | 127 | 2628
Z-2137-2011 | 58166 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | NX is the radiographer's image identification and quality control tool. It offers a broad array of features developed especially to meet the needs of the radiographer. Standard tasks are completed quickly and effortlessly with a simple touch screen. In-room integration provides a more convenient workflow for the radiographer throughout the imaging process. NX communicates seamlessly with the hospital's picture archiving and communication system (PACS), radiology information system (RIS) and hospital information system (HIS). It is DICOM compliant and meets IHE guidelines. |  |  | March 04, 2011 | 2011 | AGFA Corp. | When a radiographer is working on two open studies on an NX workstation and, in specific situations, a problem of image/annotation mix-up can occur on the NX system. | Other | AGFA Corp. sent an "URGENT FIEL SAFETY NOTICE" letter via FED-EX to the customers on March 4, 2011. The letter describes the safety alert and mitigation. Customers were instructed to complete and return the Urgent Field Safety Notice Feedback Form by fax to 864-421-1664.For questions regarding this recall call 864-421-1815. |  | 775 | 775 units | Worldwide Distribution - USA including CA, GA, MS, ND, NM, NY, OK, PA, SC, SD, TX, and VA and the country of Canada | Terminated | March 27, 2012 | 389 | 775
Z-2139-2011 | 58326 | [6 empty] | ***REF IOSS413*** OSSEIOTITE Certain Implant 4 x 13mm*** Sterile using Radiation*** Rx Only***BIOMET 3i Dental Iberica S.L.WTC Almeda Park, Ed, 1, Planta 1 Pl. de la Pau s/n08940 Cornella de Llobregat (Barcelona) Spain*** |  |  | March 10, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i is recalling their product Osseotite Certain Implant Device IOSS413. The internal thread was not manufactured correctly. | Other | BIOMET 3i sent an Urgent Medical Device Recall Letter on March 14, 2011, to all affected customers via emial or fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers were told to contact Customer Service at 1-800-342-5454.Customers were instructed to respond with the attached Response Fax or email DomesticComplaints@Biomet.com.If customers still had implant(s) in their inventory, they were to return the implant(s) to BIOMET 3i using the following address:BIOMET 3iRegulatory Compliance Department4555 Riverside DrivePalm Beach Gardens, Florida 33410Incident No. CD117886A replacement implant(s) will be sent to them one the product is returned.For questions regarding this recall call 561-776-6906 (EST 8:00am to 5:00pm). |  | 217 | 217 | Worldwide Distribution - USA including AK, FL, GA, NC, NJ, NY, SC and the countries of Canada, Europe, Japan, and South America. | Terminated | January 25, 2012 | 321 | 217
Z-2149-2011 | 53211 | [6 empty] | 4Z1c Matrix Array Transducer, model number 10033682, used with ACUSON SC2000 ultrasound systems, model number 10433816, manufactured by Siemens Medical Solutions, Ultrasound Business Unit, Mountain View, CA. |  |  | July 10, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Burn risk-- due to coolant leaks, patients may suffer burns or blisters due to increased temperature. | Equipment maintenance | A customer safety advisory notification was sent out to customers on 7/10/09, and a mandatory software update will also be conducted. |  | 42 | 42 units | Nationwide. Foreign: Japan, Italy, South Korea, Singapore, Yemen, Germany, Australia, and Greece. | Terminated | May 05, 2011 | 664 | 42
Z-2143-2011 | 58480 | No | MJA | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BACTEC(tm) MGIT(tm) 5-Tube Spare AST Set Carrier, Bag of 3, Catalog 445943. ***Becton, Dickinson and Company |  |  | August 31, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic device components are mislabeled. | Other | BD Diagnostic Systems sent an Urgent Product Recall letter on August 31, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter advised customers to discard any carriers or labels that contain the incorrect barcode for replacement. Accounts were requested to complete an attached reply form to acknowledge their receipt of the notification and fax it to 410-316-4258.For questions regarding this recall call BD Customer Service at 1-800-675-0908. |  | 4123 | 4123 total labels | Worldwide Distribution - USA including CA, MA, MD, VA and the countries of Argentina, Colombia, Belgium, China, India, Hong Kong, Peru, Singapore, Thailand, and Trinidad. | Terminated | July 08, 2011 | 311 | 4123
Z-2145-2011 | 58305 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029. |  |  | February 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the DxH 800 may fail to meet the carryover specification for the WBC Differential (DIFF) when the WBC is greater than 70 X 10 to the 3rd power cells/micro L (70 X 10 to the 9th power cells/L). Beckman Coulter indicated the DIFF% carryover is unlikely to have clinical impact. | Other | Beckman Coulter sent a Product Corrective Action (PCA) letter with attached Recall Response Form via US Postal Service on February 11, 2011, to all customers who have the affected instruments. Consignees were informed of the reasons for recall, the products affected and that a resolution for these issues would be corrected in a future release(s) for the DxH 800. Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with your laboratory staff and retain this notification as part of your Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, Consignees were told to contact Beckman Coulter, Customer Service at (800) 526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Class 2 Recall - Worldwide Distribution -- USA, and the countries of : Canada, Australia, Macao, Belgium, Malaysia, Colombia, Myanmar, Croatia, Netherlands, Czech Republic, Norway, France, Philippines, Germany, Puerto Rico, Greece, Qatar,Hong Kong, Russian Federation, Hungary, Saudi Arabia, India, Singapore, Spain, Israel, taly Sweden, Japan, Switzerland, Korea, Turkey and United Kingdom. | Terminated | September 02, 2014 | 1299 | 566
Z-2146-2011 | 58181 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Single Day Infusor 2 mL/h Portable Elastomeric Infusion System; an Rx sterile disposable drug delivery system; Baxter Healthcare Corporation, Deerfield, IL 60015; REF 2C1071KJP |  |  | March 15, 2011 | 2011 | Baxter Healthcare Corp. | Some of the Single Day Infusor System 2 mL/h have a foil over-pouch that was incorrectly labeled as the Half Day Infusor System 5 mL/h, product code 2C1073KJP, lot 10H054. | Labeling mix-ups | Baxter Healthcare Corp. sent Urgent Product Recall letters dated March 15, 2011, via first class mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The accounts were requested to examine all of their inventory locations for both product code 2C1071KJP and 2C1073KJP, lot 10H054, and return the affected lot for replacement or credit. The accounts were also requested to complete the attached customer reply form listing the number of held for return, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Center for One Baxter at 1-800-422-9837. |  | 3696 | 3,696 units | Worldwide Distribution - USA including AL and the countries of Mexico and Spain | Terminated | March 12, 2012 | 363 | 3696
Z-2148-2011 | 58307 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029 |  |  | February 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because: The DxH 800 may fail to meet the carryover specification for NRBC% when the WBC is greater than 50 X 10 to the 3rd power cells/micro L (50 X 10 to the 9th power cells/L) Beckman Coulter indicated the NRBC% carryover is unlikely to have clinical impact. | Other | Beckman Coulter sent an Urgent Product Correction letter, dated February 11, 2011, via US Postal Service to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Consignees were informed the issues would be corrected in a future release(s) for the DxH 800. Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with their laboratory staff and retain this notification as part of their Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, consignees were told to contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Worldwide Distribution - USA including AL, CA, CO, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, and WI and the countries of Australia ,Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey, and the United Kingdom | Terminated | September 02, 2014 | 1299 | 566
Z-2147-2011 | 53051 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X300 Systems with software versions 3.0.00 and 3.0.02, model numbers 10348531, 10348532 and 10348533, manufactured by Siemens Medical Solutions, Mountain View, CA. |  |  | July 24, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Mistreatment risk-- when a user defined formula is used to calculate a value in a vascular exam, data is switched laterally when it is internally transferred and displayed on the summary report within the system. | Software design | A field correction was initiated and a Customer Safety Advisory Notification was sent out on 7/29/09. The letter identified the affected product and stated that there was a potential for user defined calculations to be incorrect within the vascular package when using the listed software versions of the product. The letter also discussed the steps that customers could take to avoid the potential risk of the issue. A mandatory update software was developed and will be installed free of charge. Customers should have been contacted to schedule a time for the installation. Questions were to be directed to the local service support person. |  | 154 | 154 units | Worldwide Distribution -- Austria, Brazil, China, Columbia, Germany, Denmark, Egypt, Great Britain, Honduras, Ireland, India, Italy Japan, South Korea, Mexico, Malaysia, Peru, Poland, Portugal, Romania, Russia, Sweden, Singapore, Thailand, Turkey, Taiwan, United States, Uruguay, Venezuela, and South Africa. | Terminated | May 09, 2011 | 654 | 154
Z-2152-2011 | 57044 | No | CBL | Carbon dioxide absorbent. | Anesthesiology | Anesthesiology | 510(K) Exempt | Allied Healthcare Products Inc. CarbO2Lime Carbon Dioxide Absorbent Granules for Canister Refill, Non-Hygroscopic Soda Lime USP-NF, Type: High Moisture (12% to 19%), Granule Size: 4 to 10 Mesh (4.8-2.0mm), packaged in 5-gal. pails, #55-01-0003; prepacked canisters (12 canisters/case), #55-01-0005; and replacement bags (12 bags/case), #55-01-0006, RX. The label shows the product is Manufactured by Allied Healthcare Products Inc., St. Louis, MO. |  |  | August 12, 2010 | 2010 | Allied Healthcare Products Inc | The granules do not absorb as long as expected in the anesthesia machines | Other | Allied Healthcare Product, Inc. issued recall letters dated August 10, 2011, via certified mail beginning 8/12/10 and/or via e-mail with a read receipt. Letters were issued for both distributors and end-use customers (such as medical facilities or pharmacies) explaining the reason for recall. Distributors were requested to conduct a subrecall using the enclosed customer letter for end users. A response form was included for both distributors and end users to complete to indicate the amount of product to be returned. The customer was instructed to contact the recalling firm for return instructions.For questions regarding this recall call 314-268-1661. |  | 317915 | 949 pails, 335/12-canister cases, and 706/12-refill bag cases | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, LA, MD, MS, NC, NE, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI and the countries of Canada and Grenada. | Terminated | March 28, 2012 | 594 | 317915
Z-1971-2011 | 57383 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Precision 500D X-ray systems with under-table collimator part number 5234960. |  |  | June 10, 2010 | 2010 | GE Healthcare, LLC | Use of high force when moving the image carriage/IDD into forward/park position can cause changes to the mechanical alignment of the collimator blades and allow the blades to open wider thanspecified, consequently allowing radiation outside the permitted field without warning to the user. Thisissue manifests with symptoms wherein the user cannot fully close down the lateral blades. | Radiation Control for Health and Safety Act | GE Healthcare sent an Importan Electronic Product Radiation Warning letter to customers. The letter discussed the product defect description and related hazards, affected product details, instructions, and product correction. The firm will send a GE Field Service Representative to visit each affected customer site to inspect and correct all potentially affected units in the field following the instructions in field modification instructions No: 10846.If you have any questions please call 800-437-1171. |  | 440 | 440 US | Nationwide Distribution | Terminated | September 15, 2014 | 1558 | 440
Z-2132-2011 | 58567 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron LXi 725, Part Number: 476501. |  |  | April 05, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has received reports of incorrect electrolyte results on Synchron LX systems due to a variety of maintenance-related hardware issues. | Equipment maintenance | The recall communication was initiated on 04/05/2011 with Beckman Coulter forwarding a Product Correction letter with attached customer response form to the affected customers via US Postal Service. The letter informed customers of the products affected, the issue, the impact, and the actions to be taken.Customers were instructed to follow the steps provided in order to address the issues related to microbial contamination; and to complete and return the enclosed response form within 10 days so that receipt of this important information could be confirmed. In addition, the letter requested customers share the information with their laboratory staff and retain the notification as part of your laboratory Quality System documentation.For customers who have any questions regarding this update, the letter indicated to contact Customer Support Center Call Center via the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/. They may also contact Beckman Coulter at (800) 854-3633 in the United States and Canada or through your local Beckman Coulter representative.Beckman Coulter will issue two additional recall notification ("UPDATE TO URGENT PRODUCT CORRECTIONSynchron LX20, LX20 PRO, LXi 725 Clinical Systems") beginning on 05/20/2011, one to Laboratory Directors and one to Physicians. The notices will be mailed via US Postal Service on Friday, May 20th. Customers who have subscribed to email notifications will receive an email containing the letters on Monday, May 23th. Both letter informs the Laboratory Directors and Physicians that Beckman Coulter has received reports of erroneous/aberrant sodium, potassium, chloride and calcium results that have been released from the laboratory. Electrolyte shifts can be low or high and may occur above, below or within the normal range.Laboratory Directors were instructed to immediately disseminate the Attachment to this letter (via notification/alert procedure appropriate to your |  | 1284 | 1284 units total for all products | Worldwide Distribution -- Australia, Belgium, Canada, China, Colombia, Czech Republic, Ecuador, Egypt, Finland, France, French Polynesia, Germany, Hong Kong, India, Ireland, Italy, Japan, Republic of Korea, Kuwait, Mexico, Netherlands, Oman, Panama, Philippines, Puerto Rico, Saudi Arabia, Singapore, Slovakia, South Africa, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and the United States. | Terminated | February 06, 2013 | 673 | 1284
Z-2153-2011 | 58318 | No | HFX | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Circumcision Clamp, models 1.1 cm and 1.3 cm, |  |  | January 26, 2011 | 2011 | Surgical Design Inc | Medical device used in surgical procedure is defective and may break during use. | Nonconforming Material/Component | Surgical Design, Inc. sent a Customer letter dated February 23, 2011, informing them of the affected product and providing instructions for its specific consignees. The instructions to the User is to please check your inventory to determine whether any Circumcision Clamps bought directly match those on your shelves. Please fill out the form and mail it to Surgical Design at the address on the form. If any clamps from the affected lots remain in your inventory, please ship them back to Surgical Design, noting the number being returned on the Reply form. These are the affected Lots.100210 (1.3cm)100211 (1.1cm) and100060409 (1.3cm)If you have any further questions, please call (703) 541-0196. |  | 6100 | 6100 clamps | Class 2 Recall -- Nationwide Distribution. | Terminated | June 21, 2013 | 877 | 6100
Z-2150-2011 | 58504 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens brand ACUSON S2000, Model 10434730, ABVS ultrasound systems; |  |  | March 28, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | The weld connecting the ABVS arm to the system base has the potential to fail, causing the arm to fall off of the unit. | Process change control | Siemens Medical Solutions USA, Inc. initiated the recall via phone calls and e-mail on March 28, 2011, to all affected customers, informing them of the affected product and providing instructions on the recall. Three units S/N 250171, 250173, and 250177 were to be removed and shipped back to Siemens. Siemens Service personnel were to be dispatched to the customer sites to perform validated stress test on units S/N 250157, 250162, 250163, 250164, 250166, 250167, and 250169 at customer sites.For questions regarding this recall call 650-694-5993. |  | 10 | 10 | Worldwide Distribution - USA, Portugal, Germany, Russia, Estonia, Brazil, France, and Saudi Arabia. | Terminated | May 06, 2011 | 39 | 10
Z-2151-2011 | 58441 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | LANTIS Commander, LANTIS Commander 6.1, LANTIS 8.3 Commander, LANTIS System, LANTIS 6.1 Complete, LANTIS 6.1 Essential, LANTIS 8.3 Complete and LANTIS 8.3 Essential. |  |  | March 23, 2011 | 2011 | Siemens Medical Solutions USA, Inc | There is a potential safety issue in regard to data transfer between LANTIS OIS system and the hospital Information System (HIS) when the HL7 (ADT) interfact is used. | Other | Siemens sent an Urgent Medical Device Correction letter on March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to include the Safety Advisory Notice in their LANTIS owner manual and contact there Siemens Service Engineer or Application Specialist or contact Lantis4.healthcare@siemens.com.For questions regarding this recall 925-602-8157. |  | 47 | 47 | Nationwide Distribution. | Terminated | April 19, 2012 | 393 | 47
Z-2154-2011 | 58379 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Enzymatic Creatinine (ECREA) Flex(R) reagent cartridges, K1270 for the Dimension Vista(R) 500 System |  |  | June 10, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Results may be falsely depressed up to 0.4 mg/dL or falsely elevated up to 0.6 mg/dL across the assay range. | Device Design | Siemens sent an Urgent Field Safety Notice letter dated June 2010 to all Dimension Vista(R) ECREA customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to:Discontinue use of ECREA method and transition to the Dimension(R) Jaffe creatinine (CREA) method.To notify anyone to whom they may have distributed the product.Complete and return the Effectiveness Check Questionnaire attached to the letter by fax to (302) 631-8467.For questions regarding this recall call (302) 631-6299. |  | 1399 | 1,399 | Nationwide Distribution - including AL, AZ, CA, CO, DE, FL, GA, ID, IL, IA, MD, MA, MI, MT, NH, NJ, NY, NC, OH, OR, PA, SC, TX, VA, WA, and WI. | Terminated | December 19, 2013 | 1288 | 1399
Z-2156-2011 | 58323 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Brightview XCT, 3/4" |  |  | March 28, 2011 | 2011 | Philips Medical Systems | The current Headholder to support Bright View XCT brain imaging allowed for imaging the brain but doe allow for imaging of the associated shoulder/neck area of the body. | Device Design | Customer information sheets will be provided to all customers when the field change order commences in April 2011.Philips will implement a full change order for the installation of the new headholder to replace the old headholder, free of charge, to all affected Philips BrightView XCT SPECT/CT units.For questions regarding this recall call 408-463-3000. |  | 1 | 1 | Worldwide Distribution - USA including US distribution to OH, MA, FL, VT and KY and the countries of Norway, Saudi Arabia, Belgium, Brazil, Japan, China, Poland, Austria, United Kingdom, Spain, Indonesia, Malaysia, Taiwan, Germany, United Arab Emirates, Netherland, Switzerland, Ireland, Czech Republic and Canada. | Terminated | June 22, 2011 | 86 | 50
Z-2157-2011 | 58308 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029 |  |  | February 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because: The handheld barcode scanner does not read barcode labels using NW 7 symbology.Impact: Specimens with NW7 labels cannot be read by the handheld scanner. | Pending | Beckman Coulter initiated a Product Corrective Action (PCA) letter with attached Recall Response Form sent via US Postal Service on February 11, 2011 to all customers who have the affected instruments. Consignees were informed of the reasons for recall, products affected and resolution for these issues would be corrected in a future release(s) for the DxH 800; but were also asked to take the following actions:If they use NW7 bar code symbiology, the labels can be read using one of the following methods:-At the single-tube presentation station (using the internal bar code scanner)-Enter the bar code information manually at the user interface of the single-tube station-Use cassette presentation.Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with their laboratory staff and retain this notification as part of their Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, consignees were told to contact Beckman Coulter Customer Service at 800526- 7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Worldwide Distribution - USA (nationwide) and the countries of Australia ,Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey, and the United Kingdom | Terminated | September 02, 2014 | 1299 | 566
Z-2179-2011 | 55466 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems Inc. SSH-880CV Aplio Artida; Diagnostic Ultrasound System |  |  | April 20, 2010 | 2010 | Toshiba American Medical Systems Inc | There is a possibility of the system shutting down owing to the power supply unit causing the protection function to activate thus turning off the system prematurely. | Device Design | The recall was initiated on 4/20/10 with Toshiba America Medical Systems (TAMS) forwarding a Medical Device Correction with a Customer Reply Form (via US Postal Service for US Customers) to all customers who purchased the SSH-880CV Aplio Artida Diagnostic Ultrasound System. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to avoid this problem until corrective action is taken, in case this problem occurs; please contact their Toshiba Representative. If emergency use is necessary, use the following procedure which allows the system to be used temporarily: 1. Turn OFF the breaker at the back of the system and confirm that the error code near the breaker disappears. 2. Turn ON the breaker again to supply power to the system. Customers were instructed to complete and return the attached form and fax it to the toll free number at the top of the form. This form can also be sent via email to raffairs@tams.com. If customers had any questions they can call Paul Biggins at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 14 | 14 | USA, including Puerto Rico. | Terminated | May 25, 2011 | 400 | 14
Z-2178-2011 | 55952 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter Access Free T3 Calibrators, Part Number A13430 in vitro diagnostic. |  |  | December 17, 2009 | 2009 | Beckman Coulter Inc. | Beckman Coulter has confirmed reports that the Access Free T3 calibrator lot numbers 989902, 989903, 989904, 989905, 989906 and 989907 will show a decline in performance before reaching their stated expiration date.The affected calibrator lots will cause one or more quality control samples to produce out-of-range high result. | Nonconforming Material/Component | Beckman Coulter sent a letter dated December 29, 2009, providing the customers with an explanation of the reason for recall and asked them to take the following actions:- Discontinue use of the lots of reagent listed above.- Review your historical quality control to ensure that the assay performance of these lots has been within your laboratory specifications.NOTE: Out-of-range high quality control results may indicate invalid test results.Consignees were asked to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation; and complete and return the enclosed response form within 10 days so that Beckman can be assured that they received this important notification.Assistance or questions regarding this notification, were directed to TechnicalSupport at 1-800-854-3633 in the United States and Canada. Outside the United States andCanada, contact consignees were told to contact their local Beckman Coulter Representative. |  | 2320 | 2,320 kits | Distribution Nationwide and Canada. | Terminated | March 29, 2012 | 833 | 2320
Z-2169-2011 | 57446 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Statspin Express 4 Centrifuge |  |  | August 18, 2010 | 2010 | StatSpin, Inc., dba Iris Sample Processing | Rotors may fail due to wire spokes breaking | Device Design | Iris notified customers by letter in August 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to install the new rotors and discard the older rotors with the wires. There was no need for customers to return the rotors, but the firm requested customers fax or e-mail the enclosed form to verify that they have installed the new rotors. For any questions customers should call 1-800-782-8774. |  | 2090 | 2090 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Australia Belgium, Czech Republic, Finland, France, Greece, Kuwait, Norway, Turkey,UK, and Norway. | Open, Classified |  |  | 2090
Z-2168-2011 | 57594 | [6 empty] | Philips FloTrak Elite modules: |  |  | January 04, 2011 | 2011 | Respironics Novametrix, LLC. | Software: two displayed respiratory parameters are incorrect due to errors in the system software: MValv (alveolar minute ventilation) and Vt/kg (tidal volume/patient weight). In both instances, the displayed parameters are higher than actual. | Software design | Respironics issued Field Action Notification Letter via Fed'x on 1/4/11, notifying customers of the issue and to instruct them on conditions for continued use and return of the device. The device will be updated with revised software to correct the calculation errors. Any questions or need additional information about this recall action, Contact Respironics US Customer Service at 1-800-345-6443 for customers residing in the USA or +011 (203)697-6488 for International customers. |  | 38 | 38 units | NationwideForeign: Belgium, China, Columbia, Egypt, England, India, Italy, Saudi Arabia, and Spain | Terminated | March 07, 2013 | 793 | 220
Z-2164-2011 | 57472 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | S5-7606 Silicone Oil injection Cannula 6mm Long - 20G |  |  | December 17, 2010 | 2010 | Becton Dickinson and Company | Restricted or no fluid flow for Visitec High Viscosity Injector. | Component change control | Beaver Visitec (BVI) notified customers via telephone and by sending an Urgent Recall Customer Notification letter dated December 7, 2010. The letter identified the product, the problem, and the action to be taken by the customer. ,Customers were requested to return all unused product in their inventory for credit or refund and complete the attached Customer Response Form and return it by fax to 781-893-7957. Questions can be directed to Inside Sales Manager at 781-906-7952. |  | 1000 | 1000 units | Worldwide Distribution - USA (nationwide) and the countries of Israel, Jordan, Russia, Saudi Arabia, South Africa, Turkey, Iraq, Lebanon, Morocco, Tunisia, Syria, India, Argentina, Australia, Austria, New Zealand, China, UK, Germany, Russia, Belgium, Italy, France, Sweden, and Ireland. | Terminated | July 03, 2013 | 929 | 19129
Z-2172-2011 | 58234 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, |  |  | February 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because: Specimen tubes may fall out of a Type A cassette (PN A53551, kit of 20) when the analyzer inverts the cassette in the mixing station, or when an operator manually inverts a cassetteThe impact:- Specimen tubes may fall out of cassettes and be misplaced inside the instrument- Specimen tubes may break upon impact, causing possible biohazard exposure- Patien | Pending | Beckman Coulter initiated a Product Corrective Action (PCA) letter with attached Recall Response Form being sent via US Postal Service on February 11, 2011 to all customers who have the affected instruments. Customers were informed of the reasons for recall, products affected and resolution for these issues would be corrected in a future release(s) for the DxH 800; but were also asked to take the following actions:-Periodically inspect the mix station drip tray for misplaced tubes that may have fallen from the cassette.- Periodically examine the Type A cassettes to determine if the expandable grippers (clips) within the lavender-colored insert are stuck in the open position (tubes will be loose). Consignees were told to contact their Beckman Coulter representative for assistance with questionable grippers.- Clean any spills according to their laboratory standard operating procedures. Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with their laboratory staff and retain the notification as part of their Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, consignees were told to contact Beckman Coulter Customer Service at 800-526- 7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Worldwide Distribution - USA (Nationwide)and to the countries of Australia, Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey, &amp; United Kingdom | Terminated | November 21, 2012 | 649 | 566
Z-2165-2011 | 58060 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Instrumentation Laboratory GEM Premier 4000 |  |  | February 14, 2011 | 2011 | Instrumentation Laboratory Co. | Measurement (amperometric) spike can occur early in cartridge life on the glucose and lactate sensors during patient blood analysis, leading to erroneously highresults | Device Design | Instrumentation Laboratory notified accounts by a Field Safety NotifIcation Letter dated Feb 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use whichever of the following mitigation options works best for their institution. Both options will have the same effect to eliminate the potential impact of a rare spike on the glucose and lactate sensors. Customers were instructed to complete the enclosed Response Tracking Form and return immediately fax to 781-861-4207 or e-mail to ra-usa@ilww.com.Customers were advised to contact their local representative with any questions. |  | 4379 | 4379 units | Worldwide Distribution - USA (nationwide) and the countries of CANADA ISRAEL ITALY APAN KUWAIT MALAYSIA MEXICO NETHERLANDS NORWAY PANAMA POLAND PORTUGAL PUERTO RICO SAUDI ARABIA SLOVENIA SPAIN SWEDEN SWITZERLAND TAIWAN TURKEY UNITED ARAB EMIRATES UNITED KINGDOM and URUGUAY. | Terminated | May 23, 2012 | 464 | 4379
Z-1678-2011 | 58047 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | 2.0 mL Cartridge/Syringe Piston component of Insulin Pump |  |  | February 25, 2011 | 2011 | Animas Corporation | Action is being undertaken to address the occurrence of adverse events associated with leaking cartridges. Firm became aware of this issue through customer complaints from pump users reporting the cartridges were leaking. | Material/Component Contamination | Animas Corporation sent an Urgent Recall letter dated February 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Change their cartridge immediately using the instructions provided. Return all of the unused affected product using the enclosed box. Shipments must be returned using the United States Postal Service. For technical questions call Animas Customer Support at 1-855-254-5668 or call the toll free number at 1-8777-937-7867.For any questions about returning the product call 1-877-280-2339.Customers were instructed to notify the firm by logging on to the following website: www.animasnotification.com or call the Call Center at 1-877-280-2339 if they have no product to return. |  | 384180 | 384,180 | Worldwide Distribution - USA (nationwide) and France. | Terminated | April 03, 2014 | 1133 | 384180
Z-2182-2011 | 58466 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | KCI RotoProne Critical Care Therapy System buckles for securing patient to the surface during proning therapy, Model 209500. |  |  | February 18, 2011 | 2011 | KCI USA, Inc. | Buckles used to secure patient to surface during proning therapy can come open if not properly engaged by user. | Device Design | Firm sent a letter with the "Warning: Improper Insertion of the RotoProne Therapy System Buckle Can Lead To Unintentional Bed Exit" to all consignees on 2/23/2011. The letter identified the affected product and the reason for the recall. Customers were instructed to read the attached product warning regarding the proper method to attach the RotoProne buckles and to post it with the device to ensure continued proper use of the buckles. If customers had questions regarding the proper use of the buckles, they were to contact KCI Customer Technical Service at 1-800-275-4524. |  | 229 | 229 units | Worldwide Distribution - USA and Kuwait. | Terminated | October 24, 2011 | 248 | 229
Z-2176-2011 | 58252 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03976; Model Number: HZ1; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | March 03, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The anomaly that has been identified with the Optical Guidance Platform (OGP) software may not be always displaying the correct transfer date on the patient file. | Software design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the Optical Guidance Platform (OGP) Instructions For Use on Loading a treatment plan; before positioning the patient, compare the plan details loaded into the OGP with the plan details printed from the treatment planning systems; verify that the patient information, isocenter, gantry and couch angles are identical as requested in the Instructions For Use, and advise the appropriate personnel working in their radiotherapy department of the content of this recall letter.Varian has made changes to the OGP Instructions For Use -Varian strongly recommends you download the most recent version at www.MyVarian.com or request a copy from your local Varian Customer Support District or Regional Manager. Varian also is preparing a software correction which will be applied to affected unites to correct this issue.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager at the follow: USA and Canada - 1-888-827-4265; Europe +41 41 749 8844; E-mail-North America: support-americas@varian.com; Australia/New Zealand: support.anz@varian.com; Europe: support-emca@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com, and Latin America: soporte.al@varian.com; Internet: Oncology Systems customer site - www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 599 | 599 units | Worldwide distribution. | Terminated | June 19, 2012 | 474 | 599
Z-2175-2011 | 58182 | No | NSX |  |  | General Hospital | Contact ODE | Philips IntelliVue Clinical Information Portfolio software as follows: |  |  | March 16, 2011 | 2011 | Philips Healthcare Inc. | Edits to an order can result in unintended scheduled interventions and a patient may receive additional interventions or medication doses that are not intended. | Software design | Philips Healthcare notified accounts on 3/16/11 via the Urgent Medical Device Correction Notification. Customer notifications in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography. Customers will be asked to follow the guidelines in the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notice until they receive their software correction. The correction will be in the form of a software upgrade and will be performed by Philips service personnel.All customers potentially affected by this issue will receive the software correction as part of the field action. |  | 387 | 387 units | Worldwide distribution: USA, Canada, Australia, Austria, Belgium, China, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Italy, Kenya, Latvia, Malaysia, Netherlands, Norway, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 387
Z-2180-2011 | 58395 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Digisonics, DigiView and ERS software release 3.6.4.4 through 3.6.5.2+, 2008 Digisonics, Inc., 3701 Kirby Drive, Houston TX 77098, 800-940-3240. |  |  | January 20, 2009 | 2009 | Digisonics, Inc | Data from one patient's study had the potential for being put into another patient's report. | Device Design | Digisonics notified end users via e-mail January 20, 2009 and January 21, 2009, that Data from one patient's study was being put into another patient's report. The problem occurred when a System bug was found in DigiView Release 3.6.4.4, and higher versions in which under certain circumstances (specific workflow) a failure to switch to a patient's images occurs. Two software patch versions will be made available to sites on or after February 5, 2009, to correct the reported problem.If you have any further questions, please contact Digisonics Customer Support at (800 ) 940-3240. |  | 70 | 70 units | Worldwide Distribution -- USA including States of : OH, NC, LA, TX, CA, MO, NY, MD, MA, DC, IA, AZ, PA, MI, NC, FL, IN, SC, ID, MN, AR, OK, IL, CT, and countries of Canada, Lebanon, and New Zealand. | Terminated | May 31, 2011 | 861 | 70
Z-2166-2011 | 58400 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DigiView software release 3.4.4.4 through 3.5.5.2. , 2008 Digisonics, Inc, 800-940-3240. |  |  | November 04, 2008 | 2008 | Digisonics, Inc | Failure of software to update some measurements when tracings or points are amended after measurement calculations have been performed. | Device Design | Digisonics, Inc., notified consignees by telephone and e-mail on 10/17/2008, informing them that a problem had been found in the Save Linked Images functionality of DigiView Release 3.6 (and all releases which start with 3.6 such as 3.6.1.). The problem is the failure of the software to update some measurements when tracings or points are amended after measurement calculations have been performed. A utility program, called Linked Image Measurement Verify was being supplied to all sites as a patch to the DigiView program. Release 3.6.6 was to be provided to all customers having previous 3.6 version. Problem was fixed in version 3.6.6 and 3.7.1.If you have any questions, please contact Digisonics Customer Support at (800) 940-3240. |  | 70 | 70 units total | Worldwide Distribution -- USA including States of : TX, OH, NC, MA, NY, IL, FL, OK, MO, ID, AZ, AR, MI CA, MD, and country of Canada.*** Due to the age of this recall, district is recommending no audit checks*** | Terminated | May 31, 2011 | 938 | 70
Z-2181-2011 | 58590 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical System TrueBeam system for stereotactic radiosurgery and radiotherapy. Model number H19; |  |  | April 05, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Position sensor failure mode may result in an inaccurate position calculation. 1. The video returned by one of the two cameras inside the Spectra is all white or all black, and the Spectra stops tracking. Or 2. Certain video intensities are not available on one of the two cameras. The image has abnormally abrupt transitions from dark to light portions, without the normal shades of gray. | Nonconforming Material/Component | Varian Urgent Medical Device Correction, Urgent Field Safety Notice letters were sent to all affected facilities by certified mail on April 5, 2011. |  | 20 | 20 | Nationwide distribution. | Terminated | June 19, 2012 | 441 | 20
Z-2170-2011 | 58168 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Beckman Coulter HbA1c APT (Hemoglobin A1c, Whole Blood Application), Part Number: OSR61177, Lot Numbers: 9390 &amp; 9472 |  |  | March 01, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has identified a problem with two lots of HbA1c APT reagent OSR61177.When HbA1c APT reagent lots 9390 and 9472 are calibrated with HbA1c calibrator ODR3032,elevated recovery may be observed for patient samples. Falsely elevated HbA1c values may beinterpreted as increased risk for diabetes or poorly controlled diabetes. Consequent medicalacti | Pending | Beckman Coulter initiated the recall communication with a Product Corrective Action (PCA) letter with attached recall response form on March 1, 2011 to the affected customers. The letter explained the problem and identified the action to be taken. The PCA also instructed the consignees to:(1) Please discontinue the use of lots 9390 and 9472 and discard according to your local procedures.(2) Please contact your local Beckman Coulter support Organization for replacement reagent.(3) You may need to re-evaluate samples for HbA1c which were tested with either of the listed lots of reagent.In addition, customers were instructed to:(1) share the information with their laboratory staff and retain this notification as part of their Quality System documentation. If any of the affected product was forwarded to another laboratory, a copy of the letter should be provided to them.(2) complete and return the enclosed response form within 10 days so that the firm was assured the customer received the important communication. For any questions regarding this notice, customers were instructed to contact Customer Technical Support at 1-800-854-3633 in the U.S.A and at 00353656831111 in Europe. |  | 373 | 373 kits | Worldwide Distribution - USA (Nationwide) Austria, Croatia, France, Germany, Korea, Philippines, Russian Federation, Thailand, Venezuela, and Vietnam. | Terminated | June 28, 2012 | 485 | 373
Z-2177-2011 | 58380 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension(R) EZCR Flex(R) Reagent Cartridge. Catalog number DF270B. In vitro diagnostic test for the quantitative measurement of creatinine in human serum, plasma and urine on the Dimension Vista(R) System. |  |  | June 10, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Patient results may be falsely depressed or falsely elevated across the assay range. | Device Design | Siemens issued an Urgent Field Safety Notice dated June 2010 to all Dimension(R) EZCR customers; to notify them that Siemens confirmed an issue which many cause EZCR results to be falsely depressed or falsely elevated across the assay range. The letter asks users to discontinue use of this lab test kit. |  | 860 | 860 | Nationwide distribution. | Terminated | March 19, 2013 | 1013 | 860
Z-2171-2011 | 58216 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Ventricular Bolt Monitoring Kit W/Cranial Access, Drain, 1104HMC, Integra LifeSciences, 3498 West 2400 South, West Valley City, Utah 84119. |  |  | February 22, 2011 | 2011 | Integra LifeSciences Corp. | Outer box of Monitoring Kit was mislabeled with the incorrect part number. | Error in labeling | Integra LifeSciences Corp. sent an Urgent Product Recall Notification letter dated February 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to quarantine any affected product and arrange for return to Integra by contacting Integra Customer Service Returns at 1-800-654-2873 (choose option 2) to obtain a Return Merchandise Authorization number and place an order for replacement product. Customers were asked to complete the attached "Recall Acknowledgement and Return Form" indicating whether or not they possess any of the product identified. For any questions regarding this recall call 609-936-2485. |  | 45 | 45 kits | Nationwide Distribution including CA, CT, GA, HI, IL, KY, MA, NC, NJ, NM. NV, OH, TX, WA, and WI. | Terminated | May 10, 2011 | 77 | 45
Z-2173-2011 | 58180 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NEXGEN COMPLETE KNEE SOLUTION FEMORAL AND PROVISIONAL IMPACTOR/EXTRACTOR, NONSTERILE, REF 00-5901-026-00, Zimmer UK, Ltd., Zimmer Warsaw, IN. |  |  | March 14, 2011 | 2011 | Zimmer Inc. | The firm is initiating a recall due to the potiential for the spring clip to break off during use and fall into the surgical site. The firm has received 24 complaints of the spring breaking, four of which were reported that the patient had to undergo additional x-rays to ensure the fractured clip had not been left in-vivo. | Device Design | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE RECALL" notice dated March 14, 2011 to its Consignees/customers. The letter described the product, problem, and the actions to be taken. The customers were instructed to stop using the device and quarantine immediately; carry out a physical count of all affected product in their territory and record this data on the Inventory Return Certification Form and fax to (574) 372-4265; ensure they have provided acknowledgement via "email return receipt" upon receiving notification; return the recalled product along with the completed Inventory Return Certification Form and Sterilization Certification Form to: Zimmer Distribution Center, Attn: Product Service, 1777 West Center Street, Warsaw, IN 46580-Note: The instrument must be returned even if the spring clip is not missing from the instrument; and if they distributed these affected lots further, provide a copy of this letter to the those customers when they implement these recall instructions. The letter stated that a new or refurbished instrument with a redesigned spring clip will be provided as replacement for the recalled device. For shipping assistance, questions or assistance in notifying your accounts about the recall, please contact Zimmer, Inc. at 1-800-613-6131. |  | 802 | 802 devices | Worldwide distribution: USA (nationwide)including states of: HI, CA, NY, MI, IN, GA, FL, WI, MN, NJ, IL, VA, TX, NC, PA, MS, WA, OR, OH, KS, TN, NH, AZ, NV, UT and AK and countries of : Australia, Canada, China, Germany, India, Japan, Singapore and Sweden. | Terminated | August 28, 2013 | 898 | 829
Z-2184-2011 | 58583 | No | CFL | Human growth hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IGFBP-3 ELISA Reagent Kit; Part Number" DSL-10-6600; Lot Numbers: 992743, 992287, 991916, 991530, 991432, 990220 &amp; 890830 |  |  | March 30, 2011 | 2011 | Beckman Coulter Inc. | The recall was was initiated because patient results with the IGFBP-3 ELISA kit may be elevated by 70% compared to stated expected ranges in the Instructions for Use (IFU). | Process control | Beckman Coulter sent a Product Corrective Action (PCA) letter with attached customer response form on March 30, 2011 to all customers who have the affected product, IGFBP- 3 Enzyme-Linked immunosorbent Assay ( ELISA) kit .This kit provides materials for the quantitative measurement of IGFBP-3 in serum. The letter informed the customers of the reason for recall, the products affected and actions to be taken.Consignees were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. If they forwarded any of the affected product listed above to another laboratory, they were asked to provide a copy of the letter to them.Beckman further asked consignees to complete and return the enclosed Response Form within 10 days so they are assured consignees have received this important communication.Questions were directed to the Beckman Coulter Customer Support Center.- Website: http://www.beckmancoulter.com/customersupport/support- Phone, call 1-800-854-3633 in the United States and Canada- Outside the United States and Canada contact local Beckman Coulter representative. |  | 1212 | 1,212 units | Worldwide -- USA, Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, France, Germany, Greece, Hong Kong, India, Italy, Korea, Republic of, Lebanon, Luxembourg, Morocco, Netherlands, New Zealand, Puerto Rico, Syrian Arab Republic, Taiwan, Thailand, Turkey &amp; United Kingdom | Terminated | December 02, 2011 | 247 | 1212
Z-2183-2011 | 58479 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | TEMPUS IC PROFESSIONAL Patient Monitor, Part No. 00-1002, Manufactured by: Remote Diagnostics Technologies Ltd., The Old Coach House, Farleigh Wallop, Hants, UK RG25 2HT. |  |  | March 25, 2011 | 2011 | Remote Diagnostic Technologies Ltd. | It was discovered that when a 12 Lead ECG is recorded and monitoring is quickly restarted, the ECG can cease functioning. | Component change control | Remote Diagnostic Technologies, Ltd notified its consignees of the problem and the work-around via letter with the subject line "ECG Issue - Field Action Required" on 03/25/2011. The letter stated that customers can avoid the issues by viewing the recorded ECG for at least 13 seconds. Two diagrams were attached with the letter depicting the conditions under which the issue can occur. Customers can continue to use the device normally in all other respects. The letter states that the firm is working on a software modification which will permanently fix the issue and prevent it from occurring. Customers are to contact the Regulatory Affairs and Operations Manager if they have any questions. |  | 19 | 19 units | Nationwide Distribution -- KY, MD, VA, and FL. | Terminated | September 16, 2011 | 175 | 19
Z-2188-2011 | 57194 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 with Arterial Blood parameter module Catalog number 500A , Terumo Cardiovascular Systems Corporations, Ann Arbor, MI. |  |  | November 01, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The CDI 500 monitors may display inaccurate values. These errors include inaccurate potassium, CO2, hematocrit, O2 saturation, pH, hemoglobin and other miscellaneous inaccurate values. | Device Design | The firm, TERUMO, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 27, 2010 to its customers. The letter described the product, problem and actionst to be taken. The customers were instructed to reveiw this notice; assure that customers are ware of this notice; confirm receipt of this notification by faxing the attached Customer Response Form, and place the attached stickers in a visible place in the Operators' Manuals, such as the inside of the front cover, if additional stickers are needed, the customers should contact Terumo CVS Customer Service at 1-800-521-2818. The letter states that when a displayed value is significantly different from expectations based on the clinical situatioin, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operatros manual.The letter also states "Please note that we are not recommending that you stop using your CDI monitor system."If you have any questions or concerns, please contact Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours are from Monday-Friday, 8AM - 6PM EST. |  | 3260 | 3260 all products | Worldwide distribution: USA (nationwide) and countries of : Australia, Barbados, Belgium, Brazil, Columbia, Costa Rica, Dominican Republic, Ecuador, El Salvador, Guatamala, Hong Kong, Honduras, Jordan, Japan, Libya, Mexico, Malaysia, Panama, PH, Saudi Arabia, Singapore, South Africa, Thailand, Trinidad, Taiwan and Uruguay, Germany, Saudia Arabia, Paraguay, South Korea, Dominican Republic, Phillipines, South Africa and UAE. | Terminated | July 31, 2014 | 1368 | 3260
Z-2190-2011 | 58517 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE / RT Therapist v4.1 system. |  |  | March 24, 2011 | 2011 | Siemens Medical Solutions USA, Inc | When using the ARTISTE Linac and RT Therapist v4.1 along with an imported cropped reference image, the image maybe processed incorrectly, thus resulting in a mistreatment. | Pending | Distribution of the Urgent: Medical Device Correction Customer Advisory Notice began on March 24, 2011 via certified mail or by the Siemens Customer Service Group. The letter identified the issue and when it occurs. Users were asked to verify the isocenter position of the reference image in the Portal Review application before using a reference image imported from the treatment planning system. Siemens is currently working on a solution and will notify customers as soon as the update is available. The letter should be included with the Owner Manual of the system in the chapter for "Safety Advisory Letters" until an update is provided. |  | 228 | 228 active devices | Worldwide Distribution | Terminated | February 13, 2012 | 326 | 228
Z-2191-2011 | 58219 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Forteo [teriparatide (rDNA origin) injection] black starter kits containing Triad alcohol pads. The starter kits look like insulated lunch boxes and contain alcohols pads, needles, instruction guides, and freezer gel packs. Note: the kits do not contain the drug, Forteo. |  |  | March 17, 2011 | 2011 | Eli Lilly and Company | Eli Lilly and Company announced 3/17/2011 that patients should not use the alcohol prep pads made by the Triad Group that are contained in the black starter kits for Forteo [teriparatide (rDNA origin) injection] in the United States.The Tri-ad Group is recalling the alcohol prep pads due to potential contamination with the bacteria, Ba-cillus cereus, which could result in life threatening infecti | Material/Component Contamination | The firm, Eli Lilly and Company, sent two letters entitled "URGENT MEDICAL DEVICE RECALL TRIAD GROUP ALCOHOL PREP PADS INCLUDED IN FORTEO STARTER KITS ALL ALCOHOL PREP PAD LOTS" to its customers. The first letter was dated March 18, 2011 and second follow-up letter was dated March 28, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to examine their inventory of Forteo starter kits; if they have any Forteo starter kits with a black bag, remove and appropriately discard the Triad Group alcohol prep pads; discontinue providing the product with the starter kits; inform any patients that may have received Forteo starter kits containing the Triad Group alcohol prep pads to discontinue use; note: after removing the Triad Group alcohol prep pads, all other components of the Forteo starter kits are acceptable for use: travel bag, pen case, gel packs, needles and instructional materials, and complete and return the enclosed business reply card. "It is important to note that this recall does not affect or involve the Forteo delivery device."For questions pertaining to the recall, contact Triad Group at 262-538-2900. (Please refer to the Triad Group press release issued on January 5, 2011 on the FDA website.) If you have product specific medical questions related to Forteo please call 1-800-LILLYRX. |  | 47000 | approximately 47,000 kits | Worldwide distribution: USA (nationwide) and countries of: Czech Republic, Germany, Slovak Republic, and Switzerland. | Terminated | August 30, 2011 | 166 | 47000
Z-2192-2011 | 58451 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD Nokor Admix Needle |  |  | March 01, 2011 | 2011 | Becton Dickinson &amp; Company | Incorrect unit package labels which reflect Reorder #305216 instead of the correct item, Reorder #305213, were used. | Packaging | BD sent an Urgent: Product Recall letter with return response forms to direct consignees on 03/01/11 and 03/02/11 via UPS. The letter identified the affected product and lot along with the reason for the recall. The letter asked customers to complete the enclosed Distributor Packing List and to follow the instructions provided on the Packing List for returning the affected product. The firm will issue a credit upon receipt of the affected product. Customers can contact BD Customer Service if they need assistance or have questions regarding the recall. |  | 161000 | 161,000 units | Worldwide Distribution -- USA, including the states of NJ, IL, LA, NM, OH, NV, and VA, and country of Canada. | Terminated | November 30, 2012 | 640 | 161000
Z-2193-2011 | 58421 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Digisonics, Digiview and ERS software releases 3.6.44, 3.6.4.6, and 3.6.5, Digisonics, Inc., 3701 Kirby Drive, Houston, TX 77098, Tel: (800) 940-3240, Fax: (713) 529-7999 |  |  | February 06, 2009 | 2009 | Digisonics, Inc | Left ventricle volume was inaccurate when the volumes were measured from images having different scales. | Other | The firm, Digisonics, notified consignees/customers via e-mail entitled "Distribution of Notification of Biplane Error" dated February 6, 2009. The email described the product, problem and actions to be taken. The customers were instructed to immediately contact Digisonics if their site calculates biplane volumes-a utility will be provided which will run on their database and identify all studies in which biplane volumes are calculated. The customers were also instructed to share the following important information with all users of DigiView at their site. The email stated -until the software patch has been installed, biplane volumes should be calculated only from pairs of images having the same depth scale. Replacement DigiView/ERS software that prevents further occurrence of this problem will be provided if it appears their site could be impacted.The Distribution of New Software Release was sent on March 6, 2009.If you have any questions, please call Digisonics Customer Support at 800-344-4266 |  | 74 | 74 units | Worldwide distribution: USA (Nationwide) and countries of: Canada, Lebanon and New Zealand. | Terminated | May 31, 2011 | 844 | 74
Z-2198-2011 | 58198 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Curad Flex-Fabric/Tissu Souple, 3/4 x 3 in., 100 sterile adhesive bandages, Made in China, Manufactured for Medline Industries, Mundelein IL 60060 |  |  | February 22, 2011 | 2011 | Jiangsu Wujin Hualian Health Dressing Co., Ltd. | The sterility of the product has the potential to be compromised due to weak sealing of the bandages. | Process control | Changzhou Hualian Health Dressing Co., Ltd notified sole consignee by e-mail and a Medical Device Recall letter, dated 3/18/2011. The letter identifies the affected product and the reason for the recall. The customer was to examine their inventory and quarantine the affected product. If the product had been further distributed, those customers were to be identified and notified of the recall. The enclosed response form was to be completed and returned to Medline Industries, Inc. Questions could be directed to Customer Service, Monday through Friday between 8:30 A.M. and 4:00 P.M. |  | 1092 | 1,092 cases | Nationwide Distribution -- IL. | Terminated | April 23, 2012 | 426 | 1092
Z-2195-2011 | 58438 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac, Medical Linear Accelerator, All Varian Clinac with Model Numbers: H14, H18, H26, H27, H28, H29, Hcx; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | March 28, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly whereby, following prolonged use, the screw fastener holding the wedge body to the tray may fail. | Process control | The firm, Varian Medical, sent an "Urgent Medical Device Correction" Letter dated March 28, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately inspect all hard wedges for missing or sheared screw fasteners, if any wedge is found to have one or more missing or sheared screw fasteners, remove the wedge from use and notify Varian immediately; inspect all hard wedges for loose screw fasteners, and retighten as necessary, and implement a program of monthly inspection of wedges for loose or missing hardware. The customers were also instructed to advise the appropriate personnel working in their Radiotherapy Department of the content of this letter.Varian Service shall immediately begin an interim action to replace screw fasteners on all hard wedges for Clinacs installed prior to January 2004. Varian Service will contact the customers to arrange for the inspection visit.Varian also stated that they shall continue their Investigation of this issue to further clarify the exact conditions under which the wedge screw fasteners may fail, and to analyze possible final solutions. Varian will notify the customers when final determinations are made.If you require further clarification, please feel free to contact your local Varian Customer Support Services Manager and/or USA and Canada at 1-888-4265; Europe: +41 41 749 8844; or email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.Helpdesk@varian.com; Japan: Japan.apps helpdesk@varian.com; Latin America: soporte.al@varian.com, a d Internet: Oncology Systems-www.myvarian.com or Varian Medical Systems - www.varian.com. |  | 5616 | 5616 units | Worldwide distribution: | Terminated | June 07, 2013 | 802 | 5616
Z-2197-2011 | 58167 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Manta Ray" Anterior Cervical Plate (ACP) System, |  |  | January 10, 2011 | 2011 | Integra LifeSciences Corp. | Based upon the results of the firm's literature review, they determined that the Package Insert (Information for Use) and information in the Surgical Technique should be revised to include additional warnings, precautions and possible adverse events. | Pending | The firm, Integra Spine, sent an "URGENT: FIELD CORRECTIVE ACTION" letter dated January 10, 2011 with a Field Corrective Action Acknowledgement and Return Form sent via Fed Ex/ priority mail to all customers including physicians and distributors. The letter describes the product, problem and actions to be taken. The customers were instructed to review the current inventory to identify product from the affected part numbers; replace the existing IFU (22-PI-01 rev C) with the revised IFU (22-PI-01 rev D) and the existing Surgical Technique (ST-22-001 rev A) with the revised Surgical Technique (ST-22-001 rev B), and complete and return the Field Corrective Action Acknowledgement and Return Form via fax to: 877-558-6227 or email.If you have any additional questions, contact the Manager Quality Systems Integra Spine at 330-475-8646. |  | 244 | 244 IFU | Nationwide distribution: USA including states of: AL, AZ, CA, CO, FL, GA, ID. IL, KS, LA, MD, MO, NH, NM, NV, NY, OH, OK, OR, PA, PR, SD, TN, TX, UT, VA, WI, and WV. | Terminated | October 27, 2011 | 290 | 244
Z-2196-2011 | 57951 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Aortic Perfusion Cannula with Duraflow Coating, REF: DARH201190TA, 20 Fr. x 28 cm, Sterile EO, Edwards Lifesciences LLC, Irvine, CA 92614. |  |  | January 21, 2011 | 2011 | Edwards Lifesciences Llc | Bonding of the tip to the aortic cannula tubing may decline over time and may contribute to a drop in bond strength. Possible tip separation. | Other | The firm, Edwards Lifesciences, sent an "URGENT PRODUCT RECALL - ACTION REQUIRED" letter dated January 21, 2011 it customers via Fed-Ex. The letter described the product, problem and actions to be taken. The customers were instructed to check their inventory for any of the lot numbers; return all unused product from their stock to Edwards; contact their Customer Service Organization at (800) 424-3278 to obtain a Returned Goods Authorization number and replacement product; complete and return the Aortic Perfusion Cannula Recall Response form via fax to: Edwards at (949) 250-3489, and return unused product to Edwards Lifesciences, 12050 Lone Peak Drive, Draper, UT 84020. If you have questions that have not been answered by this letter, call Edwards Customer Service at (800) 424-3278 from 8:00AM - 4:30PM Pacific Time. |  | 20 | 20 | Nationwide distribution: CA only. | Terminated | May 13, 2011 | 112 | 20
Z-2194-2011 | 58490 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra NeuroSciences, Combo Kit, Lot 171237, Integra LifeSciences 3498 West 2400 South, West Valley City, UT 84119. |  |  | February 22, 2011 | 2011 | Integra LifeSciences Corp. | Outer packaging of Combo Kit was mislabeled with an incorrect lot number. | Labeling mix-ups | Integra LifeSciences Corp. sent a " Urgent Product Recall Notification" letter dated February 22, 2011, to the only affected consignee. It was reported the Combo Kits of SPOO98 / INS- 4500 ( Hermetic Large Style Ventricular Catheter set and drill set for NY) - Catalog # 31046 Lot Number 171237, Outer packages were labeled incorrectly. The customer was advised to return this, and they would be replaced with correctly labeled product. The issue is easily recognizable, and does not pose a safety risk to patients.An acknowledgement form was included to be returned to Integra. Further information is available at 609-936-2485. |  | 7 | 7 Kits | Nationwide Distribution -- including the state of New York. | Terminated | May 13, 2011 | 80 | 7
Z-2064-2011 | 58160 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Defibtech Lifeline AED and ReviveR AED, |  |  | April 01, 2011 | 2011 | Defibtech, LLC | Software with version 2.004 software or earlier may cause an affected AED to cancel shock during the charging process | Software design (manufacturing process) | The firm, Defibtech, sent an "URGENT MEDICAL DEVICE SAFETY INFORMATION AND CORRECTION" letter dated March 11, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Remove the pads from the storage pouch on the back of their DDU-100 series AED and inspect the 9-digit serial number label.2)Carefully note the 9-digit serial number on the back of the AED, then visit the Defibtech website at www.defibtech.com/fa11 and enter their serial number to determine if their AED is affected; email their serial number and contact information to fa@defibtech.com or call Defibtech directly at 1-877-453-4507 or 1-203-453-4501 and Defibtech will inform them if their unit is affected.3) If their AED is affected, confirm receipt of this notification with Defibtech as soon as possible by website and/or confirmation card. Complete and return the confirmation form via fax to 1-203-738-1072. Note: "it is very important for you to contact us either through the Defibtech website or return this card to confirm that you have received this notice. This information will expedite the process of upgrading your AEDs."If your DDU-100 series AED serial number DOES NOT fall within the listed range, your AED is not affected by this recall.A letter was also sent to the distributors and direct accounts that have received the affected product by first class mail with a prepaid, return written acknowledgement card. The field action consists of a software change to the AED. All affected AEDs operating with software version 2.004 or earlier are to remain in the field and in service throughout the field action process.Upon receipt of contact information from affected customers, Defibtech will send an upgrade kit to customers consisting of a software data card, instructions on how to perform the upgrade, and a pre-paid upgrade confirmation postcard that must be returned to Defibtech upon completion |  | 91472 | 91, 472 units | Worldwide distribution: USA (nationwide) and countries including: AUSTRALIA, AUSTRIA , BAHRAIN, BRAZIL, CANADA, CHILE, CHINA , CYPRUS, CZECH REPUBLIC, EGYPT, EL SALVADOR, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JORDAN, KOREA, REPUBLIC OF LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NORWAY, PHILIPPINES, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM AND URUGUAY. | Open, Classified |  |  | 91472
Z-2200-2011 | 58505 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, Version 2.6 and 2.6.1, Model Number: HZl, Manufactured and Distributed by: Varian Medical Systems Inc., Palo Alto, CA. |  |  | March 30, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Optical Guidance Platform (OGP) Software v2.6 and v2.6.1 where the transfer of datasets from treatment planning systems other than the FastPlan system result in a lateral offset error. | Software design | On 3/30/2011, Varian distributed an Urgent Medical Device Correction Urgent Field Safety Notice to all affected users, with a description of the problem and identification of the affected product. The letter also informed customers of how they can determine if the problem affects them. The letter states that there is no work around for the issue unless they possess a FastPlan system. Varian is preparing a software patch to correct the issue and it will be applied to all OGP v2.6 and v2.6.1 systems. Customers are to advise the appropriate personnel of the content of the letter. If further clarification is needed, customers are to contact their local Varian Customer Support District or Regional Manager. |  | 598 | 598 units | Worldwide Distribution | Terminated | June 19, 2012 | 447 | 598
Z-2199-2011 | 58596 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | The Operating Manual contains the Operating Instructions for AS 104 Blood Cell Separator and has a Software Version 4.71. The AS 104 Blood Cell Separator Device has a Catalog Number 9007031. |  |  | April 19, 2011 | 2011 | Fresenius Kabi, LLC | Operating Manual of the AS 104 Blood Cell Separator is being updated to eliminate any reference to Red Blood Cell (RBC), C4Y, P1Y and BMSC sets. This is a corrective action and recall expansion of Z-0616-2011. | PMA | Fresenius Kabi sent a "Urgent Notification" letter dated May 6, 2011, to the affected consignees they are recalling the Operating Manual for the Fresinius AS104-Blood Cell Separator.The Red Blood Cell (RBC) sets have been recalled according to recall number Z-0616-2011, and will permanently not be used with the AS 104 Blood Cell Separator Device. The RBC set was never approved by the FDA according to Warning Letter (Reference # GEN1001508) dated 31 January 2011. The customers are being instructed to complete the self addressed postage pre-paid " Product Response Card." Place a check mark next to the disposition choice used.Fresenius plans to send the revised Operating Manual and a notification letters to consignees at the end of April 2011.Consignees who have questions about this recall can call the firm at 1-425-242-2074. |  | 158 | 158 units | Nationwided Distribution -- AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KY, LA, MA, MI, MN, NC, NH, NJ, NM, NY, OH, PA, SD, TN, TX, VT, and WI. | Terminated | December 06, 2011 | 231 | 158
Z-2212-2011 | 55207 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate TM Sample Processing System, P/N A20745 Basic Assembly - Automate |  |  | May 07, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed an issue with the AutoMate Sample Processing Systems. If during the recovery from Error 81201001, a stuck sample carrier with a tube is moved backwards to the loading position, sample tube/input tray information will become mismatched for samples being loaded onto the AutoMate. | Other | Beckman Coulter sent an Urgent Product Corrective Action letter dated May 7, 2009, with a Fax Back Response Form (sent by US mail) to all customers who received the AutoMate Sample Processing System. The letter identified the product, the problem and a workaround to correct the problem. Customers were instructed to take the following steps:If the stuck sample carrier is moved backwards during recovery from this error, the sample tube/input tray information will become mismatched.Figure 2 in the PCA letter shows the AutoMate viewed from the rear. Sequence 1 to 3 depicts the error recovery issue. The sample is loaded, but cannot move along the track. As the sample attempts to move along the track, it is captured by the track stopper (not seen) below the track. Error 81201001 will be displayed at the console with the following description: The stopper (SL01) did not stop the sample tube carrier. If the stopper or an operator forces the sample backward along the line, it will cause the information mismatch.If you encounter this scenario, do not move the carrier backward to the loading position. See Figure 3 in the PCA letter.Recovery from Error 81201001:1. The normal recovery for error 81201001 is to move the sample carrier forward. If there is no room to do so in the Input module, extra steps must be taken for the carrier to move forward.a. Resume the module by performing error recovery on the Hardware Error Recovery screen at the console. b. Continue the console error recovery process until a carrier is released from BR01, thereby creating space for the stuck carrier to be pushed forward.c. Move the stuck sample carrier forwardd. Once again, resume the module by performing error recovery on the Hardware Error Recovery screen. (The error should not return again)Recovery if Symptoms are Observed: 1. If the incorrect processing described above is observed, the system must be stopped immediately and shut down. All samples m |  | 5 | 5 units total | Nationwide Distribution including VA KY and CA | Terminated | June 28, 2012 | 1148 | 5
Z-2204-2011 | 55542 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | CHVH-001A/1B Variable Helical Pitch (vHP) Software for TSX-101A: Aquilion 64; System, X-Ray, Tomography Computed |  |  | April 09, 2010 | 2010 | Toshiba American Medical Systems Inc | The recall was initiated by Toshiba America Medical Systems (TAMS) because the software associated with TSX-101A: Aquilion 64 can prevent scanning from being performed at the specified helical pitch (couch movement speed). | Software design | Toshiba America Medical Systems (TAMS) forwarded an Urgent: Medical Device Correction letter, dated 4/6/10, with attached Customer Reply Form, via US Postal Service to all customers who purchased the CHVH-001A/1B variable helical pitch (vHP) software for TSX-I0IA: Aquilion 64. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were informed that a Toshiba service representative would contact them to schedule an appointment in order to install a corrective measure that will prevent the occurrence of the problem. Toshiba would like customers to set a scan range different from that for the previous scan for the second and subsequent vHP scans, if multiple vHP scans with the same scan conditions are performed in Auto Mode.Customers were instructed to complete and return the attached form and fax it to the toll free number at the top of the form. The form can also be sent via e-mail to raffairs@tams.com. Customers with any questions can contact the Director of Regulatory Affairs at (800) 421-1968 or contact their local Representative at (800) 521-1968. |  | 35 | 35 | Nationwide Distribution -- MN, OH, MO, MI, NY, KS, MD, TX, NJ, IL, ID, TN, AZ, PA, HI, GA, MT, WA, FL, NC, WI, ME, &amp; CA. | Terminated | August 17, 2012 | 861 | 35
Z-2203-2011 | 55543 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Software for SSH-880CV Aplio Artida; System Diagnostic Ultrasound |  |  | January 06, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) has become aware that the FLEX M software has reversed left and right annotation and post FLEX M measurement error, which may result in incorrect values. | Software design | The recall communication was initiated on 1/6/10 with Toshiba forwarding an Urgent: Medical Device Correction letter with attached Customer Notification Reply Form (via US Postal Service for US Customers) to all customers who purchased the FLEX-M mode on Aplio Artida Diagnostic Ultrasound System. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to avoid performing M image measurement using images saved in FLEX-M mode until the corrective action is completed. If customers need to perform measurement for images saved in FLEX-M mode, they were instructed to use manual calibration. Customers were instructed to note that even though the corrective action is taken, the existing saved images (saved images before the corrective action is taken) are not modified. Therefore, when referring to the measured values on the images which are saved in FLEX-M mode before the corrective action is taken, customers were instructed to carefully handle the measurements results because they may be incorrect.Customers were instructed to complete the attached form and fax it to the toll free number at the top of the form. The form can also be sent via e-mail to raffairs@tams.com.If customers have any questions regarding the medical device correction letter, they can contact the Director of Regulatory Affairs at (800) 421-1968 or their local Toshiba representative. |  | 28 | 28 | Nationwide Distribution -- Including states of IA, MT, PA, CA, OR, GA, LA, OH, TX, MA, OH, WI, and Puerto Rico. | Terminated | April 20, 2012 | 835 | 28
Z-2201-2011 | 56389 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | Software CD is labeled in part: |  |  | July 19, 2010 | 2010 | Bio-Rad Laboratories | The EVOLIS Assay Protocol Files Software is configured in such a way that may result in a miscalculation of quantitative test results and/or an error in reporting positive/negative sample values. | Software design | Bio-Rad Laboratories sent an Urgent Product Correction letter dated July 19, 2010, to all customer sites via FedEx. There are two issues related to The EVOLIS Assay Protocol Files (APF) Software that could potentially affect patient results. Issue 1 involves quantitative sample test/calculations/results interpretation for assays with linear regression, point to point or 4PL data reduction. Issue 2 involves sample test results with " Result Report Flags''. A software patch to correct these two issues will be sent to each customer site. Bio-Rad Laboratories will work with each customer to complete documentation confirming that the software patch has been successfully installed on each instrument. If you have any questions or concerns regarding these two issues, please contact Technical Support at 1-800-224-6723, option 2 and then option 3. |  | 272 | 272 | Nationwide Distribution -- Including Puerto Rico. | Terminated | August 25, 2011 | 402 | 272
Z-2214-2011 | 57591 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | BrightSpeed Excel Select, BrightSpeed Edge Select ,BrightSpeed Elite Select and BrightSpeed Elite Computed Tomography X-ray systems with software version 09BW35.11 or 09HW30.4. |  |  | January 19, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential set of circumstances that could cause X-ray continuation during an unexpected table stop on certain BrightSpeed systems (BrightSpeed Excel/Edge/Elite Select and BrightSpeed Elite Computed Tomography X-ray Systems ) if this event were to occur on your system, patient safety may be impacted. | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION GE Ref: 22932" letter dated January 19, 2011 to its consignees/customers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, Contact Information, and Additional Dosimetric Information.A second letter dated March 11, 2011 was sent to the consignee/customer letter consisting of a cover letter and an attachment. "This letter is to notify you of an error in the January 19, 2011, Urgent Medical Device Correction Letter, that you have already received (GE Ref: 22932). In the Safety Issue section of that letter it is written: "If this issue were to occur, it would result in the remainder of the exposure being delivered at the patient position where the table stopped up to the prescribed scan time or 30 seconds, whichever is shorter." The stated 30 seconds limitation is incorrect for your particular model. This information has been corrected on the attached Urgent Medical Device Correction letter, so please discard the original letter if it is still in your possession. "GE Healthcare will provide a software update for all affected systems to address the issue. A GE Healthcare service representative will contact you to arrange for this correction.The customers were also instructed as follows: prior to GE's installation of a software update, continue to monitor patient scanning closely. Should the table stop unexpectedly and X-ray remains on, manually stop the scan, and ensure that all potential users in your facility are made aware of this safety notification and the recommended actions.If you have any questions, please contact GE immediately at the following: United States: 800-437-1171; Japan: 0120-055-919; Korea: 1544-6119; Australia/New Zealand: 800-659-465 and China: 800-810-8188. For other countries please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notifi |  | 479 | 479 | Worldwide distribution: USA (nationwide) including states: AL, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, OH, PR, RI, SC, SD, TN, TX, and WV; and countries including: URUGUAY, TURKEY, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, REUNION, PUERTO RICO, PORTUGAL, POLAND, PARAGUAY, NEW ZEALAND, NETHERLANDS, MEXICO, MARTINIQUE, LEBANON, JAPAN, ITALY, INDONESIA, INDIA, HUNGARY, GREECE, GERMANY, FRANCE, EGYPT, COLOMBIA, CHINA, CHILE, BRAZIL, BELGIUM, BELARUS, BAHRAIN, AUSTRALIA, ANGOLA, ALGERIA. TUNISIA, THAILAND, SYRIAN ARAB REPUBLIC, PHILIPPINES, PARAGUAY, OMAN, NEW ZEALAND, MOROCCO, KUWAIT, KAZAKHSTAN, ISLAMIC REPUBLIC OF IRAN, ARGENTINA. | Terminated | July 10, 2012 | 538 | 479
Z-2215-2011 | 57469 | No | MSW | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Access Immunoassay Systems Thyroglobulin, |  |  | October 18, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because In addition to Access Tg Sample Diluent, an additional diluent, Access Wash Buffer II, was added to Thyroglobulin instructions for use (IFU) for evaluation of samples having Tg levels above the reportable range of the assay. When customers interchange between diluents there is an unacceptable shift in results. A mediandifference of 16% (ranging from 14% to 23%) b | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 18, 2010, with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Access Immunoassay Systems Thyroglobulin. The letter describes the product, problem and actions to be taken. The customers were instructed to:1. Do not dilute elevated thyroglobulin samples with Wash Buffer II.2. Do not run thyroglobulin patient samples using the onboard dilution feature on the UniCel DxI instrument.3. Dilute elevated thyroglobulin samples only with Access Thyroglobulin Sample Diluent (REF 33866).4. Review your patient result archives. Manually dilute and repeat the analysis of all elevated thyroglobulin samples that had previously been diluted with Wash Buffer II, using Access Thyroglobulin Sample Diluent instead.5. A forthcoming version of the UniCel DxI Assay Protocol File (APF) will remove the Access Thyroglobulin onboard dilution assay (dTg) from the list of available onboard dilution assays.6. The Access Tg product Instructions for Use will be updated to eliminate Wash Buffer II as an acceptable diluent for elevated thyroglobulin samples. Customers should update procedure manuals and other documents that your laboratory may have created which list Wash Buffer II as an acceptable diluent for elevated thyroglobulin samples.In addition, customers were instructed to complete and return the enclosed PCA Response Form with in 10 days via fax to: 786-639-4000 or mail to: Beckman Coulter, Inc., 1000 Lake Hazeltine Drive, Chaska, MN 55318, Attn: Regulatory Affairs. Customers with any technical questions regarding the Product Corrective Action were instructed to contact Beckman's Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside the US and Canada, customers were instructed to contact their local Be |  | 6414 | 6,414 units | Worldwide distribution: USA (nationwide) and country of: Canada. | Terminated | August 10, 2012 | 662 | 6414
Z-2208-2011 | 58037 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | DePuy Mitek GRYPHON T BR DS Anchor w/ORTHOCORD |  |  | February 24, 2011 | 2011 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Metal debris in the shaft component of the device | Manufacturing material removal | DePuy Mitek sent an Urgent Voluntary Product Recall Notification letter dated February 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to immediately check their inventory to determine if they have any affected product. Customers were asked to complete the enclosed business reply form and fax it to 508-828-3750 or email to mitekcomplaints@its.jnj.com whether or not they have any of the affected product remaining in their inventory. Customers were asked to utilize the return good authorization number provided by customer service and return all identified product to:DePuy MitekATTN: Recall Coordinator50 Scotland BoulevardBridgewater, MA 02324For any questions call 1-877-379-4871 or contact DePuy Mitek Customer Service directly at 1-800-382-4682, PROMPT 3. |  | 1126 | 1126 units | Worldwide Distribution - USA (nationwide) and the countries of Austrailia Austria Belgium Canada China Estonia Finland France Germany Great Britain Indonesia Ireland Italy Japan Lebanon Lithuania Netherlands Poland Singapore Slovenia South Africa Spain Sweden and Switzerland | Terminated | February 20, 2013 | 727 | 10139
Z-2219-2011 | 58429 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | PercuNav Field Generator Stand is a Computed tomography, X-Ray System. Part Number 453561431911. |  |  | April 01, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | The PercuNav Field Generator Stand could tip over during set up, during use, or during transport. | Device Design | On 4/4/11, Philips Healthcare began sending out the URGENT-Medical Device Correction notification letter to their consignees.Philips informed their customers of a potential tipping hazard of the PercuNav Field Generator Stand. The notification letter also provides instructions to customers to ensure the articulating Mounting Arm is positioned so the Field Generator is not further than the specified distance from the back of the clamp on the stand's central pole during set up or during use. Customers are instructed to apply a provided label on stand to warn of a potential for tipping due to excessive force. For transport, customers are instructed to remove the Field Generator from the Mounting Arm and place the Field Generator in the storage compartment on the system cart. The Mounting Arm should then be collapsed and folded into a 90 degree angle.Customers are informed that a Philips field service representative will contact them to schedule installation of a service kit on the stand, free of charge, once service kits are available.Customers can call Philips Customer Service at 1-866-767-7822 and reference to FCO 79500197 for any questions about this recall. |  | 65 | 65 units total (33 units distributed to US and 32 units to foreign consignees) | Worldwide Distribution -- USA, Australia, Austria, Belgium, China, France, Germany, Hong Kong, India, Korea, Netherlands, Singapore, Sweden, Switzerland, Taiwan, Thailand, and United Kingdom. | Terminated | March 15, 2012 | 349 | 65
Z-2217-2011 | 58667 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Audio-Video Engineering PAS-210 Remote Active Speakers Kit used with the IntelliVue Information Center |  |  | May 03, 2011 | 2011 | Philips Healthcare Inc. | PAS-210 Kit Remote Active Speaker Assembly used with the IntelliVue Information CenterSpeaker cable connection may lead to intermittent or loss of audio alarm annunciation | Component design/selection | Philips Healthcare issued an Urgent - Medical Device Correction letter dated April 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the "Action to be taken by Customer/User" section of the instructions while they await the correction, which will likely come in the form of a replacement audio cable for all affected speaker assemblies.For further information or support concerning this issue, please contact Philips Healthcare Customer Care Service Center at 800-722-9377 # 3, # 1 and reference the letter or your local Philips representative in your area. |  | 15062 | 15062 units | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | April 25, 2017 | 2184 | 15062
Z-2213-2011 | 58235 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | COULTER LYSE S III diff Lytic Part Number: 8546796. |  |  | December 28, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated by Beckman Coulter because the COULTER LYSE III diff Lytic Reagent Lot number 101554F was released for distribution to customers between November 17 and December 1, 2010 prior to completion of QC testing. | Pending | Beckman Coulter sent a Product Corrective Action ( PCA) letter dated December 28, 2010, to all affected consignees with an attached PCA Response form via US Postal Service for US Customers) who purchased the Coulter Lyse S III diff Lytic Lot # 101554F. The letter provides customers with an explanation of the problem identified and an action to be taken. If customers currently had any of the lot mentioned above in their laboratory, they were asked to discontinue use and discard remaining product according to their laboratory chemical safety procedures. And to complete and return the enclosed response form within 10 days to acknowledge that they received this notice and are taking appropriate action.For the population at most risk, there is no serious consequences and only remote probability of injury and medically reversible consequences due by the small change for Hgb and WBC results.Consignees were also asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation.Technical questions regarding this Product Corrective Action were directed to Beckman Coulter Customer Service at( 800) -526-7694 in the United States and Canada. |  | 1296 | 1,296 units | Worldwide Distrubution-- USA ( nationwide) including states of : AL, AR, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NY, OH, OK, PA, RI, SC, TN, TX, VA, WI, &amp; WV and country of Canada. | Terminated | June 04, 2012 | 524 | 1296
Z-2221-2011 | 58614 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Radiesse, Injectable Implant - 1.3cc Syringe, Sterile, Sterilized Using Steam, Use Before 2013-01, Ref 8046M7, Lot 1024417, Manufactured by: BioForm Medical Inc, 4133 Courtney Rd #10 Franksville, WI 53126. |  |  | April 06, 2011 | 2011 | Merz Aesthetics, Inc. | Merz Aesthetics, Inc. is recalling 1 lot (1034417) of Radiesse 1.3 cc Implant. The lot was supposed to be scrapped because it had failed to meet its Loss On Drying (LOD) specification. Unfortunately it had been marked as approved, released and shipped. | Employee error | The firm, Merz Aesthetics, Inc., contacted its consignee/customer via e-mail on April 6, 2011. The e-mail states the product, problem and action to be taken. The customer was instructed to follow up with the concerned and inform the status time to time to the Recall Coordinator. The product was confirmed to be quarantined and has been requested to be sent back.If you need any additional information regarding this issue or if you have any questions, contact the International Regulatory Affairs Specialist at 262-835-3300 ext. 3081. |  | 695 | 695 units | International only: Taiwan. | Terminated | November 17, 2011 | 225 | 695
Z-2069-2011 | 57473 | No | RGB |  |  |  |  | Lockheed Martin Aculight Capella-r 1850 Laser. This device has the Accession number 0612339. |  |  | November 19, 2010 | 2010 | Lockheed-Martin Aculight | It has been determined that there are certain software driven operating sequences involving off-on cycling of the key-switch that do not result in a reset of the Stim Enable button, and emissions can resume after a standard 5 second countdown without pressing this button a second time. | Radiation Control for Health and Safety Act | On 11/22/10, Lockheed Martin Aculight sent a Memo to their Customers and Users informing them of the "Stim Enable" button reset error. The defect involves the function of the "Stim Enable" button on the front panel of the Capella R-1850.In December 2010, the firm sent a second lletter with instructions on how to return the Capella R-1850 Laser back for repair.The firm's plan to correct the defect involved removal of existing software, installation of revised software, and documented verification testing, per written instructions, to confirm resolution of the Stim-Emit button failure to reset. Corrections were to be made at no cost to the customer. A Lockheed Martin Aculight customer service representative will contact you shortly and provide detailed information on how to return the Capella-R 1850 laser product in your possession for testing and repair on this defect.Should you have any questions, please call (425) 482-1100. |  | 17 | 17 units | Devices were distributed Universities, Institute, and government consignees in CT, IL, MA, MD, OH, TN, TX, and UT.One unit was distributed to Germany. | Terminated | August 23, 2011 | 277 | 17
Z-2220-2011 | 58581 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, Client software versions 8.0 through 8.6; Model Number: H48; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | April 04, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly in Eclipse Treatment Planning System, versions 8.0 to 8.6, whereby Actual SSD (source surface distance) value may not correspond to the Planned SSD value after Planning Approval. | Software design | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 4, 2011, to its customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Set the Planning Approved status in Eclipse (Eclipse displays both the Actual and Planned SSD values during plan approval) and verify both SSD values prior to approving the plan.2)Verify the SSD values in the Eclipse Treatment Plan Report prior to approving the plan for treatment or proceeding with treatment delivery (this report lists both the Actual and Planned SSD values).3) Prior to approving the plan in RT Chart, verify both the Actual and Planned SSD values in the Parameters Workspace by selecting Field Parameters&gt;Geometry tab.4) Prior to treatment, they should verify intended SSD values during independent MU calculation.5)Re-familiarize themselves with the concepts of Actual SSD and Planned SSD by reviewing the listed information.Please advise the appropriate personnel working in your radiotherapy department of the content of this letter.Varian mitigated the product behavior in version 8.9.If you require further clarification, please feel free to contact the Varian Oncology Help Desk ; phone: USA and Canada - 1888-827-4265; Europe: +41 41 749 8844; email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com; Latin America: soporte.al@varian.com; Internet: Oncology Systems-www.myvarian.com; Varian Medical Systems -www.varian.com; or your local Varian Customer Support District or Regional Manager. |  | 5288 | 5288 units | Worldwide distribution. | Terminated | June 27, 2012 | 450 | 5288
Z-2216-2011 | 58534 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE , ONCOR and PRIMUS accelerators that utilize the RT Therapist v4.2 system in combination with syngo RT Oncologist 4.2. |  |  | March 31, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a potential safety issue that may result in mistreatment when customers are configured with the ARTISTE Linac RT Therapist v4.2 in combination with RT Oncologist v.4.2 and use the remote "Online 2D Image Review. | Software design | The firm, Siemens, sent an "Urgent: Medical Device Correction Customer Safety Advisory Notice" on March 31, 2011 via Siemens Customer Service Group or certified mail to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to not use the "Online 2D Image Review" on the Syngo RT Oncologist 4.2; to always calculate the offset directly on the RTT 4.2 before it is applied for the treatment with the patient on the treatment table; include this notice in their syngo RT Therapist 4.2 System Owner Manual chapter "Safety Advisory Letter" where it should remain until the measures defined in the notice are performed, and in the interest of safety, perform the preventative measures noted in this notice and inform all affected personnel immediately.Siemens is currently working on a solution for the incremental loading workflow in syngo RT Therapist 4.2 and will provide a software update as soon as possible. Siemens will notify the customers as soon as an update for their system will be available. UPDATE: The release of Safety Update UI TH010/11/S to distribute the software update for RTTv4.2.94O04 began on April 15, 2011.If you have any questions, contact Regulatory Affairs at 925-602-8157. |  | 45 | 45 active devices | Worldwide distribution: USA (nationwide) including states of: CA, FL, OH, PA and UT; and countries including: Australia, Belgium, Croatia, Germany, Japan, South Africa, Sweden and United Kingdom. | Terminated | March 30, 2012 | 365 | 45
Z-2205-2011 | 58510 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT, Computed Tomography X-Ray System, including the Brilliance iCT SP system, Model Numbers: 728311 and 728306. The affected systems are equipped with the following software version: 3.2.0.19010. |  |  | March 07, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | A program software error in the algorithm produces image artifacts that could possibly lead to a misdiagnosis of patients who have small metal implants. | Other | A 'Field Safety Notice - Urgent - Medical Device Correction' notification was sent to the firm's customers via Certified mail on 3/8/2011. The notification letter explains the nature of the Metal Artifact Reduction/Metal Artifact Correction software error; the Brilliance CT X-Ray systems subject to correction and how to identify them; and the circumstances under which this error can occur. It also asks that the user stop running the MAR application on any of the affected Brilliance systems identified with the recalled serial numbers and software versions alluded to in the medical device correction notification. Lastly, the notification states that Philips Medical will provide a solution that will fix the issue and the notification lists contact names; telephone numbers; and E-mail addresses which the customer can contact for further information concerning this device issue, or customer support assistance. |  | 8 | 8 units | Nationwide Distribution -- DC, IL, IN, MI, MO, NY, OH, OR, PA, TX, and VT. | Terminated | February 22, 2013 | 718 | 23
Z-2222-2011 | 58586 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague, Colleague CX and Colleague CXE Infusion Pumps, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2M8151, 2M8161 and 2M9161. Intended for use in a wide variety of patient environments for adult, pediatric and neonatal patients. |  |  | April 14, 2011 | 2011 | Baxter Healthcare Corp. | The Hi-Pot safety testing may not have been conducted on the affected pumps during service, prior to being returned to the end user. | Employee error | The firm, Baxter Healthcare Corporation, telephoned the customers on April 14, 2011, informing them that that the Hi-Pot test,and sent a follow up "URGENT DEVICE CORRECTION" letter dated April 27, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to locate the pumps and take them out of service as soon as possible (Baxter will then send a shipping container for retrieval, re-inspection and return of the device back to the customer at no cost.); and return their colleague pumps for full functional testing including the HiPot test to the Baxter Buffalo Grove Global Technical Services (GTS) Center at Baxter GTS, 900 Corporate Grove Drive, Buffalo Grove, IL 60089.Baxter will perform full functional testing, including the HiPot test, on affected Colleague pumps for which the customer has continuing medical necessity. Should Baxter identify any issues through repair or service prior to returning the Colleague pumps to the customer.If you have questions regarding this communication, please call The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CT, Monday through Friday. |  | 117 | 117 pumps | Worldwide distribution: USA (nationwide) including states of :Alabama, Arizona, California, Connecticut, Florida, Georgia, Illinois, Kentucky, Louisiana, Massachusetts, Maryland, Mississippi, North Carolina, New Jersey, New York, Ohio, Pennsylvania, South Carolina, Texas, Virginia, Washington and West Virginia, and country of: Canada. | Terminated | May 09, 2012 | 391 | 117
Z-2223-2011 | 58514 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | V-CATH Insertion Kits, Sterile; Product Code Number: 301-00; |  |  | April 05, 2011 | 2011 | Neo Medical, Inc. | Firm received recall notice from their contract manufacturer due to visible particulates in Bacteriostatic Sodium Chloride Injection, USP 0.9%, 30 mL vials that are included in insertion kits. | Material/Component Contamination | Neo Medical forwarded Urgent Recall Notices to Dealers and customers on April 5, 2011 via fax. Neo Medical Customer Care also contacted customers by phone on April 7, 2011, followed by Urgent Product Recall Notification letters were sent on April 11, 2011 along with a copy of the Recall Notification from Centurion. The letter requested that customers with affected lots immediately discontinue distribution and contact the end users instructing them to stop dispensing the product. Customers were to complete the accountability information even if inventory had been depleted and fax it to 510-353-9025. In order to return the affected product, customers should contact the Complaint Coordinator at 925-719-2197 or pjduvall@neomedicalinc.com. Replacement or credit would be issued for all product returned. |  | 1990 | 1990 | Worldwide Distribution -- USA, including states of MO, WA, PA, CA, OH, UT, TX and Canada. | Terminated | November 22, 2011 | 231 | 1990
Z-2218-2011 | 58549 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | AVOXimeter 1000, AVOXimeter 4000 |  |  | March 23, 2011 | 2011 | International Technidyne Corp. | An incorrect calibration code that affects the total hemoglobin measurement was applied to affected Avoximeter instruments. | Device Design | International Technidyne Corp. sent a letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to check their inventory to determine if they have any of the affected product. If they have any of the affected product, customers were instructed to stop using it and remove it from their inventory. Customers were also instructed to complete the attached Customer Account Tracking Form and return it by fax to (732-635-0144) or e-mail (techsupport@itcmed.com) or mail to:ITC Technical Support20 Corporate Place SouthPiscataway, New Jersey 08854For any questions call the Distributor or ITC Nexus Dx Technical Support at 800-631-5945 (US) or 732-548-5700 (International), ext 4707. |  | 17 | 17 | Worldwide Distribution - USA including TX, KS, OH, OR, FL, GA, NY, MA, and DC and the countries of Taiwan and Thailand | Terminated | October 25, 2012 | 582 | 17
Z-2202-2011 | 58179 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Theken L"POD" Titanium Vertebral Body Replacement System, Vu c"POD" PEEK-OptimaÂż Vertebral Body Replacement System, c"POD" Titanium Vertebral Body Replacement System |  |  | January 18, 2011 | 2011 | Integra LifeSciences Corp. | Certain Theken Spine Vertebral Body Replacement Systems were distributed (i) without certain components (spacers and set screws) that are necessary for their use as vertebral body replacements, and (ii) with a component (insertion tool) that is not relevant to the devices' intended use. | Other | The firm, Integra Spine, sent an "URGENT: RECALL NOTIFICATION" letter dated January 18, 2011, to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to review their current inventory and return all remaining kits; complete and return the RECALL ACKNOWLEDGMENT AND RETURN FORM via fax to: 877-558-6227 or email: melissa.niosi@integralife.com, and contact Integra Spine Customer Service at 866-942-8698 regarding availability of suitable replacement products.Integra Spine will be taking immediate action to restore your inventory and ensure minimal impact on availability for surgical procedures.Should you have any additional questions, please feel free to contact me at 609-936-2495 or email at jon.caparotta@integralife.com. |  | 93 | 93 Kits | Nationwide distribution: USA including states of: AL, AZ, CA, FL, ID, IL, MD, MI, NH, NM, NV, OH, OK, OR, PA, TX, UT and WI. | Terminated | October 27, 2011 | 282 | 93
Z-2229-2011 | 58675 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Kits and Trays containing alcohol swabstick/prep pad: |  |  | February 07, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | Custom surgical kits contained either one or more of the following possible contaminated recalled products: Triad lubricating jelly, alcohol prep pads, and alcohol swabs. | Other | DeRoyal notified their consignees by letter with the subject line "Re: DeRoyal Recall of Triad Medical Lubricating Jelly, Alcohol Prep Pads and Alcohol Swabs" on 02/07/2011 that kits in their possession contained recalled Triad products that were possibly contaminated. Customers were to use the attached spreadsheet to identify and quarantine the affected product. The spreadsheet should have been completed even if there was no affected product. The form should have been faxed back to 865-362-3716 or emailed to recalls@deroyal.com. Customers' DeRoyal Sales Representative would re-label the affected product in order to prevent further use. The DeRoyal Sales Representatives are also providing in-service training on effectuating this recall, including the importance of identifying, retrieving and discarding the recalled products as well as the documentation necessary to account for these products. Distributors were directed to recall down to the retail level. If there are questions or if further assistance is needed, customers can contact their DeRoyal Sales Representative or customer support department at 1-800-251-9864. |  | 340976 | 340,976 units total | Nationwide Distribution | Terminated | August 06, 2014 | 1276 | 340976
Z-2237-2011 | 58563 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Nasal Cannula (Adult) Salter Style with modified nasal prongs, 7" (2.1M) Q supply tube and CO2 sample tube with male luer connector 22mm (I.D) x 6mm (O.D) Adaptor. Adult Divided Cannula. |  |  | April 19, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | Complaint received that the product in the box was different from labeled product. | Labeling mix-ups | The firm, Salter Labs, sent an "URGENT PRODUCT RECALL" subtitled Adult Divided Cannula Product Code 4707 Lot 012411, letter dated April 19, 2011 via certified mail to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory; return any of the lot in their possession to Salter; contact Salter Customer Service at 1-800-235-4203 to arrange for the return of the product, and complete and return the Customer Reply Form via fax to 661-854-6816 or Toll Free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 22150 | 22,150 pieces | Nationwide distribution. | Terminated | August 23, 2011 | 126 | 22150
Z-2231-2011 | 58569 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | COULTERÂż Ac"T" 5diff Fix Reagent. |  |  | April 08, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that an issue has been identified with one of the raw materials used to manufacture the Coulter A T 5 diff PN 8547171. | Other | Beckman Coulter initiated a Product Corrective Action (PCA) letter on April 8, 2011, with attached PCA Response Form to all customers who purchased the COULTERÂż Ac"T" 5diff Fix Reagent. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to discard the product according to their laboratory safety procedures and use the attached response form to request replacement product. All flagged results should be reviewed as outlined in the Instructions for Use manual and your laboratory protocol. A retrospective review of results is not required if all recommendations for review have been followed as shown in the Instructions for Use.Customers were instructed to share the information with their laboratory staff and to retain the notification as part of their Quality System documentation. If the product was further distirbuted to another laboratory, customers should provide them a copy of the notification letter.The raw material issue has been corrected beginning with COULTERÂż A T 5diff Fix lot 16902E.Customers were instructed to complete and return the enclosed Response Form within ten (10) days.Customers with any questions concerning this notice, were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 4662 | 4662 units total (2278 in US) | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Angola, Australia, Botswana, Bulgaria, Czech Republic, Estonia, Finland, France, Germany, Greece, Guyana, Italy, Japan, Kenya, Mexico, Morocco, Namibia, Netherlands, Nigeria, Philippines, Puerto Rico, Russian Federation, Slovakia, South Africa, Switzerland, Taiwan, United Arab Emirates, United Kingdom, Venezuela, and Zambia | Terminated | June 04, 2012 | 423 | 4662
Z-2227-2011 | 58584 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UnicelÂż DxH 800 CoulterÂż Cellular Analysis System. |  |  | April 14, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed the auto prune will not remove patient results if the patient record includes demographics where the results are less than 30 days old. | Software design | The recall communication was initiated on 4/12/11 with Beckman Coulter forwarding an Urgent Product Correction (PC) letter with attached PC Response Form (via US Postal Service for US customers) to all customers who purchased the UniCelÂż DxH" 800 CoulterÂż Cellular Analysis System PN 629029. The letter provides customers with an explanation of the problem identified and an action to be taken. For Issue 1 &amp; 2, customers were instructed to perform manual database maintenance procedures to locate and delete unnecessary patient result and demographic information that extends beyond the life (in days) of the control lots. Ensure an alternative method of analysis is available during database maintenance.For Issue 3, customers were instructed to not attempt to transmit the files again after receiving an incomplete report. Instead, submit the missing file(s) by one of the following methods:- Submit your data at the electronic IQAP website (eIQAP) at(http://www.beckmancoulter.com/qap/index.jsp)- Fax printouts of the missing files to 786-639-3999- Email an electronic copy (.CSV) of the control file to IQAP.Reports@beckman.com.- Include the IQAP number on any communication. A corrected report will be returned to the customers' laboratories.For issue 4, customers were instructed to submit their control data to IQAP by one of the following methods:-Submit their data at the electronic IQAP website (eIQAP) at(http://www.beckmancoulter.com/qap/index.jsp)-Fax printouts of the missing files to 786-639-3999-Email an electronic copy (.CSV) of the control file to IQAP.Reports@beckman.com.-Include your IQAP number on any communication.Customers were instructed to complete and return the enclosed Response Form within ten days. Customers were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 673 | 673 units | Worldwide Distribution -- US, Canada, Australia, Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Republic of Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | May 23, 2012 | 405 | 673
Z-2225-2011 | 58333 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Universal Power Supply Units, Part Number DP0007 and DP0800 , for use with the Integra Padgett Dermatomes - Model S, SB, and S6, Device Listing #E208841, 510 (k) exempt. |  |  | March 11, 2011 | 2011 | Integra LifeSciences Corp. | Dermatome power supply units distributed by the recalling firm may not meet the required medical device electrical safety regulations which can result in an increased potential for electric shock to the device user and/or patient. | Other | Integra LifeSciences Corp. initiated their recall by contacting their Sales Representatives via telephone conference calls on 3/10/2011. Separate written recall communications, in the form of an 'URGENT PRODUCT RECALL NOTIFICATION' were provided via E-mail and FedEx to the firm's Sales Representatives on 3/11/2011; and to their hospital and distributor customers on 3/11/2011; 3/14/2011; and on 4/12/2011. The recall notification addressed to the non-Integra Sales Rep customers informs the customers of the problem with the power supply units (that they may not meet the required medical device electrical safety standards which may result in an increased potential for electrical shock). The notification goes on to ask that the customers examine their inventories and determine if they have any affected product on-hand; if so, the customers are instructed to discontinue use of the power supply unit and return the unit(s) to Integra LifeSciences Corporation immediately for reworking. The letter also provides photographs of the affected units which the customer can use to locate and verify the serial number(s) of the device (if applicable). In addition, the notification is accompanied by a 'RECALL ACKNOWLEDGMENT AND RETURN FORM' which the customer is asked to sign and complete in order to document receipt of the recall notification and whether or not the customer has any of the recalled power supply units available for return. The notification letter provided to the Integra Sales Representatives describes the device models affected; the reason for the recall; and asks that the Reps. examine their inventories for the presence of any of the affected power supply units. If found, the Reps are instructed not to use, or to further distribute any of the affected units in their possession, but to immediately make arrangements to return the unit(s) to Integra for rework. The notification supplied to the Sales Reps provides them with a synopsis of the firm's field action |  | 1047 | 892 units were distributed in the U.S &amp; 155 units were distributed in foreign countries | Worldwide Distribution - USA including FL, CO, IL, AR, MD, IN, CA, MT, AZ, NY, LA, TX, MA, MO, CT, MI, NE, OH, NC, WV, VA, PA, NJ, TN, UT, MN, OK, WI, GA, AL, WA, KS, OR, AK, DC, SC, NE, MS, RI, NV, IA, and ND and the country of Canada | Terminated | April 17, 2012 | 403 | 1047
Z-2232-2011 | 58464 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MAKO PN 204588 Camera |  |  | October 25, 2010 | 2010 | Mako Surgical Corporation | MAKO Surgical Corporation recalled ONE infra-red camera MAKO PN 204854 which was a component of their RIO Robotic Arm Interactive Orthopedic System with SSN P7-03127, on November 22, 2010. This component can function improperly and cause the accuracy of the system to be outside of the specification listed by the manufacturer. The recall was initiated after the camera manufacture, NDI, Waterloo, On | Nonconforming Material/Component | The firm, MAKO, sent a "RECALL NOTICE" letter dated November 22, 2010, to its customer. The letter described the product, problem and action to be taken. The customer was instructed to complete and return the attached Acknowledgement Form to their MAKOplasty Specialist (MPS) via mail to: MAKO Surgical Corp., Attn: Jim Pomeroy, 2555 Davie Road, Fort Lauderdale, FL 33317. The letter stated that "It is important that the camera be replaced immediately, and MAKO is currently addressing the situation with the camera manufacturer." In addition, MAKO will replace the camera with a camera that is not affected by the recall and verify the customers system to meet the functional specifications.If you have any questions or concerns, please contact your MPS or Sr. Director of Quality Assurance at 954-927-2044 ext 604, and /or Sr. Director, Customer Support at 954-927-2044 ext. 442. |  | 1 | 1 | Nationwide distribution: NY only. | Terminated | November 09, 2011 | 380 | 1
Z-2224-2011 | 58558 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac and TrueBeam, High Energy Linear Accelerator, Model Numbers: H14, H19, H29; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | March 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The coolant may leak. The resulting coolant leak presents a risk of electrical shock to any person working within the protective housing. | Process design | On 3/31/2011, all consignees were notified via an Urgent Medical Device Correction letter and on-site visit by the firm's sales representative, informing them of the affected products, the issue, and recommended actions. If non-Varian personnel service the device, then they should have been immediately notified of the content of the letter. Anyone who performs work on the device should exercise extreme care to identify any possible coolant leak and immediately locate and correct the source of the leak. Customers were to immediately notify Varian Service or their internal service personnel of any coolant leakage. Varian was to contact customers to arrange for inspection if one had not been performed. The appropriate personnel should have been advised of the content of the letter. If further clarification was needed, customers were to contact their local Varian Customer Support. |  | 718 | 718 units | Worldwide Distribution | Terminated | March 13, 2012 | 348 | 718
Z-2226-2011 | 58555 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SYNCHRON Systems Calibrator 1, Part Numbers: 468405, Lots prior to M005558 |  |  | March 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because standardization of the Synchron Systems Calibrator 1, used for both the Synchron LX and UniCel DxC systems, has been established to the new IFCC Reference Preparation for Plasma Proteins, ERM-DA470k (formerly CR470) beginning with lot number M005558.During the standardization to the new IFCC standard, ERM-DA470k, we confirmed that Complement C4 (C4) had been re | Pending | Beckman Coulter initiated a Product Corrective Action (PCA) letter on the week of March 14, 2011, to the affected customers. Though Beckman is recalling the LX Cal 1, the recall letter was sent to the C4 reagent customers, as these are the only customers who will be impacted by the problem noted in the PCA letter. The letter provided the customers with an explanation of the problem identified and an action to be taken.The letter recommended consignees take the following actions:(1)Notify their clients to review and evaluate the impact on patients whose Complement C4 recovered near the clinical decision points.(2) Re-evaluate their laboratory's reference ranges for C4.The letter further requested that consignees share this information with laboratory staff and retain this notification as part of your laboratory Quality System documentation.Questions regarding this Product Announcement letter, or requests for the correlation graphs should be directed to Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/Customer Support/Support Request or call (800) 854-3633 in the United States and Canada. |  | 96173 | 96,173 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Belgium, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Costa Rica, Czech Republic, Denmark, Eritrea, France, French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Republic of Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Namibia, Netherlands, New Caledonia, New Zealand, Oman, Panama, Poland, Puerto Rico, Qatar Reunion, Russian Federation, Rwanda, Serbia, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Viet Nam | Terminated | May 04, 2012 | 420 | 96173
Z-2242-2011 | 58564 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Salter Labs Nasal Cannula (Adult) Salter Style with 14' (4.3) supply tube. |  |  | April 19, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | One complaint received of product in box different from what was on label. | Other | Salter Laboratories sent a URGENT PRODUCT RECALL NOTIFICATION letter dated April 19, 2011, certified mail to all of the affected consignees, along with a RECALL EFFECTIVENESS CHECK FORM to acknowledge receipt of this communication.It was reported that a package labeled 1600-14, Adult Nasal Cannula actually contained a 25-foot length of tubing and has determined that additional product may have been incorrectly labeled in this manner.No illness or injury would result from this product mix-ups.Should you have any questions, please contact Salter Laboratories at 1-800-235-4203. |  | 3500 | 3500 pieces | Worldwide Distribution - USA (nationwide) and the country of South America. | Terminated | August 23, 2011 | 126 | 3500
Z-2240-2011 | 58487 | mattress, flotation therapy, non-powered | IKY | Nonpowered flotation therapy mattress. | General Hospital | General Hospital | 510(K) Exempt | Stryker, Impression, Non-Powered Mattress, Model 2980, Stryker Medical, Portage, MI. |  |  | April 29, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | Affected mattresses may not conform to federal flammability standards. | Other | Stryker Medical sent an URGENT MEDICAL DEVICE RECALL letter dated April 29, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate all affected mattresses , remove the existing mattress cover, and replace it with an enclosed mattress cover. Customers were then to notify Stryker to confirm repairs. If the mattresses were loaned or sold, Stryker requested customers to notify the new user of the recall. Customers with questions or concerns, were asked to call 269-389-6689. |  | 213 | 213 | Worldwide Distribution - USA including California and the country of Canada. | Terminated | October 07, 2011 | 161 | 213
Z-2243-2011 | 58533 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE , ONCOR and PRIMUS digital accelerators with syngo RT Therapist v4.1 and 4.2 system; |  |  | March 31, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens found a potential safety issue related to patient treatment with table offsets on syngo RT Therapist (RTT) 4.1 and 4.2 which could result in dose being delivered to wrong location. | Software design | Siemens initiated an Urgent Medical Device Correction letter on March 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:use only relataive table setup with all planned table values '0' if applicable.In cases where the treatment plan contains at least one table parameter other than '0' to consider the following:Use only one PVG plus resulting offset per fraction.Once an offset has been applied do not use 'apply offset' at RTT for any further offset.Select the PVG and treatment beams in the same auto-sequence. In casethis is not possible, the table position for each treatment beam, in the fraction, needs to be verified separately (especially for treatments with beams having different table posiitons. If an in-session resumption or individual download of beams has to be performed, make sure that the treatment table is in the correct position.In order to prevent any unintended table movement select manual table movement for all directions in the syngo RT Therapist service UI.Scenario 2 - Manual change of table posiiton followed by a PVG calculated offset.Always deliver the PVG's at the planned table position. Avoid manual change of table posiiton prior to applicaiton of PVG.Select the PVG and treatment beams in the same auto-sequence. In case this is not possible, the table position for each treatment beam, in the fraction, need to be verified separately (especially for treatments with beams having different table positions).If an in-session resumption or individual download of beams has to be performed, make sure that the treatment table is in the correct position.In order to prevent any unintended table movement select manual table movement for all directions in the syngo RT Therapist service UI.Customers were informed that a software update would be provided as soon as possible.Customers were asked to include the Customer Sa |  | 164 | 164 active devices | Worldwide Distribution - USA including CA, DE, FL, KY, LA, MA, MI, MO, NC, NE, NJ, NY, OH, PA, SD, TN, UT, WI, and WV and the countries of Australia, Belgium, Canada, Croatia, Egypt, France, Germany, Hungary, India, Ireland, Italy, Japan, Malaysia, Netherlands, New Zealand, Norway, P.R. China, Philippines, Poland, Republic of Korea, Russian Federation, Saudi Arabia, South Africa, Spain, Turkey, and the United Kingdom. | Terminated | March 30, 2012 | 365 | 164
Z-2241-2011 | 58582 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Slice CT Scanner System is a computed Tomography X-ray system, and has a Part number 989605858501. |  |  | February 09, 2011 | 2011 | Philips And Neusoft Medical Systems Co., Ltd. | Dimension measurement displays on a combined image were found inaccurate after the clinician merged multiple image series into one image. | Other | Neusoft Medical Systems sent an Urgent Device Correction Notification letter dated February 9, 2011, via certified mail to their distributor NEUISYS. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to discontinue use of combined images. Instead, use MPR Batch for diagnostic reviews until the planned upgrade is provided to the product.Neusoft Medical System and Philips Medical System released the mandatory Field Change Order (FCO) by March 31, 2011. All nine affected systems will be updated with aforementioned kits in six months after the release.Consignees were instructed to contact Service Support Department by e-mail to nms-service@neusoft.com |  | 9 | 9 units distributed in US | Nationwide Distribution including CA, GA, LA, NC, OH, TN and TX. | Terminated | August 24, 2011 | 196 | 9
Z-2239-2011 | 58685 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂż Soft Mesh 4"x6" (10cm x 15cm) Flat, Product Code: 0117010. |  |  | April 25, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeling: complaints of product labeled as BardÂż Soft Mesh 4" x 6" containing BardÂż Soft Mesh 6" x 6" | Labeling mix-ups | Davol issued an URGENT: MEDICAL DEVICE RECALL notification letter to customers dated April 25, 2011, informing them the recall involving a single lot of the BARD Soft Mesh - Lot# HUUF0530. This recall has been initiated because this lot number may contain a different size patch than is indicated. Customers were instructed to immediately examine inventory and quarantine product subject to recall. Customers are asked to identify and notify their customers to whom recalled product may have been further distributed. A Recall Effectiveness Check reply form should be completed and returned to Davol via fax (401-825-8753). For directions on returning product or for questions regarding the recall, contact Davol Customer Service Department at 1-800-556-6275 or C.R. Bard's Medical Services &amp; Support Department at 1-800-562-0027. |  | 198 | 198 units | Nationwide Distribution | Terminated | October 28, 2011 | 186 | 198
Z-2238-2011 | 58789 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Maquet Magnus Carbon-fibre table top; Device Part Number 1180.16xx (x represents a digit from 0-5 and F-for the US market) |  |  | April 15, 2011 | 2011 | Maquet Inc. | Maquet GmbH's quality monitoring detected that incorrect cover sheets were installed with some Magnus carbon fibre table tops. | Device Design | Maquet sent an Urgent Medical Device Corrective Action letter dated 4/15/2011, along with a Action Response Form, via Federal Express to their three customers informing them the incorrect cover sheets were installed with some of the affected "MAGNUS 1180.16F0 carbon-fibre table tops." The firm will have a Service Technician contact you to service the OR table, and will be exchanging the cover sheet free of charge. To the customers, please follow the instructions on the response form. At this time there have been no reports that injury has occurred.If you have additional questions, please contact yur local MAQUET Customer Service at 1-888-627-8383 |  | 3 | 3 | LA, MO, VA | Terminated | August 21, 2013 | 859 | 3
Z-2246-2011 | 58565 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX 6.5 Client Software. |  |  | April 05, 2011 | 2011 | AGFA Corp. | Software issue that occurs after dictation of a study. | Other | Agfa Corp. sent an "URGENT SAFETY NOTICE" customer letter via Agfa email to customers on 4/14/2011, along with a "URGENT SAFETY NOTICE FEEDBACK FORM," The letter identified the product , the problem, and the action to be taken by the customer.It was discovered a software issue in the IMPAX 6.5. The upgrade has already been released to address the software problem.Customers were instructed to fax back the feedback form as soon as possible to (864) 421-1664.If you have any questions, please contact your local Agfa Healthcare organization. |  | 95 | 95 units | Nationwide Distribution -- including states of WI, and SC. | Terminated | September 15, 2011 | 163 | 95
Z-2245-2011 | 58328 | [6 empty] | ***REF INT411***OSSEOTITE Tapered Certain Implant 4 x 11.5mm***Sterile Using Radiation***Rx Only***BIOMET 3i Dental lberica S.L.WTC Almeda Park, Ed, 1 Planta 1 Pl. de la Pau s/n08940 Cornella de Llobregat (Barcelona) Spain***Phone +34 934-705-500; Fax +34 933-717-849. |  |  | March 10, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i is recalling their product Osseotite NT Certain Dental Implant INT411 . The internal thread was not manufactured correctly. | Other | Biomet 3i, LLC sent an Urgent Medical Device Recall letter dated March 14, 2011, via email or fax to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A small percentage of the implant lots may have a condition affecting the internal thread. The condition, if present, will prevent the cover screw, temporary healing abutment, final restorative abutment, etc. from engaging the implant. Customers were instructed to contact Customer Service at 1-800-342-5454 if this condition is recognized upon attempting to assemble any of the screws with the implant.Customers were instructed to return the implant to the firm for replacement to the following address:BIOMET 3iRegulatory Compliance Department4555 Riverside DrivePalm Beach Gardens, Florida 33410Incident No. CC117886Customers were also instructed to respond with the attached Response Fax or email to DomesticComplaints@Biomet.com.For any questions regarding this recall call 561-776-6906. |  | 236 | 236 | Worldwide Distribution - USA (nationwide) and the countries of Europe, Japan, Asia, and South America. | Terminated | January 25, 2012 | 321 | 236
Z-2244-2011 | 58550 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERT - FA Reagent Bottle |  |  | April 13, 2011 | 2011 | bioMerieux, Inc. | The charcoal suspension in a portion of the lot does not meet the intended density. | Other | BioMerieux sent an Urgent Product Removal Notice dated April 13, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Evaluate positive results adhering to their institution's protocols for determining false versus true positives. Negative bottles should be checked by smear and/or subculture at some point prior to discarding as negative. Any remaining inventory of BacT/ALERT FA culture bottle lot 1027481 should be destroyed in accordance with their facility's policies and procedures. An Acknowledgement and Product Replacement Form was included to be completed and returned to bioMerieux or credit or replacement of unused bottles. |  | 327700 | 327,700 bottles | Worldwide Distribution - USA (nationwide) and the countries of China, Tokyo, South Korea, Canada, Sweden, Austria, Poland, and Belgium | Terminated | April 03, 2012 | 356 | 327700
Z-2293-2011 | 58545 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics version 9.0, Siemens Medical Solutions, USA, Inc., Ann Arbor, MI. |  |  | March 24, 2011 | 2011 | Siemens Medical Solutions, USA, Inc | If one or more graphs are in a table on the report, and if any of those graphs are deleted, it may lead to reports failing to save and/or marking the study as read might leave the study in an unread state or might mark the study as read but leave the report with a "Preliminary" watermark displayed. | Software design | The firm, SIEMENS, sent a "Customer Safety Advisory Notice" letter dated March 24, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to place graphs in sections and not tables in the report; avoid deleting graphs from the reports in run-time and include this safety advisory notice in their operator's manual where it should remain until the service patch is applied. SIEMENS will be issuing a service patch to fix the defect. The letter states the patch will be available in the next few weeks.If you have any questions, call +1 (734) 205-2400. |  | 29 | 29 devices | Nationwide distribution: USA including states of: CA, CT, FL, ID, MD, MI, MN, NJ, NY and PA. | Terminated | July 10, 2012 | 474 | 29
Z-2233-2011 | 58373 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | MicroStream Filterline Infant/Neonatal products: Part # 006285 - FILTERLINE H SET INF/NEO 25 UN PHYSIO (Inf/Neo Airway Adapter). The adult/pediatric airway adapter and the adult/pediatric FilterLine and VitaLine sets are not affected by this recall. |  |  | March 25, 2011 | 2011 | Oridion Medical 1987 Ltd. | fine plastic strands on inner surface of the infant neonatal airway adapter may become dislodged and inhaled by the patient | Equipment maintenance | The firm, Oridion, sent an "Microstream Filterline Field Correction Notice" letter dated March 25, 2011 to its US and foreign customers via telephone, email with the documentation discussed in the call, and letter. The letter described the product, problem and actions to be taken. The customers were instructed to place all affected product held in their inventory on ship hold; gather these units in their inventory and at the customer sites and exchange with the new product; inspect filter line starter (sample) kits shipped with Capnostream 20 and Microcap patient monitors for affected product, if affected Infant/Neonatal filter lines are found, remove them from the starter (sample) kit; send a purchase order in order to obtain replacement product for units in their stock and at customers sites, and complete and return the Inventory Reporting Form Excel Spread Sheet within two weeks of receiving the email even if they or their customers do not have the affected inventory. Oridion shall conduct a removal of the products from the related production lots from the field and an immediate replacement with product from subsequent production lots that meet all requirements. This action will be conducted by Oridion's distributors or by Oridion directly in some cases. Oridion anticipates that new product will be available for shipment during the week of April 4th.Questions regarding the recall can be directed towards the Customer Service Coordinator at: 1-888-674-3466 x1234 or NA-FCA@oridion.com. |  | 350 | 350 units in US (14 boxes with 25 units ea) | Worldwide distribution. | Terminated | March 19, 2012 | 360 | 10525
Z-2291-2011 | 58420 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LOGIX Order Entry (OE) Software, all versions, catalog 2M8410; Baxter Healthcare, Deerfield, IL 60015; |  |  | March 29, 2011 | 2011 | Baxter Healthcare Corp. | The LOGIX OE automated backup does not retain the three most recent backup files and instead will only retain the last three backups of the previous year. The software stores backup files in numeric order rather than chronological date order. | Software design | Baxter Healthcare Corp. sent an URGENT DEVICE CORRECTION letter dated March 29, 2011, via first class mail to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. It was discovered there is an issue with the LOGIX OE automated backup not retaining the three most recent backup files and instead will only retain the last three backups of the previous year. The accounts were requested to turn off the automated backup function on the LOGIX OE Software, and perform manual backups following the instructions included in the Urgent Device Correction letter, and retain a copy of the letter at each computer on which LOGIX OE Server is installed. The customer were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Medical Information Services at 1-800-422-2751. |  | 84 | 84 units | Worldwide Distribution -- USA including Puerto Rico, and the country Canada. | Terminated | January 05, 2017 | 2109 | 84
Z-2261-2011 | 58021 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Case Label: McKESSON Medi-Pak Uretheral Catheter Tray With Red Rubber Catheter Sterile Single Use Disposable 20 Trays Per Case Reorder No. 37-202 Marketed By McKesson Corporation McKesson Medical-Surgical, Richmond, VA 23228 Made in the U.S.A. |  |  | January 20, 2011 | 2011 | Nurse Assist, Inc | The trays contain Triad Group Sterile Lubricating Jelly packs recalled by Triad Group because of the possibility of inadequate sterilization. | Other | The firm, Nurse Assist Incorporated, sent a letter dated January 20, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory immediately; quarantine the product; identify their customers and notify them at once about the product recall (Nurse Assist will provide bright labels to affix to your affected inventory); provide their customers with labels as well, and complete and return the attached RECALL ACKNOWLEGEMENT form via fax to: 817-231-1500 or mail to: Nurse Assist, Inc., 3400 Northern Cross Blvd., Forth Worth, TX 76137.Only the product and lots identified on the attached list are affected.If you have any questions or to request labels, please call Nurse Assist Customer Service Monday through Friday, from 8:00 am to 5:00 pm Central Time at 1-800-649-6800. |  | 1504 | 1504 cases | Worldwide distribution: USA (nationwide) and countries including: Canada, Dubai and Saudi Arabia. | Terminated | October 25, 2012 | 644 | 124789
Z-2294-2011 | 58600 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll E Series BLS Defibrillator/Pacemaker/Monitor |  |  | April 21, 2011 | 2011 | ZOLL Medical Corporation | Zoll E Series BLS Defibrillator may fail to auto charge and deliver energy | Software design | The firm, Zoll, sent an "Urgent Medical Device Correction" letter dated April 22, 2011 via UPS to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately alert all E Series BLS users to this problem; instruct users to not touch patient and keep the patient still during ECG analysis, Instruct users that if they experience this problem, to press the Analyze button again, and contact ZOLL Technical Service for the software upgrade kit. The letter also states that affected units can remain in service until the software has been updated. If you have any questions, please contact 24/7 technical support at 1 (800) 348-9011 or +1 (978) 421-9460. |  | 874 | 874 units | Worldwide distribution: USA (nationwide) and countries of: Australia, Canada, Finland, Germany, Malaysia and United Kingdom. | Terminated | January 23, 2014 | 1008 | 874
Z-2292-2011 | 58601 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | WaveLight FS200 Femtosecond Laser System |  |  | April 18, 2011 | 2011 | Wavelight AG | The software calculating the patterns were found to has a rounding error for certain parameter ranges. These rounding errors could result in an incomplete cut in the donor or patient tissue. | Other | WaveLight AG sent an URGENT MEDICAL DEVICE CORRECTION letter April 18, 2011 to the affected customers. The letter identified the product, problem and the action to be taken by the customer. The letter advised the customers todiscontinue keratoplasty cuts with the WaveLight FS200 until the new version can be installed. Alcon will contact you to schedule a visit of a service engineer at your site, to disable the keratoplasty function in your device, until a new software version can be installed to fix this problem. If you have experienced a product problem or adverse event related to this medical device correction, please contact Alcon Medical Safety at 1-800-561-6466, option 3.Customers were requested to complete the attached reply form to acknowledge receipt of this notification and fax it to ( 817) 916-9087. |  | 55 | 55 units | Worldwide Distribution - USA including states of TX, KS, NC, CA, FL, OH, SC, MD, VA and NY and countries of Switzerland, Germany, Greece, United Kingdom, Ireland, Spain, Italy, Australia, India, Malaysia, France, Finland, Saudi Arabia, Czech Republic, Poland, China and South Korea and Iran. | Terminated | September 23, 2011 | 158 | 55
Z-2383-2011 | 58650 | No | KNY | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gyrus ACMI, REF 25 BX, PTFE Coated 3 cm Flexible Tip Guidewire, .035" x 150 cm, Sterile EO, Rx Only, Gyrus ACMI, Inc. 136 Turnpike Road,Southborough, MA 01772-2104 USA |  |  | April 15, 2011 | 2011 | Gyrus Acmi, Incorporated | Gyrus ACMI, Inc. wants to bring a recall issue to their customers attention concerning one lot of their 25BX-PTFE Coated 3 cm Flexible Tip Guidewire. As a result of a manufacturing error in the production this lot of guidewires, some units of product were loaded in the dispenser backwards. In such cases, the stiff end of the guidewire would be dispensed from the product first and, if undetected, | Process control | The firm, Olympus ACMI, sent an "Urgent: Medical Device Recall" letter dated April 15, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Inspect their stock to confirm if any of the affected lot of products remains in their possession2)Immediately cease any further use of the affected product, remove, quarantine and return it. 3) Call their Gyrus ACMI customer service representative at 1-888-542-7266 to obtain a Returned Goods Authorization 4) Return the product to: Attn: PTFE Guidewire Recall RGA#_________Gyrus ACMI LP, 6845 Wedgwood Court, Maple Grove, MN 553115) Complete and return the RECALL REPLY FORM via fax, even if they are returning no product, to: Regulator Affairs (508) 804-2624. The letter states that the customer will receive repeat mailings until they have faxed in the form. Additionally, If they have further distributed the product they are requested to identify and notify their customers at once.If you have any questions, please do not hesitate to contact customer service at 1-888-524-7266. |  | 145 | 145 | Nationwide distribution: USA including states of: CA, GA, HI, IN, KY, MT, NY, NC, PA, TX, UT, and WA. | Terminated | August 23, 2012 | 496 | 145
Z-2295-2011 | 58612 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50Âż Ventilator. Model #HT50 H1-B |  |  | January 13, 2009 | 2009 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical Instruments confirmed that the Newport HT50 ventilators that were shipped in December 2008 were without the label on the top of the ventilator that reminds users to keep HT50 ventilators plugged into an external power supply whenever possible to preserve the longterm capacity of the battery. | Other | Newport Medical Instruments, Inc. sent an FIELD CORRECTION letter dated January 13, 2009, to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. The Newport HT50 ventilators that were shipped in December 2008, were without the label on the top of the ventilator that reminds users to keep HT50 ventilators plugged into an external power supply whenever possible to preserve the longterm capacity of the battery. The customer were also requested to complete the attached customer REWORK INSTRUCTIONS form and fax it to (714) 427-0489. If you have any further questions, please call (714) 427-5811 x 344 |  |  |  | Worldwide Distribution-- USA including states of OH, VA, WA, WI and country of Turkey. | Terminated | June 01, 2011 | 869
Z-2247-2011 | 58471 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Philips MicrostreamÂż Model M1923A FILTERLINE H SET INF/NEO 25UN PHILIPS |  |  | April 07, 2011 | 2011 | Philips Healthcare Inc. | Fine plastic strands on inner surface of the infant/ neonatal airway adapter may become dislodged and inhaled by the patient. | Equipment maintenance | The firm, Philips, sent a Field Safety Notice titled "URGENT MEDICAL DEVICE RECALL" letter dated April 2011 to its customers. Customer notifications outside the US will be managed by Philips representatives in each affected geography. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify all products from affected lots; remove from inventory and dispose in accordance with local regulations (new products will be sent proactively to each affected customer). Direct accounts were instructed to notify additional customers to whom the product has been distributed. These customers are instructed to follow the instructions in the "Action to be taken by Customer/User" section of the Notice. Questions regarding the recall can be directed towards the customer's local Philips representative at 1-800-722-9377. |  | 25400 | 25,400 units (1016 boxes) | Worldwide distribution. | Terminated | March 14, 2013 | 707 | 25475
Z-2382-2011 | 58629 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50 Ventilator. |  |  | May 14, 2010 | 2010 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical has confirmed that the Newport HT50 hard shell Emergency Preparedness (EP) Storage Cases must be relabeled to caution users to follow current labeling guidelines for charging the battery for 8 hours but not longer than 24 hours at any one time. | Other | Newport Medical sent an FIELD CORRECTION NOTICE to all affected customers dated May 14, 2010. The letter identified the product, the problem and the action needed to be taken by the customer. Newport Medical also provided a revised labeling for the Emergency Preparedness (EP) Storage Case and an instruction sheet describing where to apply the new labels. Customers were instructed with the information that all the EP storage cases must be relabeled to caution users to follow current labeling guidelines for charging the battery for 8 hours but not longer than 24 hours at any one time. Customers were instructed to complete the attached acknowledgement form and fax it to " Regulatory@NewportMedical.com " Customers with any questions should call Technical Support Department at (800) 451-3111 for US customers, and (714)427-5811 x500 for international customers. |  | 1113 | 1,113 EP cases | Worldwide Distribution -- USA including states of PA, TX, AK, GA, CA, MD, KY, FL, VA, LA, NC, OK and countries of Canada, Colombia, Armenia and Japan. | Terminated | June 01, 2011 | 383 | 1113
Z-2387-2011 | 58652 | No | DJG | Opiate test system. | Toxicology | Toxicology | 510(k) | DrugCheck, Manufactured by: Express Diagnostics Int;l Inc. 1550 Industrial Drive, Blue Earth MN 56013 USA, Model numbers : |  |  | March 18, 2011 | 2011 | Express Diagnostics Int'l., Inc. | Express Diagnostics Int'l, Inc. has initiated a medical device recall on various model numbers of DrugCheck devices,This recall was due to incorrect product expiration dating after findings from an FDA inspection. Use of expired product may not work properly, and may give incorrect screening results. | Incorrect or no expiration date | The firm, Express Diagnostic Int'l, Inc., sent an "Urgent Recall of DrugCheck Product" letter dated April 19, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine the product; notify Express Diagnostic of the quantity even if no product is remaining; complete and return the DISTRIBUTOR/CUSTOMER RECALL CONFIRMATION FROM via fax to 1-507-526-2252 as soon as possible. In addition, if they have further distributed the product to identify their customers and notify them of the recall. Provided procedures to be followed for them to destroy the quarantined product. Express Diagnostics will replace the DrugCheck Drug Screen devices.If you have any questions, please contact the Quality Assurance Manager at 507-526-3951, ext. 117. |  | 6925 | 6925 devices | Worldwide distribution: USA (nationwide) including states of: FL, IA, MD, OK, and WA; and countries including: AUSTRALIA and ITALY. | Terminated | July 06, 2012 | 476 | 6925
Z-2386-2011 | 58431 | No | OHH | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | Enforcement Discretion | AC Power Adaptor (transformer), TENPAO Model U090100Q31; used with the Pump In Style Advanced family of breast pumps; Medela Inc., McHenry, IL 60051-0650; Made in China; Article #9207010. |  |  | April 04, 2011 | 2011 | Medela Inc | The power adaptor (transformer) used with the Pump In Style Advanced family of breast pumps have power cords that may easily become damaged if not properly handled. | Device Design | Medela Inc., sent an " IMPORTANT FIELD ACTION " letter dated April 4, 2011, to all affected customers via The United States Postal Service ( USPS) United Parcel Service ( UPS) and E-mail .The letter identified the product, the problem, and the action to be taken by the customer. It was discovered the power adaptor/transformer used with the Pump in Style Advanced family of breast pumps can develop damage in the form of twists, kinks, bends, bulges or exposed wires in the power cord. Use of the power adaptor with a damaged cord could result in a short in the cord resulting in sparks, smoke, burning and melting of the transformers. If the users have a damaged power adaptor or power adaptor cord, they were requested to discontinue use of the adaptor and visit Medela's website at www.poweradaptorhelp.com to request a replacement adaptor. Questions about the recall process were directed to Medela Customer Service at 1-855-367-0104 or e-mail at poweradaptor@medela.com. Distributor/retailers were requested to post the recall notice within their establishment visible to their customer base, and to notify their customers of the recall. Information on the recall was posted on Medela's website www.medelapoweradaptorhelp.com on April 4, 2011. |  | 478823 | 17,681 units alone, 461,142 units with products | Worldwide Distribution - Nationwide Distribution and Puerto Rico and countries of Canada, Bermuda, Cayman Islands, Hong Kong and Haiti.CHI-DO proposes phone call audit checks of 2 consumers, 2 retailers and 1 hospital account by each district. | Terminated | October 16, 2012 | 561 | 478823
Z-2385-2011 | 58199 | No | JEY | Bone plate. | Dental | Dental | 510(k) | ***BIOMET MICROFIXATION BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***6 HOLE REGULAR DOUBLE Y PLATE***"1.5MM SYSTEM PLATE DOUBLE Y-SHAPE, REGULAR 0.6 MM TITANIUM***NON-STERILE PRODUCT*** |  |  | July 28, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling their 1.5 6 Hole Reg Double Y, Lots 785710, 753120, 693500 and 733600. This product is labeled as 01-7110-K (1.5 6 Hole Reg Double Y), however the package contains a 01-7112-K (1.5.6 Hold Long Double Y). | Labeling mix-ups | The firm, Biomet Microfixation, sent an "URGENT MEDICAL SAFETY ALERT NOTICE" letter dated July 28, 2010 to its customers/distributors. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their stock and ensure that the plates were properly distributed into their trays and that they have the sufficient quantity of plates that they need. In addition, they were instructed to complete and return the attached Inventory Reconciliation form via fax to: Attn: Customer Relations Specialist at 904-741-9425 and notify the recipient of this notice, if they have further distributed the product.If you have any questions, please contact the Customer Relations Specialist at 800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 180 | 180 | Worldwide distribution: USA (nationwide) including states of: FL, IN, KY, MT, NC, NE, NY, OH, TX, VA, and WI; and countries including: Argentina, Canada, Colombia, Czech Republic, Japan, Korea, Mexico, Norway, South Africa, Spain and Turkey. | Terminated | June 24, 2011 | 331 | 180
Z-2384-2011 | 58632 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50 Ventilator, Model # HT-50-H, HT50-HB, HT50-H1, and HT50-H1B |  |  | July 22, 2009 | 2009 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical has confirmed that the HT50 Dual Pac Ventilator is not compatiable with the external battery pack (p/n BAT3300A) supplied with the previous single battery ventilator. The Dual Pac system has a slightly higher amperage draw than the single battery and as a result can cause the fuse in the external battery pack to blow making the battery inoperable. | Other | Newport Medical sent an IMPORTANT PRODUCT NOTICE on July 22, 2009, to all affected customers who purchased the Newport HT50 Dual Pac Ventilator with the external battery pack (p/n BAT3300A) supplied with the previous single battery ventilator. Newport Medical also sent an upgrade kit (FSK3301) which includes instructions for changing the fuse on the external battery. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to not use an external battery with the new Dual Pac Ventilator until they have completed the fuse replacement. Customers with any questions about this fuse replacement or require additional upgrade kits, were instructed to contact Newport Medical's Technical Support Department at (800) 451-3111 or (714) 427-5811 x500. |  | 11 | 11, 333 units total | Nationwide Distribution | Terminated | June 01, 2011 | 679 | 11
Z-2380-2011 | 58346 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Material ULT2175, ULTIMATE LINE TRAY, CENTURION ULTIMATE LINE TRAY, REORDER ULT2175, STERILE, LATEX FREE, CENTURION MEDICAL PRODUCTS CORP. WILLIAMSTON MI 48895, USA 800.248.4058 www.centurionmp.com |  |  | March 31, 2011 | 2011 | Centurion Medical Products Corporation | This voluntary recall was initiated because the vials of Bacteriostatic Sodium Chloride Injection, USP, 0.9% utilized to make the products may contain visible particulates. | Material/Component Contamination | Centurion Medical Products Corporation issued press on April 1, 2011, and sent an Urgent Voluntary Recall letter dated March 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the kits immediately and instructions were provided for the return of the kits.Customers were also asked to complete a form to fax or email back to the firm regarding the number of kits the firm has in their possession.Customers were instructed to provide recall notification to other firm's who received the recalled kits from their direct account.For any questions regarding this recall call 631-924-4000, ext 142. |  | 7160 | 7160 | Nationwide Distribution | Terminated | September 13, 2011 | 166 | 71471
Z-2392-2011 | 58665 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | URI-DRAIN Medium Male External Catheter With Two-Sided Adhesive Strap and Skin Protective Wipe |  |  | April 29, 2011 | 2011 | Covidien LP | A component of the Uri-Drain device recalled the Skin-Prep (manufactured by H&P; Industries-Triad) a protective wipe due to bacterial contamination | Other | The firm, Covidien, sent an "URGENT PRODUCT RECALL" letter dated April 29, 2011 to its customers via US Postal Service. The letter described the product, problem and actions to be taken. The customers were instructed to immediately stop using the product, place in quarantine, and return any inventory. Distributors are requested to notify their customers. All customers are required to respond to the notice regardless of remaining inventory. Responses will be tracked for follow-up and effectiveness checks.For Customer Service questions regarding the recall please telephone: 1-800-962-9888 option 1 (weekdays, 8am to 6:30pm Eastern) and email to: sdfeedback@covidien.com. Should you have any medical questions concerning this recall, please call 508-261-6393. |  | 5760 | 5760 units | Worldwide distribution: USA (nationwide) and countries including: Canada and Bermuda. | Terminated | May 29, 2015 | 1491 | 62352
Z-2395-2011 | 58518 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ES0101, AEM Disposable Curved Scissors, 1/2", 35cm, Sterile R, Encision Inc., Boulder, CO 80301 |  |  | April 07, 2011 | 2011 | Encision, Inc. | Sterility may be compromised due to holes in packaging. | Package design/selection | Encision, Inc. sent an Important Safety Notice letter dated April 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to either carefully inspect any remaining pouches for readily visible holes in the Tyvek pouch, or return the affected lots to Encision for inspection.Customers were also asked to fill in and fax back the attached Confirmation to 303-444-2693. |  | 598 | 574 Ten-Count Boxes, 24 Kits | Worldwide Distribution - USA (nationwide) and the countries of Australia and New Zealand | Terminated | August 09, 2011 | 124 | 2052
Z-2401-2011 | 58568 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Xpert Infinity Software, V 4.1a. |  |  | April 01, 2011 | 2011 | Cepheid | A mismatch of Patient ID or Sample ID and the Host Order from the hosted LIS and its requested test results could occur when running in batch mode and cartridges are introduced in rapid succession. | Software design | Cepheid sent an Corrective Action Notification letter dated April 13, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cepheid recommended customers using 4.1a software on the INFINITY 48 System that is interfaced to a LIS with Host Ordering to make these alterations in operationof the system.After scanning the Patient ID (optional). Sample ID and Cartridge barcode, the host test order information should apear highlighted in the Order Test workspace.Once the highlighted host order information appears select SUBMIT.Warning. If the host test order informaiton is not highlighted, cancel the order by selecting CLOSE. Re scan Patient ID (optional). Sample ID and cartridge barcode and verify highlighted host test order before SUBMIT. Contact Cepheid Technical Support at 888-838-3222 in the US or 33 5 63 82 53 19 outside of the US for any questions regarding this Corrective Action Notice. |  | 19 | 16 US, 3 ROW | Worldwide Distribution - USA including CA, UT, TX, NJ, FL, NC, CO, IN, VT, MN and the countries of the UK, Denmark and Saudi Arabia. | Terminated | June 20, 2011 | 80 | 19
Z-2403-2011 | 58636 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Capsule Neuron UMPC is a component in the DataCaptor" Connectivity System. |  |  | August 25, 2010 | 2010 | Capsule Tech Inc. | Software: Data cached on the Neuron due to server unavailability is not being delivered to the server after server availability is restored.If a compact flash failure occurs, the file system may be corrupted and result in system hangs, lock-ups or crashes.Batteries can discharge completely in a Capsule Neuron" UMPC even though it is connected to power source | Software design | Capsule Tech initiated three corrective action notifications for the Neuron" UMPC units three failure modes as follows:a) Data cached on the Capsule Neuron" UMPC due to server unavailability is not being delivered to the server after server availability is restored:i) A customer advisory was provided to a subset of consignees, along with DataCaptor software version 6.6.5 on December 1, 2010.ii) A customer advisory was provided to the remaining consignees, along with DataCaptor software version 6.6.7 on January 27, 2011.If a compact flash failure occurs, the file system may be corrupted and result in system hangs, lock-ups or crashes:i) A customer advisory was provided to all consignees on March 23, 2011, via phone call, email and FedEx;.The correction will be provided as a software update. A formal recall notification letter will be provided to all consignees alerting them to the availability of the correction once it is released, via phone call, email and FedEx. c) Batteries can discharge completely in a Capsule Neuron" UMPC even though it is connected to power. Note: the Capsule Neuron" UMPC operates normally when connected to power.i) A customer advisory was provided to all consignees on April 23, 2010, via FedEx.For questions regarding this recall contact Capsule technical support team at (800) 260-9537 or e-mail support@capsuletech.com.. |  | 3144 | 3,144 Capsule Neuron UMPC units | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | March 20, 2012 | 573 | 3144
Z-2402-2011 | 58557 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cholestech LDX Lipid Profile"GLU test cassettes, Catalog Number 10-991, Lot # 245104 |  |  | February 18, 2011 | 2011 | Biosite Inc Dba Innovacon Inc. | The recall was initiated due to an issue related to an error in the encoding of calibration coefficients in the sample cassette that cause the Cholestech LDX analyzer to report glucose values as &gt;500 mg/dL regardless of the true value of the specimen. | Pending | Alere San Diego, Inc ( formerly known as Biosite Incorporated sent an URGENT MEDICAL DEVICE RECALL letter on February 18, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Customers were requested to:(1) Discontinue use and/or sale of Cholestech LDX Lipid Profile-GLU test cassettes, Catalog Number 10-991, Lot # 245104.(2) Discard in accordance with your local regulations, all units of the above listed product.(3) Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation.(4) If you have forwarded the product listed above to another laboratory, please provide a copy of this letter to them.(5) Complete and FAX the enclosed Verification Form within 10 days to confirm your receipt of this notice, to indicate the number of boxes of test cassettes discarded from your inventory and replacement product will be sent to you.For questions about the information contained in this notification, customers were instructed to contact Alere Technical Services at 1-877-308-8289 (24 hours a day, 7 days a week) or via e-mail at service.cholestech@alere.com. |  | 5305 | 5,305 Boxes | Worldwide Distribution - USA ( nationwide ) including the countries of South America ( EL Salvador) , Europe ( Denmark, France, Norway and Switzerland), Bahrain, Kuwait, South Africa, and United Arab Emirates. | Terminated | April 02, 2012 | 409 | 5305
Z-2407-2011 | 58267 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.50.00 and above |  |  | May 27, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: For a specific clinical setup, the MLC leaf positions on the exported DRR images might not match the MLC leaf positions displayed in XiO. | Software design | The firm, Elekta, issued an "important Safety Notice" undated, to its customers. The notice indicates the problem and the clinical impact. There is no reasonable workaround for this issue. The notice states a future XiO Release will resolve the issue and the customers will be notified when this solution is available. A return postcard was included with the mailing to verify consignee received the notice. For sites not returning the postcard within one month, the recalling firm will contact them via phone or email to confirm receipt of the notice. The customers were instructed to distribute this notice to any and all users of the CMS software at their organization who are potentially affected by this issue.If you have any questions, please call 314-993-0003 or 800-878-4267. |  | 415 | 415 | Nationwide distribution, including Puerto Rico, to medical facilities. Distribution was also made to Military and Government consignees. Foreign distribution was made to Albania, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lithuania, Malaysia, Malta, Mexico, Netherlands, New Zealand, Poland, Portugal, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, and United Kingdom. | Terminated | April 10, 2014 | 1414 | 415
Z-2405-2011 | 58664 | No | EXF | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Drainage Bag, Model 600D, 600 mL, Single use, Ethylene Oxide sterilized, Remington Medical, Inc., 6830 Meadowridge Court, Alpharetta, GA 30005. |  |  | January 12, 2011 | 2011 | Remington Medical Inc. | There have been complaints of the Drainage Bags leaking from the outlet and the inlet port. | Other | Remington Medical Inc.sent an Urgent Medical Device Recall letter dated January 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine product subjec to recall. If they further distributed this product, they were to identify their customers and notify them at once of the product recall. Notification to their customers should include a copy of the recall notification letter. Customer were to receive a Return Authorization Form, reference RGA #600-D.For questions regarding this recall call 800-989-0057, ext. 213. |  | 89294 | 89294 units | Nationwide Distribution | Terminated | September 13, 2011 | 244 | 89294
Z-2406-2011 | 58619 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT70 Ventilator. Model #HT70, |  |  | September 20, 2010 | 2010 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical has confirmed a limited number of power cords recently received from their supplier, Glob-Tek, have a slightly oversized plug connector which prevents thecord from locking into place. These cords are fully functional as supplied. However, the cord can be easily disconnected. If the cord is disconnected, the HT70 will alert the user with a power di | Other | Newport Medical sent an PRODUCT NOTICE LETTER, dated September 20, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to discard the original cord and only use the replacement cord to avoid unintentional disconnection of the power cord. Each affected Customer was sent a replacement cord.Customers with any questions regarding the Product Notice were instructed to contact the Technical Support Department at (800) 451-3111. |  | 22 | 22 units total | Worldwide Distribution--USA (nationwide) including the states of MA, PA, OH, IA and Puerto Rico and the countries of Colombia, UAE, Hong Kong, Hungary, India, Japan, Nepal, China, Philippines, Singapore and South Africa. | Terminated | June 03, 2011 | 256 | 22
Z-2404-2011 | 58637 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE or ONCOR Linear Accelerators in combination with 160MLC (Multi Leaf Collimator). Linear Accelerators that may be affected are: |  |  | April 15, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety risk related to the behavior of the 160 MLC when using the "Motion Stop" button followed by "Reset" several times during the treatment of one beam or one segment. | Other | Siemens sent a Consumer Safety Advisory and Urgent Medical Device Correction letter to all affected customers beginning on April 15, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Updated instructions for using the Motion Stop button with the 160-MLC were provided with the letter. Customers were requested to use the Motion Stop button and the corresponding reset procedure only when required. In the unlikely event that the Motion Stop button needs to be used repeatedly during treatment of the same beam or segment, customers were advised to perform an "in session resumption of the beam by selecting 'abort' at the syngo RT Therapist and then re-selecting the beam again for the remaining delivery. In this case, the beam will be re-loaded with the originally prescribed leaf positions. Please consider the already delivered dose for resumption.Customers were asked to include the Customer Advisory Notice in their LINAC System Owner Manual in chapter 'Safety Advisory Letters' where it should remain until the measures defined above were performed. Customers were asked to inform all affected personnel immediately. |  | 221 | 221 active devices | Worldwide Distribution - USA including CA, DE, FL, IL, KY, LA, MA, MI, NE, NJ, NY, OH, PA, SD, TN, UT, WI, and WV. | Terminated | March 16, 2012 | 336 | 221
Z-2409-2011 | 58681 | No | EZD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SureCath Set with Accessories, Coloplast Corp.-MN 06093, Sterile EO., Manufacturer Coloplast -A/S 3050 Denmark. |  |  | April 21, 2011 | 2011 | Coloplast Manufacturing US, LLC | Coloplast has initiated a recall on SpeediCath with Insertion Supplies and SureCath Set with Insertion Supplies because they contain Iodine prep pads which may expire before the overall product expiration date. SpeediCath with Insertion Supplies and SureCath Set with Insertion Supplies are both packaged with a set of accessories which includes Iodine prep pads. The Iodine prep pads included i | Incorrect or no expiration date | Coloplast Manufacturing US, LLC sent a Coloplast "Product Field Correction Notice" dated April 21, 2011. The letter was addressed to "Dear Valued Customer". The letter described the problem and the product involved in the recall. Customers were advised to provide the current inventory levels of the lots affected and the firm will be providing them with new iodine prep pads with sufficient shelf life. This product will include customer instructions, informing the customer to discard the original prep pads and use the new prep pads instead. If they prefer, they could themselves send the affected product back to Coloplast and they in return will attach new prep pads to the product for them.If they have further distributed the product, they advised consignees to contact their customers and to include the letter provided to them to inform them of the recall. |  | 438412 | 438,412 | Nationwide Distribution including AL, AZ, CA, CO, CT, DE, FL,GA, ID, IL, IA, MD, MA, MI, MN, MS, NJ, NY, NC, OH, OK, OR, PA,PR, SC, TN, TX VA, WA, WI. | Terminated | April 24, 2012 | 369 | 564759
Z-2411-2011 | 58785 | No | FNK | Hydraulic adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Eleganza 3 Hospital Bed |  |  | May 02, 2011 | 2011 | Linet Americas Inc | Linet has been made aware that there has been breakage of the Mobi-lift handle on a few beds which may cause failure during patient assist use. | Device Design | The firm, Linet America, notified US customers only via phone on May 5, 2011. A follow up "service advisory information" letter dated May 11, 2011 was sent to all customers. The letter described the product, problem and action to be taken. The customers were instructed to remind/retrain staff to ensure that the Mobi-lift handles are always in the stowed position during patient transport and after use, and complete and return the attached Inspection procedure for each affected Eleganza 3 bed via mail to: Linet Americas, 10420-R Harris Oaks Blvd., Charlotte, NC 28269 or fax to: 704-248-5655. Linet Americas will be receiving a supply of new Mobi-lift design and will begin the process of executing service calls to replace the Morbi-lift handles with the new design.If you have any questions, contact the Director of Quality and Service Operations at 704-248-5650 or 980-406-6446. |  | 1601 | 1601 units | Worldwide distribution: USA (nationwide) including states of: KY, MA, NC, and NY; and country of: Canada. | Terminated | July 19, 2012 | 444 | 1601
Z-2408-2011 | 58205 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 1.0.2 and above. |  |  | September 10, 2010 | 2010 | Computerized Medical Systems Inc | When RT Plans are exported from XiO and RT Structure Sets and Images are exported from Monaco, the DICOM iso-center coordinates may not be accurate. | Software design | Elekta / Computerized Medical Systems Inc began mailing an Important Safety Notice to all affected customers in September 2010. The notice identified the product, the problem, and the action to be taken by the customer. Customers were advised not to mix treatment plan information from multiple planning systems. Customers were asked to return the postcard as proof of receipt and their understanding of the Safety Notice. For questions regarding this recall call 408-830-8000. |  | 1377 | 1377 | Worldwide Distribution - USA (nationwide and Puerto Rico) and the countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 08, 2014 | 1306 | 1377
Z-2412-2011 | 58735 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DuraspanTM Dialysis Catheters, Ultra and Long Term, 15.5 Fr, Sterile EO, Manufactured for r4 Vascular, Inc . 7550 Meridan Circle North, Suite 150, Maple Grove, MN 55369. |  |  | May 03, 2011 | 2011 | R4 Vascular Inc | r4 Vascular, Inc. is conducting a voluntary recall of specific lots of Duraspan Dialysis Catheters. Duraspan catheters are 15.5 Fr dual lumen chronic hemodialysis catheters. Over a six month period, seven patients at five hospitals experienced separation of the luer connector from the extension tube on the Duraspan catheter. Testing during failure analysis disclosed that these lots and other lot | Nonconforming Material/Component | R4 Vascular Inc sent an Urgent Medical Device Recall letter dated May 3, 2011, to all affected customers. The letter described the problem and the product being recalled. Consignees were instructed to stop distributing the product and to quarantine the recalled lots and if they had already distributed the affected product to seek its return from their customers. A template letter to physicians and healthcare facilities was enclosed for Distributors/Retailers to send or hand deliver to their customers.For questions regarding this recall call 1-704-905-2105. |  | 2489 | 2489 kits | Nationwide Distribution including CA, CO, CT, FL, GA, IL, LA, MD, MA, MN, MS, MO, NJ, NC, OH, PA, SC, TX, and WA. | Terminated | July 27, 2012 | 451 | 2489
Z-2415-2011 | 58677 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Pinnacle Destination Guiding Sheath, 6 French, 90 cm, Straight Tip, Tuohy-Borst Valve, Lot MD24, MD31, ME07, &amp; ME14. 510(k) K012812 |  |  | June 24, 2010 | 2010 | Terumo Medical Corporation | Molding defect in destination product that render the device unusable. | Equipment maintenance | Terumo Medical Corporation sent an Urgent Product Bulletin Recall letter dated June 24, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Terumo requested customers to immediately discontinue use, and quarantine any available stock of the listed product codes. In order to facilitate the exchange of these devices, a Terumo Territory Manager contacted all customers by July 9th to confirm receipt of the Terumo Destination Bulletin and scheduled an appointment to visit each customer in person by July 31st to personally manage the exchange. For questions or for further information, call (800) 283-7866 and ask to speak with the Product Recall Coordinator. |  | 214 | 214 units | Worldwide Distribution - USA including AL, AK, AZ, TX, LA, MS, MN, MO, CA, GA, FL, HI, KS, IA, ID, IL, TN, NM, OR, OK, WA, and WI and the countries of Belgium, Brazil, and Panama | Terminated | July 07, 2011 | 378 | 258
Z-2413-2011 | 58591 | No | LOL | Hepatitis A virus (HAV) serological assays. | Microbiology | Microbiology | 510(k) | ARCHITECT HAVAB-M Assay Reagent; a chemiluminescent microparticle immunoassay (CMIA) for the qualitative detection of IgM antibody to hepatitis A virus (IgM anti-HAV) in human adult and pediatric serum and plasma (dipotassium EDTA, lithium heparin, and sodium heparin) and neonatal serum; Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA, List 6L21-25 |  |  | April 20, 2011 | 2011 | Abbott Laboratories | The HAVAB-M reagent lot has reduced specificity and may cause an increase in false grayzone and/or false reactive results. | Nonconforming Material/Component | Abbott Laboratories issued a Product Recall letters dated April 20, 2011 to all affected customers that received ARCHITECT HAVAB-M reagent, lot 93794HN00. The letters informed the accounts of the increased grayzone and/or reactive results and were instructed to discontinue use of the lot and destroy any remaining inventory of the lot according to laboratory procedures. If the account did not have an alternate lot of reagent, they were told to continue using the recalled lot until their lab receives the replacement lot, and to perform supplemental testing with another HAVAB-M method to confirm the results for all specimens that are grayzone or reactive. The accounts were also requested to complete and fax back to Abbott the enclosed customer reply sheet, acknowledging the receipt and understanding of the recall. For further questions regarding this please contact Abbott at 1-877-422-688 |  | 788 | 788 kits | Nationwide Distribution including Puerto Rico | Terminated | February 10, 2012 | 296 | 788
Z-2420-2011 | 58658 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics OrthoPAT -Orthopedic Perioperative Autotransfusion System |  |  | April 29, 2011 | 2011 | Haemonetics Corporation | 2 Models Model 1:OrthoPAT devices with no centrifuge drain spill containment system may have potential for fluid contamination leading to potential spark or smoke may occur are recalledModel 2:OrthoPat equipped with a spill collection drainage system updated with revised Instructions for Use and Quick reference Guide Addendum | Device Design | The firm, Haemonetics, sent an "Urgent Field Recall" letter dated April 29, 2011 to all customers. The letter describes the product, problem and actions to be taken. This recall covers the 2 models. Model 1: OrthoPAT devices without drain ports (Serial numbers 00001 through 02279-T). All OrthoPAT devices without drain ports worldwide, will be removed from distribution and scrapped. Model 2: OrthoPAT devices with drain ports (serial numbers beginning with 02280-T and higher). All OrthoPAT devices with drain ports will remain in clinical use and will receive updated labeling.The customers were instructed to follow the enclosed directions to confirm their possession of any units of this model and arrange for their return to Stericycle; immediately cease use of devices not equipped with a spill collection drainage system ( to determine whether their OrthoPAT device is equipped with the spill collection drainage system, please refer to the attached illustrations/instructions); complete and return the IMPORTANT-CUSTOMER CONFIRMATION sheet via fax to: Stericycle at 1-866-792-5450, and RETURN the OrthoPAT devices NOT equipped with spill collection drainage system (for instructions on how to arrange return of OrthoPAT Devices, call a Stericycle Representative at 1-866-918-8736). A Haemonetics representative will be in contact with the customers facility within 14 working days to assist with the recall actions.The firm, Haemonetics, will update the OrthoPAT labeling. The labeling will have specific instructions to permanently deploy the centrifuge spill collection bag at all times. Device cleaning instructions in the event of a blood spill are also clarified and include specific instructions to unplug the device from the AC outlet in the event of a fluid spill on the device. If a device does experience a fluid spill, new instructions direct the user to return the device to Haemonetics for evaluation and repair. Operator Manuals will be replaced with a new revisi |  | 2976 | 2976 units | Worldwide distribution: USA (nationwide) and countries including: Austria, Australia, Belgium, Switzerland, Canada, China, Czech Republic, Germany, Spain, Finland, France, Great Britain, Greece, Guatemala, Ireland, Israel, India, Italy, Japan, Lithuania, Luxembourg, Latvia, Monaco, Netherlands, Norway, Panama, Poland, Puerto Rico, Russian Federation, Saudi Arabia, Singapore, Slovak Republic, EI Salvador, Turkey and Taiwan., | Terminated | April 21, 2014 | 1088 | 2976
Z-2467-2011 | 58703 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Hitachi HIVISION Preirus/Avius Ultrasound equipped with V03.02 Software - Doppler Trace Measurement program. 510 (k) Numbers: Preirus - K093466 (6/1/10), Avius - K102901 (10/27/10). |  |  | April 18, 2011 | 2011 | Hitachi Medical Systems America Inc | Software anomaly. The Doppler Trace Measurement feature on certain models of the firm's diagnostic ultrasound systems, may, under certain conditions whereby multiple measurements are done sequentially, record incorrect values. The recorded value is displayed to the user on a report; however depending upon the measurement area, displayed velocity values could be lower than at the true velocity lead | Software design | The firm, Hitachi, sent a "DEVICE CORRECTION NOTICE" letter dated April 15, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were inform of the nature of the software error in the Doppler Trace Measurement feature; the circumstances under which this error can occur; and the safety issue involved should incorrect values generated lead to a false diagnosis of stenosis in a patient. The letter provides the customer with a workaround solution in order to avoid this issue. The customer were also instructed to call the HMSA Applications Helpline at 877-871-2777 to reprogram their system to temporarily eliminate the noted safted issue and to immediately have all selections unchecked on the Doppler trace measurement result store setting set up screen. The notification letter also states that Hitachi Medical Systems America (HMSA) will be having their service personnel schedule a visit at each customer location and install a software upgrade on each affected system unit in order to directly address, and permanently resolve this issue. If you need any further information regarding this notice, please contact the Manager of Regulatory Affairs at 800-800-3106 x3720 or email to: thistlewaited@hitachimed.com. |  | 24 | 24 system units | Nationwide distribution: USA including states of: CA, MD, KY, FL, MN, NJ, IA, OH, TX, IL, and IA. | Terminated | February 19, 2013 | 673 | 24
Z-2468-2011 | 58620 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport e360 Ventilator, Model Number: e360 |  |  | October 27, 2009 | 2009 | Newport Medical Instruments Inc | Cregarding e360 ventilators that were shipped with a different software version from what the customers ordered. | Mixed-up of materials/components | Newport Medical Instruments Inc. initiated a field correction on October 27, 2009, to address complaints received regarding e360 ventilators that were shipped with a different software version from what the customers ordered. Newport Medical Instruments provided the correct software to the Newport sales or service providers who were responsible for the units affected. They installed the correct software into the ventilators at the customer sites. No written Correction Notice was provided since the corrections were done at the customers' request.Should you have any questions or require additional information, please call (714) 427-5811 extension 344. |  | 18 | 18 | Worldwide Distribution--USA (nationwide) including states of CA, KY, and Silver City, New Mexico and country of Quito, Ecuador. | Terminated | June 08, 2011 | 589 | 18
Z-2423-2011 | 58663 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | Ringloc + Acetabular Shell limited hole finned 50 mm Size 23 porous coat, for uncemented use, sterile, REF 16-104150, Biomet Orthopedics, Warsaw, IN. |  |  | April 14, 2011 | 2011 | Biomet, Inc. | The firm has initiated this action following an investigation which found that the locking ring may be assembled incorrectly in the acetabular shell. If the locking ring is incorrectly assembled in the acetabular shell, the liner cannot be completely seated in the shell. | Employee error | Biomet Orthopedics sent an URGENT MEDICAL DEVICE RECALL NOTICE DATED April 14, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Immediately locate and remove the affected product.Follow the instructions on the enclosed "FAX Back Response Form" and fax a copy of the Response Form to 574-372-1683 prior to return of product. Customers were to confirm receipt of the notice by calling 800-348-9500, ext 3755 or 3756. For any questions regarding this recall call 574-371-3755 or 574-371-3756. |  | 762 | 762 all products | Worldwide Distribution - USA (nationwide) and the countries of Australia, Netherlands and Japan. | Terminated | September 16, 2011 | 155 | 2069
Z-2541-2011 | 58571 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bio-modular reverse shoulder humeral tray with lock ring, 44 mm 10 degree angle, sterile, single use, REF 11-113685, Biomet Orthopedics, Warsaw, IN. |  |  | April 22, 2011 | 2011 | Biomet, Inc. | The firm is initiating this recall following an investigation which identified the rare possibility that certain units of Comprehensive Reverse Shoulder Humeral Trays may contain a locking ring that is incorrectly assembled. | Employee error | The firm, Biomet Orthopedics, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 22, 2011 to their customers. The letter describes the product, problem, and actions to be taken. The customers were instructed to immediately locate the recalled product; discontinue use; return the product back to Biomet; confirm receipt of the notice by calling 800-348-9500 extension 1368 or 1325, and complete and return the attached form via fax to: 574-372-1683 or mail to: Biomet, 56 East Bell Drive, Warsaw, IN 46580.Customers with questions regarding this recall should call 800-348-9500 extension 1368 or 1325. |  | 1575 | 1,575 units all products | Worldwide distribution: USA (nationwide) including: Puerto Rico and countries including: Australia, Canada, Chile, Columbia, Costa Rica, Korea and Netherlands. | Terminated | February 21, 2013 | 671 | 1575
Z-2552-2011 | 58263 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV (Cardio Vascular) |  |  | March 21, 2011 | 2011 | AGFA Corp. | Customer error with the angiography (XA: X-Ray Angiography) measurement calibration functionality within the Cardiovascular Review Station resulted in improper calibration. | Other | The firm, AGFA Healthcare, sent an "URGENT SAFETY NOTICE" letter dated March 21, 2011 via FED-EX to its customers.. The letter describes product, problem, and actions to be taken. The customers were instructed to distribute this information within their facility to all individuals who need to be aware and to complete and return the URGENT SAFETY NOTICE FEEDBACK FORM via fax to: 864-421-1664 or email. AGFA is supplying the customers with a CRS User Manual Addendum to emphasize the proper steps required to achieve an accurate measurement when using the calibration tool within CRS. Mitigation will provide the customer/end-user with awareness of the correct XA measurement calibration technique and the potential consequences if proper technique is not followed. Additionally, AGFA will offer customer training on the use of the calibration tool. This training will be provided remotely via the web at no cost. Customers were ask to register before May 1, 2011 by contacting AGFA Customer Support Center toll-free at 877-777-2432. Training sessions are from May-June 2011.If you have any questions about this matter, please contact your local AGFA HealthCare organization or the AGFA Customer Support Center at 877-777-2432. |  | 536 | 536 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | November 29, 2011 | 253 | 536
Z-2419-2011 | 58405 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Atlantis Âż SR Pro 2, 40 MHz Coronary Imaging Catheter, 3.6F (1.18mm) x 135 cm, USE BY 2011-05, REF 39014, UPN H749390140, Sterile R Sterilized using irradiation, Made in USA 47900 Bayside Parkway, Fremont, CA 94538 |  |  | March 28, 2011 | 2011 | Boston Scientific Corporation | Catheter tip detachment that is not associated with elongation of the catheter typically caused by excessive pulling forces. EXPANDED: Boston Scientific is initiating a recall of all iCross Coronary Imaging Catheters because they have received and confirmed 8 occurrences of catheter tip detachments due to embrittlement. | Process design | Boston Scientific " Important Customer Notification" letters dated March 14, 2011 were hand delivered to customers by Boston Scientific sales representatives starting on March 28th, 2011. The sales representatives verbally reviewed information. A tracking log will be maintained document completion of the customer notifications.The letter described the product, problem and possible root cause as well as one corrective action, a change in sterilization method from Gamma to E-beam implementation.Boston Scientific sent an Urgent Medical Device Recall Removal - Immediate Action Required letter dated May 26, 2011, to all affected customers. Customers were instructed to Identify and discontinue use of any affected product within their inventory and to segregate it immediately and return it to Boston Scientific |  | 30669 | 30,669 | Worldwide Distribution - USA (nationwide) and Puerto Rico and the countries of Japan, Trinidad, Tobago, and American Virgin Islands. | Terminated | July 30, 2012 | 490 | 96702
Z-2551-2011 | 58736 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Naturalyte@ Sodium Bicarbonate Liquid Concentrate, Dialysate Concentrate for Hemodialysis Liquid |  |  | March 23, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Mislabeled: incorrect lot number printed on case labels | Employee error | Fresenius sent an CUSTOMER NOTIFICATION LETTER letter dated April 18, 2011, to all affected customers via Certified Mail, Return Receipt Requested. Fresenius also contacted affected customers by telephone April 18, 2011, The letter identified the product, the problem and the action needed to be taken by customer. Customers were instructed to examine their stock to determine whether they have the incorrect case label and are asked to remove thecorrectly labeled product from the incorrectly labeled case and discard the case.Acknowledgement forms were Fax-back forms for clinics and mail-back forms (including self-addressed, stamped envelopes), for home patients.Question or concerns should be directed to Fresenius Customer Service Team at 1- 800.323.5188. |  | 3190 | 3,190 cases | Nationwide Distribution--AL, CA, CT, FL, GA, NJ, NV, NY, OR, PA, and WA | Terminated | September 12, 2012 | 539 | 3190
Z-2553-2011 | 58730 | No | HFS | Nipple shield. | General Hospital | General Hospital | 510(K) Exempt | Contact Nipple Shield (Sterile) - 24mm (Standard); Individually wrapped silicone contact nipple shields for immediate use for latch-on difficulties while breastfeeding the baby under professional care only; Manufactured for Medela, Inc., P.O. Box 660, McHenry, IL 60051-0660, Made in Switzerland; article #67203S |  |  | May 18, 2011 | 2011 | Medela Inc | The product sterility may be compromised due to possible pinholes in the rigid tray of the package. | Packaging | The firm, Medela, sent an "Important Device Recall" letter dated May 9, 2011 to all of their customers/direct accounts. Medala also began telephoning or visiting the direct accounts/customers on May 18, 2011. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) not to use the affected lots and place them in quarantine2) stop sale and distribution and set aside their existing inventory of the affected product 3) conduct a physical count and record the data on the Product Inventory Form inlcude with this letter4) complete and return the Product Inventory Form via fax to 800-995-7867 or email jim.hughes@medela.com, even if they have no affected product on hand5)If you have the affected product in stock, Medela Customer Service will ship substitute product to the customer at no cost. Each affected product will be replaced with (2) non-sterile nipple shields.5) and to record the number of units on hand for return and replacement on the enclosed Product Inventory Form. The completed form should be returned via fax or e-mail to Medela even if the account has no inventory of the affected lots. Questions about the recall process were directed to Medela Customer Service at 1-888-835-5968 or e-mail at medelareturns@medela.com. Distributor/retailers were requested to notify their customers of the recall, using the enclosed customer letter. |  | 4320 | 4,320 units | Nationwide distribution: USA including states of: AR, CA, CO, CT, DE, FL, GA, IA, IL, IN, MA, MD, ME,MI, MN, MS, MT, NC, ND, NH, NJ, NY, OH, OK, PA, SC, TX, UT, VA, VT, WA, WI, and WV. | Terminated | February 16, 2012 | 274 | 4320
Z-2388-2011 | 58794 | No | LKB |  |  | General Hospital | Enforcement Discretion | Churchill Medical Systems |  |  | April 19, 2011 | 2011 | Vygon Corporation | Skin-Prep wipes purchased by Churchill (a Vygon company) for inclusion into Dressing Kits and PICC Insertion Kits were recalled by the manufacturer due to potential bacterial contamination. | Nonconforming Material/Component | A recall letter printed on Vygon letterhead dated April 20, 2011 and specific to each customer was sent via overnight courier mail. The letter described the affected product, issue and actions to take. Customers were instructed to check their inventory for the recalled products, cease use and distribution, and quarantine all affected product. Each customer received a custom copy of the letter with the specific product codes and lots that were shipped to them described in the Recall Acknowledgement and Inventory Return Form. Distributors were provided with instructions on how to handle product that was further distributed by them. The firm issued an updated Recall Notice dated May 31, 2011 to two consignees effected by the expanded recall.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by e-mail at customerservice@vygonus.com. |  | 7900 | 7,900 units | United States (FL, MA, MD, and VA). | Terminated | June 18, 2012 | 426 | 15400
Z-2414-2011 | 58661 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Datascope/ Maquet IABPs. System 98/98XT Intra-Aortic Balloon Pump; CS100/CS100i Intra-Aortic Balloon Pump; CS300 Intra-Aortic Balloon Pump; |  |  | March 16, 2011 | 2011 | Maquet Datascope Corp - Cardiac Assist Division | Several models of the Intra-Aortic Balloon Pumps are being recalled due to a potential problem with the fan's assembly which may cause the power supply to detect an overheating event and shut down. | Device Design | Maquet sent an Urgent Product Field Correction letter dated March 16, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to note the WARNING in the IABP Operating Instruction. Customers were advised to refer to the IAB Instructions for Use, Section D (Manually inflating and Deflating a Catheter). Section D of the IAB instruction for Use reiterates that an IAB should not remain inactive for more than 30 minutes and also provides detailed instruction on how to keep an IAB active during IABP failure. Customers were advised that their Service Representatiave would arrange to replace their supply which would contain a new fan assembly. Upon completion of the power supply replacement, customers were asked to sign a service repair order to verify satisfactory completion of the work. For technical questions contact Technical Support Department at 1-800-777-4222 and press 4.For scheduling a service visit by a Field Service Representatiave, contact the firm's US Call Center at 1-800-777-4222 and press 3.Maquet issued a press release on June 6, 2011 advising consumers of the issue and that their recall was considered a class one. |  | 885 | 885 units | Worldwide Distribution | Terminated | October 25, 2012 | 589 | 885
Z-2557-2011 | 58455 | joint, hip, external limb component | ISL | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Accolade TMZF Plus Hip Stem: |  |  | March 01, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Accolade TMZF Plus Hip Stems are being recalled due to grit blast media which was observed in the drive hole. | Employee error | Stryker Orthopaedics sent an Urgent Product Recall letter dated March 1, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to fax back the attached Product Recall Acknowledgement Form within 5 days to 201-831-6069.Customers were also instructed to:Contact the hospitals in their territory that have the affected product to arrange return of the product.Return the affected product via and Inter-Org Transfer to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey 07430. Attach the fluorescent orange PRODUCT REMEDIATION sticker to their return, indicating the Product Remediation # RA 2009-459. Mark the outer box with the words "Product Recall."For any questions call (201) 972-2100. |  | 44 | total of all units is 44 | Nationwide Distribution | Terminated | October 10, 2012 | 589 | 44
Z-2561-2011 | 57760 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation. Monaco Release 2.03.00 and above. |  |  | January 13, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: For VMAT plans, the measured dose of radiation delivered is lower than the software calculated dose by about 5% in Monaco Release 2.03.00 and 2.04.00. | Software design | Computerized Medical Systems, Inc. sent a "USER NOTICE" on January 13, 2011, to all affected users. This notice identified the product, the problem, and the action needed to be taken by the customer. A return postcard is included for the customer to confirm receipt of the user notification. The firm intends for correction in the dose calulation engine to appear in the Monaco Release 3.00 in March 2011. Any further questions please call (403) 830-8023. |  | 94 | 94 | Worldwide Distribution-- USA (nationwide) and countries of Austria, Australia, Belarus, Canada, China, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, India, Israel, Italy, Japan, Mexico, Netherlands, Poland, Singapore, Slovenia, South Korea, Spain, Switzerland, Taiwan, Thailand and UK. | Terminated | April 03, 2014 | 1176 | 94
Z-2562-2011 | 57845 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning Workstation. Focal Release 4.51.00 and above. |  |  | November 09, 2010 | 2010 | Computerized Medical Systems Inc | Focal: If a plan involving Dynamic MLC Delivery is sent from XiO to Focal and called up there in Plan Review, the angle of the gantry is set to zero for all beams. If this plan is then sent from Focal to the R&V; System, the gantry angles all remain at zero which could lead to patient mistreatment. | Software design | Elekta CMS Software sent a USER NOTICE dated November 5, 2010, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. A return postcard was included for the customer to confirm receipt of the User Notice. The firm intends to correct the software so it does not modify anything about the plan sent from XiO.Customers were instructed to contact their local customer support office with any questions. |  | 155 | 155 | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Australia, Belarus, Belgium, Brazil, Canada, Chile, China, Egypt, Estonia, Finland, France, Germany, Hong Kong, India, Italy, Latvia, Libya, Lithuania, Mexico, Morocco, Netherlands, New Zealand, Philippines, Poland, Romania, Singapore, Slovenia, South Africa, Spain, Suriname, Switzerland, Taiwan, Turkey, Turkmenistan, and the United Kingdom. | Terminated | April 02, 2014 | 1240 | 155
Z-2559-2011 | 58226 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.34.00 through 4.50.00. |  |  | June 12, 2009 | 2009 | Computerized Medical Systems Inc | Focal Software: When a patient is defined as prone and the shifts are manually entered and the LOCK SHIFT box is checked the X coordinate (left/right) and the Z coordinate (anterior/posterior) are reversed (left to right, anterior to posterior). This could potentially lead to incorrect isocenter definitions from prone patients and thus a mislocation of dose. | Software design | Computerized Medical System sent a CUSTOMER ADVISORY dated June 12, 2009, to all affected users. This advisory identified the product, the problem, and the action needed to be taken by the customer. There were no patients mistreated as a result of this issueThe firm will send out a User Notice to all sites still runningaffected software to remind them that the issue resolution is available.For further questions, please call (403 ) 830-8023 |  | 462 | 462 | Worldwide Distribution -- USA (nationwide) including Puerto Rico, and the countries of Algeria, Australia, Austria, Azerbaijan, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Croatia, Egypt, Estonia, France, Germany, Greece, Hungary, India, Israel, Italy, Japan, Malaysia, Mexico, Mongolia, Morocco, Netherlands, Nicaragua, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sudan, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, United Kingdom, and Venezuela. | Terminated | April 09, 2014 | 1762 | 462
Z-2560-2011 | 58229 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Monaco Release 1.0.0 through 1.0.2. |  |  | June 30, 2008 | 2008 | Computerized Medical Systems Inc | Monaco: If the user selects Monte Carlo Photon as the Secondary Algorithm (to be used for final dose calculation through segments), and selects SKIP Forward button at any time during stage 2 optimization, the system may only use the pencil beam algorithm for final computation of dose even though Monte Carlo was selected. Other than specific output to the console window, there will be no other in | Software design | Computerized Medical Systems, Inc. sent a "MONACO CUSTOMER ADVISORY" letter dated June 2008 to all affected customers. The letter describes the product, problem, and the action to be taken. The customer was instructed to use caution when using the Skip Forward button in stage 2 according to written instructions provided in the letter.This defect exists in Monaco Release 1.0.0 through 1.0.2 and will be resolved as soon as possible. Note: The issue was resolved in Monaco Release 2.0.1 and above.Questions should be directed to the manufacturer's representative telephone number at 408-830-8023. |  | 47 | 47 | Worldwide Distribution-- USA (nationwide) including the states of IL, MI, MO, OR, SC, TX, VA, and Puerto Rico and the countries of Canada, China, Cyprus, Estonia, France, Germany, Greece, Italy, Japan, Netherlands, Philippines, Poland, Portugal, Slovenia, South Korea, Spain, and United Kingdom. | Terminated | April 04, 2013 | 1739 | 47
Z-2558-2011 | 58272 | No | CGZ | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OPTI LION E-Plus Cassette, Model BP7507, Opti Medical (25 Cassettes with Samples Probes per box). |  |  | March 15, 2011 | 2011 | OPTI Medical Systems, Inc | The values reported by the chloride sensors may report erroneous values below specification on aqueous controls and potentially with patient samples of whole blood and serum/plasma. | Nonconforming Material/Component | OPTIMedical Systems sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 15, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were notified by letter via e-mail and fax. Customers were instructed to discontinue use of the product and to return any unused product to OPTI Medical Systems, Inc. Customers outside the U.S. were instructed to destroy unused product and certify the destruction on the enclosed Recall Response Card. The Response Card was to be completed and returned to OPTI Medical Customer Service via fax. OPTI Medical Systems will replace the cassettes they have with product from another lot.Customers were instructed to contact the Technical Support Department at 1-800-490-6784 for questions relating to this notice. |  | 86 | 86 boxes (25 units (per box) | Worldwide Distribution-- USA (nationwide) including states of CA and VA, and countries of India, Lebanon, Philippines, Sweden, Venezuela and Vietnam. | Terminated | August 08, 2011 | 146 | 86
Z-2556-2011 | 58762 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | 804159, Terumo Coronary Ostia cannula, 14 Fr, 10 (25 cm) long |  |  | May 16, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The presence of adhesive and/or skiving of plastic material observed inside the distal tip of the cannula, which could lead to particulate and/or decrease in the internal diameter. The plastic skiving may detach and enter the patient, embolizing the coronary vasculature. The potential risk due to decrease in the internal tip diameter is inadequate delivery of cardioplegia and/or prolonged procedu | Process control | Terumo Cardiovascular Systems Corporation sent an URGENT MEDICAL DEVICE REMOVAL notice dated May 16, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken. The letter described the problem, the potiential hazard, and the action to be taken by the consignee. Consignees were instructed to discontinue use of the product and return all product to Terumo. The letter provided a list of alternate products to be used instead of the recalled products. Consignees with questions should call 1-800-521-2818 Monday through Friday 8AM to 6PM. |  | 960 | 960 | Worldwide Distribution - USA (nationwide) and the countries of BELGIUM, BRAZIL, DOMINICAN REPUBLIC, UNITED ARAB EMIRATES (UAE), SINGAPORE, MALAYSIA, AND SOUTH AFRICA. | Terminated | March 21, 2012 | 310 | 6400
Z-2563-2011 | 58676 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory Instrument Interface to Data Innovations Instrument Manager software; GE Healthcare IT Solutions, Barrington, IL 60010. |  |  | May 03, 2011 | 2011 | Ge Healthcare It | The GE Centricity Laboratory Instrument Interface to Data Innovations is processing a preliminary result instead of the final result from the Process System Manager when there is a test rerun conducted. | Software design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 3, 2011 to all affected customers. The letter describes the product, problem, and action to be taken by the customer. The letter provided the users with specific safety instructions to follow until the software was updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. |  | 14 | 14 units | Worldwide Distribution-- USA (nationwide) including states of Georgia, New York, Ohio, Oregon, and Tennessee and countries of Canada, England, and Qatar. | Terminated | October 04, 2011 | 154 | 14
Z-2565-2011 | 58747 | No | JPF | Prothrombin-proconvertin test and thrombotest. | Hematology | Hematology | 510(K) Exempt | Factor II (Prothrombin) G20210A Kit Light Cycler Instrument; Manufactured in Germany for Roche Molecular Systems, Inc. Branchburg, NJ 08876 USA DisFor use with Light Cycler 2.0 instrument; |  |  | March 31, 2011 | 2011 | Roche Molecular Systems, Inc. | Roche Molecular Systems has decided to discontinue marketing the Factor II (Prothrombin) G20210A Kit and the Factor V Leiden Kit for use with the Light Cycler 2.0 Instrument in the US Market | Other | Roche sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 31, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. The customer was instructed to discontinue use of the affected product by August 31, 2011.Complete the attached faxback form ( 4293-00-0311 ) and fax it to 1-888-345-5359.For questions or technical support, please contact Roche Molecular at 1-800-526-1247, |  | 13864 | 13,864 units | Nationwide Distribution | Terminated | August 13, 2012 | 501 | 30890
Z-2564-2011 | 58258 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Monaco Release 2.0.1 and above |  |  | May 31, 2010 | 2010 | Computerized Medical Systems Inc | Monaco: due to a software defect, the Radiological Depth (cm) and SSD to DRP (cm) parameters are reported incorrectly on the Beam Summary Report. This occurs for all Monaco IMRT plans. | Software design | Computerized Medical System sent a "USER NOTICE" letter dated June 24, 2010, to all affected users. This notice identified the product, the problem, and the actions to be taken. There is no current workaround. All customers will be mailed a postcard to be returned as proof of receipt and understanding of the notification. Consignees will be contact via telephone or e-mail after June 30, 2010 if the postcard is not returned This notice is to be distributed to any and all users of CMS software at your organization who are potentially affected. The notice states the issue will be resolved in Monaco Release 2.04, but no date for this release was given.For further questions, please call (403) 830-8023 |  | 62 | 62 | Worldwide Distribution -- USA (nationwide) and the countries of Australia, Austria, Belgium, China, Finland, France, Germany, India, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, South Korea, Spain, Switzerland, Taiwan, Thailand and United Kingdom. | Terminated | April 10, 2014 | 1410 | 62
Z-2569-2011 | 57843 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation. Monaco Release 2.04.00 and above. |  |  | October 27, 2010 | 2010 | Computerized Medical Systems Inc | Monaco: In XiO IMRT (Dynamic MLC Delivery) software, the radiation beam angle is set to zero. If different beam angles are selected via Monaco software, the beam angles will remain at zero. | Software design | Computerized Medical Systems, Inc. sent a User Notice to the affected consignees on October 27, 2010, explaining the reason for correction, the clinical impact, and the workaround method. (VERIFY USE NOTICE DATE) A return postcard was included for the customer to confirm receipt of the User Notice. The firm intends for this issue to be corrected in Monaco Release 3.00 in March 2011. |  | 11 | 11 | Worldwide Distribution - USA including CA, IL, MN, NM, OH, NY, TX, and WA and the Netherlands. | Terminated | April 02, 2014 | 1253 | 11
Z-2567-2011 | 58227 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.1 through 4.40.00. |  |  | December 09, 2008 | 2008 | Computerized Medical Systems Inc | XiO Software: Under certain conditions, when the port shape is altered, dose is not recalculated. The port-change flag is not being set correctly when the port shape is regenerated. The dose distribution will not match the blocking displayed and the mis-match might not be obvious. Monitor Unit calculations will also be based on an incorrect dose distribution. This problem has existed since Re | Software design | Computerized Medical Systems, Inc. sent a "XiO CUSTOMER ADVISORY" letter dated December 8, 2008 to all affected customers. The advisory describes the product, problem, and the actions to be taken. The customers were instructed to set the prompt to "Yes" to avoid the problem from occurring according to written instructions provided in the advisory. This problem has existed since Release 4.3.1 and is expected to be resolved with XiO Release 4.50.00. Questions should be directed to the manufacturer's representative at 408-830-8023. |  | 784 | 784 | Worldwide Distribution-- USA (nationwide) including Puerto Rico and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Israel, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Mongolia, Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Tajikistan, Thailand, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 09, 2014 | 1947 | 784
Z-2572-2011 | 58233 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.3.1 through 4.34.02. |  |  | March 31, 2008 | 2008 | Computerized Medical Systems Inc | Focal: In absolute dose mode a beam with a port may be calculated with open normalization, with the result that the displayed dose reflects open normalization while the monitor units and QA plan reflect blocked normalization. This is particularly important when these beams also include compensators. The problem is detectable by looking at the delivered dose to the weight point in the patient as | Software design | Computerized Medical Systems Inc sent a "CUSTOMER ADVISORY" letter dated March 2008 to all affected customers. The letter describes the product, the problem and the actions to be taken. The firm provided a workaround for customers that states "Alternatively the plan can be imported directly into XIO". A copy of the Advisory was shipped with software within the affected range to those requesting a software upgrade. The issue was resolved in Focal Release 4.40.00 introduced in June 2008.The firm will send out a User Notice to all sites still running affected software to remind them that the issue resolution is available.For further questions, please call (403) 830-8023. |  | 307 | 307 | Worldwide Distribution -- USA (nationwide) and the countries of Albania, Argentina, Australia, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Costa Rica, Croatia, Czech Republic, France, Germany, Greece, India, Italy, Japan, Netherlands, Peru, Portugal, Qatar, Romania, Russia, Serbia, South Africa, South Korea, Spain, Thailand, Turkey, United Kingdom, Uruguay, and Venezuela. | Terminated | April 09, 2014 | 2200 | 307
Z-2574-2011 | 58214 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 1.5.0 and above. |  |  | May 14, 2008 | 2008 | Computerized Medical Systems Inc | In XiO, when Telepathy is exited, re-entered, and the temporary plan re-called, the OLD beam weight will be displayed but the dose distribution will reflect the modified beam weight. | Software design | Computerized Medical Systems, Inc. sent an CUSTOMER ADVISORY letter to all affected customers in May 2008. The letter identified the product, the problem and the action needed to be taken by the customer.The issue was resolved in XiO Release 4.50.00 which has been available since June 2009.A User Notice was sent in March 2011, to all the customers reminding them of the availability of the correction. This notice was sent hardcopy via US Mail but there was no return receipt.For further questions, please call ( 408) 380-8023 |  | 936 | 936 | Worldwide Distribution -- USA (nationwide) including Puerto Rico, and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Israel, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 08, 2014 | 2155 | 936
Z-2568-2011 | 58215 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.0 through 4.50.00. |  |  | May 13, 2009 | 2009 | Computerized Medical Systems Inc | In XiO, when the beams are renumbered, it is possible to get into a state where the dose displayed for a beam does not match the beam placement on the patient. This can occur when dose is calculated, beams are renumbered, and the dose is re-summed only. | Software design | Computerized Medical Systems, Inc. sent an CUSTOMER ADVISORY letter dated May 2009, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.Computerized Medical Systems, Inc have created "patches" to resolve the problem. A User Notice was sent in March 2011, to all users still using the affected software notifying them of the updated software available to correct the problem. This letter was sent by US Mail but no return receipt was included.For further questions, please call (408) 380-8023. |  | 1116 | 1116 | Worldwide Distribution -- (USA) nationwide Distribution, including Puerto Rico, and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia, Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 09, 2014 | 1792 | 1116
Z-2571-2011 | 58657 | No | JFP | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Ca++ Membrane Units, Model #942-046, and Model #942-060 used on ABL505/555, EML105, ABL600, ABL700 Series, and ABL800 Analyzers. The membranes are packaged in a plastic box which contains four (4) membranes per box. |  |  | April 07, 2011 | 2011 | Radiometer America Inc | Based on internal testing and user/customer complaint reports received, the firm determined that wrinkles may arise on the outer membrane of the calcium membrane units upon installation. Retention of fluids between the wrinkles can cause a positive bias resulting in calcium measuring errors. This situation poses a risk to the safety of patients, particularly when the measurement results are used | Pending | The firm, Radiometer America, sent a "MEDICAL DEVICE FIELD CORRECTION" letter dated April 14, 2011, to all affected customers. The letter describes the product, problem and the actions to be taken. The letter provides the customers with a series of steps to be taken when replacing the Calcium Membrane (Until further notice the above procedure must be carried out when replacing the Calcium membrane). The letter also provides quality control instructions for verifying calcium performance. Additionally, the customer is required to manually adjust the upper and lower QC limits on the analyzer in relation to the limits printed on the insert sheet in order to detect the bias during quality control measurements (This must be done for every lot of QC ampoules used, and for any new lot of ampoules installed).The letter asks that each customer complete and return the attached form via fax to 440-871-0463, in order to verify that the notification letter was received. Questions should be directed to the firm's Technical Support staff at 800-736-0600, Opt. 4. |  | 30323 | 30323 boxes | Worldwide Distribution--USA (nationwide) including the states of AK, AL, AR, AZ, BM, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WV, and the countries of: Algir, argentina, Austria, Austrailia, Bahrain, Belarus, Belgium, Brazil, Brunei Darussalam, China, Columbia, Cuba, Greece, Denmark, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Georgia, Germany, Great Britain, Greenland, Guatemala, Hungary, India, Iraq, Ireland, Isreal, Italy, Jamaica, Japan, Kasakhstan, Kenya, Lebanon, Lithuania Korea, Malaysia, Malta, mexico, Moldavia, Nepal, New Zealand, Norway, Netherlands, Peru, Philipines, Poland, Portugal, Qatar, Oman, Rumania, Russia, Saudi Arabia, Singapore, Serbia and Kosvo, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Syria, Thailand, The Netherlands, Tunesia, Turkey, UK, Ukraine, United Arabia Emirates, Urguary, Usbekistan, and Venezuela. | Terminated | February 19, 2012 | 318 | 30323
Z-2579-2011 | 58197 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.0 and above |  |  | October 03, 2010 | 2010 | Computerized Medical Systems Inc | XiO: When the beam template is recalled, XiO prematurely initiates a dose calculation, re-snaps the MLC or Port to the structure of the current patient , but then fails to re-calculate dose for the new MLC or Port position. If no subsequent change is made to force a dose recalculation, XiO will display the dose from the original MLC/Port shape dose instead of the dose from the re-snapped MLC/Port | Software design | The firm, Elekta CMS Software, sent an "IMPORTANT SAFETY NOTICE" dated September 14, 2010 to its affected customers. The notice describes the product, problem and actions to be taken. The customers were instructed to distribute this notice to any all users of CMS software at their organization who are potentially affected by this issue, and complete and return the attached postcard to confirm receipt of the Safety Notice. The notice also states that this problem has existed since the XiO Release 4.3.0 and is resolved in the XiO release 4.26.00. Elekta will notify the customers when this solution is available.If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 1208 | 1208 | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, .Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 03, 2014 | 1278 | 1208
Z-2578-2011 | 58200 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation treatment Planning System, XiO Release 4.2.0 and above |  |  | September 17, 2010 | 2010 | Computerized Medical Systems Inc | When a dialog box is displayed and the dose status is changed to "Dose not current", beam edits can be made in the graphics area using the mouse (rotate collimator, reposition collimator, move isocenter, move weight point, etc.). Then if CANCEL is selected on the dialog box, any edits made on the dialog box are undone, the dose is left unchanged, but the beam edits made in the graphics area with t | Software design | The firm, Elekta CMS Software, sent an "IMPORTANT SAFETY NOTICE" dated September 6, 2010 to its affected customers. The notice describes the product, problem and actions to be taken (workaround). The customers were instructed to distribute this notice to any and all users of CMS software at their organization who are potentially affected by this issue. The notice also states that this problem has existed since XiO Release 4.2.0 and will be resolved in XiO Release 4.70.00. Elekta will notify the sites when this solution is available. If you have any questions, please call 314-993-0003 or Toll Free: 800-878-4267. |  | 1293 | 1293 | Worldwide distribution: USA (nationwide) including Puerto Rico, and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 08, 2014 | 1299 | 1293
Z-2577-2011 | 58210 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.0.0 through 4.40.00 |  |  | April 08, 2008 | 2008 | Computerized Medical Systems Inc | XiO: When the patient is positioned with their feet toward the CT gantry, the coordinates exported to the LAP IsoMark laser positioning system are not correct. | Software design | The firm, CMS, Inc., sent a "XiO Customer Advisory" letter dated April 8, 2008 to all affected customers. The letter describes the product, the problem, and the actions to be taken ( workaround). The letter also states the problem will be resolved in a future release. No release number or date was given.If you have any questions, please call 408-380-8023. |  | 75 | 75 | Worldwide distribution: USA (nationwide) and countries of: China, France, Germany, Japan, Mongolia, Netherlands, Nicaragua, Poland, Russia, South Korea, Switzerland, and Turkey. | Terminated | April 08, 2014 | 2191 | 75
Z-2529-2011 | 58910 | Yes | NIP |  |  | Cardiovascular | PMA | Boston Scientific Innova Over-the-Wire Self-Expanding Stent System, Catalog No. 39181-07403. 7 mm x 40 mm x 130 cm (stent diameter x stent length x delivery system length). Made in USA, Two Scimed Place, Maple Grove, MN 55311 USA. The Innova Self-Expanding Stent System is comprised of two components: the implantable endoprosthesis and the stent delivery system. The stent is a laser cut self-expanding stent composed of a nickel titanium alloy (Nitinol). On both the proximal and distal ends of the stent, radiopaque markers made of tantalum increase visibility of the stent to aid in placement. The stent is constrained within a 6F (2.1 mm maximum OD) delivery system. The delivery system is a triaxial design with an outer shaft to stabilize the stent delivery system, a middle shaft to protect and constrain the stent, and an inner shaft to provide a guidewire lumen. The delivery system is compatible with 0.035 in (0.89 mm) guidewires. When ready to be implanted, the stent is deployed by retracting the middle shaft of the delivery system. A radiopaque marker at the distal end of the delivery system aids in visibility during deployment. As the stent is exposed to body temperature, it expands to appose the vessel wall. |  |  | May 19, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific Corporation is conducting a recall of the Innova Self-Expanding Stent System because they have received 6 complaints involving no deployment/partial deployment of the Innova Stent. Potential health hazard events resulting from this type of failure include increased procedure time, vessel wall injury, and/or emergency surgery to remove the partially deployed stent. This recall d | Device Design | Boston Scientific sent an Urgent Medical Device Recall Removal - Immediate Action Required letter, dated May 13, 2011, to all affected customers via registered mail. The letter explained the issue, identified affected product, stated that further distribution and or use of the devices should cease immediately, and asked that devices in inventory to be segregated and immediately returned to Boston Scientific. An Account Reply Verification Tracking Form was also asked to be returned. Boston Scientific Sales Representative were assist in the return process by providing a Returned Goods Authorization (RGA) number.Customers were instructed to contact their local Sales Representative for any questions. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Product removal in US is being managed through SuperNOVA Clinical Trial suspension in accordance with IDE #G100291. OUS distribution to include: Austria, Finland, France, Germany, Italy, Netherlands, Spain, and Switzerland. | Terminated | February 04, 2014 | 992
Z-2582-2011 | 57807 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System. XiO Release 4.50.00 and above |  |  | November 16, 2010 | 2010 | Computerized Medical Systems Inc | XiO: When a new electron beam is added to a plan, the calculation Algorithm is set to Pencil Beam, and an Aperture is added to the beam, the user can select Port-Port Properties and edit the Material Thickness values. If the algorithm is then changed to Monte Carlo, the calculation will proceed and the user is not warned the Material/Thickness values just entered are not used in the dose calculati | Other | Elekta CMS Software drafted a User Notice which explains the reason for correction, the clinical impact, and the workaround method. The User Notice was distributed to all affected sites in November 2010. The issue will be resolved in the XiO Release 4.70.00, expected to be available in July 2011. |  | 556 | 556 | Worldwide Distribution - USA (nationwide, Washington, DC and Puerto Rico and the countries of Albania, Algeria, Australia, Austria, Bahamas, Belarus, Belgium, Brazil, Canada, Chile, China,Croatia, Cyrus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Phillipines,Poland, Portugal, Romoania, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Suriname, Switzerland, Taiwan, Thailand, Turkey, Turmenistan, Ukraine, United Kingdom, and Vietnam | Terminated | April 03, 2014 | 1234 | 556
Z-2580-2011 | 58787 | No | ITH | Power traction equipment. | Physical Medicine | Physical Medicine | 510(k) | MTD 4000 - Mettler Traction Decompression System, Model: ME 4000 |  |  | April 14, 2011 | 2011 | Mettler Electronics Corp | The recall was initiated because Mettler has confirmed the possible failure of internal mechanical component and software detection for potential failure regarding Electronics Traction Device - MTD4000. The firm are initiating the recall because component failure might result in patient injury. Use of the device should cease immediately. | Device Design | Mettler Electronics sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 14, 2011 to all affected consignees. The letter identified the product, the problem and the actions to be taken. Customers were instructed to immediately cease use of the recalled device, examine their inventory and quarantine product subject to recall. In addition, consignees were told if they have further distributed the affected product to please identify their customers and notify them at once of this product recall. The letter asks that that each includes a "Recall Return Response Form" that consignees are required to complete and return. If you have any questions please contact Mettler Electronics Corp 714-533-2221 x331, Monday through Friday, between the hours of 8:30 A.M. and 4:00 P.M. Pacific Time. |  | 221 | 221 | Worldwide Distribution --USA (nationwide) and countries of Taiwan, Ireland and Haiti. | Terminated | July 10, 2012 | 453 | 221
Z-2590-2011 | 58717 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ENDO-FUSE(TM) FUSION ROD, 7 mm X 70 mm IMPLANT MATERIAL: STERILE R, Ti6A14V, CPTi, REF: 500005970, LOT 088605422, Wright Medical Technology, Inc , 5877 Airline Road, Arlington, TN 38002 |  |  | April 20, 2011 | 2011 | Wright Medical Technology Inc | Two lots- one lot of Endo-Fuse Fusion Rod, 7x50mm was commingled one lot of Endo-Fuse(R) Fusion Rod, 7x70mm. Not all parts in each lot are affected. The commingled rod sizes might not be detected prior to surgery. | Process control | The firm, Wright Medical, sent an "URGENT: Medical Device Voluntary Recall" letter dated May 6, 2011 to their customers via FedEx . The letter describes the product, problem, and actions to be taken. The customers were instructed to examine their inventory for the affected products; immediately return any affected product to Wright Technology as soon as possible (contact Customer Service at 800-238-7117 for return instructions and replacement inventory), and immediately complete and return the attached response form via fax to: 901-867-7401.If you have any questions regarding this matter, please contact 800-874-5630. |  | 121 | 121 | Nationwide distribution: USA including states of: AL, AZ, CA, CO, FL, ID, IN, IL, LA, MI, MN, MO, NC, NE, NH, NY, OH, PA, SC, TN, TX, VA, VT, WA, and WI. | Terminated | December 27, 2012 | 617 | 186
Z-2586-2011 | 58784 | No | MZP |  |  | Microbiology | PMA | COBASÂż AmpliPrep / COBASÂż TaqManÂż HCV Test, US-IVD |  |  | April 14, 2011 | 2011 | Roche Molecular Systems, Inc. | The COBAS Ampliprep/COBAS TaqMan HCV Test has been shown to under-quantitate a subset of genotype 4 patient specimens by approximately 1.0-1.5 log 10 in the absence of any sequence mismatches. | Other | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 14, 2011 and faxback form to their customers UPS ground. The letter described the product, problem and actions to be taken. The customers were instructed to review prior test results if there have been indications of discrepancies between test results generated with COBAS AmpliPrep/COBAS TaqMan HCV Test as compared to other relevant clinical and laboratory findings; complete and return the attached Product Advisory Notice: Affiliate Feedback Fax-Form via fax to: 1-866-792-5445, and file this letter for future reference, in particular for use with their 2010 CAP Survey Results to explain any under-quantitation of sample HVL 2-01.Please contact Roche Molecular Diagnostics Technical Support at 1-800-526-1247 if you have any questions about the information contained in this letter. |  | 50993 | 50,993 units | USA (nationwide) and countries of: Australia, Austria, Azerbaijan, Belgium, Brazil, Canada, China, Croatia, Czech Republic, Egypt, Finland, France, Germany, Ghana, Greece, Hong Kong, India, Israel, Italy, Japan, Kuwait, Luxembourg, Malaysia, Mongolia, Netherlands, New Zealand, Norway, Poland, Romania, Saudia Arabia, Sinagpore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates and United Kingdom. | Terminated | August 27, 2012 | 501 | 339392
Z-2584-2011 | 59037 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | Philips SureSigns VS3 Vital Signs Monitor Models: |  |  | June 10, 2011 | 2011 | Philips Healthcare Inc. | Speakers on the SureSigns Patient Monitors and Vital Signs, failure may be identified by the display of a Speaker Malfunc message in the alarm message area or an AUDIO FAILED message in the main screen of the device or by the absence of audible sound | Device Design | Philips Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 2011 to all affected customers. The letter identifies the products, problem, and actions to be taken.The letter states that customers may continue to use the affected products while awaiting replacement of the speaker, provided that the user routinely checks for the display of the "Speaker Malfunct." INOP. Additionally, the letter states that if the customer experience this INOP or there is no sound from the SureSigns Monitor, the customer is to remove the device from use and contact their local Philips service representative.The letter explains that Philips is completing the design for a replacement speaker assembly, which is expected to be available within 8 weeks. When available, Philips will replace the affected speaker assemblies free of charge.Customers are instructed to contact their local Philips Customer Care Center at 1-800-722-9377 if further information or support is required. |  | 5225 | 5225 units | Worldwide Distribution--USA (nationwide) and countries of ARAB JAMAHIRI, AUSTRIA, AUSTRALIA, BAHRAIN, BANGLADESH, BELGIUM, CANADA ,CHILE, COLOMBIA, CROATIA, CYPRUS, FRANCE, GERMANY, HUNGARY, INDIA, IRELAND, ITALY, KUWAIT, LEBANON, LIBYAN, MALAYSIA, MEXICO, NAMIBIA, NETHERLANDS, NEW ZEALAND, OMAN, PHILIPPINES, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA ,SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | August 15, 2013 | 797 | 5740
Z-2591-2011 | 58853 | No | OUB | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MiraDry System consisting of the DTS3000 Console, MiraDry Handpiece |  |  | April 24, 2011 | 2011 | Miramar Labs, Inc. | Reports of skin burns on axilla after treatment with the miraDry System. | Device Design | The firm,Miramar, has a service provider in Japan that assists in dealing with customers. The service providers were provided a copy of the investigation report and communicated directly with the customers to set up appointments for the field correction. The customers were personally visited by a representative from Miramars service provider in Japan. All devices had field correction performed on April 26-28, 2011. |  | 8 | 8 units | International only: Japan | Terminated | June 17, 2011 | 54 | 8
Z-2588-2011 | 58808 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Kinamed SuperCable Grip and Plate System Driver, 3.5mm Hex, Catalog Number: 35-860-2060 |  |  | August 18, 2010 | 2010 | Kinamed Inc | The recall was initiated due to inadequate mechanical strength of these screwdrivers' hexagonal tipfeature, causing it to strip during final tightening. | Nonconforming Material/Component | The firm, Kinamed Inc., sent a recall letter dated August 18, 2010 to consignees/customers via FedEx. The letter described the product, problem and actions to be taken. The customers were instructed to Immediately examine their stock and return any Kinamed SuperCable Screwdriver, 3.5mm Hex bearing these lot numbers, along with the completed attached Effectiveness Check Form and shipment under RMA #74567 to the attention of: "Director of Quality Assurance"Kinamed, Inc.820 Flynn RoadCarnarillo, CA 93012-8701Note: The Kinamed sales agent shipped and had the product replaced for the customers.If you have any questions, please contact (303)-681-2208, (805) 384-2748 or(800) 827-5775 |  | 8 | 8 units | Worldwide distribution: USA (nationwide) state of: CO and country of: Columbia. | Terminated | October 25, 2011 | 433 | 8
Z-2585-2011 | 58837 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra CUSA NXT Ultrasonic Tissue Ablation System Service Module |  |  | May 02, 2011 | 2011 | Integra Limited | The insulated covering of a wire cable that powers the vacuum pump in the CUSA NXT Service Module can become abraded over time resulting in the exposure of bare wire. If this bare wire were to come in contact with the metal surface of the pump, it could result in a short circuit and a loss of aspiration. | Device Design | Integra sent a "FIELD ACTION NOTIFICATION" dated May 2, 2011 to all affected customers. The letter identified the product, the problem and the actions to be taken. Customers were informed of a Field Corrective Action that is being initiated. All devices (customer owned, Integra Loaner Pool, and Sales demos) will undergo rework of the vacuum pump assembly. The rework will be scheduled, at customer sites by Integra Field Service Technicians or returned to a designated Integra Service Center for rework. A loaner will be presented to customer accounts upon request. The letter asks that each customer complete and return the enclosed Recall Acknowledgement and Return Form to Integra, maintaining a copy of the letter and form for your records. For any additional information please call 609-936-2485. |  | 105 | 105 units | Worldwide -- USA (nationwide) and countries of France, Germany, Hungary, Ireland, Italy, Netherlands, Spain, Sweden and Taiwan | Terminated | March 26, 2014 | 1059 | 105
Z-2602-2011 | 58763 | No | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Encompass Hip Stem, Plasma Spray, 12/14 Taper, Standard, REF 108-XXXX, Sterile R, Ortho Development, Draper, Utah 84020. |  |  | May 10, 2011 | 2011 | Ortho Development Corporation | Tests on PVC packaging material resulted in alert levels for cytotoxicity. | Package design/selection | Ortho Development Corporation notified consignees by letter on May 11, 2011. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to return all remaining affected units to Ortho. Replacement product was to be shipped by May 31. |  | 113 | 113 units | Nationwide Distribution - USA including CA, FL, and TX. | Terminated | November 09, 2011 | 183 | 113
Z-2598-2011 | 57801 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation, Monaco Release 2.03.00 and above |  |  | January 13, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: When the isocenter location is edited in the IMRT Activity screen, the shift coordinates are not updated in the IMRT Plan report. | Software design | Computerized Medical Systems Inc drafted a User Notice which explains the reason for correction, the clinical impact, and the workaround method. The User Notice was sent January 13, 2011. A return postcard is included for the customer to confirm receipt of the User Notice. The firm intends for this issue to be corrected in Monaco Release 3.00 in March 2011. |  | 137 | 137 | Worldwide Distribution - USA (Nationwide distribution to AK, AZ, CA, CO, FL, IL, IN, MA, MI, MN, MO, NC, NM, NY, OH, OR, PA, SC, TX, VA, WA, and WI and the countries of Algeria, Australia, Austria, Belarus, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, India, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, Phillipines, Poland, Portugal, Singapore, Slovenia, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, and the United Kindom. | Terminated | April 02, 2014 | 1175 | 137
Z-2599-2011 | 58250 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.1.1 through 4.40.00. |  |  | February 06, 2009 | 2009 | Computerized Medical Systems Inc | Focal Software: if individual MLC leaf position edits are made, the original MLC aperture is retained. A subsequent aperture edit should not affect the MLC leaf positions however it was found that instead the leaves will then snap (unintentionally) back to the original aperture. Also, if the patient is sent to XiO after the MLC is edited, an isocenter shift made in XiO will also cause the MLCs to | Software design | Elekta CMS Software sent an "ADVISORY NOTICE" dated February 6, 2009 to its affected customers. The notice identified the product, the problem, and the actions to be taken. Customers were informed that the workaround method available is to (edit the MLC pattern using the aperture first and the individual MLC leaf edits second). The firm states that this issue will be resolved in Focal version 4.50.00 and above. A "User Notice" will be sent to all users still using the affected software notifying them of the updated software available to correct the problem. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 760 | 760 | Worldwide Distribution -- USA (nationwide) including the state of: Puerto Rico and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Egypt, Estonia, France, Germany, Hungary, India, Israel, Italy, Japan, Malaysia, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Sudan, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, United Kingdom, Uruguay and Venezuela. | Terminated | April 08, 2014 | 1887 | 760
Z-2597-2011 | 58502 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Peloris II Rapid Tissue Processor; Leica Biosystems Melbourne Pty Ltd, Mount Waverly, 3149 VIC, Australia; catalog number 26.0005. |  |  | April 11, 2011 | 2011 | Leica Microsystems, Inc. | The processor potentially has a faulty impellor which could lead to reagent carry-over from one processing step to another, | Nonconforming Material/Component | Leica Microsystems sent a MEDICAL DEVICE CORRECTION letter dated April 11, 2011, to the affected customers via first class mail. The letter identified the product, the problem, and the action needed to be taken by the customer.The Customers were instructed to discontinue use of the device until their local Leica Microsystems representative contact them to arrange the replacement of the faulty impellor. The accounts were asked to complete the attached Acknowledgement Form, listing the serial numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. |  | 22 | 22 units | Worldwide Distribution -- USA (nationwide) and the countries of Australia, Canada, Germany Japan, Portugal and United Kingdom. | Terminated | June 20, 2011 | 70 | 22
Z-2595-2011 | 58260 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.0.2 and above |  |  | May 28, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: when Digitally Reconstructed Radiographs (DRRs) are generated in XiO, and DICOM is used to export the DRRs (File-Print-DRR), the exported images are associated with the Beams, but NOT the Plan. The Plan association had always been done manually by the user. This is functionality that has never been provided on XiO. The problem occurs because XiO is not storing a persistent DICOM | Software design | Elekta / Computerized Medical Systems Inc sent an undated Safety Notice to all affected customers. The notice indicated the problem, the clinical impact, and the workaround. The notice stated the sites will be notified when a solution is available.No date was given for the mailing or the solution. |  | 1292 | 1292 | Worldwide Distribution - USA (Nationwide distribution, including Washington, DC and Puerto Rico) and the countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 10, 2014 | 1413 | 1292
Z-2592-2011 | 58269 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.50.00 and above |  |  | May 10, 2010 | 2010 | Computerized Medical Systems Inc | Focal: for XiO plans calculated in relative dose mode, the Focal Plan Review application will display the DVH and maximum dose correctly. However, the isodose display is not scaled correctly. If the user evaluated the plan based only on the isodose display, they might misinterpret the actual dose delivered to the patient. | Software design | Elekta / Computerized Medical Systems Inc sent an undated Safety Notice to consignees beginning on May 21, 2010. The notice identified the product, the problem, and a workaround. |  | 308 | 308 | Worldwide Distribution - USA (Nationwide including Puerto Rico) Foreign distribution was made to Albania, Australia, Austria, Belgium, Canada, Chile, China, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Lithuania, Malaysia, Malta, Mexico, Netherlands, New Zealand, Portugal, Russia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, and Turkmenistan. | Terminated | April 02, 2014 | 1423 | 308
Z-2604-2011 | 58270 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.50.00 and above |  |  | May 19, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: when a treatment machine is defined in Source File Maintenance (SFM), the MLC Model is set to Elekta, and beams using this treatment machine are added to a new plan in Teletherapy, XiO automatically adds an MLC to the beam. This is because the Elekta linac will always have an MLC in place when the plan is actually delivered. | Software design | A Safety Notice, dated 5/20/10, was sent to Elekta / Computerized Medical Systems Inc clients in May 2010. The notice indicates the problem, the clinical impact. The notice states a future XiO Release will resolve the issue. A return postcard is included with the mailing to verify consignee received the notice. For sites not returning the postcard by 5/28/10, the recalling firm will contact them via phone or email to confirm receipt of the notice. The notice also informed the consignees they would be contacted when a solution was available. |  | 57 | 57 | Nationwide distribution to medical facilities including Puerto Rico. Distribution was also made to two Military and Government consignees. Foreign distribution was made to Australia, Belgium, Brazil, Canada, Colombia, France, India, Ireland, Israel, Italy, Malaysia, Netherlands, Philippines, Poland, Spain, and United Kingdom. | Terminated | April 02, 2014 | 1414 | 57
Z-2573-2011 | 58811 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The Ultraview SL 91517 Capnography module is a mainstream and sidestream gas analyzer designed to measure the concentration of carbon dioxide in a gas mixture, to aid in determining the patient's ventilatory, circulatory, and metabolic status. |  |  | May 12, 2011 | 2011 | Spacelabs Healthcare, Llc | Capnography Module, Model 91517, resets 20 seconds after power failure, instead of after three minutes as specified. This reset clears the RAM and the module returns to default settings. | Software design | Spacelabs Medical sent a "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 12, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken. Customers were recommended to set all their Capnography Module Configuration Manager alarm settings to "ON". The letter also recommended that customers brief their staff about the possibility of losing changes to the default settings and to restore them after a power interruption.A field service engineer will be contacting customers to set up a date and time when affected modules can be upgraded at no cost to the customer. For questions about this corrective action program, call Spacelabs Medical at 1-800-522-7025, Select 2 for Technical Support. For outside the US, customers can call 1-425-657-7200, x5089. |  | 5883 | 5883 units | Worldwide Distribution--USA (nationwide) and countries of ALGERIA, ARGENTINA, AUSTRALIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA,COLOMBIA, CYPRUS, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, FRANCE,GERMANY, GREECE, GUATEMALA, HONG KONG, INDIA, ISRAEL, ITALY,JORDAN, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NEPAL,NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PALESTINE, PANAMA,PERU, POLAND, PORTUGAL, ROMANIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND,SYRIAN ARAB REPUBLIC, TAIWAN, REPUBLIC OF CHINA, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Terminated | December 02, 2011 | 204 | 5883
Z-2594-2011 | 58253 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.40.00 and above. |  |  | January 19, 2009 | 2009 | Computerized Medical Systems Inc | Focal: the couch position can be moved between scans of the same CT study. When the images are imported into Focal, the software is not checking the DICOM Image Position (Patient) tag and therefore is not aligning the images sets if there are different couch positions indicated. A shift between image sets will exist of the same amount that the CT couch was moved. | Software design | Elekta CMS Software sent an "ADVISORY NOTICE" dated January 19, 2009 to its affected customers. The notice identified the product, the problem, and the acton to be taken. Customers were informed that the workaround method available is to (use the image fusion feature already in Focal to relign the scans). The firm states that this issue is resolved in Focal version 4.62.00, which was released in March 2011. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 31 | 31 | Worldwide Distribution -- USA (nationwide) including states of: CA, CO, FL, IL, IN, KS, MO, NC, OH, OK, PA, TN, and WA. and countries of Australia, Austria, Canada, India, Netherlands, Portugal, Singapore, Spain, and United Kingdom. | Terminated | April 10, 2014 | 1907 | 31
Z-2593-2011 | 58262 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 2.1.0 and above |  |  | May 27, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: The Patient Scan position is being confused with the treatment orientation of the patient. The Patient Scan position is determined from the data entered in PFM. Subsequently, when a New Teletherapy Plan is created for a patient in Teletherapy, the user is prompted with Head/Foot Gantry on the text page. This prompt is asking for the orientation of the patient when they are treated. | Software design | Computerized Medical Systems Inc sent an undated Safety Notice to all affected customers. The notice indicated the problem, the clinical impact, and the workaround. The notice stated the sites will be notified when a solution is available. |  | 1400 | 1400 | Worldwide Distribution - USA (Nationwide distribution, including Washington, DC and Puerto Rico) Foreign distribution was made to Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 09, 2014 | 1413 | 1400
Z-2596-2011 | 58256 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.3.0 through 4.34.01 |  |  | November 30, 2007 | 2007 | Computerized Medical Systems Inc | Focal: when a plan is sent to Focal from XiO and the structure properties are edited in Focal CT Sim activity, these changes result in changes to the MLC or field size for the plan when they should not. XiO will recalculate dose when modified plan is returned to XiO however the user may not notice that it has changed form the plant they had approved earlier. Editing the structure properties (col | Software design | Computerized Medical Systems Inc sent a Customer Advisory, dated 11/30/07, was sent to consignees December 2007 to notify them of the issue and how to avoid the workaround. The advisory listed the problem, clinical impact, and a workaround. The firm will be sending each site a reminder that the issues is resolved. |  | 233 | 233 | Worldwide Distribution - USA (Nationwide AL, AR, CA, DE, FL, GA, IL, IN, MA, MD, MI, MOP, NE, NY, OH, PA, RI, TX, VA, WA, WI, and WV and the countries of Albania, Argentina, Belgium, Bosnia and Herzegovina, Brazil, China, Costa Rica, Croatia, Czech Republic, France, Germany, Greece, India, Italy, Japan, Netherlands, Peru, Portugal, Qatar, Romania, Russia, Serbia, South Africa, South Korea, Spain, Turkey, United Kingdom, Uruguay, and Venezuela. | Terminated | April 10, 2014 | 2323 | 233
Z-2608-2011 | 58906 | No | JTZ | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | BBL" OXACILLIN SCREEN AGAR, catalog #221952, box of 10 plates. |  |  | April 20, 2011 | 2011 | Becton Dickinson &amp; Co. | Oxacillin Screen Agar may not have been produced according to specification and could result in false antibiotic susceptibility test results in patient specimens. | Nonconforming Material/Component | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated April 2011 to all affected customers. The letter identifies the product, problem and actions to be taken by the customers.The letter instructed customers to discontinue use of the affected lot number and discard any remaining packages. Customers were requested to complete and return an attached response form via fax at 410-316-4258. The letter states that all discarded product will be replaced.Questions concerning the recall process were directed to BD Customer Service Department at 1-800-675-0908. |  | 5440 | 5440 plates | Worldwide Distribution--USA (nationwide) and the countries of Canada, Columbia, Jamaica, South Korea, and Taiwan. | Terminated | August 15, 2011 | 117 | 5440
Z-2601-2011 | 59000 | No | LDR | Infusion pump. | General Hospital | General Hospital | 510(k) | "***DOSI-FLOW 1 I.V. EXTENSION SET WITH I.V. FLOW REGULATOR REGULADOR DE CAUDAL PARA SETS DE INFUSION I.V.***REF 259-0250***50 UNITS***STERILE OE***LATEX-FREE***LEVENTON LEVENTON, S.A.U. C/Newton, 18-24 08635 Sant Esteve Sesrovires BARCELONA-SPAIN***Distributed by: Victus www.victus.com 4918 S.W. 74th Court Miami, FL 33155***" |  |  | March 11, 2011 | 2011 | Leventon S. A. U. | Leventon, S.A.U. Barcelona, Spain is recalling DOSI-FLOW REGULATOR, Part Number 259-0250, Lot Numbers 101444L and 101494L. Product has potential to leak. | Device Design | The firm, Leventon, S.A.U. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 11, 2011 to the affected customer. The letter identifies the product, problem, and actions to be taken by the customer.The letter instructs the customer to examine their inventory and quarantine any product subject to the recall. They were also instructed to identify their customers who have received the product, notify them of the recall, send a copy of the recall notification letter and request that the customers return any unused product. The customers were also instructed to complete and return the attached RECALL RESPONSE FORM within 5 working days via fax to 00 34 93 6532556 or e-mail to dsalvatierra@leventon.es.Questions concerning this recall should be directed to the QA Manager at 00 34 93 8176316 or e- mail to dsalvatierra@leventon.es. |  | 1000 | 1000 cases (50,000 units) | Nationwide Distribution: FL only. | Terminated | June 28, 2011 | 109 | 1000
Z-2606-2011 | 58833 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems |  |  | May 10, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that Access Ostase Calibrator and QC vials from the affected lots have shown an increased rate of fracturing and may leak the contents in the box during shipment to customers.Operator injury may occur due to contact with broken glass, exposure to infectious agents, and/or exposure by skin contact to agents that may cause sensitizatio | Package design/selection | Beckman Coulter sent an PRODUCT CORRECTIVE ACTION (PCA) letter on May 10, 2011 to all affected customers who purchased Access Ostase Calibrators and Access Ostase QC. The letter identified the customers with an explanation of the problem and the potential impact should the problem occur and the action needed to be taken. Customer were instructed to:(1) Inspect all Access Ostase Calibrator and QC vials and packaging for signs of failure prior to use, including:-Cracking, particularly along the bottom of the vial,-Absence of reagent and/or leaking of contents into the box(2) Human source material used in the preparation of Access Ostase Calibrator and QC material has been tested and found non-reactive for Hepatitis B (HBV), Hepatitis C (HCV), and Human Immunodeficiency Virus (HIV-1 and HIV-2). No known test method can offer complete assurance that infectious agents are absent. Access Ostase Calibrator and QC vials should be handled in accordance with Good Laboratory Practices, and should be handled as if capable of transmitting infectious disease.(3) Skin contact with Access Ostase Calibrator and QC material may cause sensitization. Wear suitable gloves when handling vials. If contact with skin occurs, wash immediately with plenty of soap and water.Customers were told to share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected products listed above to another laboratory, please provide a copy of the letter to them.In addition, customers were requested to complete and return the enclosed Response Form within 10 days so we are assured you have received this important communication.For replacement of defective vials:(1) In the United States, please contact Client Services at 1-800-526-3821, option 1.(2) In Canada, please contact Customer Service at 1-800-463-7828.(3) Outside of the United States and Canada, contac |  | 702 | 702 kits | Nationwide in US and worldwide to: Algeria,Argentina,Australia,Austria,Belgium,Canada,China,Colombia,Czech Republic,France,Georgia,Germany,Greece,Hong Kong,Hungary,India,Italy,Japan,Jordan,Lebanon,Liechtenstein,Mexico,Morocco,New Zealand,Panama,Reunion,Romania,Russian Federation,Saudi Arabia,Serbia,Singapore,Slovakia,South Africa,Spain,Sweden,Switzerland,Taiwan,Turkey,United Kingdom | Terminated | July 11, 2012 | 428 | 891
Z-2570-2011 | 58896 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHECK PERFORMA STRIP NO. 05987288031, 100 BANDELETTES REACTIVES, ROCHE DIAGNOSTICS GMBH, WWW.ACCUCHEK.COM glucose monitoring test strips. |  |  | May 18, 2011 | 2011 | Roche Diagnostics Operations, Inc. | Erroneous results: Internal investigations revealed that the test strips exhibited a performance issue that can cause test strip errors or a bias in test result that exceed the tolerance range according to EN ISO 15197. | Component design/selection | Direct accounts were notified by a Roche Diagnostics press release on 5/18/2011 that included the product's name, lot number, and information to withdraw the remaining inventory of lot 470049 from storage |  | 27019 | 27,019 units | No US distribution, in France Only. | Terminated | October 25, 2011 | 160 | 27019
Z-2607-2011 | 56787 | No | LGW |  |  | Neurology | PMA | Eon Mini Neurostimulation (IPG) System Model 3788 |  |  | May 24, 2011 | 2011 | Advanced Neuromodulation Systems Inc. | Defective battery was identified as cause for reports of inability to communicate or recharge the product. | Process control | St. Jude Medical sent an "IMPORTANT MEDICAL DEVICE RECALL" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem and actions to be taken by the customers.St. Jude Medical recommends no action be taken if the product serial numbers do not match the numbers listed in the letter. For product remaining in unimplanted inventory, customers are instructed to contact their SJM representative so that product may be returned and replacements be provided. For implanted product that matches the serial numbers, the Medical Advisory Board recommend not to unnecessarily explant the devices if they are functioning as intended. Replacement devices will be provided at no additional cost to the customer. A Product Reconciliation Form was attached to the letter for customers to complete and return any unused product to the St. Jude Medical Neuromodulation Division at 6901 Preston Road., Plano, TX 75024. Questions regarding this recall are directed to the customer's St. Jude Medical Neuromodulation Division Representative. |  | 17141 | 17,141 | Worldwide Distribution-- USA (nationwide) including PUERTO RICO and countries of AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, FINLAND,FRANCE, GERMANY, GREECE, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, MEXICO, NETHERLANDS, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA,SPAIN, SWEDEN, SWITZERLAND, TURKEY, and UNITED KINGDOM. | Terminated | July 21, 2015 | 1519 | 17141
Z-2575-2011 | 58995 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Integra Pain Management Disposable, Convenience Kit |  |  | May 19, 2011 | 2011 | Integra LifeSciences Corp. | Kits contain protective wipes that may be contaminated with Bacillus cereus. | Nonconforming Material/Component | Integra LifeSciences consignees were notified by Urgent Medical Device Recall letter on May 20, 2011 and told to quarantine any remaining product and to notify customers to return any remaining product to them. Further information is available at (800) 241-2210. |  | 520 | 520 kits | UT | Terminated | August 01, 2011 | 74 | 560
Z-2610-2011 | 58779 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI TM 101 Monitor, CDI TM 101 Hematocrit/Oxygen Saturation Monitoring System, |  |  | April 28, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo discovered through its own testing that the presence of electromagnetic interference (EMI) in the operating room may affect the accuracy of the valuesdisplayed by the CDI"101 Hematocrit/Oxygen Saturation Monitoring System: hematocrit (HCT), hemoglobin (Hgb) and Oxygen Saturation. | Device Design | Terumo Cardiovascular Systems, Corp issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 28, 2011, to all affected customers. The letter identifed the product, the problem and the action needed to be taken by the customer. After a solution has been verified, Terumo CVS will implement a correction for all affected units. Customers were instructed to review this Medical Device Correction notice, Assure that all users are aware of this notice. Confirm receipt of this and to fax back the Customer Response form.For further question please contact Terumo CVS Customer Service at 1-800-521-2818, Monday - Friday, 8 AM - 6 PM EST. |  | 622 | 622 | Worldwide Distribution -- USA ( nationwide) and the countries of Colombia, Vietnam, India, Indonesia, Pakistan, Canada, Chile, Vietnam, Chile, Hong Kong, United Arab Emirates, Singapore, Australia, Malaysia, Singapore, Taiwan, Belgium, Thailand, Vietnam and India | Terminated | November 06, 2012 | 558 | 622
Z-2625-2011 | 58822 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | 12-lead ECG Cable Assembly contains the following cable and attachments: |  |  | May 11, 2011 | 2011 | Physio Control, Inc. | The 12-lead ECG cables have experienced premature failures which cause the shielding and conductors to break internally creating ECG noise or "lead off" indications. | Device Design | Physio-Control, Inc. sent an "URGENT-MEDICAL DEVICE ACCESSORY INFORMATION" letter dated May 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The customers were advised to follow the instructions:-Use standard hard paddles or therapy electrodes for ECG monitoring if monitoring cannot be performed using the ECG cable. Refer to Troubleshooting Tips for ECG Monitoring in the LIFEPAK 12or LIFEPAK 15 Operating Instructions-Monitoring chapter.-Disconnect the ECG cable from the monitor after use and coil it loosely for storage. Please refer to the Storage Instructions for 12-lead ECG Cable.-Inspect cables daily for bulges which may indicate damaged internal wires. If bulges are visible, replace the ECG cable.-Perform periodic 6 month function testing: Refer to Maintenance and Testing Schedule in the LIFEPAK 12 or LIFEPAK 15 Operating Instructions- Maintaining the Equipment chapters for instructions on how to perform a patient ECG cable function check to confirm proper operation ofthe cable.Customers can visit Physio-Control's website at www.physio-control-notices.com/12LeadECG for additional information or todownload copies of current LIFEPAK 12 defibrillator/monitor or LIFEPAK 15 monitor/defibrillator Operating Instructions and Operator's checklists.Customers can also call Physio-Control's Technical Support Team at 1-800-442-1142, option 5, from 6:00 A.M. to 4:00 P.M. (Pacific), Monday-Friday. |  | 21825 | 21,825 units distributed worldwide | Worldwide Distribution--USA (nationwide) including Puerto Rico and countries of Australia, Brazil, Chile, Germany, Hong Kong, Mexico, The Netherlands, and New Caledonia. | Terminated | August 12, 2011 | 93 | 21825
Z-2612-2011 | 59072 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | 45cm Laparoscopic Instruments Sterilization tray. |  |  | June 08, 2011 | 2011 | Stryker Endoscopy | During testing and review, Stryker Endoscopy discovered that the sterilization parameters for the 132 degree C gravity steam sterilization cycle in the IFU may not consistently sterilize all instruments in the 45cm Laparoscopic Instruments Tray. | Nonconforming Material/Component | Stryker Endoscopy sent a "URGENT: DEVICE CORRECTION" letter dated June 8, 2011 to affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were instructed to discard all old Instructions for Use sheets for their 45cm Laparoscopic Tray and download the new IFU from the URL link:http://stryker.com/45cm tray. Additionally, customers were instructed to complete and return an attached "acknowledgement of receipt" form and fax to 408-754-8378 or scan and e-mail to trays@stryker.com.Questions can be answered by calling 1-800-624-4422 or by e-mail at trays@stryker.com. |  | 603 | 603 | Worldwide Distribution--USA (nationwide) and countries of Brazil, Canada, Chili, Germany , Korea, Mexico, Philippines, Poland, Portugal, Singapore, Sweden, Switzerland, and Taiwan. | Terminated | October 25, 2011 | 139 | 603
Z-2609-2011 | 58767 | Yes | MRM |  |  | Cardiovascular | PMA | Paradym DR Model 8550 and Paradym VR Model 8250, Sterile EO, Manufactured in Italy by Sorin CRM S. r.l., Via Crescentino s.n., 13040 Saluggia(VC) Italy |  |  | May 02, 2011 | 2011 | ela Medical Llc | On February 28, 2011, the company discovered that since the beginning of the US commercial distribution of Paradym DR 8550 and Paradym VR 8250 ICDs in October 2010, the devices included an incorrect version of the Patient Booklet.The company determined that there is no risk of adverse event associated with this labeling error. | Error in labeling | Sorin CRM USA instructed its field representatives on April 27, 2011 via a conference call to attach the correct Patient Booklet (ref. N582,) as an addendum to the device's external packaging, along with a cover letter "Quality Note" dated May 2, 2011 explaining the discrepancy. The field representatives were fedexed the information on May 2, 2011.The letter instructed the Field Representative to attach the correct patient booklet and an addendum, to the outside of the device box.Field Representatives were asked to complete the acknowledgement form by identifying S/N's in which they attached the pouch and their signature of completion and fax the form to 866-500-6096 or e-mail a scanned copy to sorin.ptusa.sorin.com.For questions regarding this recall call 877-663-7674." For affected units that were already implanted as of April 27, 2011, the patient booklet given to the patient may include information written by the physician or hospital staff that is related to the implanted system (ICD and leads), physician references and patient medications. For this reason and given the fact that the company determined that there is no risk for the patient associated with this labeling error, the company decided to not replace the patient booklet for this population. |  | 275 | 275 | Nationwide Distribution including AL, AZ, AR, CA, CO, CT, FL, GA, IA, MA, MI, MO, MS, NJ, NY, NC, OH, OK, PA, SC,A TN,T X, and WA. | Terminated | April 26, 2012 | 360 | 275
Z-2624-2011 | 58877 | No | HPZ | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | Product Brand Name: Coated Bio-Eye Orbital Implant w/Conformer. |  |  | May 20, 2011 | 2011 | Integrated Orbital Implants Inc | The recall was initiated because Integrated Orbital Implants has confirmed certain lots of the Coated Bio-Eye were manufactured without expiration dating visible on the packaging. All unlabeled product is currently beyond its expiration date. | Other | Integrated Orbital Implants Inc initiated a recall communication on 5/20/2011 with Integrated Orbital Implants (IOI) forwarding a Urgent Safety Notification letter to all their customers who had purchased the Coated Bio-Eye implants. The letter asked the customers to look for evidence of a printed expiration date on Coated Bio-eye package. If any Coated BioEye packages were not labeled with an expiration date the letter instructed customers to immediately sequester the products and not to use the products in surgery and contact IOI by telephone at (858) 6779990 or 1 (800) 4246537 to request an RMA number and shipping instructions.Also stated in the letter, IOI will replace Coated BioEye implants without expiration dates free of charge upon receipt of the implants labeled withoutexpiration dates. |  | 4237 | 4237 | Worldwide Distribution - USA (nationwide) and the countries of nationwide and worldwide to: Japan, Greece, Germany, Romania, UAE, Australia, Canada, Philippines, Korea, Taiwan, Chile, Denmark, UK, Italy, Pakistan, Turkey, Philippines, India, South Africa, and Singapore. | Terminated | January 11, 2012 | 236 | 4237
Z-2623-2011 | 58834 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Large Bore 4-Way Stopcock, Volume 0.65 mL, Injection Site, Rotating Male Luer Lock; a sterile, nonpyrogenic fluid pathway; product code 2N5601. Intended use: For the administration of drugs and solutions to a patient's vascular system through a vascular access device. |  |  | May 17, 2011 | 2011 | Baxter Healthcare Corp. | There is the potential for pinholes and tears in the individual packaging of the stopcocks, compromising the sterility of the product. | Package design/selection | Baxter Urgent Product Recall letters dated May 17, 2011 were mailed via first class mail to the direct accounts, to the attention of the Director of Nursing and the Director of Materials Management, on the same date, informing them that potential pinholes and tears in the individual packaging could compromise the sterility of the stopcock. The accounts were requested to locate and remove all affected inventory from their facility and return it to Baxter after calling Baxter Healthcare Center for Service at 1-888-229-0001 to obtain the returned goods authorization The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. The Baxter Field Corrective Action Gatekeepers were also notified of this action via e-mail on 5/17/11, and instructed to notify their customers and ministries of health of the recall. |  | 4800 | 4,800 units | California and internationally to Japan and Australia | Terminated | June 18, 2012 | 398 | 4800
Z-2611-2011 | 59004 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Pathfinder Holter Analyzer is a Programmable Diagnostic Computer. |  |  | May 24, 2011 | 2011 | Spacelabs Healthcare, Llc | A potential hazard associated with the use of Spacelabs Healthcare Pathfinder Holter Analyzer software. The second 24 hours of a 48 hours EVO recording will not be analyzed. | Software design | Spacelabs Healthcare sent a "URGENT-MEDICAL DEVICE CORRECTION" letter dated June 3, 2011 to all customers. The letter describes the product, problem, and the actions to be taken by the customers. Spacelabs Healthcare recommended that until software upgrades can be done, the customers should weigh the benefits versus the risks when deciding whether or not to continue using the product. Spacelabs Healthcare also recommended that customers advise their staff that the second 24 hours of a 48 hour EVO recording would not be analyzed.Field service engineers will be contacting customers to set up a time and date to install the revised version of the Pathfinder software. Spacelabs will upgrade Pathfinder software regardless of whether the consignees are using EVO recorders to ensure this hazard does not happen in the future.Consignees can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support, or outside the US, 44 1992 507740 for any questions about this recall. |  | 58 | 58 units total (5 units in US and 53 units internationally) | Worldwide Distribution--USA (nationwide) including states of CA and NC and countries of Czech Republic, Denmark, Hong Kong, Israel, The Netherlands, and United Kingdom. | Terminated | October 26, 2011 | 155 | 58
Z-2616-2011 | 58838 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Cardiovascular Procedure Kit Part number 70367-07 Mission St. Joe's - Asheville X-Coated Low Prime Pack, Sterile, Terumo Cardiovascular Systems, Ashland, MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. |  |  | May 03, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process. | Other | Terumo Cardiovascular Systems sent an URGENT MEDICAL DEVICE REMOVAL notice dated May 3, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Customers were sent 10 replacement pressure monitoring kits with the removal notice as immediate replacements. The firm will send additional replacements to customers that have more than 10 units of affected product as noted on the Customers' returned reponse form. Customers that received the custom packs were provided stickers to identify the recalled packs. The customers were instructed not to open the packs to remove the affected pressure monitoring kit prior to clinical use, but to replace the pressure domes at the time of use.For more information please contact Terumo Cardiovascular Customer Service at 800- 521-2818. |  | 1515 | 1515 all custom packs | Worldwide Distribution -- USA (nationwide )including the states AR, CA, FL, GA, MO, MT, NC, ND, OK, OR, PA, TN, TX, UT, VA, WA and WY, and the countries of Argentina, Belgium, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Guatamala, Honduras, Hong Kong, Indonesia, Kuwait, Malaysia, Mexico, Pakistan, Panama, Philippines, Singapore, South Korea, Thailand, United Arab Emirates, Uruguay and Vietnam | Terminated | July 18, 2012 | 442 | 21010
Z-2581-2011 | 58854 | No | DXQ | Blood pressure cuff. | Cardiovascular | Cardiovascular | 510(k) | Cuffable Blood Pressure Cuffs; Manufactured for Vital Signs, Inc., a GE Healthcare Company 20 Campus Road, Totowa, NJ 07512 Made in Mexico. |  |  | April 26, 2011 | 2011 | GE Healthcare, LLC | BP cuffs may not properly inflate due to a leak | Device Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 26, 2011 to all affected customers. The letter identified the product, the problem and the action to be taken. Customers were instructed to discontinue use of the product, isolate all affected products, return only unused products and distribute this letter to any other healthcare institutions that are potentially affected by this issue. The letter asks that each customer complete and return the enclosed confirmation form via fax to 800-535-7923, in order to confirm receipt of the notification letter. For any questions, contact your local Vital Signs Account Manager or the International Customer Service at 1-800-932-0760. |  | 2684996 | 2,684,996 units | Worldwide Distribution -- USA (nationwide) including Puerto Rico and countries of: Argentina, Australia, Belgium, Canada, Chile, Ecuador, Finland, Germany, Hong Kong, Israel, Italy, Latvia, Lebanon, Nicaragua, Panama, Peru, Saudi Arabia, Slovenia, UK, Uruguay and Venezuela | Terminated | May 28, 2013 | 763 | 2684996
Z-2633-2011 | 55276 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tubing Assembly #142, V2, V12 Mixer/Wash a component of |  |  | July 23, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because the Wash Concentrate Tube Assembly 142 may fail to deliver the appropriate amount of wash concentrate on some SYNCHRON DxC Systems due to an accumulation of crystals.The impact of this issue would include:(1) Quality Control and patient results can potentially be affected.(2) If system has been calibrating properly and Quality Control is recovering within es | Pending | Beckman Coulter sent out an "Urgent Product Corrective Action" letter dated July 23, 2009 and mailed the week of July 27, 2009 to all affected customers. The letter provides the cusotmers with explanation of the problem identifed and asks them to take the following actions:" If consignees are not able to resolve these issues through routine troubleshooting they were told to contact BCI Customer Technical Support at (800) 854-3633 in the United States and Canada." This issue will be corrected with an upcoming software version and tubing modification." Complete and return the enclosed response form within 10 days so that Beckman Coulter may maintain their records.Consignees were also asked to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation.Questions regarding this Product Corrective Action letter were directed to their Customer Technical Support Center (Hotline) at (800) 854-3633 in the United States and Canada. |  | 2392 | 2392 units | Worldwide Distribution - USA (Nationwide) and Canada | Terminated | July 02, 2012 | 1075 | 2392
Z-2622-2011 | 58228 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Interplant Radiation Treatment Planning System, Interplant Release 3.4.0 |  |  | December 01, 2008 | 2008 | Computerized Medical Systems Inc | Interplant Software: If the trajectory needle with a non-zero retraction is adjusted in any of the planning views (Transverse, 3D) and the program is closed and reopened, the Needle Definitions page in the Dose Plan Summary (Plan View) displays the needle with a 0.00 retraction even when the needle is displayed correctly in the planning view. If the Eliminate Leading Spacers function is applied t | Software design | Computerized Medical Systems, Inc. sent a "CUSTOMER ADVISORY" letter dated December 1, 2008, to all affected customers. The letter identified the product, the problem, and the actions to be taken. Customers were instructed to discontinue use of the product. There is no workaround at this time. This issue was resolved in Interplant Release 3.4.1 software. All possibly affected users have been shipped the Interplant Release 3.4.1 software in May 2010. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 41 | 41 | Worldwide Distribution -- USA (nationwide) including states of: CA, KS, KY, MA, MD, ME, MO, NC, NE, NY, OH, PA, TX and VA and countries of: Australia, Brazil, Chile, China, Germany and Japan. | Terminated | April 04, 2013 | 1585 | 41
Z-2627-2011 | 58873 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BACTEC" FX Instrument-Top, catalog #441385. |  |  | August 10, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic instrumentation may exhibit random operational errors that could delay effective patient treatment. | Software design | BD Diagnostics Systems sent an "URGENT PRODUCT RECALL" letter dated August 2010 to all affected customers. The letter describes the product, problem and action to be taken.The letter described the type of error messages the customers may experience. All affected units will be scheduled for on site software update as well as a disk drive replacement. A BD field service engineer will contact the customers to schedule the on site visit.Customers experiencing issues prior to the scheduled visit or have questions should contact the BD Diagnostics Service Department at 800-638-8663. |  | 907 | 907units | Worldwide Distribution--USA (nationwide) including countries of Australia, Belgium, Canada, China, Colombia, New Zealand, Hong Kong, Japan, Korea, Philippines, Singapore, Taiwan, and Thailand. | Terminated | July 05, 2011 | 329 | 907
Z-2629-2011 | 58793 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Servo-s Ventilator Systems; Article number 66 40 440. |  |  | April 20, 2011 | 2011 | Maquet Inc. | Through customer complaints, Maquet has been informed that SERVO-i ventilators have stopped to ventilate and have generated three different technical error codes. | Other | Maquet Inc. sent an URGENT DEVICE SAFETY ALERT letter dated April 20, 2011, to all affected customers via Federal Express. The letter identified the product, the problem and the action needed to be taken by the customer.The letter informs the user of the issue and provides additional instruction and how the existing manual should be updated. Customers are asked to fill out the attached response form and fax it to 1-973-807-9210. The following warning will be included in all user's manuals delivered with new ventilators:'Warning ! The Servo-i Ventilator System is not intended to be used during radiotherapy, since this may cause system malfunction.Warning! The Servo-s Ventilator System is not intended to be used during radiotherapy, since this may cause system malfunction.'Additionally the warning stated above will be included in the format of a sticker with the yearly preventive maintenance kit for Servo-i and SErvo-s to update exiting User's manuals for the installed base of Servo-i and SErvo-s ventilators. A maintenance report form will also be included in the kit which serves as an acknowledgement from the user that the sticker has been added to an affected ventilator unit's operating manual.For further questions, please call (973) 709-7660. |  | 17881 | 17039 units worldwide; 842 units US | Worldwide Distribution - USA (nationwide) | Terminated | September 17, 2013 | 881 | 75503
Z-2631-2011 | 58859 | [6 empty] | 1500T14 Cardiac Ablation System - Generator |  |  | May 09, 2011 | 2011 | St Jude Medical | The recall was initiated because St. Jude Medical has confirmed its T14 and T12 RF Generatormay malfunction upon encountering an unexpected electric current (e.g. static electricity).The resulting 'malfunction results in a failure of the generator to respond to key pressesproperly, and may cause unexpected energy delivery or inability to discontinue energydelivery with the start / stop but | Other | The firm, St. Jude Medical (SJM), sent a "Product Correction" letter dated May 9, 2011 to it's affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of the T14/T12 generators and remote controls until the units have been properly serviced; complete and return the attached field action form, and utilize the units in accordance with the Operators Manual. In case of system malfunction, the customers were instructed to shut off the RF power delivery by turning off the generator via the rocker switch, dislodging the power cord, or disconnecting of the patient cable or indifferent electrode connection. An SJM sales representative will contact the customers in the near future to facilitate repair or exchange of the impacted unit.If you have questions regarding this notice, contact the Marketing Manager, AF at (905-286-4015), Marketing Director, AF/CRM, at (905-286-4107); Product Manager (Tokyo) 03-6255-5740, and/or your SJM sales representative. |  | 10 | 10 | International distribution only: Canada and Japan. | Terminated | April 03, 2012 | 330 | 168
Z-2634-2011 | 58887 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | ***Processed at: Regeneration TECHNOLOGIES, INC. Alachua, FL, USA 32615***Distributed by: SpinalGraft Technologies, LLC 4340 Swinnea Road STE 39 Memphis, TN 38118***OSTEOFIL DBM Paste, 5cc Tissue ID: 321933 Reorder: 002405 Lot #(Donor ID): 101050353 Expiration: 02/21/2013***STERILE***" |  |  | April 09, 2011 | 2011 | RTI Biologics, Inc. | RTI Biologics issued a recall of Osteofil DBM Paste, 5cc that was distributed on 03/04/2008 to one customer. The product was labeled as 5.0cc but should have been labeled as 0.5cc. All customers were notified of this recall on 04/09/2011. | Labeling mix-ups | RTI Biologics sent a "VOLUNTARY RECALL" letter dated April 9, 2008 to the one distributor, Spinal Graft Technologies. LLC. The letter describes the product, problem, and action to be taken by the distributor.The letter states that some of the product was issued a return authorization and requests that the distributor recover the remaining product from distribution. The recovered product will be issued a return authorization. RTI states the grafts can be returned in a non frozen state.Recall effectiveness will be determined by return of the product. |  | 26 | 26 units | Domestic Distribution--TN only. | Terminated | July 24, 2012 | 472 | 26
Z-2645-2011 | 58206 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ERGO++ Stereotactic Radiation Treatment Planning System, ERGO++ Release 1.6.3 and 1.6.3.1 |  |  | December 11, 2009 | 2009 | Computerized Medical Systems Inc | ERGO Release 1.6.3 is overestimating the MU values. | Software design | Computerized Medical Systems, Inc. sent an "USER NOTIFICATION" dated December 11, 2009, to all affected customers. The notice identified the product, problem and suggested a workaround to minimize the risk. The notice stated that this workaround does NOT substitute the need for regular Quality Assurance. Customers are instructed to review page 39 of user manual for further explanation of warnings. Customers were asked to complete the Notice for Customer section and return to 3D Line Medical System. This issue has been resolved in Release 1.7.0 and above. No follow-up notice will be distributed to customers. For further information please contact: CMS Software (314) 812-4460 (St Louis Office) |  | 45 | 45 | Nationwide distribution to medical facilities in CA, MO, NC, ND, NV, NY, OH, OK, TX, and WA. Foreign distribution was made to Brazil, Cuba, Czech Republic, France, Germany, India, Italy, Japan, Malaysia, Poland, Portugal, Serbia, and Spain. | Terminated | April 01, 2014 | 1572 | 45
Z-2635-2011 | 58813 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BEAMVIEW IT, camera imaging system, Part Number: 5410389 &amp; 5852301; |  |  | November 04, 2010 | 2010 | Siemens Medical Solutions USA, Inc - Radiation Oncology | The firm has become aware of a potential issue that may result in a service engineer receiving an electrical shock should the connector cable housing become unfastened from the chassis. | Other | The firm, SIEMENS, sent a "Customer Information" letter to its customers on November 4, 2010. The letter describes the product, problem and actions to be taken. The customers were instructed to include this letter in their BEAMVIEW TI System Owner Manual chapter "Safety Advisory Letters" where it should remain. Siemens has developed an updated kit to install an additional grounding cable on the camera's central control unit to eliminate the problem. The updated kit was delivered with the Customer Information letter.If you have any questions, contact the Service Manager- Regulatory Affairs at 925-246-8407. |  |  |  | Worldwide distribution: | Terminated | November 14, 2011 | 375
Z-2654-2011 | 57906 | No | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | First Breath Âż Adult Face Tent with Tubing 188cm (5ft.) (Non-latex) |  |  | February 11, 2011 | 2011 | Smiths Medical | A sticky substance on the inside and/ or outside of the Mask | Nonconforming Material/Component | Smiths Medical notified consignees by letter on 2/14/11 titled Urgent Medical Device Correction Notice via U.S. Mail (USA only) This Notice includes an Urgent Medical Device Correction Notice Confirmation Form that provides instructions to consignees to return unused affected product to Smiths Medical and to complete the effectivity tracking form. On 3/17/11 Smiths Medical issued an additional recall letter to extend the recall to 3 additional lots not included in the original recall letter dated 2/14/11. Also, the firm issued a recall letter on 3/17/11 to include the Value Pack (VP1302) General Anesthesia Kit to include 7 kit lots. Contact Recall Coordinator 1-800-348-6064 prompt 4, ext 2273 |  | 600 | 600 units | Worldwide Distribution - USA (nationwide) and the countries of United Arab Emirates, Armenia, Brazil, Canada, France, Hungary, Italy, Jordan, Japan, Libya, Poland, Russian Federation, and Saudi Arabia. | Open, Classified |  |  | 88653
Z-2657-2011 | 58828 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Silicone CTS (Close to Shaft) Cuff Adjustable Neck Flange Tracheostomy Tube (Wire Reinforced), Arcadia Medical Silicone Air Cuff Adjustable Tracheostomy Tube, Sterile, Latex-Free, |  |  | May 08, 2011 | 2011 | Arcadia Medical Corporation | The product does not contain a punched hole above the cuff access port (a lumen that enables suctioning of secretions proximal to the inflated cuff) as intended by the devices design. This eliminates the user's ability to suction secretions through this suction port, necessitating that the user revert to standard of care methods for suctioning secretions from around the cuffed trach tube shaft. | Process control | Arcadia Medical Corporation sent a "MEDICAL DEVICE RECALL" letter on May 9, 2011 and email on May 8, 2011 to all affected costumers. The letter included; description of the product name, afffected lot numbers, description of the problem, instructions to discontinue use of the recalled devices and to return them to Arcadia Medical for replacement . The letter instructed the customers to email Arcadia Medical with the number of devices in their inventory. For additional information contact Arcadia Medical Corporation at 219-779-9431. |  | 118 | 118 | Worldwide Distribution --USA including the state of Ohio, and the countries of Denmark, Italy, Germany, Turkey, Slovakia, Isreal, France, Sweden, and Romania | Terminated | August 30, 2011 | 114 | 118
Z-2661-2011 | 59064 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | 0200086, Triad Alcohol Prep Pads, 100 box |  |  | February 28, 2011 | 2011 | Alere San Diego | The recall of the Triad Group alcohol prep products is due to potential contamination of these products with the bacteria Bacillus cereus, which could lead to life-threatening infections. | Other | Alere sent a CUSTOMER NOTIFICATION letter on February 28, 2011 to all affected custumers.The letter included; description of product, affected lots, statement of potential for contamination which could lead to life-threatening infections. Alere asked its customers to discard the Triad Alcohol Prep Pads provided and complete and FAX the enclosed Verification Form within 10 days to confirm receipt of notice and to indicate the number of discarded/required replacement alcohol prep pads.The letter instructed customers to call Alere Technical Services for additional information at 888-246-7483. |  | 4829 | Estimated 4829 total units | Worldwide Distribution--(USA) Nationwide and countries of Canada, Australia, Japan and Puerto Rico | Terminated | February 29, 2012 | 366 | 4829
Z-2658-2011 | 58875 | No | MKZ | Chlamydia serological reagents. | Microbiology | Microbiology | 510(k) | BD ProbeTec" CT/GC/AC Reagent Pack, catalog #440450, labeled in part ***Becton Dickinson and Company, 7 Loveton Circle, Sparks, MD 21152*** |  |  | September 29, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic reagent packs may be underfilled and contribute to invalid quality control testing or false positive patient results. | Process control | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated September 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers were instructed to discontinue use of the product and discard any remaining product. BD will replace any discarded products. Additionally, customers were requested to complete and return an attached form whether or not they have any inventory remaining, to confirm receipt of the notification. The letter instructed the customers to fax the completed form to: Attention: Regulatory Compliance at 410-316-4258.For assistance regarding replacements, contact BD Customer Service Department at 1-800-675-0908. All other inquiries should contact BD Technical Services Department at 1-800-638-8663. |  | 130 | 130 kits | Worldwide Distribution--USA (nationwide) including the states of AL, CA, CO, CT, FL, GA, KY, LA, ME, MN, MO, MS, MT, NE, OH, PA, RI, TX, and WA and the countries of Australia, Canada, Europe, and Japan. | Terminated | June 24, 2011 | 268 | 130
Z-2663-2011 | 58442 | No | KZD | Pressure infusor for an I.V. bag. | General Hospital | General Hospital | 510(K) Exempt | Infusable and Infusa Scan Single Patient Use Pressure Infusors; Vital Signs Inc. |  |  | March 04, 2011 | 2011 | GE Healthcare, LLC | Infusable and InfusaScan Pressure Infusors may leak and not apply desired pressure to their IV solution bag and subsequently the rate of infusion may be inadequate. | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 25, 2011 with an attached return response form to its customers via Fed Ex. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) DO NOT use infusable or InfusaScan pressure infusors (listed lot numbers).2) Isolate all affected product and contact Vital Signs Customer Service at +1-973-956-5300 to arrange for the return and replacement of the infusor.3) If they have forwarded any of the affected product to any other healthcare institution, forward a copy of this letter to those institutions.4)Return only unused product.In addition, the customers were instructed to ensure that all potential users in their facility were made aware of this safety notification and the recommended actions, and to complete and return the attached confirmation form via fax to +1-973-956-5440.Please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other countries: 262-896-2890. |  | 954888 | 954,888 units | Worldwide distribution: USA (nationwide) states including: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, WV, WY; and countries of: Argentina, Australia, Chile, Colombia, Canada, Denmark, Dominican Republic, Finland, Germany, Greece, Hong Kong/Macao, India, Israel, Italy, Korea, Lebanon, Lativia, Mexico, Malaysia, Nicaragua, Netherlands, New Zealand, Paraguay, Philippines, Portugal Panama, Puerto Rico, Peru, Russia, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Taiwan, Turkey, United Kingdom, Uruguay, Venezuela, and Vietnam. | Terminated | October 31, 2012 | 607 | 954888
Z-2660-2011 | 58864 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Condyle Kit with Hexalobular, REF 114700 DISCOVERY ELBOW SYSTEM, HUMERAL CONDYLE SET - HEXALOBULAR, CO-CR-MO / YI-GAL-4V ALLOY, IF USED FOR REVISION SURGERY, INSTRUMENTS 414926 &amp; 414923 MAY BE REQUIRED LOT 183300, 1 SET, BIOMET ORTHOPEDICS 56 EAST BELL DRIVE, P.O. BOX 587 WARSAW, IN 46581 USA EXPIRY DATE: 2021-03 |  |  | April 29, 2011 | 2011 | Biomet, Inc. | The firm initiated this recall after becoming aware that a Discovery Condyle Kit with Hexalobula contained two male condyles instead of a male and female condyle. | Packaging process control | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated May 25, 2011 to all affected customers. The notice describes the product, problem, and actions to be taken by the customers.The notice instructed customers to locate and remove the product. Customers were asked to complete a response form and fax to 574-372-1683, within 3 business days. Additionally, customers were instructed to ship the package priority mail with a copy of the response form.Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday 8 AM to 5 PM. |  | 10 | 10 KITS | Worldwide Distribution--USA (nationwide) including the states of AR and SD and the country of Japan. | Terminated | August 29, 2011 | 122 | 10
Z-2662-2011 | 58186 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Impac Software Sequencer Verify &amp; Record System, for MOSAIQ |  |  | March 10, 2011 | 2011 | Impac Medical Systems Inc | There's a problem in the database conversion program that is used when upgrading from the Multi-ACCÂżS to MOSAIQ or from one version of MOSAIQ to another. If the gantry start angle is incorrect prior to pressing beam-on, the patient will be mistreated. | Software design (manufacturing process) | The firm, Elekta, sent a "USER NOTICE" to its customers with an attached confirmation of receipt form. The notice describes the product, problem and actions to be taken. The customers were instructed to follow the Interim mitigation instructions noted in the notice -verify the Arc start angle of the Simple ARC with MLC treatment field definition prior to delivery and if the customer identifes a field with an errant gantry start angle, to cease using that plan immediately and consult with the original plan for that field and make the correction to the gantry start angle. In addition, the customers were instructed to complete and return the Confirmation of Receipt via fax to: North America and Asia at 702-992-5002 or Europe and Africa +44 (0)1293 654401, Attn: Elekta OIS Su[pport and/or scan and email to: support@impact.com, europe.support@impact.com or by mail; and distribute this notice to any all users of IMPAC software at their organization who are potentially affected by this issue.For questions or requests for upgrade, email support@impac.com (North America/Asia) or europe.support@impac.com (Europe) or call USA 1-800-488-4672. |  | 840 | 840 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | September 16, 2011 | 190 | 840
Z-2659-2011 | 58770 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Precision Ice One-Patient-Use Control Unit and Wrap System (Model # 4BB02BK). |  |  | April 18, 2011 | 2011 | Pro Trainers' Choice Company | The control panel of the Precision Ice One Patient Use Control Unit and Wrap System has a overheating issue. The control panel printed circuit boards were missing a jumper connection across resistors R28 A, B, C, and D. | Other | Pro Trainers' Choice sent a Mandate-Medical Device Recall on April 18, 2011 and April 20, 2011 to all affected costumers.The letter included a description of the product, affected model and lots. It instructed the costumers to discontinue using the Precision Ice, Model 4BBK02BK until the device is returned and corrected by Pro Trainers' Choice. The device should be returned to PO Box 27, Kingston, WA 98346.The customers were instructed to notify their sub-account customers of this recall if the devices were further distributed.Customers can call Pro Trainers' Choice Company at 360-297-3902 or visit the website www.precisionice.com for any questions about this recall. |  | 336 | 336 units | Nationwide distribution including CA, CO, FL, IN, KS, NY, OR, TX, and WI. | Terminated | March 01, 2012 | 318 | 336
Z-2646-2011 | 58894 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Smith &amp; Nephew Hip Arthroscopy Repair Instrument Tray |  |  | May 19, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | Hip Arthroscopy Repair Instrument Tray shipped with an incorrect Instructions For Use (IFU) | Labeling mix-ups | Smith &amp; Nephew sent an "URGENT FIELD CORRECTION NOTIFICATION" letter dated May 19, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers were instructed to confirm receipt of the corrected IFU document (10600729 rev B.) and communicate to appropriate facility individuals. Once appropriate individuals are informed, customers were instructed to complete and return information requested at the bottom of the letter. The completed information should be returned to Smith&amp; Nephew via fax at 978-749-1654 or by mail at Smith &amp; Nephew Endoscopy, 150 Minuteman Road, Andover, MA 01810.Questions may be answered by contacting the Manager, QA and Compliance at 978-749-1397. |  | 11 | 11 units | Worldwide Distribution--No USA Distribution; countries of Italy, Portugal, United Kingdom, and Venezuela. | Terminated | May 09, 2012 | 356 | 11
Z-2670-2011 | 58716 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | SCREENERS Dip Drug Test, Catalog Number: SCD-7MB300 |  |  | April 26, 2011 | 2011 | Ameditech Inc | The recall was initiated because Ameditech's internal review of regulatory files found that the two affected lots were manufactured without proper regulatory clearance. | Process control | Ameditech, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 18, 2011 to the affected customer. The letter describes the product, problem, and actions to be taken by the customer.The letter requested the customer to:(1) Cease distribution of the affected lot and provide remaining inventory count to Ameditech.(2) Destroy any remaining inventory of these lots and document destruction per your site requirements.(3) Ensure that all users of the device receive a copy of the Urgent Medical Device Recall letter.A recall verification form was attached to the letter to be completed and returned by the customer via fax at 1-858-677-0243.Any questions about the information contained in the recall notice should be directed to Ameditech at amd.complaint@alere.com or call 858-535-1968. |  | 6000 | 6000 test cards (120 kits) | Nationwide Distribution | Terminated | August 11, 2011 | 107 | 6000
Z-2665-2011 | 58884 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | HALO Catheter with Auto 10 Technology, Part Number D-1160-43-S, Catalog Number |  |  | May 23, 2011 | 2011 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster, Inc. has recently become aware that 20 electrodes catheters with Auto ID Technology are not being recognized when connected to the CARTO 3 System. | Pending | Biosense Webster, Inc. sent a Customer Notification Letter, dated May 24, 2011, with an Acknowledgement Form and Instructions for Manually Entering the Catheter Identification in the CARTO 3 System to all US and Outside US (OUS) customers who purchased these products. For customers in the US, two letters and forms will be sent via express mail, addressed to the EP/Cardiology Lab and Risk Management Office, respectively. Instructions for manual entry of the catheter identification in the CARTO 3 System are being included in the communication to customers. These instructions will allow customers to input the identity of the catheter into the CARTO 3 System and be able to utilize the catheter.Customers were instructed to complete and return the attached Acknowledgement Form in accordance with the instructions listed on the form. Customers were instructed to contact their BWI representative for any questions. |  | 431 | 431units (298 in US) | Worldwide Distribution - USA (nationwide) and the countries of Belgium, Japan, and Singapore | Terminated | October 12, 2012 | 508 | 1531
Z-2672-2011 | 58966 | No | FAQ | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | RUSCH Urinary Drainage Bag, 2000 ml |  |  | June 02, 2011 | 2011 | Teleflex Medical | Sterile integrity cannot be guaranteed. | Other | Teleflex Medical sent an " URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 2, 2011. The letter identified the product, the problem and the actions needed to be taken. The customers were instructed to immediately discontinue distribution, quarantine any affected products, communicate this recall to any of their customers using the provided customer letter, and complete and fax the enclosed Recall Acknowledgement Form to Customer Service at 1-866-804-9881. Once the fax is received instructions for product return will be provided to customers. In the event that an alternative approach is needed, customers are instructed to contact Teleflex Medical Customer Service at 1- 866-246-6990. For further questions contact your sales representative or Customer Service at 1-866-246-6990. The recall was expanded and an additional letter was sent out on July 20, 2011 to the affected consignees with the above instructions.*****RECALL EXPANDED***** Product 2. Teleflex sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated July 20, 2011, to all affected customers.The letter identified the product, the problem, and the action needed to be taken by the customer. On July 19, 2011 this recall was expanded to include additional product. Please see above strategy information for details |  | 14488 | 14,488 eaches | Nationwide Distribution: AZ, CA, FL, GA, IL, KS, KY, LA, MA, MS, NC, NJ, NM, NY, OH, OK, PA, TN, TX, VA, WI including Puerto Rico. | Terminated | April 05, 2013 | 673 | 28978
Z-2671-2011 | 58718 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | B. Braun CelsiteÂż Venous Access Ports , Distributed by Aesculpa Inc. |  |  | April 13, 2011 | 2011 | B. Braun Interventional Systems | The affected products contain the incorrect winged SurecanÂż needle accessory component in certain CelsiteÂż Access Port package configurations. The winged SurecanÂż needle is used to deliver medications and fluids into implanted CelsiteÂż Access Ports.Two winged SurecanÂż needle sizes are provided in CelsiteÂż U.S. package configurations: 20G (PN 4058631) or 22G (PN 16006593A). A 20G or 22G needle | Process design | B.Braun Interventional Systems Inc, (BIS) sent a "Urgent Medical Device Recall" letter on April 16, 2011 to all affected customers.The letter described the problem and the product involved in the recall. Informed consignees that their local B.Braun Interventional System Inc Product Specialists will contact them with the specifics for return/replacement of the product upon return of qualified removed product. They also requested that a CelsiteÂż Access Port Patient Implant Form be filled out for each patient implanted with one of the CelsiteÂż Access Ports from the affected lots. This Form to be provided to them by their local BIS Product Specialist.For additional information contact 800 523-9676 x4459. |  | 83 | 83 | NY, ND, OK ,PA | Terminated | July 06, 2012 | 450 | 83
Z-2673-2011 | 58878 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Instrumentation Laboratory GEM Premier 3500 Blood Gas Analyzer |  |  | May 11, 2011 | 2011 | Instrumentation Laboratory Co. | Software upgrade: misuse of the barcode gun by repeatedly clicking can cause the instrument to attach test results to an incorrect patient ID. | Software design | Instrumentation Laboratories sent a "FIELD SAFETY NOTIFICATION" letter dated April 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers.Included with the letter was an upgrade package containing Software V7.2.2 disk along with installation instructions. Also included with the upgrade kit was a Upgrade Tracking Form to be completed and returned via fax at 781-861-4207 or via e-mail at ra-usa@ilww.com.Customers with questions were directed to contact their local representative. |  | 840 | 840 units | Worldwide Distribution-- USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE , INDONESIA, ISRAEL, ITALY, JAPAN , KOREA, REPUBLIC OF CHINA, MALAYSIA, NETHERLANDS , PHILIPPINES, POLAND, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, THAILAND ,TURKEY, and UNITED KINGDOM. | Terminated | March 15, 2013 | 674 | 840
Z-2668-2011 | 58832 | No | MIP |  |  | Gastroenterology/Urology | PMA | AMS ActiconÂż Neosphincter Pump, Product Number 72402287, Sterilization Method Steam, American Medical systems Inc., 10700 Bren Road West, Minnetonka, MN 55343. |  |  | May 09, 2011 | 2011 | American Medical Systems, Inc. | American Medical Systems is recalling all unexpired AMS 800 and AMS Acticon Neosphincter Control Pumps.Based upon a review of our product test procedures, we are unable to confirm that all Control Pumps have met our requirements. Test errors could potentially result in a failure to identify a malfunction in the Control Pump activation mechanism. This malfunction may lead to inadvertent activa | Nonconforming Material/Component | American Medical System sent an "URGENT MEDICAL DEVICE RECALL letter dated May 9, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customerCustomers were instructed to return any existing stock . Please contact American Medical System at 1 (800) 328-3881 to obtain a Return Authorization number. Customer Service will also provide a FedEx shipping number for your use to return affected product.For further questions please contact American Medical Systems at 1- (800)328-3881. |  | 225 | 225 | Worldwide Distribution--USA (nationwide) including Washington D.C. and Puerto Rico and the countries FRANCE, GERMANY, CANADA, ISRAEL, SOUTH AFRICA, LEBANON, AUSTRIA, ARGENTINA, SAUDI ARABIA, SPAIN, BELGIUM NETHERLANDS, PERU, VENEZUELA, SRI LANKA, PORTUGAL, TUKEY, BRAZIL AUSTRALIA, CZECH REPUBLIC, TUNISIA. LUXWMBOURG, QATAR, SOLOMON ISLANDS, BOSNIA &amp; HERZEGOVIA, CHILE, CHINA, COLOMBIA, COSTA RICA, DENMARK, ECUADOR, ESTONIA, FINLAND, GREECE, HONG KONG, HUNGARY, ICELAND, INDIA, IRELAND, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NEW ZEALAND, NORWAY, PANAMA, POLND, ROMANIA, RUSSIAN FEDERATION, SINGAPORE, SLOVAKIA, SLOVENIA, SWEDEN, SWITZERLAND, TAIWAN PROVIDENCE OF CHINA, THAILAND, UNITED ARAB EMIRATES, UNITED KINGDOM and URUGUAY. | Terminated | July 06, 2012 | 424 | 15691
Z-2685-2011 | 56408 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter UniCel DxI 800 Access Immunoassay System with Unicel DXI system software version 4.3 |  |  | July 06, 2010 | 2010 | Beckman Coulter Inc. | Test results may have incorrect ranges. If a default sample type setting on the tests screen is changed while running Unicel DXI system software version 4.3, the new setting is not saved to the system database. The changed sample type setting will reset to the previously saved setting when power is interrupted to the system, or if an operator reboots the system. If the default sample type i | Software design | Beckman Coulter forwarded on 7/6/10 a Product Corrective Action letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the UniCel Dxl 600 and Dxl 800 Access Immunoassay systems. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to:A revised version of UniCel DxI System Software that corrects this issue is currently in development. Until the new software version is available, Beckman Coulter has provided an alternate method for changing the default sample type setting. See the appendix included with this letter for additional information.Refer to the attached Questions and Answers document for additional details and instructions, and to determine if their laboratory is running UniCel DxI Sytem Software version 4.3. Customers may wish to confirm that proper ranges and units of measure are being reported for any Access assays with sample type settings that were changed while running software version 4.3.Customers were asked to share this information with their laboratory staff, and retain the notification as part of their laboratory Quality Ststem documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that consignees received the notification.If they need assistance or have any questions regarding this notification, customers were told to contact Technical Support at 1-800-854-3633 in the United States or Canada. Outside the United States and Canada, contact your local Beckman Coulter Representative. |  | 626 | 626 units | Nationwide in the US and to Canada | Terminated | May 07, 2012 | 671 | 794
Z-2676-2011 | 58807 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Jet-Ventilator Catheter for Adults, 14G, STERILE, NO LATEX, REF 30-02-918-1, VBM Medizintechnik GmbH, Einsteinstrasse 1 7212 Sul z a.N. Germany |  |  | May 09, 2011 | 2011 | VBM Medical Inc. | Product labeled with incorrectly as being appropriate for use in adults when the intended population for use of the product is children and infants. | Error in labeling | VBM Medical Inc sent an "URGENT FIELD SAFETY NOTICE" dated May 9, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.The reference and lot numbers, details of the error, along with the incorrect and correct box label, and instructions to the customers to discard the box with the incorrect label or send the complete product back to VBM Medical Inc. for re-packaging. For any further or detailed questions, please contact VBM Medical Inc. at (317) 776-1800.On 7/12/11 upon receiving the recall classification letter, the firm sent out an additional letter to its direct accounts informing them of the recall's classification and reiterating the instructions listed above. |  | 29 | 29 | Worldwide Distribution -- USA ( nationwide ) including the country Montreal QC | Terminated | March 14, 2012 | 310 | 417
Z-2677-2011 | 59005 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Remel Spectra MRSA, Ref R01821, 10 plates per package, and Ref R01822, 100 plates per package. The firm name on the label is Remel, Lenexa, KS. |  |  | May 11, 2011 | 2011 | Remel Inc | The product failed to prevent growth of methicillin-sensitive Staphylococcus aureus (MSSA) | Process control | Remel, Inc. sent an "IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" letter, dated May 10, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to review their inventory and discard all remaining product lots. Additionally, customers are to complete the Product Inventory Checklist and indicate whether or not they desire replacement of the product or credit. A self-addressed , stamped envelope was enclosed in the letter to use to return the checklist. Questions can be answered by contacting the Technical Services Department at 800-255-6730 or 913-888-0939. |  | 48807 | 187/10-packs and 261/100-packs | Nationwide Distribution--including Puerto Rico. | Terminated | June 18, 2013 | 769 | 48807
Z-2693-2011 | 58975 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | THD155428- 15.5F X 28cm Titan HD(TM) Catheter Kit |  |  | May 19, 2011 | 2011 | Medical Components, Inc dba MedComp | Product was packaged with the incorrect valved peelable introducer. The catheter sized is 15.5F and the introducer packaged in the kit is a 15F. The kit should contain a 16F introducer. | Mixed-up of materials/components | Medcomp sent a "PRODUCT ALERT" letter dated May 19, 2011 to all affected customers. The letter described the product, problem, and actions to be taken. The letter is requesting the return of all affected product. Customers are instructed to contact their customer service representative for an RGA (returned good authorization) number. If you have any questions or concerns contact Medcomp at 215-256-4201 ext. 2225. |  | 90 | 90 | Worldwide Distribution -- USA (nationwide) including states of AZ, CA, CO and RI and countries of Australia, Brazil, and Netherlands | Terminated |  |  | 90
Z-2692-2011 | 58560 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Software versions |  |  | April 13, 2011 | 2011 | Eos Imaging | Print true size, print requests from the EOS acquisition station or the sterEOS image review station may, in certain cases, be affected by a different enlargement factorduring printing. T | Software design | EOS imaging sent a "FIELD SAFETY NOTICE" letter dated April 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to wait for the printing queue to empty before sending a second "print true queue" request with different print settings. Customers should then click the button shown on the instruction sheet to view the print queue.EOS imaging will provide a software update to correct all affected systems at no cost to the customers. An EOS representative will contact customers to schedule a time to upgrade the software.Questions can be answered by calling 33 1 55 25 63 90. |  | 15 | 5 EOS Systems, 10 sterEOS workstations | Worldwide Distribution--USA (nationwide) including the states of CA, IA, MO, VA, and WI, and countries of Canada and Europe. | Terminated | January 18, 2012 | 280 | 15
Z-2687-2011 | 59019 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | The product is the "FASTx Sternal Intraosseous Device". |  |  | November 05, 2010 | 2010 | Pyng Medical Corporation | Device failed to operate; the insertion of the infusion tube of the FASTx Sternal Intraosseous Device was not achieved in live subjects. | Other | Pyng Medical Corp. sent a "FSN/RECALL FOR FASTx STERNAL INTRAOSSEOUS DEVICE, P/N FASTx -H" letter dated November 5, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers are to cease distribution and use of all devices. All devices remaining in inventory should be quarantined. Customers are asked to complete the information at the bottom of the letter and return the product by FEDEx. Questions regarding this notice should contact the Director of Operations at 1-604-303-7964 ext. 211. |  | 402 | 402 devices | Worldwide Distribution--USA (nationwide) and the countries of Australia, Austria, Canada, England, Italy, Mexico, Poland, and Sweden. | Terminated | August 10, 2011 | 278 | 402
Z-2688-2011 | 58444 | No | KNF | Unipolar endoscopic coagulator-cutter and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | LiNA Loop Endoscopic Monopolar Loop for LSH (Laparoscopic Supracervical Hysterectomy). Large loop, 4mm connector, 1 X REF: EL-200-4 and 10 X REF: EL-200-4, Manufactured by LiNA Medical ApS, Surgery Division. |  |  | December 01, 2009 | 2009 | LINA Medical ApS | Potential for the LiNA Loop to break during operating procedure. | Labeling design | Lina Medical Aps notified customers by newsletter in December 2009 followed by a letter to distributors including warnings added to instructions for use in January 2010. |  | 200 | 200 pieces | GA, MN, PA, WA, TN, MA, IN, Ontario Canada | Terminated | February 23, 2012 | 814 | 2720
Z-2678-2011 | 58532 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Systems, CDI blood parameter monitoring system 500 with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe, Catalog AVHCT. |  |  | April 21, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Inaccurate readings may result after Methylene Blue, an intravascular dye, is administered to the patient. The presence of Methylene Blue may interfere with the readings from the shunt sensor for the CDI 500 system in a way that would cause inaccurate pH values. Because the pH value is used to calculate other values, these parameters could also be affected by the presence of Methylene Blue: K | Device Design | Terumo sent URGENT MEDICAL DEVICE CORRECTION letters to Chief of Perfusion, Department of Cardiothoracic Surgery or Director of Operating Room Services, dated 4/21/2011. The letter explained the reason for the correction and the potential hazard. The letter stated that Terumo was updating the Operators Manual for all CDI systems and would provide the manuals to all users when available. Consignees with questions should call 1-800-521-2818. Monday- Friday 8 am- 6 pm. |  | 3767 | 3767 units (all CDI 500 products) | Worldwide distribution, including USA, Australia, Barbados, BELGIUM, BRAZIL, CANADA, Chile, Colombia, Costa Rica, Dominican Republic, ECUADOR, El Salvador, France, Germany, Guatemala, Holland, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Korea, Libya, MALAYSIA, MEXICO, PAKISTAN, PANAMA, Paraguay, Saudi Arabia, SINGAPORE, South Africa, Spain, TAIWAN, THAILAND, The Phillipines, TRINIDAD, UNITED ARAB EMIRATES, Uruguay, Venezuela and Vietnam. | Terminated | August 06, 2012 | 473 | 5582
Z-2691-2011 | 59039 | No | MND | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Disposable Hemostasis Clip, DCH-7-230, Rx only, Single Use only, Cook Medical |  |  | June 07, 2011 | 2011 | Cook Endoscopy | The disposable clip may not deploy after being positioned inside the patient. | Device Design | Wilson-Cook Medical, Inc. (Cook Endoscopy) sent an "URGENT: PRODUCT RECALL" notice dated June 7, 2011 and a "PRODUCT RECALL EFFECTIVENESS SURVEY' dated June 21, 2011 to all affected consignees. The notice described the product, problem, and actions to be taken. Medical users were instructed to quarantine any affected products and return to Cook Medical. The distributors were instructed to contact their customers and coordinate the return of the affected products to Cook Medical. The recalled notice and survey should be completed and returned to Cook Medical via fax to 812-339-7316 or scanned and emailed to CustomerRelationsNA@cookmedical.com. If you have any questions regarding this matter, please contact the Customer Relations Department at toll free-1-800-457-4500, press 4 then enter extension 2146 or by email at CustomerRelationsNA@cookmedical.com . |  | 307 | 307 units | (USA) Nationwide Distribution including the states of: CA, HI, MA, ME, NC and NY. | Terminated | May 15, 2012 | 343 | 307
Z-2674-2011 | 58909 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Spectra Optia Apheresis System, REF: 61000, CaridianBCT, Lakewood, CO 80215. |  |  | December 09, 2010 | 2010 | Caridian BCT, Incorporated | An additional air detection system is being added to all existing systems to detect air in the system's return line. | Software design | CaridianBCT sent a "MEDICAL DEVICE RECALL" letter dated December 2010 to all customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were advised to continue using their systems as long as they followed the instructions in the Operator's Manual. An Acknowledgement of Receipt form was included with the letter to be completed and faxed to 1-303-876-9277 or e-mailed to Regulatory.Affairs@caridianbct.com. A CaridianBCT representative will schedule a visit to facilities to upgrade their systems within the next 6 months.Further information is available by contacting the CaridianBCT Support Center at 877-339-4228 or 303-231-4357. |  | 594 | 594 units | Worldwide Distribution--USA (nationwide) and countries of Australia, Belgium, Canada, China, Hong Kong, Japan, Malaysia, Mexico, Pakistan, Saudi Arabia, Singapore, South Korea, Taiwan, Thailand, and Vietnam. | Terminated | September 12, 2011 | 277 | 594
Z-2708-2011 | 58913 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Solution Set for Epidural Use, 110" (2.8 m), Male Luer Lock Adapter; a sterile, single use, nonpyrogenic fluid pathway for use with 6200 and 6300 series Flo-Gard Volumetric Infusion Pumps; Baxter Healthcare corporation, Deerfield, IL; product code 2C7554s |  |  | May 31, 2011 | 2011 | Baxter Healthcare Corp. | The lot of Baxter Solution Set for Epidural Use is being recalled because the position of the slide clamp and roller clamp may have been reversed during assembly. The reversed orientation of the slide clamp and roller clamp will cause the set to be incorrectly loaded into an infusion pump, resulting in fluid flow from the patient to the pump, rather than from the pump to the patient. | Employee error | Baxter Healthcare sent Urgent Product Recall letters dated May 31, 2011, via UPS overnight delivery to all affected customers. The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and remove all affected inventory from their facility and return it to Baxter after calling Baxter Healthcare Center for Service at 1-888-229-0001 to obtain the returned goods authorization The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. The Baxter Field Corrective Action Gatekeepers were also notified of this action via e-mail on 5/31/11, and instructed to notify their customers and ministries of health of the recall. *** The recall was expanded to include two additional lots of product code 2C7554s on July 11, 2011 via letter of the same date. The recall strategy remains the same as that of May 31, 2011. *** |  | 13920 | 13,920 units | Worldwide Distribution - USA including Alabama, Arizona, California, Illinois, Kentucky, Louisiana, New Mexico, North Carolina, Ohio, Oregon, Tennessee and Washington, and internationally to Canada and South Africa. | Terminated | June 18, 2012 | 384 | 13920
Z-2703-2011 | 58936 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | "***LINVATEC SRS SHOULDER RESTORATION SYSTEM CrossFT 5.5mm CrossFT Suture Anchor with Two#2 (5 metric) Hi-Fi Sutures***SRS.Linvatec.com***ConMed Linvatec Biomaterials, Ltd. Hemiankatu 6-8L 33720 Tampere, FINLAND***STERILE EO***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA Made in USA***REF CFP-5502***" |  |  | January 26, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Largo, FL is recalling CrossFT Suture Anchor with Two #2 HI-Fi sutures, Lot 211581, Product Number CFP-5502. There is a possibility that the product may have been incorrectly assembled. The disposable driver may have the incorrect color handle (red versus black). | Employee error | ConMed Linvatec sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated January 27, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers were instructed to check their inventory for affected lots and return for credit. A Reply Form and FedEx Air bill was enclosed with the letter for customers to complete and return with the product.Questions regarding the recall can be answered by contacting the Customer Service Department via e-mail at Custserv1@linvatec.com. |  | 45 | 45 each | Worldwide Distribution--USA (nationwide) and countries of Canada and South Africa. | Terminated | July 18, 2011 | 173 | 45
Z-2698-2011 | 58702 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Oxford Femoral Knee System REF 154600, LOT 2272736, WARNING DO NOT RESTERILIZE, STERILE UNLESS PACK DAMAGED, BIOMET UK Ltd, (510K# K011138) |  |  | April 22, 2011 | 2011 | Biomet U.K., Ltd. | This recall was initiated due to a customer complaint which reported that while peeling off the Tyvek lid on the blister pack, very little effort was required to remove it and this could lead to contamination of the device. | Packaging | Biomet sent a "FIELD SAFETY NOTICE" dated April 21, 2011, to all affected accounts. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product, locate and remove from inventory and quarantine pending return to their Bioment distributor. Customers were also instructed to complete and return the attached "Fax Back Response Form" to their local Biomet distributor.For questions customers should contact their local Biomet Contact. |  | 12 | 12 devices | Outside the US to Austria, Belgium, and Germany | Terminated | August 03, 2011 | 103 | 12
Z-2701-2011 | 58145 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | CODMAN ENTERPRISEÂż Vascular Reconstruction Device and Delivery System 4.5mm x 28mm |  |  | March 21, 2011 | 2011 | Codman &amp; Shurtleff, Inc. | Mislabeled: One of the stick on stickers on the sterile pouch inside of the product box indicates the stent size to be 22mm in length, when in fact the product is 28mm in length. | Labeling design | Codman &amp; Shurtleff, Inc.sent an "Urgent Medical Device Correction Notice" on March 21, 2011.The letter recommend the incorrect label be removed and discarded. There are 3 additional peel-off labels, 2 on the outer box and one on the pouch. All 3 of these contain the correct printed size information (28 mm) that can be used for applying to the inventory sheet, patient record or patient card, as needed. It is NOT required to return any of this product to Codman.Customer Service at 1-800-225-0460 or contact your local Codman Neurovascular Representative |  | 380 | 380 units | Worldwide - Nationwide (USA), Canada, Korea. | Terminated | February 17, 2016 | 1794 | 380
Z-2705-2011 | 58376 | No | HAO | Nonpowered neurosurgical instrument. | Neurology | Neurology | 510(K) Exempt | Sequoia final driver closure top retention, non-sterile, REF 3384-2, Zimmer Spine, Austin, Texas, The device is used to apply final torque to closure tops of the implant. |  |  | April 04, 2011 | 2011 | Zimmer Inc. | A rounded or partially rounded tip on the screwdriver may not provide the required locking torque to lock the closure top to the polyaxial screw. When the hex tip of the screwdriver is not fully engaged and/or the screwdriver is used off axis, the hex tip of the screwdriver can become rounded or partially rounded. | Device Design | The firm, ZIMMER, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 4, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that each instrument tray had both final drivers (with and without retention feature); contact their customer service representative if they do not have both versions; ensure proper seating of the drive by moving the driver side to side, and complete and return the Distributor Certification Form via fax to: Zimmer Spine at 512-258-0995. The letter states the firm will be distributing the version of the screwdriver with no retention feature beginning in March 2011 and distributing the new surgical technique in April 2011. The firm is including an additional final driver in the instrument tray for final tightening of the closure tops. Please keep Zimmer Spine informed of any adverse events associated with the device or any other Zimmer spine product by calling 800-326-0635. For any further questions or concerns contact the Vice President Quality Assurance &amp; Regulatory Affairs at 800-655-2614. |  | 505 | 505 devices | Worldwide distribution: USA (nationwide) including the states of: AL, AZ, CA, CT, CO, FL, GA, LA, MI, MN, MO, NC, NY, OH, OK, PA, TN, TX, WA and WI; and countries of: Czech Republic, Denmark, France, Singapore, Spain, Switzerland, and United Kingdom. | Terminated | August 28, 2013 | 877 | 505
Z-2706-2011 | 59028 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Brand Name |  |  | May 20, 2011 | 2011 | Staar Surgical Co. | The recall was initiated by Staar Surgical due to an error in the Directions for Use (DFU) for Staar Surgical's Collamer Ultraviolet Absorbing Posterior Chamber Single Piece Foldable Intraocular Lens (Collamer IP IOL), alsoknown as the nanoFLEX IOL. The Device Description erroneously indicates that the lens could beimplanted in the ciliary sulcus. | Other | Staar Surgical send out an URGENT - Medical Device Correction Notice to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to acknowledge receipt of the letter by signing, dating, and faxing the form to Staar Surgical. For any questions call (800) 292-7902, ext 2345. |  | 16750 | 16,750 devices distributed to the field | Worldwide Distribution - USA (nationwide) and the countries of South Africa, Slovakia, Mexico, Hong Kong, and Sri Lanka | Terminated | March 04, 2014 | 1019 | 16750
Z-2707-2011 | 58576 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | MaxPlus bifuse extension set (minibore) needleless connectors, |  |  | April 20, 2011 | 2011 | Carefusion 303, Inc. | Carefusion has initiated a product recall for the MaxPlus bifuse extension set (minibore) needleless connectors Model MP9232-C as a result of a potential for occlusion. The product may be occluded at the bifused component. | Device Design | CareFusion initiated on May 3, 2011, two distinctive "URGENT MEDICAL DEVICE RECALL NOTIFICATION" via mail delivery, to all affected customers and distributors. The notices identified the product, the problem, and the actions to be taken. Customers were instructed to 1.) Immediately complete and return the enclosed pre-addressed and postage paid Recall Response Card to CareFusion. Note on the card your distributor, the MaxPlus bifuse extension set lot numbers and quantities that will be returned.2.) All recalled product should be returned directly to the distributor from whom it was purchased. Customers will receive their return goods number through their distributor or CareFusion. 3.) Also, enclosed is a Frequently Asked Question (FAQ) Letter to address any questions or concerns related to the use of the MaxPlus" bifuse extension sets.The second notification was specifically addressed to distributors. The distributors were provided specific instructions for product returned. For customers who may have questions, please contact your distributor directly, or call the CareFusion Support Center at 1-800- 562-6018. The Support Center will be in operation from 7 a.m. - 4 p.m. PST. For distributors who may have any questions or require additional materials, please contact the CareFusion Support Center at 1800- 562-6018. Additionally, instructions were given to report adverse reactions experienced via the FDA's Med Watch Program. |  | 26150 | 26,150 units | Nationwide Distribution -- USA including states of: CA, IA, IN, and TX | Terminated | November 28, 2011 | 222 | 26150
Z-2695-2011 | 58592 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | RUSCH CrystalClear Tracheostomy Tube, Cuffed, Sterile Pack |  |  | April 04, 2011 | 2011 | Teleflex Medical | Recalled due to air leakage from the main ventilation lumen through air passageway, resulting in potential injury requiring medical intervention. | Other | Teleflex Medical sent an URGENT FIELD SAFETY NOTICE dated January 4, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their stock for product, cease use and distribution of stock from the affected lot number and quarantine all products from the affected lot number immediately. Customers were instructed to complete the field safety corrective action acknowledgement and stock status form and fax to their local Teleflex Medical Customer Service Sales Rep.If the customers did have stock from the affected lot number, they were to contact their local Teleflex Medical Customer Service/Sales Rep or distributor as indicated on the attached field safety corrective action acknowledgement and stock status form for a Return Autorization Number.Distributors were to communicate the field safety corrective action notice to their customers and forward the completed Acknoweldgement Form to orders.intl@teleflexmedical.com or fax to +353 1 437 0773.For questions regarding this recall call +353 906 460 838 |  | 328727 | Total 328,727 for all units | Worldwide Distribution - Arabia, Asia, Austria, Belgium, Czech Republic, Denmark, Estonia, Germany, Greece, Ireland, Italy, Spain, Saudi Kuwait, Portugal, Russia Federation, Sweden, Slovenia and South America. | Terminated | March 02, 2012 | 333 | 328729
Z-2699-2011 | 58810 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | LIFECATH S PICC |  |  | April 26, 2011 | 2011 | Vygon Corporation | Peripherally inserted central catheter has the potential to leak or break at the hub while in use, which could lead to a loss of blood or catheter migration. | Other | Vygon Corporation sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 26, 2011. The letter identified the affected products, problem, and actions to be taken. The customers were instructed to check inventory and quarantine all affected products; immediately cease use and distribution of all affected products; return any affected products using the instruction from Customer Service involving product pickup (contact Customer Service at 1-800-473-5414 for a Return Merchandise Authorization number); and immediately complete and return the attached form via fax to 215-390-2019. Distributors were provided with instructions on how to handle products that were further distributed.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by e-,ail at customerservice@yygonus.com. |  | 570 | 570 | Nationwide Distribution--USA including states of: AZ, CA, FL, MO, NE, NJ, NY, PA, TX, UT, VA, and WA. | Terminated | June 01, 2012 | 402 | 1170
Z-2702-2011 | 58836 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Radiation Treatment Planning System, model H48; |  |  | May 03, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Analytical Anisotropy Algorithm versions when modeling the second source, the collimator jaw positions are read from the first control point of the field. This can cause the delivered dose to be higher than planned locally by up to 10%. | Software design | Varian Medical Systems, Inc. Oncology Systems sent a "Urgent Medical Device Correction" letters on May 3, 2011***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter contained recommended actions. For additional information contact Varian Medical Systems, Inc. at 888- 827-4265. |  | 10649 | 5201 units ***12-12-12 AMENDED TO: 5448 units*** | Worldwide distribution; Nationwide (USA) and the countries of Vietnam, Venezuela, United Kingdom, United Arab Emirates, Ukraine, Turkey, Tunisia, Trinidad and Tobago, Thailand, Taiwan, Syria, Switzerland, Sweden, Sri Lanka, Spain, South Korea, South Africa, Slovenia, Singapore, Saudi Arabia, Russia, Romania, Qatar, Portugal. Poland, Philippines, Panama, Pakistan, Norway, New Zealand, Netherlands, Morocco, Mexico, Malaysia, Macau, Lithuania, Libya, Latvia, Kuwait, Kazakjstan, Japan, Italy, Israel, Ireland, Iraq, Indonesia, India, Hungary, Hong Kong,, Guam, Greece, Germany, France, Finland, Egypt, Ecuador, Denmark, Czech Republic, Colombia, China, Canada, Brazil, Belgium, Belarus, Bangladesh, Bahrain, Austria, Australia, Argentenia, Algeria and Canada. | Terminated | July 09, 2014 | 1163 | 10649
Z-2704-2011 | 58857 | No | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Broselow/Hinkle Pediatric Emergency System, Vital Signs Inc., 20 Campus Road, Totowa, NJ 07512 USA |  |  | April 26, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | Broselow Pediatric Emergency System Kit may contain oxygen masks that may contain a potentially sticky substance on the inside and/or outside of the mask. | Other | GE issued an URGENT MEDICAL DEVICE CORRECTION letter dated April 26, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use the affected product, isolate all affected product, complete the attached confirmation form and fax back per the instructions and return any unused product.For any product forwarded to any other healthcare institution, customers were instructed to forward a copy of the URGENT MEDICAL DEVICE CORRECTION letter to those institutions. For any questions contact Vital Signs Customer Service at 1-800-932-0760. |  | 33002 | 33,002 units | IL | Terminated | January 14, 2013 | 629 | 33002
Z-2694-2011 | 58912 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Tobii C-Series Lithium batteries Type TB8A, TB8B, TB12A used with the Tobii C8 and Tobii C12 Speech Generating Communication Devices. |  |  | May 31, 2011 | 2011 | Tobii Assistive Technology, Inc. | Lithium Batteries may overheat and become combustible | Environmental control | Tobii Technology AB sent an "URGENT: MEDICAL DEVICE RECALL" letter dated May 25, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were advised to stop using the batteries and to remove them from their devices. They were instructed to dispose of the batteries in an environmentally responsible manner as hazardous materials. Tobii will replace the affected batteries. Customers were asked to complete and return an acknowledgement form via fax at 781-461-8213 or e-mail at recall@tobiiati.com.Contact the firm by E-mail: recall@tobiiati.com or telephone: 1-800-793-9227. |  | 2430 | 2430 units | USA (nationwide) | Terminated | October 01, 2012 | 489 | 2430
Z-2711-2011 | 58809 | No | ETA | Total ossicular replacement prosthesis. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Kurz, Heinz Kurz Gmb, Medizintecnik, Tuebinger Strasse 3, 7214 Dusslingen, Germany, Surgical Implant, Clip - Piston MVP Titanium, REF: 1006 711, Lot 2 09 13 06. |  |  | April 27, 2011 | 2011 | Heinz Kurz GmbH Medizintechnik | Outer package labels product .75 mm longer than actual size. | Labeling False and Misleading | Kurz Medical, Inc. sent an "URGENT VOLUNTARY MEDICAL DEVICE RECALL" letter dated May 9, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to return unused product for replacement by using an enclosed FedEx label. If the product has been implanted , the letter requests the customers send the date of surgery and size of the implant. |  | 40 | 32, only 8 consigned with U. S. hospitals | Nationwide Distribution--USA (nationwide) including the states of MN, NC, and WA. | Terminated | November 23, 2011 | 210 | 40
Z-2709-2011 | 59043 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Product Brand Name: Versafitcup Double Mobility Liner Inserter |  |  | May 23, 2011 | 2011 | Medacta Usa Inc | Medacta International has recently realized that the base of the Liner Inserter Ref 01.26.10.0018Lot 105549 of the Versafitcup Double Mobility system could become rusty after the washingprocedures. This is due to an incorrect raw material used by one of our suppliers during theproduction phase of this lot. For this reason, Medacta International has made the decision toremove all the pieces | Labeling False and Misleading | Medacta USA sent an "URGENT SAFETY NOTIFICATION-VERSAFITCUP DM LINER INSERTER" letter dated May 23, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory for the affected device, immediately sequester it and not use the device in surgery. The letter instructs customers to pack the affected device for shipment to Medacta USA under Return Material Authorization 1590. Customer will receive a replacement within a few days.Questions regarding this recall are directed to the Director of Regulatory and Quality at 805-437-7085 ext. 26. |  | 27 | Worldwide: 18 pieces. In the USA: 9 pieces | Worldwide Distribution--USA (nationwide) including the states of CA, ID, IL, and OH, and 9 products to foreign countries. | Terminated | July 12, 2011 | 50 | 27
Z-2713-2011 | 58544 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Comprehensive Shoulder Instruments Fracture Stem Inserter w/Version Control, REF 31-406901, Biomet Orthopedics Warsaw, IN, Instrument used to introduce the stem into the canal. |  |  | April 07, 2011 | 2011 | Biomet, Inc. | When the surgeon attempts to remove the inserter from the stem, the inserter may become stuck not allowing it to disengage from the stem. The stem may have to be removed from the patient in order to disengage the inserter from the stem. If the stem is removed from the patient, it could potentially cause a significant delay in the procedure or damage/fracture of the humerus which may require inter | Nonconforming Material/Component | The firm, Biomet Orthopedics, sent two "URGENT MEDICAL DEVICE RECALL NOTICE" letters one dated April 7, 2011 and the second letter dated April 12, 2011 which included an additional lot number, to its customers. The letters described the product, problem and actions to be taken. The customers were instructed to immediately locate and discontinue use of the noted product; carefully follow the instructions on the enclosed "FAX Back Response Form"; complete and return the Fax Response Form within three (3) business days via fax to 574-372-1683; use priority carrier for their shipment; If the product was further distributed, they MUST notify and give letter to Hospital personnel, and confirm receipt of this notice by calling, 800-348-9500, extension 3755 or 3756.Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8AM to 5PM. |  | 469 | 469 both products | Worldwide distribution: USA (nationwide) including Puerto Rico; and countries of: Australia, Canada, Chile, Columbia, Costa Rica, Italy, Japan, Korea, Netherlands, New Zealand, Spain and Sweden. | Terminated | April 08, 2013 | 732 | 469
Z-2710-2011 | 58695 | No | MEQ |  |  | Gastroenterology/Urology | PMA | The following kits contain the UROLOGIX, Coolant Bag, CA1111B, Sterile EO, Manufactured for: UROLOGIX, Inc., 14405 Twenty-first Avenue North, Minneapolis, MN 55447. |  |  | April 28, 2011 | 2011 | Urologix, Inc. | Urologix is recalling certain lots of TargisÂż System Coolant Bag devices distributed over a two week period. The devices are packaged into disposable Procedure Kits containing a microwave catheter, a Coolant Bag and a RTU PlusÂż balloon. The reason for this Recall involves three lots of Coolant Bags that contain an increased risk of leakage of water at the Sensor Module. This leak is detectabl | Nonconforming Material/Component | Urologix, Inc. sent an "IMPORTANT RECALL INFORMATION" letter dated April 28, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs the customers to identify the affected kit lots, remove the Coolant Bag pouch(s) and to complete and return a Return Material Authorization form with the product(s). Additionally, an enclosed Customer Acknowledgement Form should be completed and faxed to 763-475-1443.Customers with questions should contact the Director of Regulatory and Quality at 763-475-1400. |  | 119 | 119 | Nationwide Distribution--including the states of AL, CA, CO, FL, GA, MA, MI, MN, NJ, NY, OR, PA, TN, WV, and including Puerto Rico. | Terminated | April 26, 2012 | 364 | 119
Z-2714-2011 | 56848 | No | LNX |  |  | General Hospital | 510(k) | SoftReports versions 1.1.6.x and 1.1.7.x |  |  | June 23, 2010 | 2010 | SCC Soft Computer | SCC Soft Computer is issuing a correction for SoftReports Versions 1.1.6.x released 02/2008 and 1.1.7.x released 06/2008. Potential hazard - when using SoftReports Layout Designer, results displayed in the report may not match the patient information in the report header. | Software design | SCC Soft Computers sent a "RISK TO HEALTH NOTIFICATION TASK" letter dated June 9, 2010 to affected customers. The letter identified the product, problem, and actions to be taken by the customers.Customers were offered an upgrade or hot fix. The letter requested acknowledgement of receipt of the task notification and permission to install the hot fix. |  | 49 | 49 | Worldwide Distribution--USA (nationwide) and the country of Canada. | Terminated | July 26, 2012 | 764 | 49
Z-2715-2011 | 59174 | No | FPB | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Pall Total Nutrient Admixture Filter, 1.2 micron filter |  |  | June 15, 2011 | 2011 | Pall Life Sciences Puerto Rico Llc | Incorrect filter media used (0.2 um filter used instead of the required 1.2 um) within the TNA1 filter used with total parenteral nutrition mixtures. | Labeling mix-ups | Pall Life Sciences LLC, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 15, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers.The letter instructed customers to maintain the product in quarantine. If distributed, the letter request that customers notify the user to discontinue use. Additionally, customers were asked to complete and return an attached Recall Response Form. The letter also included instructions on how to return the product.Please contact the Customer Service Department at 516-801-9923, if you have questions. |  | 15575 | 15,575 units | Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, IL, KS, LA, MA, MO, MS, NC, NY, OK, SD, TX, WA, and WY, and Puerto Rico and the countries of Colombia and Costa Rica. | Terminated | January 10, 2013 | 575 | 15575
Z-2720-2011 | 58971 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | ECMO Heater Model 333, Cardiopulmonary bypass temperature controller. |  |  | May 02, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | The probe jack contacts (terminals) on certain of the firm's ECMO Heater devices are located too close to the solid state relay contacts (terminals) which can lead to the terminals coming into contact with each other. If this situation occurs, the service technician, user, and/or patient could be exposed to electric shock during servicing and/or use. | Device Design | Cincinnati Sub-Zero Products Inc issued an "Urgeant Medical Device Field Action" letter on May 2, 2011 to all affected customers.The letter includes; description of product and problem, affected lots and instructs customers to discontinue its use. The product is to be returned to the Service Department for an update and the customer is instructed to complete and mail Field Action Form.For additional contact the firm at 1-800-989-7373 or (513) 772-8810. |  | 116 | 116 units | Worldwide Distribution and countries of The Netherlands, Kuwait, India, Saudi Arabia, Taiwan and Australia. (USA) Nationwide including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, LA, MA, MD, MI, MIN, MO, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, WA, and WI. | Terminated | March 07, 2013 | 675 | 744
Z-2731-2011 | 58733 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DYONICS POWER Shaver Blades, 4.5mm Long Curved Full Radius Convex, package of 3. |  |  | May 06, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | Product has insufficient weld between the shaft tube and tip, could result in the tip separating from the shaft tube during use. | Process control | Smith &amp; Nephew notified customers by an "URGENT - PRODUCT RECALL 1st NOTIFICATION" letter dated May 6, 2011. The letter describes the product, problem, and actions to be taken. The letter asked customers to call 800-343-5717 (option 3) or send an email to endo.andreturn@smith-nephew.com for a Return Authorization (RA) Number to return product, obtain credit or order replacement blades. Customers are to complete the form and return to the address provided being sure to reference the RA number. For any questions regarding this recall customers were instructed to call 508-337-3731. |  | 744 | 744 units | Worldwide Distribution -- USA (nationwide) and countries of: Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Germany, Hong Kong, Italy, Japan, Mexico, Netherlands, Turkey, and United Kingdom. | Terminated | May 09, 2012 | 369 | 2598
Z-2716-2011 | 58855 | No | MYA | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | ultraView Universal DAB Detection Kit, Part number: 760-500 (05269806001) |  |  | May 16, 2011 | 2011 | Ventana Medical Systems Inc | The recall was initiated because Ventana Medical Systems, Inc. (Ventana), hasidentified the presence of precipitate in some DAB chromogen dispensers in two lots of ultraViewUniversal DAB Detection Kit (part number 760-500, lot A05885 and lot B00988) and one lot of iViewDAB Detection Kit (part number 760-091, lot B01057A.) | Pending | The firm, Ventana, sent two letters dated May 16, 2011, an "URGENT Medical Device Recall" notice to customers who received the affected lots and an "URGENT Medical Device Correction" to customers of unaffected lots. The letters described the product, problem and actions to be taken. The customers were instructed to discontinue use of any kits from the lots identified above; follow their internal standard operating procedures for assessing potential impact to reported patient results, If their laboratory had already used kits from these lots; conduct visual inspections of any DAB chromogen dispensers in use as part of their daily maintenance routine; inspect their current inventory and contact their local customer support team (800-227-2155) to replace any unused or partially used kits from the affected lots.Ventana highly recommends use of same-slide controls, especially when testing for targets of therapy, to mitigate any impact on patient results.If you find a dispenser containing precipitate in lots not identified in this letter or have any questions and concerns, please notify/contact your local customer support team (800-227-2155) . |  | 3990 | 3990 units (2981 in US) | Worldwide distribution: USA (Nationwide) and countries of Canada, Germany and Israel. | Terminated | February 10, 2012 | 270 | 8030
Z-2719-2011 | 58465 | No | HXI | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PERPOS PLS SYSTEM, CATALOG NO. 9045-02 |  |  | March 23, 2011 | 2011 | Interventional Spine Inc | This recall has been initiated due to K-wires with a distal tip diameter that may be greater than specified causing an interference fit with the canulated screw. Removing the K-wire during a procedure may dislocate the implant resulting in the implant loosening or being pulled out or position.Use of this product may result in a delay of the surgical procedure and/or pull-out of the facet screw | Pending | Interventional Spine sent an RECALL NOTIFICATION letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. The customer was instructed to take the following actions: Immediately examine inventory and quarantine product subject to recall. In addition, if this product was further distributed, please notify the Vice President Quality Assurance at once. Consignee were asked to also complete and return the enclosed response form as soon as possible. Consignee with questions were directed to the Consignees' Regional Director or the Vice President Quality Assurance.For further questions, please call (949) 472-0006 |  | 864 | 864 units total for all products | Nationwide Distribution including the states of to NC, NY, TN &amp; VA | Terminated | February 14, 2013 | 694 | 864
Z-2734-2011 | 58678 | No | HTQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Foundation Knee System Instrumentation, Tibial Broach Assembly. |  |  | April 21, 2011 | 2011 | Encore Medical, Lp | Strike plate may dislodge from the broach stem handle during impaction of the device. | Process control | Encore Medical, Lp (d.b.a. DJO Global) sent out a "Urgent Field Safety Notice" dated April 20, 2011 to all affected customers. The firm also e-mailed the consignees as well. The letter described the product, affected lots, reason for recall and provided recommendations including; asking customers contact firm for a replacement, provided a "Field Safety Response Form" to fill out and return. The firm also e-mailed the consignees as well. For additional information please contact the Director of Commercial Logistics at (512) 834-6330 or the International Customer Service at 512 834-6275. |  | 15 | 15 units | Louisiana and Arkansas and the countries of Germany, Italy, United Kingdom and Japan. | Terminated | April 06, 2012 | 351 | 15
Z-2733-2011 | 58686 | No | MXC | Medical magnetic tape recorder. | Cardiovascular | Cardiovascular | 510(k) | SJM Confirm, Implantable Cardiac Monitor (ICM) Models: DM2100, DM2102 (IDE). |  |  | March 31, 2011 | 2011 | St Jude Medical CRMD | The recall was initiated because St. Jude Medical made new software available for the SJM Confirm DM2100 implantable cardiac monitor (ICM) that incorporates enhancements to the devices sensing algorithm. These enhancements are intended to provide improved sensitivity and specificity for diagnosing arrhythmias. The user is presented with the option to upgrade new software into the SJM Confirm ICM | Software change control | St. Jude Cardiac Rhythm Management Division sent an "Important Product Correction" letter on March 25, 2011 to all affected customers.The letter explains the reason for the recall and provided recommendations including a software upgrade by SJM Field Clinical Engineer or Technical Services. Customers were recommended to fill out and return a "Physician Acknowledgement Form".For additional information please contact your local St.Jude Medical Sales Representative, Field Clinical Engineer or the Technical Services Department at 800-722-3774 |  | 83 | 83 units | Worldwide distribution: (USA) Nationwide including the States of CA, WA, FL, WV, NY, MD and LA, and the country of Germany | Terminated | July 19, 2011 | 110 | 83
Z-2748-2011 | 58525 | No | KMF | Liquid bandage. | General Hospital | General &amp; Plastic Surgery | 510(k) | smith&nephew; NO-STING SKIN-PREP No-Sting Protective Wipes Effective protection between tape and skin Reduces risk of tape stripping Improves tape, film and appliance adhesion No-sting, non-irritating Reduces friction:Packaged 50 wipes per box. |  |  | April 04, 2011 | 2011 | Smith &amp; Nephew Inc. | Smith &amp; Nephew, Inc., St. Petersburg, FL is recalling multiple lot numbers of the following products: REMOVE Universal Adhesive Remover Wipes, UNI-SOLVE Adhesive Remover Wipes, SKIN-PREP Protective Wipes, PERI-PREP Protective Wipes, and NO-STING SKIN-PREP Protective Wipes. These products were manufactured by H&P; Industries dba The Triad Group who has initiated a recall of products manufactured i | Material/Component Contamination | Smith &amp; Nephew Inc. sent an "URGENT DEVICE CORRECTION" letter dated April 8, 2011 to all customers who received the recalled products. The letter described the product, problem and the appropriate actions to be taken. Customers are instructed to immediately discontinue using the recalled products, examine their inventory, and follow the return or disposition instructions provided. The letter instructed customers to seek an alternative product for use. Distributors are instructed to provide the same instructions to their customers supplied with the recalled products. If you have any questions contact the Recall Hotline Center at 1-888-613-0271. |  | 108490 | 108490 cartons (471800 wipes); 192,945(kits containing wipes) | Worldwide Distribution -- USA (nationwide) including the states of: NC, PA, NM, OR, VA, TN, NY, OR, AL, IL, MN, TX, and Puerto Rico and the countries of: Australia, Canada, England, France, Germany, Hong Kong, Japan, Mexico, New Zealand, Singapore, South Africa, Taiwan, and United Kingdom | Terminated | July 18, 2011 | 105 | 55600840
Z-2747-2011 | 58845 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Quinton BETA-CAP Adapter; a Peritoneal Dialysis accessory to the PD catheter used to connect the catheter to the PD transfer set; Covidien product code 8814-661001, Baxter product code 5K4560 |  |  | May 19, 2011 | 2011 | Baxter Healthcare Renal Div | A change made in the Quinton BETA-CAP adapter used to connect the patient PD catheter to the Baxter transfer set has resulted in a reduction in thread engagement which may lead to an increased risk of separation and the potential to develop peritonitis. | Component change control | Baxter sent an URGENT PRODUCT RECALL letter dated May 19, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The letter described the problem of loose connections when using the Quinton BETA-CAP with Baxter transfer sets. Baxter requested the dialysis centers to contact their patients currently using a Quinton BETA-CAP adapter to schedule an office visit with their attending physician as soon as practical tohave the connection inspected or replaced. If they choose to continue use of the Quinton BETA-CAP adapter, Baxter recommends the clinician exercise extra caution and vigilance to ensure the connection between the Quinton BETA-CAP adapter and the Baxter Transfer Set is secured and tight. The clinicians were instructed to also check their inventory for the Quinton Beta Cap adapter 8814-661001 purchased directly from Baxter, and contact Baxter at 888-229-0001 to arrange for the return and credit of the unused product. Baxter recommends that the Covidien BETA CAP adapter no longer be used and to instead exclusively use the Baxter Titanium Adapter (5C4129) which has sufficient thread geometry for maintaining a secure connection with the Baxter Transfer Set. The customers were also instructed to complete and fax the enclosed customer reply form to (847) 270-5457, indicating the quantity of Quinton Beta Cap adapters being returned and acknowledging the understanding and dissemination of the information in the letter. |  | 13942 | 13,942 units | Class II Recall -- Nationwide Distribution | Terminated | May 21, 2012 | 368 | 13942
Z-2744-2011 | 58919 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Cambridge Endoscopic Autonomy Laparo-Angle 5mm Metzenbaum Scissors with Monopolar Cautery, 45 cm |  |  | May 20, 2011 | 2011 | Cambridge Endoscopic Devices, Incorporated | PET shrink sleeve insulation material dislodgingfrom the distal tips of Autonomy Laparo-Angle (ALA) Metzenbaum Scissors | Process control | Cambridge Endoscopic Devices, Inc. sent an "IMPORTANT MEDICAL DEVICE RECALL" letter dated May 20, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers are instructed to return unused inventory using RGA Number 301. Returned devices will be replaced.Questions regarding this recall, call 508-405-0790 for assistance. |  | 255 | 255 units | Worldwide Distribution--USA (nationwide) and the countries of Armenia , Australia, Austria, Chile, China, Denmark, Hong Kong, India, Ireland , Italy, Japan, Saudi Arabia, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | July 24, 2012 | 431 | 902
Z-2741-2011 | 58655 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | JASMINE Full Body (Patient) Lift, electrically-powered device with a 500 lb. weight capacity. |  |  | April 01, 2011 | 2011 | Invacare Corporation | Component Defect. The rear mast connector on certain patient lifts distributed by the firm, may break. Failure of the rear mast connector, while in use would result in a patient being dropped, possibly resulting in serious injury or death to the patient or care provider. | Other | INVACARE contacted their customers via telephone on 4/8/2011 and 4/11/2011, and informed them of this recall. During these telephone conversations, the customers were asked to immediately cease use of the JASMINE patient lifts and to quarantine all product on-site until further notice from Invacare. The firm followed up on the telephone recall notification with a recall letter dated: April 18, 2011. The recall letter was sent via registered mail or Federal Express so as to verify customer's receipt of the recall notification. The recall letter refers the customer/care provider back to the telephone conversations. and notifies the customer of the possibility of failure of the rear mast connector which could result in a patient's possibly being dropped from the patient lift. The letter asks that the customer remove the suspect patient lift(s) from service and quarantine the device in order to prevent further use until such time as an Invacare Rep. contacts the customer by telephone and makes arrangements to pick-up the affect product. The letter is accompanied by a Response Form and a Product Tracking Sheet which lists all of the JASMINE devices shipped to the customer's facility. The customer is asked to complete the attached Response Form and to fax the form back to Invacare at the telephone number provided on the bottom of the form. The signed Response Form provides information to Invacare documenting the customers receipt of the form and understanding of the instructions contained therein, and also; the number of affected devices remaining on-hand at the customer's location, and the number of devices further distributed by the customer. In those instances in which the products have been further distributed, the customer is asked to 'carry out the instructions of the recalling firm and extend the recall to their sub-account customers. |  | 274 | 274 units | Worldwide Distribution - USA including DC, FL, GA, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NJ, NV, NY, OH, OK, PA, RI, SC, TN, VA, WA, and WI and the countries of Canada, Ireland, United Kingdom, Norway, Austria, Germany, and Australia. | Terminated | February 28, 2013 | 699 | 274
Z-2745-2011 | 58908 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HALO Ultra Ablation Catheter: |  |  | May 25, 2011 | 2011 | Barrx Medical, Inc. | Complaints of malfunction of the HALO ULTRA Ablation Catheter related to the device pivot mechanism's inability to return the electrode cap to a flat neutral position. | Component design/selection | BARRX Medical, Inc. sent a "PRODUCT RECALL NOTIFICATION" letter dated May 25, 2011 via Federal Express to all of their customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of the recalled product. BÂżRRX Medical Territory Managers will be in contact with each facility and arrange for pick up of the recalled product. The "Steps for Voluntary Recall" form that is enclosed with each letter provides customers with complete instructions on segregating the product, completing the Tracking/Verification Form, obtaining Returned Good Authorization (RGA) number and preparing product for pick up. Customers are to complete and return the Tracking/Verification Form by fax to 408-738-1741 which should include amount of product on hand, contact name and hospital/distributor information. Questions should be directed to BÂżRRX Medical Customer Service at 888-662-2779. |  | 936 | 936 devices | Worldwide Distribution -- USA (nationwide) and the countries of the United Kingdom and China | Terminated | September 07, 2011 | 105 | 936
Z-2746-2011 | 58907 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Wire Caddy, Stryker Instruments Wire Caddy, REF 505-505, STERILE, QTY 1, MANUFACTURED FOR STRYKER INSTRUMENTS, KALAMAZOO, MI USA 49001 (269) 323-7700 (800) 253-3210. |  |  | May 27, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | The towel clamps packaged with the wire caddy have been found to have blood stains on them which could result in infection or blood borne pathogen transmission. | Material/Component Contamination | Stryker Instruments sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter to all affected customers. The letter described the product, problem, and actions to be taken by the customers.The letter instructs customers to quarantine affected product and return the recalled product using the shipper supplied to them by Stryker. Replacements will be issued to customers upon receipt of the product.For questions regarding this recall please call 1-800-800-4236 ext. 3808. |  | 522 | 87 boxes, 435 untis | Nationwide Distribution-- including the states of Alabama, Arizona, California, Florida, Illinois, Kansas, Maryland, Michigan, Missouri, Pennsylvania, Texas and Virginia. | Terminated | August 03, 2011 | 68 | 522
Z-2723-2011 | 58536 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Pathfinder NXT Closure top driver - short, REF 3566-2, Zimmer Inc., Warsaw, IN |  |  | May 20, 2011 | 2011 | Zimmer Inc. | Zimmer Spine, Inc. is initiating a voluntary recall regarding an increased risk of rounding or partial rounding of the hex tip on the 3566-1 (long) and 3566-2 (short) Closure Top Driver used for insertion, provisional tightening, and final locking of Closure Tops during the screw and rod construct assembly of the Pathfinder NXT Pedicle Screw System.When the hex tip of the Closure Top Driver is | Device Design | Zimmer, Inc. issued an "URGENT MEDICAL DEVICE RECALL" letters dated April 2011, to all affected customers. The letter identified the product, problem and actions to be taken. The customers were instructed that (THE GOLD CLOSURE TOP DRIVERS 3566-1 and 3566-2 ARE NOT TO BE USED FOR FINAL TIGHTENING); an alternate instrument for final construct tightening will be included in the instrument set 3570-1 short and 3570-2 long (These instruments are silver instead of gold for easy identification and do not include a retention feature for holding the Closure Top); to ensure final drivers are added to each PathFinder NXT set and inspect Closure Top Drivers for rounded hex wear; contact their customer service representative if they do not have both versions; and complete and return all attached forms via fax to: Zimmer Spine at 512-258-0995. The letter states that the firm will revise the surgical technique to show the updated uses of the screwdrivers. Please keep Zimmer Spine informed of any adverse events associated with the device or any other Zimmer spine product by calling 800-326-0635. For any further questions or concerns contact the Vice President Quality Assurance &amp; Regulatory Affairs at 800-655-2614. |  | 135 | 135 devices | Worldwide Distribution -- USA (nationwide) including states of: AZ, CA, WI, MI, KS, OK, TX, GA, TN, FL, NJ, CT, PA and the countries of: Australia and Italy. | Terminated | October 07, 2011 | 140 | 269
Z-2739-2011 | 58562 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Reverse Shoulder Prosthesis System RSP Humeral Socket Shell DJO 9800 Metric Blvd. Austin, TX 78758 STERILE R Model numbers: 508-00-000 and 508-00-008. |  |  | April 18, 2011 | 2011 | Encore Medical, Lp | The device has the potential to not appropriately mate with the humeral stem. | Nonconforming Material/Component | DJO Surgical sent an "URGENT FIELD SAFETY NOTICE" letter dated April 18, 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers.The letter instructed customers to quarantine affected products and to contact Customer Service to obtain replacement products and a RMA.Questions are directed to the Director of Commercial Logistics at 512-834-6330. |  | 81 | 81 units | Worldwide Distribution--USA (nationwide) including the states of CA, FL, IL, LA, MA, MD, ME, NM, NY, OK, PA, SC, TN, TX, UT, and WA and the countries of Germany, Italy, and United Kingdom. | Terminated | January 25, 2012 | 282 | 81
Z-2738-2011 | 58547 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Large Roller Pump 6" Diameter, Catalog 801041, Terumo Cardiovascular Systems, Corp., Ann Arbor, MI |  |  | March 31, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has received multiple reports of the lower roller pump tube clamp mechanism failing to function in a small population of devices, making it difficult to remove or insert the tubing. If the failure occurs, the clamp dial will no longer be able to control opening/closing of the tube clamp mechanism. | Device Design | Terumo Cardiovascular Systems sent an URGENT MEDICAL DEVICE CORRECTION letter to all affected customers dated March 21, 2011. The letter identified the product, the problem and the action needed to be taken by the customer. An additional URGENT MEDICAL DEVICE RECALL correction notification and cover letter was sent to all consignees on XX/XX/XX to explain the reason for the expansion of the recall and will provide an update to the timing of implementation of the field correction.Expansion now includes devices manufactured between 2006 and March 2012.The correction strategy is to issue the customers a letter and addendum to the Operator's Manual that specifies what the user should do in the event of a tube clamp failure for any reason. The letter states that the firm is still investigating the issue and the correction. The letter states that in the coming months when a resolution is identified, consignees may be contacted to make arrangements for a upgrade to the system. There is no specific timeline for the correction to be made.If you have any further questions, please call a Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours Monday - Friday, 8 Am - 6 Pm EST. |  | 111 | 111 devices. | Worldwide Distribution -- USA (nationwide) and the countries of Argentina, Australia, Belgium, Canada, Chile, Columbia, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, Indonesia, Japan, Jordan, Korea, Kuwait, Malaysia, Mexico, Pakistan, Republic of Georgia, Russia, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand, The Phillipines, Turkey, United Arab Emirates and Vietnam. | Terminated | August 28, 2013 | 881 | 3679
Z-2740-2011 | 58792 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | NucliSens EasyQ Enterovirus serotype 68 v1.1: Catalog number: 200300 |  |  | May 10, 2011 | 2011 | bioMerieux, Inc. | The NucliSENS EasyQ Enterovirus Reagents are unable to detect Enterovirus serotype 68, in respiratory samples. | Device Design | BioMerieux Vitek, Inc. sent an "Urgent Product Correction Notice" letter on/about May 10, 2011.The letter included; a description of the product, problem, possible impact to the patient and recommendations. Recommendations included instructing customers to discontinue the use of the kit for the detection of Enterovirus 68 and Rhinovirus 87 serotypes and noted the corrected IFU will be available the first week of April 2011. An Acknowledgement form was attached to be returned via fax.For additional questions please contact your local BioMerieux Clinical Customer representative at (800) 682-2666, option 3, option 2. |  | 42 | 42 kits | (USA) Nationwide distribution including the states of MD, MI, MN, MT, NY and PA | Terminated | October 06, 2011 | 149 | 42
Z-2742-2011 | 58847 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | TEG Hemostasis System Level II Control: |  |  | March 04, 2011 | 2011 | Haemoscope Division of Haemonetics Corp | Haemoscope received a report that during the process of Validation, Level II QC was running with an out-of-range low MA (Maximum Amplitude). | Other | Haemoscope, a Division of Haemonetics Corp, sent a letter dated March 2, 2011 to all of the affected customers. The letters informed the customers that Haemoscope had received a higher than normal inquiry rate for the Level II Quality Control Kits, part 07-008, lots HMO196 and HMO 197. The accounts were requested to return their stocks of those two lots for further testing and evaluation and to complete the enclosed acknowledgement form, confirming receipt of the letter and the number of kit boxes being returned, and fax it back to Haemoscope at 847-588-0455. The returned kits will be replaced at no charge.For additional information please contact Haemoscope at 847-588-0453, ext-217 . |  | 568 | 568 kits | Worldwide Distribution: (USA) Nationwide and the countries of Canada, Czech Republic and the Philippines | Terminated | September 25, 2012 | 571 | 568
Z-2760-2011 | 58335 | [6 empty] | J-Vac Drain Adaptor Manufactured for: Ethicon, Inc. Somerville, New Jersey, 08876-0151. |  |  | March 24, 2011 | 2011 | Ethicon, Inc. | There is a potential for the sterile barrier to be compromised in multiple lots of Blake silicone drains, Blake silicone drain kits, Blake cardio connectors, J-VAC reservoirs, J-VAC drain adaptors. | Packaging process control | Ethicon, Inc. sent an "URGENT VOLUNTARY PRODUCT RECALL" letter dated March 24&amp; 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue use of the products and to return the products to Stericycle. The letter includes shipping labels for product return. Ethicon will send replacement product. Customers are asked to return a completed Business Reply Card to Stericycle regardless of whether or not they have the products subject to this recall.Customers with questions regarding this recall are directed to call 1-866-664-1401. |  | 846990 | 846, 990 | Worldwide Distribution-- USA (nationwide) and Puerto Rico and the countries of Armenia, Australia, Austria, Azerbaijan, Columbia, Cyprus, Czech Republic, Denmark, French Guinea, French Polynesian, Germany,Israel, Italy, Japan, Kuwait , Netherlands, New Caledonia, New Zealand, Norway, Singapore, Slovenia, South Africa, Spain, Belarus, Belgium, Brazil, Canada, Chile, Dominica Republic, Egypt, Estonia, Finland, France, Greece, Guadeloupe, Hong Kong, Hungary, Ireland, Latvia, Lithuania, Malaysia, Martinique, Mozambique, Portugal, Reunion, Russia, Saudia Arabia, Sweden, Switzerland, Ukraine, and United Kingdom. | Terminated | January 10, 2014 | 1023 | 846990
Z-2753-2011 | 59024 | No | NEU | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bariatric Target Ring (Nitinol) (a thin, circular nitinol ring) Model # K-9536-N, Lot #08-001, Lot #08-002 &amp; Lot #08-003. One ring per plastic pouch, three rings are then placed into a plastic box with lid. This box is then placed into a solid white box. Item is shipped and sold non-sterile. |  |  | May 24, 2011 | 2011 | Kapp Surgical Instrument Inc | During an FDA Inspection on 3/21/11 - 3/22/11 at Kapp Surgical an uncontrolled instruction for use was discovered which contained sterilization parameters which differed from the instructions contained in the device master record. Kapp Surgical decided to conduct the recall because we cannot be sure that incorrect instructions were not shipped. | Labeling mix-ups | Kapp Surgical Instrument, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to check their inventory and quarantine any product affected by this recall. The firm will make arrangements for returning the products. Replacements will be made by Kapp Surgical.Questions about this recall are directed to the QA Manager at 800-282-5277. |  | 222 | 222 units (rings) | Product was distributed to the following states: OH, PA &amp; SC. Product was also shipped to Saudia Arabia. | Terminated | February 19, 2013 | 637 | 222
Z-2727-2011 | 59110 | [6 empty] | Stainless Steel, Progressive Lower, .016: Arch Wires. |  |  | January 07, 2008 | 2008 | Highland Metals | Boxes of stainless steel orthodontic wires were mislabeled as size.018" | Employee error | Highland Metals sent out "Highland Metals Advisory Notice Mislabeld Products Alert" and "Advisory Notice" letters from January 11, 2008 to April 18, 2008 to all affected costumers.The letter included: product name and affected lots. It advised them of the mislabeling and instructed them to return the product for replacement.For additional information contact Highland Metals at (408) 271-2955 |  | 2200 | 22 packs of 100 | Worldwide distribution including the countries of Colombia, France, Scotland, Germany, Australia, Mexico, England, United Kingdom, Thailand; and (USA) Nationwide including the states of OK, CA, NJ, AZ, MD, LA, IL, PR, FL, NE, IN, TX, NH, NY, CO and PA. | Terminated | July 07, 2011 | 1277 | 14600
Z-2755-2011 | 58850 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ElevateÂż System with InteProÂż Lite", Anterior and Apical Prolapse Repair System, REF720093-01, Sterile EO, Manufactured by:American Medical Systems, Inc., 10700 Bren Road West, Minnetonka, MN 55343 USA |  |  | May 09, 2011 | 2011 | American Medical Systems, Inc. | This is an important notification regarding the ElevateÂż Anterior and Apical Prolapse Repair Systems, Part Number 720093-01, manufactured between August 3, 2010 and January 17, 2011. We have recently received complaints stating that the eyelets were missing from the center graft of some units within the scope of this issue.The complaint rate to-date for the affected product is 0.045%, however th | Process control | American Medical Systems (AMS) sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 9, 2011 to all customers who received the recalled product. The letter identified the product, problem, and emphasize that the recalled product should not be implanted. The customers were instructed to examine their inventory for the recalled product and call AMS Customer Service at 1-800-328-3881 (select option 1) for an exchange of the product. For assistance in evaluating inventory the customers are instructed to contact their local AMS sales representative. |  | 8563 | 8563 | Worldwide Distribution -- USA (nationwide) including DC and Puerto Rico and countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, FRANCE, GERMANY, PORTUGAL SPAIN, UNITED KINGDOM, ARGENTINA, COLOMBIA, ESTONIA, FINLAND, GREECE, IRELAND, ISRAEL, ITALY, LEBANON, NETHERLANDS, NEW ZEALAND, SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SWEDEN, SWITZERLAND, TAIWAN REPUBLIC OF CHINA, VENEZUELA | Terminated | June 28, 2012 | 416 | 8563
Z-2761-2011 | 55227 | No | CGR | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active Cortisol EIA, 96 Wells, DSL-10-2000 |  |  | July 20, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the listed lot numbers of Active Control ELISA reagent will result in falsely high values for patient samples due to a shift in the calibration of the assigned standard values.Beckman Coulter recommends that patient results, generated with the affected lots, be evaluated in the context of other diagnostic tests and the clinica | Other | Beckman Coulter sent a Product Corrective Action (PCA) letter dated July 13, 2009. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use until further notice and discard remaining affected product. Customers were instructed to contact Beckman Coulter at 1-800-231-7970.Newer lots (lot number 991069 or higher) of the product have been recalibrated to restore the performance as described in the Instructions for Use (IFU). For method correlation information, please contact Technical Support at the phone number listed below.Consignees were asked to share this information with laboratory staff and retain the PCA letter for their Quality System documentation, and complete and return the enclosed Response Form within 10 days so that the firm can be assured that they received the notification.Questions regarding this notification were directed to the DSL Technical Support Center at 1-800-231-7970 in the United State and Canada, or their local Beckman Coulter Representative. |  | 1369 | 1369 Kits | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | April 27, 2012 | 1012 | 1369
Z-2764-2011 | 58976 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac Linear Accelerator; |  |  | May 18, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The Coolant System input water supply manifold may leak in some Clinac Linear Accelerators, posing a risk of electrical shock to any person working within the protective housing. | Other | Varian Medical Systems, Inc. Oncology Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 18, 2011, via certified mail, return receipt requested. The letter identified the product, problem and recommended actions to be taken. The letter states that customers will be contacted by Varian Service to arrange for replacement of the recalled product. For additional information contact Varian Medical Systems, Inc at 888- 827-4265. |  | 93 | 93 | Worldwide Distribution -- USA (nationwide) and countries of: Australia, Columbia, Denmark, France, Ireland, Italy, Japan, Morocco, Poland, Portugal, Saudi Arabia, South Korea, Spain, and United Kingdom. | Terminated | March 13, 2012 | 300 | 93
Z-2763-2011 | 58974 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA, versions 8.0 and 8.1; |  |  | May 17, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue has been detected when using the RT Chart function in ARIA versions 8.0 and 8.1. When users select the "Convert to Setup Field" option, a Set-up field is created, that removes the DMLC plan components in order to create an open collimator field for setup purposes. | Software design | Varian Medical Systems, Inc. sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 17, 2011, via Certified Mail to US; Federal Express to Canada. Foreign distribution will be handled by Varian Managers. The letter contained recommended actions. For additional information contact Varian Medical Systems, Inc at 888- 827-4265.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** |  | 357 | 357 | Worldwide Distribution -- USA (nationwide) and countries of Algeria, Australia, Austria, Belgium, Brazil, Chile, China, Canada, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, India, Iraq, Ireland, Italy, Japan Jordan, Kuwait, Malaysia, Morocco, Netherlands, Oman, Pakistan, Panama, Portugal, Romania, Russia, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Tunsia, Turkey, Ukraine, United Arab Emirates, and United Kingdom, | Terminated | March 15, 2012 | 303 | 357
Z-2762-2011 | 58830 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Picture Archiving Communication System |  |  | May 10, 2011 | 2011 | Siemens Medical Solutions, USA, Inc | For systems upgraded to syngo Dynamics version 9.0, if a legacy report generated by version 3.x or earlier is opened in the version 9.0 portal, the legacy report may be deleted and replaced with a new report from the version 9.0 reporting system or a blank report. | Software design | Siemens Medical Solutions, USA, Inc sent a letter entitled "Customer Safety Advisory Notice" to all direct accounts on May 10, 2011. The letter stated the name and version of the software, problem noted, a work-around to be implemented until a service patch is created to fix the problem, and that the letter should be placed in the software's operator's manual to be seen by all who utilize the software. |  | 38 | 38 | Nationwide Distribution | Terminated | July 09, 2012 | 426 | 38
Z-2765-2011 | 58566 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Diagnost 94, Diagnost 96, Diagnost 97 Universal X-Ray System |  |  | April 18, 2011 | 2011 | Philips Healthcare Inc. | Updated Instructions For Use: During an examination where the patient is positioned on a -rotatable- chair with knees pointed to the rear of the system, the patients lower legs or ankles may be injured/broken | Device Design | Philips Healthcare notified accounts by letter on April 18, 2011. The letter advised users of the problem and provided updated instructions for proper placement of the patient during the examination. An Updated Instruction For Use was provided.For further information and support customers were instructed to contact the Philips representative. |  | 52 | 52 units | Nationwide Distribution | Terminated | February 15, 2013 | 669 | 52
Z-2766-2011 | 58994 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam Linear Accelerators (aka Trilogy Mx) |  |  | May 19, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Imaging arms of the TrueBeam Accelerator may have loose encoder pulleys that could lead to inaccurate readout of arm geometry. | Device Design | Varian Medical Systems, Inc. Oncology Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 19, 2011, via certified mail, return receipt requested in the US. The letter identified the product, problem and recommended actions to be taken. For additional information contact Varian Customer Support at 1-888- 827-4265. |  | 77 | 77 | Worldwide Distribution -- USA (nationwide) and countries of: Canada, Denmark, France, Germany, Isreal, Italy, Netherlands, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 13, 2012 | 299 | 77
Z-2767-2011 | 58981 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | VITEK 2 System, VITEK 2 60 110V (Catalog #27225) and VITEK 2 60 220V (Catalog #27226). |  |  | June 02, 2011 | 2011 | Biomerieux Inc | Cards may not load into the incubator correctly causing them to be loaded in the wrong order which could result in the results being linked to the wrong accession number and reporting results erroneously. | Device Design | Biomerieux, Inc. sent an "URGENT PRODUCT CORRECTION NOTICE" dated June 2, 2011 to all affected customers. The letter describes the products, problem, and actions to be taken by the customers.The letter explains how to recognize the issue and instructs customers to take certain steps to resolve the problem. Additionally, an Acknowledgement Form is attached to the letter for customers to complete and return via fax to 919-314-4539. A firmware update will be applied by Biomerieux, Inc. technical support representatives.Questions regarding this recall are directed to the Customer Service Department at 1-800-682-2666, option 3. |  | 627 | 627 instruments | Worldwide Distribution-- USA (nationwide) and the countries of Austria, Australia, Belgium, Brazil, Canada, China, Colombia, Germany, France, India, Israel, Italy, Japan, Korea, Mexico, The Netherlands, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, United Arab Emirate, and United Kingdom. | Terminated | June 06, 2013 | 735 | 1734
Z-2769-2011 | 58993 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The SIU (System Interface Unit) is a component in the Tool Connection Unit of the PercuNav system. It is part of the position sensor system used to calculate the position of the instrument. |  |  | June 01, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | PercuNav Tool Connection Unit may contain an electronic component that does not meet specification, which could cause registration difficulties and/or inaccurate instrument tracking. | Nonconforming Material/Component | Philips Healthcare sent an "MEDICAL DEVICE CORRECTION" letter dated May 27, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The notification letter instructs customers to follow all directions in the Instructions for Use. A Philips Healthcare representative will schedule an appointment to replace the Tool Connection Unit at no charge to the customers.If further information is needed concerning this recall please call Philips Customer Service at 866-767-7822. |  | 10 | 10 units (6 in US and 4 internationally). This was later changed to 7 US and 3 International) all serial codes remain the same | Worldwide Distribution--USA (nationwide) including the states of CA, GA, OH, WA, and Washington DC, and the countries of Australia, India, Netherlands, and South Korea. | Terminated | February 15, 2012 | 259 | 10
Z-2781-2011 | 54621 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | DICOM Query/Retrieve, version 1.1.4. Individual part # 9733808, or as part of a system order, DICOM Q/R Software Kit, part # 9733933. DICOM Query/Retrieve software is installed on Medtronic StealthStation systems. |  |  | June 02, 2011 | 2011 | Medtronic Navigation, Inc. | Software allows incorrect patient or patient history to be displayed. | Software design | Medtronic Navigation sent a "FIELD SAFETY NOTICE" letter dated June 2, 2011 to all affected customers. The letter identified the product, problem and description of the corrective action. The letter states that Medtronic field representatives will visit each consumer/user and upgrade the version of DICOM Q/R to 1.1.5. Further information is available by contacting Medtronic Navigation Technical Services at 1-800-595-9709. |  | 90 | 90 | Worldwide distribution - USA (nationwide) -- including states of: TX, NC, ID, FL, OH, UT, CA, TN, IL, AZ, NE, LA, PA, OR, AR, MI, NJ, MD and countries of: Canada, Denmark, Finland, Great Britain, Hungary, Ireland, Japan, Korea, Netherlands, Poland, South Africa, Spain, and Taiwan. | Terminated | July 12, 2011 | 40 | 90
Z-2782-2011 | 57012 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Site Rite Vision Ultrasound System, REF 9770032, |  |  | June 21, 2011 | 2011 | Bard Access Systems | Purchasing erroneously ordered the incorrect fan using the distributor's part number instead of using the specified manufacturer part number. | Process control | Bard Access Systems, Inc sent out an "Urgent" Medical Device Recall (User Level)" dated June 2011 to all affected customersThe letter included: product name, description of problem and requested the return of units bearing "B50" on the fan. Bard Access Systems, Inc. will replace the incorrect fan with the correct fan. For additional information call 1-724-779-3000. |  | 86 | 86 units | Nationwide (USA) distribution including the states of AL, AR, AZ, CA, CO, FL, IA, IN, MD, MI, MO, MS, NC, NJ, NM, NY, OH, PA, RI, SC, TX, UT, VA, VT, and WA. | Terminated | July 11, 2012 | 386 | 86
Z-2788-2011 | 59006 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System; |  |  | May 20, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Eclipse Monte Carlo (eMC) algorithm where not-equidistant data points in the open beam profile used during configuration may lead to inaccurate results. | Software design | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter via certified mail dated May 20, 2011, to all affected customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the product from use and advise users to manually initiate the configuration process for resampling to be applied when a previous calculation model is copied. Customers were asked to sign and return the notice to confirm receipt. For questions call USA and Canada 1-888-827-4265For Europe call +41 41 749 8844 |  | 10356 | 10,356 | Worldwide Distribution - USA (nationwide) | Terminated | July 13, 2012 | 420 | 10356
Z-2780-2011 | 58969 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System; Release 6.5 through 10 |  |  | May 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Eclipse versions, user editing of the MLC Leaf positions may not automatically invalidate calculated dose distributions. | Software design | Varian Medical Systems, Inc sent a "Urgent Medical Device Correction Urgent Field Safety Notice" letter dated May 13, 2011 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter included the description of the product, problem and recommendations and includes the release of a service pack in the future. For additional information please contact the Varian Oncology Help Desk at 888-VARIAN5 (888- 827-4265) |  | 10664 | 10,097***12-12-12 AMENDED TO: 10664 units*** | Worldwide distribution: Nationwide (USA) and the country of Canada. | Terminated | August 15, 2014 | 1190 | 10664
Z-2787-2011 | 58897 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens YSIO - solid state x-ray imager |  |  | April 28, 2011 | 2011 | Siemens Medical Solutions USA, Inc | There is a potential hazard to staff or patients when lowering the table height on the Ysio system. If staff or patients lower extremities are placed underneath the table and the user lowers the table top, it is possible that parts of the table may collide with extremities and cause injuries. | Other | Siemens sent a Safety Advisory Notice to all affected customers. The Safety Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to place the Safety Notice in the System's instruction for use. If product was distributed further, customers were to forward the safety notice to the new owner and inform the firm of the new owner. Customers were instructed to complete the Acknowledgement of Receipt Form and return to the firm. |  | 141 | 141 | Nationwide Distribution including Washington, DC and the following States: AR, CA, CO, FL, ID, IL, IN, IA, KS, KY, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI. | Terminated | April 10, 2014 | 1078 | 141
Z-2771-2011 | 55164 | No | LIG | Vitamin B[bdi1][bdi2] test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCel DxC 880i, 860i,680i and 660i Synchron Access Clinical Systems |  |  | May 07, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports of inaccurate Access Intrinsic Factor Ab test results displayed on, or printed from, the UniCel Dxl portion of an integrated workstation listed above. The problem is limited to numerical Intrinsic Factor Ab results produced on integrated systems running UniCel Dxl system software version 4.0, 4.1, or 4.2. The DxC portion of your inte | Software design | Beckman Coulter, Inc. sent out an "Urgent: Product Corrective Action" letter dated May 19, 2009 to all affected customers.The letter explained the reason for recall and provided recommendations including: -System software currently under development will correct this problem, and will be made available soon. Until then, please refer to UniCel DxC reports, user interface information, and LIS transmissions, which contain accurate Access Intrinsic Factor Ab test results. -Refer to the UniCel DxC system documentation for information on displaying and printing results from the DxC portion of an integrated workstation, or from your laboratory information system (LIS) -Past results for these assays only need to be reviewed if they were reported from the Dxl portion of an integrated workstation -Complete and return the enclosed response formFor additional information please contact the Customer Technical Support at 1-800-854-3633, or your local Beckman Coulter representative. |  | 5 | 5 units | Worldwide Distribution: (USA) Nationwide including KY; and the country of Canada. | Terminated | May 07, 2012 | 1096 | 5
Z-2783-2011 | 59115 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Bard PowerLine 6 Fr. Dual-Lumen Polyurethane Catheter with SureCuff Ingrowth Cuff, REF 0700610, Sterile, Contents...1 Each - PTFE Introducer and Dilator, 6.5 FR. x 10 cm..., Assembled in Mexico. |  |  | June 20, 2011 | 2011 | Bard Access Systems | Kit label indicates the wrong dimension for an enclosed introducer. | Labeling design | Bard Access Systems, Inc. notified customers by letter dated June 16, 2011. The letter notified customers of the mislabeling issue and offered a replacement product if desired. For additional information contact Bard Access Systems at 1-800-443-3393. |  | 6485 | 6,485 units | Worldwide distribution - USA (nationwide) including the states of: HI, KY, TX, VT and the country of Canada. | Terminated | March 01, 2012 | 255 | 6485
Z-2775-2011 | 58720 | No | NDR | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter - Manual Inhibin A ELISA Reagent Kit; Part Number DSL-10-28100t-4 |  |  | March 31, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because patient results are falsely elevated by approximately 30% because the reference standard of the Manual Inhibin A ELISA assay has shifted from the assays original calibration.Erroneously elevated results may lead to unnecessary ultrasound testing to confirmgestational age, patient psychological stress until diagnostic testing is completed, and/ormedical decis | Process design | Beckman Coulter sent an "Urgent Product Correction" letter dated March 31, 2011 to all affected customers.The letter included an explanation of the problem and recommendations including; discontinue use of the kit and a Response Form to fill out and return.For additional information contact the Beckman Coulter Customer Support Center at 1-800-854-3633 or your local Beckman Coulterrepresentative. |  | 5521 | 5521 units total for all products | Worldwide distribution: (USA) Nationwide and the countries of Armenia, Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Jordan, Republic of Korea, Lebanon, Malaysia, Mexico, Morocco, Netherlands, Peru, Portugal, Romania, Russian Federation, Serbia, Slovenia, Spain, Thailand, Tunisia, Turkey, United Arab Emirates, and United Kingdom. | Terminated | July 11, 2012 | 468 | 5521
Z-2786-2011 | 58885 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Systems Symphony, Sonata and Symphony a TIM System |  |  | March 31, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm has discovered during product monitoring that on a few systems the gradient cable connections did not meet their specifications after installation. | Nonconforming Material/Component | Siemens initated their recall of the device on March 31, 2001, by sending a Customer Safety Advisory Notice/Letter to affected consignees. The letter identified the product, the problem, and the action to be taken by the customer.Customers were notified that detectors will be installed in appropriate locations on the system, which will switch off the system's power electronics in the event of a malfunction in the power supply of the gradient coils, preventing damage to the system and its environment. |  | 440 | 440 | Nationwide Distribution including AL, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, and WI and Puerto Rico. | Terminated | January 08, 2014 | 1014 | 440
Z-2778-2011 | 58970 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac iX, 23EX, CX, DHXC, Novalis Tx, Trilogy and TrueBeam Linear Accelerators Drive Stands |  |  | May 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Some drive stand bolts may fail to meet mechanical specifications. | Nonconforming Material/Component | Varian Medical Systems, Inc sent a "Urgent Medical Device Correction Urgent Field Safety Notice" letter dated May 13, 2011 to all affected customers.The letter included a description of product, problem and recommendations. This includes performing daily QA to detect isocenter shift and the use of independent patient position system. Customers are advised to discontinue its use if a shift is detected and to contact Varian Medical Systems for a replacement stand bolts and adjustment if necessary.For additional information please contact the Varian Oncology Help Desk at 888-VARIAN5 (888- 827-4265) |  | 36 | 36 | Worldwide distribution: (USA) Nationwide including the states of PA. NJ, NY, MA, VA, TX, KS, FL, OH, SC, MN, CA, TN and NV; and the countries of Canada, Colombia, France, Italy, Japan, Netherlands and United Kingdom. | Terminated | July 13, 2012 | 427 | 36
Z-2770-2011 | 59003 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Langston Dual Lumen Catheters, REF 5515, 7F, Sterile EO, Rx ONLY, 1 x Catheter and 5 x Catheter, Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369. |  |  | June 09, 2011 | 2011 | Vascular Solutions, Inc. | Vascular Solutions, Inc investigation of a recent Device Experience Report has made them aware of a potential problem with their 7F Langston dual lumen pigtail catheter (Model 5515) with lot number 548264.Recently it was reported that pouches containing Langston catheters (Model 5515) were missing a seal,posing a risk of product contamination. It is possible that a compromised sterile barrier | Packaging process control | Vascular Solutions, Inc. sent an "URGENT MEDICAL DEVICE FIELD ACTION- LOT SPECIFIC" letter dated June 7, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructed customers to check their inventory for the affected product and provided information on how to examine the product. Customers are to complete and return a Field Action Customer Inventory Form via fax at 763-656-4250 or e-mail at www.vascularsolutions.com. VSI Customer Service will contact customers to set up return shipment of affected product. All affected product will be replaced upon receipt of the shipped product.Customers should contact their local Sales Representative for questions concerning this recall. |  | 23 | 23 (5 pack) boxes | Worldwide Distribution--USA (nationwide) including the states of CA, FL, IL, MA, MI, MO, NM, NC, OK, and WA and the countries of Austria and Canada. | Terminated | April 05, 2012 | 301 | 23
Z-2777-2011 | 59135 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Steerable Coronary Sinus Catheter, Model SCS01, 9 Fr x 20.4", Sterile R, Edwards Lifesciences Port Access Systems. |  |  | June 15, 2011 | 2011 | Edwards Lifesciences, LLC | Possibility of separation of the introducer sheath from the hub on Coronary Sinus Catheters. | Employee error | Edwards Lifesciences, LLC sent an Urgent Product Recall - Action Required letter dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cusotmers were instructed to check their inventory for any affected product and return all unused product from their stock to Edwards at:Edwards Lifesciences12060 Lone Peak PkwyDraper, UT 84020Attention: Suzanne Carpenter, RGA #Customers were instructed to contact Customer Service to obtain a Returned Goods Authorization number and replacement product . Once customers verified their inventory, they were to complete the attached form and fax it to Edwards at (949) 250-3489. For any questions customers were instructed to call Edwards Customer Service at (800) 424-3278. |  | 97 | 97 units | Nationwide Distribution. | Terminated | June 01, 2012 | 352 | 2895
Z-2789-2011 | 59055 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR/PRIMUS Linear Accelerator systems with RTT v2.x AND an ARTISTE Linear Accelerator/RT Therapist v4.1 system with the LANTIS OIS. |  |  | May 25, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Potential safety issue of possible data loss for the transfer of the same patients treatment data between the LANTIS OIS and the RTT v2.x software and the RT Therapist version 4.1 system | Software design | Siemens Customers Service group began providing Update Instructions and installing the software update on affected machines beginning May 25, 2011. A customer information letter was also provided. |  | 82 | 45 customers with affected configuration in US, 37 customers ROW. | Worldwide Distribution - USA including distribution to LA, CA, PA, MA, NE, NJ, DE, WI, NC, NY, WI, MD, FL, KY and to Poland, Netherlands, Ireland, G=Belgium, Germany, New Zealand, P.R. China, Saudi Arabia, Australia, Malaysia, Egypt, Canada, and Spain. | Terminated | March 01, 2012 | 281 | 82
Z-2779-2011 | 59100 | No | KBW | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Labeled in part:"***BIOMET MICROFIXATION***BIOMET MICROFIXATION***1520 TRADEPORT DRIVE***JACKSONVILLE, FL 33218 USA***www.biometfixation.com***2.0 MM X 7.0 MM TITANIUM***#01-6577***LOT XXXXXX***QTY 00001***X-DRIVE FOSSA SCREW 5/PKG***2.0 MM X 7.0 MM***". |  |  | June 25, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling part number 01-6577, 2.0x7 mm Fossa X-drive screws, 5 pack, lot # 233270. This product is being recalled due to the possibility that the pack may have contained another part, 01-6581 2.0x11 mm Fossa x-drive screws, 5 pack. | Mixed-up of materials/components | Biomet Microfixation, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 25, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to complete and return an Inventory Reconciliation form via fax to 904-741-9425. The form requests that customers indicate whether they want to keep the product or return for credit.Customers can contact the Customer Relations Specialist at 1-800-874-7711 or 904-741-4400 ext. 468 for questions regarding this recall. |  | 25 | 25 units | Product was distributed to one distributor in Argentina. | Terminated | November 09, 2011 | 502 | 25
Z-2794-2011 | 59057 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker 4.0 mm. Angled Aggressive 12 Degree Blade, Precision Cutters, Stryker Instruments, Kalamazoo, MI 49001 USA, (269) 323-7700 (800) 253-3210, Sterile, Single Use, Rx ONLY |  |  | June 13, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | Recalled device was packaged with the incorrect angled blade. The labeling indicates a 12 degree angle but the product has a 40 degree angle. | Labeling mix-ups | Stryker Instruments Division of Stryker CorporationOn 6/13/11 sent an 'URGENT MEDICAL DEVICE RECALL NOTIIFICATION" letter to all affected customers that included the device name, lot number, dates of distribution, reason for the recall, risks associated with use of the device, a photo of the device, instructions to quarantine and return the devices, and a form to fax back to the firm to confirm receipt and adherence to the recall letter. |  | 95 | 95 Units | Worldwide Distribution - USA including CA, LA,, MI, NJ, OH, OR, AR, IN, and Internationallly to Australia, Duisburg, Newbury and Hong Kong. | Terminated | August 03, 2011 | 51 | 95
Z-2790-2011 | 59035 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | HUDSON RCI, Humidifier Adaptor, Rx only, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | June 06, 2011 | 2011 | Teleflex Medical | The packaging may not have been properly sealed, and therefore the sterility of the product cannot be guaranteed. | Packaging | Teleflex Medical sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 6, 2011 to all affected customers. The letter described the product, problem, and the action to be taken.The letter instructs customers to discontinue use and quarantine the affected product. Customers must complete and return a Recall Acknowledgement Form via fax to 1-866-804-9881, Attn: Customer Service. A representative will contact customers with instructions on how to return the affected product. Customers with no affected product are also instructed to complete and return the form to verify receipt of the notification.Contact Customer Service at 1-866-246-6990 if you have questions regarding this recall. and a record of returns. A third letter was addressed to the Risk Manager with the above instructions. |  | 218880 | 228, 960 eaches | Worldwide Distribution--USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, IL, IN, KS, KY, MI, MN, MO, MS, NM, NE, NY, OH, OK, PA, TX, VA, WI and WM and the countries of Japan, Philippines, and Thailand. | Terminated | May 01, 2014 | 1060 | 218880
Z-2792-2011 | 58979 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Abbott Vascular COPILOT Kit, COPILOT Bleedback Control Valve |  |  | January 07, 2011 | 2011 | Abbott Vascular | Five lots of Guide Catheters and Accessory products were processed through Ethylene Oxide (ETO) sterilization with additional plastic material covering the pallet, which is not consistent with validated packaging configuration for ETO sterilization. | Process control | Abbott Vascular sent an Urgent Device Recall letter, dated January 7, 2011, noting the products being recalled and the reasons for recall. The firm also provided the following reconciliation instructions to its territory manager:-Help customer identify aflected product.- CaIl Abbott Vascular Customer Service (800) 221-9902 to receive RGA number. Record RGA number below.- Fax this completed form to (951) 914-3826.- Return a copy of this completed form with the returned product |  | 389 | 389 units total | Nationwide distribution, and Ireland. | Terminated | June 26, 2012 | 536 | 389
Z-2795-2011 | 58824 | No | MSB | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | XVIVO Perfusion AB (Part of Vitrolife), THAM Solution, 3.3 mmol/ml, 50 mL, REF 99019, Sterile, Used to increase the pH of the Perfadex used for perfusion of donor lungs. |  |  | May 12, 2011 | 2011 | Vitrolife, Inc. | Precipitation observed in some bottles of Tham buffer used for adjusting pH of Perfadex prior to perfusion of donor lungs. | Process control | XVIVO Perfusion, part of the Vitrolife Group, sent out a "Partial recall of Perfadex with THAM" letter on May 13, 2011 to all affected customers.Customers were instructed to to quarantine all Perfadex with THAM. Replacement THAM was to be shipped by Vitrilife within a few days and the exchanged THAM was to immediately sent to Vitrolife, Englewood, CO.For additional Information please call (303) 762-1933 |  | 1457 | 1457 bottles | Worldwide distribution including USA and Europe. | Terminated | July 15, 2011 | 64 | 1457
Z-2799-2011 | 58768 | No | NSQ |  |  | Dental | Enforcement Discretion | "***REF IWTH56***Certain EP HEALING ABUTMENT 5mm(D) X 5mm(P) X 6mm(H).*** Healing abutment*** STERILE.***Rx Only.*** Made in USA*** Implant Innovations, Inc. Global Headquarters 4555 Riverside Drive, Palm Beach Gardens, FL 33410 USA Tel: (800) 342-5454 www.3i.online.com. Outside US 561-776-6700***Implant Innovations Siberica S.L., WTC Almeda Park, Ed. 1 Planta 1a Pl. De la Pau s/n 08940-Cornella de Llobregat (Barcelona) Spain. Phone: +34 934 705500. Fax +34 933 717 849. Lot Number: 602691.***" |  |  | February 23, 2010 | 2010 | Biomet 3i, LLC | Biomet 3i of Palm Beach Gardens, FL is recalling their Certain EP Healing Abutment (IWTH56). The Healing Abutment dimensions did not match outside label and laser etching on part. The product was actually 8mm but the package indicated 6mm. | Error in labeling | The firm, BIOMET 3i, initially contacted each US customer by telephone using a prepared phone script. Following telephone contact, the firm followed up each call with a faxed "Urgent: Medical Device Recall" letter dated March 1, 2010. International customers were notified first by email then by fax. The letter and email described the product, problem and actions to be taken. The customers were instructed to return the healing abutment(s) for replacement if they have not already opened and used the product; respond with the attached response fax and return the product(s) to the: Regulatory Affairs Department, BIOMET 3i, 4555 Riverside Drive, Palm Beach Gardens, FL 33410, Incident No. IC95863. BIOMET will send a replacement healing abutment(s) once the returned product is received.If you have any questions or concerns, please contact the Recall Coordinator at 561-776-6906. |  | 49 | 49 | Worldwide distribution: USA (nationwide) and countries of: Australia, Europe and Japan. | Terminated | November 09, 2011 | 624 | 49
Z-2798-2011 | 58863 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Terumo(tm) Hypodermic Needle, sterile, individually packaged into a "blister-pack" flexible plastic bottom-web thermo-formed to create a pocket, heat-sealed with paper top-web,1,000 needles per case. |  |  | June 08, 2010 | 2010 | Terumo Medical Corporation | Hypodermic needle packaging defect could compromise sterility. Affected lot numbers are: MC1725, MC2025, MC2025 | Packaging | The firm, Terumo Medical Corporation sent an "URGENT PRODUCT BULLETIN: RECALL OF TERUMO PRODUCT" letter dated June 8, 2010. Consignees were contacted via telephone on June 7 thru 9, 2010 to confirm receipt of the Product Recall Bulletin. The letter describes the product, problem and the actions to be taken. The letter instructed customers to immediately discontinue use and quarantine any available stock of the listed product codes. The letter stated that distributors must not distribute any of the listed product codes, quarantine any available stock and send the recall bulletin immediately to each customer that received the recalled products to inform them of this issue. Additionally, Terumo Medical Corporation sent a "Terumo Customer Instructions for Returning Recalled Product and Customer Confirmation of Product Recall Notification" letters dated June 8, 2010 to all affected facilities. Customers were instructed to fill-out the requested information and return to Terumo Medical Corporation (TMC) Product Recall Coordinator. Consignees were contacted via telephone on June 8 and 9, 2010 to confirm receipt of the recall information and compliance with the requested action. For questions or need for further information, call (800) 283-7866 and ask to speak with the Product Recall Coordinator. |  | 453000 | 453,000 needles | Nationwide Distribution (USA) -- AZ, CA, CO, FL, GA, IL, KS, KY, MA, MO, NC, NJ, NY, OH, PA, TN, TX, WA, UT, and VA | Terminated | July 14, 2011 | 401 | 453000
Z-2796-2011 | 58806 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague 3 and 3 CX Infusion Pumps, Triple Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2 M8153 and 2M8163. |  |  | May 04, 2011 | 2011 | Baxter Healthcare Corp. | The Air In Line (AIL) test was not properly performed by the service technician at the Service Center. | Employee error | Baxter telephoned all seven affected customers beginning May 4, 2011, informing them that that a post repair test was not performed correctly on their Colleague pumps during the last service. As a result, the device Air Sensor may not function as designed and could lead to undetected air in the IV line. Baxter wants to ensure that the devices are properly tested, safe and effective. Baxter requested that they locate the pumps and remove them from use as soon as possible. Baxter will coordinate with the affected customers, the return of the device for the necessary testing to be completed at their facility and return of the device back to the customer at no cost.Please call ( 800 ) 437-5176 for a complaint that includes any indication of the failure of the product to meet a customer's expectations for quality or to meet performance specifications. |  | 12 | 12 pumps | Nationwide Distribution including the states of NY, Oh, KY,SC, PA, FL and Puerto Rico. | Terminated | January 10, 2017 | 2078 | 12
Z-2797-2011 | 58916 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | Sysmex Stromatolyser-IM; a lytic reagent for use with the XE-2100, XE-2100L and XE-5000 Series Sysmex Automated Hematology Analyzers; the reagent solution is supplied in a 10-liter cubetainer (cardboard cube with a 10-liter plastic liner); Sysmex America, Inc., Mundelein, IL 60060; catalog number SIM-220A. |  |  | June 23, 2009 | 2009 | Sysmex America, Inc. | The affected lots of Stromatolyser-IM lysing reagent was found to be contaminated with Pseudomonas bacteria, which resulted in rising IMI channel background counts, increased false "Platelet Clump" flags and a strong sulfur-like odor when used on the XE-2100, XE-2100L and XE-5000 analyzers. | Environmental control | Sysmex America, Inc. sent a "MEDICAL DEVICE RECALL" letter dated June 2009 to all affected customers. The letter identifies the product, problem, and actions to be taken. The letter instructs customers to discontinue use of the affected SIM lots, and to remove affected lots from their analyzers as soon as possible to avoid long-term performance problems on the analyzer. A Stromatolyser -IM (220-A) Voluntary Recall form was attached for customers to complete and return via fax at 800-218-2711. The form will be used to place orders to replace the affected product. Any questions concerning the information contained in the recall letter are directed to the Sysmex Technical Assistance Center at 1-888-879-7639.The firm sent a second notice to non-responding accounts on September 14, 2009. A copy of the June 26, 2009 recall letter and the fax back form were included with the letter. The customers were informed that Sysmex had not received their completed fax back form and were requested to respond to confirm the notification was received. |  | 4005 | 4,005 cubetainers | Worldwide Distribution--USA (nationwide) and the countries of Brazil, Canada and Honduras. | Terminated | July 13, 2011 | 750 | 4005
Z-2801-2011 | 59207 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | iSymm Intraocular Lens (Model FC-60AD) |  |  | June 17, 2011 | 2011 | Hoya Surgical Optics, Inc. | The recall was initiated because Hoya Surgical Optics has confirmed that the iSymm Intraocular Lens (Model FC-60AD) and iSert Intraocular Lens (Model PC-60AD) have been sold and marketed with out an FDA approval. | PMA | Hoya Surgical Optics Inc. sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated June 22, 2011 to all affected customers. The letter identifies the product, problem, and the actions to be taken.The letter instructs customers to quarantine affected product and return to HSO at their earliest convenience. Shipping instructions are included in the letter. Customers are to complete an Effectiveness Check Questionnaire and fax it to 1-909-680-3986.Questions regarding the recall process are directed to Hoya Customer Service at 1-866-750-5870 or e-mail at customercare@hoyasurgopt.com. |  | 61617 | 61,617 units total | Nationwide Distribution | Terminated | November 29, 2012 | 531 | 61617
Z-2800-2011 | 59018 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Model 134, WarmAir Convective Warming Device, 115V, Part #86180, Device Listing #D111865, The WarmAir 134 is used in conjunction with a single-patient use blanket to warm a patient's body temperature. The device draws in room air through a filter by means of an electric blower. It heats the air to the desired temperature by controlling the air temperature entering a flexible connecting hose, and conveys it to the disposable blanket (i.e. FilteredFlow Blankets, Warming Tube) in contact with the patient, in order to warm the patient. |  |  | May 19, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | A component defect. The recalling firm determined that one of the connectors attached to the firm's WarmAir 134 device's Printed Circuit Board (PCB) can overheat. Should this event occur, the excess heat build up within the device may result in damage to the device's external case and/or melting of some of the internal components of the unit. | Nonconforming Material/Component | CSZ Cincinnati Sub-Zero sent an "URGENT MEDICAL DEVICE RECALL FIELD ACTION" notice to all affected customers. The letter identifies the product, problem, and actions to be taken.CSZ instructs customers to complete and return an enclosed recall response form and to state whether or not their facility is in possession of the device. The completed form should be returned via fax at 513-772-9119 or via e-mail to WA134Recall2011@cszinc.com. Additionally, customers are instructed to discontinue use of the device and return to their Biomedical Engineering and/or Service Department for device exchange and return.Questions regarding this notice are directed to the WarmAir 134 field action (recall) administrator at 1-800-989-7373. |  | 463 | 463 units were distributed | Worldwide Distribution--USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, LA, MA, MI, MO, NC, NJ, NM, NV, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI and the country of Canada. | Terminated | February 26, 2013 | 649 | 463
Z-2725-2011 | 56016 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IQecg IQdevices SDK, Part Numbers: 4-100-1403, Software Version 8.3.2 |  |  | May 18, 2010 | 2010 | Brentwood Medical Technology Corp | The recall was initiated because when a user of the IQmanager, IQdevices SDK, and IQiC (Centricity) software versions listed above accesses the QT dialogue box to either accept or edit the QT value, the QTc value is recalculated using a Hodges QTc equation, regardless of patient's age. The software does this without notifying the user of the change from use of the Bazett equation, which is utiliz | Software design | The firm, Midmark, sent an "URGENT: MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter on May 19, 2010 with an attached Corrective Return Material Authorization (RMA) form via US Postal Service Certified Mail to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their IQecg, IQmanager, or lQiC Centricity software to determine the software version currently in use. If it is a listed product, it is subject to this field corrective action. A sample software label was attached to aid the consignees/customers in locating and identifying the software CD. For customers who have already installed the software, the version number can be identified in the "About" page.Consignees/customers were instructed to complete the enclosed Field RMA form and return it with the affected software to:Quality Assurance DepartmentMidmark Diagnostics Group1125 West 190 S t.Gardena, CA 90248In addition, if they further distributed this product, identify customers and notify them at once of this product field corrective action.Replacement software was to be shipped to the consignees/customers upon receipt of the returned software CD. If they require a UPS return shipping label or an on-line download of replacement software, contact Midmark Diagnostics Group Support Services at 800-624-8950 extension 5145 or 5108. If you have any questions, call Midmark Diagnostics Group Support Services at 800-624-8950 extension 5145 or QA/RA Manager, Midmark Diagnostics Group 310-516-5116. |  | 8 | 8 units | Worldwide distribution: USA (Nationwide) and countries of: Algeria, Australia, Canada, France and Saudi Arabia. | Terminated | August 22, 2011 | 461 | 1621
Z-2833-2011 | 59096 | No | JCT | Tracheal prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific, WALLGRAFTÂżÂż OVER-THE-WIRE, Tracheobronchial Endoprosthesis, Made in IRELAND, Ballybrit Business Park, Galway, IRELAND, STERILE, EO. |  |  | May 25, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is initiating an important field correction related to WALLSTENT", WALLSTENT" RP, and WALLGRAFTÂż Tracheobronchial Endoprosthesis products. This letter is intended to clarify and re-emphasize that U.S. WALLSTENT", WALLSTENT" RP, and WALLGRAFTÂż Tracheobronchial Endoprosthesis stents are NOT currently cleared by the FDA for use in benign tracheobronchial strictures. | Labeling False and Misleading | The firm, Boston Scientific, sent an "Urgent Medical Device Field Correction" letter dated May 25, 2011 to its customers. The letter described the product, the problem and actions to be taken. The letter states that no product is being recalled and asks consignees not to return the product. The customers were instructed to read the enclosed information and post it in a way to ensure it is easily accessible to all users of this device; complete and return the Reply Verification Tracking Form via fax to: CV Field Action Center, 1-866-213-1806, and if they are a distributors to forward this notification to their customers. If you have any questions or would like assistance with this field correction, please contact your local Sales Representative and/or call 763-494-1133. |  | 23218 | 23,218 FOR ALL PRODUCTS | Worldwide distribution. | Terminated | August 10, 2012 | 443 | 23218
Z-2835-2011 | 59078 | [6 empty] | OneTouch Delica Lancing Device; |  |  | May 27, 2011 | 2011 | Lifescan Inc | Individuals attempting to remove a lancet from the OneTouch Delica lancing device without following instructions for use may experience an accidental needle stick. | Other | Lifescan sent an Urgent Medical Device Correction letter dated May 27, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were told that no action on their part is needed and were asked not to return any of the affected product. Customers were instructed to continue to distribute the product. Healthcare professionals and pharmacies were asked to share the information with their patients who use the OneTouch Delica Lancing Device. Patients with questions should contact Customer Care Line at 1-800-574-6138. For questions regarding this recall call 1-800-600-7226. |  | 1032515 | 1,032,515 in US | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | January 05, 2012 | 223 | 1032515
Z-2818-2011 | 59104 | No | GAP | Natural nonabsorbable silk surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Syneture Sofsilk 3/0 BK 12X18" PRE-CUT WAX-RACE |  |  | June 03, 2011 | 2011 | Covidien LP | Potential for the sterile barrier to be breached compromising the sterility of the product | Equipment maintenance | Covidien sent an URGENT MEDICAL DEVICE RECALL letter dated June 3, 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately identify and quarantine all affected inventory. Distributors who further distributed the affected product were instructed to forward the letter to their customers who received the affected product. Customers were instructed to contact Covidien Customer Service at SDFeedback@Covidien.com or (800) 722- 8772, option 1, to obtain a Return Material Authorization (RMA) prior to returning the affected product. Customers were also asked to complete the attached product recall form in its entirety and fax the form to 203-822-6009 and include a copy with products being returned. If customers did not have any of the affected product they were to simply return the recall form via fax or email indicating they had zero units. For any questions regarding this recall customers were to contact their Covidien Representative or Covidien Quality Assurance at 203-492-5232. |  | 960 | 960 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Denmark, France, Germany, Italy, Japan, Korea, and Singapore | Terminated | February 10, 2014 | 983 | 93600
Z-2834-2011 | 59099 | No | EAZ | Dental operating light. | Dental | Dental | 510(K) Exempt | LuxStar Dental Light have the following Agency model or Catalog Model numbers as ML1000, OL1000, CL1000, WL1000, CB1000, TL1000, and UL1000. |  |  | May 27, 2011 | 2011 | Dental Equipment LLC | Lens heat shield (birdcage) of the Luxstar Dental Light could fall and injure a patient. | Device Design | Dental Equipment, LLC sent an "URGENT MEDICAL DEVICE RECALL" notification letter via USPS 1st class to US consignees and via Federal Express or Fax to foreign consignees on June 24, 2011. The letter informed customers about the product, problem and tether upgrade corrective action. A Field Correction Acknowledgement Form and the Field Correction Completion Form was attached for customers to complete and return via the fax number 503-537-3617. The letter states that customers are to review their records in order to identify where the affected devices are located and which are still in service. Customers can call the Customer Service at 800-304-5332 and order the self-install Upgrade Kits with the Part Number 30792 at no charge. Customers are instructed to complete the Rework Correction Form once all of the lights have been updated and fax or mail to Marcus, attention Bonnie Dundas, at (503) 537-3617. |  | 13226 | 13,226 units distributed worldwide | Worldwide Distribution -- USA (nationwide) including states of: CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, and Puerto Rico and the countries of: Australia, Belgium, Canada, France, Germany, Ireland, Italy, Mexico, Norway, Poland, Russia, United Kingdom, Vietnam, and Trinidad. | Terminated | October 24, 2012 | 516 | 13226
Z-2831-2011 | 59184 | sheath, corrugated rubber, for non-indwelling catheter, non-sterile | NNY | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Rusch, Sure Seal Golden Drain, One Piece Urinary Incontinence Device, Teleflex Medical. |  |  | June 27, 2011 | 2011 | Teleflex Medical | The Kit contains Skin-Prep protective wipes that were manufactured by the Triad Group and are being recalled by Smith &amp; Nephew, due to possible bacterial contamination. | Other | Teleflex Medical sent an "Urgent Medical Device Recall Notification" dated June 27, 2011 to all affected customers.The firm provided recommendations including all affected product be discontinued and quarantined. A Recall Acknowledgement Form was provided to be completed and returned. If the product was further distributed, they were urged to contact their customers and communicate the recall.For additional information please contact Customer Service at 866-246-6990 or your local sales representative. |  | 6880 | 6880 units (total for both products) | Nationwide Distribution (USA) including the states of FL, GA, OH and SC. | Terminated | March 20, 2012 | 267 | 6880
Z-2837-2011 | 59263 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Non-DEHP I.V. Catheter Extension Set, 6.6 (17 cm), Volume 0.8 mL, Injection Site, Male Luer Lock Adapter; an Rx sterile, single use nonpyrogenic fluid pathway; 200 sets per case; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 2N3375 |  |  | July 06, 2011 | 2011 | Baxter Healthcare Corp. | Some of the individual extension set packages were improperly sealed, compromising the sterility of the product. | Packaging | The firm, Baxter, sent an "Urgent Product Recall" letter dated July 6, 2011 via first class mail to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to locate and remove all affected product from their facility; call Baxter Healthcare's Center for Service at 1-888-229-0001 between 7:00 am and 6:00 pm Central Time, Monday through Friday to return affected product and receive credit; complete and return the attached CUSTOMER REPLY FORM to Baxter via fax at 1-847-270-5457 or scan and email to fca@baxter.com and if you are a dealer, wholesaler or distributor/reseller, notify your customers of this recall. If you have questions regarding the communication, call The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 9800 | 9,800 sets | Worldwide distribution: USA (nationwide) including: AL, LA, MO, SD, TX and WV; and country of: Canada. | Terminated | April 18, 2012 | 287 | 9800
Z-2838-2011 | 58769 | No | NSP |  |  | Dental | Enforcement Discretion | "***3i Incise Custom-Made Label. QP157 Approved 07/09/10*** 3i inc se Custom-made Dental Framework***Biomet 3i 4555 Riverside Drive Palm Beach Gardens, FL 33410 800-342-5454 www.biomet3i.com Please consult instructions for use and recommendation***For Dental Lab use ONLY Non-sterile/For single Use***" |  |  | December 01, 2010 | 2010 | Biomet 3i, LLC | On December 1, 2010 Biomet 3i recalled their product Incise Zirconia Coping (CBZR0102). The Coping contains a manufacturing defect. Coping is oversized. All affected direct accounts were notified by letter on December 14, 2010. | Nonconforming Material/Component | Biomet 3i sent an "MEDICAL DEVICE RECALL" letter dated December 14, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken.The letter instructs customers to complete and fax the Product Recall Response form to 561-514-6316. In addition, a shipping label is enclosed with the letter for returning the affected product.Customers with questions and concerns should contact Biomet 3i recall coordinator at 561-776-6906. |  | 16 | 16 | Worldwide Distribution--USA (nationwide) and the country of Canada. | Terminated | July 18, 2011 | 229 | 16
Z-2840-2011 | 59155 | [6 empty] | ENCORE(R) UNDERGLOVE, Latex, Powder-Free, Smooth Surgical Gloves, REF 2018455, Size 5 1/2, Sterile, Manufactured for Ansell Healthcare, 1635 Industrial Road, Dothan, AL 36303, Made in Sri Lanka |  |  | June 10, 2011 | 2011 | Ansell Healthcare Products LLC | Gloves do not meet the tensile requirements of ASTM D 3577 reducing the shelf life of the product from 5 years to 2 years. | Other | Ansell Healthcare Products LLC sent an Urgent Voluntary Recall letter dated June 9, 2011, to all affected customers on June 10 and 13, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconstinue use and distribution of any of the affected product and return them to the firm. If the customer distributed gloves further, customers were advised to forward a copy of the notification letter along with their tracking number to their customers. Customers were instructed to complete and return the attached Response Form to Ansell by fax at 1-800-722-8155. Returned product would be replaced or credited to the customer. |  | 9580 | 9580 pairs | Nationwide Distribution | Terminated | January 10, 2013 | 580 | 1176176
Z-2839-2011 | 57518 | No | LNR |  |  | Gastroenterology/Urology | PMA | Therakos Cellex Procedural Kits. |  |  | December 02, 2010 | 2010 | Therakos Inc | Cellex Procedural kits may have minor tears in their Tyvek covers. | Employee error | Therakos, Inc. sent an "Important Product Notification" dated December 2, 2010 to all affected customers.The letter included description of problem, affected lot and recommendations. Customers were asked to dispose any product with affected lot and to contact them for for replacement, and to complete and return the attached Fax reply form.For additional information contact Therakos Customer Technical Services at 877-865-6850. |  | 110 | 110 units | Nationwide (USA) distribution including the state of MI, MO, WI, CA, OH, NH, MA, and OR. | Terminated | February 19, 2015 | 1540 | 110
Z-2849-2011 | 58980 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | "***REF: 922070 TMJ SYSTEM ARTHROSIMPLICITY KIT***STERILE PRODUCT***STERILE EO***DISTRIBUTED BY BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***" |  |  | October 08, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling TMJ System OnPoint Scope Procedure Kit REF 24-3050; TMJ System ArthroSimplicity Kit REF 922070 &amp; 24-4050; TMJ System InnerVue Scope Procedure Kit REF 922090. The OnPoint Scope Kits and ArthroSimplicity Kits contain Chloraprep Swabsticks that do not contain any warnings regarding the use of the product on eyes, ears or mouth. | Pending | Biomet Microfixation (formerly W. Lorenz) sent an "URGENT MEDICAL DEVICE SAFETY ALERT" letter to the consignees/customers on October 7, 2010. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to fill out the attached Inventory Reconciliation sheet and fax to 904-741-9425 and if the consignees/customers further distributed this product, they should notify their customers and provide them with a copy of this safety alert (the letter must be given to the personnel responsible for receiving field action notices). If you have any questions, please contact the Customer Relations Specialist at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 82 | 82 | Worldwide distribution: USA (nationwide) including states of: AL, AR, CA, CO, CT, DC, CR FL,GA, HI, IA, IL, IN, KS, KY, LA, MA, MD,ME, MI MN, MO, MS, MT, NC, ND, NJ, NY, OH, O K, OR, PA, RI, SC, TN, TX, UT, VA and WA; and countries of: Argentina, Bahrain, Canada, Colombia, Costa Rica, and Finland. | Terminated | April 19, 2012 | 559 | 828
Z-2853-2011 | 59270 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Code: 2M8161. |  |  | July 05, 2011 | 2011 | Baxter Healthcare Corp. | While replacing a defective pump head mechanism and performing calibration, Baxter's Service center may have failed to apply GLYPTOL (a red insulating varnish) on a potentiometer that helps maintain pump accuracy for detecting an air bolus. | Process control | Baxter Healthcare Corporation telephoned its customers on July 1, 2011 and sent a follow up "URGENT PRODUCT RECALL" letter dated July 6, 2011, informing them that the Glyptol (a red insulating varnish) may not have been applied on the Air-In-Line (AIL) circuit board. The letter described the product, problem and actions to be taken. The customers were instructed to locate the affected pumps, remove from service and return to Baxter for inspection as soon as possible. If GLYPTOL is found to be missing, recalibration will be performed and GLYPTOL will be applied per requirement before returning the pumps back to the customer at no cost. The customer was instructed to complete and sign the attached customer reply form. The letter states that if the COLLEAGUE Pumps were distributed to other facilities, please forward communication to the pump owner as appropriate. For any questions contact The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central Time, Monday through Friday. |  | 2 | 2 units | Nationwide (USA) Distribution -- including the state of Pennsylvania. | Terminated | January 20, 2012 | 199 | 2
Z-2858-2011 | 59051 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Runthrough NS Tapered PTCA Guidewire, item code number TW*DP418FH |  |  | June 03, 2011 | 2011 | Terumo Medical Corporation | Baloon catheter guidewires may become uncoiled during angioplasty procedures. | Device Design | Terumo Medical Corporation initiated contacting consignees on June 3, 2011 via an in-person visit to hand deliver an "URGENT PRODUCT BULLETIN: TERUMO PRODUCT SAFETY ALERT" letter dated June 3, 2011. The letter emphasized to consignees the importance of following the current instructions and recommendations provided to minimize the potential for damage during use and removal of the recalled product. For further questions, please contact your Terumo Territory Manager or call (800) 283-7866 and ask to speak with the Runthrough NS Guidewire Product Manager, at extension 4129. |  | 70 | 70 units | Nationwide (USA) Distribution -- including the states of: CT, GA, and SC. | Terminated | July 27, 2011 | 54 | 70
Z-2803-2011 | 58996 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GEM PREMIER 4000 CARTRIDGE K+ REPORTING: |  |  | May 31, 2011 | 2011 | Instrumentation Laboratory Co. | K+ results on the GEM Premier 4000 are too low when compared to a reference analyzer, withbiases exceeding allowable error claim of aÂż 0.5 moUL | Nonconforming Material/Component | The firm, Instrumentation Labs, issued a Revised "Urgent Field Safety Notification" letter dated July 14, 2011, to GEM Premier 4000 customers. The letter describes the product, problem, updated issue description and actions to be taken. The customers were instructed to follow the instructions to Disable K+ Parameter ,provided in the letter, to eliminate the potential of erroneous K+ result reporting and to complete and return the MANDATORY RESPONSE TRACKING FORM immediately via fax at 781-861-4207 or email to ra-usa@ilwww.com. The letter also noted an Important Change to Customer Actions from June Urgent Field Safety Notification.Prior strategy :Instrumentation Laboratories (IL) notified U.S and Canadian customers via the Urgent Field Safety Notification mailed USPS between June 2 - 3,2011. International Affiliates / Dealers notified through e-mail on June 1, 2010. A response (tracking) form was included. A Worldwide Regulatory Notice was e-mailed on May 31.2011.to IL country affiliates and distributors, notifying them of the issue, providing a risk summary and the support documentation for the notification. The customer notification will be handled per their local regulatory requirements.If you have questions regarding this notification, please contact IL at 800-678-0710, press 2 at prompt. |  | 2369 | 2369 | Worldwide distribution: USA (nationwide) and countries of: CANADA, ARGENTINA, AUSTRALIA ,AUSTRIA, BELGIUM, BRAZIL, CHILE, CHINA, CROATIA (local name: Hrvatska), CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, MALAYSIA, MEXICO, NETHERLANDS, NORWAY, PANAMA, POLAND, PORTUGAL, PUERTO RICO, SAUDI ARABIA, SLOVENIA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VIET NAM. | Terminated | November 27, 2012 | 546 | 28616
Z-2859-2011 | 59103 | No | IZQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiomat Illumena Syringe, 150mL Linden Luer Syringe with Handi-Fil Straw (p/n 900103), in a sealed molded plastic tray. |  |  | May 20, 2011 | 2011 | Mallinckrodt Inc | Syringes packaged and labeled as the Illumena 150mL Linden Luer Syringe with Handi-Fil Straw, product 900103, contained the Illumena150mL Syringe with Handi-Fil Straw, product 900101. | Other | COVIDIEN sent an Urgent Device Recall letter dated May 20, 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use the attached business response form and report the quantity of lot number 0253304 currently in their stock. If the customer did not have any of the affected lot they were to indicate by checking the appropriate box on the form. When completed, customers were asked to fax to (314) 654-8206 or email to productmonitoringhazelwood@covidien.com. All product from lot number 0253304 was to be returned to Covidien. Once the business notification form is received a Return Goods Authorization for the quantity of product reported will be sent to the customer. Customers would receive credit for all return product from affected lot.For any questions regarding this recall customers were asked to contact the Product Monitoring department at 800-778-7898. For questions about credit for the product, call Customer Service at 888-744-1414, menu option 1, and then menu option 2. |  | 550000 | 11,000 units (200 cases/50 units per case) | Worldwide Distribution - USA including CA, FL, IL, IN, MI, MO, NY, OH, TN &amp; TX. and the country of Australia. | Terminated | February 19, 2013 | 641 | 550000
Z-2836-2011 | 59161 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Kimberly-Clark, Three Quarter Drape, Non sterile/Disposable, |  |  | June 06, 2011 | 2011 | Kimberly-Clark Corporation | Metal shavings were found within the folded portion of the surgical drape. | Process control | Kimberly-Clark Corporation sent an "Urgent: Voluntary Product Recall" letter dated June 6, 2011 to all affected customers and a subsequent letter dated June 22, 2011.The letter notified them of the recall and provided recommendations. They were instructed to evaluate their facilities unused inventory of the affected drapes to determine if any of the impacted product lots remain with their facility. If present, they were to quarantine and discontinue use. A Customer Product Recall Response Form was included to be completed and faxed to KC. They will be contacted by a representative regarding the return and replenishment of the product. If the product was further distributed, the consignee was advised to contact the end user, make them aware of the recall and to retrieve any of the affected product in inventory. For additional information please contact Kimberly-Clark Regulatory Affairs at 770-587-7400. |  | 223 | 223 cases (300 units per case) | Nationwide (USA) distribution, including the states of CA, FL, GA, IA, IL, MO, NC, NY, OH, TN, TX, VA and WI. | Terminated | October 14, 2011 | 130 | 223
Z-2855-2011 | 59291 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MatriStemÂż Plastic Surgery Matrix XS |  |  | July 07, 2011 | 2011 | ACell, Inc | Surgical mesh for implantation may have elevated endotoxin levels. | Nonconforming Material/Component | ACELL, Inc. sent an "Urgent: Medical Device Safety Alert" letter dated July 07, 2011 to all affected customers.The letter notified customers (physicians) of the possibility of elevated endotoxin levels in the product and affected lots. They were instructed to monitor their patients receiving the affected implants.For additional information please contact the Director of Regulatory Affairs &amp; Quality Assurance at (410) 953-8508. |  | 30 | 30 units | Nationwide (USA) Distribution including the states of CA, CT, FL, ID, IL, MD, MI, MS, NC, NJ, NV, NY, OH, OR, PA, TX and VA | Terminated | October 21, 2011 | 106 | 30
Z-2856-2011 | 59259 | No | LGB | Gonococcal antibody test (GAT). | Microbiology | Microbiology | PMA | PALL Medical LeukoGuardÂż LGB Arterial Line Leukocyte Reduction Filter with Bypass Loop, STERILE. CATALOG/MODEL/Reorder Number: LGB. CLASSIFICATION NAME: Filter, Blood, Cardiopulmonary Bypass Arterial Line. COMMON/USUAL NAME: Arterial Line Filter with Bypass Loop. -- Manufactured by: PALL NEWQUAY, CORNWALL, TR9 6TT, UK -- Device Listing # D000745. --- Designed to reduce the levels of circulating leukocytes and remove microemboli greater than 40 microns in size from the perfusate during extracorporeal circulation including: gas emboli, fat emboli and aggregates composed of platelets, red blood cells and other debris. |  |  | May 26, 2011 | 2011 | Pall Corporation | Potential for the housing of the arterial line Leukocyte Reduction Filter to separate before or during the priming procedure. | Process control | Pall Medical sent an "Important Customer Notification" letter dated May 26, 2011 to all affected customers.The letter included a description of the problem, affected lots and provided recommendations including the return of the product.For additional information, please contact the firm at (516) 484-0263. |  | 232 | 232 units | Nationwide (USA) Distribution including the states of CA, LA, MI, TX and WI. | Terminated | October 03, 2012 | 496 | 232
Z-2854-2011 | 59189 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | VirtuoSaph Plus Endoscopic Vessel Harvesting System , REF VSP550, Sterile, Terumo Cardiovascular System |  |  | June 10, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems received four (4) complaints from a single customer about tunnel breathing/inadequate insufflation with the VirtuoSaph Plus Endoscopic Vessel Harvesting System during cases in a limited production launch situation. Teumo's investigation identified issues with the devices's capability to maintain pressure within the tunnel. | Device Design | Terumo Cardiovascular Systems contacted the one customer to this device on June 10, 2010 by telephone.There was one consignee. The firm contacted this one consignee by phone on 6/10/2010. There were 2 contacts at this one consignee. The recalling firm contacted both of the contacts on 6/10/2010. The consignee was asked to dispose of remaking inventory of devices. The consignee said had disposed of the product.For information on this recall call (734) 741-6173 |  | 35 | 35 units | Nationwide (USA) Distribution including the state of New York. | Terminated | July 27, 2011 | 47 | 35
Z-2158-2011 | 58419 | No | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | GE Healthcare DISCOVERY MODEL NM/CT 670 |  |  | January 26, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a missing laser aperture warning label (i.e. Laser Aperture - DoNot Stare Into Beam) affecting certain Discovery Systems. | Radiation Control for Health and Safety Act | GE Healthcare will bring affected systems into compliance. A GE Healthcare Service Representative will provide a product label correction to address this issue. This will be carried out via a Field Modification Instruction (FMVI 408311 planned to be released in June 2011 with an estimated completion date of August 2011.For questions call 262-513-4122. |  | 4 | 4 | Nationwide Distribution | Terminated | August 31, 2012 | 583 | 4
Z-2865-2011 | 59201 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Prime 5th Wheel Stryker Medical Stretchers Model 1105 |  |  | July 19, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | The firm has determined that the brake rod assemblies of various stretchers that were built with a 5th wheel were manufactured with roll pins that did not meet the materials specifications. Some roll pins may have fractured during the assembly process and could fall out of the brake rod assemblies, rendering the brake or steer systems inoperable from certain pedals or causing a reduction in overa | Nonconforming Material/Component | Stryker Medical sent an "Urgent Medical Device Correction" dated July 19 ,2011 to all affected customers.The letter included; the name and model numbers of the stretchers, the reason for the correction and dates of manufacture. It contained the instructions to check the braking system and if found to be inoperable, to remove it from service, and if operable to check prior to each use. Also, information regarding contacting the recipients of any loaned or sold stretchers to relay the information in the letter and Medwatch reporting information. A self addressed postcard to confirm receipt of the letter was also included.For additional information call (269) 389-6604. |  | 80 | 80 | Worldwide Distribution: Nationwide (USA) including the states of NY, MI, LA and CA ,and the countries of Canada, China, Spain, Japan, Korea, United Kingdom, South Africa. | Terminated | August 29, 2013 | 772 | 112
Z-2873-2011 | 59002 | [6 empty] | "***REF: 91-6704***THINFLAP SYSTEM CROSS-DRIVE, SELF-DRILLING SCREW 1.5X4MM (.059X.159") (WXL)***TITANIUM***NON-STERILE PRODUCT***BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA*** www.biometmicrofixation.com***QTY 00001***". |  |  | May 17, 2010 | 2010 | Biomet Microfixation, Inc. | On May 17, 2010, Biomet Microfixation issued a Safety Alert on their Thinflap System Cross Drive, Self Drilling Screw, 1.5x4mm, Model #: 91-6704, Lot#: 173770. The product is labeled incorrectly. Some packages of 91-6704 lot 173770 are labeled as cross drive screws, but contain center drive screws. All affected customers were notified of the recall on May 17, 2010. | Labeling mix-ups | Biomet Microfixation notified distributors and customers by letter, email and telephone calls. Distributors and customers were instructed to fill out and return the Inventory Reconciliation sheet, via Fax to 904-741-9425. Distributors were also instructed to notify their customers if the product was distributed further.For questions regarding this recall customers were instructed to call 1-800-874-7711 or 904-741-4400, ext. 468. |  | 927 | 927 | Worldwide Distribution - USA including AL, CA, GA, HI, IL, MO, NC, NY, TX and the country of Japan. | Terminated | October 12, 2011 | 513 | 927
Z-2867-2011 | 59077 | [6 empty] | Philips brand BrightView Gamma Camera System, |  |  | June 03, 2011 | 2011 | Philips Medical Systems | The gap between the two parts of the table may permit pinches to occur on either patients skin or other body parts; it has the potential for serious injury if a patients finger is caught between the pallet and patient table resulting in broken finger bone. | Device Design | Philips Healthcare issued a Customer Information letter (Field Change Order (#88200402)) dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the instruction for use for their system, which states "Monitor the patient and the system vigilantly during any system motions to make sure that ...accessories, equipment, and all ojbects such as hair, jewelry, or neckties-remain clear of any moving camera and imaging table parts. For additional information, customers were istructed to see the BrightView SPECT instructions for Use, section 2.2 - Safety Warnings and Precautions and Safety and Information Labels. Philips Healthcare will notify all affected customers and will install a sticker through an Action for Performance - Proactive Field Change Order (FCO) to resolve the issue. For customers in North America who require further information or support contact Customer Care Solutions Center at 1-800-722-9377, Select Option 5. In all other countries the local Philips Healthcare office should be contacted. |  |  |  | Worldwide Distribution - USA (nationwide) and Puerto Rico and the countries of Australia, Austria, Bahrain, Belgium, Brazil, China, Chile, Columbia, Czech Republic, Denmark, Dominican Republic, Egypt, France, Greece, India, Iran, Israel, Italy, Japan, Lithuania, Malaysia, Mexico, Netherlands, Peru, Portugal, Russia, Saudi Arabia, Singapore, South Korea, Spain, Switxerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, and the United Kingdom. | Terminated | November 06, 2012 | 522
Z-2866-2011 | 59167 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Avanta Fluid Management Injection System, Catalog numbers: AVA 500 TABL, AVA 500 PEDL |  |  | June 02, 2011 | 2011 | Medrad Inc | Event was initially intended to be a product enhancement of the MEDRAD Avanta Fluid Management Injection System. Upgrades were initiated in October, 2007. Observance of a possible increase in complaints and adverse event reporting resulted in the firm issuing "Important Product Use Reminder" to remind users of proper procedures for using the system and the need for diligence in reducing air embo | Device Design | Medrad Inc., sent two types of "MEDICAL DEVICE FIELD CORRECTION" notifications on June 2, 2011 to all affected customers. 1. To those customers who Avanta Systems have already undergone market withdrawal and thus have been upgraded, the notification provided customer awareness that prior action should be considered a field correction. The letter stated that NO action is required since their system(s) has already been upgraded. If you have any questions, concerning these activities contact the Director of Quality Systems at (724) 940-6905. 2. To those customers whose Avanta Systems have not yet been upgraded through market withdrawal, the notification informed customers that the recalling firm is undertaking a Field Correction to upgrade their system(s). The letter included a step by step guide that will allow customers to verify that their system(s) has NOT been upgraded. Customers are instructed to complete and fax this form to MEDRAD at (412) 406-0941. If you have any questions, concerning these activities contact the Product Manager at (412) 767-2400, extension 8401. |  | 842 | 842 | Worldwide Distribution -- (USA) (nationwide) including the state of Puerto Ricoand the countries of: AE, AL, AT, BE, BG, BR, CA, CH, CO, CY, DE, DK, DO, ES, FI, FR, GB, HK, HR, HU, IE, IL, IS, IT, JO, JP, KR, KW, LY, MA, MT, MX, MY, NL, PA, PH, PK, PT, QA, RS, RU, SA, SE, SG, SI, TR, VI, and ZA. | Terminated | June 05, 2012 | 369 | 842
Z-2878-2011 | 59021 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ATLANTIS (R) Self-Tapping Variable Angle Screw, Rx only, Medtronic Sofamor Danek USA, 1800 Pyramid Place, Memphis, TN 38132 |  |  | June 28, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | Screws are longer then size etched on the screw. | Other | Medtronic sent a Voluntary Recall letter dated June 28, 2011, directly to Risk Managers and Sales Representatives, with Recall Questionnaires and Confirmation forms included. The letter identified the product, the problem, and the action to be taken by the affected customers. If appropriate, personal follow up with Sales Representatives will be made to assure that the recall communication is understood by Risk Managers and to facilitate return of product and the associated Recall response documentation. Foreign consignees will be notified by email on June 29, 2011.For any questions or concerns regarding this recall customers were to contact their Medtronic Sales Representative or the Global Quality Department at 800-876-3133, ext 3173. |  | 101 | 101 units | Worldwide Distribution - USA (nationwide) and the countries of MN, AK, FL, PA, ME, GA HI, CA, TX, OH, IN, MO, WI, MI, NY, IN, LA, MD, MA, NC, OR, NE WA, ID, CO, OR, UT, IL, WV, Canada, Singapore, Netherlands, Brazil, Columbia, Hong Kong, and Mexico. | Terminated | December 24, 2012 | 545 | 460
Z-2880-2011 | 59178 | No | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | eSwallow Multi-Program Dysphagia Therapy Test unit |  |  | June 22, 2011 | 2011 | eSwallow USA LLC | It was discovered that the firm was distributing medical devices without FDA approval. | Other | eSwallow USA LLC sent an URGENT MEDICAL DEVICE RECALL letter dated June 22 2011, via US mail and email notification to each distributors. The letter identified the product, the problem, and the action to be taken. Distributors were instructed to immediately examine inventory, quarantine all product subject to recall, and identify customers that received these products and notify them at once of the recall. The letter advised distributors to forward a copy of the recall notification letter to their customers. The letter states that the distributors are to return all inventories to the eSwallow Corporate Office eSWALLOWusa, 106 Hidden Drive, Scottsboro, Al 35769. The recall should be carried out to the customer/user level. Customers are instructed to complete and return the attached Response Form as soon as possible. If you have any questions, call (256) 571-0443. |  | 123 | 123 units | Nationwide (USA) Distribution - including states of AL, AZ, CA, GA, FL, LA, MS, and VA. | Terminated | July 03, 2012 | 377 | 123
Z-2879-2011 | 59146 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Arthrex Reciprocating Saw Attachment |  |  | May 24, 2011 | 2011 | W &amp; H DentalWerk Buermoos GMBH | Reciprocating saw attachment V 600 Sr (REF: 17306210), screws can become loose or come off (the attachment) and fall into the surgical wound and may not be noticed due to its small size. | Equipment maintenance | Madrimed GmbH (Distributor partner) of W&H; Dentalmechanik GmbH sent a letter dated May 24, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were ask to advise their users to check the adapters thoroughly before and after each application, until the conversion of all attachments have been completed and to send all products that need to be converted to the nearest service centers.W&H; decided to convert all products that might be affected in their service centres and equip them with an additional security device.If you have any questions, call +43627462360 or email: regulatory@wh.com. |  | 194 | 194 units | Worldwide distribution: USA (nationwide) and countries of: Austia, Australia, Canada, Czech Republic, Denmark, France, Germany, Italy, Kuwait, Malaysia, Netherlands, Russia, Slovakia, South Africa, Spain, Taiwan, Turkey and United Kingdom. | Terminated | October 09, 2015 | 1599 | 194
Z-2882-2011 | 57968 | No | JKC | Dehydroepiandrosterone (free and sulfate) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DHEA RIA Reagent |  |  | September 09, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports that DHEA measurements using DHEA RIA (REF DSL-8900) kit lots identified may overstate actual DHEA levels in patient samples. | Process control | Beckman Coulter sent an "URGENT: PRODUCT CORRECTION ACTION" letter dated November 15, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to discontinue using the affected products and discard them according to local regulations. Customers were asked to review patient DHEA results ran on the affected lots and interpret them in light of additional tests (such as LH, FSH, testosterone, DHEA-S, cortisol, 17-hydroxyprogesterone, and estradiol) and the patient's clinical status. A response form was enclosed for customers to complete and return.Contact the Beckman Coulter Technical Support Center at 1-800-854-3633 in the United States, or their local Beckman Coulter Representative for questions regarding this notification. |  | 3542 | 3,542 units (1181 in US) | Worldwide Distribution--USA (nationwide) and the countries of Argentina, Australia, Belgium, Brazil, Canada, Chile, China, Czech Republic, France, Germany, Hungary, India, Ireland, Italy, Japan, Mexico, Netherlands, Poland, Spain, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | May 04, 2012 | 603 | 3542
Z-2885-2011 | 59109 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 500 Automated External Defibrillator is a semi-automatic defibrillator that uses a patented Shock Advisory System. |  |  | June 06, 2011 | 2011 | Physio Control, Inc. | The potential reliability issue is the result of a design characteristic of a SRAM (Static Random AccessMemory) electronic component on the main printed circuit board assembly. | Component design/selection | Physio-Control, Inc. sent an "URGENT- ACTION REQUIRED" letter dated June 2011 to all affected customers. The letter describes the product, problem, and the action to be taken by the customer.The letter recommends that users who store their device(s) in hard-shelled carrying cases remove them from the case(s)immediately and discontinue the use of the case(s). Additionally, the letter recommends that customers perform regular maintenance in accordance with the Operating Instructions. A Confirmation Sheet was enclosed with the letter for customers to complete and return via fax at 1-866-448-9567. Customers can call the Technical Support Team at 1-800-442-1142, Option 5. |  | 12197 | 12197 units distributed worldwide (5169 units outside US; 7028 units in the US) | Worldwide Distribution--USA (nationwide) and the countries of Argentina, Australia, Bahamas, Brazil, Canada, Chile, Columbia, Germany, Hong Kong, Iran, Japan, Mexico, Netherlands, Peru, and Uruguay. | Terminated | February 23, 2012 | 262 | 12197
Z-2884-2011 | 58921 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | LPCCC2: Low Profile Abutment Non-Hexed Castable Cylinder by Biomet 3i |  |  | March 14, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i, Palm Beach Gardens, FL is recalling the Model LPCCC2 Low Profile Abutment, Non-Hexed Castable Cylinder, Lots 2010101036 and 2010101037  Exp. 10/2015. Package labeling may not correctly describe the package contents. Product was distributed in Europe only. | Labeling mix-ups | Biomet 3i sent a "Medical Device Advisory Notice" letter dated February 17, 2011 to all affected customers.The letter described the problem, instructed customers to returned any unused product and to confirm receipt of the notification.For additional information, please contact Customer Service at 800-342-5454. |  | 232 | 232 unit | Worldwide distribution: http://webapps.cdrh.fda.gov/division-tracking/findTrackable.do?docNum=RCL110656&reviewType;=RCL | Terminated | July 27, 2011 | 135 | 232
Z-2881-2011 | 59101 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Calloway LabsMulti-Drug 9 Panel Assay |  |  | June 10, 2011 | 2011 | Alere San Diego | Alere is initiating a voluntary recall on the listed lots of MULTI-DRUG 9 PANEL INTEGRATED CUP manufactured by Alere for sole use by Calloway Labs due to a mistake that occurred while processing the order, these products were incorrectlylabeled as CLIA Waived, when in fact the devices are classified as Moderately Complex. Therefore these devices are misbranded. | Vendor change control | Alere San Diego sent an "Urgent Medical Device Recall" letter dated June 9, 2011 to all affected customers.The letter included: products affected, reason for recall and instructed them to return all affected product. An acknowledgement form was also included for the customer to return to Alere.For questions contact Technical Services at 888.246.7483, option 2 or by e-mail at InnovaconComplaints@alere.com. |  | 2400 | 2400 kits (60,000 devices) | (USA) Nationwide Distribution including the states of MO and MA. | Terminated | April 04, 2012 | 299 | 2400
Z-2887-2011 | 59016 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent, Part Number: 447070, Lot Number: M908398 &amp; M911529. |  |  | May 31, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the listed lots of Rheumatoid Factor (RF) Reagent might produce false positive results. A false positive result might lead to additional unnecessary testing. | Pending | Beckman Coulter sent an "URGENT FIELD SAFETY NOTICE/PRODUCT CORRECTION" letter dated May 31, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue use of the affected products. Customers were asked to complete and return an attached FAX BACK Response Form to 714-961-4234. Replacement products may be obtained by calling Order Assistance at 800-526-3821 in the United States, (800) 463-7828 in Canada or contacting your local Beckman Coulter representative.For any questions regarding this notice, contact the Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/CustomerSupport/Support or call (800) 854-3633 in the United States and Canada; or contact your local sales representative. |  | 6302 | 6302 units (850 in the US) | Worldwide Distribution--USA (nationwide) and the countries of Australia, Belgium, Brazil, Canada, Chile, China, Czech Republic, France, Germany, Greece, Hong Kong, India, Ireland, Italy, Japan, Kuwait, Macao, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Poland, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, and United Kingdom. | Terminated | June 03, 2013 | 734 | 6302
Z-2857-2011 | 59288 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Global Focus Marketing and Distribution, Silencer S2200 Centrifuge |  |  | July 15, 2011 | 2011 | Global Focus Marketing &amp; Develop | Potential for bucket and insert not seated on the rotor to cause the lid to break and buckets to be ejected. | Nonconforming Material/Component | Global Focus Marketing and Distribution, Ltd issued a press release on 07/19/11 and a URGENT: DEVICE RECALL letter was mailed to the firm's customers on 07/15/11 and 07/21/11. Discontinue use. For questions, contact: (800) 323-4306 or silencer@gfmd.com |  | 859 | 859 units | Worldwide to USA, United Kingdom, Ireland, Canada, and China. | Terminated | November 06, 2013 | 845 | 859
Z-2889-2011 | 59346 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10 Dual Program Reorder Pack; |  |  | July 11, 2011 | 2011 | Bio-Rad Laboratories Inc | Reports of calibration failure and late retention time using Elution Buffer AA10434 and AA10435. | Other | Bio-Rad Laboratories Inc sent "URGENT: MEDICAL DEVICE CORRECTION" letters to all affected D-10 Dual customers and Bio-Rad subsidiaries on July 11, 2011 via FedEx or fax. The letters identified the product, the problem, and the action needed to be taken. The letters ask that inventory be examined and affected reorder pack lots be quarantined and destroyed according to local, state and federal regulations. All D-10 Dual customers were instructed to review their chromatograms generated from these lots. If the calibration has passed with good chromatography, the controls are within specification, and the F, A1c, and A2 peaks elute in their specified retention time windows for the patient samples, the customer may continue to use and report results with the affected buffer lots. The letters include a chart showing Good Patient Chromatogram with Correct Retention times and Failed Calibrator Chromatogram with late retention times. If calibration fails and/or late retention times cause one or more peaks to be misidentified, affected reorder packs should be destroyed. The correction letters and response form should be translated, if necessary, and distributed to all affected customers. Tracking documentation should be compiled for all customers who receive the Medical Device Correction letter. The Medical Device Correction Response Forms should be completed and returned by fax to Bio-Rad CSD Regulatory Affairs at 510-741-3954 or a PDF copy to Jackie_Buckley@bio-rad.com |  | 418 | 418 | Worldwide -- USA (nationwide) including the states of: AR, CA, CT, FL, HI, IN, MD, MN, MO, MS, NJ, NY, OH, PA, RI, TX. and the countries of: Brazil, Canada, China, England, France, India, Italy, Mexico, New Zealand, and S. Korea. | Terminated | March 12, 2012 | 245 | 418
Z-2888-2011 | 59023 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | WarmAir Model 135 Convective Warming Device, |  |  | June 01, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | Due to its design the circuit breaker on the device could trip causing it to shut down and not provide heated air to the blanket. | Device Design | Cincinnati Sub-Zero Products, Inc. sent an 'URGENT Medical Device Recall Field Action Notice" dated July 12, 2011 to all affected customers.The letter includes: description of product, problem, and information to request a repair tubing kit free of charge. It also asks customers to complete and return a "Recall Response Form".For additional information e-mail WA135-2011@cszinc.com or call (800) 989-7373 or (513) 772-8810. |  | 5474 | 5474 units | Worldwide Distribution: (USA) Nationwide, and the countries of Australia, Brazil, Bulgaria, Canada, Chile, Ecuador, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Malaysia, Mexico, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Korea, Spain, Sri Lanka, Taiwan, Thailand, The Netherlands, Turkey, U.A.E., UK, Uruguay, Venezuela, and Vietnam. | Terminated | August 02, 2013 | 793 | 5474
Z-2890-2011 | 59036 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BACTEC(tm) MGI (tm) System Software, Catalog 445922, Lot V5.02A. |  |  | March 28, 2011 | 2011 | Becton Dickinson &amp; Co. | Software contained three additional set configurations which caused the system to reinitialized the AST set carrier defaults. | Software design (manufacturing process) | BD Diagnostic Systems verbally notified all affect customers via the telephone on March 28, 2011. The notification identified the product, the problem, and the action to be taken by the customer. An "URGENT PRODUCT RECALL" follow-up letter was mailed to all customers on April 8, 2011 via UPS. Customers were instructed to complete and fax the attached reply form to acknowledge their receipt of the notification to Regulatory Compliance, 410-316-4256. The letter instructs customers to refer to the modified release notes and IMPORTANT NOTICES outlined in the attached notification letter. If further assistance is needed regarding replacements and all other inquires, please contact BD Technical Services and Support at 1-800-638-8663. |  | 421 | 421 disks | Nationwide Distribution (USA) including states of: AK, AL, AR, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | July 29, 2011 | 123 | 421
Z-2893-2011 | 59088 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac Linear Accelerators; |  |  | May 24, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The throat cover on the High Energy Clinac may detach if not properly installed and possibly strike a patient. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter provides directions on how to re-install the gantry center throat cover.Customers can contact their Varian Customer Support Manager at 1-888-827-4265 for further clarification. |  | 278 | 278 | Worldwide Distribution-including USA (nationwide) | Terminated | March 14, 2012 | 295 | 278
Z-2860-2011 | 59148 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | I-Flow ON-Q with Select-A-Flow Pumps, |  |  | April 15, 2011 | 2011 | I-Flow Corporation | The recall was initiated because I Flow became aware that a small percentage of ON-Q Pump with Select A Flow may develop a minor leak with a flow rate controller. There may be consequences for the patient, such as under delivery of pain medication or the need to prematurely remove the pump, which would result in loss of the pain treatment provided by the pump. | Process control | I-Flow Corporation sent an "Important Voluntary Market Withdrawal Notice" dated May 12, 2011 to all affected customers.The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to quarantine and return all impacted products within their facilities. Customers were instructed to respond with in 5 days of receiving the Market Withdrawal Notice and fax it to (949) 206-2663. Customers with additional questions, were instructed to call I-Flow Customer service at (800) 448-3569 (8-5pm, Pacific Standard time) or I-Flow 24 hour technical support line (800) 444-2728. |  | 96892 | 96,892 units total (95,637 units in US) | Nationwide (USA) Distribution and worldwide to: Australia, Canada, Cyprus, Italy, Germany, Finland, Greece, Israel, South Africa, Sweden, UAE, and UK. | Terminated | August 09, 2012 | 482 | 96892
Z-2892-2011 | 59377 | [6 empty] | Aquamantys 2.3 Bipolar Sealer |  |  | July 18, 2011 | 2011 | Salient Surgical Technologies, Inc | Insulating material on the shaft of the device may separate and expose the electrical conductor and result in burning the patient skin | Nonconforming Material/Component | Salient Surgical Technologies, Inc. sent an "URGENT MEDICAL DEVICE RECALL" notification letter dated July 18, 2011 to all affected customers and distributors. The letter describes the product, problem and the actions to be taken. The letter requested that customers and distributors immediately identify and discontinue use of all Aquamantys 2.3 Bipolar Sealer devices in inventory and quarantine them until they can be returned to Salient. The letter states that a Salient sales representative will arrange return of the recalled product to Salient. The letter provides information for customers about the replacement of the recalled product. If you have any questions, please contact Customer Service, at 866-777-9400, or your sales representative. |  | 48963 | 48, 963 units | Worldwide Distribution - USA (nationwide) and countries ofChina, Denmark, France, Germany, Greece, Italy, Japan, , South Africa, Spain, and Turkey. | Terminated | March 09, 2012 | 235 | 48963
Z-2891-2011 | 59272 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DeRoyal (R) Incision and Drainage Tray, RES 47-763.02, Rx Only, STERILE EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA |  |  | April 28, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | Custom surgical kits contain recalled Triad PVP prep pads | Other | DeRoyal sent a Customer letter dated July 6, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.The customer was instructed to reference the attached spreadsheet to identify your accounts kits/trays and the corresponding lot numbers affected by this recall.Please complete and return the Acknowledgement Form by fax ( 865-362-3716) confirming you have been notified of this action.For further questions, please call ( 865) 362-2334. |  | 11856 | 11,856 kits/trays | Nationwide Distribution--including the states of CA, FL,GA, HI, IL, IN, LA, MA, MO, MS, NY, PA, and RI. | Terminated | August 07, 2014 | 1197 | 11856
Z-2897-2011 | 57856 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Aplio Artida Ultrasound Diagnostic System SSH-880CV. |  |  | October 01, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems, Inc. is initiating a field correction on the Aplio Artida Ultrasound Diagnostic System because the panel operation for the Aplio Artida may become disabled (i.e., system may hang up) when processing to display patient information (heart rate, time, VCP counter, ect.). | Software design | Toshiba America Medical Systems, Inc. sent out an "Urgent Medical Device Correction" dated December 21, 2010 to all affected customers.The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were notified that Toshiba will issue a field modification to correct the software, and if the problem occurs, they were instructed to shutdown the system holding down the power supply button for 5 seconds for 5 seconds and then restart the system. A form to send back to Toshiba is included. For questions on this recall call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 46 | 46 units | Worldwide Distribution: (USA) Nationwide distribution including the states of MD, OH, PA, NY, MA, GA, AL, TX, IL, WI, IA, MN, KY, MT, OR, LA, and CA; and the country of Puerto Rico. | Terminated | April 20, 2012 | 567 | 46
Z-2894-2011 | 59113 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Arrow ASK-20608-CMC StimuCath Continuous Nerve Block Kit. |  |  | June 17, 2011 | 2011 | Arrow International Inc | Some of the outer trays used in the packaging were found to have an insufficient seal. Therefore, packaging may have been compromised, and sterility of the product cannot be guaranteed. | Packaging | Arrow International, Inc. sent an "Urgent Medical Device Recall" letter dated June 17, 2011 to all affected customers.The letter describes the problem and provides recommendations including; quarantine affected product, return of product, and send an acknowledgement form back to manufacturer.For additional information call 800 233-3187. |  | 95 | 95 kits | Nationwide Distribution including the state of North Carolina. | Terminated | October 01, 2012 | 472 | 95
Z-2895-2011 | 58031 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion One: Aquilion TSX-301A, 64 System, X-Ray, Tomography. |  |  | January 07, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) is initiating a field correction for the TSX-301A and TSX-301B: Aquilion 64 System, X-Ray, Tomography because communication between reconstruction units in the system console may be broken due to failure of the optical conversion connector, resulting in a break of communication between units. | Software design | Toshiba America Medical Systems, Inc. sent an "Urgent Medical Device Correction" letter dated January 4, 2011 to all affected customers.The letter provides the customers with an explanation of the problem and the firms plan to install preventive measures and provided them with a Customer Reply form to send Toshiba. A representative will contact the customers for an appointment to install the measures to their CT system. For question on this recall call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 6 | 6 | (USA) Nationwide Distribution including the states of WI, MO, TN, IA, AL, AZ, TX, and GA. | Terminated | August 17, 2012 | 588 | 11
Z-2848-2011 | 59047 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | Vital Signs- A GE Healthcare Company - Hygroscopic Condenser Humidifier (HCH)/Anesthesia Breathing Circuit; GE Medical Systems SCS 283 rue de la Miniere 78530 BUC, France p +33.1.30.70.40.40; Vital Signs Colorado Inc 11039 E. Lansing Circle Englewood, CO 80112 P +1.973.956.5300. |  |  | April 29, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | The Hygroscopic Condenser Humidifier (HCH) device may have an occlusion that can prevent proper flow of the medical gases or oxygen, possibly resulting in insufficient oxygen delivered to the patient. | Device Design | An "Urgent Medical Device Correction" letter with the return confirmation form was sent to affected customers. Phone calls were also placed to customers. Consignees were informed of the affected device, issue, and actions to take. Customers were instructed to isolate all affected product, arrange return by contacting Customer Service at 1-800-932-0760, and return only unused product. Customers were asked to complete the attached confirmation form and fax back as per the instructions. For questions, contact your local Vital Signs Account Manager or call Customer service at 1-800-932-0760. |  | 222009 | 222,009 | Worldwide Distribution including the United States. | Terminated | August 06, 2014 | 1195 | 222009
Z-2883-2011 | 59212 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Arrow NextStep Antegrade Chronic Hemodialysis Catheter. |  |  | June 28, 2011 | 2011 | Arrow International Inc | Firm has received complaints indicating breakage and/or separation of the stylet within the Arrow(R) NextStep(R) Antegrade Catheters. | Device Design | An "Urgent Medical Device Recall" letter dated June 28, 2011 was issued to consignees. The letter requested that customers check their stock, cease use and distribution, and quarantine all affected product. Consignees were further directed to complete the "Recall Acknowledgement &amp; Stock Status Form", return the form to their sales representative, and return affected product as instructed on the letter. For questions or concerns, call 1-800-233-3187. |  | 41 | 41 | United States (states of CA, DE, FL, MI, NC, and TN). | Terminated | April 16, 2012 | 293 | 41
Z-2900-2011 | 59046 | No | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | AxioBionics Wearable Therapy BioShorts that includes 2-16 BioGel Velcro Electrodes in sizes: 2"x2", 2"x2.5", 2"x3", 2"x3.5", 2"x4", 2"x5", 2"x6", 2"x8", 1.5"x3", 1.5"x5", 1.5"x6", and 1"x1.5". |  |  | May 27, 2011 | 2011 | Axiobionics, LLC | Premature Device Failure : Failure of the electrode sites to conduct and provide muscle stimulation due to silver fabric component failure. | Nonconforming Material/Component | Axiobionics, LLC is notifying all affected customers by phone. Three attempts will be made to contact the customer by phone. Contact by telephone will be made to prescribing physician to determine if Axiobionics has latest patient contact information. If information is different, new contact information will be recorded and 3 additional attempts will be made with the new information.If the patient cannot be reached by telephone, a certified letter will be sent to the patient's address. If contacting by mail, the letter template will be provided to the FDA recall office . All patients should have been contacted by telephone, certified mail, or in person by June 10, 2011.All patients affected by the field correction will be asked the following questions to determine if their electrodes are functional and resistances are not exceeding the output capabilities of the stimulator they are using. 1. Are you using your garment?2. Is your garment delivering the stimulation required for your therapy?3. If the patient is using an NP4 stimulator: What are the resistance readings on each channel?4. If the patient answers No to question 2 or states that the resistance is reading greater than 200 ohms on at least one channel in question 3, the garments will be returned to Axiobionics for repair.Garments that are not functioning correctly will be repaired first. Functioning garments will be tested by immersion into a salt bath under the given protocol to determine if the silver fabric is properly coated.Questions regarding this correction should be directed to the Axiobionics representative at 800-552-3539. |  | 9 | 9 | Nationwide Distribution--USA (nationwide) including the states of AZ, GA, IL, IN, MI, MO, NY, TX, and WY. | Terminated | August 29, 2013 | 825 | 106
Z-2913-2011 | 58920 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***BIOMET MICROFIXATION***CENTER DRIVE STAINLESS STEEL***2.0MM SYSTEM SCREWDRIVER BLADE***Non-Sterile Product***BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***" |  |  | March 15, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling Center Drive Blade, 2.0, Model Number 01-9047, Lot # 064980. Product is non-conforming (blades do not retain screws). | Nonconforming Material/Component | The firm, Biomet Microfixation, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" undated letter to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to remove the affected products and immediately return them to Biomet Microfixation, Attn: Global Complaint Coordinator, 1520 Tradeport Drive, Jacksonville, FL 32218-2480; complete and return the Inventory Reconciliation sheet via Fax to the attention of: Global Complaint Coordinator at 904-741-9425, and notify their customers if product was further distributed. If you have any questions, contact the Global Complaint Coordinator at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 35 | 35 | Worldwide distribution: USA (nationwide) and countries of: Canada, Israel and South Africa. | Terminated | November 09, 2011 | 604 | 35
Z-2902-2011 | 59136 | No | CBR | Nitrous oxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Ultraview Spacelabs Medical Multigas Analyzer, Model 91518, V1.00.14. |  |  | June 08, 2011 | 2011 | Spacelabs Healthcare, Llc | Spacelabs received a report of the Model 91518 Multigas Analyzer trending low levels of an anesthetic agent that is not present during a particular case. The error is limited to only the trend table data. | Software design | Spacelabs Healthcare, Llc. sent a "Customer Bulletin" dated June 20, 2011 to all affected customers. The customers were notified of the trend problem and are advised to follow the instructions in the Perioperative section and the ADMIT/DISCHARGE section of the users manual.For questions contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support or 1-425-657-7200 x5089. |  | 481 | 481 units worldwide (333 units internationally and 148 units in the US) | Worldwide Distribution: (USA) Nationwide distribution, and the countries of ARGENTINA, BAHRAIN, BOLIVIA, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, ECUADOR, EGYPT, FRANCE, INDIA, MEXICO, NICARAGUA, PARAGUAY, PERU, POLAND, Slovakia, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, CHINA, THAILAND, UNITED ARAB EMIRATES, U.K., URUGUAY, VENEZUELA, and VIETNAM. | Terminated | September 21, 2011 | 105 | 481
Z-2915-2011 | 59210 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF Big Bore computer tomograph software version 3.6. Catalog #882476, Gantry S/N 9001-9023. |  |  | August 05, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | Two non-safety related defects have been found on the Philips GEMINI TF Big Bore running version 3.6; if these defects were to recur, it could impact the performance or reliability of the system. Philips is issuing a software update to all Philips GEMINI TF Big Bore customers to correct these defects. | Software design (manufacturing process) | Philips Medical Systems (Cleveland) Inc. sent an "GEMINI TF BIG BORE CPE CUSTOMER LETTER" dated October 20, 2011 to all affected customers. The letter describes the product, problem, and updates that are available to the customers. The letter provides additional information about acquisition workflows for SUV Calibrations and Validations.For support concerning this upgrade, contact the Customer Care Solutions Center at 1-800-722-9377. |  | 23 | 23 units | Worldwide Distribution-USA (nationwide) including the states of CO, LA, MD, MI, MO, MN, PA, TX, VA, VT and WI and the countries of Belgium, Germany, Italy, Saudi Arabia and Switzerland. | Terminated | September 21, 2012 | 1143 | 23
Z-2914-2011 | 59166 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Vertical Brake Hubs component of the Philips MX8000 Series Dual v. Exp. Computed Tomography System, Serial Numbers: 8008-8413, 80001-800022, Model #728130, 510 (k) #K010817. |  |  | June 03, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Philips Medical Systems decided to conduct a recall/field correction of the patient support vertical brake hub component in all of their CT MX8000 Dual v. Exp Computed Tomography X-Ray System scanners due to a component defect which could cause an unexpected downward motion of the table (couch) while the patient is on the table. This event could occur because the short key connecting the vertical | Device Design | Philips Medical Systems (Cleveland) Inc. issued an "URGENT - Medical Device Correction - Field Safety Notice" letter dated June 1, 2011 to all affected customers, via Certified Mail for tracking purposes. The notification identified the product, the problem involved, and the action(s) to be taken by the customers with regard to this device issue. The notification instructs the customers to monitor their units and if the recalled product malfunctions, to immediately notify their Phillips service representative. The letter also states that a Philips Field Service Engineer will visit each customer location and perform a brake repair or replacement on their system to resolve the issue. For further information or support, the customers are instructed to contact their local Philips representative or local Philips Healthcare office. For North America and Canada the customers are instructed to contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompt(s). |  | 412 | 412 system units were distributed. | Worldwide Distribution -- USA (nationwide) including the states of: AR, AZ, CA, CO, FL, FA, HI, IL, IN, KY, MA, MI, MO, MS, NC, NV, NY, OH, OK, SC, SD, TN, TX, VA, VT, WA and Puerto Rico and the country of Canada. | Terminated | August 07, 2013 | 796 | 412
Z-2903-2011 | 58698 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | "***REF MCA5-15SB QTY: 1***1.5 mm Fluted Ball, 5.6 cm. Use with Micro Curved Attachment (MCA).***Sterilized using Irradiation***RX Only***ANSPACH Manufacture The Anspach Effort, Inc. 4500 Riverside Drive |  |  | October 13, 2010 | 2010 | The Anspach Effort, Inc. | On July 30th, 2010 the Anspach Effort, Inc. initiated a voluntary recall on their Anspach Single Use, Sterile Bone Cutting Burrs. A complaint was recieved stating that the product had failed during a procedure. | Nonconforming Material/Component | The Anspach Effort, Inc. sent an "URGENT: MEDICAL DEVICE PRODUCT REMOVAL" letter dated October 20, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.The letter instructs customer to screen their inventory for the affected product and not to include any product received after October 18, 2010. Customers are instructed to contact Anspach Customer Support at 800-327-6887 to arrange for return and replacement of product. Customers are to complete and return a Customer Reply Form via fax at 800-327-6661.Questions regarding this notice should be directed to Anspach Customer Support at 800-327-6887.Customer notification began on 10/20/2010 in writing (Fed Ex) with instructions to return the required devices. Anspach utilized their sales representatives to facilitate the product replacement and return activities. |  | 46 | 46 | Worldwide Distribution-USA (nationwide) and the countries of Australia, Belgium, France, Italy, Japan, South Africa, Spain, Sweden, and United Kingdom. | Terminated | November 07, 2011 | 390 | 1816
Z-2754-2011 | 59111 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic 6F Zuma Guide Catheter with AL 1.5 Curve with side holes |  |  | June 08, 2011 | 2011 | Medtronic Vascular | Mislabeled; Product labeled as 6F Zuma Guide with AL 1.5 curve with side holes may contain a 5F Zuma Catheter with SAL 1.0 curve | Labeling Change Control | The firm, Medtronic, sent an "Urgent Medical Device Recall Notice(for US)" letter dated June 7, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to segregate the product so that it will not be used. A Medtronic Sales Representative will visit the customers to account for all units they received and initiate a credit as appropriate.If you have any questions, please call (978) 739-0371 or (978) 739-3259. |  | 5 | 5 units | Nationwide distribution: USA including states of: CA, FL, and MA. | Terminated | April 17, 2012 | 314 | 5
Z-2916-2011 | 59441 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Vital Sign elance Monitor, Model 93300 |  |  | July 06, 2011 | 2011 | Spacelabs Healthcare, Llc | Model 93300 Vital Signs Monitor may stop operating while in use, and when this occurs, the monitor may also fail to power ON. The resulting situation may cause a loss of monitoring for the patient. | Process design | Spacelabs Healthcare, Llc sent an "Urgent-Medical Device Correction" letter dated 18 July 2011 to all affected customers. The letter notified the customers of the problem and recommended customers make their staff aware. It also informed them a Field Service Engineer will contact them to check, and if necessary make adjustments.Customers can contact Spacelabs Healthcare at 800-522-7025 and select 2 for Technical Support or outside the U.S. at +86512-8717-1888, extension 304. |  | 1970 | 1970 units | Worldwide Distribution: (USA) Nationwide distribution including the state of WI, and the countries of Puerto Rico, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Chile, China, Colombia, Cyprus, Dominican Republic, Ecuador, Egypt, El Salvador, France, Germany, Greece, Guatemala, Honduras, Hong Kong, India, Indonesia, Iraq, Israel, Italy, Kenya, Korea, Kuwait, Lebanon, Malaysia, Napel, Netherlands, Nicaragua, Oman, Palestine, Panama, Paraguay, Peru, Philippines, Republic of Macedonia, Romania, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Switzerland, Syrian Arab Republic, Thailand, Trinidad and Tobago, Tunisia, Turkey, UAE, United Kingdom, Venezuela, and Vietnam. | Terminated | January 25, 2012 | 203 | 1970
Z-2928-2011 | 58684 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | remel Haemophilus Test Medium (Agar), 100mm plates, 10/pk, Ref. #R01503. The firm name on the label is Remel, Lenexa, KS. |  |  | April 21, 2011 | 2011 | Remel Inc | Organism strains may fail to consistently produce confluent growth | Other | The firm, Remel, Inc., sent two "IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" letters one dated April 20, 2011 to: Microbiology Lab Supervisors and the second dated April 25, 2011 to:Distributors via regular mail. The letters described the product, problem and actions to be taken. The customers were instructed to to review their inventory for the listed product lot(s); discard all remaining units of the product lots; and complete and return the Product Inventory Checklist by May 9, 2011 regardless if they have affected product in the enclosed self-addressed, stamped envelope. The distributors were further instructed to notify their customers if they further distributed the product. If you have inquiries concerning this advisory, please contact our Technical Services Department at 800-255-6730 or 913-888-0939. |  | 223 | 223/10-pack boxes | Nationwide distribution. | Terminated | June 21, 2013 | 792 | 761
Z-2919-2011 | 59102 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO versions equal to or greater than 2.4.0 and less than 4.63.0 |  |  | June 09, 2011 | 2011 | Computerized Medical Systems Inc | The XiO Software currently prompts for entry of CT to ED data for a studyset. No indication is given that the Electron Density data should only be used for Photon and Electron calculations and either Relative Stopping Power or Mass Density should be used for Proton calculations (depending on the algorithm). | Software design | The firm, Eletka, Inc., issued an "Important Safety Notice" dated May 18, 2011 via mail to all affected sites. The notice explains the issue and identifies a work around to the problem until the new release is available. The customers were instructed to complete and return the enclosed "Confirmation of Receipt" form via FAX to 314-993-1175, attn: QA; e-mail or mail: CMS Software, Attn: QA, 13723 Riverport Drive, Suite 100, Maryland Heights, MO 63043.The notice states that "this problem has existed since XiO Release 2.4.0 and has been resolved in XiO Release 4.63.0 by adding warning messages and explanations to the XiO On-Line Help. Your site will be notified when this release is available."If you have any question, call 314-993-0003, 800-878-4267 or 408-830-8023. |  | 7 | 7 | Worldwide distribution: USA (nationwide) including: IL, IN, MA, NJ and OK; and country of Germany. | Terminated | April 04, 2013 | 665 | 7
Z-2918-2011 | 58978 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Aruba Wireless Network Access Points, Model 65 utilized by the Acuity Central Monitoring System. |  |  | May 16, 2011 | 2011 | Welch Allyn Protocol, Inc | Aruba Wireless Network Access Points, Model Number 65, experienced a higher than expected incidence of dropouts due to external RF interference. | Software design | Welch Allyn Protocol, Inc. sent a letter dated May 16, 2011 to affected customer.The letter stated there were concerns over drop-outs in the system recommended the system at this hospital to revert back tothe Symbol wireless infrastructure. For questions call Welch Allyn at 315-685-4100. |  | 12 | 12 units were affected | Worldwide distribution: (USA) Nationwide distribution including the state of California; and the countries of Singapore; Philippines, France, Italy, Malaysia, Poland, UK, England, UAE, Germany, Spain, Ireland, Portugal, New Zealand, Hong Kong, Thailand, Australia, Canada, and the Netherlands. | Terminated | August 12, 2011 | 88 | 12
Z-2931-2011 | 58852 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Pinnacle Pelvic Floor Repair Kit--Anterior/Apical STERILE, and Pinnacle Pelvic Floor Repair Kit--Posterior STERILE. |  |  | May 10, 2011 | 2011 | Boston Scientific Corporation | The device may exhibit low tensile strength between the needle and suture and lead to needle detachment during mesh leg placement. | Device Design | Boston Scientific Corporation sent on May 10, 2011 an "URGENT MEDICAL DEVICE RECALL-IMMEDIATE ACTION REQUIRED" letter to all affected customers.The letter included; reason for recall, instructions for the quarantine, discontinued use, and return of the devices, and information for Medwatch reporting. The letter also includes a form that is to be faxed to the firm.For information call (508) 683-4678. |  | 540 | 540 | Worldwide Distribution; (USA) Nationwide distribution, and the countries of Puerto Rico, France, Great Britain, Ireland, Italy, Spain, Sweden, Canada, and Australia. | Terminated | November 15, 2011 | 189 | 540
Z-2926-2011 | 59214 | No | INK | Powered wheeled stretcher. | Physical Medicine | Physical Medicine | 510(k) | 1125 Prime Zoom Stryker Medical Stretcher |  |  | July 19, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | Stryker Medical has determined that some Prime Brake Pad Assembly components were manufactured with brake shoes that did not meet specifications. Accordingly some brake shoes used within the Prime Brake Pad Assembly may wear prematurely, over the life of the product. Reduced braking force may result in a patient fall or caregiver muscle strain, either of which may require medical intervention. | Nonconforming Material/Component | Stryker Medical sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated July 18, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.The letter instructs customers to locate the affected stretchers and to test the brakes. If the brakes are found to be inoperable, customers are instructed to remove the stretchers from service and contact the firm at 800-STRYKER, option 3. Additionally, customers are instructed not to put the stretchers back in service until they are repaired by the firm's field service representative. For units with properly functioning brakes, the letter also instructs the user to test the brakes before each use until the unit has been serviced. A post card was enclosed to confirm receipt of the letter. Questions or concerns may be addressed by calling Stryker at 269-389-6604. |  | 7 | 7 | Worldwide Distribution-USA (nationwide) including the states of AL, CA, FL, GA, HI, IL, MA, MN, MO, NH, NJ, NV, NY, PA, RI, TX, UT, and VA and the country of THE NETHERLANDS. | Terminated | August 28, 2013 | 771 | 164
Z-2932-2011 | 59356 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RUSCH, GB: Female Catheter, All Purpose PVC, 08 Fr. 18 cm, Distributed by Teleflex Medical, Research Triangle Park, NC |  |  | June 30, 2011 | 2011 | Teleflex Medical | Pin holes were detected in the packaging therefore sterility of the device can not be guaranteed. | Other | Teleflex Medical sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 30, 2011 to all affected consignees. The letter described the product, problem and actions to be taken. The letter instructs the customers to discontinue use and quarantine any of the affected products. Customers were asked to complete the attached Recall Acknowledgement Form and return as per the instructions. A separate letter was sent to distributors instructing them to communicate the recall to any of their customers who may have receive the affected products, using the recall letter and form as a template. For questions, contact your local sales representative or call Customer service at 1-866-246-6990 |  | 14378890 | 14,378,890 units | Worldwide Distribution and (USA) Nationwide Distribution. | Terminated | April 29, 2014 | 1034 | 14378902
Z-2930-2011 | 59300 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical da Vinci Si Surgical System (IS3000); |  |  | July 05, 2011 | 2011 | Intuitive Surgical, Inc. | Potential failure of the retention component of the Master Tool Manipulator (MTM) of the Intuitive Surgical da Vinci Si Surgical System which could cause uncontrolled movement. | Device Design | Intuitive Surgical sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated July 1, 2011 to all affected customers. Customers were asked to complete the attached response form and return as per the instructions. The letter identifies the product, problem, and actions to be taken by the customers. The letter reinforce to customers to follow the product labeling WARNING. The letter states that all affected customers sites will be receiving a call from a Customer Support staff to schedule a retrofit of the component. A letter was also sent to non affected customers to inform them of the urgent medical device correction and actions to be taken. If you have any questions, please contact Customer Service at 1-800-876-1310 or for Outside United States call +41 21 821 2020. |  | 183 | 183 | Worldwide Distribution -- USA (nationwide) and the countries of: Belgium, Czech Republic, Denmark, France, Germany, Greece, India, Ireland, Israel, Italy, Netherlands, Norway, Philippines, Portugal, Saudi Arabia, Singapore, South Korea, Spain, Switzerland, Turkey and United Kingdom. | Terminated | September 02, 2011 | 59 | 183
Z-2920-2011 | 59419 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Solution Set with Duo-Vent Spike; 109" (2.8 m), Male Luer Lock Adapter 10 drops/mL. |  |  | July 21, 2011 | 2011 | Baxter Healthcare Corp. | The Solution Sets are being recalled because the position of the slide clamp and roller clamp may have been reversed during assembly. The reversed orientation of the slide clamp and roller clamp will cause the set to be incorrectly loaded into an infusion pump, resulting in fluid flow from the patient to the pump, rather than from the pump to the patient. | Employee error | Baxter Healthcare Corporation sent a "Urgent Product Recall" letter dated July 21, 2011 to all affected customers.The letter included a description of the problem and requested they contact Baxter at 1-888-229-0001 to return all affected product. They were asked to complete and return the attached customer reply form, and to contact and notify their customers.For questions contact Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 4752 | 4,752 units | (USA) Nationwide Distribution including the states of Alabama, California, Iowa, Illinois, Indiana, Massachusetts, Maryland, Michigan, North Carolina, New Hampshire, Ohio, Pennsylvania, Tennessee and Virginia. | Terminated | June 18, 2012 | 333 | 4752
Z-2917-2011 | 59301 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Automated External Defibrillator (non-wearable). Model number 9200, 9200N, 9200DN, 9200RD, 9210N, 9210RD, and CardioVive 92530. |  |  | June 29, 2011 | 2011 | Cardiac Science Corporation | Cardiac Science Corporation (CSC) intent to end service support and discontinue accessory sales of Powerheart G2 Automated External Defibrillators. Customers are recommended to choose a replacement AED prior to the current AED displaying any "Service Required" self test messages. | Nonconforming Material/Component | Cardiac Science Corporation (CSC) sent two "9200 SERIES CARDIAC SCIENCE AED NOTICE" letters dated August 2010 and March 2011 to all affected customers. The letters identifies the product, problem, and actions to be taken by the customers.Customers are encouraged to upgrade their devices through special offers by calling CSC at 800-426-0337, Option 2.For questions about the end of factory support for the affected devices, call the Technical Support team at 800-426-0337, Option 1.*****In August 2011, Cardiac Science sent out another updated recall notification letter to their consignees. This letter stated that once a device has entered the "Service Required" state, please contact the firm at 800.426.0337 Option 2. Per prior notifications, the 9200 series AEDs have reached the end of service life and can no longer be repaired. Information about these notifications may be found as an announcement in the Services and Support section of Cardiac Science website.Cardiac Science advised the customers to upgrade to a Powerheart AED G3 (this device supports the 2010 AHA/ERC Guidelines). It is stronglyrecommended that customer to choose a replacement AED prior to the current AED displaying any "Service Required" self test messages. For information about special offers to help you upgrade please call 800.426.0337. Option 2. |  | 48700 | 48,700 units (41396 units in US; 7304 units Internationally) | Worldwide Distribution-USA (nationwide) and the countries of AE, AT, AU, AW, BE, BR, CA, CH, CL, CN, CO, CY, CZ, DE, OK, EG, ES, FI, FR, GB, HK, HR, HU, 10, IE, IL, IN, IR, IS, IT, JP, KR, KW, KY, LB, LT, LV, MT, MX, NG, NL, NO, NZ, PE, PF, PL, PR, PT, QA, RO, RU, SA, SG, SK, SY, TH, TR, TT, TW, VG, ZA, and ZW. | Terminated | December 20, 2011 | 174 | 48700
Z-2945-2011 | 59058 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard PS Open Box Femoral 70 MM Left Interlock, " FOR CEMENTED USE ONLY" CO-CR-MO ALLOY. |  |  | July 06, 2011 | 2011 | Biomet, Inc. | The recalled device was manufactured with two distal pegs but package labeling indicates the device is pegless. | Employee error | Biomet Orthopedics sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter to all affected customers.The letter included: reason for the recall, risks associated with use of the device, and alternate instructions for use of the recalled device. Customers are instructed to quarantine and return the devices, and a return form.For information or questions on this recall call Biomet at (574) 371-3756. |  | 8 | 8 | Worldwide Distribution: (USA) Nationwide including the states of NJ, PA, LA, SD, TX, and KS; and the country of The Netherlands. | Terminated | September 26, 2012 | 448 | 8
Z-2948-2011 | 59252 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT Gamma Camera, 3/8". |  |  | June 15, 2011 | 2011 | Philips Medical Systems | Non-conformances with BrightView X/XCT systems running version 2.0.1 including data transfer, XCT acquisition, Pixel Mask for image reconstruction and table moving incorrectly. | Software design | Philips Medical Systems, Inc. set a "Customer Information" letter dated July 8, 2011 to all affected customers.The letter informed them of an upcoming software updated (version 2.5).For questions please contact 800-722-9377 (option 5). |  | 213 | 213 units, all systems | Worldwide Distribution: (USA) Nationwide distribution and the countries of Greece, Australia, Belgium, China, Czech Republic, France, Germany, Indonesia, Israel, Japan, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Puerto Rico, Saudi Arabia, Slovakia, South Korea, Spain, Switzerland, Taiwan U.A.E., and the UK. | Terminated | February 15, 2013 | 611 | 213
Z-2950-2011 | 59454 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | CAPIOXÂż Flexible Venous Reservoirs, 1500 mL, product codes: CXFVR1500X and CXFVRI500XC, used singly or packaged as components in cardiovascular procedure trays labeled in part ***Terumo Caerdiovascular Systems Corporation, Elkton, MD 21921*** |  |  | July 21, 2011 | 2011 | Terumo Cardiovascular Corporation | Reservoir blood bag used in cardiovascular procedures may leak. | Nonconforming Material/Component | Terumo Cardiovascular Systems Corporation sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated July 20, 2011, to all affected consignee. The letter described the problem, the potential hazard, and the actions to be taken. The letter instructed customers to identify affected inventory and segregate from non-affected product. The recalling firm advises users to not use affected product and return for replacement. The customers are instructed to call the recalling firm Customer Service for a Returned Goods Authorization (RGA) number. Customers were asked to complete the attached Response Form and return as per the instructions. Consignees with questions should call 1-800-521-2818 Monday through Friday 8AM to 6PM EST. |  | 253 | 253 reservoirs | (USA) Nationwide Distribution including states of: FL, MI, MT, NY, TN and VA. | Terminated | February 10, 2012 | 204 | 253
Z-2946-2011 | 59368 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions, ProntoÂż V4, Extraction Catheter, Sterilized using ethylene oxide, Rx Only. The following sizes/model numbers are affected: 5.5F( 4005), 6F(4006, 7F(4007, 8F(4008) are involved in the recall. |  |  | July 11, 2011 | 2011 | Vascular Solutions, Inc. | Vascular Solutions , Inc recently discovered that there was a manufacturing defect with Pronto V4 extraction catheters (Models 4005, 4006, 4007, 4008). Use of this defective product could lead to a tip separation. This maycause injury and/or procedural delay, resulting in a deteriorating hemodynamic condition of the patientwhich may require medical intervention. | Process control | Vascular Solutions, Inc. sent an "URGENT: MEDICAL DEVICE FIELD ACTION" letter dated July 8, 2011 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers.The letter instructs customers to identify the affected product and remove from current inventory. Included with the letter was a Field Action Customer Inventory Form for customers to complete and return. VSI will contact customers for replacement product upon receipt of the form and returned product.Contact your local representative for questions concerning this notice. |  | 1086 | 1086 | Nationwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, FL, ID, IL, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK ,OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Terminated | April 03, 2012 | 267 | 1086
Z-2952-2011 | 59382 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Turbo-Booster Laser Guide Catheter, 8 Fr, REF 518-063, Sterile EO, Spectranetics Corporation, 96 Talamine Court, Colorado Springs, CO 80907. |  |  | July 18, 2011 | 2011 | Spectranetics Corporation | Packaging integrity may be compromised, thus sterility can not be assured. | Package design/selection | Spectranetics sent an "PRODUCT RECALL NOTICE: EU SHIPMENTS OF TURBO-BOOSTER LASER GUIDE CATHETERS" letter dated July 15, 2011 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory and set aside affected product. Customers are to contact their sales representative for information on how to return the affected product and receive replacement product.Further information is available at +011.49.719.447.2539. |  | 123 | 123 units | Foreign Distribution-including the countries of Belgium, Cyprus, Germany, Greece, Italy, Spain, Switzerland, and United Kingdom. | Terminated | August 12, 2011 | 25 | 180
Z-2953-2011 | 58043 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Hemochron JR. Microcoagulation Systems Prothrombin Time (PT). |  |  | May 10, 2011 | 2011 | International Technidyne Corp. | Certain lots of Hemochron Jr. PT Cuvettes (J201) demonstrate a higher bias than historically observed when compared to laboratory reference instruments. | Device Design | International Technidyne Corp. (ITC) sent "Urgent Medical Device Recall" letters dated May 10, May 25 and June 09, 2011 to all affected customers.The letter included a description of the problem, asked customers fill out the attached form, and return affected product for replacement.For questions, please contact ITC Technical Support at (800) 631-5945. |  | 10931 | 6,218 boxes in US ; 4713 boxes outside US | Worldwide Distribution: (USA) Nationwide Distribution including the countries of; Austria, Australia, Canada, Switzerland, Germany, Denmark, Estonia, Spain, Finland, France, United Kingdom, Hong Kong, Ireland, Israel, India, Italy, Korea, Liechtenstien, Libya, Netherlands, Norway, Oman,Portugal, Qatar, Saudi Arabia, Sweden, Turkey, Taiwan, and South Africa. | Terminated | November 02, 2012 | 542 | 10931
Z-2956-2011 | 59379 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics cardioPAT -Cardiovascular Perioperative Autotransfusion System |  |  | July 27, 2011 | 2011 | Haemonetics Corporation | Updated Operations Manual for the deployment of the spill collection system and cleaning | Component design/selection | Haemonetics sent a "FIELD NOTIFICATION: UPDATE OF CardioPAT OPERATIONS MANUAL FOR DEPLOYMENT OF SPILL COLLECTION BAG AND CLEANING" letter dated July 27, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter requests that customers complete an enclosed Customer Confirmation sheet indicating that the affected devices are in their possession and to return the sheet via fax to Haemonetics at 781-987-9252. Additionally, instructions for deployment of the spill collection drainage system, revised Operations Manual and Quick Reference Guide Addendum are enclosed with the letter.Contact Haemonetics at 1-800-537-2802 for questions regarding this notice.. |  | 316 | 316 units | Worldwide Distribution--USA (nationwide) and the countries of Canada, France, Germany, Great Britain, Italy, Japan, Spain, and The Netherlands. | Terminated | June 09, 2014 | 1048 | 316
Z-2955-2011 | 59392 | No | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | Supreme Electrophysiology Extension Cable, REF 401984, For use with Hexapolar Catheters, Sterile EO, Rx only, St Jude Medical, 14901 DeVeau Place, Minnetonka, MN 55345. |  |  | June 01, 2011 | 2011 | St. Jude Medical | St. Jude Medical is recalling a single lot (3334664) of Supreme Electrophysiology Extension Cable which were packaged with a Quadripolar cable instead of a hexapolar Cable. St. Jude Medical has determined that the incorrectly labeled product poses no safety risk since the extension cable can be safely used with a corresponding Quadripolar catheter but cannot be incorrectly connected to a Hexap | Labeling mix-ups | St. Jude Medical sent an "IMPORTANT MEDICAL DEVICE FIELD ADVISORY NOTICE" letter dated June 8, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to locate the affected product and to complete and return an attached SJM AFD Reconciliation Form. The form will be used to replace product and as documentation that the customers received the notice. Fax the completed form to 877-497-031.Contact the Director of Product Marketing at 512-329-1054 for questions regarding this notice. |  | 32 | 32 | Nationwide Distribution-USA (nationwide) including the states of FL, NC, NY, and OR. | Terminated | August 16, 2012 | 442 | 32
Z-2959-2011 | 59119 | No | FJS | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Aspira Drainage Kit, Product Code 4991502. |  |  | May 21, 2011 | 2011 | Bard Access Systems | Labeling for drainage kit contained incomplete information. | Error in labeling | The firm, BARD, sent a letter entitled "Labeling Discrepancy on Aspira Drainage Accessory Kits" dated June 12, 2011 to its customers. The letter provides the customers with the information missing from the labeling and/or IFU stating that the product can be safely used for peritoneal and pleural applications.If you have any questions please do not hesitate to contact BARD at 1-800-290-1689. If there are any clinical questions please contact Medical Services and Support at 1-866-893-2691. |  | 39290 | 39,290 units | Nationwide distribution. | Terminated | November 02, 2011 | 165 | 68530
Z-2957-2011 | 59352 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | HOLDEX Single-Use Holder PP closed pack. |  |  | June 29, 2011 | 2011 | Greiner Bio-One North America, Inc. | At removal of the tube from the holder, the needle may dislodge and blood spill can occur. | Other | Greiner Bio-One North America, Inc. sent a "Recall HOLDEX Single Use Holder PP" letter dated June 29, 2011 to all affected customers.The letter instructed customers stop distributing the product immediately. A Product Disposition form was provided to be completed and returned to Greiner Bio-One at which time upon receipt, arrangements for product return to Greiner Bio-One will be made for replacement or credit. They were requested to provide a list of their customers to Greiner Bio-One that received the affected item.For additional information call Greiner Bio-One at (888) 266-3883. |  | 142400 | 142,400 pieces | Nationwide Distribution including the states of CA, GA, IL, KS, MA, IL, PA and VA. | Terminated | November 03, 2011 | 127 | 142400
Z-2958-2011 | 59449 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015. Product Code: 2M8161. |  |  | July 22, 2011 | 2011 | Baxter Healthcare Corp. | Baxter's Service Center tested and returned to the user the Colleague infusion pump that did not meet PSIG (Pounds-force per Square Inch) specifications | Other | Baxter Healthcare Corp. notified the consignee by telephoned on July 22, 2011, with the call documented per the approved call script. A follow-up "URGENT PRODUCT RECALL" notification was sent dated July 25, 2011, reiterating the telephone communication informing them that the COLLEAGUE pump may not have met certain PSIG (Pounds-force per Square Inch) specifications after testing by the Baxter Global Technical Service Center. Both forms of notification identified the affected product, problem description, hazard involved, and actions needed to be taken. The consignee was instructed to locate the device, remove it from service and return it to Baxter by contacting Baxter Global Technical Services at 1-800-843-7867 option 1, to obtain a prepaid UPS label and shipping box. Upon receipt of the returned unit, Baxter will perform the required testing and address any issues through repair or service. The notication instructs the consignee to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter as per the instructions provided. If the recalled product have been distributed to another facility, or if you are a dealers, wholesalers and distributors, the letter request that this communication be forward to the pump owner as appropriate. For any questions regarding this communication contact the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central time. |  | 1 | 1 unit | (USA) Nationwide Distribution including the state of Minnesota. | Terminated | February 14, 2012 | 207 | 1
Z-2965-2011 | 59330 | No | LZS |  |  | Ophthalmic | PMA | STAR S4 IR Excimer Laser System; Software Version 5.00 |  |  | July 18, 2011 | 2011 | AMO Manufacturing USA, LLC | Reports of lower success rates with Iris Registration (IR) image capture with the STAR S4 IR Excimer Laser System v 5.3 that with previous version. | Software design | Abbott Medical Optics (AMO) sent an "ADVISORY NOTICE" dated July 18, 2011 to all customers using the AMO WAVESCAN Wavefront System. The letter identifies the product, problem, and actions to be taken by the customers.The letter instructs customers to create a separate patient file for each eye on the WAVESCAN System (STAR S4 IR Excimer Laser Systems,software versions 5.00-5.30) according to the directions in attachment A. Additionally, customers are to complete and return an attached form to AMO via fax at the regional fax number.Contact an AMO representative for questions regarding this notice. |  | 673 | 673 total v 5.00 - 5.22 | Worldwide Distribution-USA (nationwide) including Puerto Rico and countries of Africa, Asian Pacific, Canada, Europe, Japan, Latin America, and the Middle East. | Terminated | September 10, 2013 | 785 | 673
Z-2964-2011 | 59426 | No | HJB | Corneal radius measuring device. | Ophthalmic | Ophthalmic | 510(K) Exempt | Lenstar LS 900 EyeSuite Biometry version i2.000, Four (4) Protective Cardboard Containers Per Each System. |  |  | June 20, 2011 | 2011 | Haag-Streit USA Inc | A patient record may be linked to an incorrect/inconsistent set of calculation results under specific circumstances when using EyeSuite Software version i2.000. | Software change control | Haag-Streit USA, Inc. sent an "Urgent Field Safety Notices" dated June 20, 2011 to all affected customers.The letter included a description of the affected product and problem. The firm instructed their customers to make sure all new and current templates are saved and upgrade to EyeSuite version i4.000. Customers are asked to complete and return the Software Update Verification Form.For additional information contact the Haag-Streit at (513) 398-3937. |  | 216 | 216 | Nationwide (USA) Distribution. | Terminated | March 08, 2013 | 627 | 216
Z-2963-2011 | 59349 | [6 empty] | Heat Activated Niti Arches, Natural, Upper, .019x 25", Arch Wires. |  |  | July 12, 2011 | 2011 | Highland Metals | Heat Activated Niti, Natural, Upper, .019 x 25" Arch wires mixed with size 021x0.25" Arch Wires. | Mixed-up of materials/components | Highland Metals, Inc. notified their affected customers by e-mail and followed with an "Important Alert" letter dated July 12, 2011.The letters informed customers of the mix up of wires. Customers were asked to contact Highland Metals at 800-368-6484 to schedule a call tag to return any remaining product. Upon receipt credit will be issued.For additional information contact Highland Metals at (800) 448-7971 or (408) 271-2955. |  | 1000 | 10 packages of 100 | Worldwide Distribution; Nationwide (USA) distribution including the states of NV, CT, and FL; and the countries of Australia and Turkey. | Terminated | September 14, 2011 | 64 | 1000
Z-2981-2011 | 55941 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Aplio XG model SSA-790A |  |  | June 02, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems Inc (TAMS) is initiating a field correction on the the Aplio XG model SSA-790A because it did not meet the requirement of clause 15b, IEC60601-1 which requires that the residual voltage shall not exceed 60V after 1 second from taking off the AC plug. | Device Design | Toshiba America Medical Systems (TAMS) sent an "URGENT:MEDICAL DEVICE CORRECTION' letter dated June 2, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were instructed to be careful not to touch the contacts or prongs on the AC plug within two seconds after unplugging the system from a wall socket. This precaution is to be adhered to until a Toshiba Service Representative replaces the power supply. Customers are to complete and return a Customer Reply Form and fax to 877-349-3054.Customers with questions regarding the notification should contact the Director of Regulatory Affairs at (800) 421-1968 or call their local Toshiba Representative at (800) 521-1968. |  | 78 | 78 | Nationwide Distribution | Terminated | April 20, 2012 | 688 | 78
Z-2977-2011 | 59107 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE/RT Therapist v4.2 and/or Oncologist v4.2 systems. |  |  | June 03, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety rick related to transferring reference images for offset calculation between a syngo RT Therapist v 4.2 or a syngo RT Oncologist 4.2 and other versions of the COHERENCE/syn go RT Therapist or Oncologist or vice versa. An incorrect offset may be applied for patient positioning and result in dose being delivered to the wrong location. | Software design | Siemens sent an "URGENT: MEDICAL DEVICE CORRECTION/CUSTOMER SAFETY ADVISORY NOTICE" letter dated June 3, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter provides preventive measure instructions for customers to use until an update for the system is available. |  | 121 | 121 devices | Worldwide Distribution-USA (nationwide) including the states of CA, FL, NC, NY, OH, PA, SD, UT, and WI and countries of Australia, Belgium, Croatia, Czech Republic, Germany, India, Ireland, Japan, Poland, Saudi Arabia, South Africa, Spain, Sweden, and United Kingdom. | Terminated | March 30, 2012 | 301 | 121
Z-2980-2011 | 58918 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia Radiation Therapy Pallet. |  |  | June 22, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | There a potential patient safety issue when using the Symbia Radiation Therapy Pallet. When a patient is not iso-centered before starting a scan with the Symbia T Series Systems and a non-circular orbit is selected, it is possible that one of the detectors may come into contact with the patient. | Device Design | Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated June 06, 2011 to all affected customers.The letter informs customers of the potential safety issue if the patient is not iso-centered before starting a scan and a non-circular orbit is selected and provided recommendations including ensure the patient is iso-centered before starting a scan. If the patient cannot be positioned at the iso-center of the RTP, instructions were provided. The customers are asked to complete the Mandatory Fax-Back Form acknowledging receipt and understanding of the letter.For questions or assistance contact Siemens Medical Solutions USA, Inc. at (800) 888-7436 (USA). |  | 18 | 18 pallets | Worldwide Distribution: Nationwide (USA) including the states of Florida and Texas; and the countries of Canada, China, Italy, Mexico, Poland, Portugal and Russia. | Terminated | November 27, 2012 | 524 | 18
Z-2978-2011 | 59087 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac, Trilogy. Trilogy Tx and Novalis Tx Linear Accelerators. |  |  | May 26, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | On Varian Linear Accelerators where the customer may calibrate the collimator angle position readout in reverse, switching the 90 degree and 270 degree positions. | Device Design | Varian Medical Systems sent an "Urgent Medical Device Correction-Urgent Field Safety Notice" letter dated May 26, 2011 to all affected customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letters identifies the product, describes the problem and provides recommendations. Varian recommends maintenance only be performed by qualified personnel with appropriate training, always verify collimator calibration and anytime a collimator posthion readout calibration is performed, and always verify that the collimator angle is not calibrated in reverse by performing the appropriate checks. For questions contact Varian at (888) 827-4265. |  | 5685 | 5685 | Worldwide Distribution: Nationwide (USA) Distribution and the countries of; Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Belorussia, Bolivia, Bosnia and Herzegoniva, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Rep., Denmark, D.R. Ecuador, Egypt, Estonia, Finland, France, French Poly., Georgia, Germany, Greece, Guam, Guatemala, H.K., Honduras, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Macedonia, Malaysia, Martinique, Mexico, Moldova, Monte Negro, Morocco, Nepal, Netherlands, N.Z., Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, P.R., Qatar, Reunion, Romania, Russia, S.A.,Serbia, Singapore, Slovakia, South Africa, S.K., Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad/Tobago, Tunisia, Turkey, Ukrain, U.A.E., U.K., Uruguay, Venezuela, Vietnam, V.I., and Zimbabwe. | Terminated | July 13, 2012 | 414 | 5685
Z-2979-2011 | 59170 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Aera/Skyra System - Nuclear magnetic resonance imaging |  |  | June 20, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm discovered during product monitoring that on a few systems the gradient cable connections did not meet their specifications after installation. Locally, the cable connections overheated, making servicing necessary.The potential exists for a cable connection to become overheated, leading to the emission of smoke and possible material damage. | Nonconforming Material/Component | Siemens sent a "CUSTOMER SAFETY ADVISORY NOTICE" letter dated June 20, 2011 to all affected customers. The letter describes the product, problem, and actions that will be taken.The letter informs customers that an inspection of the gradient cable connections and an emergency shutdown mechanism will be installed by a field service representative during a visit to all affected locations. |  | 26 | 26 | Nationwide Distribution-USA (nationwide) including the states of AK, FL, GA, IL, MA, MI, MN, MT, NC, NJ, NY, OH, TX, UT, VA, WA, and WV. | Terminated | January 14, 2014 | 939 | 26
Z-2994-2011 | 58744 | No | LNR |  |  | Gastroenterology/Urology | PMA | XT125 PHOTOPHERESIS PROCEDURAL KIT FOR USE WITH THE UVAR XTS INSTRUMENT; XT001 PHOTOPHERESIS PROCEDURAL KIT FOR USE WITH THE UVAR XTS INSTRUMENT |  |  | March 31, 2011 | 2011 | Therakos, Inc. | Increased complaints were received of tubing leaks at the junction between the tubing and Fluid Logic Module plastic ports of the kits. | Device Design | The firm, Therakos Photopheresis, sent a "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated March 31, 2011 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to immediately discontinue using the affected products; deface and discard their remaining inventory of affected procedural kits, following their internal procedures for destroying and discarding of medical waste, and complete and return the attached Confirmation of Receipt Form no later than 08-April-2011 via fax to: Safety, Risk Management and Surveillance at 1-888-557-3759 or 1-585-453-4110. If you have any additional technical questions, please contact Therakos Customer Service Center at 1-800-828-6316. If you have technical questions regarding this communication, please call Therakos Technical Support at 1-877-865-6850. |  | 13239 | 13239 kits | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD,TN, TX, UT, VA, VT, WA, and WI; and countries of: Argentina, Austria, Australia, Belgium, Canada, Denmark, Estonia, Finland, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Korea, Kingdom of Saudi Arabia, Jordan, Netherlands, Norway, Philippines, Poland, Portugal, Russia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | March 18, 2015 | 1448 | 13239
Z-2990-2011 | 59042 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss Model DV5x Series CPAP with Heated Humidifier |  |  | March 11, 2011 | 2011 | Sunrise Medical HHG Inc. dba DeVilbiss Healthcare | Some units were manufactured with a defective humidifier chamber o-ring. Unit could leak, allowing water to escape during use, causing damage to furniture. | Nonconforming Material/Component | The firm, DeVilbiss Healthcare, issued a "URGENT: MEDICAL DEVICE CORRECTION" letter dated March 11, 2011 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to verify their inventory of CPAPs with Heated Humidifiers and isolate any units that are included in the list; replace the humidifier chamber o-rings with the new ones provided; discard the original o-ring removed from the humidifier chamber; complete and return the URGENT: MEDICAL DEVICE CORRECTION Confirmation letter via mail to Quality Manager, DeVilbiss Healthcare LLC, 100 DeVilbiss Drive, Somerset, PA, 15501; affirm that all units have been corrected and that all end users have been sent replacement parts, when all of the CPAPs on the serial number list have been corrected, and return a copy of the sign serial number sheet as confirmation.If you have any questions regarding the information provided in this communication, please contact DeVilbiss Healthcare Customer Service at 1-800-388-1988 from 8:00AM-5:00PM EST. |  | 210 | 210 | Worldwide distribution; USA (nationwide) including states of AL, AR, CA, FL, IL, IN, IA, KS, KY, MD, MA, MI, MS, NE, NJ, NY, OH, PA, WA, WV, and WI; and countries of Europe. | Terminated |  |  | 210
Z-2993-2011 | 59401 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Plastic Needle with Mandrin; 2.0mm diameter, length 320 mm. |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Needle tip may separate from the needle shaft such that the tip may detach in the patient upon withdrawal. | Device Design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated July 13, 2011 to all possibly affected customers. This notification supplements and modifies the previous "Customer Technical Bulletins" originally sent in September 2008. The initial Customer Technical Bulletins informed customers of the possibility that needle tips cans crack and separate. It also reminded users that the product has a limited life and to check needles prior to use as per instructions. Current correction notice letter instruct customers to stop using any plastic needles in inventory with lot numbers C01 to D01 and those with lot numbers starting with H or I. The letter also request that the recalled products be returned to Varian BrachyTherapy for replacement. The letters include information as to where the lot number is located on the device. All appropriate personnel should be notified of the correction notice. Contact information is provided for all geographical areas. USA contact for questions or concerns is 1-800-360-7909. |  | 601 | 601 - all sizes. | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Austria, Bangladesh, Belgium, Brazil, Canada, Columbia, Cyprus, Czech Republic Denmark, Dominican Republic, Estonia, France, Germany, Greece, Hungary, India, Japan, Kazahastan, Kuwait, Lebanon, Macedonia, Mexico, Moldova, Nepal New Zealand, Norway, Pakistan, Philippines, Poland, Russia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain Sweden, Switzerland, Taiwan, Thailand, Trinidad, Tobago, Turkey, Ukraine, United Arab Emirates Uruguay, United Kingdom, and Venezuela. | Terminated | June 20, 2012 | 343 | 601
Z-2989-2011 | 59362 | No | JAD | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | INTRABEAM Shielding Material Flat 20 x 20 cm, |  |  | July 01, 2011 | 2011 | Carl Zeiss Meditec, Inc. | Carl Zeiss Meditec received a report that the INTRABEAM 20 x 20 cm sterile shield's packaging looked wet inside. Investigation found the substance to be Triethylene Glycol, | Process control | Carl Zeiss Meditec, Inc. sent an "Important Field Correction Action" letter dated July 13, 2011 to all affected customers.The letter included description of affected product and problem. Customers are asked to notify the firm if they have experienced any problems with the product and to return any remaining product for replacement upon completion of required testing. Customers are asked to fill out and return an Acknowledgement Form. For additional information contact Carl Zeiss Meditec at (925) 570-4844. |  | 80 | 8 boxes of 10 sterile packages. | Nationwide (USA) Distribution including the states of FL, CA, NY, VA and AL. | Terminated | October 13, 2011 | 104 | 80
Z-2983-2011 | 59347 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | LANTIS Oncology Information System (OIS) |  |  | June 30, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety issue exists for data transfer between LANTIS OIS system (containing PRIMEVIEW or PRIMEVIEW Data Gateway) and the Sequencer V&R; system via a third party interface is used. | Software design | Siemens sent an "Update Instructions/Customer Safety Advisory Notice" to all affected users, beginning June 30, 2011.The letter includes description of the product and problem and provides updated instructions to ensure data is stored in the system.If you have questions contact Siemens at (925) 602-8157. |  | 59 | 59 devices | Worldwide Distribution: Nationwide (USA) including the states of MI, DE, WI, FL, CA, TN, NJ, PA, NE, IL, WV , MA and VA; and the countries of Poland, France, Germany, Canada, Netherlands, India, Ireland, Egypt, P.R. china, Australia, Saudi Arabia, Turkey, Republic Korea, Belgium, Hungary, New Zealand, Italy, Malaysia, and Spain. | Terminated | March 09, 2012 | 253 | 59
Z-2987-2011 | 59367 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Direct Drive Grasper, Model number C4140, 5MM X 45CM DIRECT DRIVE |  |  | June 17, 2011 | 2011 | Applied Medical Resources Corp | The recall was initiated because Applied Medical has confirmed specific lot numbers of the DirectDriveÂż Grasper (model numbers C4130 and C4140) due to a potential weakness in the locking trigger, which could result in the trigger fracturing and the jaw locking in the closed position. | Device Design | Applied Medical sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 22, to all affected consignees. The letter described the product, problem and actions to be taken. Consignees were instructed to check their inventory for the recalled products and returned immediately as per the instructions. The letter instructs consignees to complete the attached Voluntary Recall Notification Confirmation Form and fax it to Regulatory Affairs at 949-713-8832 or email it to recall60342148@appliedmedical.com. For questions pertaining to product return contact Chris Jacobson, Sales Operations Administrator, at (949) 713-8652 or by email at cjacobson@appliedmedical.com. For regulatory questions contact Susan Fehrenbacher at (949) 713-8041 or by email at sfehrenbacker@appliedmedical.com. |  | 38864 | 38,864 units total (24,868 units in US) | Worldwide Distribution - USA (nationwide) including the state of Puerto Rico and the countries of: Australia, Bahrain, Canada, Chile, Colombia, Ecuador, Europe, India, Israel, Jordon, Kuwait, Lebanon, Libya, Malaysia, Mexico, Peru, Qatar, Singapore, , South Africa, South Korea, Saudi Arabia. Taiwan, Thailand, Turkey, and Venezuela. | Terminated | March 26, 2012 | 283 | 38864
Z-2985-2011 | 59221 | [6 empty] | ATLANTIS(R) Anterior Cervical Plate System Screw Caddy; Self Drilling Fixed Screw Caddy (976-516). |  |  | July 06, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | Lines on the screw gauge may not placed on the caddy properly which could result in improper measurement of screws. | Process control | Medtronic Sofamor Danek USA, Inc. sent "Voluntary Recall - Medtronic" letters on July, 2011 to all affected customers.The letter included; description of product and problem, and asked customers to stop using the product. Customers are asked to contact their Sales Representatives for replacement.For question on this recall call Medtronics Global Quality Department at (800) 876-3133. |  | 191 | 191 units | Worldwide Distribution: Nationwide distribution including the states of MN, OK, IL, AK, CA, PA, UT, LA, SC, NY, MI, GA, WI, KY, AL, MO, AZ, TN, TX, MA, NE, OH, FL, CO, MT, CT, VA, NC, NH, WA, WV, WY, IN, IA, KS, MT, ME, DE, DC, NJ, MD, MS, NV, HI, ID, ND, and AR; and the country of Uruguay. | Terminated | April 26, 2013 | 660 | 1064
Z-2982-2011 | 59386 | [6 empty] | Biopsy Trays (CT, Puncture &amp; Prep), Packaged individually in header bags, 10 per case. |  |  | May 13, 2011 | 2011 | North Coast Medi-Tek Inc | During routine bioburden testing, unusually high results were found for a Bioburden Test done on 4/19/11. Bioburden Level reported as &gt;150,000 CFU. Our Alert Level is 17,000 CFU and the Action Level is 23,000 CFU. | Other | North Coast Medi - Tek Inc. sent a URGENT PRODUCT RECALL letter dated May 13, 2011 to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.The customers were instructed to stop shipment of all remaining product and return to recalling firm. They were also requested to contact their customers and notify them of the recall and retieve all remaining product.For further questions please call (440) 974-0750. |  | 2190 | 2,190 trays | Nationwide Distribution CA, CO, FL, GA, ID, IL, MA, NC, NE, NY, OH, OK, PA, TN, NC and TX | Terminated | February 16, 2012 | 279 | 2190
Z-2976-2011 | 59296 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper Systems Models: |  |  | July 08, 2011 | 2011 | Philips Healthcare Inc. | System freezes from 20 seconds up to one minute. No fluoroscopy exposure and no geometry movements during this time period. | Software design | Philips Health Care North America sent a FIELD SAFETY CORRECTION ACTION letter dated July 4, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.This FCD will be identified as FCD72200200 and will be free of charge. This FCD consists a CD with the correct firmware and a installation procedure to implement the FCD correctly.Please contact your local Philips representative. For further follow up please call (978 ) 687-1501. |  | 1423 | 1423 units | Nationwide Distribution | Terminated | May 14, 2014 | 1041 | 1423
Z-3001-2011 | 58322 | [6 empty] | Micro-Fuel Cell Oxygen Sensor, Class RI5 |  |  | November 29, 2010 | 2010 | Teledyne Instruments - Analytical Instruments | The recall was initiated because during a test on a batch of sensors on the company shipping dock, Teledyne Instruments discovered electrolyte loss from the sealing cylinder on a few sensors. This test was done pursuant to a complaint from one customer of label discoloration caused because of the electrolyte loss. The riskof a leak form the oxygen sensor while in use is the disruption of therape | Process design | Teledyne Instruments sent emails and "Urgent--Device Field Corrections/Removal" letters in December 2010 to all affected customers.Th letter includes description of device and problem. Customers are advised to check them before use for any signs of electrolyte loss and if discoloration is noted to discontinue its use. and provides recommendation. Teledyne-desplaines to generate a product field corrections/removal log that lists names/addresses of the consignees, quantity of sensors shipped, sensor serial numbers and the date of shipment.For additional information or questions contact Teledyne Instruments at (626) 934-1504. |  | 2813 | 2813 | Worldwide Distribution: Nationwide (US) distribution including the states of CA, TN and WI; and the countries of , Canada, Brazil, Chile, Saudi Arabia, and the U.K. | Terminated | August 22, 2011 | 266 | 2813
Z-3011-2011 | 59112 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker EMS Stair - Pro Evacuation Chair Model 6252 |  |  | June 28, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | This recall was initiated as the firm ascertained the supplier may have reworked some of the red release bars used during the time frame of 22 December 2010 through 31 January 2011. The rework subjected the parts to excessive heat and as a result, the bars can bend during the action of squeezing the release bar against the black cross tube to release the track. The red release bars may bend to | Nonconforming Material/Component | The firm, Stryker Medical, sent an "Urgent-Medical Device Recall" letter dated June 28, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate the listed product; inspect each unit prior to use; return the enclosed post card to confirm receipt of this notification; if they have loaned or sold any of the chairs listed, forward a copy this letter to the new users and advise Stryker of new users names and locations, and if they have disposed of any of the affected chairs and they are no longer in use, advise Stryker of their obsolescence by providing them with the serial numbers. Stryker stated in the letter that, their goal is the timely upgrade of all affected units within the next 90 days.If you have any questions or concerns, please contact us at 800-STRYKER, option 2. Please reference RA-2011-027. Normal business hours are Monday-Friday 8am-5pm EST. |  | 214 | 214 | Worldwide distribution: USA (nationwide) and countries of: Canada, China, France, Germany, India, Netherlands, and Sweden. | Terminated | August 29, 2013 | 793 | 434
Z-3003-2011 | 59122 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE Linear Accelerator with RT Therapist v4.1 system. |  |  | June 03, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety issue was discovered when a pause for an imaging segment is inserted (where an imaging segment does not exist) with the ARTISTE and RT Therapist 4.1 system will continue to the next image. | Software design | Siemens sent an "Urgent: Medical Device Correction/ Customer Safety Advisory Notice" letter to all affected customers starting June 3, 2011. The letter describes the product, problem, and instructions to prevent mistreatment. The letter states that Siemens is working on an updated version of RTT 4.1 that will fix the issues. Users will be notified as soon as the update for the system is available. The customers were instructed to include Customer Safety Advisory notice in their syngo RT Therapist 4.1 System Owner Manual chapter "Safety Advisory Letters" where it should remain. |  | 223 | 223 active devices | Worldwide Distribution - USA (nationwide) including the states of: DE, FL, KY, LA, MA, MI, MO, NE, NJ, OH, PA, TN, WI, UT, WV and Puerto Rico and the countries of: Australia, Belgium, Brazil, Canada, China, Egypt, France, Germany, Hungary, India, Ireland, Italy, Lebanon, Malaysia, Netherlands, New Zealand, Philippines, Poland, Republic of Korea, Russian Fed, Saudi Arabia, South Africa, Norway, Spain, Turkey and United Kingdom. | Terminated | February 17, 2012 | 259 | 223
Z-2996-2011 | 59404 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Titanium Fletcher-style Applicator Set - Defined Geometry - CT Compatible with Titanium intrauterine probe with stopper length 60 mm, 30 degree angle. |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Potential of weld failure of the Titanium intrauterine probe at the cervical stopper. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated July 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to return all affected product to Varian Brachy Therapy by using the contact information provided in the letter. Replacement product will be provided free of charge to the customers.Customers should contact their local Varian Customer Support District or Regional Manager for questions regarding this notice. |  | 437 | 437 - all sizes | Worldwide Distribution. | Terminated | June 20, 2012 | 343 | 437
Z-3002-2011 | 59378 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System; |  |  | June 30, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Eclipse versions, the treatment field energy may change from the current selection to a previous selection without use intervention and without invalidating the MU or calculated dose distribution. | Software design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 30, 2011 to all affected consignees. The letter identifies the affected product, problem, and actions to be taken. The letter states that Varian will contact all affected customers to apply a database configuration script to the database and is preparing a correction for Eclipse Version 10 to address the issue. A Varian representative will contact each site to schedule correction when it is available. The letter also provides RECOMMENDED USER ACTIONS to follow. The Letters go on to reiterate caution and warnings from Eclipse instructions for use. All appropriate personnel should be notified of the correction notice. For future reference, this document will be posted to the Varian customer support website: http://www.MyVarian.com. For further clarification, contact your local Varian Customer Support District or Regional Manager. |  | 2436 | 2436 | Worldwide Distribution. | Terminated | May 09, 2014 | 1044 | 2436
Z-3008-2011 | 59445 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Deerfield, IL; product code 2M8161 |  |  | July 22, 2011 | 2011 | Baxter Healthcare Corp. | A review of Baxter service records indicates that the Baxter Colleague CX Infusion Pump has not met certain voltage specifications. This issue could potentially cause a delay in initiation of therapy. | Other | The firm, Baxter, called by telephone on July 22, 2011, and sent a follow-up "Urgent Product Recall" letter dated July 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate and take the affected pump out of service; call Baxter Global Technical Services at 1-800-THE-PUMP (1-800-843-7867), Option 1, between the hours of 6:00 AM and 7:00 PM Central Time, Monday through Friday, if arrangements have not yet been made to return the pump, and promptly complete and return the CUSTOMER REPLY FORM via fax to 847-270-5457, or scan and email to fca@baxter.com.If you have any questions regarding this communication, please call the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central Time. |  | 1 | 1 pump | Nationwide distribution: Missouri only. | Terminated | October 19, 2011 | 89 | 1
Z-3007-2011 | 59448 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pump. Baxter Healthcare Corporation, Deerfield, IL. Product Code: 2M8161. |  |  | July 22, 2011 | 2011 | Baxter Healthcare Corp. | The Colleague infusion pump may not have met certain Air in Line (AIL) specifications after testing by the Baxter Global Technical Service Center. This issue could potentially cause a delay in initiation of therapy. | Process design | The firm, Baxter, called by telephone on July 22, 2011, and sent a follow-up "Urgent Product Recall" letter dated July 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate and take the affected pump out of service; call Baxter Global Technical Services at 1-800-THE-PUMP (1-800-843-7867), Option 1, between the hours of 6:00 AM and 7:00 PM Central Time, Monday through Friday, if arrangements have not yet been made to return the pump; promptly complete and return the CUSTOMER REPLY FORM via fax to 847-270-5457, or scan and email to fca@baxter.com, and forward this communication to the pump owner as appropriate, if they distributed the pump to another facility or if they are dealers, wholesalers, or distributors/resellers.The letter states "Upon receipt of the returned device, Baxter will perform required testing on the affected COLLEAGUE pump for which you have indicated there is continuing medical necessity. Should Baxter find any issues during testing, Baxter will address those issues through repair or service prior to returning the pump to you at no charge".If you have any questions regarding this communication, please call the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central Time. |  | 1 | 1 Pump | Nationwide distribution: MN only. | Terminated | November 09, 2011 | 110 | 1
Z-3009-2011 | 59275 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Artis U mobile x-ray system |  |  | June 29, 2011 | 2011 | Siemens Medical Solutions USA, Inc | When the heat unit of the angiography system reaches or exceeds 65%, the cool down time display will appear with an acoustic signal. If users ignore the warning and continue the radiation, the tube may be autoprotected when it reaches its temperature limitation causing the X-Ray to lock until the tube cools down. Following a tube auto-protection, an error message "Door is Open" or "Err 421/240: | Labeling False and Misleading | Siemens sent a "Customer Safety Advisory Notice" letters to all affected users dated June 13, 2011. The letter identified the product, the problem, and the actions to be taken. Siemens included in the notice an addendum to the instruction for use to each system and request customers to take special attention. Customers were instructed to place the Safety Notice in the System's instruction for use. All appropriate personnel should be notified of the correction notice. If product was distributed further, customers were to forward the safety notice to the new owner and inform the firm of the new owner. |  | 18 | 18 | Nationwide Distribution (USA) including the states of: CA, CT, FL, KS, MN, MO, NJ, NY, NC, PA, and UT. | Terminated | January 24, 2014 | 940 | 18
Z-3005-2011 | 59416 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Surgical Clamp Inserts, Model#G-8650, 86MM Insert Soft/Fibra 15/Bx |  |  | March 07, 2011 | 2011 | Applied Medical Resources Corp | The recall was initiated because Applied Medical has confirmed that Surgical Clamp Inserts, Model Numbers: G-8600, G-8650, G-8655, has the potential packaging nonconformance which could result in a compromise to the sterile barrier. | Other | The firm, Applied Medical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 11, 2011 to its customers. The letter provides customers with an explanation of the problem identified and actions to be taken. The customers were instructed to check their inventory; complete and return the attached Voluntary Recall Notification Confirmation Form via fax to Regulatory Affairs at 949-713-8832, and return the recalled Standard 86mm Inserts immediately to Applied Medical, Attn: RGA# 60308783, 22872 Avenida Empresa, Rancho Santa Margarita, CA 92688.Customers with product return questions, were instructed to contact the Sales Operations Administrator, at (949) 713-8652 or by email at cjacobson@appliedmedical.com. Customers with regulatory questions, were instructed to contact the Director, Regulatory Affairs, at (949) 713-8041 or by e-mail at sfehrenbacher@appliedmedical.com. |  | 20247 | 20,247 units total | Worldwide distribution: USA (nationwide) including states of: Arizona, California, Colorado, Florida, Georgia, Hawaii, Iowa, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, and Michigan; and countries of: Japan, Israel, and Singapore. | Terminated | March 26, 2012 | 385 | 20247
Z-3018-2011 | 57583 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | i-CAT Classic (X-Ray Tube head and Power Supply) |  |  | December 15, 2010 | 2010 | Imaging Sciences International, LLC | Recalling firm received a customer complaint regarding a defective X-ray tube head and power supply | Other | The firm, Imaging Sciences, sent an "Urgent Medical Device Field Correction" letter dated December 15, 2011 via first class USPS to its customer. The letter noted the product and described the problem and action to be taken. The customer was instructed to complete and return the ACKNOWLEDGEMENT FORM via fax to 215-997-5665 or email: darin.oppenheimer@imagingsciences.com. Imaging Sciences stated in the letter that an X-ray tube head and power supply will be provided and installed by one of their technicians. Should you have any questions regarding this information, call 215-622-1545. |  | 1 | 1 | X-ray tube and power supply were distributed to one consignee: Dr. Alan D'Allessandro, 2500 W. Higgins Road, Suite 665, Hoffman Estates, IL 60169 | Terminated | March 01, 2013 | 807 | 1
Z-3014-2011 | 59463 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Metha(R) Hip System Cobalt Chrome (CoCr) Metha Modular Neck Adapters used with the Metha Modular Short Hip Stems. |  |  | July 19, 2011 | 2011 | Aesculap Inc | Aesculap Implant Systems received a notification from the manufacturer, Aesculap, AG of a fractured CoCr neck adapter. | Unknown/Undetermined by firm | AESCULAP Implant Systems sent an "Important Recall Notification" letter dated to all affected customers.The informed customers of a complaint received where fractured CoCR adapter was found and to return any affected product to the firm. The reason for the fracture is currently under investigation. While the firm is trying to determine the root cause of the fracture, they will be offering the Metha non-modular short hip implants to their customers. Customers are asked to return an Inventory Sheet to AESCULAP.For additional information on this recall call (610) 984-9239 or (610) 984-9300. |  | 149 | 149 | Nationwide (USA) Distribution including the states of AL, FL, GA, MD, NM, NY, OK, and TN. | Terminated | March 19, 2013 | 609 | 149
Z-3012-2011 | 59059 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage 2.1 and 2.2 Workstations with RIS CCOW integration, GE Healthcare 3000 North Grandview, Waukesha, WI 53188 |  |  | May 31, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware that patient mismatch could occur on some Seno Advantage 2.1 and 2.2 Workstations with RIS CCOW integration. This issue may impact patient safety.When Seno Advantage workstation is synchronized with RIS using CCOW (an HL7 standard protocol designed to enable disparate applications to synchronize in real-time), the following issues could occur: 1.Two different | Software design | GE Healthcare sent a URGENT MEDICAL DEVICE CORRECTION letter to all affected customers dated May 31, 2011. The letter identified the product , the problem and the action needed to be taken by the customer.GE Healthcare will upgrade all affected workstations at no cost. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative |  | 391 | 391 | Worldwide Distribution--USA (nationwide) including the states of: AZ, CA, CO, CT, FL, GA, HI, IL, IN, IA, LA, MD, MA, MI, MN, MO, MT, NE, NJ, NY, NC, OK, OR, PA, SC, TN,TX, and WI, and the countries of UNITED KINGDOM, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, PORTUGAL NORWAY, NEW ZEALAND, LUXEMBOURG, LEBANON, KOREA, JAPAN, ITALY, IRELAND, GERMANY, FRANCE, DENMARK, CANADA, BELGIUM AUSTRIA, AUSTRALIA FINLAND. | Terminated | June 29, 2012 | 395 | 391
Z-3017-2011 | 59517 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas 2020 Insulin Infusion Pump. |  |  | July 29, 2011 | 2011 | Animas Corporation | Incorrect component (1nF capacitor) was used to produce 191 printed circuit boards used in six pump lots of which 380 pumps were distributed. This issue may result in the pump displaying a "Replace Battery" alarm not related to the condition of the battery. | Mixed-up of materials/components | Animas Corporation sent an "Urgent: Insulin Pump Recall" letter dated July, 29 20011 to all affected customers.The letter included affected product and problem. They requested anyone who received the pump to send it back to Animas and a replacement pump will be sent at no charge.For additional information contact the Customer Support Department at (877) 937-7867. |  | 380 | 380 | Worldwide Distribution: Nationwide (USA) distribution including the states of CA, CO, MA, and MD; and the countries of Australia, Finland, France, Israel, Italy, Sweden and the U.K. | Terminated | December 07, 2011 | 131 | 380
Z-3013-2011 | 59388 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Flexible Applicator probe; 3.2mm diameter, braced PVDF. |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Certain GammaMed Flexible applicator probes may not have been heat annealed during manufacturing to prevent shrinkage after autoclave. | Process control | Varian Medical Systems, Inc. sent out Customer Technical bulletins in August 2008 and Urgent Medical Device Correction/Urgent Field Safety Notices on July 13, 2011 to all affected customers.The letter included description of device and problem. The current notice supplements and modifies previous instructions and ask that they stop using any flexible applicator probes in inventory. They are to be returned to Varian Brachy Therapy for replacement. Letter go on to list possible scenarios if the length deviation is not detected and advised all personnel working in the radiotherapy department should be notified of the recall. For additional information contact (800) 360-7909. |  | 1343 | 1343 | Worldwide Distribution: Nationwide (USA) and including the countries of France, Germany, India, Trinidad and Tobago, United Kingdom, Canada, Nepal, Ukraine, Czech Republic, Bangladesh, Thailand, Turkey, Brazil, Estonia, Switzerland, Russia, Greece, Philippines, Spain, Belgium, Kuwait, Romania, Morocco, Mexico, Columbia, Austria, Poland, Kazahstan, Vietnam, Cyprus, New Zealand, Dominican Republic, Norway, Uzebestan, Hungary, Lebanon, South Africa, South Korea, Slovenia, Venezuela, Serbia, Tunisia, Pakistan, Macedonia, Ireland, Azerbajian, Ukraine, Morroco, Italy, Singapore, Nambia, Turkey, Sudan, and Taiwan. | Terminated | June 20, 2012 | 343 | 1343
Z-3016-2011 | 59376 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | AirLife Nebulizer Heater, Catalog Code 2M8021 |  |  | July 13, 2011 | 2011 | Carefusion 211 Inc | The recall was initiated because CareFusion's Investigation has confirmed the Nebulizer Heater can, over time, develop a failure mode that results in an internal electrical short. The short can result in smoke and sparks being emitted from the device. In the presence of oxygen enriched environments and a fuel source, this failure represents a potentialrisk of flame. CareFusion is voluntarily rec | Other | Carefusion sent a URGENT PRODUCT RECALL letter dated July 15, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to examine their inventory for the affected product as indicated on the sample shown in the enclosed Attachment A-1 and remove any affected heaters present in their facility. Confirm the total number of affected units in their possession and enter the quantity and respective serial number on the enclosed Customer Response Card (Attachment A-2). Customers were instructed to contact CareFusion's Technical Support Department at 800-554-8933, (6:30 am to 5:00 pm PDST), to obtain an RMA (returned materials authorization) number for the return of the affected product. Upon receipt of their response, CareFusion will ship to their customers new replacement Nebulizer Heaters in exchange for the affected units.Customers were instructed to acknowledge receipt of this communication by completing the enclosed Customer Response Card and faxing or sending a copy by return mail to the address indicated on the Customer Response Card, (Attachment A-2). |  | 22670 | 22,670 units total | Nationwide Distribution | Terminated | April 30, 2012 | 292 | 22670
Z-3015-2011 | 59398 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Fluoroscan Mini C-arm, Models: FSIII, FSIIIC and Officemate |  |  | July 21, 2011 | 2011 | Hologic, Inc. | The Flex arm connector to C-arm can break and the C-arm that is supported by the flex arm may fall down | Nonconforming Material/Component | The firm, Hologic, Inc., sent an "Urgent: Medical Device Recall" letter dated July 25, 2011 via FedEx to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to visually check the the Flex arm/C-arm on their unit for any sign of damage or cracking, if any damage or cracks are observed, remove the unit from service until they receive the retrofit kit; and to complete and return the Confirmation of Notification form via fax to Hologic at 781-280-0662 or mail using the enclosed pre-addressed, pre-paid return envelope. After the customer confirms that they have the unit in their possession, Hologic will expedite the user installable repair kit. The kits will be available starting the first week of August. The kits will also be dispatched via FedEx. Confirmation of installation will be requested for all dispatched kits.Should you have any questions regarding this recall, require further information or assistance with the software installation call Hologic's Help Desk at 1-800-321-4659. |  | 540 | 540 units | Worldwide distribution: USA (nationwide) and countries of: Argentina, Spain, and Viet Nam. | Terminated | June 19, 2012 | 334 | 540
Z-3019-2011 | 57543 | No | IXQ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | CIVCO Pro-Lok Platform, REF MTSBRT002. Firm name on label is CIVCO, Orange City, Iowa |  |  | December 30, 2010 | 2010 | Civco Medical Instruments Inc | Systems sold before December 2009 displayed a scale of 0-40cm going from left to right, while systems after December 2009 reversed this and were displayed as 40-0cm from left to right. | Nonconforming Material/Component | CIVCO issued a "FIELD SAFETY NOTIFICATION" letter dated 12/20/2010 beginning 12/30/2010, via FedEx or UPS to distributors and customers. The notification identifies the affected device, problem, and field safety corrective action. The letter states that if customers have more than one system purchased from different time periods that they should examine their systems and assess if there is any mismatch. The recalling firm will rework the system to correct the difference and send it back to the customer. To take advantage of this correction customers are advised to contact the recalling firm within thirty days of receiving the safety notification. The letter includes a response form to be completed and returned as per instructions, to indicate that the notice has been read and understood by all operators of the equipment. The letter to distributors requested that a copy of the notification letter be forward to all customers who have received this component. For assistance contact your CIVCO insides sales representative by calling CIVCO Medical Solutions, Radiation Oncology, at 800-842-8688. |  | 20 | 20 platforms | Worldwide Distribution - USA ( nationwide) including the state of PA. and countries of: Canada, China, England, Israel, Italy, Japan, Korea, The Netherlands, Russia, Spain, Sweden, Switzerland and Turkey. | Terminated | June 28, 2013 | 911 | 134
Z-3024-2011 | 56839 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical elance Vital Signs Monitor, Model 93300. |  |  | September 23, 2010 | 2010 | Spacelabs Healthcare, Llc | The monitor has a 3 second delay in the waveforms and respiration alarm; the delay is not disclosed in the IFU. | Error in labeling | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated September 24, 2010 to all affected customers. The letter describes the product, problem, and the action to be taken.The letter includes the revised Operator's Manual as a replacement for customers current manual.For Technical Support call Spacelabs Medical at 1-800-522-7025 and select 2. |  | 2517 | 2517 | Worldwide Distribution-USA (Puerto Rico) and countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, FRANCE, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, KENYA, KOREA, REPUBLIC OF, KUWAIT, LEBANON, MACEDONIA, MALAYSIA, NEPAL, NETHERLANDS, NICARAGUA, OMAN, PALESTINE, PANAMA, PARAGUAY, PERU, PUERTO RICO, PHILIPPINES, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SRI LANKA, SWITZERLAND, SYRIAN ARAB REPUBLIC, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Terminated | August 17, 2011 | 328 | 2517
Z-3023-2011 | 59235 | No | CGI | Estriol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AccessÂż Immunoassay Systems |  |  | December 08, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports that a single lot of Access Unconjugated Estriol reagent (PIN 33570) may produce suppressed patient results and an increased incidence of IND (Indeterminate) flags. These issues have been associated with reagent lot 913348. This reagent met all testing specifications prior to release. | Other | Beckman Coulter sent a PRODUCT CORRECTIVE ACTION (PCA) letter dated December 8, 2009, to all affected customers. The letter was sent via US Postal Service for all US customers with attached PCA response form. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to discontinue use of Access Unconjugated Estriol reagent lot number 913348. " No further orders of Access Unconjugated Estriol reagent lot number 913348 will be shipped." Abnormally low Access Unconjugated Estriol results from samples tested previously with the affected lot should be reviewed in the context of other diagnostic tests." Access Unconjugated Estriol results should be interpreted in light of the total clinical presentation of the patient, including: symptoms, clinical history, data from additional tests, and other appropriate information. Rerun any tests that produce a questionable result.In addition, customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with any technical questions regarding this Product Corrective Action were instructed to call Beckman Coulter Customer Service at (800) 854-3633 in the US and Canada. Outside of Canada and the US, customers were instructed to contact their local Beckman Coulter Representative. |  | 2786 | 2786 units total (2776 in the US) | Worldwide Distribution -- USA (nationwide) including the states of CA, CT, MA, NC, NM and TX and the country of Canada. | Terminated | April 16, 2012 | 860 | 2786
Z-3025-2011 | 59661 | No | HDC | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | GYNDisposables Inc, IUD INSERTION KIT 935K, STERILE, SINGLE USE, Tenaculum 356T, Uterine Sound 906S, Scissors, Povidone-Iodine Solution, (3) Swabsticks, Paper Towel, STERILE EO, REF 935K, QTY 1, |  |  | July 29, 2011 | 2011 | GYN Disposables Inc | Lack of sterility assurance. | Process design | GYN Disposables Inc. sent out an "Urgent: Medical Device Recall" letter dated July 29, 2011 to all affected customersThe letter includes: product affected, description of problem and asks customers quarantine all affected product. They can contact GYN Disposables to have the product picked up and to receive credit. A Response Form is also included.For questions on this recall call (888) 372-2203, extension 101. |  | 3987 | 3987 units | Nationwide Distribution. | Terminated | October 29, 2012 | 458 | 3987
Z-3022-2011 | 59394 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, 4D Structure Merge Tool Option in Eclipse Release 8.5 -10 (Client Build 8.2.2x through 10.0.34) with 4D Planning; Model Number: H48; |  |  | July 07, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified in the 4D Structure Merge option of certain versions of Varian's Eclipse Treatment Planning System. The treatment field energy may change from the current selection to a previous selection without user intervention and without invalidating the MU or calculated dose distribution. | Software design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/FIELD SAFETY NOTICE" letter dated July 7, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to cease using the 4D Structure Merge tool option in the Eclipse until further notice. Additionally, it is recommended that customers ensure that a qualified physician reviews the accuracy and placement of all patient structures used for plan development and evaluation and to conduct standard quality assurance procedures prior to patient treatment. Customers will be contacted by a Customer Service representative when the technical correction for this issue is available.Contact your local Varian Customer Support District or Regional Manager if you have questions regarding this notice. |  |  |  | Worldwide Distribution | Terminated | April 04, 2014 | 1002
Z-3028-2011 | 59218 | No | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | BD Vacutainer Push Button Blood Collection Set; REF 367324 Rx only, 23 G x 3/4" x 12 " (0.6 x 19 mm x 305 mm); |  |  | April 22, 2011 | 2011 | Becton Dickinson &amp; Company | BD Vacutainer Push Button Blood Collection Set can potentially undergo cannula separation from the hub. | Device Design | BD Diagnostics sent a "Lot Specific Product Recall" letter dated April 22, 2011 to all affected customers.The letter asked customers return any affected product and provides instructions for its return. The letter ask customers return the attached Packaging List with the shipment.For question on this recall contact BD Diagnostics at (201) 847-5117. |  | 1744000 | 1,744,000 | Nationwide (USA) Distribution. | Terminated | June 06, 2012 | 411 | 1778960
Z-3027-2011 | 59239 | No | KLI | Total thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AccessÂż Immunoassay Systems Total T4 Calibrators, |  |  | August 26, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the stability claims for the Access Total T4 calibrators may not be met. This issue is associated with calibrator lots 917074 and 008534. | Nonconforming Material/Component | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated September 13, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue use of the affected lot numbers and review historical quality control performance. Additionally, a response form was attached to the letter for customers to complete and return. Alternate lots will be provided for current and future replenishment orders or requests.Customers with questions regarding this notification can call Technical Support at (800) 854-3633 or contact their local Beckman Coulter representative. |  | 2495 | 2495 units total (2492 in US) | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | July 11, 2012 | 685 | 2495
Z-3026-2011 | 59408 | No | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | AlterG Anti-Gravity Treadmill; |  |  | July 15, 2011 | 2011 | Alter-G, Incorporated | Potential installation error on AlterG Anti-Gravity Treadmill that could cause greater than expected unweighting. | Device Design | Alter-G, Inc. notified all affected customers by telephone, e-mail and followed by a "Medical Device Customer Advisory Notification" letter dated July 15, 2011. Customers were provided information on the installation error, the potential harm to patients, and the actions to be taken. The letter instructs customers to discontinue use of the Alter G until the recalling firm can evaluate the machine and correct any installation error. The letter states that the recalling firm is planning to visit affected customer's facility as quickly as possible and will contact customers to schedule this visit. Questions or concerns should be directed to David Tjon at 510-270-5369. |  | 16 | 16 potentially affected units | Nationwide Distribution (USA ) - including the states of: OK and TX. | Terminated | August 16, 2011 | 32 | 16
Z-3032-2011 | 59456 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Patient Monitor is a Physiological, Patient monitor With Arrhythmia Detection or Alarms. The Models are 91367, 91369, 91370, 91387, &amp; 91393 are sold with with Perioperative Mode (option D), which includes Start case/End case functionality. |  |  | July 08, 2011 | 2011 | Spacelabs Healthcare, Llc | The clinician did not realize the monitor was in END CASE mode and was expecting audible alarms. Visual alarms continue to work normally. | Device Design | Spacelabs sent an "URGENT - MEDICAL DEVICE CORRECTION" letter dated July 25, 2011 along with the Perioperative section of the Operations Manual to all consignees in the US. This letter includes recalled model 91369, 91370, 91387, &amp; 91393 (option D). On July 27, 2011, Spacelabs sent out an updated recall notification letter to include one additional model 91367 to all consignees in the US. The letter identifies the products, problem, and actions to be taken. On August 2, 2011, a recall notification letter (translated as necessary) and the Perioperative section of the Operations Manual was emailed to all international subsidiaries and distributors of record. The notification letter stated that the Operating Room staff may not be aware that alarm tones are off if the device has been left in the END CASE state. If the customers make use of the Perioperative START CASE / END CASE feature, they should advise the clinicians to familiarize themselves with this feature as described in the Perioperative section of the Operations Manual. In particular, they should know how to confirm the state of the alarm tones prior to starting a case. The letter states that the firm is in the process of upgrading the software and will be contacting customers to update monitors at no cost. If customers have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support, or outside the United States 1-425-657-7200 x5089. |  | 2812 | total 2812 units distributed worldwide | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Australia, Bahrain, Bolivia, Brazil, Canada, Caymen Island, Chile, China, Colombia, Cyprus, Ecuador, Finland, France, Germany, Guam, India, Indonesia, Israel, Italy, Jordan, Mexico, Morocco, New Zealand, Nicaragua, Pakistan, Paraguay, Peru, Philippines, Poland, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, Uruguay, and Venezuela. | Terminated | May 23, 2012 | 320 | 2812
Z-3030-2011 | 59488 | No | JPD | Hemoglobin A[bdi2] assay. | Hematology | Hematology | 510(k) | Hemoglobin F &amp; A2 Calibrator, |  |  | May 25, 2011 | 2011 | Tosoh Bioscience Inc | The value after reconstitution appeared to be lower than the assigned value on the label. The measured value for Lot BCAL05R was 5.3%, whereas the value on the label is 5.7%. The criteria of the value is +/- 0.3%. | Process control | TOSOH Bioscience sent an "Urgent Recall Notification" letter dated July, 25, 2011 to all affected customers.The letter described the problem and asks customers to immediately discontinue use of the product and return all remaining product back to the firm. Customers can contact TOSOH for replacement.For additional information or question on this recall please contact TOSOH Technical Support at (800) 248-6769. |  | 24 | 6 boxes/4/2 mL btls | Worldwide Distribution: Nationwide (USA) distribution including the states of FL and IL; and the country of Canada. | Terminated | February 24, 2012 | 275 | 24
Z-3036-2011 | 56806 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Product is the Ultraview Spacelabs Compact monitor, models 91367, 91369, and 91370, with touch screen display and wireless option. |  |  | September 17, 2010 | 2010 | Spacelabs Healthcare, Llc | Monitor may reset repeatedly which can result in a loss of patient monitoring for several minutes. Cause is a software problem. | Software design | Spacelabs Medical sent an "URGENT - MEDICAL DEVICE CORRECTION" notification letter to U.S. consignees by mail on September 17, 2010 and to international consignees by e-mail on September 24, 2010. The letter identified the recalled product, warns customers of the potential hazard and identifies some ways to mitigate the risk if they decide to continue use. The letter states to correct this situation, the recalling firm Field Service Engineers will be contact customers to schedule a time to install a free software update. If customers have any questions about this corrective action program, please contact Spacelabs Medical at 1-800-522-7025 and select 2 for Technical Support, or outside the United States 1-425-657-7200 x5089. |  | 579 | 579 total (517 to US acc'ts; 62 to international acct's. | Worldwide Distribution - USA (nationwide) and the countries of: International distribution to: Canada, China, France, Germany, Italy, and the United Kingdom. | Terminated | August 22, 2011 | 339 | 579
Z-3034-2011 | 58381 | No | MRW |  |  | Orthopedic | 510(k) | Catalog # 9024-00 or LSW-S-3040 for 4.5mm BONE-LOK PLS Implant |  |  | March 15, 2011 | 2011 | Interventional Spine Inc | This recall has been initiated due to Pull Pins that may disengage from the Facet Screw during the compression step, requiring that compression be completed with a device driver rather than the compression tool.Use of this product may result in less than optimal compression leading to pseudoarthrosis, non-union or facet pain. | Other | Interventional Spine, Inc. sent a recall letter dated March 15, 2011, informing consignees of the product, the problem, and the action to be taken. Consignees were asked to immediately examine their inventory and quarantine product subject to recall. In addition, if they distributed the product further they were asked to notify Jane Metcalf, the Vice President of Quality Assurance at once.Interventional Spine further informed consignees that the recall was being carried out to the individual consignee level. And their assistance was appreciated and necessary to prevent use of this product.They were then asked to complete and return the enclosed response form as soon as possible. Questions were directed to call their regional directors.For questions regarding this recall call 949-472-0006. |  | 407 | 407 units | Nationwide Distribution including CO, CT, FL, KY, LA, MA, MO, NC, NY, OK, OR, TN, TX, VA &amp; WI | Terminated | February 11, 2013 | 699 | 960
Z-3035-2011 | 59365 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging versions 11.5, 11.5.1, 11.5.2, 11.6, 11.7 and 11.8. |  |  | February 02, 2011 | 2011 | Mckesson Information Solutions LLC | Images for one patient may appear in the study of another patient. | Other | Mckesson Medical Imaging Group sent a URGENT FIELD SAFETY NOTICE letter dated June 13, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.On February 2, 2011, the McKesson Medical Imaging company initiated a process for applying an image display software update to the HMI systems at the hospital site that initiated the complaint. They were advised to contact McKesson Customer Support at 1-800-663-2533, to indicate whether they would like to receive the software update. They were advised to distribute the notification to all those within their organization who need to be aware and to alert other parties affiliated with their organization that are affected by this advisory notice. |  | 438 | 438 devices | Worldwide Distribution--USA (nationwide) Australia, Canada, Guam and the United Kingdom. | Terminated | October 04, 2012 | 610 | 438
Z-3038-2011 | 59465 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Duo-Decapolar Electrophysiology Catheter, Manufacturing Part Number D-1171-35-S, Catalog Number D728260RT |  |  | April 29, 2011 | 2011 | Biosense Webster, Inc. | The recall was expanded because between February and April 2011, Biosense Webster, Inc. notified customers that a voluntary Customer Notification had been Initiated due to the inaccessibility of the electronic version of the Instructions for Use (e-IFU) for the "10 ft. Interface Cable, 10 Pin - Shielded Tip Pins to 12 Pin - Blue", Catalog Number 0128709,through the J&J; Gateway (i.e., e-IFU websi | Labeling Change Control | Biosense Webster sent an "URGENT FIELD SAFETY NOTICE" letter dated April 29, 2011 to all affected customers. The letter describes the product, problem, and the action taken by the firm.The letter notifies and advises customers that the updated IFUs for ISMUS Catheter with Auto ID Technology and the Duo-Decapolar catheters have been released and posted on the e-IFU website as of April 22, 2011 and April 21, 2011. Customers were instructed to complete the Acknowledgement Form and return the form to Biosense. Customers with questions or require additionial information were instructed to contact their Biosense Webster representative. |  | 1530 | 1530 units | Nationwide Distribution | Terminated | October 12, 2012 | 532 | 1855
Z-3042-2011 | 59525 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Equate Muti-Purpose Solution, For Soft Contact Lenses, Sterile, 12 Fl oz. (355 mL), Distributed by Wal-Mart Stores, Inc., Bentonville, AR 72716 |  |  | June 30, 2011 | 2011 | Bausch &amp; Lomb Inc | Muti-purpose solution bottles, lot GC1107 were packaged in a carton that was marked with an incorrect lot code and expiration date. | Packaging | Bausch &amp; Lomb Inc. notified the affected customer via teleconference and e-mail on June 30, 2011. The e-mail describes the product, problem, and actions to be taken by the customer.The customer was instructed to hold and stop shipments of any inventory of the affected product. An official recall was initiated with the customer on July 8, 2011 via e-mail and teleconference.Customer was asked to review their inventory and store returns, and send any of the affected product to Bausch &amp; Lomb Returns Dept., 130 Commerce Drive, Greenville, SC 29615. |  | 1200 | 1200 bottles | Distribution-in the state of AK only | Terminated | November 16, 2011 | 139 | 1200
Z-3041-2011 | 59520 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | AmscoÂż V-PRO 1 and AmscoÂż V-PRO 1 Plus Low Temperature Sterilization Systems Model #MLR, Serial # 033250701 - 031581115. |  |  | August 01, 2011 | 2011 | Steris Corporation | STERIS learned that AmscoÂż V-PRO 1 and V-PRO 1 Plus Low Temperature Sterilization Systems users may experience premature wear of the vacuum pump and filtering assembly components. Users may also experience incomplete aeration of the device's chamber following a failed leak test. | Other | Steris Corporation sent an URGENT VOLUNTARY FIELD CORRECTION NOTICE dated August 5, 2011, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer.A STERIS Service Representative will contact your facility to arrange for STERIS to upgrade your V-PRO System(s). You can continue to use your system prior to the installation of the system upgrades. In the interim period should you experience any alarms or aborted cycles, the load mustbe handled in accordance with the AmscoÂż V-PROTM 1 and V-PROTM 1Plus Low Temperature Sterilization Systems Operators Manual and your internal hospital procedures.For further information or if you have questions regarding this matter, please contact Field Service Dispatch at 1-800-333-8828. |  | 663 | 663 units | Worldwide Distribution -- USA (nationwide) including the states of: AK, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY, and countries of: Australia, Bangladesh, Belgium, Belarus, Brazil, Canada, Chile, China, Colombia, Costa Rica, Denmark, Ecuador, Egypt, Finland, France, Germany, Hungary, India, Iraq, Ireland, Italy, Japan, Kuwait, Lebanon, Mexico, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, South Korea, Spain, Switzerland, Turkey, Vietnam, United Arab Emirates, and United Kingdom. | Terminated | March 27, 2013 | 604 | 663
Z-3039-2011 | 59391 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Calcium Flex Reagent is used on the Dimension clinical chemistry system as an in vitro diagnostic test intended for the quantitive determination of calcium activity in human serum, plasma and urine. |  |  | July 19, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostic has confirmed a high frequency of "Abnormal assay" flags and imprecision. Test results across the assay range may be falsely depressed or elevated, as much as, 1.93 mg/dL at a concentration of 9.26 mg/dL. | Other | An Urgent Field Safety Notice datey July 2011 was issued to all affected customers. Customers are being advised to immediately discontinue use and discard any remaining inventory of Dimension(R) CA Flex(R0 reagent cartridge - lot EA2153. Replacement order for the affected lot is in process, and will be sent to impacted customers free of charge.Customers were instructed to forward the Field Safety Notice to anyone whom they may have distributed the affected product.For questions regarding this recall call 302-631-6299. |  | 2500 | 2500 cartons | Worldwide Distribution to Guam, Puerto Rico, Chile, Costa Rica, Venezuela, South Korea, Vietnam and domestically to AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MN, MI, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, WA, WV, WI, and WY. | Terminated | December 27, 2013 | 892 | 2500
Z-3044-2011 | 59375 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TSRH(R) Spinal System SCREW, REF 837-645, Size: 6.5mm x 45mm, Mat'l: TITANIUM ALLOY, Rx only, QTY: 1 EA, Medtronic Sofamor Danek USA. |  |  | July 12, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | Screws are 5 mm longer than the size etched on the screw. | Error in labeling | On 07/12/2011, the recalling firm, Medtronic notified the Sales Managers and Sales Force by conference call and email regarding the imminent notification and follow-up activities. A "VOLUNTARY RECALL" notification letter was sent to Risk Managers and Sales Representatives on 07/13/2011 with a delivery date of 07/14/2011. The letter states to immediately remove the recalled product from use and return to Medtronic. The letter includes a questionnaires form to be completed and returned as per the instructions. If appropriate, personal follow up with Sales Representatives will be made to assure that Risk Managers understand the recall communication and to facilitate return of product and the associated questionnaire documentation. For questions, contact your Medtronic Sales Representative or the Global Quality Department at 800-876-3133, extension 3173. |  | 17 | 17 units | Nationwide Distribution (USA) including the states of: OH, NY, and WI | Terminated | July 20, 2012 | 374 | 17
Z-3040-2011 | 59688 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | enFlow IV Fluid Warmer Part number 980105VS; |  |  | June 30, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | enFlow IV Fluid Warmer may have an issue which results in excessive external surface temperature on the underside of the warmer. | Device Design | GE Healthcare sent UMDC Letters/UMDC Confirmation Letters to customers on June 30, 2011. GE Healthcare sent UMDC Letters/UMDC Confirmation Letters on October 24, 2011 reflecting a new recall strategy. |  | 1307 | 1307 units | Worldwide Distribution - USA (nationwide) and the country of Vietnam | Terminated | June 14, 2013 | 715 | 1307
Z-3051-2011 | 56251 | No | DHN | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | "BIO-RAD Kallestad Mouse Stomach/Kidney 8 Well Slides, Bio-Rad Laboratories, STORE FROZEN" |  |  | June 30, 2010 | 2010 | Bio-Rad Laboratories | KALLESTAD Mouse Stomach/Kidney (MSK) test kits contained Crithidia slides instead of the MSK slides. | Employee error | Bio-Rad Laboratories notified affected customers by phone on 6/30/2010 and followed up with a 2nd call on 7/13/2010.The call informed customers of the problem with the kit and requested they destroy the Crithidia slide packs. Bio-Rad send out replacement slide packs. For information or questions on this recall call Bio-Rad at (425) 498-1709. |  | 7840 | 224 kits were distributed; 35 of those kits were determined to contain the wrong slides. | Worldwide Distribution: (USA) Nationwide distribution including the states of CT, OK; and the country of Puerto Rico. | Terminated | August 22, 2011 | 418 | 7840
Z-3049-2011 | 59466 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Churchill Medical Latex-Free IV Start Kit |  |  | July 15, 2011 | 2011 | Churchill Medical Systems, Inc. | Outer Kit labeled expiration date longer than the expiration of the shortest expired product | Employee error | Churchhill Medical Systems, Inc A Vygon Company issued an URGENT MEDICAL DEVICE RECALL letter was via overnight courier mail on July 15, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The product was requested to be returned. Distributors were provided with instructions to communicate the recall to customers who received the affected product by providing a copy of the notification and a copy of the Recall Acknowledgement and Inventory Return Form. The form should be completed in its entirety, signed and returned to the Distributor. Contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by email to customerservice @vygonus.com |  | 130750 | 130,750 units | Nationwide Distribution | Terminated | March 27, 2013 | 621 | 130750
Z-3045-2011 | 59713 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | VerifyÂż SixCess Indicator Strips, Model #: PCC005, PCC006, PCC007, PCC008, PCC031, PCC033, PCC034, PCC037 &amp; PCC038, Mfr. By: Albert Browne, LTD, Chancery House 190 Waterside Rd. Hamilton Industrial Park, Leicester, Leicester, LE5QZ, United Kingdom. |  |  | March 16, 2011 | 2011 | Steris Corporation | During an inspection the firm was notified that language contained in labeling and promotional literature directs Customers that the products can be used to replace biological indicators and/or release loads containing implants. | Labeling False and Misleading | STERIS sent a Customer Notification letter dated March 16, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review and follow the Challenge Packs' revised instructions for use enclosed as Attachment A (270-4) and Attachment B (275F 3).For any questions regarding this recall customers were instructed to contact their STERIS Account Manager or call 440-392-7421. |  | 2500 | 2,500 (one per customer) | Worldwide Distribution - USA (nationwide) and the countries of Lebanon and Qatar. | Terminated | March 19, 2013 | 734 | 2500
Z-3050-2011 | 59329 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Smith &amp; Nephew, Tendon Stripper, Slotted, 5.0 mm I.D. x |  |  | July 07, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | Shaft tip (casting) does not meet the SNE 416specification, which could result in rusting of the Tendon Strippers | Nonconforming Material/Component | Smith &amp; Nephew sent an Urgent - Product Recall Notification letter dated July 7, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the form included in the letter and return to Hoangthi.le@smith-nephew.com or fax to 508-261-3620. Each affiliate must then notify each facility affected by the recall. To return product and obtain credit or order replacement product customers should contact the Smith &amp; Nephew Endoscopy Returns at 800-343-5717 (option 3) for a Return Authorization (RA) or send an email to, endo.andreturns@smith-nephew.com. For questions, contact Hoangthi by phone at 508-337-3731. |  | 570 | 570 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Chile, China, Columbia, Denmark, Dubai, Finland, France, Germany, Greece, Hong Kong, India, Italy, Israel, Japan, Korea, Malaysia, Mexico, Netherland, Norway, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, United Kingdom, and Venezuela | Terminated | July 23, 2015 | 1477 | 570
Z-3057-2011 | 59576 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet HL Console for Heart Lung Machine HL20: |  |  | June 20, 2011 | 2011 | Maquet Inc. | HL 20 Consoles for 4-pump and 5-pump configuration can malfunction if an excessive amount of liquid is used causing circuit board malfunction. | Device Design | Maquet sent an URGENT PRODUCT RECALL NOTIFICATION letter dated June 17, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to acknowledge receipt of the notification by completing and sending back the attached response form whether or not they have any affected product, fax to 973-807-9210 no later than June 30, 2011. For any questions customers were instructed to contact thier local MAQUET representative. |  | 229 | 229 units | Nationwide Distribution | Terminated | October 16, 2012 | 484 | 229
Z-3048-2011 | 59568 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity Gamma X XL Portable Patient Monitor. |  |  | July 29, 2011 | 2011 | Draeger Medical Systems, Inc. | Monitor keys may spontaneously become inoperative or active. This may cause a membrane switch panel malfunction, which could result in the monitor discharging a patient automatically. | Device Design | Draeger Medical sent an "Urgent Medical Device Recall Notification" letters dated July 2011 to all affected customers.The letter included product and problem affected. The letter informed the customer a Draeger representative will contact them to upgrade the system. Until the solution is implemented, Draeger recommends the Delta and/or Gamma XXL may be used. If the affected monitors become inoperative or activate without user interaction, Draeger advises removing the monitor from service and contacting a Draeger representative.For questions or for information on this recall please call Draeger Service at (800) 543-5047 (press 1 at the prompt, then 2, then 32349). |  | 29 | 29 | Worldwide Distribution - (USA) Nationwide distribution including the states of AK, FL, LA, MN, MS, MO, NJ, NC, OH, OK, PA, TN, TX, NC, NH, NJ, NY, WV, and WI; and the countries of Australia, Austria, Argentina, Bahrain, Brazil, Bulgaria, Canada, China, Colombia, Czech Republic, France, Germany, Georgia, Greece, Hungary, Italy, Ireland, India, Kazakhstan, Latvia, Macedonia, Malaysia, Mexico, Netherlands, Norway, Russian Fed., Romania, Saudi Arabia, Slovakia, Spain, South Africa, Switzerland, United Arab Emirates, United Kingdom, Uruguay, Vietnam, and Yemen. | Terminated | August 16, 2012 | 384 | 61
Z-3056-2011 | 59476 | No | FHN | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | The Resolution II Clip, 155 cm size and 235 cm size, Boxes of 1, 10, and 20 clips consists of a pre-loaded, radiopaque single use clip on a flexible delivery system. |  |  | July 19, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific has received reports where, following clip deployment, the alignment of the clip and delivery system does not allow for the withdrawal of the delivery system away from the clip, resulting in tissue trauma, increased bleeding and the clip being pulled off tissue as the delivery system is withdrawn. | Other | Boston Scientific sent an URGENT MEDICAL DEVICE RECALL letter dated August 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the Medical Device Recall to their customers. Customers were instructed to:Immediately discontinue use of any affected product.Complete and return the Reply Verification Tracking Form by fax to 508-683-5578.Package and return any affected product to:Boston Scientific CorporationDistribution CenterAttn: QA Returns500 Commander Shea BlvdQuincy, MA 02171For questions regarding this recall call 508-683-5578. |  | 1302 | 14, 093 boxes of both 155 cm and 235 cm size clips | Worldwide Distribution - including USA (nationwide) | Terminated | June 22, 2012 | 339 | 1302
Z-3055-2011 | 59402 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Flexible Probe with Blocking Washer (part of Segmented Cylinder Set) |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Blocking washers on the Flexible Probe with Blocking washers may shift and cause potential for treatment length error. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to locate old design applicator probes and guiding tubes and to cease use of the affected products. Information in the letter included instructions on how to return the products to Varian BrachyTherapy.Customers should contact their local Varian Customer Support District or Regional Manager for questions regarding this notice. |  | 1019 | 1019 total | Worldwide Distribution | Terminated | June 20, 2012 | 343 | 1019
Z-3052-2011 | 59226 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | High-Definition 120 Multileaf Collimator |  |  | June 06, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | When using the HD120 Multi Leaf Collimator for beam shaping, and under specific treatment conditions, dose may be delivered outside of the intended treatment field that exceeds the expected transmission dose. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 6, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter describes various jaw position recommendations to be followed when planning and treating dynamic fields on the HD120 MLC. Contact the Varian Oncology Help Desk at 1-888-827-4265 for questions regarding this notice. |  | 106 | 173***12-20-12 AMENDED TO: 106 units*** | Worldwide Distribution. | Terminated | December 23, 2011 | 200 | 106
Z-3059-2011 | 59524 | No | MJJ | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Adjustable Vent Valve, TruChoice Alliant Healthcare Products REFERENCE#213-030NS QUANTITY: 100 DESCRIPTION: ADJUSTABLE VENT VALVE, NON-STERILE; and 3T Medical Systems Catalog# 030NS Adjustable One Way Vent Valve, Quantity: 100, NON-STERILE |  |  | July 21, 2011 | 2011 | Alliant Enterprises LLC | The firm is initiating this recall due to the receipt of 2 complaints of malfunction of the joint holding the 2 halves of the vent valve together. | Other | Alliant Healthcare sent an URGENT: MEDICAL DEVICE RECALL letter on 7/21/2011 and 7/25/2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately review the recall notification and quarantine any affected product. If product was further distributed, customers were to forward the letter with the Recall Return Response Form immediately to all affected locations and indicate which locations on the Recall Return Response Form. Customers were also instructed to discontinue use immediately, dispose of the affected product per their normal disposal protocol, and fax the Recall Return Response Form to Alliant Healthcare Products at 269-629-5770.Upon receipt of the Recall Return Response Form, a substitute vent valve (non-adjustable only) will be sent to their location for each affected product for which disposal has been properly documented in writing. For questions call 269-629-0300, ext 226. |  | 14900 | 149 Boxes of 100 | Worldwide Distribution - USA including Michigan, Tennessee, New York, and California and the countries of Chile and Korea | Terminated | April 23, 2012 | 277 | 41475
Z-3054-2011 | 59276 | No | GTY | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | @ Home Strep Test |  |  | July 06, 2011 | 2011 | Jay Bird Investments Inc | Home Strep Test does not qualify as a Class I device and needs pre-market approval. | PMA | Jay Bird Investments, Inc. sent a notification letter dated July 6, 2011 to customers. The letter describes the product, problem and the actions to be taken. The customers were instructed to notify retailers and remove the product from the shelves. For questions regarding disposition of the product please contact Myron Bouwhuis at 801-721-3528. |  | 7541 | 7,541 | Nationwide Distribution (USA) including the states of OR and UTAH. | Terminated | June 20, 2012 | 350 | 7541
Z-3063-2011 | 59523 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | Sysmex Stromatolyser-IM; a lytic reagent for use with the XE-2100, XE-2100L and XE-5000 Series Sysmex Automated Hematology Analyzers; the reagent solution is supplied in a 10-liter cubetainer (cardboard cube with a 10-liter plastic liner); Sysmex America, Inc., Mundelein, IL 60060; catalog number SIM-220A. |  |  | April 11, 2011 | 2011 | Sysmex America, Inc. | The affected lots of Stromatolyser-IM lysing reagent were found to be contaminated with Pseudomonas bacteria, which resulted in rising IMI channel background counts, increased false "Platelet Clump" flags and a strong sulfur-like odor when used on the XE-2100, XE-2100L and XE-5000 analyzers. | Process change control | Sysmex America, Inc. sent a "PRODUCT NOTIFICATION" letter dated April 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken. The letter instructs customers to discontinue use of the affected SIM lots, and to remove affected lots from their analyzers as soon as possible to avoid long-term performance problems on the analyzer. A Stromatolyser -IM (220-A) Voluntary Recall form was attached for customers to complete and return via fax to Sysmex Customer Service at 800-689-8296. The form will be used to place orders to replace the affected product. Any questions concerning the information contained in the notification are directed to the Sysmex Technical Assistance Center at 1-888-879-7639. |  | 1887 | 1,887 cubetainers | Worldwide Distribution--USA (nationwide and Puerto Rico) and the countries of Canada, Colombia, Ecuador, Mexico, Netherland Antilles and Peru. | Terminated | March 14, 2012 | 338 | 1887
Z-3053-2011 | 59693 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Katzen" Infusion Wire, 0.035 in x 180 cm, |  |  | August 01, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is voluntarily recalling one lot/batch of Katzen" Infusion Wires because we have identified that the label on the carton and the inner pouch indicated an incorrect device length. The labels indicate a device length of 180cm while the actual device length is 145cm. | Labeling mix-ups | Boston Scientific sent an "Urgent Medical Device Recall Removal" letter dated August 1, 2011 to all affected customers.The letter described the problem and product being recalled and advised customers to segregate the affected product immediately and return it to Boston Scientific. Boston Scientific requested the completion and return of the Reply Verification Tracking Form.For additional information on this recall call Boston Scientific Quality Systems at (763) 494-1133. |  | 33 | 33 | (USA) Nationwide distribution including the states of AZ, AR, CA, CO, FL, GA, IL, IN, NV, NY, NC, OH, PA, RI, and TX. | Terminated | August 10, 2012 | 375 | 33
Z-3074-2011 | 59680 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Habib, 4x Kit, Bipolar Resection Device RITA (Radiofrequency Ablation Devices). |  |  | August 30, 2010 | 2010 | Angiodynamics, INC | The labeling contains errors which result in the products being misbranded | Labeling False and Misleading | AngioDynamics, Inc. sent out a "Medical Device Recall" letter dated August 30, 2010 to all affected customers.The letter included; products affected, problem and asked them to return the product back to AngioDynamics. Customers were asked to fill out and return the included Recall Reply Form. For question on this recall contact AngioDynamics at (800) 772-6446. |  | 132 | 132 units | Worldwide Distribution: (USA) Nationwide distribution and the countries of Austria, China, Denmark, Egypt, France Germany, India, Italy, Latvia, Netherlands, Poland, Russian Federation, South Africa, Switzerland and the United Kingdom | Terminated | March 12, 2012 | 560 | 769
Z-3070-2011 | 59588 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Entroy Pool Lift; a patient lift used in conjunction with a chair or stretcher to lift them in/out of a swimming/therapy pool; Arjo Hospital Equipment AB, Eslov, Sweden; Models GAB1000-01-xx, GAB1000-11-xx and GAC100x |  |  | August 04, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | The Entroy Pool Lift boom may rapidly lower un-commanded to the bottom of its stroke, with a potential for injury to the resident or caregiver. | Device Design | ArjoHuntleigh sent an Urgent Field Correction Notice I.D. 110701AHE to the affected end user via UPS 2nd Day on August 4, 2011. The notice informed the account of the potential for lift boom to rapidly lower un-commanded to the bottom of its stroke. The account was requested to immediately cease use of the affected lift, and disconnect the power to the lift to prevent further use until corrected. The account was requested to complete and sign the Customer Response Form, and return it to ArjoHuntleigh to schedule a service call to correct the lift. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 54485. |  | 311 | 1 unit in the U.S., 311 units in the rest of the world | Worldwide Distribution - USA including CT | Terminated | November 28, 2011 | 116 | 311
Z-3069-2011 | 59560 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical 15x100 mm Kii Optical Access Systems, Non-Threaded 6/Box, Model C0R39 |  |  | December 18, 2009 | 2009 | Applied Medical Resources Corp | The recall was expanded on December 22, 2009 to include four models: C0R36, C0R37, C0R38, and C0R39. Initially, Applied Medical conducted a voluntary recall on June 9, 2009 of two lot numbers, 1084426 and 1085506, of the 15 mm Kii non-threaded optical access system, model C0R36, due to a potential weakness in the cannula tip, possibly resulting in a tip breakage. | Component design/selection | Applied Medical Resources Corp. notified customers via an overnight "URGENT: MEDICAL DEVICE RECALL" letter dated December 22, 2009. The letter informed customers of the recalled products, problem, and actions to be taken. The letter asked that the listed recalled products be returned immediately. The customers were asked to check their inventory, complete the attached Voluntary Recall Notification Confirmation Form and fax as per the instructions. The recall letter provided customers with a contact and telephone number for product return questions (1-949-713-8652) and a separate contact for regulatory questions (1-949-713-8369). |  | 13791 | 13,791 units total | Worldwide Distribution - Nationwide in the USA and to countries of: Australia, Austria, Belgium, Canada, Chile, Columbia, Cyprus, Denmark, Finland, France, Germany, Great Britain, Hong Kong, India, Ireland, Italy, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Saudi Arabia, Singapore, South Africa, and Switzerland. | Terminated | September 06, 2011 | 627 | 13791
Z-3075-2011 | 58489 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, DASH 3000, 4000, 5000 Patient Monitors |  |  | July 05, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware of an issue with the Dash 3000/4000/5000 patient monitors configured with the 802.11b wireless LAN option that may impact patient safety. The Dash monitor may exhibit reboots on some networks when the 802.11b wireless LAN option is used to connect to the hospital wireless LAN infrastructure. This reboot results in a loss of patient monitoring for approxima | Software design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. GE recommended actions that should be taken by a qualified service representative. For any questions regarding the medical device correction or identification of affected items contact Technical Support at 1-800-558-7044 or customers should contact their local Service Representative.UPDATE: A second letter dated 12/13/11 was sent on 1/20/12 to consignees. The revised letter provides an update to the safety instructions and product correction. |  | 8029 | 8029 devices | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, WI and the countries of UNITED KINGDOM, TURKEY, THAILAND, TAIWAN, SWEDEN, SPAIN, SLOVENIA, SINGAPORE, SAUDI A ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, POLAND, NORWAY, NEW ZEALAND, NETHERLANDS, MEXICO, KUWAIT, JAPAN, ITALY, ISRAEL, INDIA, GERMANY, FRANCE, FINLAND, ESTONIA, ECUADOR, DENMARK, CANDA, BELGIUM AUSTRIA, AUSTRALIA, and ALGERIA. | Terminated | September 11, 2012 | 434 | 8029
Z-3085-2011 | 59469 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microline Surgical ReNew Laproscopic Long Fenestrated Grasper Forceps Tip |  |  | July 18, 2011 | 2011 | Microline Surgical, Inc. | Mislabeled: ReNew product number 3262 (Long Fenestrated Grasper) was packaged and labeled asproduct number 3642 (Super-Atrau Fenestrated Grasper). ReNew product number 3642 (Super-Atraumatic Fenestrated Grasper) was packaged andlabeled as 3262 (Long Fenestrated Grasper | Packaging process control | Microline Surgical sent a RECALL NOTICE letter dated July 18, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine product subject to the recall. If customers further distributed the product, customers should identify their accounts and notify them immediately of the product recall. Customers were instructed to contact Microline Customer Service to obtain a Return Authorization Number for return of the recalled item. Customers were asked to complete and return the enclosed Recall Return Response Form as soon as possible. For any questions customers were instructed to call the Regulatory Affairs Manager at 978-867-1726. |  | 170 | 170 units | Worldwide Distribution - USA including FL, MI, MN, TN and WI and the countried of Canada, The Netherlands, Thailand, UK, Norway, and Italy. | Terminated | September 09, 2011 | 53 | 470
Z-2600-2011 | 58804 | No | REO |  |  |  |  | Prism Laser 7642 |  |  | April 27, 2011 | 2011 | American Stress Technologies, Inc. | It was discovered The Prism 7642 was not certified at the time of delivery. The product lacks the required labeling and the product was sold and delivered without the submission of the Initial Product Report. The product lacks User In formation | Radiation Control for Health and Safety Act | American Stress Technologies sent a customer notification letter dated April 1, 2011, informing them of the action they have taken to correct this issue. (1) ASTI conducted laser safety tests and checks on May 3,2011 to verify the Prism was compliant with FDA regulations. This was confirmed in a June 17, 2011, electronic mail message.(2) ASTI placed the required labels described in the initial product report on the Prism on May 3, 2011.(3) ASTI delivered updated user manuals to UNO on May 3, 2011, that describe potential laser hazards and the required labeling in the same way as described in the initial product report.(4) ASTI completed the above corrective actions at no cost to UNO as confirmed in a June 20, 2011, electronic mail message.If you have any further questions please call (724) 410-1030 |  | 1 | 1 | Nationwide Distribution | Open, Classified |  |  | 1
Z-3079-2011 | 58874 | [6 empty] | Elekta Synergy |  |  | May 06, 2011 | 2011 | Elekta, Inc. | A small number of the Modular Reverse Diodes (D1A), a component in the equipment room of the Elekta Linear Accelerator, have ruptured. | Other | ELEKTA sent an IMPORTANT NOTICE (FCO 20000103022), dated May 6, 2011. The notice identified the product, the problem, and the aciton to be taken by the customer. Elekta recommended that service personnel always wear protective safety glasses, when they do work on the modulator assembly with the covers removed. Elekta also recommended that users make sure that the modulator cover is installed when the digital accelerator is in operation. Elekta will release a Field Change Order when a corrective action is identified. |  | 368 | 368 | Nationwide and Puerto Rico | Terminated | October 02, 2012 | 515 | 368
Z-3043-2011 | 59317 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II, Model 8637, (The SynchroMedÂż II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The implantable components of the SynchroMed II Infusion System include the pump, catheter, and catheter accessories. |  |  | July 05, 2011 | 2011 | Medtronic, Inc. - Neuromodulation | Medtronic is updating information regarding the potential for reduced battery performance that can lead to sudden loss of therapy in a small percentage of Medtronic Model 8647 SynchroMed II pumps that was communicated with Healthcare providers in July 2009. The purpose of the current communication is to provide updated information regarding the scope and occurrence of this issue and to emphasize | Component design/selection | Medtronic Neuromodulation notified physicians with an Urgent Medical Device Correction letter beginning July 05, 2011. A press was issued July 08, 2011. The letter described the issue, severity, recommendations if they notice issue, and patient management recommendations. UPDATE: in November 2011, MDT began to exchange unused pumps in inventory with pumps that contained the new battery. A letter dated November 2011 was left behind at locations in which the MDT representative swapped out the devices. |  | 139653 | 139,653 (105,002 US, 34,61 OUS) | Worldwide including USA, Puerto Rico, Algeria, Aruba, Australia, Austria, Belarus, Belgium, Brazil, Chile. Colombia. Costa Rica. Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Hungary, Iceland, India, Iran Iraq, Ireland, Israel Italy, Jordan, Kuwait, Lebanon, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Norway, Panama, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | August 24, 2012 | 416 | 139653
Z-3084-2011 | 59577 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet HL Console for Heart Lung Machine HL20 |  |  | June 20, 2011 | 2011 | Maquet Inc. | Maquet HL 20 Twin Pump Modules may unintentionally increase speeds for a 1.5-2 second period of time before returning to its original speed. | Device Design | Maquet sent a PRODUCT RECALL NOTIFICATION LETTER/return response forms, dated June 17, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product until a software update was performed by their trained technicians. Customers were also asked to forward the information contained in the notification letter to all users and to all appropriate Biomedical Staff. Customers were instructed to acknowledge receipt of the notification letter by completing and faxing back the attached response form to 973-807-9210 whether or not they have any of the affected product. For questions customers were instructed to contact their local MAQUET representative. For questions regarding this recall call 973-709-7994. |  | 299 | 299 units | Nationwide Distribution | Terminated | October 16, 2012 | 484 | 299
Z-3077-2011 | 59370 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MagNA Pure LC 1.0 (software version 3.0.11). |  |  | July 21, 2011 | 2011 | Roche Diagnostics Operations, Inc. | Roche has detected a problem with the MagNA Pure LC 1.0 (software version 3.0.11) and the MagNA Pure LC 2.0 (software versions 1.1.23 and 1.1.24) instruments. Under certain circumstances, only a fraction of each sample is processed which may lead to false negative results, and an impact to patient health if used in a homebrew diagnostic application. | Other | Roche Diagnostics Operations, Inc. sent a letter dated July 21, 2011 entitled "URGENT MEDICAL DEVICE CORRECTION" to all affected customers.The letter included the name product, software versions, and reason for recall and informed them the problem is under investigation. In the meantime, customers are asked to wait 30 seconds after the finalization of the decontamination cycle before initiating the next operation, in order to avoid this issue.For questions on this recall please call Roche at (800) 428-5074. |  | 422 | 422 | Nationwide Distribution; including the islands of Puerto Rico and Hawaii. | Terminated | July 02, 2012 | 347 | 570
Z-3083-2011 | 59483 | No | ONW | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008K@home |  |  | May 04, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Upgrade software to implement new on screen graphics and instructions, the addition of a bracket on the side of the machine to temporarily hold the bloodlines during set up in order to conform to approved 510(k) | Device Design | Fresenius Medical NA issued a formal field correction notification dated May 4 ,2011, sent by certified mail with return receipt requested. The letter identified the product the problem and the action needed to be taken by the customer. The customer was notified of the retrofit that would be done to their machines. The patients' facilities were made aware of the changes that would be made to each of their patient's 2008K@home machines on May 4, 2011. On May 13, 2011, all facilities with active 2008K@home machines were sent a notification letter of the retrofit.Please return the receipt for confirmation that you have received this notice.Please call (781) 699-2053, to schedule a certified technician to conduct the machine retrofit. |  | 705 | 705 units | Nationwide Distribution.Canada - FMCNA is not conducting the recall to Canada, as Health Canada did not require the same modifications to the machines | Terminated | August 14, 2014 | 1198 | 705
Z-3081-2011 | 59462 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Mini MaxLock Extreme Plating System Case &amp; Tray, Model Number(s): MXM-901, |  |  | July 06, 2011 | 2011 | Orthohelix Surgical Designs Inc | The surgical tray does not meet our Instruction for Use (IFU) for dry time requirement. | Nonconforming Material/Component | Orthohelix Surgical Designs Inc sent a RECALL NOTIFICATION letter dated July 6, 2011 and an UPDATED URGENT RECALL 2ND NOTIFICATION letter dated July 13, 2011 to all affected customers. The letter identified the recalled product, the problem, and the Field Correction actions to be taken. Customers were instructed to remove the blue rubber mat in the utility bay of the tray base immediately, tag the mat with the tray number, and return it to recalling firm immediately. The letter states once customers have removed the blue rubber mat from the tray base, they may continue to use the tray and system normally. No further actions or precautions are needed. Distributors were also asked to notify their customers of the recall. If you have any questions, please call 330-247-1444 or 330-247-1445. |  | 61 | 61 trays | USA Nationwide Distribution including the states of: AR, AZ, CA, CO, FL, GA, IL, IN, KY, MA, MD, MI, MO, NC, NE, NJ, NM, NY, OH, OR, PA, SD, TN, TX, UT, VA, WA, WI &amp; WY. | Terminated | September 10, 2012 | 432 | 61
Z-3080-2011 | 59400 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model 754/754M Disposable MRI Ventilator Circuit |  |  | May 27, 2011 | 2011 | Impact Instrumentation, Inc. | Model 754/754 M Disposable MRI Ventilator Circuits are being distributed without market clearence (510k). | Labeling False and Misleading | Impact Instrumentation Inc. sent an "Urgent Product Recall" dated May 27, 2011 to all affected customers.The letter informed customers the product is not a cleared product and to return the product back to Impact Inst. for credit. A notification form is included in the letter to be return to the firm.For information on this recall contact the firms Customer Service Department at (800) 969-0750, 8:30 am to 5:00 pm EST. |  | 492 | 492 units | Nationwide Distribution; including the states of CA, FL, IL, NJ, OH, TX, UT, MN, MO, MS, OH, OR, TX, AR, and PA. | Terminated | May 23, 2013 | 727 | 492
Z-3082-2011 | 59399 | No | JDB | Elbow joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Latitude", Tige Humerale, Humeral Stem, Medium 77 mm, Humeral Stem Right, For Cemented use only, Sterile R, REF 0030402, Tornier In. Edina, MN 55435 USA. |  |  | June 23, 2011 | 2011 | Tornier, Inc | Tornier has discovered that some models of the Tornier latitude Elbow Prosthesis (Small and Medium Humeral stems) were manufactured with a humeral screw that may not function as designed. | Process change control | Tornier, Inc. sent an "Urgent Product Correction" letter dated June 23, 2011. The letter was addressed to the doctors. The letter described the product and the problem. Doctors were insructed to monitor patients that have received the units for evidence of humeral screw loosening.For questions customers were instructed to contact the Customer Service Department at 1-888-494-7950. |  | 3 | 3 | Nationwide Distribution including FL and MO | Terminated | July 17, 2012 | 390 | 3
Z-3091-2011 | 57806 | No | LGW |  |  | Neurology | PMA | Software version 3.3 for Rapid Programmer devices, Model 3832. |  |  | June 02, 2009 | 2009 | Advanced Neuromodulation Systems | Product programmed with unapproved software was distributed. | Software change control | St. Jude Medical Neuromodulation contacted all affected customers by telephone between June 2, 2009 and June 5, 2009. Customers were given information about the product, problem and action to be taken.All affected devices were returned to the firm as instructed. |  | 14 | 14 units | Distributed in New Hampshire and Puerto Rico. | Terminated | September 27, 2012 | 1213 | 14
Z-3092-2011 | 59754 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Evident MWA Generator, REF VTGEN, Output: 60 W / 50 Ohm, Valleylab, a division of Tyco Healthcare Group LP Boulder, CO 80301. |  |  | February 06, 2009 | 2009 | Covidien, PLC | Generator may deliver full power when a lower power is selected. | Component design/selection | Covidien sales reps began notifying affected customers on February 6, 2009, followed by a letter. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to return the affected product to:CovidienEnergy Based Devices5920 Longbow RoadBoulder, Colorado 80301Customers were asked to label the exterior of the carton with their Return Authorization number. Covidien EbD would address the issue and return the unit to them.For any questions customers were asked to call Covidien EbD's Service Department toll-free at 1-800-255-8522, x 4923 or contact their local sales representative for additional information. |  | 357 | 357 units | Worldwide Distribution - USA (nationwide) and the countries of Belgium, Denmark, France, Germany, Hong Kong, Hungary, Israel, Italy, Netherlands, Poland, Portugal, Romania, Singapore, Slovakia, Spain, Switzerland, and the UK. | Terminated | August 30, 2011 | 935 | 357
Z-3100-2011 | 59351 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand 4D Integrated Treatment Console (4DITC) Radiation Therapy, Interpretation of a RapidArc Treatment Plan as a 3D Conformal Arc Treatment Plan; version v8.6 or 8.8 in combination with treatment planning systems; Model Number: H51; Reference/PSCA Identifier: CP-03063; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | June 27, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In certain configurations, the 4DITC can interpret a RapidArc or VMAT treatment field as 3D Conformal Arc Treatment Plan. If a RapidArc or VMAT treatment field is incorrectly interpreted as Conformal Arc, the field will be delivered with a constant dose rate and gantry speed for the entire duration of the delivery, which can lead to delivered dose being different than planned. | Software design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 27, 2011 to all affected customers. The letter identified the product, problem and recommended user actions to follow. Customers were instructed to notify all appropriate personnel of the correction notice. The letter states that Varian Service will contact customers to arrange to perform update to 4D Integrated Treatment Console (4DITC) software. For further clarification, contact your local Varian Customer Support District or Regional Manager. The letter also provides Varian Oncology Help desk contact information 1-888-827-4265 for USA and Canada and +41417498844 for Europe. |  | 645 | 645 | Worldwide Distribution - USA (nationwide) | Terminated | December 05, 2013 | 892 | 645
Z-3093-2011 | 59781 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Encision AEM Suction Irrigation Electrode, L-Diamond, |  |  | August 24, 2011 | 2011 | Encision, Inc. | Electrode tips have been reported breaking off during use or cleaning. | Device Design | Encision Inc. sent out an "Urgent Product Recall" letter dated August 24, 2011 to all affected customers.The letter included: affected product, problem and instructed customers to discontinue use and return all affected electrodes to Encision. An enclosed reply form was to be faxed back to Encision indicating status of firm's inventory.For information on this recall call Encision at (303) 339-6917. |  | 169 | 169 units | Worldwide Distribution: Nationwide distribution, and the country of New Zealand. | Terminated | September 04, 2012 | 377 | 1078
Z-3099-2011 | 59238 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PFC Sigma Round Dome Patella 38 MM 3 Peg |  |  | June 28, 2011 | 2011 | DePuy Orthopaedics, Inc. | DePuy initiated a voluntary recall effective immediately, for one lot of P.F.C. SIGMA RD Dome Patella 3 Peg (38mm) due to packaging labels that did not correctly reflect the size of the enclosed patellar component. In the recalled lot, the packaging is labeled for a P.F.C. SIGMA 38mm 3 peg Round Dome Patella, but the product enclosed in the package is a P.F.C. SIGMA 35mm 3 peg Oval Dome Patella. | Employee error | The firm, Depuy , sent an email that included a letter entitled "URGENT INFORMATION - RECALL NOTICE" dated June 28, 2011 to DePuy Area customers ( Directors, Distributors, Office Managers, Hospital Representatives, and Doctors). The letter described the product, problem and actions to be taken. The letter also included implications of implanting the recalled device and patient communication information. The customers were instructed to immediately cease further implantation/use or distribution of the product, immediately return any recalled devices in their inventory, notify their customer of the recall and complete, sign and return the Reconciliation Form(s) via fax to DePuy at 574-372-7567 within seven business days.The letter also states that DePuy is not recommending prophylactic revision or additional follow up in the absence of symptoms. Additionally, DePuy recommends that surgeons continue to follow patients as per their standard care.Questions regarding the recall may be directed to the following DePuy representatives--questions about recall information call the Manager of Customer Quality at 574-372-7333; Salesforce questions, call the Group Product Director, 574-372-5046 or Associate Product Director at 574-372-7154 and for Clinical/Surgeon questions, contact DePuy's Scientific Information Office at 1-888-554-2482. |  | 29 | 29 | Nationwide distribution: CA, MI, NJ and WI. | Terminated | August 28, 2013 | 792 | 29
Z-3090-2011 | 59662 | [6 empty] | Bausch &amp; Lomb Stellaris 1.8mm Premium Vacuum Phaco Pack with MICS Needle, Ref: BL5114, packaged 6 units/case, RX, Sterile. |  |  | July 27, 2011 | 2011 | Bausch &amp; Lomb Inc | The aspiration channel on the affected vacuum cassettes may be blocked due to a molding malfunction | Nonconforming Material/Component | Bausch &amp; Lomb Inc. notified all affected customers by direct telephone contact on July 27, 2011 and a follow-up FAX or e-mail dated July 27, 2011The letter included; affected product, problem, and requests customers return the product. An acknowledgement form was included to complete and return with the product.For question on this recall contact Bausch &amp; Lomb at (949) 521-7895. |  | 6 | 6/6-unit cases | Nationwide Distribution. | Terminated | June 21, 2013 | 695 | 246
Z-3102-2011 | 57590 | No | EZO | Urethrotome. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | OLYMPUS Pediatric Knives and Resection Electrodes |  |  | December 22, 2010 | 2010 | Olympus America Inc. | The device sterile package seal shows evidence of being breached or was breached | Packaging process control | Olympus America Inc. sent an "URGENT PRODUCT RECALL" letter dated December 22, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory for the affected product and to destroy or discard, if found. A questionnaire was attached to the letter for customers to complete and return via fax to Olympus at 484-896-7128.Contact the firm at 484-896-5688 for questions regarding this recall. |  | 1305 | 1305 devices | USA (nationwide) including the states of AL, AR, CA, CO, CT, FL, IA, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OR, TX, VA, VT, and WI. | Terminated | September 28, 2012 | 646 | 1305
Z-3101-2011 | 59262 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | On-Board Imager (OBI) 1.3, 1.4, 1.5; and Trilogy Mx, TrueBeam 1.0, Offline Review 1.0 - 2.0; |  |  | June 21, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified a software anomaly in the 3D Match environment, regarding a slight mismatch between CT image and Structures in On-Board Imager w/s. | Nonconforming Material/Component | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 21, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter recommends several user actions. The firm is developing a correction for the issue and will contact customers when the correction is available.Questions regarding this notice may be addressed by contacting the Varian Oncology Help Desk at 1-888-827-4265. |  | 2268 | 2268 | Worldwide Distribution | Terminated | March 15, 2012 | 268 | 2268
Z-3103-2011 | 59217 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BrachyVision, versions 8.2, 8.5, 8.6 and 8.9. |  |  | August 11, 2010 | 2010 | Varian Medical Systems, Inc. | Medical device for brachytherapy treatment may cause users to enter incorrect syntax and lead to potential misinterpretation of patient treatment plan. | Software design | Varian Medical Systems, Inc. sent a "Medical Device Recall" letter dated August 05, 2010 to all affected customers.The letter included description of product affected and problem. Customers were advised to review template files to ensure that there are no characters immediately after the string in the customized dose quality parameters. Device users were requested to return a response form.For information on this recall contact Varian Technical Support at (800) 360-7909. |  | 1252 | 1252 units | Worldwide and Nationwide Distribution. | Terminated | November 25, 2011 | 471 | 1252
Z-3113-2011 | 59632 | No | JXN | Propoxyphene test system. | Toxicology | Toxicology | 510(k) | DrugCheckÂż, Distributed by Express Diagnostics Int'l Inc., 1550 Industrial Drive, Blue Earth, MN 56013. Following cup Model numbers involved in the recall: 60670, 60700, 60713, 60717, 60870, 60923, 61012, 61019, 61023, 61024, 61040, 61042, 61057, 61102, 61105, 61203, 61204, 61206, 61208, 61301, 61303, 61401, 70603, 70809, 71032, 71043, 71055, 71065, 71202, 71204, 71205, 71401, 71402, 60718-3, 60720-6, 60925-A, 60927-3, 60930-A, 60980-3A, 61002-3, 61010-4, 61023-A, 61028-3A, 61039-4, 61040-A, 61055-A, 61056-3, 61057-6, 61100-3, 61102-3,61104-3,61105-5 CRL,61120-3,61202-4,61204-4,61204-A,61206-3A,61207-4,61302-3, 61303-A, 61401-A, 70660-3, 70716-3, 71045-A, 71055-6, 71102-3A, 71204-4, 71206-3A. |  |  | October 04, 2010 | 2010 | Express Diagnostics Int'l., Inc. | Express Diagnostics Int'l, Inc (EDI), is recalling DrugCheckÂż Drug Screen Cups containing buprenorphine (BUP) and propoxyphene (PPX) because they must be labeled as "forensic use only", meaning they cannot be distributed or used for treatment or diagnostic purposed in the United States. | Labeling False and Misleading | The firm, Express Diagnostics Int'l, Inc., sent a "RECALL NOTICE" letter dated October 4, 2010 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to notify their customers of the labeling change, and complete and return the DISTRIBUTOR/CUSTOMER CONFIRMATION FORM via fax to Express Diagnostics Int'l, Inc., at 1-507-526-2252. The letter also noted that "Further shipment of product containing BUP and/or PPX cannot continue until Express Diagnostics Int'l, Inc. receives the signed form."If you have any questions, call 507-526-3951. |  | 196750 | 196,750 | Nationwide distribution:AZ, CA, CT, GA, MD, MA, MN, MO, NH, NJ, OR, SC, and WA. | Terminated | July 06, 2012 | 641 | 196750
Z-3106-2011 | 59229 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Depuy Mitek Ratchet Handle with Quick Connect |  |  | June 27, 2011 | 2011 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Quick Disconnect Ratchet Handle may slip up the driver shaft or disengage leaving the shaft in the screw head. | Process change control | Depuy Mitek issued Original Customer Notification Letter on 6/27/11 and an Updated letter to clarify the lot number information sent 6/29/11. Hospitals requested to immediately check all inventories for affected product and return for repalcement product. Also notify surgeons aware of this potential issue and potential risk of a delay in surgery. questions or concerns with regard to this program, please contact your local DePuy Mitek Representative or DePuy Mitek Customer Service directly at 1-800-382-4682, PROMPT 3. |  | 75 | 75 units | Nationwide distribution. | Terminated | January 19, 2012 | 206 | 75
Z-3108-2011 | 59439 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | HiÂżArt versions 4.0.x, and HD versions 1.0.x, H-0000-0003. |  |  | July 13, 2011 | 2011 | Accuray Incorporated | As a result of an internal review, TomoTherapy has identified an issue with HiÂżÂżArt versions 4.0.x, and HD versions 1.0.x that we would like to bring to your attention.During DICOM export of plan level images with a non square exported Field of View (FOV), an anomaly in the process of squaring the plan level image may cause the image to shift with respect to ROIs and dose. When the anomaly occu | Software design | TomoTherapy sent a "Urgent Medical Device Correction" letter dated July 13, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This issue will be remedied with HiÂżÂżArt software version 4.1 and HD software version 1.1 and later.If you have any questions, please contact the TomoTherapy Technical Solutions Center by e-mail "support@tomotherapy.com" . Please refer to the letter for what phone number you should call depending on where you are. |  | 346 | 346 | Worldwide Distribution-- USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, MT, NE, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WV, WI and WY and the countries United Arab Emirates United Kingdom Turkey Switzerland Sweden Spain Saudi Arabia Poland Netherlands Italy Germany France Belgium Taiwan Singapore Malaysia Korea Japan India Hong Kong China Australia Mexico El Salvador and Canada | Terminated | March 18, 2014 | 979 | 346
Z-3104-2011 | 59567 | No | DIC | Alcohol test system. | Toxicology | Toxicology | 510(k) | Drug Check, Distributed by Express Diagnostics Intl Inc., 1550 Industrial Drive, Blue Earth, MN 56013. |  |  | November 01, 2010 | 2010 | Express Diagnostics Int'l., Inc. | Express Diagnostics Intl'l Inc is recalling DrugCheck devices because they were informed via letter from the alcohol test strip manufacturer on October 27, 2010 that the product was not 510 K cleared for in vitro diagnostic use and could not be labeled as such in the US. | Error in labeling | Express Diagnostics Intl., Inc. sent an Updated Recall Notice dated November 1, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed if they have open orders of the affected product to contact customer service about their order. Customers were asked to complete and fax the form on page 3 of the document to 507-526-2252. Further shipment of products could not continue until Express Diagnostics received the signed form.For questions regarding this recall call 507-526-3951. |  | 56650 | 56,650 | Nationwide Distribution including AZ, CA,K CT, FL, GA, IL, IA, LA, MD, MA, MN, MS, MO,K MT, NE, NJ, NY, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WY. | Terminated | July 06, 2012 | 613 | 56650
Z-3111-2011 | 59297 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Philips Medical: |  |  | July 08, 2011 | 2011 | Philips Healthcare Inc. | Under certain circumstances, edits to an order can result in unintended scheduled interventions on the IntelliVue Clinical Information Portfolio. A patient may receive interventions or additional medication doses that are not intended. | Software design | Philips Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers are asked to follow the guidelines in the "Action to be taken by Customer/User" section of the Urgent Medical Device Correction Notice. Philips has developed a software upgrade to address this problem in the short term. This upgrade will be available immediately. This software upgrade will be performed by Philips service personnel. While customers await the correction of their devices, customers are instructed to follow the instructions for use for the device, which includes the directions to:1. Always use the medication administration record (MAR) in determining patient care.2. Always review orders and pending administrations in the administration record after any changes to an order are stored.For the longer term, Philips is developing safeguards intended to minimize the potential for similar errors resulting from order changes and upgrade to be available in September 2011.If you need any further information or support concerning this issue, please contact your local Philips representative. |  | 454 | 454 units | Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY HONG KONG, IRELAND, ITALY, KENYA, LATVIA, MALAYSIA, NETHERLANDS, NORWAY, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 454
Z-3105-2011 | 59572 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Treatment version 6.6.5052. |  |  | July 20, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Potential for Varian Treatment application failing to load the physical wedge filter for a treatment field on a Siemens Mevatron linear accelerator. Delivery of a treatment field without the planned wedge would deliver more dose than intended and would change the shape of the dose distribution for that field. | Other | Varian Medical Systems sent out "Urgent Medical Device Correction/Urgent Field Safety Notice" letters dated July 20, 2011 to all affected customers.The letter identified product, problem and provided recommended user actions when setting up and running the system. It also asked customers to notify appropriate personal provided instructions.For information or questions on this recall please call Varian at (888) VARIAN5 (888- 827-4265). |  | 118 | 118 | Worldwide Distribution: Nationwide distribution including the states of AZ, CA, FL, GA, IN, MA, MD, MI, MN, MO, NJ, NY, PA, SC, TN, VA and WV; and the countries of Canada, Austria, Belgium, Brazil, Denmark, Finland, France, Germany, India, Ireland, Italy, Lithuania, Poland, Portugal, Romania, South Africa, South Korea, Spain, Sweden, Switzerland and UK. | Terminated | December 05, 2013 | 869 | 118
Z-3098-2011 | 59350 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand 4D Integrated Treatment Console (4DITC) Radiation Therapy, version 8.1 SP2 to 10.0 SP1; Model Number: H51; Reference/PSCA Identifier: CP-03940; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | June 23, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has scheduled for an upgrade of the device from current version to 10.2. There are circumstances in which the 4DITC may be involved in a failure to correctly record treatment history information. | Software design | Varian Medical Systems, Inc. sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 23, 2011 to all affected customers. The letter identified the product, problem and recommended user actions to follow. Customers were instructed to notify all appropriate personnel of the correction notice. The letter states that the recalling firm will issue a Customer Technical bulletin to inform users of the mandatory upgrade and related information. For future reference, this document will be posted to the Varian customer support website: http://www.MyVarian.com. Any questions regarding this Product Notification letter, contact Varian Medical Systems Help desk at 1-888-827-4265 for USA and Canada and +41417498844 for Europe. |  |  |  | Worldwide Distribution - USA (nationwide) | Terminated | March 15, 2012 | 266
Z-3107-2011 | 59470 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Li-Ion Docking Station; a lithium ion battery powered docking station for use with GemStar Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; List Numbers 13075-04-01, 13075-04-03, 13075-04-05, 13075-04-07, 13075-24-01, 13075-24-03, 13075-24-05 and 13075-24-07, Note: list numbers with xxxxx-04-xx are for U.S., Canada, Japan, Netherlands and Singapore, while those with xxxxx-24-xx are for international distribution only. |  |  | August 01, 2011 | 2011 | Hospira Inc. | Hospira has received reports of sparking, smoking, charring and electrical shock when using the GemStar Docking Station due to fluid ingress into the docking station. | Device Design | Hospira Inc. sent an "URGENT DEVICE FIELD CORRECTION" letter dated August 1, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The requests that customers do not hang or place fluid containers over the docking station, not spray fluid directly on the docking station, and to clean the docking station using a cloth dampened with an apporoved cleaning solution. A reply form was attached to the letter for customers to complete and return via fax at 262-577-6921.Contact Hospira at 1-800-441-4100 for questions concerning this notice.*** Hospira sent an update letter dated September 26, 2012 to the accounts, informing them of the availability of the improved design replacement docking stations, and requesting them to complete and return the enclosed reply via fax or e-mail to Hospira, indicating the number of docking stations to be returned for replacement at no charge if the old docking stations are returned within 45 days of receipt of the replacement stations. *** |  | 18000 | 18,000 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Belgium, Canada, Chile, France, Germany, Hong Kong, Ireland, Italy, Japan, Jordan, the Netherlands, Portugal, Saudi Arabia, Singapore, Spain, Sweden and Switzerland. | Open, Classified |  |  | 18000
Z-3114-2011 | 59777 | No | LZC | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | MICRO-TOUCH(R), S (indicating size small), NITRILE, Powder Free, Medical Examination Gloves, REF 6034301, 200 Pieces, Manufactured for: Ansell Healthcare - Made in Malaysia |  |  | August 04, 2011 | 2011 | Ansell Healthcare Products LLC | Gloves were inadvertently released for distribution prior to completion of testing required by the U.S. Food and Drug Administration. | Release of Material/Component prior to receiving test results | Ansell account managers were notified on August 4, 2011 to contact (by phone, fax, e-mail) their respective accounts and notify them of the recall situation. The Ansell account managers were provided with a copy of the recall notification to provide to their accounts. The recall notification includes a request that if these gloves have been further distributed to other customers, please advise them of the recall situation and have them return any outstanding stocks. |  | 3700 | 370 cases of 10 | Nationwide distribution: NC, GA, IL, MO, LA, and PA | Terminated | November 27, 2012 | 481 | 14420
Z-3112-2011 | 59345 | No | NTI | Drug metabolizing enzyme genotyping system. | Toxicology | Toxicology | 510(k) | xTAG CYP2D6 Kit v3 IVD For use with Luminex 100/200 Instrument Luminex Molecular Diagnostics, Inc. 439 University Ave. Toronto, Ontario Canada Catalog # I030B0373 |  |  | May 24, 2011 | 2011 | Luminex Corporation | Inaccurate test results, due to possible defective primers, which could lead to incorrect dosage on manufactured drug assays. | Other | Luminex sent a "Customer Advisory" letter dated May 24, 2011, by email and courier to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to destroy any remaining product and to submit a return postcard to confirm that the consignee had been notified of the recall and had destroyed the product.For further questions please call Luminex Technical Support at 1-877-785-2323 ( US ) and 1-512-381-4397 ( Outside US ) Monday through Friday. |  | 45 | 45 kits | Worldwide Distribution -- USA ( nationwide) and the countries of Belgium, UK and Canada. | Terminated | March 23, 2012 | 304 | 45
Z-3110-2011 | 59436 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Cell Marque antibody CD99 (H036-1.1) |  |  | July 12, 2011 | 2011 | Cell Marque Corporation | Incorrect label on product bottle, although labeling on outer box is correct. | Labeling mix-ups | The firm, Cell Marque, notified customers by telephone on July 12, 2011. The customers were instructed to return the product. Cell Marque immediately issued replacement of the product.If you have any further question or require additional documentation, contact the VP of Regulatory and Quality Affairs at (916) 746-8900. |  | 2 | 2 units | Worldwide distribution: USA (nationwide) include state of : CA and country of: Great Britain. | Terminated | August 31, 2011 | 50 | 2
Z-3109-2011 | 59366 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | REF 91200, Juggerknot system, Soft Anchors, -Two implants size 1, Biomet Sports Medicine, STERILE. |  |  | July 21, 2011 | 2011 | Biomet, Inc. | The device packaging was punctured by the device, resulting in loss of implant sterility. | Package design/selection | Biomet Sports Medicine notified distributor accounts by telephone on July 08, 2011 with instructions to quarantine all lots of the recalled product. Urgent Medical Device Recall Notice letters dated July 21, 2011 were sent to all affected customers. The letter identified the recalled product, problem, and actions to be taken. The customers were instructed to immediately locate and discontinue use of the recalled product. The distributors are advised to communicate this notice to their customers if the recalled products were further distributed. The letter also states that distributors are responsible for the return of the recalled product. Customers were asked to complete a response form and fax to 574-372-1683, within 3 business days. Customers were instructed to provide phone confirmation at 800-348-9500, extension 3755 or 3756, when the recall notice is received. Further questions should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8 AM to 5 PM. |  | 4925 | 4925 | Worldwide Distribution - USA (nationwide) and the countries of: Columbia, Europe and Korea. | Terminated | April 26, 2013 | 645 | 4925
Z-3118-2011 | 59570 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Reliance Vision Single Chamber Washer/Disinfector, Model # 1802983, |  |  | May 09, 2011 | 2011 | Steris Corporation | STERIS learned that Vision SC Washer users were experiencing nuisance alarms and/or aborted cycles that are adversely affecting the reliability and operation of their unit. | Pending | The firm, STERIS sent an "URGENT FIELD CORRECTION NOTICE" letter dated May 9, 2011 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed that they can continue using their Vision SC Washer(s) prior to installation of the system upgrades. The letter states that STERIS is correcting the issues by implementing a series of quality improvements to the customers Vision SC Washer(s) to ensure satisfaction. Additionally, a STERIS Service Representative will contact the customers facility to arrange for STERIS to upgrade their Vision SC Washer(s).For further information or if you have questions regarding the servicing of your unit, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 404 | 404 Units | Worldwide distribution: USA (nationwide) including states of: L, AR, AZ, CA, CO, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI and WY; and countries of: Australia, Canada, Hong Kong and South Korea. | Terminated | March 01, 2013 | 662 | 404
Z-3125-2011 | 58705 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Unicel DxH 800 Coulter Cellular Analysis System |  |  | May 09, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the probe wash collar tubing may become stretched or trapped at the Sample Aspiration Module (SAM) during routine use of single-tube presentation station. When the tension on the tubing is severe, the tubing may not be able to return to its home state and may pull the wash collar out of alignment with the aspirate probe. The misal | Pending | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated May 11, 2011 with an attached Product Correction Response Form to all affected customers. The letter provides the customers with an explanation of the problem identified, potential impact should the problem occur and actions to be taken. The letter states that this issue will be corrected in a future hardware release. Customers were instructed to share this information with laboratory staff and retain this notification as part of Quality System documentation. If ownership or location of the analyzer(s) has been transferred to another laboratory, customers were instructed to provide a copy of the letter to that party. Customers were also asked to complete and return the enclosed Response Form within ten (10) days, so Beckman can be assured they have received this important communication. Questions concerning this notice, were directed to the Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 680 | 680 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Australia, Belgium, Canada, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Republic of Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | June 18, 2012 | 406 | 680
Z-3121-2011 | 59232 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, 1.5T Optima MR450w |  |  | July 29, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of two potential safety issues associated with MR products (Optima MR450w, Discovery MR450, Discovery MR750, Brivo MR355, Optima MR360) that may impact patient safety.1) Inline post-processed (i.e. automated post-processing) image locations are not correctly annotated under some conditions. 2) Reference images displayed in Functool FiberTrak do not match with | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated July 27, 2011 to customers with affected systems. The letter described the recalled product, safety issues and recommended actions to be taken. The letter provided immediate mitigation (Safety Instuctions) that address each issue. The letter states that the recalling firm will contact customers to arrange to provide a software upgrade to all affected systems at no cost. For any questions concerning this notification contact your local service representative. |  | 159 | 159 | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, LA,ME, MD, MA, MI, MN, MS, MO, MT, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VT, VA, WA, WV, WT, WY.and the countries of: AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CROAT IA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KUWAIT, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, REPUBLIC OF KOREA, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TUKEYand UNITED KINGDOM. | Terminated | August 07, 2013 | 740 | 770
Z-3124-2011 | 59752 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | SYSTEM 1E Liquid Chemical Sterilant Processing System, Model # P6500, S/N: 400001 through 405005; Mfr. By: STERIS Corporation, Mentor, OH 44060. |  |  | August 11, 2011 | 2011 | Steris Corporation | STERIS learned that SYSTEM 1E users may experience aborted cycles due to a failure within the UV light system. | Other | STERIS sent an "Urgent Field Correction Notice" letter dated July 2011 to all affected customers. The letter describes the product, problem and actions to be taken. The letter states that STERIS Service Representative will contact customers to schedule the repair of the recalled product. For further information or questions contact STERIS Field Service Dispact at 1-800-333-8828. |  | 2235 | 2,235 units | Worldwide Distribution - USA (nationwide) and the country of Bahrain. | Terminated | March 19, 2013 | 586 | 2235
Z-3119-2011 | 59691 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Operations and Technical Manuals for the firm's WarmAir Model 135 Hyperthermia System (115V, 100V, and 230/240V), 510 (k) #K101148, Serial Numbers: 994-1350001 through 104-1356894. The WarmAir 135 Hyperthermia System is a patient temperature management device which provides forced air warmed by the controller to ... a blanket that is placed over or around adult, pediatric or neonatal patients in order to warm them. The heated air is blown through a connecting flexible hose to a disposable blanket to provide patient therapy by the means of warmed air. The system can be used in long-term care facilities, surgical facilities, hospitals including the Post-anesthesia Care Unit (PACU), Intensive Care Unit (ICU), Surgical Intensive Care Unit (SICU), Emergency Room (ER), Operating Room (OR), medical and surgical floors, or any other department or hospital facility requiring patient temperature management. |  |  | July 20, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | Cincinnati Sub-Zero is conducting a field correction due to changes that have been made to the WarmAir Model 135 devices Operation and Technical Manuals in order to comply with FDA device labeling regulations. | Labeling design | Cincinnati Sub-Zero Products, Inc. sent an 'URGENT Medical Device Recall Field Action Notice" dated July 27, 2011, to all affected user/customers. The letter includes: description of device; the problem with the suspect manuals, and provides contact information for the customers to use to request new versions of the manuals. Upon receipt of the new manuals, the customers are asked to immediately collect and destroy/discard all of the old manuals at their facility and to replace them with the new versions. After the manuals have been discarded and replaced, the customers are asked to complete and return the attached "Recall Response Form" via fax to (513) 772-9119 or scan and email to WA135manual@cszinc.com or mail to Cincinnati Sub-Zero, 12011 Mosteller Road, Cincinnati, OH 45241.For additional information the customers are provided with an e-mail address: WA135-2011@cszinc.com as well as a toll-free telephone number and fax number to reach the firm and obtain additional information as necessary. The contact numbers are: 1 (800) 989-7373, or (513) 772-8810, and/or fax # (513) 772-9119.. |  | 6561 | 6561 units were distributed. | Nationwide distribution: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NM, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, and WI. | Terminated | July 05, 2013 | 716 | 6561
Z-3031-2011 | 59335 | No | JKD | C-peptides of proinsulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT C-Peptide Reagent Kit; a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of C-Peptide in human serum, plasma and urine on the ARCHITECT i System; Produced by Biokit S.A., 08186 Barcelona, Spain for Abbott Diagnostics Division, Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA and Abbott, 65205 Wiesbaden, Germany; list 3L53-25 |  |  | July 11, 2011 | 2011 | Abbott Laboratories | C-Peptide lots 03610K000 and 01711A000 have the potential to generate falsely elevated results with certain patient samples and with non-Abbott Controls. | Nonconforming Material/Component | Abbott Laboratories sent a "PRODUCT RECALL" letter dated July 11, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to discontinue use of the lots and destroy any remaining inventory of the lots according to laboratory procedures. A Customer Reply form was provided to customers to complete and return via fax at 1-800-777-0051.Customers can contact Abbott Customer Service at 1-877-4ABBOTT for questions regarding this notice. |  | 3049 | 3,049 kits | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Algeria, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, India, Indonesia, Ireland, Italy, Japan, Jordan, Latvia, Libya, Lithuania, Luxembourg, Malaysia, Mauritania, Mexico, Netherlands, Oman, Pakistan, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Korea, Spain, Switzerland, Taiwan, Trinidad &amp; Tobago, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | February 10, 2012 | 214 | 3049
Z-3128-2011 | 59685 | No | NDR | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Inhibin A ELISA, Part Number: DSL-10-28100-4 |  |  | November 10, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed the microbial contamination in the conjugate diluent bottles included in the Manual Inhibin A Elisa kit (Part Number DSL-10-28100-4), kit lot number 110761. | Other | Beckman Coulter initiated an "URGENT PRODUCT CORRECTION" on June 15, 2011, via a telephone script to the only one customer impacted by the recalled product. The telephone script provided an explanation of the problem identified and the actions to be taken. Beckman Coulter recommended that the customer discontinue use of the recalled product as a preventative measure. Beckman will send a representative to the customer's facility to replace the recalled product in all remaining kits of the affected lot. The customer was instructed to share this information with laboratory staff. For any questions concerning this information, contact Beckman Coulter at 952-368-7874. |  | 342 | 342 units | (USA) Nationwide Distribution in the state of CA | Terminated | July 11, 2012 | 609 | 342
Z-3129-2011 | 59601 | No | LJC | Chlamydia serological reagents. | Microbiology | Microbiology | 510(k) | VIDAS Chlamydia test kits |  |  | July 08, 2011 | 2011 | bioMerieux, Inc. | Complaints were received of high rates of false positive results, leading to unnecessary antibiotic therapy being prescribed to patients. | Other | bioMerieux sent an "URGENT PRODUCT REMOVAL NOTICE" letter dated July 8, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue using and to destroy any of the affected product. The firm advises customers to confirm all positive results obtained with the affected lot number with another method prior to reporting results. An Acknowledgement and Product Replacement form was attached for customers to complete and return to the firm.Customers should contact their local bioMerieux Customer Service representative for questions regarding this notice. |  | 283 | 283 units | Nationwide Distribution- including CA, DC, FL, IL, IL, IN, MD, MI, MN, MO, MT, NC, NE, NJ, NY, PA, UT, VA and WI. | Terminated | October 23, 2012 | 473 | 283
Z-3130-2011 | 58613 | No | MIH |  |  | Cardiovascular | PMA | Cook Medical Zenith Flex AAA Endovascular Graft Main Body, REF TFFB-24-82-ZT, REF G48407, 18 Fr ID/ 7.1 mm OD, Sterile |  |  | April 15, 2011 | 2011 | Cook, Inc. | The firm initiated this action in response to information that there may be potential damage in the braided suture which is used to attach the external stents to the graft material. | Process control | The firm sent URGENT PRODUCT RECALL notices dated 4/15/2011 to Risk Managers/Recall Administration (Chief of Vascular Surgery and or Interventional Radiology). The letters explained the problem with the devices and the action to be taken. Consignees were instructed to return the recalled devices to Cook Medical for credit or replacement. Customers with questions or concerns can contact the firm at 800-457-2235. |  | 4 | 4 units | US, Canada, Japan | Terminated | September 13, 2011 | 151 | 71
Z-3159-2011 | 59659 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT II TURBO Link Hemoglobin A1c Program Reorder Pack. |  |  | July 28, 2011 | 2011 | Bio-Rad Laboratories, Inc. | The Elution Buffer B, lot no. GA 10430, was shipped in a 2.0 L bottle instead of the standard 2.5 L bottle. | Employee error | The firm, BIO-RAD, sent a "Medical Device Correction" letter dated July 28, 2011 to all consignees/customers via Fed Ex on August 8, 2011. The letter described the product, problem and actions to be taken. The customers were informed that they could continue to use and report results with the affected buffer lot and the stability studies will continue until lot expiration date. The consignees were instructed to distribute this letter to all customers that have received the affected lot and to distribute the MEDICAL DEVICE CORRECTION RESPONSE FORM only to those customers that request replacement Elution Buffer B. In addition, the customers were to complete and return the MEDICAL DEVICE CORRECTION RESPONSE FORM via fax to (510) 741-3954, Attention: BIO-Rad CSD Regulatory Affairs Department.If you have any questions, please contact your regional Bio-Rad office or 510-724-7000. |  | 40 | 40 | Nationwide distribution: MD, MI, MN, NY, OH, SC, VA and WI. | Terminated | November 01, 2011 | 96 | 40
Z-3120-2011 | 59692 | No | FKG | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Custom dialysis prep kits labeled as follows: |  |  | March 22, 2011 | 2011 | Lee Medical International Inc | Firm manufactured custom dialysis kits with Povidone Iodine (PVP) Prep Pads which were recalled by H &amp; P industries. Those pads were subsequently recalled due to possible bacterial contamination. | Material/Component Contamination | The firm, Lee Medical International Inc., notified their consignees by email and telephone on 03/22/2011. Consignees were instructed to pull product from stock, quarantine for re-processing, and return or destroy product. Lee Medical arranged for replacement product from alternate manufacturer to be sent to customers.For questions call Production Manager or Operations Manager at 504-734-9336 or email melissal@eleemedical.com or chriss@eleemedical.com. |  | 19625 | 19,625 units | Nationwide Distribution (USA) - including states of: AZ, CA, FL, GA, MD, NJ, NM, NY, PA, and TX. | Terminated | January 30, 2012 | 314 | 19625
Z-3160-2011 | 59737 | No | NSX |  |  | General Hospital | Contact ODE | Picis ED PulseCheck-Software, transmission &amp; storage, patient data |  |  | August 01, 2011 | 2011 | Picis Inc. | Software does not display appropriate allergy interaction warning | Software design | Picis issued notification on 08/01/11 to affected clients, explaining the potential for the error, with instructions to monitor documented patient allergies closely when prescribing medication orders in ED PulseCheck, using Muitum drug database content. Picis will track the progress on acknowledgement of notification letter receipt through the implementation of software upgrade, for all potentially affected Customers.For questions regarding this recall call 781-557-3000. |  | 62 | 62 | Nationwide Distribution | Terminated | June 27, 2013 | 696 | 62
Z-3162-2011 | 58473 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Ratcheting Torque Limiting Handle, Catalog Number 52-1012, Orthofix Spinal Implants, Firebird Spinal Fixation System, Firebird Spinal Fixation System DDD Instrument Case. |  |  | April 01, 2011 | 2011 | Orthofix, Inc. | Upon reaching torque limit, there is a possibility the product may jam or exhibit a ratchet direction switch from forward to reverse. | Device Design | Orthofix, Inc. sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated April 1, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to identify and remove the affected product from use. Customers were advised to use the combination of the Ratching R Handle and the Set Screw Driver if they experience problems with the Ratching Torque Limiting Handle. A Tracking and Verification form was attached for customers to complete and return to the Orthofix Inc. Regulatory Affairs Department via fax at 214-937-3322.Contact your Account Services Representative at 888-298-5700 for questions regarding the removal and return of the affected product. |  | 287 | 287 units | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Australia and Germany. | Terminated | March 23, 2012 | 357 | 287
Z-3161-2011 | 59305 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TiLite ZR and ZRA Series 2 are mechanical wheelchairs. Devices are sold under brand name TiLite. |  |  | July 05, 2011 | 2011 | TiSport, Llc | The Pivot Stud Bolt comprising part of the backrest assembly of TiLite ZR Series 2 and ZRA Series 2 mechanical wheelchairs have shown a tendency to fail due to the head of the bolt being sheared off. | Component design/selection | An Urgent Medical Device Correction letter, Product Correction Kit , and the Installation Instructions were sent to domestic suppliers, international suppliers, Veterans Administration hospitals and international distributors by US Priority Mail and by FedEx three days service.The letter identified the affected product and the reason for the recall. Customers were instructed not to send the product to their end user until the correction was made or to identify their end user and notify the end user of the product correction. Replacement parts and instructions to correct the issue were enclosed with the letter. Customers were to destroy the removed parts once they had completed the correction. Customers were to also return the Medical Device Correct Report that was provided, by fax to 509-586-2413, or email to ProdCorrection@tilite.com.Customers can call Technical Support at 800-545-2266, ext. 239 with any questions about recall. |  | 1037 | 1,037 units | Worldwide Distribution | Terminated | July 27, 2012 | 388 | 1037
Z-3172-2011 | 59453 | No | MCY | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Wound Dressing Kit - Tier I, 10 MM Flat Drain; product code 0877061. |  |  | July 27, 2011 | 2011 | Medela Inc | Medela determined that some of the individual kits or bags of lot #370118 of "0877061 Wound Dressing Set, Tier I, 10 MM Flat Drain", were labeled incorrectly as "0877067 Wound Dressing Set, Tier III 15 Fr Round Channel Drain". The outside label on the case is correctly labeled as "0877061 Wound Dressing Set, Tier 1,10 MM Flat Drain". | Error in labeling | Medela Inc. sent an "Urgent Device Recall" letter dated July 26, 2011 to all affected customers.The letter described the product, problem and actions to be taken. The customers were instructed to do the following: 1) not to use the affected lot and place them in quarantine 2) stop sale and distribution and set aside their existing inventory of the affected product 3) conduct a physical count and record the data on the Product Inventory Form included with this letter 4) complete and return the Product Inventory Form via fax to 800-995-7867 or e-mail suction@medela.com, even if they have no affected product on hand 5)if they have the affected product in stock, Medela Customer Service will ship substitute product to the customer at no cost. 6) and to record the number of units on hand for return and replacement on the enclosed Product Inventory Form. The completed form should be returned via fax or e-mail to Medela even if the account has no inventory of the affected lots. If you have any questions about this recall please call (877) 735-1626 or e-mail at suction@medela.com. |  | 3000 | 3,000 kits | Nationwide Distribution including the states of California, Florida, Georgia, Illinois, Kansas, Kentucky, Massachusetts, Michigan, Mississippi, New Jersey, New Mexico, North Carolina, Ohio, Pennsylvania, South Carolina, Tennessee and Texas. | Terminated | February 16, 2012 | 204 | 3000
Z-3164-2011 | 59798 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DUET TRS 45 3.5mm Universal Straight Single Use Loading Unit Surgical Stapler with Implantable Staples |  |  | August 15, 2011 | 2011 | Covidien LP | Sterility of the device is compromised due to a breach in the sterile packaging | Storage | The firm, Covidien, sent an "URGENT MEDICAL DEVICE RECALL" letter dated August 15, 2011, to its customers via federal express. The letter describes the product, problem, and actions to be taken. The customers were instructed to immediately identify and quarantine all affected inventory (Note: Units from the affected lots may have been incorporated into Covidien BEST PRACTICE KITS. If you have purchased BEST PRACTICE KITS, please evaluate whether these kits contain units from the affected lots.); complete and return the DUET TRS Universal Loading Unit Recalled Product Return form via fax to (203) 822-6009; and return all inventory from the affected lots (note: All affected units must be returned) with the complete form to: Field Returns Department, located at 195 McDermott Road, North Haven, CT 06473. Questions or concerns, contact Covidien Representative or Covidien Quality Assurance at (203) 492-5232. |  | 4362 | 4362 units | Worldwide distribution: USA (nationwide) and countries of: Asia, Australia, Canada, Europe, Japan, Latin America, and Pacific. | Open, Classified |  |  | 457461
Z-3173-2011 | 59830 | [6 empty] | BardPort Titanium Implanted Port. |  |  | August 30, 2011 | 2011 | Bard Access Systems | Implanted port has discrepant product information on package labeling. | Labeling Change Control | Bard Access System, Inc. sent a "Label Discrepancy Notification dated August 30, 2011 to all affected customers.The letter included affected product and informed them of the labeling discrepancy. Customers could either use the device or return it to Bard for replacement. For information on this recall call (801) 595-5568. |  | 1375 | 1375 units | Worldwide Distribution: Nationwide Distribution including the states of; AL, AR, AZ, CA, CT, DC, FL, GA, HI, IL, KY, LA, MA, MD, MI, MO, NJ, NV, NY, OK, PA, RI, SC, TN, TX, UT, VA, WI, and WV; and the countries of Canada and Belgium. | Terminated | March 30, 2012 | 213 | 1375
Z-3175-2011 | 58425 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Kyphon brand KyphoPak Express Tray First Fracture, 15/2 with KyphX Osteo Introducer System; Catalog Number: KPE1003; Product is manufactured and distributed by Medtronic Spine LLC (formerly known as Kyphon Inc), Sunnyvale, CA |  |  | March 23, 2011 | 2011 | Medtronic Spine LLC (formerly Kyphon Inc.) | Device was shipped past the labeled expiration date. | Employee error | On 3/23/2011, the firm initiated the recall and sent Urgent Notice: Device Recall letters identifying the affected product and explaining the reason for the recall. The letter asks customers not to use the affected product with the lot number 0001653390. The firm will replace any affected product at no cost to the customer. A Medtronic Spine Sales representative will assist with the segregation and return of the affected product. They will also assist the customer with contacting Medtronic Spine Customer Service to return the affected product. Included with the letter was a reconciliation form with instructions regarding the affected product. Questions regarding the letter should be directed to Customer Service at 866-959-7466. |  | 44 | 44 | Nationwide Distribution | Terminated | September 09, 2011 | 170 | 44
Z-3183-2011 | 59421 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Patient Management System Communicator, Model 6482. |  |  | May 25, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has identified two occurrences where the LATITUDE system did not process a data payload, which is the information uploaded to the LATITUDE system from a remote interrogation of the implanted device. These occurrences resulted in non-displayed payloads because the data is not posted to the website. If the payload is not displayed, the pertinent data will not be available on the | Device Design | Boston Scientific notified all affected customers by telephone on May 25, 2011. The customers were informed that there was an interrogation for one of their patients that was not able to be displayed on the website. Boston Scientific detected that a payload failure had occurred, but could not conclude what information was contained within the payload. Each patient has had previous and subsequent payloads that have processed and displayed on the LATITUDE website. |  | 2 | 2 | Nationwide Distribution-in the states of Florida and Texas. | Terminated | December 04, 2012 | 559 | 2
Z-3176-2011 | 59455 | No | KRI | Cardiopulmonary bypass accessory equipment. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Sarns Level Sensor II Pads. |  |  | August 12, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Level Sensor II mounting pads may not stick to the venous reservoir and may detach either at case setup or during a case. | Incorrect or no expiration date | Terumo Cardiovascular Systems Corp. sent an "Urgent Medical Device Removal" notice dated August 10, 2011 to all affected customers.The notice instructs customers to discard all remaining level sensor pads that do not have an expiration date on the package, and notify Terumo of the number of affected pads via a response form. Terumo will replace recalled product.For questions on this recall call Terumo Customer Service at (800) 521-2818, 8 am-6 pm EST. |  | 1531440 | 1,531,440 devices | Worldwide Distribution -- Nationwide Distribution, and the island of Puerto Rico. | Terminated | June 20, 2012 | 313 | 1531440
Z-3181-2011 | 59677 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Leica Bond Ready-to-Use Primary Antibody CDX2 (AMT28); an in vitro diagnostic immunohistochemistry primary antibody kit; Catalog No. PA0535; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom. |  |  | July 26, 2011 | 2011 | Leica Microsystems, Inc. | There was an issue with the controlled temperature storage unit where the products were stored prior to shipment to the customers. | Employee error | Leica Microsystems, Inc. notified all affected customers by telephone on July 26, 2011. The product, problem, and actions to be taken by the customers were discussed. Customers were instructed to discontinue use of the product and to return the product. All affected product has been returned to the firm. |  | 3 | 3 units | Nationwide Distribution-in the states of Indiana and Kentucky. | Terminated | September 09, 2011 | 45 | 4
Z-3177-2011 | 59759 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Monomid Non-Absorbable Suture are blue, monofilament synthetic absorbable suture, with different sizes and different needle sizes. |  |  | June 29, 2011 | 2011 | C P Medical Inc. | Unapproved dye was used in manufacturing the non-absorbable blue nylon suture material. | Process control | CP Medical sent a "RECALL NOTICE" letter dated June 29, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to discontinue use of the affected products in the human market and to return all their existing stock. A Recall Checklist was enclosed for customers to complete and return.Contact Customer Service at 1-800-950-2763 for a Return Good Authorization and shipping account number. The firm will pay for shipping cost and replace product at no cost to customers. |  | 776 | 776 dozens boxes | Worldwide Distribution-USA (nationwide) and the countries of Italy, Romania, South Africa, and Taiwan. | Terminated | May 24, 2012 | 330 | 776
Z-3174-2011 | 59564 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Coloplast Corp.-MN 06093, Country of origin: DK ( SenSura 19011 Post-Op, ostomy bag with window, 100 mm , Sterilized using ethylene oxide; SenSura 19010 Post-Op ,ostomy bag with window, 70 mm , Sterilized using ethylene oxide; SenSura 19001 Drainable bag with soft tap 1 piece/Opaque with inspection windows, 10-76 mm Sterilized using ethylene oxide, Manufacturer: Colopast A/S 3050 Denmark.) |  |  | July 27, 2011 | 2011 | Coloplast Manufacturing US, LLC | On July 19, 2011 Coloplast became aware of the failure or routine sterility testing for one lot of SenSuera and Assura Post-Op Ostomy Bags. Based on these results it was determined that all affected lots be recalled and quarantined. This affects lots manufactured and distributed from January 2010 until July 19, 2011. | Process control | Coloplast sent a "RECALL OF STERILE SENSURA AND ASSURE POST-OP POUCHES" letter dated July 27, 2011 to all affected customers. An enclosed Field Safety Notice described the product, problem, and actions to be taken by the customers.Contact your Customer Service representative at 800-533-0464 for questions regarding this recall. |  | 4020 | 4020 | Nationwide Distribution--USA (nationwide) including the states of AL, AZ, CA, FL, GA, LA, MA, MD, MI, MN, MO, NJ, NM, NY, NC, OH, OK, PA, RI, TN, TX, VA, WA, and WI. | Terminated | June 05, 2012 | 314 | 4020
Z-3179-2011 | 59656 | No | KQT | Gastroenterology-urology evacuator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Olympus Fluid Aspiration Set; a sterile set of silicone tubing and connectors for use with the Olympus SP-2 Fluid Irrigation &amp; Aspiration Pump; 10 sets per box; Manufactured for: Olympus America Inc., Two Corporate Center Drive, Melville, NY 11747-3157; catalog #72-00114-0 |  |  | July 25, 2011 | 2011 | Northgate Technologies, Inc. | The Red Tag was assembled to the wrong end of the silicone tubing. The Red Tag indicates how the fluid aspiration tubing set is loaded into the Olympus SP-2 Fluid Irrigation and Aspiration Pump. | Employee error | The firm, Northgate Technologies, sent an "URGENT: MEDICAL DEVICE RECALL" letters dated August 18, 2011 to the Olympus America customers who ordered the Aspiration Tubing Set, 72-00114-0. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the product subject to recall for subsequent return to Northgate Technologies; if they further distributed this product, they were to identify and notify their customers of the recall, and complete and return the enclosed Return Response Form via fax to (847) 608-9405 or mail to: Northgate Technologies Inc., 1591 Scottsdale Court, Elgin, IL 60123. If you have any questions, please call (224) 856-2222 or e-mail to: ckurek@northgate-tech.com. |  | 10680 | 10,680 sets | Nationwide distribution: USA including states of: CA, FL, ID, IL, MI, NY, PA, SD and WV. | Terminated | December 14, 2011 | 142 | 10680
Z-3185-2011 | 55319 | No | MHZ |  |  | Neurology | Contact ODE | The 91220 mCare 300 Vital Signs monitor has a 10.4 inch resistive touchscreen display. |  |  | April 14, 2010 | 2010 | Spacelabs Healthcare, Incorporated | There have been reports that the Spacelabs Medical Vital Signs Monitor, Model 91220, will not power up with AC power and will no longer charge the batteries. | Component design/selection | SpaceLabs Healthcare sent an URGENT - MEDICAL DEVICE CORRECTION letter dated April 15, 2010, to all affected customers. International consignees were notified by e-mail on April 22, 2010, with the customer notification letter. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to weigh the benefits versus the risks when deciding whether or not to continue to use the device until their monitors can be upgraded. Customers were asked to brief their staff about the possibility of the monitor failure to help prevent an adverse event from happening. |  | 614 | 614 units | Worldwide Distribution - USA (nationwide) and the countries of UNITED ARAB EMIRATES, ARGENTINA, AUSTRALIA, BELGIUM, BAHRAIN, BRAZIL, CANADA, CHINA, COLOMBIA, GERMANY, ECUADOR, EGYPT, FINLAND, FRANCE, UNITED KINGDOM, GREECE, HONG KONG, JORDAN, REPUBLIC OF KOREA, KUWAIT, MEXICO, MALAYSIA, OMAN, POLAND, PUERTO RICO, PARAGUAY, ROMANIA, SAUDI ARABIA, SINGAPORE, SYRIAN ARAB REPUBLIC, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, and URUGUAY. | Terminated | September 12, 2011 | 516 | 614
Z-3187-2011 | 59764 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ENDOPATHÂż ETS Compact Flex (240 mm shaft length) Linear Cutters, Model Number(s): SCW45; CTS45; SCB45; CTS45NK; 6CB45 and SCG45. Mfr. By: Ethicon Endo-Surgery, LLC, Guaynabo, PR 00969 |  |  | July 29, 2011 | 2011 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated a voluntary global recall for specific production lots of ENDOPATHÂż ETS Compact Flex 45mm articulating linear cutters because the possibility exists that damage to the blister pack may have compromised the sterility of the device. | Packaging | Ethicon Endo-Surgery Inc sent an URGENT DEVICE RECALL letter dated August 11, 2011 to all affected customers. The letter identified the product, the problem, the action to be taken by the customer. Customers were instructed to examine their inventory immediately to determine if they have any of the affected product. Fill out the Business Reply Form and fax it to 1-866-853-2802 if they have any of the affected product. To ensure adequate identification of the return product at their facility customers should complete the Customer Purchase Order Field on the Business Reply Form. For any questions regarding this recall call 1-866-708-2330 and reference Event 2471. For clinical or product support call 1-800-USE-ENDO, option 6, or contact your local sales representative. |  | 27582 | 27,582 instruments | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI and WV and the countries of ARGENTINA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, CZECH REPUBLIC, GREECE, INDIA, ITALY, JAPAN, REPUBLIC OF KOREA, LIBYAN ARAB JAMAHIRIYA, MALAYSIA, MEXICO, POLAND, PUERTO RICO, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, THAILAND and UNITED ARAB EMIRATES | Terminated | February 15, 2013 | 567 | 27582
Z-3186-2011 | 59477 | No | DRA | Steerable catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic FlexCath, model 3FC12, Steerable Sheath. The FlexCath Steerable Sheath is a percutaneous catheter introducer fitted with a valve to allow for introduction, withdrawal, and swapping of catheters and wires while preventing air ingress and minimizing blood loss. A side-port with stopcock is integrated to allow continuous drip infusion, injection through the center lumen, flushing, aspiration, blood sampling and pressure monitoring. Rx only. Sterilized using ethylene oxide. Manufacturer: Medtronic CryoCath LP, Kirkland, Quebec H9H 5H3, Canada. |  |  | July 15, 2011 | 2011 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic CryoCath would like to inform you of an observation related to a potential leak in the hemostatic valve of the Flex Cath 12 Steerable Sheath, Model 3FC12. There wave been reports of blood and saline solution leaking out of, and air ingress into the hemostatic valve. This includes reported instances of the introduction of air bubbles through the valve during aspiration of the side port | Pending | Medtronic Sales Reps began on July 15, 21011 hand delivering an "IMPORTANT MEDICAL DEVICE iNFORMATION" letter dated July 2011 to consignees. The letter identified the product, the problem, and the actions needed to be taken. The letter stated the rate of occurrence, described the potential clinical implications, and reinforced the importance of following the information found in the instructions for use as consignees continue to use the device. The letter also stated that Medtronic CryoCath is working urgently on a number of fronts to mitigate this observation and will continue to communicate with consignees as appropriate.For further information, please contact Medtronic VP of Quality at (763) 526-0576. |  | 6712 | 6712 | Worldwide Distribution - USA (nationwide) including the states of: AZ, CA, CO, FL, GA, IA, IL, IN, KY, LA, MA, MI, MO, MN, MS, NC, ND, NE, NJ, NM, NY, OH, OK, PA, SC, TN, TX, VA, WA, WI. and the countries of: Australia, Austria, Belgium, Canada, Czech, Denmark, France, Germany, Greece, Hong Kong, Hungary, Ireland, Italy, Netherlands, Norway, Poland, Qatar, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | June 07, 2012 | 328 | 6712
Z-3188-2011 | 58478 | No | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Orthovita(R) Cortoss(TM) Delivery Gun |  |  | March 31, 2011 | 2011 | Orthovita, Inc. | Recalling firm has discovered through internal testing that the Cortoss(TM) Delivery Gun may abrade its packaging material under certain transportation and handling conditions. This abrasion may cause small holes or tears in the packaging material, which may not be immediately visible to the user. Although t the Cortoss(TM) Delivery Gun is provided in double pouches and there have been no report | Packaging | Orthovita sent an "URGENT FIELD SAFETY NOTICE" letter dated March 31, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to Identify, quarantine and discontinue use of the affected product. The letter included an acknowledgement form to be completed and returned as per instructions.. The letter states that the recalling firm will contact customers to arrange for the return or destruction of any affected product. A pre-paid shipping label is provided for product return. The letter also states that the recalling firm will provide replacement product at no cost. For any questions regarding this notice call 484-323-8860. |  | 3174 | 3174 | Product was distributed within the US to medical facilities in AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NY, NC, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, and WI. Government accounts include Bethesda Naval Center, 8901 Rockville Pike, Bethesda, MD; Lackland Air Force Base, 2200 Berquist Drive, Suite 1, Lackland AFB, TX; Hunter Holmes McGuire VA Hospital 1201 Broad Rock Blvd., Richmond, VA; Madigan Army Medical Center, Madigan Army Medical Center, Tacoma, WA; Ralph H. Johnson VA Medical Center, 109 Bee Street, Charleston, SC; VA Medical Center, 1481 W. 10th St., Indianapolis, IN; VA Hospital of Palo Alto, 3801 Miranda Ave., Palo Alto, CA; VA Loma Linda Healthcare System, 11201 Benton Street, Loma Linda, CA; and VA Medical Center/NY Harbor Healthcare, NY, NY. | Terminated | September 25, 2012 | 544 | 3174
Z-3178-2011 | 59331 | No | LPB |  |  | Cardiovascular | PMA | STOCKERT 70 Radio Frequency Generator, Catalog Number 57001 |  |  | July 07, 2011 | 2011 | Biosense Webster, Inc. | Biosense Webster has initiated a voluntary field notification to all customers that have a Stockert 70 Radio Frequency Generator, to communicate undetected software anomalies in the form of software release note in order to amend the labeling and thereby enhance the safety of the product. | Software design (manufacturing process) | The recall communication was initiated on 7/8/11 and 7/11/11 with Biosense Webster forwarding A Customer Notification Letter, Acknowledgement Form and Software Release Note (Attachments 3, 4 and 1) via FedEx to all U.S. customers who purchased the Stockert 70 Radio Frequency Generator. Two letters and forms were sent to U.S. customers via express mail, addressed to the EP/Cardiology Lab and Risk Management Office, respectively.The Release Note provides tips for the avoidance of potential events when working with the EP-SH UTTLE, STOCKERT 70, STOCKERT J50 or STOCKERT J70 RF generators with SW versions 1.035 / 1.035J /1.035J70 to 1.037/1.037J70. |  | 3397 | 3,397 units total (1,217 in U.S.) | Worldwide Distribution -- Algeria, Argentina, Australia, Austria, Belarus, Belgium, Brazil, Brunei Darussalam, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Egypt, England, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Korea, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tahiti, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, United States, Uruguay, and Vietnam. | Terminated | October 12, 2012 | 463 | 3397
Z-3200-2011 | 59175 | [6 empty] | Philips GEMINI TF64 Diagnostic Imaging Systems, Model Number: 882471, 510(k) #K052640, Serial Numbers: 7502, 7504, 7506, 7514, 7083 and 7164. The recalled GEMINI TF64 system units equipped with software versions: 3.5.1.1 and 3.5.2.1 are subject to recall/software correction. |  |  | May 31, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The firm is conducting a recall/software correction to several of their GEMINI PET/CT X-ray system units in the field. The firm believes that if these software defects were to recur, it could impact the performance of reliability of the these PET/CT system units. | Software design | The firm , Philips Healthcare, issued a 'Customer Information Letter' dated August 9, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were not required to perform any action. A Philips' Field Service Engineer will visit each customer site and install the necessary software upgrade and Philips will track the progress of the software corrections as they take place.For customers in North America, if you need any further information or support concerning this update, please contact our Customer Care Solutions Center at 1-800-722-9377. Selection option 5 for "All Imaging Systems". Enter your site ID number. If you do not know your site ID, please hold and then select option 5 for "Nuclear Medicine" and finally select option 2 for "PET" support. In all other countries, the local Philips Healthcare office should be contacted at 440-483-3000. |  | 6 | 6 units | Worldwide distribution: USA including states of: FL, MA, NE, NC, OH, TX, WA and WV; and countries of: France, Germany, Italy, Kuwait, The Netherlands, Spain and Turkey. | Terminated | February 22, 2013 | 633 | 19
Z-3191-2011 | 58985 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Sterile Bone Cutting Burrs model QD8-4OVD labeled in part:***QD8-4OVD***QTY: 1***4 mm x 8 mm Diamond Oval. Use with***QD8 Attachments.***Use by 2016-03***ANSPACH***Manufacturer***Made in USA***The Anspach Effort, Inc.***4500 Riverside Drive***Palm Beach Gardens, FL 33410, USA***Tel: (800) 327 6887 / + 1 561 627 1080***EC REP***Anspach Europe Ltd.***Aston Court, Kingsmead Business Park***Frederick Place***High Wycombe, HP 11 1LA***United Kingdom***Tel: +441494 616126***Rx Only***LOT E173049563***. |  |  | May 09, 2011 | 2011 | The Anspach Effort, Inc. | The Anspach Effort, Inc., Palm Beach Gardens, FL, is recalling their Anspach Single Use, Sterile Bone Cutting Burrs, catalog numbers, IMRI-S-8NS-M, M-6P, QD8-4OVD, S-5B-4, and SIL-7-9, lot #'s C323023789, C403026551, C473028567, C403026563, D153034013, D173034385, and C363024789, because the bar code which defines the expiration date may be incorrect. This could cause the bar code reader system to | Employee error | The Anspach Effort, Inc. sent an "URGENT: MEDICAL DEVICE REMOVAL" letter dated May 27, 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory and to segregate any affected product found. A Customer Response Form was attached for customers to complete and return to the firm via fax at 1-800-327-6661.Contact Anspach Customer Support at 800-327-6887 for questions regarding this recall. |  | 17 | 17 | Worldwide Distribution--USA (nationwide) including the states of Alabama, California, and Wisconsin and the countries of Columbia, France, Germany, Great Britain, Italy, Panama, Poland, Saudi Arabia, South Africa, Spain, and Switzerland. | Terminated | February 15, 2012 | 282 | 114
Z-3195-2011 | 59690 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301B: AQUILION Premium; System, X-Ray, Tomography, Computed |  |  | June 22, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on their TSX-301B: AQUILION Premium; System, X-Ray, Tomography, Computed,TSX-301A: AQUILION One; System, X-Ray, Tomography, Computed,TSX-20IA: AQUlLlON LB; System, X-Ray, Tomography, Computed, and TSX-1O1A: AQUILION 16; System, X-Ray, Tomography, Computed due to several software issues that cause the scanners to function incorr | Software design | Toshiba America Medical Systems (TAMS) sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 22, 2011, to all affected customers. Customers were instructed that a revised software to prevent occurrence of the problem would be installed on their system. Their Toshiba Service Representative would contact them for an appointment to install the software to their system when it became available.Customers were requested to take temporary measures until corrective measures were taken.1) In dual scanoscopy, do not press the short key between the completion of the first scanoscopy and the start of the second scanoscopy.2) In dual scanoscopy, do not click the (Scano Skip) button immediately before the completion of the first scanoscopy.3) When the SplitCase function is set to ON, do not select two or more orders simultaneously from the appointment list.4) When an image comment is to be preset by making a selection from the preset memo, do not select any image comment with more than 10 characters. Customers were asked to share the information with all users and reviewing radiologists as well as clinical engineering or Biomedical group at the facility.Customers were asked to complete and return the attached form and fax it to 877-349-3054.For any questions call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 155 | 155 units total | Nationwide Distribution | Terminated | August 11, 2014 | 1146 | 155
Z-3184-2011 | 59797 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Straumann coDiagnostiX, version 8.0, used to guide a clinician in the proper implant selection, oral placement and the proper selection of an oral drill. |  |  | August 04, 2011 | 2011 | Straumann Manufacturing Inc. | Software version 8.0 has a software anomaly, in which the length is not adjusted accordingly when it is changed after the surgical sleeve position planning step has been completed. | Software design | Staumann USA notified accounts by an Urgent Field Safety Notice letter dated August 4, 2011 and delivered via Federal Express Delivery Service.The letter identified the affected product and explained the reason for the recall. Customers were asked to review the letter, take appropriate actions, including completing and faxing the enclosed. Customer Confirmation Form. Straumann released a corrected software upgrade (coDiagnostiX 8.0.2.3672), which is available for download when users go online and activate current software. If the user does not have internet access a replacement CD can be ordered from Technical Support at 866 - 531 -7365. Questions should be directed to Technical Support Monday through Friday from 8:00 am to 8:00 pm Eastern Time. |  | 272 | 272 | Nationwide Distribution | Terminated | May 23, 2012 | 293 | 272
Z-3076-2011 | 59760 | No | RBS |  |  |  |  | Hamamatsu model L10951-35 Micro Focus x-ray generator. |  |  | July 07, 2011 | 2011 | Hamamatsu Corporation | It was discovered about a potential issue of X-ray leakage with our Model # LI1 0951-35 Micro Focus X-Ray generator. | Radiation Control for Health and Safety Act | Hamamatsu Corporation sent a letter dated July 19, 2011, to the manufacturer and their subcontractor informing them of the affected product and what they propose to do. They will repair the units both in the field and those returned to our plant by installing additional lead shielding in the units. The lead shielding will be fixed to the unit by soldering the addition to the existing lead shield (see attachment 2). We expect to complete the repair of all units in the field by October 1, 2011. All new units to be shipped from our plant will be repaired before shipment.For further questions, please call (908) 231-0960 X 2300 |  | 8 | 8 | Worldwide Distribution -- USA ( nationwide ) the countries of Aichi, India and Japan. | Open, Classified |  |  | 8
Z-3206-2011 | 57889 | No | DPZ | Ear oximeter. | Cardiovascular | Anesthesiology | 510(k) | DATEX-OHMEDA TRUSAT PULSE OXIMETER AND ACCESSORIES, MODEL 3500 (GE Healthcare) |  |  | April 08, 2011 | 2011 | GE Healthcare, LLC | The plastic enclosure on the external power supply of the TruSat Pulse Oximeter is prone to breakage, and overheating. The issue may impact patient or user safety. Potential for system or components to start on fire or explode (general case). | Device Design | On 4/8/11 GE Healthcare sent consignees an "Urgent Medical Device Correction" letter dated March 30, 2011. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, Director of Biomedical Engineering &amp; Home Healthcare Provider.The letter listed the Safety Issues, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 31846 | 31,846 | Worldwide distribution: USA, ZAMBIA, VIETNAM, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, UGANDA, TURKEY, TUNISIA, TRINIDAD &amp; Tobago, THAILAND, UNITED REPUBLIC OF TANZANIA, PROVINCE OF CHINA TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SUDAN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, QATAR, PORTUGAL, POLAND, PHILIPPINES, PERU, PANAMA, OMAN, NORWAY, NICARAGUA, NEW ZEALAND, NETHERLANDS, NEPAL, MOZAMBIQUE, MEXICO, MALTA, MALAYSIA, MACEDONIA, LITHUANIA, LIBYAN ARAB JAMAHIRIYA, LEBANON, LATVIA, KUWAIT, REPUBLIC OF KOREA, KENYA, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HUNGARY, HONG KONG, HONDURAS, GUATEMALA, GUADELOUPE, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, KOREA, EL SALVADOR, EGYPT, ECUADOR, DENMARK, CZECH REPUBLIC, CYPRUS, CROATIA, COSTA RICA, COLOMBIA, CHINA, CHILE , CANADA, BRITISH INDIAN OCEAN TERRITORY, BRAZIL, BOLIVIA, BENNIN, BELGIUM, BANGLADESH, BAHAMAS, AUSTRIA, AUSTRALIA, ARGENTINA, and ALGERIA. | Terminated | March 04, 2016 | 1792 | 31846
Z-3213-2011 | 58820 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Crystal Metal Miller Disposable Blades (sizes 0 art no: 50676 and 00 part no 50647 |  |  | March 10, 2011 | 2011 | Penlon, Ltd. (Fmly East Healthcare) | Disposable laryngoscope blades incorrectly labeled with wrong size. | Other | Penlon Ltd ( frmly East Healthcare ) sent an " Urgent Field Safety Notice" dated March 7, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed to check their stock with expiry dates between August 2010 to July 2015, and ensure that the sizes written on the blades, inner plastic bags, and the carton all match.Please complete the reply form below and send to the address provided. If you have any stock where the sizes on the bags, blades, and carton do not match, quarantine the blades and contact Penlon using the information below.For further questions please call 44 1235 547093. |  | 5 | 5 cartons (20 blades each) | Worldwide Distribution -- USA (nationwide) including the states of CA, VA and WI., and the countries of Belgium, Ireland, France, Netherlands, Norway, Portugal, Switzerland, UK, Hong Kong, Greece, Bahrain and China. | Terminated | September 21, 2012 | 561 | 5
Z-3211-2011 | 59438 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Victory 10, 3 Wheel (SC610) |  |  | June 22, 2011 | 2011 | Pride Mobility Products Corp | Contract manufacturer error. Component part(s) (fuses) specifications were not adhered to as outlined in Pride's Engineering drawings. | Nonconforming Material/Component | Pride Mobility Products Corp. sent a "RECALL OF VICTORY 3 AND 4 WHEEL UNITS; REPLACEMENT OF FUSES" letter dated June 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers are asked in the letter to locate their affected Victory unit(s) and inform users of the recall. A letter has been provided by the firm to distribute to users alerting them to the recall. Additionally, customers are to set up an appointment with users to replace the fuses. Customers are instructed to contact Pride's Upgrade department at 866-222-6212 to order the kit and for questions regarding this recall. . |  |  |  | Worldwide Distribution-USA (nationwide) including the District of Columbia and Puerto Rico and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and countries of Australia, Canada, France, and United Kingdom. | Terminated | November 20, 2014 | 1247
Z-3202-2011 | 59182 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Cutting Burr labeled in part:"***S-1504TD QTY:1***Twist Drill for 1.5 mm x 4 mm Screw. Use with SHORT Attachments.***ANSPACH***Manufacturer Made in USA***The Anspach Effort, Inc***4500 Riverside Drive***Palm Beach Gardens, FL 33410, USA***Tel (800) 327 6887 / +1 561 627 1080***Anspach Europe Ltd.***Aston Court, Kingsmead Business Park***Frederick Place***High Wycombe HP11 1LA***United Kingdom***Tel: +441494616126***Sterilized using Irradiation***Do Not Re-use***Keep Dry***Caution; consult accompanying documents***Sterile unless damaged or opened***Rx Only***". |  |  | March 17, 2011 | 2011 | The Anspach Effort, Inc. | The Anspach Effort, Palm Beach Gardens, FL, is recalling their Single Use, Sterile Bone Cutting Burrs, Catalog #, S-1504TD, Lot #s D363039140. D323038110, D293037349 and Catalog # TURQ-434-4, Lot #s D343038845, D323038131, due to the burr being 2.0 mm longer than stated on the label. As a result, a hole drilled using this burr would be 2 mm deeper than the screw. | Mixed-up of materials/components | Anspach Effort, Palm Beach Gardens, FL sent a Customer letter dated March 31, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were instructed to screen their inventory, contact Anspach Customer Support at (800) 327-6887 if any of the recalled product was found, and complete and return the attached reply form. If the product had been further distributed, they were asked that the recall information be forwarded as appropriate. |  | 599 | 599 | Worldwide Distribution -- USA (nationwide) including the states of CA, CO, IN, KY, MD and TX. and the countries of Germany, Japan, Taiwan, United Kingdom and Sweden. | Terminated | October 11, 2011 | 208 | 628
Z-3210-2011 | 59826 | No | IRP | Powered inflatable tube massager. | Physical Medicine | Physical Medicine | 510(k) | Game Ready Professional Therapy System, |  |  | August 23, 2011 | 2011 | CoolSystems, Inc. | It was discovered that some Game Ready Professional Therapy System units that were to have been quarantined pending final inspection and testing were inadvertently moved to finished goods inventory and shipped. | Employee error | The firm, CoolSystems, Inc., contacted customers by phone on August 23, 2011. The customers were informed that some Game Ready Professional Therapy System units have been quarantined pending final inspection and testing were inadvertently moved to finished goods inventory and shipped. Instructions were given to immediately cease fulfilling orders from that inventory. Loaners were shipped to the customers overnight. The customers were instructed to replace the shipped device with the loaner that was sent and return it to the company ASAP. If you have any questions, contact the Vice President, Quality Assurance and Regulatory/Clinical Affairs at 510-868-5378. |  | 26 | 26 units | Nationwide distribution: CA, MA, MO, NH, OH, TX, UT, and WA. | Terminated | November 14, 2011 | 83 | 26
Z-3205-2011 | 59473 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | PCR Eleva Radiological Image Processing System with software R.1.2 |  |  | July 27, 2011 | 2011 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be interpreted as a RIGHT marker for side indication if the image is mirrored again within a PACS system. This could potentially lead to a misdiagnosis | Software design | Philips Healthcare Inc. sent out "Urgent --Medical Device Correction" notices dated June 10, 2011 to all affected customers.The letter described the product and problem. Customers were informed the "burned-in Mirror-Icon" may be misinterpreted as a Right marker. Customers are advised to preferably use lead markers during each X-ray acquisition or alternatively electronic annotations on each image for side indication. The presentation of these markers within each image will clearly show the difference between the intended Left/Right indication and the little mirror ICON. Philips will install improved software that prevents the misinterpretation of the symbol and will schedule the installation of the updated software.For questions on this recall call Phillips Healthcare Inc. at 978-687-1501. |  | 197 | 197 units | Worldwide Distribution: Nationwide distribution; and including the countries of Albania, Angola, Argentina, Australia, Austria, Azerbaijan, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Croatia, Denmark, Egypt, France, Germany, Ghana, Greece, Greece, Haiti, Hungary, Indonesia, Iraq, Israel, Italy, Kazakhstan, Kenya, Latvia, Libya, Liechtenstein, Lithuania, Malawi, Malta, Martinique, Mexico, Mozambique, Netherlands, New Zealand, Nigeria, Norway, Oman, Philippines, Portugal, Puerto Rico, Romania, Russia, Saudi-Arabia, Serbia, Slovakia, South Africa, Spain, Sudan, Sweden, Switzerland, Tanzania, Thailand, Tschech. Rep, Turkey, Turkmenistan, UAE, Uganda, and United Kingdom. | Terminated | January 12, 2017 | 1996 | 197
Z-3208-2011 | 59681 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-Ray System, 510 (k) #K060937, Model #728306, Serial Numbers: 100019, 100023, 100087, 100103, and 100170. The recalled system units are equipped with operating software version; 3.2.0.19011. |  |  | June 15, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. Philips Medical's Field Service Engineers are installing software upgrades to certain of their Brilliance iCT and Brilliance ICT SP CT X-ray system units in order to correct non-conformances evidenced in these systems due to software errors. | Other | Philips Medical ( Cleveland) Inc. indicated in the Correction and Removal Report dated June 21, 2011, that a Philips Field Service Engineer will be contacting each site in order to schedule a time to implement the Field Change Order. After which time, the Field Service Engineer will visit each customer site to install the software upgrade and the firm will be tracking the progress of this corrective action.For further questions please call (440) 483-7000. |  | 5 | 5 units | Worldwide Distribution -- USA (nationwide) including the states of DC, IL, IN, OH, PA and VT and the countries of Belgium and Israel. | Terminated | February 22, 2013 | 618 | 8
Z-3204-2011 | 58474 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Inc., Millennium and HD-120 MultiLeaf Collimators (MLC), a component of the Clinac, All Varian Clinac using the collimator cover P/N 1106900-xx, Model Numbers: |  |  | March 17, 2011 | 2011 | Varian Medical Systems, Inc. | An improperly installed collimator cover may detach and fall from the machine, potentially striking the patient or nearby personnel. | Component design/selection | On 3/17/2011, Varian Medical Systems sent an Urgent Medical Device Correction Letter to all affected users, with a description of the problem and user corrective action steps. |  | 3791 | 3791 units | Worldwide distribution including USA, Australia, Argentina, Italy, Morocco, South Korea, Belgium, Canada, Netherlands, Denmark, Japan, Hong Kong, Malaysia, Hungary, New Zealand, Sweden, Taiwan, Switzerland, Romania, Brazil, Slovakia, India, Colombia, UK, Ireland, Spain, Singapore, Martinique, Reunion, Venezuela, Iceland, Portugal, Turkey, Israel, Bahrain, Macau, Chile, Saudi Arabia, Latvia, China, Finland, Kazakhstan, Thailand, France, Poland, and Germany. | Terminated | June 20, 2012 | 461 | 3791
Z-3194-2011 | 59586 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console v 8.6 with ARIA RT Chart v8.5, all of 4DITC v 8.3; |  |  | July 21, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The 4D Integrated Treatment Console may erratically fail to load the planned physical wedge accessory without warning or notification when receiving an approved plan from the oncology information system. | Pending | Urgent Medical Device Correction/Field Safety Notification letters were sent to all US consignees by Certified Mail on July 21, 2011., as well as Varian Sales, Marketing and Service Organizations. Varian International Customer Support managers will notify international customers by appropriate methods. The letter informed customers of the issue and the corrective action steps. Letters request that all appropriate personnel working in the radiotherapy department be notified of the contents of the letter. A new release version is being created to correct the issue. All sites will be notified when it is available.Questions or concerns should be directed to 1-888-827-4265. Contact information for international customers is also provided as well as a list of all affected serial numbers. |  | 1125 | 1125 | Worldwide Distribution -- USA, China, Austria, Belgium, Bulgaria, Czech Republic, Denmark, Estonia, Finland, France, Germany, Guadeloupe, Ireland, Italy, Latvia, Lithuania, Martinique, Netherlands, Poland, Portugal, Slovakia, Slovenia, Spain, Sweden, Turkey, United Kingdom, Singapore, Algeria, Bangladesh, Belarus, Egypt, Guam, Hong Kong, India, Indonesia, Iraq, Israel, Kenya, Kuwait, Macau, Malaysia, Philippines, Qatar, Russia, South Africa, South Korea; Taiwan, Thailand, Tunisia, Ukraine, United Arab Emirates, Japan, Brazil, Chile, Columbia, Dominican Republic, Ecuador, Mexico, Peru, Venezuela, New Zealand, and Australia. | Terminated | August 01, 2013 | 742 | 1125
Z-3214-2011 | 59528 | Yes | DWB | Roller-type cardiopulmonary bypass blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns 8000 Cardioplegia Monitor, Catalog Number 16414; also known as PROG COMP BD AMPRO 420 CPG 8K, Catalog Number 816373 |  |  | August 29, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The Cardioplegia Monitor for the Sarns Modular Perfusion System 8000 may fail to track cardioplegia delivery volume. | Component change control | Terumo Cardiovascular Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated August 29, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. The letter advises customers to make sure all users are aware of the notice. Customers are instructed to complete and return an attached Customer Response Form. Terumo will replace the software on all monitors of the affected product.Customers with questions or concerns should call Terumo CVS Customer Service 1-800-521-2818. |  | 85 | 85 | Worldwide Distribution--USA (nationwide) including the states of AL, AR, CA, CO, GA, IL, IN, MS, NC, OH, TN, VA, and WI and countries of Belgium, Columbia, Hong Kong, India, Singapore, and Taiwan. | Terminated | March 08, 2013 | 557 | 85
Z-3207-2011 | 59290 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Synthes T1 Matrix Locking Cap of the Synthes Matrix System, Part Number 04.632.000. |  |  | March 11, 2011 | 2011 | Synthes USA (HQ), Inc. | Labeling Correction -- The locking cap was not able to be loosened, resulting in the breakage of the screwdriver or the rod to be cut and the assembly to be removed (screw, cap, and rod). This poses a potential risk of significant prolongation of OR time. | Labeling Change Control | The firm initiated this action on March 11, 2011 by issuing an Urgent: Medical Device Labeling Correction letter to all consignees. The letter identified the affected product and stated the reason for the correction. It also instructed the consignee to review and attach the supplementary technique page to any MATRIX Spine System Technique Guides that they may have had in their facility. Customers are to complete the Verification Section of the letter and return it to the firm via mail, fax, or e-mail. If customers have questions, they can call 610-719-5063 or their Synthes Spine Sales Consultant. |  |  |  | Nationwide Distribution -- AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, and WI. | Terminated | December 18, 2014 | 1378
Z-3221-2011 | 59519 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-101: Aquilion 32/64; System, X-Ray, Tomography, Computed. |  |  | December 08, 2010 | 2010 | Toshiba American Medical Systems Inc | There is the potential that the current may not be controlled correctly in examinations using the ECG modulation function, which adjusts the x-ray output in synchronization with the ECG waveform. | Software design | Toshiba America Medical Systems (TAMS) issued an Urgent: Medical Device Correction notification to consignees via USPS return receipt mail. The letter identified the affected product, the problem, and corrective action. The firm plans to install revised firmware to prevent the occurrence of the problem. A Toshiba Service Representative will contact customers to schedule an appointment for installation. The firm asked customers to stop using the Power Reduction option during ECG modulated scanning until the new firmware has been installed. Customers are to share the information in the letter with all users, reviewing radiologist, and the clinical engineering or Biomedical group at their facility. Customers are also to complete and return the attached reply form via fax or e-mail to TAMS as proof of notification. Questions pertaining to the letter should be directed to 800-421-1968 or customers' local Toshiba Representative at 800-521-1968. |  | 528 | 528 units | Nationwide Distribution. | Terminated | March 18, 2013 | 831 | 528
Z-3217-2011 | 59359 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BD MAX DNA Extraction Kit DNA-2 (Whole Blood) Catalog Number 437502. Kit contains 24/2D barcode sample preparation reagent tubes *** For use with the BD MAX System ***to extract DNA from Gram-negative bacteria and viruses, which may be present in clinical specimens. The specimen is mixed with BD MAX Sample Preparation Reagent and processed using the BD MAX System. |  |  | May 24, 2011 | 2011 | Becton Dickinson &amp; Co. | Sample results maybe switched between time it was scanned with handheld barcode reader and time it was placed in the sample rack, associating sample results to wrong patient. | Nonconforming Material/Component | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated May 2011, to all affected customers. The letter identified the products, the problem, and the action to be taken by the customers. The letter advised customers to discontinue use of the Sample Preparation Reagent tubes in the referenced kits and discard any remaining products. Customers were requested to complete an attached reply form to acknowledge their receipt of the notification and fax it to (410) 316-4258. For further assistance regarding replacements contact BD Customer Service Dept. 1-800-675-0908. For all other inquiries, contact BD Technical Services Dept. 1-800-638-8663. |  | 9 | 9 kits | Worldwide Distribution-USA (nationwide) including the states of CA, KY, MD, MI, NC, NY, TX, and WI and the country Belgium. | Terminated | October 12, 2011 | 141 | 392
Z-3254-2011 | 59770 | No | HFX | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Circumcision trays and packs listed as follows: |  |  | May 05, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | The firm distributed surgical kits which contained circumcision clamps that were subsequently recalled by Surgical Design. | Other | DeRoyal Industries, Inc. sent a recall notification letter dated 05/05/2011. A second notice was sent on 07/26/2011 to affected customers who did not respond to the initial notice. Questions regarding this recall may contact the Corporate Attorney at (865) 362-1037. |  | 390 | 390 units | USA (nationwide) including the states of AR, AZ, FL, GA, LA, MI, NJ, NM, and OH. | Terminated | January 30, 2013 | 636 | 490
Z-3255-2011 | 59683 | No | LPB |  |  | Cardiovascular | PMA | 7 F Therapy Ablation Catheter, |  |  | February 10, 2011 | 2011 | St Jude Medical | The recall was initiated because St. Jude Medical has confirmed that the Therapy Ablation Catheter, 4 mm Tip Thermistor Quadripolar, were packaged with the incorrect Instructions for Use (IFU). To ensure accuracy of device labeling, St. Jude Medical is requesting that the incorrect IFU be replaced with the enclose, correct version. | Process control | St. Jude Medical sent an "important Medical Device Safety Information" letter dated February 22, 2011 to all affected customers.The letter includes: affected problem, reconciliation form and corrective action. Customers are asked to remove the incorrect Instructions for Use (IFU) and use the IFU provided in the letter, and to fill out and return the reconciliation form.For questions on this recall please contact the firm at (512) 329-1054 or (651) 756-2869. |  | 179 | 179 units total | Nationwide Distribution including the states of TN, MS, IN, MO, FL, WI, CA, and MT. | Terminated | April 03, 2012 | 418 | 179
Z-3215-2011 | 59743 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific 25ga (0.52mm) Expect Endoscopic Ultrasound Aspiration Needle REF 5002 Use By 1999-12 LOT 13615651C1 STERILE EO |  |  | August 11, 2011 | 2011 | Boston Scientific Corporation | The firm has determined that a single lot of 25 ga Expect EUS/FNA devices were mislabeled. The carton label incorrectly described the devices as 22 ga. The inner pouch label incorrectly describes the device as 25 ga. | Other | Boston Scientific Corporation sent a "URGENT MEDICAL DEVICE RECALL letter dated August 11, 2010, to all affected customers. The letter identified the product, the problem and the action needed to be taken by customer.Customers were instructed to immediately remove the affected lot of Recall products from your inventory ( whether in Labs, Central Supply, Shipping and Receiving or any other location). Also, segregate this product in a secure location for return to Boston Scientific. Complete and return the Reply Verification Tracking Form ( even if you do not have any product to return) following the directions on this page and Reply Verification Tracking Form. Fax the Form to (508) 683-5578. Further questions please call (508) 683-4678. |  | 30 | 30 units | Nationwide Distribution including the states of VT, FL, SC, NC, WA, MA, and MN. | Terminated | November 10, 2011 | 91 | 30
Z-3222-2011 | 59704 | No | MJJ | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Cardiovascular Custom Procedure Kits, REV OA |  |  | August 11, 2011 | 2011 | Terumo Cardiovascular Systems Corp | 3-T Medical Systems one way vent valves, included in the Custom Terumo Cardiovascular kits, may separate at the bond that holds the two halves | Nonconforming Material/Component | Terumo Cardiovascular Systems initiated contact on August 12, 2011 by phone, fax, email and Urgent: Medical Device Recall mail to all affected customers.The firm advised customers of the affected product, problem and required recall actions. Customers can remove whole customized packs that are not in urgent medical need and return packs. Or the User may opt to use a sterile LH 130 vent valve from Terumo Cardiovascular Systems. A response form was included to return back to the firm.For questions on this recall contact Terumo Customer Service at (734) 525-9580. |  | 34 | 34 units | Nationwide Distribution, including the states of FL and MI. | Terminated | March 20, 2012 | 222 | 4233
Z-3259-2011 | 59794 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS Angiographic Catheter, Soft-Vu Cobra (2) BRAIDED, 5 F x 90 cm x .035 in., Catalog No./REF 10709105, STERILE, MANUFACTURED IN USA, AngioDynamics, Inc. 603 Queensbury Avenue, Queensbury, NY 12804 |  |  | June 07, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The lot of Soft-Vu Catheters was labeled 90 cm in length; however the product was cut to a 65 cm length. | Employee error | The firm ANGIODYNAMICS, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 7, 2011 to the consignees via Certified Mail, Return Receipt Requested. The letter described the product, problem and action to be taken. The customers were instructed not to use the product and to return it to AngioDynamics, Inc. in Queensbury, NY. The letter also instructed the consignees to (1.) Identify and segregate the recalled lot that is in your possession, (2.) Complete the enclosed Soft-Vu Angiographic Catheter Recall Reply Form and fax it to the attention of the Soft-Vu Angiographic Catheter Recall Coordinator at 518-798-1360. It is important that even if you do not have any product remaining in your possession that you fill out the attached form noting zero quantity to be returned and fax the form to ANGIODYNAMICS. (3.) Ship the recalled product to ANGIODYNAMICS. Replacement product will be shipped upon receipt and confirmation of the returned product. Reference Return Authorization Number XXXXXXXXXX on the outside of the shipping box and include a copy of the Soft-Vu Angiographic Catheter Recall Reply Form with your shipment. If you have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please do not hesitate to contact your local sales representative or call the Customer Service Manager, at ANGIODYNAMICS Customer Service at 1-800-772-6446. You may also e-mail your questions to customerservice@angiodynamics.com. |  | 30 | Domestic - 25 units; Foreign - 5 units | Worldwide distribution: USA (nationwide) including states of: California, Tennessee, and Wisconsin, and country of: Spain. | Terminated | October 18, 2012 | 499 | 30
Z-3266-2011 | 59820 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS UNI*FUSE INFUSION SYSTEM WITH COOPER WIRE, Catalog No./REF 12401818, CATHETER SIZE 4F, CATHETER LENGTH 90 cm, INFUSION LENGTH 10 cm, OCCLUSION WIRE SIZE .035 in., STERILE, AngioDynamics Inc., 603 Queensbury Avenue, Queensbury, NY 12804 |  |  | January 28, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The particular lots of UniFuse Catheters were packaged with the wrong occluding ball guidewires. | Employee error | URGENT MEDICAL DEVICE RECALL LETTERS (dated 1/28/11) and Recall Reply Forms were sent to the consignees via Certified Mail, Return Receipt Requested. The letters identified the scope of the recall, the reason for the recall, the status, and action to be taken. Customers were to identify and segregate the recalled lot in their possession. They were also to complete and fax back the Reply Form. The recalled product should be returned along with a copy of the Reply Form. Replacement product would be shipped upon receipt and confirmation of the returned product. Questions or concerns should be directed towards customers' local sales representative or call Julie Blair, Customer Service Manager at 1-800-772-6446. |  | 23 | 23 units | Worldwide Distribution -- USA and The Netherlands. | Terminated | October 18, 2012 | 629 | 47
Z-3263-2011 | 59726 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | COULTER DxH Cell Lyse, Part number 628018(1L) and 628019 (5L). |  |  | July 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter identified unexpected high particle counts in the lots of DxH Cell Lyse listed. These lots may generate high background counts;could result in a falsely elevated Total Nucleated Count (TNC) of up to 50 cells/ÂżL; or may also result in recovery of TNC values that were higher than expected for Level 1 of the Body Fluid Control. | Pending | The recall communication was initiated on 08/01/2011 with Beckman Coulter forwarding a Product Corrective Action (PCA) letter with attached PCA Response Form (via US Postal Service for US customers) to all customers, who purchased Coulter DxH Cell Lyse. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to contact Beckman Coulter for product replacement. In addition, customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with questions regarding this Product Corrective Action were instructed to call Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their Beckman Coulter representative. |  | 4913 | 4913 units total (2245 in US) | Worldwide Distribution -- Australia, Belgium, Brunei Darussalam, Canada, France, Germany, Hong Kong, Hungary, India, Italy, Japan, Macao, Malaysia, Netherlands, Norway, Philippines, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and the United States. | Terminated | June 04, 2012 | 329 | 4913
Z-3262-2011 | 59831 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS PROFILER BALLOON DILATATION CATHETER, Catalog No./REF 16801105, BALLOON DIAMETER 5MM, BALLOON LENGTH 4 CM, CATHETER LENGTH 40 CM, STERILE, MANUFACTURED IN USA AngioDynamics 603 Queensbury Avenue, Queensbury, NY 12804 |  |  | February 25, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The particular lots of Profiler PTA Balloon Catheters were labeled incorrectly. Lot 532109 contains product that is actually 5 mm x 4 cm x 40 cm and Lot 532794 contains product that is actually 4 mm x 2 cm x 40 cm. | Employee error | Angiodynamics, Inc sent a "Urgent Medical Device Recall letter" dated February 25, 2011 to all affected customers via Certified Mail Return Receipt Requested. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed the following steps1. Identify and segregate the recalled lot that is in your possession.2. Complete the enclosed Profiler PTA Balloon Catheter Recall Reply Form and fax it to the attention of the Profiler PTA Balloon Catheter Recall Coordinator at 518-798-1360.The form lists the catalog numbers, lot numbers, and quantity our records indicate your facility has received.For further questions please call 1-800-772-6446. |  | 14 | 14 units | Nationwide Distribution including the states of CA, GA, ID, MI, OH and PA. | Terminated | October 18, 2012 | 601 | 29
Z-3268-2011 | 59467 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Contour, Contour TS Blood Glucose Test Strips and Bonus Packs |  |  | August 12, 2011 | 2011 | Bayer Healthcare, LLC | All 10- or 25- count glucose test strip vials, either loose vials or found in patient starter kits of Contour, Contour TS and Contour" USB blood glucose meters are being recalled a packaging issue which, under certain conditions, can cause some test strips packaged in some small-count vials to operate outside of US FDA-cleared performance specifications and produce a low-bias reading during blood | Packaging | Bayer sent an "URGENT DIABETES TEST STRIP RECALL" letter dated August 12, 2011, to its customers with the name of the test strips, sizes recalled, reason for the recall, potential injury, and instructions to discontinue use and return the test strips to Bayer. The recall letter also included a recall response form that was to be faxed back to Bayer at 1-800-876-2243. For questions call Customer Logistics at 1-800-348-2637. |  | 197342756237020 | 3,617,287 vials/ 54,555,460 test strips | Nationwide Distribution and Puerto Rico | Terminated | July 02, 2012 | 325 | 197342756237020
Z-3265-2011 | 59303 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Powerheart Automated External Defibrillator (AED), G3 is indicated for emergency treatment of victims exhibiting symptoms of sudden cardiac arrest who are unresponsive and not breathing. Post-resuscitation, if the victim is breathing, the AED should be left attached to allow for acquisition and detection of the ECG rhythm. If a shockable ventricular tachyarrhythmia recurs, the device will charge automatically and be able to deliver therapy. |  |  | June 30, 2011 | 2011 | Cardiac Science Corporation | Cardiac Science Corporation (CSC) has identified two errors, 0x48 and 0x99, that have repeatedly occurred during a device self-test in the G3 installed base over the last 6 years. These self test errors cause the device to enter into a service required state.CSC remind owners/operators of the Powerheart AED devices to verify the "Rescue Readiness" of their equipment and follow maintenance pro | Device Design | Cardiac Science began mailing out the recall notification letter on August 1, 2011, to all their US consignees via USPS first class mail. Cardiac Science's distribution partners will receive a notification letter and be provided with a template letter to send to their customers directly.CSC remind owners/operators of the Powerheart AED devices to check the status of of the AED every day to ensure it is Rescue Ready, and perform the regularly schedule maintenance as outlined in the Powerheart AED "Operator and Service Manual" provided with every AED.Customers can call Cardiac Science at 800-426-0337 or e-mail at aed205@cardiacscience.com. |  | 195739 | 195,739 units in US | Worldwide Distribution - US (nationwide)This action applies only to domestic | Terminated | February 13, 2014 | 959 | 195739
Z-3271-2011 | 59865 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VASOVIEW 7xB Endoscopic Vessel Harvesting System; |  |  | December 17, 2009 | 2009 | Maquet Cardiovascular, LLC | During routine production quality monitoring process, it was found that UV curing adhesive was not used in a limited number of devices. | Nonconforming Material/Component | MAQUET Cardiovascular, LLC, sent an "Urgent Device Removal/Immediate Action Required" letter dated December 17, 2009, to all affected customers. The letter identified the recalled products, the problem, and the actions needed to be taken. The letters asked customers to examine their stocks (inventory) immediately to determine if they had any affected lot numbers and to discontinue dispensing (distributing) the lot. The letter states that customers are to complete the attached Field Action Response form regardless of whether or not the affected products are located in their inventory. Distributors are instructed to provide the same instructions to their customers supplied with the recalled products. The letter instructed all users to return all unused affected products to MAQUET. For customer's convenience, all affected products that are returned may be exchanged with other unaffected VASOVIEW Endoscopic Vessel Harvesting System by overnight delivery. Questions should be directed to MAQUET Cardiovascular sales representative or Customer Service at 1-888-880-2874. |  | 177 | 177 units | Nationwide Distribution (USA) including the states of: AL, AZ, CA, CT, FL, KS, KY, IL, IN, MI, MO, ND, NY, NV, OH, OK, OR, PA, SC, TN, TX, UT, WA, and WI. | Terminated | September 21, 2011 | 643 | 390
Z-3269-2011 | 59664 | No | ODX | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CombiSetÂż True Flow" 2008K@home Bloodlines |  |  | August 11, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Blood lines coiling tabs #1 and #5 on the arterial line are not in the correct positions | Device Design | Fresenius Medical Care sent out an "Urgent" recall notice dated August 11, 2011 to all affected customers .The letter included description of the product name, problem and asked them to examine their stock immediately and quarantine bloodlines from the recalled lots. Customers were asked to return the product to Fresenius.For information on this recall call the Fresenius Customer Service Team at 1-800-323-5188. |  | 118 | 118 cases | Nationwide Distribution. | Completed |  |  | 118
Z-3273-2011 | 57865 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Synthes Bone Marrow Aspiration System, Sterile, 11 gauge, 15cm needle (without side holes) |  |  | January 26, 2011 | 2011 | Synthes USA (HQ), Inc. | Firm became aware that some systems were released to stock without going through the sterilization process. | Nonconforming Material/Component | Synthes (USA) sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 4, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. Affected product removed from inventory will be replaced by Synthes. Contact the Synthes Spine Sale Consultant at 1-800-620-7025 ext.6883 for questions concerning this recall. |  | 96 | 96 parts | Nationwide Distribution-including the states of AR, CO, FL, GA, ID, LA, MA, MD, MO, MT, NY, NC, OH, SC, SD, TN, TX, and VA. | Terminated | April 16, 2015 | 1541 | 96
Z-3274-2011 | 59791 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker SmartPump Extension Tubing (Pneumatic tourniquet). |  |  | August 25, 2011 | 2011 | Stryker Instruments Div. of Stryker Corporation | Three complaints have been reported indicating that the product inside the package did not match the label description. The units had quick connectors instead of luer lock connectors as stated in the label description. | Mixed-up of materials/components | Stryker Instruments sent out "Urgent Medical Device Recall" notices dated August 22, 2011 to all affected customers.The letter included the name of the recalled product, lot numbers, and the reason for the recall. Customers were instructed to send back any product to Stryker that is within the affected lots. The letter further stated Stryker will quarantine the product and replace any product that is returned. For information on this recall contact Stryker Instruments at (269) 389-4534. |  | 5920000003 | 5920-000-003 (30 bx) and 5920-000-005 (30 bx) | Worldwide Distribution: Nationwide Distribution including the states of HAWAII, NORTH DAKOTA, NEBRASKA, PENNSYLVANIA, and IOWA, and the country of FRANCE | Terminated | October 12, 2011 | 48 | 5920000003
Z-3277-2011 | 53897 | No | NSX |  |  | General Hospital | Contact ODE | CareSuite" - Critical Care Manager, PACU Manager and Anesthesia Manager Software, transmission &amp; storage, patient data. |  |  | August 12, 2009 | 2009 | Picis Inc. | Electronic health record software applications malfunctioned in a manner which failed to display the discontinued status for medication orders, as entered by the attending physician | Software design | Picis Inc. sent a customer notification letter dated August 11, 2009 to all affected customers. The letter describes the product, problem, and action to be taken. Customers informed of the potential patient safety error were advised to make the appropriate configuration change, within current software versions 7.1.4 or later. Contact your local Picis support representative for questions regarding this notice. |  | 8576 | 8576 UNITS | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Brussels, Belgium, Denmark, Finland, France, Germany, Malaysia, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, and UK. | Terminated | September 27, 2011 | 776 | 8576
Z-3275-2011 | 59908 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OPMIÂż PenteroÂż, OPMIÂż PenteroÂż C, OPMIÂż PenteroÂż 900, |  |  | September 12, 2011 | 2011 | Carl Zeiss Meditec, Inc. | Some third party surgical drapes have optically active protective lenses that could affect the ZEISS surgical microscope position information that is passed on to third party navigation systems. | Pending | The firm, Zeiss, sent an "IMPORTANT SAFETY ALERT FOR ZEISS SURGICAL MICROSCOPES" letter dated September 7, 2011, to all affected customers. The letter described the product, problem and action to be taken. The customers were instructed to use surgical drapes compatible with ZEISS surgical microscopes and always follow the instructions for use of the third party navigation system manufacturer for calibration, calibration verification and operation. In addition, the firm supplied customers with a supplement to their Zeiss surgical microscope user documentation and requested that they read the information and keep a copy with their original instructions also the customers were instructed to complete and return the Acknowledgement of Receipt form within 5 business days via fax to 1-866-809-6044 or USPS to Stericycle, Inc., Attn: Event 2480, 2670 Executive Dr. Suite A, Indianapolis, IN 46241.Should you have any questions about this Safety Alert, please contact your Carl Zeiss Customer Care representative at 1-888-871-7120. |  | 1677 | 1677 consignees | Nationwide distribution: USA only. | Terminated | February 24, 2012 | 165 | 1677
Z-3276-2011 | 59734 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LOGIX Order Entry (OE) Software, all versions, catalog 2M8410; Baxter Healthcare, Deerfield, IL 60015; The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding, creating and authorizing nutritional solutions intended for IV administration. |  |  | August 15, 2011 | 2011 | Baxter Healthcare Corp. | There is the potential for an order entry error during use of the LOGIX Order Entry (OE) software. When adding ingredients to a TPN order, the user can accept the pre-populated "per units" of "Bag" or select other "per units". A dosing error may occur if the user does not select the prescribed "per units" for the added ingredient. | Software design | Baxter Healthcare Corp. sent an "Urgent Device Correction" letters dated August 15, 2011 which included LOGIX OE screen illustrations of a representative neonate and adult patient, via first class mail to all affected customers. The letter identified the product, problem and actions to be taken. The customers were instructed to follow safe compounding practice, which includes verifying that the ingredient per units entered match the TPN prescription prior to preparation. Baxter requests that customers communicate this issue to all personnel responsible for LOGIX order entry and order authorization at their facility. Customers are instructed to retain a copy of the letter at each computer on which LOGIX OE software is installed. The customers were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Medical Information Services at 1-800-422-2751. |  | 95 | 95 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the country of Canada. | Terminated | January 05, 2017 | 1970 | 95
Z-3258-2011 | 59825 | No | OAB | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ANGIODYNAMICS NANOKNIFE IRE GENERATOR, Catalog/REF No. 20300101, ANGIODYNAMICS, 603 Queensbury Ave., Queensbury, NY 12804 |  |  | April 04, 2011 | 2011 | Angiodynamics, Inc. | Some units of Nanoknife have the potential to deliver a shortened pulse width without indicating an error. The affected units contain a "GOMe" circuit board which may cause the generator to drift from the calibration points set at the time of manufacture. | Component design/selection | AngioDynamics, Inc. sent a "NOTIFICATION OF MEDICAL DEVICE CORRECTION" letter dated April 4, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. The letter instructs customers to review the measurement graphs (displayed on generator) after each delivery of energy, and to contact AngioDynamics Customer Service at 1-800-772-6446 if shortened pulses are detected. Contact your local sales representative or call Customer Service for any additional concerns or questions regarding this notice. |  | 21 | 21 units | Worldwide Distribution--USA (nationwide) and the countries of Australia, Germany, Israel, Italy, and Spain, | Terminated | January 29, 2016 | 1761 | 21
Z-3260-2011 | 59877 | No | FKG | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Custom Dialysis prep kits labeled as follows: |  |  | August 15, 2011 | 2011 | Lee Medical International Inc | The firm used alcohol triple swabstick and prep pads manufactured by H &amp; P Industries in the manufacture of their custom dialysis kits. Those swabsticks and pads were subsequently recalled due to possible bacterial contamination. | Nonconforming Material/Component | The firm, Lee Medical International Inc., notified their consignees by email and telephone on August 17, 2011. Consignees were instructed to remove the affected product from stock, quarantine for re-processing, and return or destroy product. Lee Medical arranged for replacement product from alternate manufacturer to be sent to customers. For questions call Production Manager or Operations Manager at 504-734-9336 or email melissal@eleemedical.com or chriss@eleemedical.com. |  | 4075 | 4,075 units | Nationwide Distribution (USA) including the states of: CA, FL, GA, IA, LA, MD, NY and TX. | Terminated | January 30, 2012 | 168 | 4075
Z-3278-2011 | 59862 | No | CJE | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Alkaline Phosphatase, List Number 7D55-21 &amp; 7D55-31. |  |  | August 26, 2011 | 2011 | Abbott Laboratories, Inc | The reagent has the potential to be contaminated with mold, the Cladosporium fungal species. | Pending | The firm decided to recall and notified consignees by a Product Recall - Immediate Action Required letter on 08/26/11. The letter identified the affected product, the reason for the recall, the patient impact, and necessary actions to be taken by the customer. Customers are to determine if they are using or have the affected product in their inventory. Customers are to discontinue use and destroy any remaining product of the specified lots in accordance with their facility's policies and procedures. Alternative lots of the product should be ordered and the enclosed Customer Reply form should be completed and returned. The letter also states that customers may want to decontaminate their systems by following the steps in Attachment A. If the product from the affected lots were further distributed, a copy of the letter should be provided to those customers as well. Questions regarding the letter should be directed to Customer Service at 1-877-4ABBOTT for US customers, and customers outside of the US should contact their local area Customer Service. |  | 6269 | 6269 units | Worldwide Distribution | Terminated | August 17, 2012 | 357 | 6269
Z-3282-2011 | 59527 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | ADVANTAGE WINDOWS (Workstation) WITH FUNCTOOL OPTION |  |  | August 18, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware that image orientation of reference series displayed in Functool Diffusion Tensor/FiberTrak protocol on Advantage Workstation may not match with the orientation annotations displayed in the viewport. This issue may impact patient safety. | Software design | GE Healthcare sent "Urgent Medical Device Correction" letters dated August 18, 2011 to all affected customers.The letter included name of product affected and described problem. Recommended actions where provided to customers and they were notified a GE Healthcare representative will contact them for a software system upgrade.For question of this recall contact the GE Service Representative at (800) 437-1171. |  | 697 | 697 | Worldwide Distribution: Nationwide distribution including the states of AZ, CA, CT, DC, FL, ID, IL, KY, LA, MD, MI, MN, NJ, NY, OR, PA, PR, TN, TX, VA, WA, and WI; and the countries of VIETNAM, MEXICO, UNITED KINGDOM, TURKEY, TUNISIA, THAILAND, TAIWAN, SWITZERLAND, SWEDEN, ST HELENA, SPAIN, SLOVENIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FÂżDÂżRATION, ROMANIA, REUNION, REPUBLIC OF CHINA, QATAR, PORTUGAL POLAND, PARAGUAY, NETHERLANDS, MOROCCO, MALAYSIA, MACEDONIA, LUXEMBOURG, LIBYA, KUWAIT, KOREA, JAPAN, ITALY, ISRAEL, INDONESIA, INDIA, HUNGARY, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, CZECH REPUBLIC, CROATIA, COSTA RICA, COLOMBIA, CHILE, CANADA, BELARUS, BRAZIL, BOSNIA, BELGIUM, BANGLADESH, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | August 09, 2012 | 357 | 697
Z-3280-2011 | 58556 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Dual Blue-Release Adult Walkers - Wheeled, and Blue Release Junior Walkers with 5 inch Fixed Wheels, Model Numbers: 6240-5F and 6240-JR5F. The recalled product is packaged four (4) units per shipping box. |  |  | March 02, 2011 | 2011 | Invacare Corporation | The firm is conducting a recall/correction of their adult and junior wheeled walkers due to a component defect. The firm determined that the caster bolt on the wheeled extension was too short to adequately maintain the Nylex fastener. Consequently, over time, this condition incurs the potential for the caster to loosen and detach, causing the wheels to fall off the walker. This scenario would res | Device Design | The firm, Invacare Corporation notified their customers of this recall/corrective action via an 'URGENT RECALL INFORMATION' letter dated April 7, 2011. The letter was sent to the customers by registered mail or Federal Express. The letter informs the customer(s) of the component defect involved and the inherent possibility that over time, a walker wheel(s) could fall off resulting in instability of the product and potential for serious injury to the user of the device. The letter provides a list of the affected units sold to each facility on an enclosed tracking sheet. It also asks that the customer: 1) Examine their stock for the presence of any of the affected walkers and to quarantine any such products found; 2) Complete an attached recall response form indicating the customers receipt and understanding of the instructions provided in the letter and to fax the completed form to the telephone number listed on the bottom of the form. The letter goes on to state that once the faxed back form is received by Invacare, the company will send the customer replacement wheels kits for all affected walkers purchased; 3) the customer is asked upon receipt of the replacement wheel kits to replace the wheels of any units still in stock and discard all of the old wheel parts removed; 4) Lastly, the customer is asked to contact their sub-account customers and notify them of the recall. If the customer have any questions concerning the instructions, call the recall support line at 1-866-567-7899. |  | 12764 | 12,764 walkers were distributed. | Worldwide Distribution - USA (nationwide: AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, including Puerto Rico and countries of Canada, Columbia, and Israel. | Terminated | July 02, 2013 | 853 | 12764
Z-3281-2011 | 59383 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography System, Model #78231, Serial Number: 9592. |  |  | June 02, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Component defect. Philips Medical decided to conduct a recall field /correction (rework) of the mounting key locking threaded insert of the Detector Module system (DMS) bolt on one (1) Brilliance 64 CT X-ray unit. The firm initiated this action following Philip's receipt of the customer's complaint reporting that the retaining inserts were pulling out of the rotor during the removal of the UDMS | Other | Philips sent a Field Change Order (FCO) 72800519 on June 13, 2011 to the affected customer. A Field Service Engineer visited the customer site and installed a hardware upgrade to the affected Brilliance 64 device system (Serial Number: 9592). The correction action was completed on August 9, 2011.For questions regarding this recall call 1-440-483-7000. |  | 1 | 1 unit | Nationwide Distribution including PA | Terminated | March 15, 2013 | 652 | 1
Z-3279-2011 | 59503 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Trabecular Metal Modular Acetabular System is a modular acetabular cup system consisting of a shell with a preassembled locking ring, a liner, and optional screws. The Shells are available in a variety of sizes and are designed to be used with Trilogy Acetabular System liners. The shell implants are used to replace the hip socket during total hip arthroplasty. |  |  | September 01, 2011 | 2011 | Zimmer Inc. | Lot # 61132698 and Lot 001122467 (added 9/ 18/2012) may have a manufacturing non conformance that could prevent the devices from functioning as intended. Specifically, devices in the affected lot may have a locking ring groove depth that prevents the corresponding liner from snapping into the shell. The groove provides the locking ring room to displace when inserting a liner into the shell. Th | Nonconforming Material/Component | Zimmer, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated July 27, 2011 to all affected customers of lot 61132698 and 9/18/2012 to only affected customer of lot 001122467. The letters describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to stop using and to quarantine the affected product. An Inventory Return Certification Form was included with the letter for customers to complete and return to Zimmer, Inc. via fax at (574) 372-4265. Replacement product will be provided for all returned product. Contact Zimmer, Inc. at 1-800-613-6131 for questions regarding this recall. |  | 62255181 | 24 devices lot 61132689; 1 device lot 001122467 | Worldwide Distribution-USA (nationwide) including the states of CA, NC,NY, PA, WA, and WI and countries of Australia, Germany, The Netherlands, and Switzerland. | Terminated | August 29, 2013 | 728 | 62255181
Z-3283-2011 | 59686 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific, LATITUDEÂż Patient Management System Communicator, Model 6476, Distributed by Boston Scientific, 4100 Hamline Avenue North, St Paul, MN 55112-5798 |  |  | August 01, 2011 | 2011 | Boston Scientific CRM Corp | A limited number of Model 6476, 6482, and 6468 (Not available in the US) LATITUDE Communicators are unable to authenticate, and thus cannot connect to the LATITUDE System. Customers are instead receiving Latitude Temporarily Unavailable (LTU) message. | Other | Boston Scientific CRM Corporation sent a " Dear Patient", letter dated August 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A new communicator is being shipped separately to you with a corrected security certificate that will allow it to be activated and function normally.1. Please activate this new equipment at your convenience to begin home monitoring. Activation instructions are enclosed with the communicator.2. We also ask that you return the original communicator previously delivered to you:" Place the original equipment in the shipping box used for the new equipment." Pack newspaper or other filler around the equipment to prevent damage during shipping, and seal the box." Enter your address in the FROM section of the prepaid Federal Express label, and apply the return label to the outside of the box directly over the original label." Call Federal Express at 1.800.463.3339 to schedule a pick-up.We apologize for this inconvenience. If you have further questions, please call 1.866.484.3268. |  | 470 | 470 | Worldwide Distribution -- (USA) nationwide including the states of AL, AK, AZ, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NE,NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI,SC, SD, TN, TX, UT, VA, WA, WV, WI, WY., and the countries of United Kingdom, France | Terminated | November 06, 2012 | 463 | 470
Z-3284-2011 | 59953 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; the pumps were sold under the following configurations: |  |  | September 16, 2011 | 2011 | Hospira Inc. | Hospira has received reports of incorrect seating of the regulator closer on the Plum A+ Infusion Pumps. If the regulator closer does not close the flow regulator actuator when the pump door is opened, and if the clinician has not engaged the clamp prior to opening the cassette door, unrestricted flow may occur. | Nonconforming Material/Component | Hospira, Inc. sent an "URGENT DEVICE RECALL" letter dated September 16, 2011 to all affected customers. The letter describes the product, problem and actions to be taken by the customers. The letter instructs customers on how to prevent occurrences of continuous recycling/rebooting during start-up. A reply form was attached to the letter for customers to complete and return via fax to 1-888-345-5358. Contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for technical assistance or information regarding this notice. |  | 265105 | 265,105 pumps | Worldwide Distribution-USA (nationwide) including Puerto Rico and the U. S. Virgin Islands and countries of Argentina, Australia, Bahamas, Barbados, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United kingdom, United Arab Emirates, Uruguay and Vietnam. | Open, Classified |  |  | 290217
Z-3288-2011 | 59312 | No | BZK | Monitoring spirometer. | Anesthesiology | Anesthesiology | 510(k) | FloTrak Elite modules |  |  | July 29, 2011 | 2011 | Respironics Novametrix, LLC. | Pressure lines maybe occluded and display incorrect tidal volumes. | Manufacturing material removal | Respironics Novametrix, LLC on July 29, 2011 notified accounts via e-mail or tracked receipt, an Urgent - Medical Device Recall letter to notify customers of the issue and to instruct them on conditions for continued use. Monitors will need to be returned for functional performance testing at their facility. Customers can contact Respironics US Customer Service at 1-877-387-3311 for customers residing in the USA or +011 (203)697-6348 for International customers, if they have any questions or need additional information. |  | 41 | 41 units | Worldwide Distribution -- USA, Belgium, Canada, Chile, Columbia, Egypt, India, Italy, Spain, Saudi Arabia Thailand, GB, and Russia. | Terminated | May 23, 2012 | 299 | 295
Z-3289-2011 | 59724 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Comprehensive Reverse Humeral Fracture Inserter |  |  | August 03, 2011 | 2011 | Biomet, Inc. | Recall initiated in response to a Biomet Canada report, which noted that while testing the instruments, the pin used to secure the implant on the device did not engage on the effected lots and were distributed. | Nonconforming Material/Component | On 8/3/11 Biomet Orthopedics sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter to all affected customers.The letter included: reason for the recall, risks associated with use of the device, and alternate instructions for use of the recalled device. Customers are instructed to locate the affected products, discontinue use, return the devices, and a fax a copy of the response form.For information or questions on this recall call Biomet at (574) 372-1570. |  | 6 | 6 units (internationally) | International Distribution Only: Canada, Australia, and Columbia. | Terminated | October 26, 2011 | 84 | 6
Z-3126-2011 | 58719 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Part Number: 81600 |  |  | April 15, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because certain Access immunoassays can produce inaccurate results when performed within the published UniCel DxI Immunoassay Systems* room temperature operating specification of 18 degrees C to 32 degrees C (64.4 degrees F to 89.6 degrees F). An increase in room temperature causes assay results for some assays to decrease, while results for other assays increase with an i | Other | Beckman Coulter sent an "URGENT FIELD SAFETY NOTICE" letter dated April 18, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter includes information on the new room temperature operating range for all UniCel DxI Immunoassay Systems and Access Immunoassay Systems during system operation. Customers were instructed to review their historical room temperature records to determine if their Access Immunoassay System and UniCel DxI Immunoassay System were operating outside the revised temperature range. An attached Response Form was included for customers to complete and return.Please contact Beckman Coulter Customer Support Center for questions regarding this notice at 1-800-854-3633. |  | 11266 | 11,266 units | Worldwide Distribution--USA (nationwide) and Puerto Rico and the countries of Algeria, Angola, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Lebanon, Liechtenstein, Malaysia, Mayette, Mexico, Morocco, Nambia, Netherlands, New Zealand, Norway, Philippines, Poland,, Romania, Russian Federation, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, United States, Venezuela and Vietnam. | Terminated | December 21, 2012 | 616 | 14648
Z-3163-2011 | 59427 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoMobile, Hi-Art System, H-0000-0003, TomoTherapy 1240 Deming Way, Madison, WI 53717 |  |  | July 07, 2011 | 2011 | Accuray Incorporated | TomoTherapy Inc. is sending this Field Safety Notice to make you aware of an anomaly which may affect the performance of the TomoMobile Hi-Art System.TomoTherapy has discovered that when attempting to open the TomoMobile shielding doors, while the door hinges are in the locked position, the hinges may fail allowing the door to disengage from the shielding. | Device Design | The firm, TomoTherapy, sent an "Urgent Medical Device Correction" letter dated July 7, 2011 to its consignees/customers. The letter described the Issue, Product Affected, Recommended Action (advised consignees to not attempt to open the shielding doors when the hinges are in the locked position. The Hinges should always be in the unlocked position when opening the doors) and Resolution. A TomoTherapy field service engineer or an authorized service representative will contact the customers when a resolution is available.If you have any questions, please contact the TomoTherapy Technical Solutions Center by email at support@tomotherapy.com or phone at: North America: 1-866-368-4807; Belgium: 0800 38783; France: 0805 631 565; Germany: 0800 000 7541; Italy: 800 986 399; Netherlands: 0800 020 1364; Spain: 800 300049; Switzerland: 0800 001927; United Kingdom: 0808 238 6035; China/CNC (Northern): 10 800 712 1701; China (Southern): 10 800 120 1701; Hong Kong: 800 967912; Japan: 0044 22 132374; Singapore: 800 1204 683; South Korea: 0079 81 4800 7204, all other locations: +1 608 824 2900 or +32 2 400 44 44. |  | 4 | 4 | Worldwide distribution: USA including states of: AZ, CA and OK; and country of: United Arab Emirates | Terminated | June 29, 2012 | 358 | 4
Z-0001-2012 | 57858 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Probe, Tactile, Left, ASM, Long, |  |  | January 20, 2011 | 2011 | Medtronic Navigation, Inc | Distal end of tactile probe may break off during use. | Labeling design | Medtronic sent an "Urgent Product Safety Notice" to all affected customers dated Jan 20, 2011.The letter included affected product and updated instructions for use. On-site training was given for the use of the probes by Medtronic Representatives.For information on this recall call Medtronics at (720) 890-3409. |  | 6 | 6 probes | Worldwide Distribution: Nationwide distribution including the states of CO and WI; and the country of South Korea | Terminated | October 03, 2011 | 256 | 6
Z-0007-2012 | 59847 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES Safety Infusion System, Model 621-100ES, 621-200ES, 621-300ES, and 621-400ES |  |  | July 20, 2010 | 2010 | B Braun Medical, Inc | The Outlook ES Dose Guard Care Area can inadvertently be exited when the pump enters the KVO (Keep Vein Open) state. The same condition can occur when a sequence of HOLD-HOLD key strokes are performed on the pump key panel. | Software design | The firm, B. Braun Medical Inc., sent an "URGENT: FIELD CORRECTION" letter dated September 8, 2011 to its customers. The letter described the product, problem and actions to be taken. The firm informed the customers that a B. Braun representative will be contact ing them to schedule a field software upgrade of the affected pumps. The customers were instructed to complete and return the FIELD CORRECTION INFORMATION FORM via mail using the enclosed envelop and contact B. Braun Customer Service for shipping instructions at 1-800-627-7867 if product is to be returned.Should you require any additional information, please do not hesitate to contact B. Braun Medical Inc., Technical Product Support at 1-800-627-7867 (800-627-PUMP). |  | 11525 | 11,525 pumps | Worldwide distribution: USA (nationwide) including states of: AR, CA, FL, GA, IA, IL, KS, KY, LA, MA, MI, MO, NC, NJ, NY, OH, PA, SC, TX, VA, WA, WI, and WV; and country of: Canada. | Terminated | January 07, 2015 | 1632 | 11525
Z-0003-2012 | 59979 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Cochlear Nucleus CI512 Cochlear Implant, REF Z209051, New with Nucleus 5, Sterile EO, Made in Australia, Cochlear Ltd 14 Mars Road, Lane Cove 2066, Australia. |  |  | September 16, 2011 | 2011 | Cochlear Americas Inc. | Cochlear implant may shut down and cease to function. | Component design/selection | Cochlear Americas, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 16, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their inventory and quarantine the affected product. A recall response form was attached to the letter for customers to complete and return. Contact Cochlear Clinic Customer Service at 877-883-3101 for questions concerning this notice. |  | 33645 | 33,645 units (global); 25,516 (implanted); 1,741 remaining in U.S. Market for implantation. | Worldwide Distribution-USA (nationwide) and countries in Europe and Australia. | Terminated | July 10, 2012 | 298 | 33645
Z-0006-2012 | 59139 | No | LYJ |  |  | Neurology | PMA | Model 104 VNS Therapy Demipulse Duo Generator Cyberonics, Inc. Houston, Texas, Model 104 Generators are provided in a sterile single use packages as part of the VNS Therapy System. Product labeling may be found at http://us.cyberonics.com/en/vns-therapy-for-epilepsy/healthcare-professionals/vns-therapy/manuals-page. |  |  | May 06, 2011 | 2011 | Cyberonics, Inc | An investigation was initiated based on a report from the field in which an Intensive Follow-up Indicator (IFI) message was unexpectedly received by a medical professional when using Model 250 version 8.0 software to interrogate a patient's Model 103 Generator. | Software design (manufacturing process) | The firm, Cyberonics Inc., initiated containment actions on March 17, 2011 to prevent additional affected devices from being shipped. The firm initiated contact with implanting facilities in the United States on May 6, 2011 and international facilities and distributors on May 10, 2011 to remove product from the field that has not yet been implanted. Additionally, a "Field Safety Alert" letter was sent to all affected medical professionals. The letter identifies the affected product, the issue, and the action(s) to be taken by the physicians. Consignees are instructed to complete and return an attached Effectiveness Check Form. If assistance is needed physicians may contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com |  | 13 | 13 units | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, ID, IL, KS, OH, AR, MA, MI, MN, MO, NC, NH, NJ, NM, NY, OH, OR, PA, TN, TX, UT, WA, WI and WV. Products were also distributed to WASHINGTON DC and Puerto Rico and the countries of: Austria, Belgium, Finland, Germany, Ireland, Italy, Netherlands, Norway, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | December 04, 2012 | 578 | 105
Z-0008-2012 | 59213 | No | MRW |  |  | Orthopedic | 510(k) | OSTEOMED SPINE Facet Fixation System PrimaLOK FF Threaded Implant Driver Shaft OsteoMed L.P. 3885 Arapaho Road, Addison, Texas 75001 |  |  | June 07, 2011 | 2011 | Osteomed, Lp | Two of the Implant Driver Assembly tips were reported to break during surgery. | Device Design | OsteoMed LP sent an Urgent --Product Recall letter dated June 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately for affected inventory and promptly return any affected inventory to:OsteoMed LPAttan: Rebecca Ellis3885 Arapaho RoadAddison, Texas 75001If the customer further distributed any of the affected product customers were instructed to immediately contact their accounts and advise them of the recall situation and have them return any devices they may have to OsteoMed at the above address.Customers were asked to complete and return the enclosed response immediately in the enclosed, postage paid envelope even if they do not have the device. For any questions please call (800) 456-7779. |  | 20 | 20 units | Nationwide Distribution including PA, TX, CA, IL IA, FL, CT, and MO | Terminated | December 04, 2012 | 546 | 20
Z-0004-2012 | 58823 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, CADDÂż-Solis ambulatory infusion pumps, Models 2100, REF 21-2101-0200-51 and 2110, REF 21-2111-0100-51, Rx Only. |  |  | May 10, 2011 | 2011 | Smiths Medical ASD, Inc. | Reports of non-delivery/ significant under-delivery of fluids when the CADDÂż Medication Cassette Reservoir is not correctly latched to the pump.EXPANDED: On 7/14/11, Smith Medical expanded this recall to include all consignees with affected pumps. | Nonconforming Material/Component | Smiths Medical sent a customer's bulletin notice dated April 11, 2011, only to the customers who reported this issue. Smiths Medical will work directly with the five Consignees who have reported complaints for this issue.The customer was instructed to read the attached Customer Information Bulletin. If you have additional questions on how to correctly latch a cassette to the pump, than please visit the following website: https://www.smiths-medical-education.com.Smiths Medical will also provide replacement pumps, as requested. This also will be considered an extension of Smiths Medical's investigation into this issue and to further validate the effectivity of the corrective and preventative actions in actual use.For further questions please call (651) 628-7280.EXPANDED:Consignees were sent a Smith Medical " Urgent Safety Alert" letter dated June, 30, 2011. The letter was addressed to Risk/ Safety Managers, Clinicians/ Biomedical Engineering, Nursing, Pharmacy, Pain Management and Anesthesia Professionals, Medical Equipment Rental Companies,Distributors and other users of these devices. The letter described the problem and the product. It provided "Advice on Action to be Taken by the User". This included to read the attached Customer information Bulletin and to complete and return the attached Confirmation Form. |  | 4536 | 4536 (3792 USA, 744 OUS) | Worldwide Distribution -- (USA) nationwide including the states of AR, CA, MO, NY and WI., and the countries of BELGIUM, CANADA, SWITZERLAND, GERMANY, DENMARK, FRANCE, UNITED KINGDOM, IRELAND, ITALY, LUXEMBOURG, MALAYSIA, NETHERLANDS, NEW ZEALAND, PORTUGAL, QATAR, SAUDI ARABIA, SINGAPORE. *** EXPANDED*** AL, AK, AZ, CO, DC, FL, GA, ID, IL, IN, KY, LA, MD,MA, JI, MN, MT, NE, NH, NJ, NM, NC, OH, OK, OR, PA, RI, TN, TX, UT, VT, VA, WA, WY. UNITED ARAB EMIRATES, AUSTRIA, AUSTRALIA. | Terminated | July 19, 2012 | 436 | 4536
Z-0011-2012 | 59956 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Hyperbaric Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; list 11005 |  |  | September 16, 2011 | 2011 | Hospira Inc. | Hospira has received reports of continuous recycling and/or rebooting of Plum A+ infusion pumps when the backlight intensity and/or display contrast settings for the LED display has been adjusted from the original default setting. This condition could cause a delay in therapy if the clinician is unable to start the device due to the continuous recycling and/or rebooting. | Software design | The firm, Hospira sent "URGENT DEVICE RECALL" letters dated September 16, 2011 to all affected customers. The letters identified the products, problem and action(s) to be taken. The customers were instructed to read the information provided carefully, complete the attached Reply Form and return it to Hospira via fax at 1-888-345-5358. The letter states that a software upgrade to address the issue is currently being developed. Customers were instructed to notify their consignees if they have further distributed the affected product. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for technical assistance. *** Follow-up letters dated June 26, 2012 were sent to the accounts listing the device modules with software versions v13.2, v13.3, v13.40 and v13.5 that will require the software upgrade. *** |  | 120 | 120 pumps | Worldwide Distribution - USA (nationwide) including Puerto Rico and the US Virgin Islands, and countries of: Argentina, Australia, Bahamas, Barbados, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United kingdom, United Arab Emirates, Uruguay and Vietnam. | Open, Classified |  |  | 290217
Z-0013-2012 | 59952 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Natura Durahesive ConvaTec Moldable Technology Post-Operative /Surgical System (57 mm) |  |  | June 22, 2011 | 2011 | Convatec Inc. | Post-operative kits 57 mm may contain 70 mm pouches. | Labeling mix-ups | The firm, ConvaTec, sent a "MEDICAL DEVICE RECALL" letter dated June 22, 2011 via UPS 2nd day air to all customers and sales representatives that received product. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately stop distributing and quarantine the recalled lot; perform a count of recalled product currently in inventory; complete and return the Recall Response Form via fax to (908) 904-2120 or mail to 200 Headquarters Park Drive, Skillman, New Jersey 08558; return the Recall Response Form even if no recalled product is in inventory; call the toll free number 1-800-582-6514 for return instructions and RGA number-(recall product must be returned in a separate container from any other product returns); mark all return shipments "Recalled Product" and provide supplied RGA number, and forward this letter to customers, if they distributed the product.Should you have any questions, call 1-800-582-6514. |  | 283 | 283 units | Nationwide distribution: AL, AZ, CA, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, NC, NJ, NV, NY, OH, OK, OR, PA, RI, TN, TX and UT. | Terminated | January 11, 2012 | 203 | 283
Z-0012-2012 | 59943 | No | ECI | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Biomet Microfixation outer packaging labeled in part:"REF: SP-2411***RUBBER BAND, LATEX FREE SMALL, HEAVY, 300/PACKAGE 4MM, (3/16" INNER DIAMETER) NON LATEX ELASTICS LOT 307850***DISTRIBUTED BY: BIOMET MICROFIXATION 1520 Tradeport Drive, Jacksonville, FL 32218 USA www.biometmicrofixation.com***." |  |  | August 18, 2011 | 2011 | Biomet Microfixation, Inc. | On 09/18/2011, Biomet Microfixation, Jacksonville, FL, recalled their SP-2411, Rubber Band, Lot 307850 because it was labeled and sold as "LATEX FREE" although it contains "LATEX". The firm's HHE identified this as a potential high health risk to the patients or end users with latex allergies. | Error in labeling | Biomet Microfixation, Jacksonville, FL sent a Urgent Medical Device letter dated August 19, 2011, to all affected customers. They were also notified by phone, and followed-up by e-mails starting on 08/18/2011. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed If you have any of the affected products in your facility, please remove it and return it immediately to Biomet Microfixation for credit. In order to return these products, please utilize the Return Goods Authorization (RGA) number provided on the reconciliation form.Please follow the instructions on the form and fill out the attached Inventory Reconciliation sheet and FAX to (904) 741-9425. If you have any further questions, please call 1-800-874-7711 or (904) -741-4400, extension 9468, Monday through Friday, 8am-4pm ET. |  | 22 | 22 packs | Worldwide Distribution-- USA (nationwide) including the states of FL, KY, MN, NJ, SD, and VA., and the country of Australia. | Terminated | February 06, 2012 | 172 | 22
Z-0016-2012 | 59854 | No | OFD | Catheter introducer. | Cardiovascular | Cardiovascular | Enforcement Discretion | DeRoyal (R) Heat/Moist Exchange, REF 83-000042, Rx Only, NON-STERILE, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA: |  |  | August 23, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | The firm manufactured kits containing a HCH device, manufactured by Vital Signs, that was subsequently recalled because the product may have an occlusion that could prevent proper flow of medical gases or oxygen to the patient. | Nonconforming Material/Component | The firm, DeRoyal, sent a "URGENT: MEDICAL DEVICE RECALL" letter dated September 13, 2011 to its customers. The letter describes product, problem, and actions to be taken. Customers are instructed to identify inventory being recalled and place in quarantine to prevent further use. Constomers were asked to return affected product they have in inventory. DeRoyal will replace this product free of charge. Customers were asked to complete the attached Recall Response Spreadsheet and return as per the instructions. Constomers were asked to return affected product they have in inventory. DeRoyal will replace this product free of charge. Distributors were instructed to perform a sub-recall to their customers.Customers with questions should call 865-362-1037 between 9am and 5pm EST or email recalls@deroyal.com. |  | 200 | 200 units | Nationawide Distribution USA - including the states of: CA, FL, GA, NJ, and WA. | Terminated | January 30, 2013 | 526 | 1140
Z-0020-2012 | 59868 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT II TURBO Hemoglobin A1c Program Reorder Pack, 1600 tests, |  |  | September 12, 2011 | 2011 | Bio-Rad Laboratories, Inc. | A previous recall related to reports of calibration failure and late retention times using Elution Buffer led to an examination of other buffers manufactured during the same time period and a decision was made to field correct these products until a definitive root cause and corrective action have been found and implemented. | Process control | Bio-Rad sent an Urgent Medical Device Correction letter on September 12, 2011, to all affected subsidiaries and customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discard all affected product in accordance with all local, state and federal disposal procedures. If customers distributed the product further they should direct their customers to discard all unused product.Customers were asked to please translate (if necessary) the attached "Customer Medical Device Correction Response Form," customers should add their local contact information and distribute to all affected customers in their region by mail or fax.Compile tracking documentation (e.g. Excel spreadsheet) for all customers who received the Medical Device Correction Letter.Complete the "SUBSIDIARY MEDICAL DEVICE CORRECTION RESPONSE FORM" and fax to Bio-Rad Regulatory Affairs at (510) 741-3954, or e-mail PDF to CSD_RA/Hercules/US/BIO-RAD.Customers would be issued credits for the affected products based on their request. For questions call (510) 741-3954. |  | 82 | 82 | Nationwide and to New Zealand, Canada, France, England, China, India, Mexico, Italy, Brazil, Australia, Singapore, Thailand, SriLanka, hong Kong, Sweden, Germany, Peru, Jamaica, Uruguay, Columbia, Argentina, Chile, Paraguay and Korea. | Terminated | May 11, 2012 | 242 | 287668114
Z-0030-2012 | 59957 | No | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Minutex D-Dimer ; Trinity Biotech plc, |  |  | June 13, 2011 | 2011 | Tcoag Us, Incorporated | There is a performance issue with Minutex D-Dimer Kit which may cause false negative results for patient samples containing levels of greater than 250ng/mL D-Dimer. | Nonconforming Material/Component | The firm,Tcoag US, Incorporated, sent an "Urgent Device Recall" letter dated June 13, 2011 via UPS to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the products subject to recall; complete and return the attached Product Recall Notification Response Form by June 22, 2011 via fax at 973-775-3085; scrap the products consistent with their laboratory policy and if they have further distributed these products, notify their customers to quarantine and scrap the recalled product.For questions or technical assistance, please call the Tcoag hotline at 1-888-291-0415, option 2 (7.30 am to 6.00 pm Eastern Time). |  | 4409 | 4409 Units | Worldwide distribution: USA (nationwide) and country of: Canada. | Terminated | December 11, 2012 | 547 | 4409
Z-0032-2012 | 59853 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | MDC PACS - release R2.3 SP1 . |  |  | September 01, 2011 | 2011 | Philips Healthcare Inc. | Clinical measurements on true size printouts may be inaccurate | Software design | Philips Medical notified all affected customers with a "Field Safety Notice" dated August 2011.The letter provides instructions to the customer for how to avoid the problem. Customers are informed that until the software upgrade is recieved from Philip, do not use the true-size printing feature if the system is configured for more than one DICOM printer. Customers will be instructed that they will receive a software upgrade when it is available. For questions on this recall contact Philips representative at (800) 722-9377. |  | 20 | 20 units | Worldwide Distribution: Nationwide distribution including Puerto Rico; and the countries of Argentina, Brazil, China, France, Greece, India, Israel Mexico, and Turkey. | Terminated | March 13, 2017 | 2020 | 20
Z-0029-2012 | 59787 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | OLYMPUS ViziShot EBUS Aspiration Needle |  |  | August 25, 2011 | 2011 | Olympus America Inc. | Firm has received complaints from users of the ability to wipe material from the needle's stylet and the presence of material in patient cytology samples obtained by the needle following endobronchial ultrasound fine needle aspiration of a lesion. Olympus Medical Systems Corporation (OMSC) in Japan has confirmed that metal particles can occur on the stylet due to friction between the stylet and n | Device Design | OLYMPUS sent an URGENT MEDICAL DEVICE CORRECTION letter dated August 26, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A revised Instruction Manual with the new WARNINGS will be sent to the customer's attention.Customers were instructed to examine their inventory to determine if they have the affected product.To share the information on the new WARNINGS with all of the appropriate staff at their facility and any other organization where the product has been transferred.Complete and return the enclosed questionnaire even if they do not have affected product. Customers were instructed to fax the completed questionnaire to 484-896-7128.For any question call 484-896-5688. |  | 22498 | 22,498 | Nationwide Distribution including the District of Columbia and the following States: AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Terminated | September 28, 2012 | 400 | 22498
Z-0031-2012 | 59746 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista(R) Cuvettes |  |  | August 08, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed a low frequency defect in the molding of the Cuvettes that may result in Cuvette Ring and Cuvette Loader jams. It has been determined that the issue is related to a molding defect with the Cuvette flange or "wing" breaking, resulting in an error condition. Therefore, patient results are not impacted. | Nonconforming Material/Component | Siemens Healthcare Diagnostics sent a Urgent Field Safety Notice letter dated August 2011, to all affected customers. The letter identifed the product the problem and the action needed to be taken by customer.Customers were instructed to discard all cuvette cartridges of the identified lots. Customers were also informed that until a replacement lot is available, if an error is encountered, they may contact the Technical Solutions Center for assistance at (800) 441-9245. For replacement product , please call the Siemens Atlanta Customer Service Center at (800) 241-0420. |  | 1786 | 1,786 | Worldwide Distribution -- USA (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and Puerto Rico., and the countries of Austria, Belgium, France, Germany, Italy, The Netherlands, Portugal, Spain, Switzerland, Australia, Japan, South Korea, and Canada. | Terminated | December 24, 2013 | 869 | 1786
Z-0028-2012 | 59535 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coulter PrepPlus 2, Part Number: 378600, Serial Numbers: 1001008698, 1001008699, and 10018700. |  |  | August 12, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed an issue with the Coulter PrePlus 2. The reagent level sensing is impaired and will give no warning when a reagent vial is low or out of reagent. Due to the size difference in the vertical mounts, the reagent rack is not aligned optimally for the probe to aspirate or level sense the antibody. This misalignment yields impaired reagen | Mixed-up of materials/components | Beckman Coulter initiated an URGENT: PRODUCT CORRECTIVE ACTION (PCA) letter dated August 24, 2010, with attached PCA Response Form to all affected customers who purchased the Coulter PrepPlus 2. The letter provided the customers with an explanation of the problem identified and an action to be taken. The customers were instructed to not use the PrepPlus2 until their Beckman Coulter representative corrects the problem. A review of all results generated from samples prepared from the system should be conducted. Any unexpected dim fluorescence and/or decreased number stained cells should be evaluated for re-analysis, including preparing a new sample. Customers were also instructed to complete and return the enclosed response form within 10 days. Customers with any questions regarding this PCA, were instructed to call Beckman Coulter Customer service at (800) 526-7694. Outside of USA, customers were instructed to contact their local Beckman Coulter Representative. |  | 391 | 391 units total (193 in US) | Worldwide distribution. | Terminated | April 04, 2012 | 601 | 391
Z-0036-2012 | 59937 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System, Synchrony Respiratory Tracking System (Robocouch Patient Positioning System), a Radiation Therapy Device. Model Numbers: 026430, 028186. |  |  | September 12, 2011 | 2011 | Accuray Incorporated | A report and investigation determined that the set screws, which were part of the retrofit of a previous recall were incorrectly engaged at one facility which allowed rotation and subsequent descent of the assembly. | Process control | Urgent Device Correction Letters were sent to all sites affected by the retrofit on September 12, 2011 via Fed Ex. The letter identified the affected product and the description of the potential problem. Customers were asked to inspect their Synchrony System by following the instructions provided. If the lower extension feels loose, customers should contact Accuray Customer Support immediately to schedule service of the system. The letter also states that all affected systems will be inspected and corrected, if necessary by Accuray to prevent the potential problem. Accuray Customer Support will contact customers to schedule the inspection. Questions or concerns regarding the issue should be directed to Accuray Customer Support at 1-877-8667 for US customers and 1-408-716-4700 for customers outside of the US, or customersupport@accuray.com. |  | 236 | 236 total devices | Worldwide Distribution | Terminated | February 27, 2012 | 168 | 236
Z-0034-2012 | 59756 | No | FZS | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Axle Interspinous Fusion System Torque Driver Handle, Model #: X060-0323, Lot # 525941, S/Ns 001-017, Lot # 525942, S/Ns 001-050, Mfr. By: Bradshaw Medical Inc., Kenosha, WI 53144 |  |  | July 15, 2011 | 2011 | X Spine Systems Inc | The firm learned that the Torque Driver Handles were not actuating at a proper torque. It was determined that the handles were speed determinant and could output a higher torque if the handles were turned at a very slow speed. | Component design/selection | X-spine Systems, Inc. sent an URGENT DEVICE RECALL letter dated September 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their stock immediately to determine if they have any affected product. If so, customers were instructed once the set screw is tightened onto the locking plate, and resistance is detected in the handle, it will take less than 2 seconds to hear the handle "click." The "click" indicates the torque limit was reached. If further resistance is felt after 2 seconds, the surgeon should release the handle from tension, and try again.If the product was distributed further, they should contact their accounts, and advise them of this recall.The firm will send replacement handles.For any questions call (937) 847-8400, ext 115. |  | 39 | 39 Torque Driver Handles | Nationwide Distribution including CA, CO, FL, IA, IN, LA, NY, PA, TN, TX and UT. | Terminated | March 01, 2013 | 595 | 39
Z-0033-2012 | 59879 | No | FJP | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sodium Bicarbonate Mixing and Delivery System Software v 7.1. Model # MD-380, Serial Numbers: 5103031102, 5101101101, 5102161007, 5103221007, 5104221007, 5105101008, 5109131010, 5111231011, 5114301012, 5102131101 &amp; 5113091012. Manufactured by: Isopure Corp, Simpsonville, KY |  |  | May 19, 2011 | 2011 | Isopure Corp | Software problem causes transfer line between mixing and storage tank to not be rinsed after disinfection, leaving residual chlorine in the line. | Software design | Isopure sent a letter dated May 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Isopure devised an immediate solution which must be implemented at their faclity in conjunction with the disinfection procedure. Customers were asked to inform all appropriate personnel, clinical and technical, who may be involved in operating the system during the normal disinfection schedule.On 5/19/2011, the firm sent letters to their customers via US mail and email. They also telephoned their customers to notify them of the problem. |  | 11 | 11 units | Nationwide Distribution including AR, CT, GA, MA, ME, NJ, PA, TN and WV. | Terminated | February 19, 2013 | 642 | 11
Z-0043-2012 | 58746 | stains, hematology | KQC | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Harleco Giemsa Stain 25 gram glass bottle, Item #228-12 Manufactured for EMD Chemicals Inc., Gibbstown, NJ 08027 1-800-222-0342 |  |  | April 04, 2011 | 2011 | EMD Chemicals Inc. | Stains may yield lower then expected absorbance levels resulting in under stained hematology specimens. | Employee error | EMD Chemicals Inc. sent an "IMPORTANT PRODUCT RECALL" letter dated April 4, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. An attached response form is included for customers to complete and return via fax to EMD Chemicals at 856-423-6314. Any questions concerning this recall notice should be directed to EMD Chemicals Technical Service Department at 1-800-222-0342. |  | 36 | 36 units | USA Distribution in the states of PA and NY. | Terminated | October 13, 2011 | 192 | 36
Z-0037-2012 | 59898 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Paceart Integration Software, version 1.00 Model 6472. |  |  | August 12, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has determined that under certain conditions, LATITUDE Paceart Integration (LPI) softwaremay not properly transfer data resulting in the incorrect summary of programmed parameter values in the ICD Summary tab display and Current Programming Tachy summary section of the Paceart report. | Software design | Consignees were sent on 8/12/11 a Boston Scientific letter dated August 2011. The letter was addressed to "Dear Doctor". The letter described the product and problem. It also gave Background Information, Description of Potential Data Translation Error, and Additional Information and Action information.A letter, including an HRS form, was also sent to all US LATITUDE Clinics that have active LPI software to notify them of the need to install the corrected LPI Software. |  | 398 | 398 | Nationwide Distribution -- AL, AK, AZ, AR, CA, CO, CT DE, DC, FL, GA, IL,IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OJ, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA,WV, WI, and WY. | Terminated | November 06, 2012 | 452 | 398
Z-0042-2012 | 59891 | No | MSQ | Campylobacter fetus serological reagents. | Microbiology | Microbiology | 510(k) | IDkit:HP" Breath Test for Detection of H. pylori, Manufactured by: Exalenz Bioscience Ltd 4 Ha' maayan St. Modiin, Israel 71700, Distributed by: Exalenz Bioscience Inc., 1313 N. Market St., Sute 5100, Wilmington, DE 19808 USA |  |  | June 07, 2011 | 2011 | Exalenz Bioscience Ltd | The product is in violation of the FD&C; Act in that it was illegally marketed without proper clearance (no 510(k) submitted). | No Marketing Application | The firm, Exalenz Bioscience Ltd. notified consignees located in the United States of the recall by telephone, electronic mail, or written correspondence between June 9, 2011 and June 30, 2011. The letter informed consignees that Exalenz will immediately cease marketing the IDkit:HP TWO test in the United States. Consignees were instructed to return shipment of the IDkit:HP TWO test kits that they have in their possession to the companys US warehouse. For question call the VP of North American Sales at 704-770-6162 or VP of Marketing at 617-794-6936. |  | 1240 | 1240 units | Nationwide Distribution (USA) - including the states of: FL, IL, LA, MA, MS, NY, TN, TX and VA. | Terminated | July 06, 2012 | 395 | 1240
Z-0039-2012 | 59316 | microtome, rotary | IDO | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Thermo Scientific Microm: Rotary Microtome: REF HM355S-3 |  |  | August 12, 2011 | 2011 | Richard-Allan Scientific Company | The firm has received two complaints of finger lacerations in operators of this device. | Device Design | ThermoFisher Scientific sent "URGENT MEDICAL DEVICE CORRECTION, RE: Thermo Scientific HM355S-3 User Safety Advisory Notice" to all affected customers.The letter states the manufacturer identified a need to enhance the safety instructions and the description of the safety devices in the user manual. It mentions the HM3555S-3 User Manual has been modified to include more specific safe work practices and a new Quick -Reference Card summarizing the manual's safety enhancements and that the changes have been included on a CD that is enclosed with the letter. The firm requests confirmation of receipt of the new User Manual and Quick-Reference Card by October 31, 2011 to be sent by FAX to (269) 372-2674For questions on this recall call the firms Technical Support representatives at 1-800-522-7270 08:00 AM- 4:30 PM. |  | 786 | 786 | Worldwide Distribution: Nationwide distribution; including the country of Canada. | Terminated | July 25, 2013 | 713 | 786
Z-0038-2012 | 59888 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) EXL LOCI Cardiac Troponin I Calibrator (RC621) |  |  | August 31, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Internal testing has confirmed that recovery of levels 2 and 3 of LOCl Cardiac Troponin I Calibrator is lower than expected bottle value. Following calibration with LOCl Cardiac Troponin I Calibrator, lot 1AD045 an upward shift in patient and QC results may be observed. | Other | Siemens sent an Urgent Field Safety Notice dated August 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and discard any remaining inventory. Replacement orders would be processed and sent free of charge. Customers were also instructed to complete the attached FIELD CORRECTION EFFECTIVENESS CHECK form and fax it to (302) 631- 8457.Customers were instructed to forward the Safety Notice to anyone to whom they may have distributed the affected product. For any technical questions customers were to contact the Siemens Solutions Center at 800-441-9250. |  | 998 | 998 | Worldwide Distribution - USA (nationwide) including AL, AK, AZ, AR, CA, CO, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of GUAM, Austria, Canada, China, Czech Republic, Egypt, France, Germany, Italy, Poland, Portugal, Puerto Rico, Spain and Switzerland. | Terminated | May 20, 2013 | 628 | 998
Z-0040-2012 | 59712 | No | IWJ | Manual radionuclide applicator system. | Radiology | Radiology | 510(K) Exempt | Prostrate Seeding Applicator Needle. Product code PSS1820AT |  |  | August 12, 2011 | 2011 | Worldwide Medical Technologies, LLC | Length of the cannula hub is longer and may not permit a fit with the Applicator | Nonconforming Material/Component | BrachySciences, a division of Biocompatibles, Inc., sent out "Urgent Recall Device Recall" letters Dated August 12, 2011 to all affected customers.The letter included affected product and problem description. Customers are asked to quarantine and return unused product. A Product Recall &amp; Retrieval Form was included to send back to the firm.For questions on this recall contact the firm at (203) 262-0571, extension 105. |  | 1872 | 1872 units | Worldwide Distribution: Nationwide distribution including the states of AR, CO, DE. FL, KY, MD, MI, MN, NY, TX, and VA; and the countries of Belgium, and Germany. | Terminated | September 19, 2012 | 404 | 1872
Z-0041-2012 | 59418 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | ECG Cable Part Number: 01.57.471017 for use with the SE-1 series, SE-3 series and SE-601 series Smart Electrocardiograph |  |  | September 10, 2010 | 2010 | Edan Instruments, Inc. | The 3 leads of the cable, RL, LA and LL were labeled in wrong sequence. The left arm (LA) is mislabeled as right leg (RL). The left leg (LL) is mislabeled as left arm (LA). The right leg (RL) is mislabeled as left leg (LL). There have been no reports of injury as a result of this issue. | Vendor change control | Edan Instruments, Inc. sent an Urgent - Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that wrongly labeled cable may lead to incorrect ECG waveforms during ECG test, customers were reminded to review previous diagnoses made from this cable to mitigate any inappropriate recordings or treatment made.Customers were instructed to replace the affected ECG Cables once they receive the new ECG Cable, and collect all the affected ECG Cables and then destroy and dispose of them according to local regulations. Questions or concerns were directed to the following numbers: Oceania and Asia: +86 (755) 2689 8322North America, South America:+86 (755) 2689 9914Europe: +86 (755) 2669 4684 |  | 615 | 615 Ccables | Worldwide Distribution - USA (nationwide) and the countries of Dominica, Ecuador, Colombia, Honduras, Peru, Mexico, El Salvador, Guatemala, Venezuela, Uruguay, Chile, Panama, Nicaragua, Ecuador, Australia, Philippines, Malaysia, Indonesia, Chinese Taipei, &amp; Spain | Terminated | October 13, 2011 | 398 | 615
Z-0045-2012 | 58461 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PERPOS PLS System, Single Use; CATALOG #'s 9045-01 &amp; 9045-02. |  |  | February 03, 2011 | 2011 | Interventional Spine Inc | This recall has been initiated due to drivers that may contain a burr in the distal tip preventing full seating of the driver onto the head of the facet screw, resulting in difficulty screwing the facet screw into the bone.Use of this product may result in a delay of the surgical procedure | Nonconforming Material/Component | Interventional Spine, Inc. sent an "Urgent Recall" letter dated March 15, 2011 to all affected customers.The letter included; affected product, reason for recall and recommendations.Customers were asked to immediately examine inventory, quarantine affected product and to notify International Spine, Inc if any product remain.For information on this recall contact the firm at (949) 472-0006. |  | 2620 | 2,620 units | Word wide Distribution: (USA) Nationwide including the states of CO, CT, FL, KY, LA, MA, NC, NY, OK, OR, TN, TX, VA, and WI; and the countries of Italy, Spain, Turkey, and the U.K. | Terminated | February 05, 2013 | 733 | 2692
Z-0055-2012 | 59526 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 8F Safe Sheath Introducer, Catalog number: V18; |  |  | August 01, 2011 | 2011 | Angiodynamics, INC | The product may contain an incorrect size catheter and introducer. | Device Design | AngioDynamics, Inc contacted affected customers initially via telephone and a follow up "Notification of Medical Device Correction" letter was sent dated August 1, 2011. The letter informed them of the problem and provided instructions. Customers were instructed to utilize the 8F Introducer Sheath provided within the 8F Smart Port CT Implantable Port during port implantation. The 8F Valved Sheath (V18) may continue to be utilized during the placement of the 7.5F Smart Port Implantable Port (for example: CT75STSD; CT75STSD-NF). For question on this recall contact AngioDynamics, Inc at (800) 772-6446. |  | 1095 | 1095 total units | (USA) Nationwide distribution including the states of CA, FL, GA, IA, IL, IN, KY, MD, MI, MO, NC, NE, NJ, NY, OH, OK, OR, TN, UT, WA and WI. | Terminated | March 07, 2012 | 219 | 1095
Z-0048-2012 | 59768 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000rt System software; an automated system for performing fluorescence-based PCR that results in quantitative and qualitative detection of nucleic acid sequences; Abbott Molecular Inc., Des Plaines, IL 60018 USA; catalog #09K15-01 |  |  | September 08, 2011 | 2011 | Abbott Molecular | The Abbott m2000rt System software has software bug which can affect customers who manually enter a Test Order onto the m2000rt and sorts a column on the calibration or control grids of the Test Order screen by double clicking the column header. This causes a mismatch between specific concentration values and the designated calibrators and controls in the final results, which could cause an inval | Software design | Abbott Molecular Inc. sent a Field Correction Recall letter dated September 8, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until an upgrade to your m2000rt software can be installed. Abbott Molecular estimates that this software version will be available the first quarter of 2012.Please complete the enclosed reply sheet and fax it to (847) 775-6728 at your earliest convenience to acknowledge receipt. We appreciate your assistance in this matter. |  | 926 | 926 units | Worldwide Distribution -- USA (nationwide) including the countries of Algeria, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Belgium, Benin, Bolivia, Bosnia and Herzegovia, Brazil, Bulgaria, Burkina Faso, Cameroon, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Dem. Rep. of the Congo, Denmark, Egypt, Estonia, Ethiopia, Finland, France, Gabon, Gaza &amp; Jericho, Germany, Greece, Guatemala, Guinea, Haiti, Honduras, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Libya, Lithuania, Malawi, Malaysia, Mali, Mauritius, Mexico, Moldova, Morocco, Mozambique, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Togo, Tunisia, Turkey, United Arab Emirates, Uganda, Ukraine, United Kingdom, Venezuela, Vietnam, Zambia and Zimbabwe | Terminated | March 08, 2016 | 1643 | 926
Z-0044-2012 | 60008 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | AVAIRA Toric (enfilcon A) TORIC soft (Hydrophilic) Contact Lenses, 6 SOFT CONTACT LENSES FOR ASTIGMATISM, CooperVision, Scottsville, NY 14546 USA; Juana Diaz, PR 00795 USA; Hamble SO31, 4RF, UK; Device Listing # D033736. |  |  | August 19, 2011 | 2011 | CooperVision Inc. | This recall was initiated because of the unintended presence of a residue (silicone oil) on lenses from certain lots of Avaira Toric contact lenses. The company received complaints of temporary hazy or blurry vision and discomfort, most likely caused by transient or reversible swelling of the surface layer of the cornea, which is known as corneal epithelial edema. Corneal epithelial edema results | Material/Component Contamination | The firm, CooperVision sent an "Urgent: Company Initiated Voluntary Medical Device Recall - Contact Lenses" notification letter to all affected consignees. The notice identified the product, the problem, and the actions to be taken. The letter includes an attached "Recall Response Form" that consignees are required to complete and fax to (585) 889-5688. Consignees were instructed to immediately examine their inventory, stop any further distribution, quarantine and return product subject to recall. In addition, consignees were told if they have further distributed the affected product to notify customers at once of this product recall. The letter states that accounts will either be credited or exchanged for an unaffected product. If you have any further questions, call 1-800-341-2020. |  | 778301 | At least ( 778,301 lenses: Domestic - 659,847 lenses; Foreign - 118,454 lenses; including EXPANSION: Domestic - 69,047 lenses) ( to be updated) | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Australia, Canada, Europe, Germany, Hong Kong, Italy, New Zealand, and Spain. | Open, Classified |  |  | 778301
Z-0049-2012 | 59978 | [6 empty] | ArthroCare brand MiniMagnum Knotless Fixation Implant Device with Inserter Handle Plus SpeedStitch Magnum Wire Suture Cartridge; Catalog Number: OM-6007; |  |  | May 11, 2011 | 2011 | ArthroCare Corporation | ArthroCare initiated the recall of one lot of the MiniMagnum Knotless Implant Fixation Device where the incorrect diver block was packed with the implant. | Packaging process control | The firm, ArthroCare Corporation, sent a letter dated May 10, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were asked to immediately return the product; check their inventory, refer to the copy of the Mini Magnum outer label to help identify the product, and if the affected product lot is in their inventory, notify ArthoCare Product Support at (800) 797-6520, press 2 and return the product immediately. The letter also stated, "In addition, if you do not have any product from this lot remaining in your inventory, we request that you contact ArthroCare at the number above, or by e-mail at ra@arthrocare.com, to inform us of the lack of remaining product. This will ensure that you will not be contacted again and that we have accounted for product shipments from this lot." If you have any questions, please call the Chief Regulatory Officer and Senior Vice President, Quality Systems and Assurance at (408) 735-6229. |  | 32 | 32 units | Nationwide distribution: USA including states of: CA, FL, NY, OH and OK. | Terminated | October 24, 2011 | 166 | 32
Z-0050-2012 | 53046 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems Total ÂżhCG. |  |  | March 05, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that certain Access Total ÂżhCG reagent packs show decreased stability from the date of manufacture, which may lead to increased imprecision. | Nonconforming Material/Component | The recall communication was initiated on March 18, 2009 with Beckman Coulter forwarding an Urgent: Product Corrective Action (PCA) letter dated March 13, 2009 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Access Immunoassay Systems total ÂżhCG. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to discontinue use of the reagent lot numbers 824542 and 825652, and discard any unused reagent packs.Customers were asked to review their historical records for low level quality control to ensure that the assay performance for the referenced lot numbers has been within their laboratory specifications. They were also asked to share the information with their laboratory staff and retain the letter for their Quality System documentation, and then complete and return the enclosed response form within 10 days so that they may be assured that consignees received the notification.Finally they were told to contact Customer Technical Support Center at 1-800-854-3633 in the US and Canada, or their local Beckman Coulter Representative with any questions. |  | 14764 | 14,764 | Worldwide Distribution -- US, Algeria, Angola, Argentina, Armenia, Australia, Bangladesh, Belgium, Brazil, Bulgaria, canada, Chile, China, Colombia, Czech Republic, Egypt, El Salvador, Ethiopia, Finland, France, French Polynesia, Germany, Ghana, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Kenya, Korea, Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macao, Mexico, Monaco, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Reunion, Russian Federation, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Zambia. | Terminated | April 27, 2012 | 1149 | 14764
Z-0054-2012 | 59912 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | SARNS INTRACARDIAC SUCKER, BULK, NON-STERILE |  |  | September 13, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) recently discovered three instances of metal fragments inside the stainless steel tube of the Sarns" Adult Rigid Intracardiac Sucker. | Nonconforming Material/Component | On 9/13/2011 Terumo sent a letter entitled "URGENT MEDICAL DEVICE REMOVAL" to its customers that included the name and description of the device, lot numbers, the potential hazard, correction, the reason for the recall and instructions. Customers are to review the notice, ensure that all users are aware of the issue, and a recall response form to be completed and returned. Customers are to call Terumo CVS Customer Service to obtain a Return Goods Authorization number. The letter also included MedWatch reporting information and stated that Customers with questions should call 1-800-521-2818. |  | 5700 | 5700 | Worldwide Distribution -- USA, Belgium, Australia, Mexico, and Japan. | Terminated | December 10, 2012 | 454 | 15470
Z-0052-2012 | 59961 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Impac Medical Systems Sequencer |  |  | September 09, 2011 | 2011 | Impac Medical Systems Inc | Varian reported a problem related to certain collimator rotation values which are incorrect on TrueBeam compared to planned values in Mosaiq. | Nonconforming Material/Component | "Impac Medical Systems Inc., sent an USER NOTICE letter to all affected consignees on September 20, 2011. The letters identified the source characterization error, clinical impact, and final resolution. The letter states that Elekta Support will follow up with each consignee and remotely correct the source characterization. Customers were instructed to sign, date and return the attached Confirmation of Receipt Form by one of the three ways indicated. The letter instructs consignees to distribute this notice to any and all users of IMPAC software that may potentially be affected by the issue. For questions, concerns, and requests for upgrade, email support@impac.com or call. The letter provided consignees with a complete list of contact information.. |  | 11 | 11 devices | Nationwide Distribution (USA) - including the states of: AZ, CA, FL, NJ, NY, NV, TX, VA and WA. | Terminated | March 06, 2012 | 179 | 11
Z-0051-2012 | 59848 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-300ES, 621-400ES, 621-400ESR. |  |  | September 07, 2011 | 2011 | B Braun Medical, Inc | B Braun Outlook ES Safety Infusion System "Run" Light Emitting Diodes (LEDs) pumps stop infusing and a backup alarm sounds, but the "Run" LEDs advance as if the pumps were infusing. | Software design | B. Braun Medical, Inc. sent an "URGENT: FIELD CORRECTION" letter dated September 7, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. A firm representative will contact all affected customers in order to schedule the field correction. The correction consists of revising the Door and Main Processor software, inspecting the main board, and replacing the board, if warranted. A Field Correction Information Form was included in the letter for customers to complete and return. Contact Technical Product Support at 1-800-627-7867 for questions regarding this notice. |  | 6321 | 6321 | Nationwide Distribution | Terminated | May 04, 2015 | 1335 | 6321
Z-0057-2012 | 59442 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | MicroMaxx ultrasound system and M-Turbo ultrasound system is Medical Diagnostic Ultrasound. |  |  | July 18, 2011 | 2011 | Sonosite, Inc. | This is a software anomaly. In some circumstances, if the Systolic/Diastolic (S/D) ratio value is interpreted directly from the display, the clinician could potentially use an incorrect value of the S/D ratio in his/her assessment. | Software design | The firm SonoSite, Inc., sent an "Urgent - Medical Device Correction" letter dated September 29, 2011 to its foreign and domestic customers. The letter described the product, problem, and actions to be taken. The customers were instructed to follow the three (3) steps provided in the letter to verify if the extended OB/GYN Calculations (EOBC) is enabled on their system. Additionally, SonoSite will perform a software upgrade to M-Turbo and/or MicroMax ultrasound system in order resolve the issue. The customers were instructed to contact Sonosite Technical Support via website: www.sonosite.com/sdratio; email: regulatory@sonosite.com (include contact information and address), or call in U.S. and Canada: 1-877-590-4982. If you have any questions or concerns regarding this matter, do not hesitate to contact in US and Canada at 1-877-657-8118, Monday-Friday, 6 am-5 pm PST. Customer can also send an email to Technical Support at service@sonosite.com |  | 4236 | 4236 units (US) and 4089 units (internationally) | Worldwide distribution: USA (nationwide) including Puerto Rico; and countries of: Afghanistan, Algeria, American Samoa, Argentina, Australia, Bolivia, Botswana, Brazil, Canada, China, Costa Rica, Cyprus, Dominican Republic, Estonia, France, Gambia, Germany, Great Britain, Guam, Guatemala, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Italy, Japan, Jordan, Kenya, Korea, Kuwait, Mauritius, Mexico, Netherlands, New Zealand, Nigeria, Northern Mariana Islands, Oman, Panama, Paraguay, Peru, Philippines, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Tanzania, Uganda, United Arab Emirates, Viet Nam, and Zambia. | Terminated | September 14, 2012 | 424 | 4236
Z-0060-2012 | 59998 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ANCHORLOK(R) LEADING EDGE(R) SOFT TISSUE ANCHOR SYSTEM, 2.5mm, 2.0 SUTURE, REF 5100-006, 1 each, STERILE R, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002 |  |  | September 20, 2011 | 2011 | Wright Medical Technology Inc | The product is labeled with an expiration date without adequate justification. | Incorrect or no expiration date | The firm, Wright Medical Technology, Inc., sent an "URGENT: PRODUCT RECALL NOTIFICATION" letter dated September 20, 2011 to all affected consignees. The letters identified the products, the problem, and the actions to be taken. The letter instructed consignees to identify affected inventory. If recalled products are located consignees are instructed to contact customer service at 1-800-238-7117, for specific instructions on how to return the recalled products. The letter asked that this communication be forward to any associated sites if the recalled products were further distributed. Consignees are to complete and return the enclosed notification form by fax to 901-867-7401 or by email to: cathy.park@wmt.com. For questions regarding this recall contact 901-867-4324. |  | 596 | 596 units | Worldwide Distribution - USA (nationwide) and countries of: Argentina, Canada, Denmark, France, Great Britain, Ireland, Japan, and Spain. | Terminated | July 17, 2012 | 301 | 11960
Z-0071-2012 | 58302 | No | LPA |  |  | Cardiovascular | PMA | ***REF #200***TAPSCOPE 550. Each pouch contains one (1) 18 French esophageal stethoscope with bipolar pacing and recording electrodes with a YSI 400 series male connector thermistor. LOT xxxxxxx 400095-0200-C 100902. www.cardiocommand.com. |  |  | October 25, 2010 | 2010 | Cardiocommand Inc. | CardioCommand, Inc. is recalling their TAPSCOPE 550 esophageal pacing/recording catheters because there maybe electrical discontinuity, which can result in the inability to reliably pace and/or record ECG. | Nonconforming Material/Component | On October 25, 2010, an Urgent Medical Device Recall Notice was sent by certified mail to all facilities advising them of the recall. The letter identified the affected product and reason for the recall. Users were asked to discontinue use and return the affected product to the firm. If all product was already used, then this should be indicated on the form provided. The firm will send replacement product in quantities equal to the returned product. The attached form should be completed and returned via fax. Customers were instructed to report any suspected incidents related to this issue to CardioCommand Customer Service at 1-800-231-6370. Also questions or comments should be directed to CardioCommand Customer Service. |  | 567 | 567 total | Nationwide Distribution -- GA, OR, TX, CO, KY, MT, MA, AL, CA, FL &amp; TN. | Terminated | August 28, 2012 | 673 | 567
Z-0068-2012 | 60000 | No | KGT | Drape adhesive. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STRIP T's STERILE ADHESIVE ORGANIZER, Model #SC5000-4R, Lot #090710, 510(k) exempt. One (1) adhesive organizer is packed in a Tyvek pouch and 60 pouches are placed into a 'Shelf Box'. The shelf box is placed into a shipping container. |  |  | March 05, 2008 | 2008 | Kapp Surgical Instrument Inc | A portion of the labeled expiration date is missing. The expiration date on approximately half of the pouch labels distributed are missing the last digit of the year of expiration. However, the shelf box containing the pouches is labeled correctly. | Incorrect or no expiration date | Surgical Concepts, Inc ( A subsidiary of Kapp Surgical Instrument Inc) sent a notification letter dated March 5, 2008, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm ceased distribution of the suspect product on December 5, 2007. The letter informs the firm's customers (both distributors and medical user/customers) of the expiration date labeling omission of the last digit of the year of expiration. It goes on to explain that the incomplete expiration date is present on the individual packages only and that the shelf boxes are labeled correctly with the complete/correct expiration date. The letter provides the customers with a contact person and a toll free telephone number 1-(800) 282-5277, to call in order to make arrangements for receive replacement labels for any misbranded packages received. |  | 59 | 59 boxes (3540 units) | Nationwide Distribution-- including the states of AK, CA, CO, FL, GA, ID, IL, IN, LA, MA, MD, MI, MO, MT, NC, NM, OH, PA, TN, TX, UT, and VA | Terminated | February 19, 2013 | 1812 | 59
Z-0070-2012 | 60007 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre 9F Pruitt F3 Outlying Carotid Shunt with T-port |  |  | October 04, 2011 | 2011 | Lemaitre Vascular, Inc. | Extended Expiration on the outer box label 2016-05, the correct expiration date is on the tray label: 2014-05 | Incorrect or no expiration date | The firm, LeMaitre Vascular, Inc., sent an "URGENT: 9F PRUITT F3 Carotid Shunt Device Field Safety Notice" dated October 4, 2011 to its customers. The notice describes the product, problem and action to be taken. The customers were instructed to identify their inventory; return the affected products for exchange with a properly labeled device; contact Le Matire customer service for replacement at 800-628-9470, and complete and return the RECALL NOTICE RECONCILIATION FORM via fax to 781-425-5049, even if no product is on hand. If you have any questions concerning this safety notice, please contact the Quality Assurance Engineer at 781-221-2266 ext. 183. |  | 144 | 144 units | Nationwide distribution: USA including states of: CA, FL, IA, IL, MA, MI, NE, NJ, NY, OH, SC, TX, and WI. | Terminated | March 14, 2012 | 162 | 144
Z-0069-2012 | 59889 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Inc. Dimension(R) EXL Cardiac Troponin I Flex(R) Reagent Cartridges |  |  | August 31, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed customer complaints for low recovery of QC and patient samples. | Process control | Siemens Healthcare Diagnostics, Inc. Urgent Field Safety Notice was sent to customers on August 31, 2011 regarding the issue. The letter instructs customers to immediately discontinue use of the affected lots and discard all remaining inventory. |  | 6072 | 6,072 cartons | Worldwide distribution: USA, Puerto Rico, Guam, Austria, Canada, Croatia, France, Greece, Germany, Italy, Japan, Mexico, Portugal, Spain and Switzerland. | Terminated | January 14, 2014 | 867 | 6072
Z-0076-2012 | 59911 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MICROPLASTY TAPERLOC 7.5 MM RESECTION GUIDE, STAINLESS STEEL, BIOMET ORTHOPEDICS 56 EAST BELL DRIVE PO BOX 587 WARSAW IN 46582 USA. |  |  | September 16, 2011 | 2011 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that the blueprint calls forengraved measurements on both sides on the Microplasty Taperloc 7.5MM Resection Guide. Theabove lots, as supplied to Biomet by a vendor, are engraved on only one side. To date, Biomet hasreceived 2 reports involving 2 units where the instrument was missing the engraving on one side. | Nonconforming Material/Component | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated September 16, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to locate and discontinue use of the affected product. Arrangements will be made for customers to return the affected product to Biomet. To confirm receipt of this notice call Biomet at 800-348-9500 ext. 3755 or 3756. Questions related to this notice should be directed to 574-371-3755 or 3756, Monday through Friday, 8 AM to 5 PM. |  | 22 | 22 | Worldwide Distribution-USA (nationwide) including the states of AL, CA, IN, MI, OK, and VA and countries of Australia and Chile. | Terminated | March 07, 2013 | 538 | 22
Z-0075-2012 | 59968 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Total Hip Joint Replacement |  |  | July 14, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that a size 3 Accolade stem was in a box labeled as a size 5 Accolade stem. | Mixed-up of materials/components | Stryker Orthopaedics sent an "URGENT: PRODUCT RECALL" letter dated July 14, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and return an attached Product Recall Acknowledgment Form via fax to 201-831-6069. Additionally, information is provided for customers to return the affected products. Please contact the Divisional Regulatory Compliance Manager at 201-972-2100 for questions regarding this notice. |  | 21 | 21 units | Nationwide Distribution | Terminated | November 15, 2013 | 855 | 21
Z-0079-2012 | 59948 | No | DXS | Cardiopulmonary bypass coronary pressure gauge. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Custom Tubing Pack, which are designed as specified by the user. The following models contain a Pressure Display Set. Sterilized by Ethylene Oxide. Non-pyrogenic. Manufacturer: Medtronic, Inc. 710 Medtronic Parkway, Minneapolis, MN 55432-5604. |  |  | September 09, 2011 | 2011 | Medtronic Perfusion Systems | Pressure Display Set (models 61000, 61006, 62000, 62006) DFU will be updated to clarify that the intended use of the disposables is only for cardiopulmonary bypass surgery (procedures up to 6 hours). Use beyond 6 hours may result in air leakage in the system resulting in a pressure reading that is not representative of the actual line pressure. | Device Design | Medtronic Perfusion Systems, sent an "Urgent Medical Device Notification letter dated September 9, 2011, via UPS 2-day delivery to all affected customers. The letter identified the affected devices, described the issues, and the actions needed to be taken. The letter indicates when approved, the updated Directions for Use for the Pressure Display Box and Pressure Display Sets will be sent to customers. Customers were asked to complete the Customer Confirmation Certificate and return to Medtronic per the instructions on the form. The letter states that if the recalled products have been further distributed, forward this letter immediately to all customers who received the affected products, and gather proof of notification. For further questions concerning this notification, contact Medtronic Cardiovascular Technical Support at 877-526-7890 or Medtronic Sales Representative. |  | 74804 | 74,804 | Worldwide Distribution - USA (nationwide) including Puerto Rico, and Virgin Islands (U.S.) and the countries of: Argentina, Australia, Austria, Belarus, Belgium, Bulgaria, Canada, Chile, China, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, El Salvador, France, Georgia, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Latvia, Lebanon, Libya, Malaysia, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Poland, Portugal, Romania, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom, Venezuela, and Yemen. | Terminated | October 11, 2012 | 398 | 270088
Z-0083-2012 | 52105 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instrument Manager -Software Versions 8.05, 8.06, 8.07, 8.08, and 8.09. |  |  | May 11, 2009 | 2009 | Data Innovations, Inc. | The "previous result" retrieved from the Specimen Management Database may not be the most recent "previous result". | Software design | Data Innovations sent an "IMPORTANT PRODUCT NOTICE" dated May 11, 2009 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter offers temporary workarounds for customers to use until corrected versions of the software is complete. Versions of the corrections will be provided on a staggered schedule. Customers will be notified by e-mail. Questions regarding this notice are directed to Support at 802-658-1955. |  | 188 | 188 | USA (nationwide) and the country of Canada. | Terminated | October 25, 2011 | 897 | 188
Z-0085-2012 | 59892 | No | FCM | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | O'Neil Sterile Field Intermittent Urinary Catheter Kit labeled in mart:"***ORIGINAL CLOSED SYSTEM BY DR O'NEIL O'NEIL STERILE FIELD INTERMITTENT URINARY CATHETER KIT WITH IODINE SWABSTICKS (3)*** REF: ONSM 16KPES SIZE:16Fr Product No: AAM 8014KPES***SINGLE USE ONLY***Original design by: Dr. George O'Neil 200 Churchill Ave, Subiaco, WA 6008, Australia Tel.: +61 (0)8 9388 1700 Fax: +61 (0)8 9382 4952 European Representative: Advena Ltd Hereford, HR4 9DQ, UK DISTRIBUTED IN THE USA BY: American Australian Medical, Inc. 1-888-880-8871***" |  |  | September 03, 2011 | 2011 | Go Medical Industries Pty., Ltd. | O'Neil Sterile Field Intermittent Urinary Catheter Kits With Povidone Iodine Swabsticks are being recalled due to a recall on the Triad Povidone Iodine Swabsticks contained within the kit. H&P; Industries are recalling the Triad Povidone Iodine Swabsticks because it was determined that the swabsticks were manufactured without having a system for microbial testing in place at the time of release. | Nonconforming Material/Component | The firm, Go Medical Industries Pty Ltd, sent an e-mail on September 3, 2011 and follow-up with an "URGENT: MEDICAL DEVICE RECALL" letter dated September 07, 2011 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to quarantine and not sell the product in the market place; immediately perform a sub-recall and notify their customers and complete and return the RETURN RESPONSE FORM via fax at 716-883-8099 or mail to: American Australian Medical Inc, 1083 Delaware Avenue, Buffalo, NY 14209.For further details you can contact QA Manager at Go Medical Industries at +61 89 381 3226 or pmoretta@gomedical.com.au. |  | 14700 | 14,700 | Nationwide distribution. | Terminated | January 30, 2012 | 149 | 235243
Z-0104-2012 | 60004 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Spine Zero-P Instruments and Implants. |  |  | September 23, 2011 | 2011 | Synthes USA (HQ), Inc. | LABELING CORRECTION - Following review of labeling and promotional materials associated with the Zero-P Implant system, it was found that the materials reference a study where PEEK-CR is utilized as a cervical interbody spacer. PEEK-CR is not currently indicated for use as a cervical interbody spacer. | Labeling False and Misleading | Synthes (USA) sent an "URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated September 23, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to review the updated supplemental labeling page included with the letter and replace/attach the technique guide page in any Zero-P technique guide that customers may have. The Verification Section of the letter is to be completed and returned to Synthes via fax at 610-251-9005. For questions regarding this notice call 1-800-620-7025 ext. 5403. |  |  |  | USA (nationwide) including Puerto Rico, Washington, DC and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Terminated | December 18, 2014 | 1182
Z-0090-2012 | 59837 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***REF QD11-6B***6 mm Fluted Ball. Use with QD11 Attachments. ***ANSPACH***Sterile, Sterilized using Irradiation***Rx Only***Made in the USA***Manufacturer: The Anspach Effort, Inc. 4500 Riverside Drive, Palm Beach Gardens, FL 33410 USA Tel: (800) 327-6887 / + 1 561 627 1080***Anspach Europe Ltd. Aston Court, Kingsmead Business Park Frederick Place High Wycombe, HP11 1LA United Kingdom Tel: +441494 616126***" |  |  | May 19, 2008 | 2008 | The Anspach Effort, Inc. | The Anspach Effort, Inc. initiated a recall on May 6, 2008 for their Sterile Bone Cutting Burrs. All affected customers were notified of the recall by letter dated 5/29/2008. | Employee error | The firm, Anspach sent an "URGENT PRODUCT REMOVAL" letter on May 29, 2008, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) Screen their inventory immediately for all products listed on the attached "Product Replacement Form." Upon visual inspection if the pouch is sealed on both ends it is sterile. If one end of the pouch is not sealed it is non-sterile and should be returned for credit or replacement. 2) Complete the "Product Replacement Form" by indicating the quantity returned, and whether the customer would prefer to have their returned devices replaced, or their account credited. 3) Place products to be returned in a box and mark the box with the Return Authorization Number on the "Product Replacement Form." 4) Use the Anspach FedEx account number listed on the form to return the unused devices with the completed "Product Replacement Form." Following receipt of the form, Anspach will process the request immediately. For questions related to the product replacement, please contact Customer Service at 800-327-6887. For questions related to the Voluntary Removal, please contact Anspach Regulatory Affairs and Quality Assurance at 800-327-6887. |  | 7 | 7 | Worldwide Distribution - USA (nationwide) including the states of: CA, CO, FL, GA, IA, ID, SC, MA, MD, MI, MO, MN, MS, NY, OH, OR, PA, TN, TX, and WI, and the countries of: Africa, Netherlands, Panama, South Korea, Spain, and United Kingdom. | Terminated | October 24, 2011 | 1253 | 905
Z-0103-2012 | 60047 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hamilton brand Disposable Syringe Pack for use with Microlab F.A.M.E; Part Number: 147692 (case containing 96 syringes); Manufactured by: Treff AG, Degersheim, Switzerland for Hamilton Bonaduz AG, Bonaduz, Switzerland (Manufacturer of record); Product is distributed by Hamilton Company, Reno, NV |  |  | September 26, 2011 | 2011 | Hamilton Co | Certain lots of syringes have caused an increase in error message "DIS541", which indicates "improper dispensation". | Nonconforming Material/Component | The firm, Hamilton, sent an "Urgent: Product Recall Notification" letter dated September 26, 2011 to its customers. The letter describes the product, problem and action to be taken. The customers were instructed to return the syringes with a completed Hamilton Company Product Recall Questionnaire form via fax to: (775) 856-7259 or with an emailed signed pdf copy to wayne.mcauliffe@hamiltoncompany.com. The firm's Customer Service Department will contact customers with complete Returned Material Authorization information. If you have any questions, please do not hesitate to contact the Director of Quality at (775) 858-3000. |  | 1096 | 1096 units | Worldwide distribution: USA(nationwide) including states of: NJ, NV, TX and WA; and countries of: Brazil, P.R. China, France and Germany. | Terminated | March 06, 2012 | 162 | 1096
Z-0082-2012 | 60054 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew, HD1200 Autoclavable Camera Head |  |  | September 23, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | HD1200 camera heads may cause image issues | Process design | Smith Nephew sent an "URGENT  PRODUCT RECALL NOTIFICATION" letter dated sent September 23, 2011 to all affected consignees. The letter identified the product, problem and actions to be taken. Consignees were instructed to complete requested information, return affected product and order replacement product. To return product and order replacement product, consignees were instructed to contact the Smith &amp; Nephew Endoscopy Returns Group at 800-343-5717 (option 3) for a Return Authorization (RA) or send an email to endo.andreturns@smith-nephew.com. For questions call 405-917-8556. |  | 15 | 15 units | Nationwide Distribution (USA) including the states of: MI, NH, OK. | Terminated | February 16, 2016 | 1607 | 15
Z-0081-2012 | 59981 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Medpro AccuFlo Elastomeric Infusion Device, Model #CT-0020-270C, packaged 12/box. The box label shows the product is Manufactured By: Medpro International (Thailand) Ltd. for Medpro Corporation Pte. Ltd., Singapore. |  |  | July 28, 2011 | 2011 | Progressive Medical Inc | The product may have a higher flow rate than specified resulting in over administration of drug solutions to the patients. | Nonconforming Material/Component | The firm, Progressive Medical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated July 28, 2011 via email and regular mail to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue their use; remove the products from inventory and quarantine them; inform their customers to return unused products to Progressive; return the product, and complete and return the RECALL RETURN RESPONSE FORM via fax to 314-961-5786. Progressive customer service will contact the customers to coordinate the return of the goods in question.Should you have any questions, please call 800-969-6331 or email to dsullivan@progressivemedinc.com. |  | 180 | 180 units | Nationwide distribution: USA including states of: FL, KY, PA, VA and WY. | Terminated | December 19, 2011 | 144 | 180
Z-0080-2012 | 59984 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | "***Vipat REF H0020-100***Lot 100533***2015-04***STERILE EO Elastomeric pump***2mL/Hr***100mL***Medpro International (Thailand) Ltd. 360, Moo 6, Chonburi Industrial Estate, Tambon Bowin, Amphur Sriracha, Chonburi 20230, Thailand***Victus Inc. 4918 SW 74 Court 33155 44400***" |  |  | September 20, 2011 | 2011 | Victus, Inc. | Victus, Inc. Miami, FL is recalling Vipat Brand Elastomeric Infusion Pump, Model H0020-100, Lot 100533. The product may have a higher flow rate than specified. The use of this product may lead to over-administration of drug solutions to patients. | Nonconforming Material/Component | Victus, Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated July 15, 2011 to affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use, remove the products from inventory and quarantine the products. Additionally, the firm indicated that their customer service department would contact customers to coordinate the return of affected products. Customers are directed to call 305-663-2129 ext. 110 for questions regarding this notice. |  | 63 | 63 units | USA (nationwide) including Puerto Rico and the states of FL and PA. | Terminated | January 12, 2012 | 114 | 63
Z-0105-2012 | 60017 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OPMI VARIO 700 Surgical Microscope. |  |  | September 26, 2011 | 2011 | Carl Zeiss Meditec, Inc. | An incorrect nut and bolt set was used in the axis three (3) spring assembly of the OPMI VARIO 700 surgical microscope distributed to three US locations. | Nonconforming Material/Component | Carl Zeiss Meditec had Field Service Engineers directly contact all affected customers. The Field Service Engineers will replace the defective component. Questions regarding this Field Corrective Action should be directed to the Director of Clinical/Regulatory Affairs at 925-557-4616. |  | 3 | 3 | USA (nationwide) including the states of FL, GA, and LA . | Terminated | November 04, 2011 | 39 | 3
Z-0106-2012 | 58984 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | "***Assembled for Invivo Corporation 12501 Research Parkway Orlando FL 32826***Made in U.S.A.***MRI Patient Monitoring System Model 865214***Use only with Model 989803169201 Power Adapter***REF 453564155341***SN US11401652***". |  |  | August 30, 2010 | 2010 | Invivo Corporation | Philips Invivo Corporation of Orlando, FL is recalling their Expression MRI Patient Monitoring System. During transport of the system's cart with a docked Display Controller Unit (DCU), a front caster (wheel) may become loose and fall off causing the cart to become unbalanced and fall. | Device Design | Philips Invivo Corporation issued a Field Safety Notice "URGENT - MEDICAL DEVICE RECALL" letter to all affected customers notifying them that the Expression MRI Patient Monitoring System could pose a risk for patients and users. The notice identifies the following: what the problem is and under what circumstances it can occur; the actions that should be taken by the user in order to prevent risks for patients or users and the actions planned by Philips Invivo to correct the problem. This notice contains important information for the continued safe and proper use of the equipment and customers were advised to retain a copy with the equipment Instruction for Use. Customers were instructed to contact Philips Invivo Customer Support at (877) 468-4861 (option 1, then option 3) for questions or support concerning this issue. |  | 285 | 285 units | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI., SC, SD, TN, TX, VA, WA, and WI and the countries of: Australia, Austria, Canada, Denmark, Germany, Israel, Malaysia, Netherlands, Norway, Singapore, Spain, Switzerland and, United Kingdom. | Terminated | November 15, 2012 | 808 | 285
Z-0089-2012 | 59969 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | CORFLO 300 Enteral Pump; a volumetric rotary peristaltic intermittent enteral feeding pump; CORPAK MedSystems, Wheeling, IL 60090; catalog/reorder #20-0500. |  |  | September 19, 2011 | 2011 | CORPAK MedSystems | Corpak is removing the Corflo 300 Enteral Pump from the market due to the possibility that with fluid ingress or other damage, the keypad may not function, and the possibility of motor mount failure. | Process control | Corpak MedSystems sent Recall Notification letters dated September 19, 2011 to all CorFlo 300 Enteral Pump customers on the same date. The customers were informed that there is a possibility that with fluid ingress the keypad may become inoperable, and that there is a possibility of motor mount failure. Either event could possibly affect the rate at which formula is delivered to the patient, and fluid ingress could cause a potential of audible alarm failure. Corpak determined that the removal of the pumps from the market was the best option for their customers, and requested the return of all CorFlo 300 pumps to Corpak by no later than December 1, 2011. The customers were requested to contact Corpak at 1-800-323-6305 to arrange for the return of the pumps, and to complete and return the enclosed acknowledgement form, indicating the number of pumps on hand, via fax at 847-541-9526 or e-mail to questions@corpak.com. Dealers/distributors were requested to notify their customers of the recall. |  | 13000 | 13,000 pumps | Nationwide Distribution | Open, Classified |  |  | 13000
Z-0112-2012 | 58454 | No | BYF | Venturi mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Mask adult, elastic strap style, Percent-O2 Lock air entrainment system, w/7' tube (Linde) - 50/cs; |  |  | March 25, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | A small number of oxygen mask components were found to have 100% obstructed airways rendering them incapable of delivering oxygen to patients undergoing respiration therapy. | Equipment maintenance | The firm, Salter Labs, called all their customers on March 25, 2011 and asked that they hold all affected product. In addition, the firm sent an "URGENT PRODUCT RECALL" letter dated April 4, 2011 and a follow-up letter dated April 7, 2011 via e-mail and certified mail to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory, return any affected lots listed, contact Salter Customer Service at 1-800-235-4203 to arrange for return of affected product and complete and return the Customer Reply Form via 661-854-6816 or toll free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 800 | 800 all varieties | Worldwide distribution: USA (nationwide) and countries of: Belgium, Canada, Colombia and Iceland. | Terminated | October 26, 2011 | 215 | 800
Z-0115-2012 | 59326 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE DS FULL FIELD DIGITAL ESSENTIAL MAMMOGRAPHY SYSTEM. |  |  | September 22, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of inaccurate measurement on GE magnified images provided by Senographe DS and Essential Mammography systems when reviewed on IHE compliant workstations. This issue may impact patient safety. | Software design | GE Healthcare sent an " URGENT MEDICAL DEVICE CORRECTION" letter dated September 21, 2011 to all affected customers. The letter describes the products, problem, and actions to be taken by the customers. The letter provides safety instructions and advises customers to complete and return a Customer Reply Form via fax to +33(0)1 30 70 39 30. A GE Healthcare service representative will contact customers if the updated software version is required. Contact the Call Center at 800-437-1171. |  | 1344 | 1344 | Worldwide Distribution-USA (nationwide) including Puerto Rico and Washington, D.C. and the countries of ALGERIA, AUSTRALIA, AUSTRIA, BELGIUM BOTSWANA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA,GERMANY, GREECE, GUADELOUPE, GUYANA, HONDURAS, HONG KONG, HUNGARY, INDIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PHILIPPINES, POLAND, PORTUGAL, REPUBLIC OF KOREA, RUSSIAN FEDERATION SAINT PIERRE AND MIQUELON, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE FORMER YUGOSLAV REPUBLIC OF MECEDONIA, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, AND VENEZUELA. | Terminated | October 16, 2012 | 390 | 3457
Z-0107-2012 | 59522 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE HEALTHCARE, MSK 1.5T EXTREME MR SCANNER, MODEL AA5000. |  |  | August 29, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware that an exposed high voltage gradient wire internal to the MSK 1.5T Extreme system equipment may cause electrical shock, which may impact service personnel safety. This issue does not impact patient or operator safety. | Nonconforming Material/Component | GE Healthcare LLC sent a "Urgent Medical Device Correction letter" dated August 29, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to please ensure that all potential service personnel at your facility are made aware of this safety notification and the recommended actions.Contact your local service representative if you have any questions concerning this notification. For further questions please call ( 262 ) 513-4122. |  | 52 | 52 | Worldwide Distribution -- USA (nationwide) including the states of AL, CA, FL, ID, NY, NH, NJ, NV, OH, PA, TX, VA and WI., and countries of UNITED ARAB EMIRATES, UKRAINE, UNITED KINGDOM, SWITZERLAND | Terminated | May 29, 2012 | 274 | 52
Z-0074-2012 | 59933 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion EnVe Ventilator, Model 19250-001. Manufactured by: CareFusion, 22745 Savi Ranch, Yorba Linda, CA 92887 USA. EnVe is a portable critical care ventilator. It is suitable for a wide range of life support and critical care applications. The ventilator normally operates from external DC power. When the ventilator is portable, a Removable Battery Pack powers the unit. |  |  | September 12, 2011 | 2011 | CareFusion 203, Inc. | CareFusion has identified potential risks associated with the EnVe ventilator manufactured between December 2010 and May 2011. There are 3 issues that can result in ventilation delivery to the patient being interrupted. CareFusion will be contacting facilities to coordinate implementation of the necessary corrective actions. | Device Design | The firm, CareFusion, sent an "URGENT PRODUCT RECALL" letter dated September 12, 2011 to its customers.The letter identified affected product, described the 3 issues providing the description of the event, potential risk, and provided required action for Users. The letter also stated that CareFusion would be contacting the facility to coordinate the corrective action.If you have any questions and/or need support, please contact CareFusion Technical Support at 1-800-554-8933 from Mon. - Fri. 8:00am -5:00pm Pacific Time. |  | 128 | 128 units | Nationwide distribution: USA including states of: CA, CO, FL, IL, MI, MN, MT, ND, NJ, NM, NY, OK, PA, TX, UT, WI, and WV. | Terminated | November 22, 2012 | 437 | 128
Z-0131-2012 | 59147 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SYNCHRON CX7 Delta Clinical System Part #: 468205 |  |  | August 23, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed software limitations which affect only manually-loaded reagent cartridges on the CX4/5/7 CE, CX4/5/7 Delta and CX9 ALX Clinical Systems. | Software design | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated September 20, 2010 with an attached Fax Back Response Form (via US Postal Service for US Customers) to all customers who purchased the SynchronÂż CX4/5/7 CE, CX4/5/7 Delta and CX9 ALX Clinical Systems. The letter provides the customers wtih an explanation of the problem identified and actions to be taken. Customers were instructed to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded the affected product listed above to another laboratory, they were asked to provide a copy of this letter to them. Questions regarding this Product Corrective Action letter, were directed to Customer Support Center (Call Center) on the Beckman Coulter website athttp://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. Outside the United States and Canada, they were directed to their local Beckman Coulter Representative. |  | 1070 | 1070 systems total in US | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | June 28, 2012 | 675 | 1070
Z-0122-2012 | 60041 | No | LRR |  |  | General &amp; Plastic Surgery | Enforcement Discretion | Mountain First Aid Kit, Part Number 4118; UPC 0-91966-04118-3, contains 88 Pieces. |  |  | July 07, 2011 | 2011 | Lifeline First Aid LLC | The first aid kits are recalled because the kits contain Iodine Prep Pad single use wipes which were recalled by Triad. | Material/Component Contamination | Lifeline First Aid, LLC sent a " PRODUCT RECALL NOTICE-LIFELINE FIRST AID KITS WHICH INCLUDE H &amp; P [TRIAD] PVP IODINE WIPES" letter dated July 7, 2011 to all affected customers. The letter states that Lifeline First Aid, LLC has become aware of a recall of Povidone Iodine Wipes manufactured by H&P; Industries Inc. (TRIAD GROUP) and marketed under various brand names, including Triad and Swift. The recall of the Povidone Iodine Wipe is due to potential microbial contamination of the product. The letter instructs customers to collect and quarantine all affected products and to complete and return an attached Customer Response Form via fax to (503) 783-3721.Affected products will be replaced without charge and acceptable Lifeline first aid kits will be re-worked and returned using Lifeline's standard RGA process. Customers may call Lifeline customer service at 1 (877) 377-8243 for further instructions. |  | 498 | 498 kits | Nationwide Distribution | Terminated | March 29, 2012 | 266 | 26150
Z-0118-2012 | 60064 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Verardius Image Intensified X-ray |  |  | October 04, 2011 | 2011 | Philips Healthcare Inc. | Front plate broke in to two pieces causing the C-Arm to suspend on the single central bolt | Nonconforming Material/Component | Philips Healthcare Inc., sent an "URGENT - Field Safety Notice" dated September 30, 2011 to all affected customers. The notice identifies the following: what the problem is and under what circumstances it can occur; the actions that should be taken by the customer/user in order to prevent risks for patients or users and the actions planned by Philips Medical Systems to correct the problem. The notice contains important information for the continued safe and proper use of the equipment and customers were advised to retain a copy with the equipment Instruction for Use. The notice stated that Philips will replace the front plates for systems on the Unit Affected List by the end of June 2012. The notice instructed customers to contact Philips Healthcare representative for further information or support. |  | 120 | 120 units in the US | Worldwide Distribution - USA (nationwide) and the countries of: Australia, Austria, Belgium, Bharain, Canada, Croatia, Denmark, France, Germany, Greece, Italy, Japan, Netherlands, Norway, Portugal, Russian Fed., Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, and United Kingdom. | Terminated | January 12, 2017 | 1927 | 266
Z-0113-2012 | 60025 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Radio used in the Precess MRI Patient Monitoring System. 3160 (Precess) and HE97B, 94044, 94045 (radio and radio kits). |  |  | September 23, 2011 | 2011 | Invivo Corporation | Philips Invivo has determined that it is possible for the system's alarm limits and other device settings to spontaneously change without user intervention. This problem is associated with intermittent and frequent loss of communication between the wireless Processing Unit and the Display Controller Unit. | Device Design | A Field Safety Notice entitled Urgent - Medical Device Correction letter, dated September 2011, with attached Field Safety Notice for customers were sent to distributors. The letter described the reason for the recall and the intention of the Field Safety Notice. Also, the hazards involved and instructions on how to identify units that are affected by the recall were provided . The actions to be taken by the user and the actions planned by the Philips Invivo were explained. Users were to follow the instructions provided on how to safely use the device until the unit can be serviced by Philips Invivo. For further information or support concerning this issue, customers are to contact Philips Invivo Customer Support at (877)-468-4861 or their local Philips Invivo representative. |  | 875 | 875 | Worldwide Distribution -- Australia, Austria, Brazil, Canada, China, France, Germany, Malaysia, Norway, Portugal, Spain, Sweden, The Netherlands, United Arab Emirates, United Kingdom, and US. | Terminated | August 07, 2013 | 684 | 875
Z-0073-2012 | 59536 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mizuho OSI Modular Table System; |  |  | July 29, 2011 | 2011 | Mizuho Orthopedic Systems Inc | Reports of injury related incidents while using the Mizuho OSI Modular Table System. | Use error | Mizuho Orthopedic Systems Inc sent a "Field Advisory Notice (Recall)" letters dated July 29, 2011 to all customers using the Modular Table System via Federal Express. The Field Advisory Notice reference the devices involved and explain that there is no need to return any product or to discontinue the use. The letter provides WARNINGS and RECOMMENDATIONS for all users to consider when using the devices and request that the response form be completed as instructed in the advisory communication. The letters also recommend that awareness, training, inspection prior to use, and proper maintenance be established and practiced. The letter recommend that the Field Advisory Notices be posted in prominent areas and communicated to all staff and potential users of this device. Additional support and training can be obtained by contacting Mizuho OSI's Customer Resource Group at 1-800-777-4674. For further questions concerning this notification, contact the Director of Customer Resource Group for assistance (1-800-777-4674, outside the USA 00+1-510-476-8199).Revised letters were approved by CDRH and sent out in January 2012. |  | 15011 | 15,011 units | Worldwide Distribution - Nationwide USA | Terminated | August 27, 2013 | 760 | 15011
Z-0138-2012 | 59866 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Stryker Communications, 1410 Lakeside Parkway, # 100, Flower Mound, TX 75208 |  |  | August 23, 2011 | 2011 | Stryker Communications Corp | Installation records for the SwitchPoint Infinity, Booms, Lights, and Flat Panel systems are incomplete or inadequate. | Process control | Stryker sent an URGENT FIELD SAFETY NOTICE: RA2010-245 dated August 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the attached Customer Response Form within 5 business days to Stryker. A Stryker representative would contact them at their convenience to schedule a re-inspection for all of the units identified and still in use. In the event the inspected equipment does not meet Stryker's established specifications, their representative will visit their facility and hopefully resolve the issue. If a subsequent visit is required the date will be scheduled in full consideration of their services.For questions regarding this recall call 972-410-7310. |  |  |  | Worldwide Distribution including Puerto Rico, Brazil, Ecuador, Venezuela, Costa Rico, Bolivia, Peru, Mexico, China, India, Korea, Philippines, Singapore, Malaysia, Taiwan, Middle East, Russia, Africa, Poland, South Africa, Australia, New Zealand, Canada, Japan, Denmark, Finland, Norway, Sweden, Belgium, The Netherlands, France, Germany, Austria, Italy, Portugal, Spain, Romania, Switzerland, and United Kingdom. | Terminated | August 20, 2012 | 363
Z-0135-2012 | 60175 | No | FFX | Gastrointestinal motility monitoring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AR ManoShield, MSS-3599, 10 pack. |  |  | October 12, 2011 | 2011 | Sierra Scientific Instruments Inc | The recall was initiated after Given Imaging received a test report during a Canadian licensing process where the AR ManoShield exhibited pinholes and did not pass the test, the Field Safety Corrective Action (FSCA) Committee was convened. As a result of the data discussed during this meeting, a decision was made to voluntarily recall this product, and cease its further commercialization or distr | Nonconforming Material/Component | Given Imaging sent an "Urgent Product Recall 1st Notice" letter with an attached Device Recall Confirmation Form (via Fed Ex with signature for tracking purposes) to all customers who purchased the AR ManoShield, part numbers MSS-3599. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were instructed to locate the affected product and removed from use immediately. The letter stated that if the affected product have been further distributed, forward this letter immediately to all customers who received the affected product. Customers were asked to complete and return the enclosed confirmation form as per the instructions. Customers with any questions were instructed to contact Karla Guerrero at (310) 641-8492 ext. 2138 or email at karla.guerrero@givenimaging.com. |  | 292 | 292 units | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Belgium, Canada, Chile, China, Columbia, Ecuador, France, Germany, Israel, Mexico, Poland, Portugal, Sweden, and Switzerland. | Terminated | August 30, 2012 | 323 | 292
Z-0133-2012 | 59521 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, InnovaÂż 3100IQ (Optima Edition) Fluoroscopic Vascular X-ray System. |  |  | August 24, 2011 | 2011 | GE Healthcare, LLC | A potential weakness on the monitor suspension end-stop device may impact user safety. The end-stops may break if the monitor-suspension is repeatedly pushed with excessive force on the end rails. | Nonconforming Material/Component | A GE Healthcare "Urgent Medical Device Correction" letter dated August 22, 2011 was sent out 8/24/11. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology &amp; Radiologists/Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Users were told to operate the suspension without excessive force. If one of the end-stops breaks, then the system should not be used and their local GE Healthcare Service Representative should be contacted. GE will contact customers to replace all end-stops at no cost to the customer. Questions or concerns should be directed to 800-437-1171 in the US and 0120-055-919 in Japan. For other countries, customers can contact their local GE Healthcare Service Representative. |  | 15 | 15 | Worldwide Distribution -- USA, including the states of TX &amp; NC and the countries of TURKEY, POLAND, MEXICO, CHINA, CAMBODIA, BRAZIL, BOLIVIA, and SAUDI ARABIA. | Terminated | May 29, 2012 | 279 | 15
Z-0134-2012 | 60168 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Flo-Gard 6301 Dual Channel Volumetric Infusion Pump; Baxter Healthcare Corporation, Deerfield, IL 60015; Product Code # 2M8064. |  |  | August 26, 2011 | 2011 | Baxter Healthcare Corp. | The Flo-Gard infusion Pump was released to the customer with OOS values for the Air Sensor Calibrations with the Air in Tubing test. | Employee error | A Baxter Healthcare technician initially notified the sole customer on August 26, 2011 of the defective product and requested that the product be quarantined. On September 14, 2011 a technician visited the site to correct the problem observed on the Flo-Gard Infusion Pump. All required tests were performed and the results were found acceptable. |  | 1 | 1 unit | Distributed only in Puerto Rico. | Terminated | October 31, 2011 | 66 | 1
Z-0141-2012 | 59909 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Presource Custom Sterile Surgical and Procedure Kits; surgical trays; Cardinal Health, McGaw Park, IL 60085; The kit catalog numbers listed in the code information section below contain one or more of the following listed components for which Cardinal Health has been unable to confirm the manufacturer's compliance with section 502(o) of the FD&C; Act: |  |  | September 03, 2011 | 2011 | Cardinal Health | Some of the components contained in the Presource kits are being recalled because Cardinal Health has determined that the component manufacturer(s) does not have all necessary FDA registrations or filings needed for this product. | Component design/selection | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated September 2, 2011 to all affected consignees via UPS next day air, informing them that their Presource kits listed in the attached spreadsheet contain various components listed on the attachment, which are being recalled because the component manufacturer(s) do not have all the necessary FDA registrations or filings needed for these components. Enclosed with the letter were recall labels to affix to the front side of each affected Presource kit in their inventory. The labels have an ID code which coincides with the code on the component attachment to the letter. If the consignee does not wish to accept overlabeled product or to overlabel the product in their possession, they were asked to contact Presource Sales Operations at 800-766-0706 or their Sales Representative for further instructions. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules labeled. For component credit, the consignees were instructed to contact the following: Hospitals - Presource Sales Operations at 800-766-0706, Distributors - Distributor Management Group at 800-635-6021, and Federal Government facilities - Government Service at 800-444-1186.For questions regarding this recall call 847-887-6412. |  |  |  | Nationwide Distribution | Open, Classified
Z-0139-2012 | 60006 | [6 empty] | Picis Anesthesia Manager, PACU Manager,Critical Care Manager-8.2 Service Pack 4 or higher |  |  | August 12, 2011 | 2011 | Picis Inc. | Periop/Critical Care Manager EHR Software Application displayed one patient's demographic and health information in place of another patient that was in current view. | Software design | Picis Inc. sent a "Safety Notification letter" dated August 12, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed until this service pack is applied, we recommend: * For patients who have been transferred from other facilities within the same VISN, you should discontinue the use of these Picis applications. Manually check the patient's name with that displayed in the Picis patient blue banner to ensure you catch all of these cases. * For all other patients, you may safely continue the use of these Picis applications.For further questions please call (781) 557-3000. |  | 10 | 10 | Nationwide Distribution -- AZ, DE, FL, IL, NY, OH, PA, WI and WV. | Terminated | March 23, 2012 | 224 | 10
Z-0142-2012 | 59582 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 including all Catalog Numbers to include: Catalog # 500AVHCT with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe. |  |  | September 23, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo CVS identified a defective circuit board, which could potentially cause the CDI monitor to not function after a power loss. Once power is lost, the monitor will not power up again, even if AC or battery power are restored. | Component design/selection | Terumo Cardiovascular Systems Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" notice dated September 23, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter provides information on what to do in the event of a failure of the device. Additionally, a Customer Response Form was attached to the letter for customers to complete and return. Terumo CVS will schedule the return of affected products for inspection and replacement of defective parts. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this notice. |  | 432 | 432 | Worldwide Distribution-USA (nationwide) and the countries of Canada, Chile, Japan, Philippines, and Thailand. | Terminated | April 11, 2013 | 566 | 432
Z-0143-2012 | 55537 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Bioplate Titanium Fixation System, Sterile Kit. |  |  | March 31, 2010 | 2010 | Bioplate Inc | There is potential defect in seal of the sterile barrier containing the device. | Process design | BIOPLATE sent an URGENT - PRODUCT RECALL NOTIFICATION letter dated March 18, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to return any recalled product they have at their siteand if distributed further they should forward the information contained in the letter to all customers. If a device from one of the lots noted above has already been used, Bioplate would like to be advised of that use.They asked that consignees complete the enclosed Fax form and transmit the requested information to them as soon as possible at (310) 815-2126.Returned recalled product was to be sent to: Bioplate3643 Lenawee AvenueLos Angeles, CA 900163643 Lenawee Avenue, Los Angeles, CA 90016For any questions regarding this recall call (310) 815-2100. |  | 18246 | 16,020 Individual Kits; 2,226 Boxes. | Worldwide Distribution - USA (nationwide) and the countries of Canada, Europe, Japan, Taiwan, Middle East and Australia. | Terminated | April 04, 2012 | 735 | 18246
Z-0144-2012 | 60057 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂż Ventralex" ST Medium Circle with Strap Hernia Patch Diameter 6.4cm /2.5", Catalog Number: 5950008. |  |  | October 04, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: BardÂż Ventralex" ST Medium Circle was packaged in a pre-printed carton identified for a BardÂżVentralex" ST Large Circle. | Packaging process control | Davol Inc. issued an Urgent; Medical Device Recall letter on 10/4/11 for the mislabeled Bard Ventralex ST Medium Circle. The letter identified the affected product and explained the reason for recall. Customers are to immediately examine their inventory and quarantine any affected product. In addition, if any product was further distributed then those customers should be notified of the recall as well. Accounts are to contact Davol Customer Service at 1-800-556-6275 for instructions on how to return product to Davol. The attached recall effectiveness check form should be completed and returned. Questions should be directed towards Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Services &amp; Support Department at 1-800-562-0027. |  | 96 | 96 units | Nationwide Distribution | Terminated | September 20, 2013 | 717 | 96
Z-0145-2012 | 60046 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | MERZ AESTHETICS, RadiesseÂż Volume Advantage 1.5 CC, Injectable Implant, Part Number 8071M0H1 and 8071M0K1, Sterile, Rx ONLY. |  |  | August 18, 2011 | 2011 | Merz Aesthetics, Inc. | There is a potential leakage of product that may occur with Merz Aesthetics 1.5cc syringes of Radiesse Dermal Filler. When using the syringe, Health Care Professionals should be alerted to the potential for product that my leak pass the plunger at the proximal end of the syring. The defect may result in infection requiring medical intervention to prevent impairment or damage. To date Merz has n | Nonconforming Material/Component | On 8/18/11, consignees were sent a Merz Aesthetics "Urgent Medical Device Recall" letter dated August 15, 2011. The letter was addressed to "Dear Health Care Professional". The letter described the product and the problem and listed the details regarding this issue. |  | 311582 | 311,582 | Worldwide distribution including USA, Puerto Rico, Dominican Republic, Hong Kong, Israel, Netherlands, Germany, South Korea, and Canada. | Terminated | May 01, 2012 | 257 | 311582
Z-0146-2012 | 58162 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Labeling for RIO Robot Unit is comprised of three main components: 20399 RIO SURGICAL ARM, 201251 RIO GUIDANCE MODULE, 200294 RIO CAMERA STAND ASSEMBLY***PN 203999 SN ROB 125 2010-12 V 100/120/230 A 9.6/8.0/4.2 Hz 50/60 Class I Equipment. Conforms to IEC 60601-1/A2: 1995, EN 60601-1/A2: 1995 UL 60601-1: 2003, CAN/CSA-C22.2 No. 601.1-M90.***Manufactured in USA***MAKO SURGICAL CORP. 2555 DAVIE ROAD, FT. LAUDERDALE, FL 33317. |  |  | January 20, 2011 | 2011 | Mako Surgical Corporation | MAKO Surgical Corp. is recalling their RIO Robotic Arm Interactive Orthopedic System (RIO) due to software issue that exist that could potentially result in a bone resection. No adverse events reported. | Software change control | The firm, MAKO Surgical Corp., sent an "RECALL NOTICE 2.4. SITES" letter dated February 1, 2011 to its customers. The letter describes the product, problem, actions to be taken which includes a workaround and an updated user guide identifying the appropriate workflow. The customers were instructed to follow the user guide and workflow and to complete and return the enclosed confirmation memo to their MAKOplasty Specialist or return it in the pre-addressed, pre-paid envelope provided. MAKOplasty included with the letter new steps in the surgical workflow in the use of the RIO system to ensure error value accuracy; worked on a software update that will not allow the user to proceed with error values that do not support the required accuracy-this software was to be supplied to customers in February 2011, and trained MAKOplasty Specialist will make available written updates to the RIO use instructions as well.If you have any questions or concerns, please do not hesitate to speak with your MAKOplasty Specialist or contact the Manager of Customer Relations at 954.927.4022 x447. |  | 70 | 70 | Worldwide distribution: USA including states of: AZ, CA, CO, FL, GA, IA, IL, LA, MI, MO, MS, NC, NJ, NV, NY, OK, OH, OR, PA, RI, TX, UT, VA, WA, WI, and WV; and countries of: Scotland, Italy and Korea. | Terminated | December 13, 2011 | 327 | 70
Z-0014-2012 | 59949 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Dynamic Multileaf Collimator (DMLC) |  |  | May 25, 2011 | 2011 | Elekta, Inc. | It is possible to load different patients in the record and verify (R&V;) system and in the multilead control software (MCS). | Device Design | Elekta, Inc. sent an "IMPORTANT NOTICE" letter dated May 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter recommends several precautions customers should use until Elekta finds a solution. Customers will be contacted when a solution is found. |  | 27 | 27 | Nationwide Distribution-USA including the state of CT, FL, GA, IA, MN, MO, NJ, NY, ND, OH, OK, PA, RI, TX, VA , and WA. | Terminated | January 14, 2014 | 965 | 27
Z-0147-2012 | 60240 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Medrad(R) Sterile Disposable Stellant Syringe Kit, Catalog # SDS-CTP-SPK, Lot # 112865. |  |  | October 24, 2011 | 2011 | Medrad Inc | Packaging may exhibit open seals - sterility compromised | Packaging | Medrad sent a " Urgent Medical Device Recall letter" dated October 24, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed if they have any of the affected product lot, please contact MEDRAD customer support to make arrangements for the return of the product. Call 1 -800-633-7231, select option 5, then select option 1. Indicate that you are calling about the affected syringe kit recall.Complete and sign the attached response form and fax it back to MEDRAD at (412) 406-0942.Please complete and fax this form even if you do not have any of the referenced product so we can maintain proper accountability for all syring kits that have been shipped. |  | 17380 | 17,380 | Nationwide Distribution -- including the States of : AL, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, VS, WA, WV, and WI. | Terminated |  |  | 17380
Z-0149-2012 | 60034 | Yes | OAE |  |  | Cardiovascular | PMA | Medtronic Arctic Front Catheter, Sterile EO, Medtronic CryoCath LP, Kirkland, Quebec, H9H 5 H3, Canada, Model 2AF282 and 2AF231 ( not distributed in US) |  |  | August 29, 2011 | 2011 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic CryoCath has identified specific lots of Arctic Front CryoAblation Catheters that were distributedwith incorrect programming that may result in false expired catheter notifications. Impacted lots are 21571,25375, 25376, 25377, and 25381 of the Arctic Front Catheter Model Number 2AF282. Medtronic CryoCathis communicating this information to the FDA.On affected catheters, the incor | Employee error | Medtronic sent an "Important Device Recall" letter dated August 2011, hand delivered by Medtronic Field Representatives on August 29, 2011 to all affected customers. The letter described the product affected, the problem, and the actions needed to be taken. The customers were instructed to return all unused catheters from affected lots for credit and replacement. The letter states that Medtronic AF Solution representative will assist customers with the return of affected product. For further questions contact your local Medtronic AF Solutions representative. |  | 263 | 263 | Worldwide Distribution - USA (nationwide) including the states of: AK, AZ, CA, CO, FL, GA, IL, IN, KY, LA, MA, MN, MS, MO, NE, NY, NC, OH, OK, PA, SC, SD, TN, TX, VA, WA, and WI. and the countries of: Austria, United Kingdom. | Terminated | June 07, 2012 | 283 | 263
Z-0150-2012 | 59991 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RSP Humeral Socket Reamer (Size 32 SML): |  |  | September 28, 2011 | 2011 | Encore Medical, Lp | DJO Surgical has determined that the alignment pins cannot be consistently cleaned using the technique provided in the cleaning/sterilization section of the instrument IFU. | Device Design | DJO Surgical sent an URGENT FIELD SAFETY NOTICE letter dated September 28, 2011, to all affected customers. The letter identified the product. the problem, and the action to be taken by the customer. Customers were instructed to return the RSP Humeral Socket Reamers and RSP Baseplate Rim Planers. The RSP Humeral Socket Reamers will be reworked and returned to the customer. The Baseplast Rim Planer will be replaced when a suitable design is available. During the interim until instruments are returned, the use of FMP acetabular reamers size 42mm, 43mm, 45mm (DJO PIN 803-05-226, 227, 200) for the humeral preparation was validated to be effective to prepare the humerus for the implant.Customers were instructed to immediately notify the hospitals and surgeons using these devices and document it on the Field Safety Notice Response Form. For questions regarding this recall call 512-832-9500. |  | 399 | 399 Reamers | Worldwide Distribution - USA including AR, AZ, CA, CO, FL,GA, ID, IN, IL, LA, MA, MD, ME, MI, MO, NJ, NM, NY, OH, OK, PA, PR, SC, TN, TX, UT, VA, WA, and WI and Internationally toGermany, South America, Italy, Saudi Arabia, United Kingdom, and Canada. | Terminated | March 29, 2013 | 548 | 532
Z-0158-2012 | 60254 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | "***BIOMET 3i***4555 Riverside Drive Palm Beach Gardens FL 33410 USA***REF AC4425***Conical 25 ANGLED ABUTMENT***4.1mm(D) X 4mm(H)***Abutment; Pilier: Abutment; Abutment; Pilastro; Pilar***RX Only***Sterile***BIOMET 3i Dental Iberica S.L. WTC Almeda Park. Ed. 1. Planta 1'***". |  |  | September 28, 2011 | 2011 | Biomet 3i, LLC | On 09/28/2011 Biomet 3i, Palm Beach Gardens, FL initiated a recall of their Conical Angled Abutment, Model # AC4425 Lot 887628-5 and their Standard Conical Abutment Cylinder Model # SWCA62 Lot 848478. The packaging for the referenced products may not have been completely sealed prior to shipment. | Employee error | Biomet 3i, LLC sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 28, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were asked to inspect their inventory and return any affected product for replacement.A Product Recall Response Form was attached for customers to complete and return via fax to 561-514-6316. Customers are directed to call 1-800-342-5454 for questions regarding this recall. |  | 122 | 122 | Worldwide Distribution-USA (nationwide) and the countries of Chile, Italy, Japan, Korea, Spain, and Sweden. | Terminated | March 09, 2012 | 163 | 130
Z-0152-2012 | 60100 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | ProScreen PSCup-5M Waived, Catalog Number: PSCup-5M Waived |  |  | June 07, 2011 | 2011 | Ameditech Inc | The recall was initiated because Ameditech, Inc. received a complaint from US Diagnostics (USD) that First Check Drug Screen Cups were packaged in the ProScreen Cup boxes. Some of US Diagnostics' customers received ProScreen 5-panel cups that contain First Check THC cups. The recalled products are mislabeled. If the USD customer receives the FirstCheck tes, there will be no instructions for use | Packaging process control | The recall communication was initiated on June 9, 2011 with Ameditech Inc. forwarding an Urgent Medical Device Recall notification with an attached Verification Form (via phone and written notification) to the affected distributor. The letter provides the distributor with an explanation of the problem identified and actions to be taken. The distributor was required to: cease distribution of the affected lot and provide remaining inventory count to Ameditech; destroy any remaining inventory of these lots and document destruction per site requirements; and ensure that all user of the device have received a copy of the Urgent Medical Device Recall. In addition, the distributor was instructed to complete and return the enclosed Verification Form with in 10 days.For any questions about the information contained in this notice, contact Ameditech at amd.complaint@alere.com or call 858-535-1968. |  | 2000 | 2000 test cups (800 kits) | Nationwide Distribution (USA) including the state of Alabama. | Terminated | January 18, 2012 | 225 | 2000
Z-0154-2012 | 60114 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | 37% Etching Gel-1ml syringe 20 pk, part # T06A |  |  | March 31, 2011 | 2011 | Kerr/pentron Dba Kerr Corporation And Pentron Clinical | Pentron Clinical is voluntarily recalling all lots of products labeled as 37% Phosphoric Acid Etching Gel produced since April 2007 because the product continues to extrude out of the syringe even after pressure is no longer applied, and it contains a higher percentage of phosphoric acid than that indicated on the label. | Nonconforming Material/Component | The firm, Pentron Clinical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated May 31, 2011, with attached Return Form (via USPS 1st class mail) to all their consignees/customers. The letter describes the product, problem and action to be taken.All consignees/customers were instructed to complete the Recall Return Form via fax to: 1-877-677-8844 and to return any affected product in their inventory. Affected products will have expiration dates between April 2011 and April 2015.Consignees/customers with questions can contact Pentron Clinical Customer Service at (800) 551-0283 directly to handle the arrangements of a quick return and replacement. |  | 41744 | 41,744 units | Worldwide distribution: USA (nationwide) and countries of: Armenia, Austria, Australia, Antigua &amp; Barbuda, Barbados, Belarus, Bulgaria, Canada, Czech Republic, Croatia, Cyprus, Ecuador, El Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kenya, Latvia, Lebanon, Malaysia, Mexico, Morocco, Netherlands, Poland, Romania, Russia, St. Vincent &amp; Grenadine, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Taiwan, Thailand, Turkey, UAE, Ukraine, UK, and Vietnam. | Terminated | August 09, 2012 | 497 | 41744
Z-0160-2012 | 57234 | No | EKM | Gutta percha. | Dental | Dental | 510(K) Exempt | Microseal Master Cone Part Number: 815-9033 |  |  | September 09, 2010 | 2010 | Ormco Corporation | The recall was initiated because SybronEndo has confirmed that the Microseal Master Cones has been mislabeled. | Error in labeling | Sybron Dental Specialties, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated September 20, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return the affected lots of product.. Products, would be replaced at no charge or credited to their account. Customers were instructed to contact SybronEndo Customer Care at 1-800-346-3636 directly to handle the arrangements of a quick return and replacement. The affected product should be returned to SybronEndo at the following address. Please label returned product "RECALLED PRODUCT Attention: Customer Returns"SybronEndo 1332 South Lone Hik AvenueGlendora, CA 91740ADDITIONALLY, THEY KINDLY REQUEST CONSIGNEE COOPERATION IN COMPLETING AND FAXING BACK THE ENCLOSED ACKNOWLEDGEMENT/RETURN FORM. THIS FORM MUST BE COMPLETED AND RETURNED WHETHER OR NOT YOU HAVE ANY PRODUCT TO RETURN.Customers were also asked to identify and recover the affected product lots listed above that may have been shipped to their customers.For questions regarding this recall call 909-962-5600. |  | 200 | 200 units | Worldwide Distribution - USA including AR, CA, FL, PA, GA, IL, TX, WA, NY, IN, AL &amp; MT; Internationally to Canada, CIS (Commonwealth of Independent States, Israel, United Kingdom, Sweden, France, Germany, and Italy. | Terminated | November 21, 2011 | 438 | 200
Z-0161-2012 | 60003 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Spine T-PAL Spacer System. Transforaminal posterior atraumatic lumbar spacer system. |  |  | September 23, 2011 | 2011 | Synthes USA (HQ), Inc. | LABELING CORRECTION for Medical Device - previous labeling and promotional materials associated with the T-PAL Spacer system reference the use of T-PAL spacer with allograft material. T-PAL Spacer is currently indicated for use with autograft material. All accounts with record of receiving T-PAL Implant will receive notification of the labeling correction and a copy of the updated brochure. All | Error in labeling | SYNTHES sent an URGENT MEDICAL DEVICE LABELING CORRECTION letter dated September 23, 2011, to all affected customers. The labeling has been revised. Updated information packets have been sent to all accounts which included the final labeling and communications to accounts and sales force. Customers were instructed to review the updated supplemental labeling provided in the packageCustomers were instructed to attach the supplementary technique page to an T-PAL Spacer Brochure that may be in their facility. They can access an electronic version of the Brochure by accesssng Product Literature at the following link:http://us.synthes.com. The new version can be identified by code J10120 located at the bottom right corner of the back cover.To confirm receipt and understanding of the information, customers were requested to complete the Verification Section of the letter and return it to Synthes by either of the following methods:Mail provided postage paid envelopeFax 610-251-9005Scan/e-mail to FieldAction@synthes.comFor any questions call 1-800-420-7025 x5403 or contact their Synthes Spine Sales Consultant. |  |  | N/A - Labeling Correction | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, HI, KS, KY, LA, MD, MA, MI, MN, MS, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. and Puerto Rico and Washington, D.C; Internationally to Canada and Japan | Terminated | December 18, 2014 | 1182
Z-0162-2012 | 60094 | No | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | ARCHITECT Cyclosporine; an in-vitro diagnostic chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of cyclosporine in human whole blood Distributed by Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064 USA; list 1L75-25. |  |  | October 07, 2011 | 2011 | Abbott Laboratories | The ARCHITECT Cyclosporine assay has the potential to generate falsely elevated or depressed results with patient samples due to precision exceeding the Package Insert claim of &lt;15% total CV. Calibrators and Controls also have the potential to be impacted by the assay imprecision. | Nonconforming Material/Component | Abbott Laboratories sent a "PRODUCT RECALL" letter dated October 7, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to only run ARCHITECT Cyclosporine using Reaction Vessel (RV) lots manufactured with the same resin as follows: a) RV lots beginning with lot 06083P100 and higher can be used together or, b) RVs lots lower than 06083P100 can be used together. Any RV lot number that does not begin with a zero (e.g. 90032P100) should also be considered lower than 06083P100. Further instructions were included for the accounts to follow when running ARCHITECT Cyclosporine assays. A Customer Reply form was provided to customers to complete and return via fax at 1-800-777-0051. Customers can contact Abbott Customer Service at 1-877-4ABBOTT for questions regarding this notice. A second letter was sent to the accounts on 11/23/11 to simplify the instructions for identifying the Architect Reaction Vessel lots that can be used together for use with the Architect Cyclosporine assay. |  | 8913 | 8,913 kits | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Bahrain, Bangladesh, Barbados, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Brunei, Canada, Chile, China, Columbia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Vietnam. | Terminated | November 07, 2012 | 397 | 8913
Z-0165-2012 | 60241 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | Abbott brand IntraLase Femtosecond Laser System, Model 2; Catalog Numbers: 20003D/J/K, R20003D/J/K, 20005D/J/K; Product is manufactured and distributed by AMO Manufacturing USA, LLC, Milpitas, CA |  |  | October 12, 2011 | 2011 | AMO Manufacturing USA, LLC | The firm has identified a discrepancy in cut depth that may affect deep corneal incisions (e.g. 300 Âżm or deeper) created using the Laser System. | Process design | Abbott Medical Optics Inc. (AMO) issued an "ADVISORY NOTICE" letter dated October 12, 2011 to all affected customers. The letter identifies the affected products, the problem, and the actions needed to be taken. Customers were instructed to exercise caution when making cuts that approach the corneal endothelium. The letter includes a 1-page Operator's Manual addendum that clarifies the recommendations. The letter states that AMO will conduct a field correction to adjust the EOF on the affected FS Laser Systems. The letter includes a response form to be completed and returned as per the instructions. Customers were instructed to notify all appropriate staff/personnel of the Advisory Notice. For any further questions regarding the information or recommendations, contact an AMO representative at 1-800-266-3375 |  | 1226 | 1226 total laser systems | Worldwide Distribution - USA (nationwide) | Terminated | May 01, 2013 | 567 | 1226
Z-0164-2012 | 60143 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | BioPro Femoral Head 28 mm +18, Item 19010. |  |  | September 21, 2011 | 2011 | Biopro, Inc. | BioPro Femoral Head Outer Packaging was found lacking proper seal ,compromising sterility. | Packaging | Biopro, Inc notified all affected customers via telephone on September 21, 2011 with instructions to quarantine and return affected product to the manufacturer. The firm sent a follow up "RECALL NOTICE" dated September 29, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter gives instructions on how to return affected product. Contact Customer Service at 1-800-252-7707 for questions regarding this recall. |  | 8 | 8 units | USA (nationwide) including the states of FL, MI, and TX. | Terminated | February 15, 2013 | 513 | 12
Z-0174-2012 | 60255 | No | KZH | Syringe needle introducer. | General Hospital | General Hospital | 510(k) | 1.0mL 28 G x 1/2" blister packaged insulin syringe Catalog Number 329424; Lot Number 0039920 |  |  | June 15, 2011 | 2011 | Becton Dickinson &amp; Company | Experiment labels may be on shelf cartons of 1.0ml 28G x 1/2 blister packaged insulin syringes. | Labeling mix-ups | The firm, BD, sent an "PRODUCT RECALL NOTIFICATION" letter and response form dated June 15, 2011 to their consignee/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately review their inventory for the affected product, remove the product from inventory, quarantine the product and return it to BD; if the product was further distributed, they should identify their customers and notify them at once, and complete and return the attached Tracking/Verification Form via fax to: 201-847-4853.If you have any questions, contact Customer Service at 1-800-237-2762, Monday -Thursday, 8:00am-6:30pm and on Friday, 8:00-5:30pm Eastern Standard Time. |  | 10230 | 10,230 shelf cartons (1,0230,000 syringes) | Nationwide Distribution. | Terminated | July 27, 2015 | 1503 | 10230
Z-0173-2012 | 60183 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter System 1000 Tina Single Patient Hemodialysis Delivery System; Baxter Healthcare Corporation, Renal Division, McGaw Park, IL 60085; product code S1000L3P. |  |  | September 19, 2011 | 2011 | Baxter Healthcare Renal Div | The Tina System was released to the customer with OOS values for the Pressure Displays Accuracy and Leakage Rate Test. | Employee error | Baxter Healthcare Corporation sent a service technician to the lone customer site on September 19, 2011 to correct the affected product. The technician executed all tests as established in the Service Manual, with acceptable results. |  | 1 | 1 unit | Distributed in Puerto Rico only. | Terminated | November 07, 2011 | 49 | 1
Z-0169-2012 | 60069 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Presource Custom Sterile Surgical and Procedure Kits; general surgery trays; Cardinal Health, McGaw Park, IL 60085; The kit catalog numbers listed in the code information section below contain one or more of the following listed components for which Cardinal Health has been unable to confirm the manufacturer's compliance with section 502(o) of the FD&C; Act: |  |  | September 30, 2011 | 2011 | Cardinal Health | Some of the components contained in the Presource kits are being recalled because Cardinal Health has determined that the component manufacturers do not have all necessary FDA registrations or filings needed for these components. | Component design/selection | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated September 30, 2011 to the affected wholesale consignees via UPS next day air, informing them that their Presource kits listed in the attached spreadsheet contain various components listed on the attachment, which are being recalled because the component manufacturer(s) do not have all the necessary FDA registrations or filings needed for these components. Enclosed with the letter were recall labels to affix to the front side of each affected Presource kit and shipping carton in their inventory. The labels have an ID code which coincides with the code on the component attachment to the letter. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules labeled. For component credit, the consignees were instructed to contact Distributor Management Group at 800-635-6021. This wholesale recall was expanded to include additional kits via letter dated October 5, 2011, with the same instructions as listed above.For questions regarding this recall call 800-292-9332. |  |  |  | Nationwide Distribution including California, Florida, Illinois, Kentucky, Minnesota, Missouri, Ohio, Tennessee, Washington, West Virginia and Wisconsin | Open, Classified
Z-0171-2012 | 58766 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | ***REF IOSS511***OSSEOTITE Certain Implant 5 x 11.5mm***LOT 910695***Sterile using Radiation***Rx Only***BIOMET 3i Dental Iberica S.L.WTC Almeda Park, Ed, 1, Planta 1 Pl. de la Pau s/n 08940 Cornella de Llobregat (Barcelona) Spain***Biomet 3i Dental Implants are intended for surgical placement in the upper or lower jaw to provide a means for prosthetic attachment. The Thread form titanium screws (cylindrical dental implants) designed to be placed into mandibular or maxillary jaws in order to affix a dental prosthesis (edentulous or partially edentulous segments) that replaces 1 or more missing teeth. |  |  | November 18, 2010 | 2010 | Biomet 3i, LLC | The product may contain a manufacturing defect, that if present, may prevent the driver; cover screw or abutment from fully engaging the implant. | Process control | The firm sent an Urgent: Medical Device Recall letter, dated November 2010, to notify customers that received the Osseotite Certain Implant (IOSS511) of the problem. The customers were instructed to return the unopened device(s) to Biomet 3i for replacement. If customers notice an issue, they were to contact Customer Service at 1-800-342-5454 for assistance. Customers were to respond via fax and return the implant(s) to BIOMET 3i using the address and incident number provided. Customers with questions or concerns were instructed to contact Mark Mashburn at 561-776-6906. |  | 241 | 241 | Worldwide Distribution -- USA , BR, JP, TW, CA, Germany, and Guam. | Terminated | November 09, 2011 | 356 | 241
Z-0176-2012 | 59684 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Microtainer MAP Microtube for Automated Process with K2EDTA. |  |  | February 11, 2011 | 2011 | Becton Dickinson &amp; Company | BD Microtainer MAP Microtube is being recalled due to the potential clotting of blood specimens resulting from a reduced amount of EDTA additive. | Device Design | A Becton Dickinson Sales Consultant conducted a "on-site" visit to the single Customer account, Florida Hospital, on February 11, 2011. All units were successfully recovered. All product recovered was quarantined and destroyed. Becton Dickinson believes that all reasonable efforts to communicate the issue to the affected consignee have been completed. If you have any questions or request additional information please do not hesitate to call (201) 847-5636. |  | 10800 | 10,800 units | Distributed only to Florida. | Terminated | November 10, 2011 | 272 | 10800
Z-0175-2012 | 59996 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker Mantis Rod Inserter Assembly Manufactured by: Stryker Spine SAS Z. I. Marticot - 33610 - Cestas France; Distributed in the USA by: Stryker Spine 2 Pearl Court. Allendale, NJ 07401-1677 USA |  |  | June 24, 2011 | 2011 | Stryker Spine | There have been reports of difficulty in rod rotation using the newly designed Mantis Rod Inserter. | Device Design | Stryker notified branches/agencies, surgeons and hospitals beginning June 24, 2011 with an Urgent Product Recall letter, dated June 24, 2011, via Fed Ex. The letter identified the affected product, the issue, potential hazards, risk mitigations, and the return process. Customers were asked to examine their inventory and reconcile any inserters and shafts with the catalog numbers provided using the Customer Response Form. Customers are to fax the form to Regulatory Compliance at 201-760-8370 within 5 days of receipt of the letter. The affected product should be retrieved and returned. Questions should be directed to 201-760-8298. |  | 1138 | 471 units US, 667 units international | Nationwide Distribution | Terminated | November 13, 2012 | 508 | 1138
Z-0177-2012 | 59425 | No | KPT | Radionuclide dose calibrator. | Radiology | Radiology | 510(k) | Capintec CRC 25 PET, item number 5130-3217 with Software Revisions 2.01 and 2.02 |  |  | July 08, 2011 | 2011 | Capintec Inc | Software error in Rev. 2.01 and Rev 2.02 for Capintec CRC 25PET | Software design | Capintec, Inc. sent a Software Error Alert letter dated July 8, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contac their Capintec Customer Support Representative to receive a software upgrade. For questions call Capintec Customer Support at 1-800-631-3826. |  | 69 | 69 | Worldwide distribution - United States including CA, FL, IL, MA, NJ, OK, PA, and VA; and internationally to Brazil, China, India, Indonesia, Italy, Japan, Korea, South Korea, and Turkey. | Terminated | December 10, 2013 | 886 | 69
Z-0181-2012 | 60022 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Clinac, Trilogy, Trilogy Tx and Novalis linear accelerators. |  |  | September 09, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has received reports in which a user has remotely rotated the gantry into contact with the couch or with the patient, in both the manual mode and the automate mode. | Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated September 9, 2011 to all affected customers via Federal Express with attached instructions. The letter provides the description of problem; references, details, recommended user actions, and varian actions.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***In addition, the letter also states "THIS LETTER IS NOT A SUBSTITUTE FOR THE CAREFUL REVIEW OF ALL USER DOCUMENTATION ASSOCIATED WITH YOUR PARTICULAR EQUIPMENT AND SOFTWARE."If you have any questions, contact Varian Oncology Help Desk at USA and Canada: 1.888.VARIANS (888.827.4265); Europe +41 41 749 8844 and/or Internet: Oncology Systems customer site-www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 5823 | 5823 devices | Worldwide Distribution. | Terminated | July 24, 2012 | 319 | 5823
Z-0179-2012 | 59918 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Quick Shield with Snappy Safety Holder, Greiner bio-one. |  |  | August 11, 2011 | 2011 | Greiner Bio-One North America, Inc. | Reports of the Vacuette QuickShield Safety Holders breaking during use due to a crack at the holder hub. | Nonconforming Material/Component | Greiner Bio-One sent a "QUICKSHIELD WITH SAFETY HOLDER" notification letter dated August 11, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to stop distributing and using the affected products. A Product Disposition Form was attached for customers to complete and return for replacements. Contact the firm at 888-266-3883 for questions concerning this recall. |  | 443950 | 443,950 pieces | Nationwide Distribution-USA (nationwide) including the states of CO, IL, KY, NY, PA and VA. | Terminated | March 20, 2012 | 222 | 443950
Z-0178-2012 | 60104 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LifeCell Strattice Reconstructive Tissue Matrix for Stoma Reinforcement |  |  | July 13, 2011 | 2011 | LifeCell Corporation | The use of Strattice for stoma reinforcement at the time of stoma creation is not within the product's cleared indications for use in the US market. | PMA | LifeCell began notifying their customers via phone on 7/13/2011. Recall Notice Letters and return response forms, dated July 12, 2011, were sent out on the same day. The letter identified the affected product and informed customers of the reason for the recall. Customers were asked to examine their inventory and quarantine the affected product. If customers have affected product, they were to contact Customer Solutions at 1-866-423-2433 to arrange for return of the product to LifeCell. Also, the attached Recall Notification form should have been completed and returned no later than August 17, 2011. Questions should be directed to Customer Solutions at 1-866-423-2433. |  | 111342 | 2651 sublots (of 42 lots) | Nationwide Distribution | Terminated | March 04, 2013 | 600 | 111342
Z-0180-2012 | 60048 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-Stat cTnI cartridges |  |  | September 29, 2011 | 2011 | Abbott Point Of Care Inc. | i-Stat cTnI cartridges beginning with letter "T" may exhibit a higher than expected variability in reported results. This increased variability may result in falsely elevated or falsely depressed results. | Material/Component Contamination | APOC issued Urgent Recall Notices/Business Reply Forms dated September 2011 to consignees. The letter identified the affected product and explained the background of the problem. Customers are to discontinue use of the cartridges from the lots identified. All unused product will be credited and should be returned as instructed on the attached Business Reply Card. If the affected product was further distributed, a copy of the Recall Notice should be provided to those customers as well. APOC shipped replacement product in order to minimize any disruption. If replacement product has not been received then customers should contact Customer Service at 1-800-366-8020, option 5. Questions regarding this information should be directed to Abbott Point of Care Technical Support at 800-366-8020, Option 1 or customers' local APOC representative. |  | 1688755 | 1,688,755 cartridges | Worldwide Distribution | Terminated | March 11, 2013 | 529 | 1688755
Z-0172-2012 | 59829 | No | ISA | Therapeutic massager. | Physical Medicine | Physical Medicine | 510(K) Exempt | Shoulder Flex Massage Device |  |  | August 30, 2011 | 2011 | King International | The covered rotating mechanism can catch on user's clothing, jewelry (e.g. necklace), and hair. There have been reports of one death and a strangulation associated with using the device. | Component design/selection | An Urgent Product Recall letter was sent via regular mail to customers. The letter identified the affected product and described the reason for the recall. Customers were asked to stop using Shoulderflex massagers immediately and to safely dispose of them in the trash. The device components should be disposed of separately so that the device cannot be reassembled and used. This means that the power supply should be separated from the massager unit, and the massage fingers should also be removed and disposed of separately. Customers are to complete and return the enclosed the postage-paid card to King International in order to confirm receipt of the recall letter. If additional information is needed regarding this recall, customers are to contact King International LLC at 503-524-7046 or visit their website at www.shoulderflex.com or via mail to PO Box 2384, Beaverton, Oregon 97075.On 8/25/2011 FDA published a health alert on the FDA website. On 8/31/2011 the firm issued a nationwide press release. |  | 11934 | 11934 | Nationwide Distribution | Terminated | March 06, 2014 | 919 | 11934
Z-0182-2012 | 60036 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Extended Brilliance Workstation-NM (EBW-NM) |  |  | September 15, 2011 | 2011 | Philips Medical Systems | A problem related to a software error has been detected in the Philips Multiple Gated Radionuclide Angiocardiography (MUGA) application. A false negative interpretation due to a falsely elevated ejection fraction or a false positive interpretation due to a false low ejection fraction could occur. | Software design | Philips sent all affected sites an Urgent Medical Device Correction Letter on September 27, 2011 for the MUGA issue. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. Customers were asked to retain a copy of the communication with the equipment instruction for use. Additionally, a field change order was released on September 26, 2011, customers will be receiving release notes as part of their correction informing them of the changes and corrections to the software. The new software will be installed on affected units by Philips Field Service Engineers.For customers in North America contact Customer Care Solutions Center at 1-800-722-9377. Select option 5. In all other countries contact the local Philips Healthcare office. |  | 1181 | 1181 systems | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | January 30, 2013 | 503 | 1181
Z-0186-2012 | 60242 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Breeze Self-Adhesive Resin Cement, Part Number: N97A, |  |  | July 08, 2011 | 2011 | Kerr/pentron Dba Kerr Corporation And Pentron Clinical | Pentron is voluntarily recalling the Breeze Self Adhesive Resin Cement because it was discovered that there may be a curing problem that may cause cemented crowns to fall off after use of the product. | Pending | Pentron Clinical sent an "Urgent: Medical Device Recall" letter dated July 18, 2011 with an attached Recall Return Form (via USPS First Clas mail) to all affected customers. The letter identified the affected product, explanation of the problem, and actions to be taken. Customers were instructed to discontinue use and return any affected product in their inventory. The letter asks that each customer complete and return the enclosed Recall Return Form via fax to 714-288-4609, in order to confirm receipt of the customer notification.Customers were instructed to contact Pentron Clinical Customer service at (800) 551-0283 (option 1) directly to handle the arrangements of a quick return and replacement. |  | 336 | 336 units | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Belgium, Canada, Cypress, Greece, Germany, Italy, Jordon, Poland, Romania, Turkey and Vietnam. | Terminated | June 28, 2013 | 721 | 336
Z-0187-2012 | 59461 | No | LLI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Invacare 600 lb. Bariatric Bed, Head motor extension cable used on component numbers BAR5490 and BAR5490IVC. These components are part of the following Bed Model Number(s): BAR600, BARPKGCA, BAR600IVC, BARPKGIVC-1633, and BARPKG. |  |  | June 23, 2011 | 2011 | Invacare Corporation | Invacare Corporation decided to recall the products because risk analysis reveals that miskeying or disconnection of the head motor extension cable during bed set-up may result in product damage, or in rare instances, the potential for fire. | Labeling design | On 9/19/11, Invacare sent Recall Notificaiton Letters to their customers. They instructed their customers to quarantine the product, complete and return a response form. |  | 35747 | 35,747 | Distribution Nationwide and to United Arab Emirates, Dominican Republic, Egypt, Guam, Mexico, &amp; Virgin Islands. | Terminated | July 02, 2013 | 740 | 35747
Z-0184-2012 | 59919 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Greiner bio-one, 4ml K2E EDTA K2 Vacuette tube. Used for venous blood collection. Item number 454209. |  |  | August 12, 2011 | 2011 | Greiner Bio-One North America, Inc. | Vacuette 4ml EDTA DK2 tubes have been reported for short fills to no fills. | Process control | Consignees were notified by Greiner bio-one letter on August 11, 2011. They were instructed to: Stop distributing the material, complete the product Disposition form attachÂż and to fax it to Greiner bio-one and to destroy or return any of the affected products for replacement or credit. Distributors were asked to provide a list of their customers that have received the item, preferably in an Excel file, so that they could be notified of the recall. |  | 504000 | 504,000 pieces | Nationwide. | Terminated | March 12, 2012 | 213 | 504000
Z-0183-2012 | 60122 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | Pulstar(R) DVT PREVENTION SYSTEM, Anti-Embolism Wrap System, albahealth(R), A Subsidiary of Encompass Group, LLC. |  |  | October 12, 2011 | 2011 | Albahealth LLC | The wrap portion of the device was tearing at the seam. | Component change control | Albahealth, L.L.C. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 19, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to check their inventory and to quarantine all affected product. A response form was enclosed for customers to complete and return. The firm provide information on how to return the affected product and receive replacement product. Questions concerning this recall are directed to call 865-354-5229 for assistance. |  | 2582 | 2,582 pairs | Nationwide Distribution- including the states of CA, CO, FL, GA, IN, KY, LA, NY, OH, PA, TX, and TN. | Terminated | March 30, 2012 | 170 | 2582
Z-0190-2012 | 60366 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Step 2, 3000-22000 GPD (Sodium Hydroxide 2.0N) cleaning solution. 4 oz plastic bottle. Lot # 11IC002, Serial #'s: 11IC00210 - 11IC002294. |  |  | October 24, 2011 | 2011 | Isopure Corp | The cleaning solution Step 2, 3000-22000 GPD labeled bottles (NaOH) were filled with Step 1, 3000-22000 GPD chemical (HCl). | Employee error | Isopure sent a Product Recall Notice dated October 24, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue the use of any Step 2 chemical from the affected product, place any remaining chemical into the original box, seal it with tape, and segregate the product to prevent it from being placed back into stock. The firm will pick up the product for return on a call tag. Customers would receive replacement product for all affected cases of the Step 2 chemical the firm received.For questions call 502-722-1000. |  | 48 | 91cs/12/4 oz plastic bottles | Nationwide Distribution including the following states: AL, AZ, CA, FL, GA, IN, KS, LA, MA, MO, MS, NC, NH, NJ, NM, NV, OH, OR, PA, SC, TN, TX, and VA | Terminated | February 24, 2012 | 123 | 48
Z-0188-2012 | 60370 | No | LKB |  |  | General Hospital | Enforcement Discretion | PDI Alcohol Prep Pads packaged in the INRatio kit. |  |  | October 11, 2011 | 2011 | Alere San Diego | Alere has initiated a recall of the Professional Disposables International (PDI) alcohol prep products due to the potential presence of low levels of the microorganism Bacillus cereus. | Material/Component Contamination | Alere sent a recall notification letter dated October 6, 2011, and Verification Form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use an alternate prep pad that is not involved with the PDI recall, or use a sterile gauze pad with isopropyl alcohol. If customers purchase alcohol pads from a store or pharmacy, customers can confirm with their pharmacist that the pads they are purchasing are not associated with the PDI recall. INRatio/INRatio2 kits at Alere are currently being reworked to include a sterile PDI alcohol pad.Customers were instructed to complete and FAX the enclosed Verification Form within 10 days to confirm their receipt of the notice and to indicate the number of discarded/required replacement alcohol prep pads.Customers with questions about INRatioÂż/ INRatioÂż 2 were instructed to consult their healthcare provider, or call Alere Technical Service at 877-866-5313 (24 hours a day, 7 days a week). |  | 14248 | Estimated 14,248 Total units | Nationwide Distribution | Terminated | February 29, 2012 | 141 | 14248
Z-0189-2012 | 60059 | No | CDQ | Urea nitrogen test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Urea Nitrogen List Number 7D75-21 and 7D75-31 used for the Quantitation of urea nitrogen in human serum, plasma, and urine. |  |  | October 03, 2011 | 2011 | Abbott Laboratories, Inc | The reagents are being recalled due to particulate matter in some of the reagent cartridges. The particulate matter has been identified as Aureobasidium fungus species. | Process control | The firm decided to recall and they sent a Product Recall Letter to their customers on 10/03/2011. Questions regarding the letter should be directed to Customer Service at 1-877-4ABBOTT for US customers, and customers outside of the US should contact their local area Customer Service. Customers should determine if you have any remaining inventory. Please discontinue and destroy any remaining inventory of these lots. Order alternative lots of Clinical Chemistry Urea Nitrogen. Complete and return the enclosed Customer Reply form. |  | 5704 | 5704 units of Urea Nitrogen | Worldwide Distribution. | Terminated | May 11, 2012 | 221 | 5704
Z-0193-2012 | 60097 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Restoration ADM System X3 Acetabular Insert Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; Stryker France ZAC Satolos Green Pusignan Av de Satolas Green 69881 Mey Sieu Cedex France |  |  | June 03, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become ware that there is the potential for interpretation of the product labeling which may lead to an incorrect implant being used. | Labeling False and Misleading | The firm, Stryker, sent an "URGENT PRODUCT CORRECTION" letter and Product Correction Bulletin dated June 3, 2011 via Fed Ex to its customers. The letter and bulletin describes the issue, potential hazards, risk mitigation and product correction instructions. The customers were instructed to complete and return the enclosed PRODUCT CORRECTION ACKNOWLEDGMENT FORM within 5 days via fax to: 201-831-6069.If you have any questions, feel free to call (201) 972-2100 or Hip Marketing Team at 201-831-5280. |  | 1677 | 1677 units | Nationwide distribution. | Terminated | November 13, 2013 | 894 | 16891
Z-0198-2012 | 60145 | No | DZJ | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | Right IBO Blade , for use with Stryker TPS Reciprocating Saw, REF 5100-37-901S2, Stryker Instruments, Kalamazoo, MI. |  |  | October 18, 2011 | 2011 | Stryker Instruments Div. of Stryker Corporation | There is a potential for these blades to break where the blade is welded on to the arbor. This hazard can result in the following harms: tissue damage due to the broken blade being aspirated or swallowed, additional surgery to remove a broken blade, infection due to an unretrieved device component, complications associated with increased time under anesthesia to retrieve a broken blade, osteomyeli | Device Design | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated October 18, 2011 to Risk Managers, Office Managers, and Oral Surgeons. The letter described the products under recall, the problem and a series of actions to be taken. Customers were asked to immediately locate recalled products and quarantine. Products were to be shipped back to Stryker. The letter included a Business Reply Form to be faxed back to Stryker Instruments Regulatory Department, 866-521-2762. The letter instructs to forward this notification to all affected locations if the recalled products were further distributed. Customers with questions can call Stryker at 1-800-253-3210. |  | 113 | 113 | Worldwide Distribution - USA (nationwide) and the countries of Brazil, Canada, China, Germany, Greece, Hong Kong, India, Jordan, Miramar, Saudi Arabia, Singapore, Spain, Switzerland, United Kingdom. | Terminated | July 12, 2013 | 633 | 12631
Z-0191-2012 | 60174 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Spectra Optia Apheresis System, REF: 61000, Caridian BCT, Lakewood, CO 80215. |  |  | September 21, 2011 | 2011 | Caridian BCT, Incorporated | Possible higher than expected platelet loss following therapeutic plasma exchange or plasmapheresis. | Software design | CaridianBCT sent a "Safety Alert" letter dated September 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm will be sending a recall letter to consignees in November 2011, stating they are voluntarily recalling the software and will provide an upgrade once the software has been approved.If you have any further questions please contact your CaridianBCT Representative, the CaridianBCT Support Center at 1-877-3-FYU-BCT ( US Toll free 1-877-299-4228 ) or (303) 231-HELP ) ( 1 (303) 231-4357, or your local CaridianBCT Customer Service office. |  | 852 | 852 units (US: 333; Foreign: 519) | Worldwide Distribution -- USA ( nationwide ) and the countries of Australia, Belgium, Canada, Chile, China, Colombia, Hong Kong, India, Japan, Malaysia, Mexico, Pakistan, Republic of Korea, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand and Vietnam. | Terminated | February 01, 2013 | 499 | 852
Z-0185-2012 | 60202 | Yes | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | The LIFEPAK 20 defibrillator/monitor is an AC powered external defibrillator with battery backup used by healthcare providers in hospital and clinic settings. |  |  | December 30, 2008 | 2008 | Physio Control, Inc. | When users do not disconnect the test plug and reconnect the QUIK-COMBO electrodes to the therapy cable, the test plug signal can be misinterpreted as a patient waveform. | Employee error | Physio Control sent a " Urgent Medial Device Safety Alert" letter dated January 2009, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.The letter instructed the customers to follow the Operating Instructions for AED and Manual Defibrillation mode (Section 4-Therapy, pages 58 and 61) that instructs users to "Connect therapy electrodes to the therapy cable and confirm cable connection to the defibrillator" to prevent users from mistakenly leaving the test plug and reconnect the QUIK-COMBO electrodes to the therapy cable, the test plug signal can be misinterpreted as a patient waveform.Recommendations to users who use a Therapy Cable:-Disconnect the QUIK-COMBO test plug prior to patient use-Confirm therapy electrodes - to - therapy cable connection-Confirm therapy cable - to - defibrillator connection-Confirm therapy electrodes - to - patient connectionAdding 2005 American Heart Association software that includes audible and visual prompts to LIFEPAK 20 defibrillators manufactured prior to January 2007 is optional and not required. If the customers would like to have this optional kit installed in your defibrillators, please contact Physio Control, Inc.Customers can call Technical Support at 1-800-442-1142 for any questions about this recall or or visit our website at www.physio-controlnotices.com/LP20testplug. |  | 23467 | 23,467 units worldwide | Worldwide Distribution--USA ( nationwide ) and the countries of Bahrain, Lebanon , Cyprus, Netherlands, Egypt, Pakistan, Germany, Saudi Arabia, India, South Africa, Iran, Syria, Israel, Turkey, Jordan, UAE,Kuwait,Yemen, Asia Pacific, Australia, New Caledonia, Guam, New Zealand, Hong Kong, Latin America, Anguilla, Jamaica, Argentina, Mexico, Bahamas, Nicaragua, Brazil, Panama, Chile, Peru, Colombia, Trinidad, Costa Rica,Uruguay, El Salvador, Venezuela, Guatemala, Virgin Islands and Canada. | Terminated | November 21, 2011 | 1056 | 23467
Z-0201-2012 | 60118 | No | DYW | Orthodontic plastic bracket. | Dental | Dental | 510(k) | Spirit MB Brackets, part #494-1211. |  |  | May 18, 2011 | 2011 | Ormco Corporation | Ormco Corporation is voluntarily recalling three lots of Spirit MB Brackets due to a manufacturing error. | Device Design | Ormco Corporation sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product. If any of the affected product is found, customers are directed to contact Ormco Customer Care at 1-800-854-1741 to receive a RMA number. The RMA number will allow a quick return and replacement or credit. Additionally, a Recall Return/ Acknowledgement Form was attached for customers to complete and return via fax to 909-962-5605. |  | 135 | 135 packs | Worldwide Distribution-USA (nationwide) including the states of FL, MN, TX, and VA, and the countries of Australia, France, Germany, Italy, Japan, Mexico, Morocco, Netherlands, New Zealand, South Africa, Turkey, and Ukraine. | Terminated | April 24, 2012 | 342 | 145
Z-0209-2012 | 60420 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath GUI versions: 3.17.5 to 3.17.9, 4.2.1 to 4.2.3, and 4.3.0 to 4.3.0.4. |  |  | July 13, 2011 | 2011 | SCC Soft Computer | Soft Computer, Clearwater, FL, initiated a correction on the following SoftPath GUI versions: 3.17.5 to 3.17.9, 4.2.1 to 4.2.3, and 4.3.0 to 4.3.0.4. SoftPath architecture does not correctly handle AP macros and data fields inside report sections that are stored in the database. Improper information can be displayed in all places where the diagnosis from the database is read and displayed. | Software design | The firm, SCC Soft Computer, sent a "CORRECTION COMMUNICATION" notice on July 13, 2011 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to restrict security access to the template set-up files and to acknowledge receipt of the notice; grant SCC personnel permission to check their report templates and correct any that are affected; indicate if they wish to schedule the utility to be run and if they would like to have the SoftWeb parameter PathPDFLinkOnly set to "Y". SCC will review the templates on the customers systems; discuss any irregularities requiring change and move or remove AP macros and data elements as needed.If you have any questions, contact the Director of Quality Management at (727)789-0100. |  | 142 | 142 | Worldwide distribution: USA (nationwide) including states of: AR, AL, AZ, CA, CT, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MO, MN, NH, NJ, NY, OH, OK, PA, SD, TX, VA, WA, WI, WV including Puerto Rico; country of Canada. | Terminated | July 17, 2012 | 370 | 142
Z-0207-2012 | 59990 | No | DXH | Telephone electrocardiograph transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | LifeStar ACT III Platinum Sensor Ambulatory Cardiac Telemetry; an automatically activated system with a high performance triple channel EKG, office hookup and auto send/auto detect that requires no patient intervention to either capture or transmit an arrhythmia when it occurs; Manufactured by Card Guard, 2 Pekeris St, Rehovot, 76100, Israel; Model FG-00084 |  |  | September 26, 2011 | 2011 | Lifewatch Services Inc | LifeWatch received complaints of minor shocks and burns during use of the LifeStar ACT systems with shorts in the leads. | Nonconforming Material/Component | The firm, LifeWatch, sent an " Urgent Device Recall" letter dated September 27, 2011 to the affected customers. The letter describes the product, problem and action to be taken. The customers were instructed to immediately refrain from providing these devices to their customers; stop using the affected sensors, return their entire kit in the pre-paid mailer, and use the new device kit included with the recall letter. In addition, the letter informed the users that the replacement device has upgraded software and firmware that allows for detection of shorts in the cable, and if detected, cuts off current flow to the patient to prevent minor shocks and burns.If you have any questions, please call (847) 813-4258. |  | 442 | 442 sensors | Nationwide distribution. | Terminated | May 15, 2012 | 232 | 1094
Z-0202-2012 | 59960 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed paradigm mio Infusion Sets MMT-921 |  |  | May 16, 2011 | 2011 | Medtronic Minimed | The consignee received four (4) boxes of mio infusion sets (MMT-921) that were shipped to them in error. These infusion sets were shipped in a package that included a MiniMed Paradigm REAL-TIME Insulin Pump among other items.The mio infusion sets were not authorized to be shipped to the consignee based on Medtronics internal product handling procedures and they may not function properly.T | Employee error | The firm, Medtronic, sent an "IMPORTANT RECALL INFORMATION" letter dated May 16, 2011 to consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue use right away; use the replacement mio infusion sets provided to them in this package; return all unused Lot 3700255 mio infusion sets using the emailed pre paid return label, and drop the package off at any UPS location or call UPS at 800.742.5877 to schedule a pick up."Please note that only the Lot 3700255 mio infusion sets were shipped in error. The other products included in the package were not shipped in error and do not need to be returned."If you have any additional questions relating to this notification or any questions in general call 800.999.9859 ext.63704. |  | 4 | 4 boxes | Nationwide distribution: AL only. | Terminated | December 14, 2011 | 212 | 4
Z-0205-2012 | 60392 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | SUPER TORQUE MB Angiographic Catheter product is labeled in part: "***5F (1.65 mm)***110 cm .038" (0.97 mm)***MULTIPURPOSE A SUPER TORQUE MB***Max. 1200 psi (8274 kPa)***Markers***OPEN end***Sideholes***SPECIAL***Angiographic Catheter*** REF Cat No.***Lot***Use By***Sterile EO***Assembled in Mexico Cordis Corporation, 14201 NW 60th Ave. Miami Lakes, Florida 33014, USA***EC***REP*** Cordis Cashel, Cahir Road Cashel, Co Tipperary, Ireland***Cordis***a Johnson&Johnson; company***" |  |  | November 21, 2011 | 2011 | Cordis Corporation | Cordis Corporation is issuing a Correction for their SUPER TORQUE MB Angiographic Catheter after receiving a report of marker bands dislodged during a procedure. Model Numbers 532598A, 532598B, 532598C, 532598D, SRD5724MB and SRD5727MB are affected by the Cordis Correction. | Error in labeling | Cordis Corporation sent a Correction letter dated November 21, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers are instructed to read the "Description" and "Recommendations" sections, sign and return an enclosed Acknowledgement Form, and to maintain awareness of the communication until the information has been incorporated into the Cordis Super Torque MB Catheter labeling. Contact Cordis at 800-781-0282 for questions regarding this notice. |  | 310 | 310 units | Worldwide Distribution-USA (nationwide) including Puerto Rico. | Terminated | July 23, 2013 | 610 | 119000
Z-0203-2012 | 60124 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | EHEC Kit Lot# 608096 Containing: Sample Diluent 18.8 ml Lot# 8602EXP.1-BUF WASH 20X II 100.0ml.Lot# 2607EXP.2-EHEC Detection Antibody 10.0ml Lot# 8603EXP.3-EHEC-Enzyme Conjugate 10.0ml Lot# 8604EXP.4-Substrate II 12.5ml Lot# 1601EXP.5-Stop-Solution II 10.0 ml Lot# 1601EXP.6-EHEC Positive Control 3.4 ml Lot# 8601EXP.7-EHEC Negative Control 3.4 ml Lot# 1603EXP.8-96 transfer pipettes-1 Microwell holder-3 Microwell strip sealers |  |  | October 10, 2011 | 2011 | Meridian Bioscience Inc | Internal testing has shown that the indicated kits produce false negative results when tested with some direct stool specimens. These false negative reactions are not observed for all devices tested, however, the rate of occurrence will likely impact clinical sensitivity as defined by product registration and associated instructions for use. | Nonconforming Material/Component | Meridian Bioscience, Inc. sent an Urgent Medical Device Recall letter dated October 11, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop use of the affected product, evaluate testing performed with the affected product., and segregate and destroy remaining inventory.Customers were asked to contact Meridian's Technical Service Department by phone or facsimile as necessary. For replacement product phone 800-343-3858, outside North America 513-271-3700, Fax 513-272-5432.Customers were also asked to complete and return the attached Customer Response Form. In the event that they have no remaining inventory, customers were instructed to indicate such on the Customer Response Form and return to Meridian Technical Service. For questions phone 800-343-3858, outside North America 513-271-3700. |  | 3236 | 3236 Kits | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, NC, NH, NJ, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI and WV and the countries of AUSTRALIA, CANADA, CHILE, COSTA RICA, ITALY, JAPAN and NEW ZEALAND. | Terminated | February 19, 2012 | 132 | 3236
Z-0229-2012 | 57585 | No | OJV | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Angiotech***Custom Bone Marrow Tray for University of CA, Davis*** 1 - 10cc Syringe L/L; 1 - 20cc Syringe L/L; 1 - 19ga x 1-1/2" Filter Needle; 1 - 21ga x 1-1/2" Needle; 1 - 25ga x 5/8" Needle; 1 - 8ga x 4" Bone Marrow Needed "J" Type; 1 15ga x 4" Bone Marrow Needle "I" Type; 1 - T- Lok Extraction Cannula; 1 - 10% Providone Iodine Swab Sticks (3/Pkg.); 1 - Probe Guide; 1 - Obturator; 2 - 1% Lidocaine (5mL); 1 - #11 Mini Scalpel; 2 - 4" x 4" Sponges in stacks of 5; 1 - 2" x 3" ADH Dressing; 1 - Towel; 1 - Fenestrated Drape; 1 - Male Luer Cap; 1 - CSR Wrap***Catalog Number: CB0140TL***Sterile - For single patient use. STERILE EO. Caution: Federal (USA) Law restricts this device to sale by or on the order of a physician |  |  | December 22, 2010 | 2010 | Medical Device Technologies, Inc. | Medical Device Technologies, Inc. dba Angiotech, is recalling HSG Procedure Trays and Bone Biopsy Trays distributed between December 2007 and December 14, 2010 for package integrity. | Packaging | The firm, Angiotech, sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated December 22, 2010 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to review their current inventory, segregate the affected lot numbers and complete and return the attached Customer Acknowledgement Form via fax to our Quality Assurance department at 1.352.338.0662 or 1.800.333.0440 to arrange for return of the identified product.If you have any questions concerning this notification, call (352) 338.0440 ext. 355 or ext 358. |  | 225 | 225 boxes | Worldwide distribution: US (nationwide) and countries of: Canada, Columbia, Costa Rica, Japan, Mexico, Peru and Saudi Arabia. | Terminated | November 30, 2011 | 343 | 32491
Z-0210-2012 | 59818 | No | LYJ |  |  | Neurology | PMA | VNS Therapy Aspire HC Generator and VNS, Model 105 |  |  | August 12, 2011 | 2011 | Cyberonics, Inc | The devices are being recalled because the output current delivered to the vagus nerve is less than the design intent and there is a potential charge imbalance at the lead cathode and generator-can during stimulation. | Nonconforming Material/Component | Cyberonics sent a Field Safety Alert on August 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The firm decided to recall and initiated contact with implanting facilities in the United States on 08/15/2011. A Field Safety Alert was delivered on 09/19/2011 to implanting facilities. Customers were instructed not to prescribe AspireHC pulse generators for implant until Cyberonics resumes shipment of AspireHC pulse generators that have the issue remedied. Cyberonics recommends that physicians continue their current practice of titrating programming parameters for patients implanted with an AspireHC pulse generator to achieve appropriate clinical outcomes, paying particular attention to report of:Lack or loss of efficacyPain, inflammation, swelling/edema in the neck area; and/orPatient perception of stimulation or muscle twitching in the area surrounding the implant. Cyberonics recommended routine systems diagnostic testing to aid in device and patient monitoring. Physicians may contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com.For questions regarding this recall call 281-228-7323. |  | 340 | 340 units | Worldwide Distribution - USA including TN, TX, IL, OH, FL, CA, CO, OR, AZ, NY, PA, MO, NJ, WI, OK, MI, ID, KY, IA, MA, NE, GA, LA, NY, and Norway. | Terminated | August 20, 2012 | 374 | 340
Z-0213-2012 | 60101 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | DICOM Option for the Norland Illuminatus Software Revs 4.2.0 though 4.3.1. Models Norland XR-600, XR-800, XR-46, EXCELL, XR-36 and Eclipse DXA Bone Densitometers. |  |  | September 23, 2011 | 2011 | Cooper Surgical, Inc. | The DICOM Software, Norland Illuminatus DXA Software with the DICOM Option (Part Number 434A133) version 4.2.0 through 4.3.1, which is installed on the XR-600, XR-800, XR-46 and EXCELL DXA Bone Densitometers, could lead to improper patient information getting into a previous patient's record. | Software design | The firm, CooperSurgical, sent a "Notice of RECALL-URGENT" letter dated September 23, 1011 to its customers. The letter described the product, "Problem/Situation", and "When can this problem occur?". CooperSurgical will send a new revision of Illuminatus software to the customers once available.If you have any concerns about your existing database of patients, please contact our Norland Technical Support team at (203) 601-5200 Extension 9888 or 9889. |  | 184 | 184 | Worldwide distribution: USA (nationwide) including states of: AZ, AR, CA, FL, GA, IN, KS, KY, LA, MD, MA, MI, MO, NJ, NY, NC, OK, PA, TN, TX, UT, and WI ; and countries of: ARGENTINA, AUSTRALIA, BELGIUM SPAIN, CANADA, CHINA, DENMARK, EMIRATES, ENGLAND, FRANCE, GERMANY, GREECE, IRELAND, MALTA, POLAND, SAUDI ARABIA, SOUTH KOREA, and TAIWAN. | Terminated | April 09, 2013 | 564 | 184
Z-0212-2012 | 60165 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR) |  |  | June 06, 2011 | 2011 | AGFA Corp. | Incorrect entries in mapping files could lead to inaccurate measurement display in adult and pediatric echo reports when using the IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR). | Pending | The firm, AGFA HealthCare" sent an 'URGENT FIELD SAFETY NOTICE" dated May 27, 2011 via FED-EX by June 6, 2011 to all customers using the affected product. The letter describes the product, problem and actions to be taken. The customers were instructed to distribute this information within their facility to all those who need to be aware and to complete and return the attached FEEDBACK FORM via fax to 864-421-1668 or email to debbie.norris@agfa.com. AGFA will provide software corrections to all affected sites and will contact the customers to schedule the analysis and provide any corrections if required.If you have any questions about this matter, please contact your local Agfa HealthCare Service at 877-777-2432. |  | 280 | 280 | Worldwide distribution: AK, AZ, CA, CO, DC, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, and WV; and country of: Canada. | Terminated | May 08, 2014 | 1067 | 280
Z-0211-2012 | 60154 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Cardiovascular Results Management (RM) Software Versions RM 2.04.37.04 to RM 7.8 SU2 |  |  | June 06, 2011 | 2011 | AGFA Corp. | Improper merging of IMPAX Cardiovascular Results Management (RM) reports causes the signed report to become unavailable. | Software design | AGFA Healthcare sent an "URGENT FIELD SAFETY NOTICE" letter dated May 27, 2011, via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to carry out a merge post finalization of an RM reports. If it is necessary to modify a finalized report, AGFA recommends that they make the modification through one of the following actions:Amend the report through the RM interface, orCall AGFA Service to change the status of the report to "preliminary" and then proceed with the finalization of the report.AGFA issued an Acknowledgment, via FAX- Back that the information was received and understood to affected customers. For any questions customers were instructed to contact their AGFA Healthcare Service at 877-777-2432. |  | 323 | 323 | Worldwide Distribution - USA (nationwide) and Canada | Terminated | July 02, 2013 | 757 | 323
Z-0258-2012 | 60021 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Precedence System; |  |  | September 27, 2011 | 2011 | Philips Medical Systems | A mechanical malfunction of the button on the hand controller on the Philips Precedence and SKYLight SPECT systems may get stuck and cause unintended motion of a detector. | Device Design | Philips Medical Systems sent a "Customer Information" letter dated September 27, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter lists additional precautions the customers should use while operating the device. Philips will notify customers for replacing the affected product through an Action for Performance-Proactive Field Change Order to resolve the issue. Contact Philips Customer Care Solutions at 1-800-722-9377, option 5, for questions concerning this notice. |  | 658 | 658 - all units in recall | Worldwide Distribution-USA (nationwide) including Puerto Rico. | Terminated | December 21, 2012 | 451 | 658
Z-0260-2012 | 60173 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Product is a synthetic absorbable suture with needle and labeled as sterile. The product is sealed inside an inner foil pouch with an outside Tyvek pouch. The outer sterile barrier seal is intact. The product is boxed and shrink wrapped for protection against contamination. |  |  | January 02, 2009 | 2009 | C P Medical Inc. | The seal integrity of the outer product pouch may be compromised due to partial delamination of the two layers that make up part of the pouch. | Nonconforming Material/Component | CP Medical, Inc. sent a "RECALL NOTICE" letter dated January 5, 2009 to all affected customers.The letter describes the product, problem, and the actions to be taken by the customers. Customers are instructed to identify affected product and to return the product to the firm. Customers should call 1-503-232-1555 for a return goods authorization number and shipping account number. Contact your customer service representative at 1-800-950-2763 for questions concerning this recall. |  | 13632 | 13632 sutures | Nationwide Distribution | Terminated | December 13, 2011 | 1075 | 13632
Z-0261-2012 | 60405 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Software: SA BASE 3.1.4.x, 3.1.6.x and SA LabMic 4.0.0.x, 4.0.1.x, 4.0.2.x, 4.0.3.x, 4.0.4.x. |  |  | August 08, 2011 | 2011 | SCC Soft Computer | On 08/08/2011 SCC Soft Computer, Clearwater, FL initiated a correction on following SoftLab used with: SA BASE 3.1.4.x, 3.1.6.x and SA LabMic 4.0.0.x, 4.0.1.x, 4.0.2.x, 4.0.3.x, 4.0.4.x. A client reported a flagging error where the result was flagged as critical in the Instrument Menu but the printed report did not show the corresponding flag. All affected clients were notified of the issue on 0 | Software design | SCC Soft Computer notified all affected customers on August 8, 2011, of the recall via the SCC Soft Computer proprietary Task management system to manage client communication. Clients use this tool to log in requests for services and receive and provide updates to any issue. Each client user has a unique password and every action is date and time-stamped. Each Task has a unique identifier and a full audit trail.SCC recommended to the customer until Live on the correction avoid making edits in the instrument menu.For further questions please call (727) 789-0100. |  | 255 | 255 | Worldwide Distribution - USA ( nationwide) and Canada. | Terminated | November 22, 2013 | 837 | 255
Z-0265-2012 | 60301 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Endoscopy 3.55 MM Super 90 S-SERFAS Energy Probe; electrosurgical device. Model number: 279-351-300; |  |  | October 24, 2011 | 2011 | Stryker Endoscopy | An increased number of complaints for tip breaks were observed for the 3.5mm Super 90-S SERFAS Energy Suction Probes manufactured from June to August 2011. | Component design/selection | Stryker Endoscopy Urgent Medical Device Removal notifications datyed October 25, 2011, were sent to all consignees via traceable mail. Sales representatives and regional managers were notified by e-mail. Product is to be returned to Stryker Endoscopy using the pre-paid shipping label. Replacement product may be obtained by calling Stryker at 1-800-624-4422. The enclosed Acknowledgement receipt form is to be completed and returned. Questions or comments may be directed to Stryker at 1-800-624-4422 or e-mail at 90S@stryker.com. |  | 6307 | 6307 | Worldwide distribution including USA, Australia, Canada, China, Greece, Hong Kong, India, Italy, Japan, France, Netherlands, Poland, South Africa, Spain, Sweden, Switzerland and the United Kingdom | Terminated | February 29, 2012 | 128 | 6307
Z-0262-2012 | 60205 | No | KOY | Surgical skin degreaser or adhesive tape solvent. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PDI Electrode Prep Pads, Skin Prep B339 are non-sterile alcohol prep pads. They are included in the Spacelabs Healthcare Hook-Up kits. Each kit contains 2-4 pieces of PDI Electrode Prep Pads . |  |  | October 12, 2011 | 2011 | Del Mar Reynolds Medical, Ltd. | The non-sterile alcohol prep pads (B339) are included in Spacelabs Healthcare hook-up kits may have the presence of a bacterium, Bacillus cereus. These alcohol prep pads were manufactured and recalled by Professional Disposables International (PDI). | Material/Component Contamination | SpaceLabs Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated October 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop using the affected product and dispose of them per their hospital or clinic protocol. For questions contact Spacelabs Healthcare at 1-800-522-7025 and select 2for Technical Support. |  | 25000 | 25,000 pieces of PDI alcohol prep pads-15 different Spacelabs hook-up kits | Worldwide Distribution - USA (natiowide) and the countries of South Africa, Turkey, Bangladesh, Dominican Republic, Peru, Belgium, Ecuador, Romania, Austria, Honduras, Guatemala, Denmark, Republic of Korea, Argentina, Thailand, Italy, Austria, Chile, Bolivia, Mexico, Sweden, Netherlands, Switzerland, India, Serbia, Australia, Morocco, Turkey, ISpain, Czech Republic, Slovakia, Poland, China, Columbia, Portugal, Canada, Germany, United Kingdom, Singapore, France, Russian Federation, Bahrain, and Kuwait | Terminated | February 23, 2012 | 134 | 25000
Z-0264-2012 | 60365 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Super Sheath and Super Sheath R/O Introducer Sheaths, Sterile EO, Manufactured by Togo Medikit Co., Ltd. 17148-6 Aza Kamekawa, Oaza Hichiya, Hyuga City, Miyazaki Prefecture 883-0062, Japan, Distributed by Boston scientific Corporation One Boston Scientific Place, Natick, MA 01760-1537 USA. |  |  | November 03, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is initiating a medical device recall removal regarding nine lots/batches of Super Sheath and Super Sheath R/O Introducer Sheaths. Boston Scientific was informed through product complaints that the 0.035" ID dilator may be labeled incorrectly as a 0.038" ID dilator. | Packaging process control | Boston Scientific sent an URGENT MEDICAL DEVICE RECALL REMOVAL -IMMEDIATE ACTION REQUIRED letter dated November 21, 2011. The letter described the product and the problem. Customers were instructed to segregate the affected product immediately and return it to Boston Scientific. Distributors were advised that the recall removal depth is to the hospital level and the recall removal notification should be forwarded to their customers.For questions regarding this recall call 763-494-1133. |  | 2660 | 2660 | Worldwide Distribution - USA including AZ, AR, CA, CO, FL, GA, IL, KY, LA, MD, MA, MN, MO, NE, NJ, NM, NY, OH, OK PA, RI, TN, TX, UT, VA, WY, WI, and WV and the countries of BELGIUM, BRAZIL, CZECH REPUBLIC, FRANCE, GERMANY, GREAT BRITAIN, IRAQ, ITALY, NETHERLANDS, RUSSIAN FEDERATION, SOUTH KOREA, SPAIN, SWEDEN, and UKRAINE. | Terminated | January 15, 2013 | 439 | 2660
Z-0270-2012 | 60278 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | "***UPLIFT TECHNOLOGIES INC.***Bringing quality to life.***Model No.***CA200***Description***Lifting commode chair***Weight Range***80-300+ lb-136+ kg***" |  |  | June 27, 2011 | 2011 | Uplift Technologies Inc. | On June 26, 2011 Uplift Technologies Inc. Nova Scotia, Canada initiated a recall of the Uplift Commode Assist Model No. CA200, Lot No. 1C0101. The plastic seat does not meet manufacturing specifications. The seat may not fit the unit properly and may interfere with the unit's normal locking mechanism. | Nonconforming Material/Component | Uplift Technologies, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 27, 2011 to all affected customers The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their inventory and to quarantine the affected product. Customers are to call Uplift's Customer Service Department at 1-800-387-0896 to request a Return Materials Authorization number in order to return the product.. |  | 219 | 219 chairs | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | May 31, 2012 | 339 | 219
Z-0266-2012 | 60005 | No | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | Rapid Strand Rx; the RAPID Strand Rx kit consists of custom loaded absorbable seeding spacers and I-125 radionuclide brachytherapy sources (seeds) spaced at prescribed distance and configuration within a sleeve (tube) made of absorbable suture material, stiffened, loaded into prostate seeding needles, packaged and then sterilized by Gamma sterilization method, accompanied by a calibration strand, a sealed source calibration certificate (100% assay report), a radiograph of the loaded needles in the needle-holding card, and a source certification technical data sheet.; Products Provided by: Oncura, 3350 North Ridge Avenue, Arlington Heights, IL 60004 USA, Prescription Filled by: AnazaoHealth, 5710 Hoover Blvd., Tampa, FL 33634 USA; product identifier: I125RSRX, model 7000. |  |  | September 20, 2011 | 2011 | Medi-Physics Inc. dba GE Healthcare | There is a potential brachytherapy over or under dose with the RAPID Strand Rx kit that may impact patient therapy. | Process control | GE Healthcare/Oncura sent a "Urgent Medical Device Correction" letter dated September 20, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were advised to assess the pre and post plan dosimetry for their patient, taking into account the possible variation of the seed strength. GE Healthcare/Oncura will contact the customers in the near future to confirm their receipt of the letter and to discuss any questions they may have.Please call if you have any further questions (847) 398-8400 and 011-44-1494-737251 (Amersham, United Kingdom, GMT) |  | 23 | 23 orders | Worldwide Distribution -- USA (nationwide) including the States of Arizona, Florida, Louisiana, Mississippi, New Jersey, New York, Pennsylvania, Puerto Rico, Virginia and Washington, and the countries of Canada and the United Kingdom | Terminated | January 20, 2012 | 122 | 23
Z-0271-2012 | 60374 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | TUM-E-VAC REF 2075, Rx ONLY, NON STERILE --- Ethox International, Buffalo, NY 14204, USA --- Intended use: gastric lavage |  |  | March 25, 2011 | 2011 | Ethox International, Inc. | Ethox Tum-E-Vac Gastric Lavage Kit #2075, Lot Number 031123255 was labeled with expiration date symbol and date of 2010/10 on the unit product label. The carton was labeled with a manufacturing date symbol and date of 2010/10. The Ethox Tum-E-Vac Gastric Lavage Kit #2075 does not have an expiration date. | Employee error | Ethox International, Inc. sent an "URGENT DEVICE RECALL" letter dated March 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to examine their stock immediately to determine if they have the affected product. Customers are to return the affected product to the firm by using UPS account 128959. Customers will be reimbursed for the returned goods. Additionally, a response form was enclosed for customers to complete and return. Contact the firm at (716) 842-4000 for questions regarding this recall. |  | 130 | 130 units | Nationwide Distribution-including the states of California, Connecticut, Illinois, Ohio, and Virginia. | Terminated | November 28, 2011 | 248 | 130
Z-0263-2012 | 60155 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) - Physiological Patient Monitoring Solution |  |  | October 18, 2011 | 2011 | Draeger Medical, Inc. | During product training on the Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) it was discovered that the weight-based drug dosage calculation may indicate incorrect recommended values. In addition, a misalignment of ECG and ART waveforms delivered from an Infinity(RO Acute Care System (IACS) Monitoring Solution (M540) was observed on the Infinity Central Station. | Software design | The firm initiated their recall of this device on October 17, 2011 by issuing a letter to the consignee instructing customers to follow the Instructions for Use and use the Infinity Central Station only for remote assessment of a patient's status. They recommend that you do not consider the waveforms on Infinity Central Station for primary diagnosis purposes. In case of a high priority alarm situation, "Check the patient and treat if necessary".If the misalignment issue is observed, it can be cleared by undocking and redocking the M540 or by changing the order of waveforms displayed on the Cockpit screen (or any other change to the auto display screen layout). |  | 476 | 476 | Nationwide distribution - IL | Terminated | April 01, 2013 | 531 | 476
Z-0269-2012 | 60316 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Wolf Medical Supply Inc., WOLF-PAK REDI-FLO ELASTOMERIC INFUSION PUMP, FILL VOLUME 270 mL, FLOW RATE 2mL/Hr. |  |  | October 20, 2011 | 2011 | Wolf Medical Supply Inc. | Redi-Flo Elastomeric Infusion Pumps may have a higher flow rate than specified (more than +15%). | Process control | On 10/20/2011 Wolf Medical Supply Inc., customers were notified by telephone and email. Product was distributed to end users. |  | 24 | 24 cases (576 units) | Nationwide distribution, including Puerto Rico. | Terminated | April 25, 2012 | 188 | 40
Z-0272-2012 | 60164 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CardioVascular (CV) Admin Tool |  |  | June 06, 2011 | 2011 | AGFA Corp. | Potential for misattribution of data and for a report to become unavailable when the user requests for it to be moved to another study that already has a report associated with it. | Software design | An 'URGENT FIELD SAFETY NOTICE" was sent via FED-EX by June 6, 2011 to all the sites using the affected product. The letter described the issue and mitigation. The letter recommends that users not move reports to studies where a report already exists. If a report is being moved because the existing report is irrelevant to the study, then the existing report should be deleted prior to moving the new report to that study. If the report is being moved because the user wants to combine the content of the reports, then the "merge" function should be utilized. The letter also states that a software enhancement is currently available. Acknowledgment via FAX-Back was requested from the sites. Questions regarding the issue should be directed to customers' local Agfa HealthCare Service at 877-777-2432. The firm has identified 185 additional sites and an "URGENT FIELD SAFETY NOTICE" was sent on October 12, 2012. |  | 323 | 323 units | Worldwide Distribution -- USA, Virgin Islands and Canada | Terminated | June 26, 2012 | 386 | 323
Z-0273-2012 | 60123 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | IBC INTERNATIONAL BIOPHYSICS CORP. 2101-2 EAST ST. ELMO, SUITE 275, AUSTIN, TX 78744 SUCTION HANDLE (STERILE) PART NUMBER: 1990S LOT NUMBER 041811-2065 EXPIRATION DATE: 18/APR/2015 |  |  | October 05, 2011 | 2011 | International Biophysics Corp | Potential for the sterility to be compromised. | Packaging | International Biophysics Corp. sent a" Medical Device Recall Notice" dated October 5, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Use of the product in the referenced lot number should cease immediately. Customers were asked to follow the instructions in the letter to return the affected product.If you have any further questions please call (512) 814-0070. |  | 1200 | 1200 units | Nationwide Distribution including the states of AZ, CA, CO, DE, MS, OR, PA, TN, and TX | Terminated | March 23, 2012 | 170 | 1200
Z-0278-2012 | 60430 | No | LNX |  |  | General Hospital | 510(k) | SCC Soft Computer, Softlab Laboratory information System with DORA versions: 2.0.0.4 with Hot Fix 1.2808, 1.2809, 1.2956, or 1.3476; 2.0.0.5; 2.0.0.6; 2.0.0.7 with Hot Fix 1.6595; and 2.1.0.3. |  |  | March 22, 2011 | 2011 | SCC Soft Computer | Potential for patient data to be sent to the wrong patient record. A software change affected how the Softlab handles recovery from certain types of table access conflicts. When such conflicts occur, this function has the potential to cause faulty writing to the database. | Software change control | Clients were notified through SCC Soft Computer's proprietary Task Management System On 03/22/201. All clients have received a hot fix that corrected the issue or made the hosparam change that prevents the issue. |  | 12 | 12 | Sold throughout US and Canada. | Terminated | October 29, 2012 | 587 | 12
Z-0282-2012 | 59913 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Flexible Probe with Blocking Water, GM11002420, (Part of the Segmented Cylinder Set, GM11004150) |  |  | May 20, 2010 | 2010 | Varian Medical Systems, Inc. | The Instructions For Use ( IFU) distributed with product prior to 2004 lacks the correct limit to the sterilization cycles and checks regarding the firm seating of the blocking washer. | Labeling Change Control | Varian Medical Systems sent a Medical Device Recall (PNL-GM-42697-2) in May 2010 for the blocking washer on the flexible probe, Varian part number GM11002420, a component of the Segmented Cylinder Applicator Set. The letter was to inform users that a flexible probe with an unstable or shifted blocking washer must not be used due to the potential for treatment length error. The blocking washer may move from rough handling or sterilization. Users were told to discard any instances of the obsolete Instructions for Use (IFU) in use and replace it with the IFU (PGM11004150 4 March 24, 2010.pdf). A Receipt Verification Card was included requesting consignees to fill out, fax or e-mail immediately to Varian BrachyTherapy. |  | 81 | 81 parts | Worldwide Distribution: USA, Germany, Belgium, Canada, Mexico City, Spain, Brazil, Switzerland, Uzbekistan, India | Terminated | December 05, 2011 | 564 | 81
Z-0281-2012 | 60269 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity(R) Delta, Delta XL, Gamma XXL and Kappa Monitors are used to monitor, physiologically, the vital signs of patients (with Arrhythmia Detection or Alarms). |  |  | October 31, 2011 | 2011 | Draeger Medical, Inc. | Firm received 2 complaints reporting Ventricular Fibrillation (Vfib) arrhythmias were not recognized by the monitor and did not alarm when the Masimo SET pod is used for SpO2 measurements. | Software design | Firm initiated their recall of this device on October 25, 2011 by sending an "Urgent Medical Device Recall" letter to consignees. A Technical Service Bulletin will be generated and distributed globally that instructs the affected International Draeger subsidiaries to inform customers in their countries. The software anomaly has been corrected in software version 7.3. Draeger Service Representatives will schedule time with customers to upgrade the software in affected monitors. The software upgrade will be provided free of charge. |  | 14894 | 14,894 | Worldwide distribution- USA including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SD, TN, TX, UT, VA, WI and WV and International distribution including the countries of Algeria, Angola, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Botswana, Brasilien, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Deutschland, Dominican Republic, Ecuador, Egypt, El Salvador, Equatorial Guinea, Estonia, Finland, France, Frankreich, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indien, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Jordanien, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Madagascar, Malaysia, Malta, Mauritius, Mayotte, Mexico, Moldavia, Moldova, Morocco, Netherlands, Norway, Norway, Peru, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Rumania, Russian Fed., Saudi Arabia, Serbia, Seychelles, Singapore, Slovenia, Slowenien, South Africa, South Korea, Spain, Sri Lanka, SÂżdafrika, Sweden, Switzerland, Syria, Taiwan, Thailand, Tschechische Republic, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and White Russia. | Terminated | July 01, 2013 | 609 | 14894
Z-0279-2012 | 60305 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STERIS 5085 and 5085SRT Surgical Tables, Catalog Numbers: 5085- ST01-410-1; 5085SRT- ST0-1420-6. |  |  | October 17, 2011 | 2011 | Steris Corporation | Some hydraulic column cylinders installed in certain 5085 and 5085SRT Surgical Tables were assembled by their supplier with incorrect snap rings and an anomaly was present in the tilt cylinder. This error may affect the user's ability to move the table top out of the full right tilt position. | Nonconforming Material/Component | On October 17, 2011, Steris sent 'Urgent Voluntary Field Correction Notices' to their customers via Fedex with tracking numbers for delivery confirmation. The letter identifies the affected product and notifies the user/customers of the problem with the surgical tables. The notification letter goes on to inform the customers that a STERIS Service Representative will be contacting them to arrange for the field correction of the affected unit(s). This involves the replacement of the hydraulic column cylinder of each surgical table(s). |  | 167 | 167 tables were distributed - 5085=108 and 5085SRT=59 | Nationwide distribution. | Terminated | March 19, 2013 | 519 | 167
Z-0276-2012 | 59980 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Sensar Foldable IOL with OptiEdge. Model number AR40e 21.5 Diopter. |  |  | September 20, 2011 | 2011 | Abbott Medical Optics Inc (AMO) | The firm is initiating this voluntary recall because of the possibility that certain Sensar IOLs are labeled with an incorrect diopter power. | Employee error | Abbott Medical Optics, Inc. sent an URGENT FIELD SAFETY NOTICE dated September 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using and remove the affected product from their inventory, complete and return the included facsimile form to Abbott Medical Optics and indicate the status of each affected serial number on the form. Customers were asked to return the form even if they have no inventory affected by the recall. For any questions customers were instructed to contact their Abbott Medical Optics Customer Service Representative.For any questions regarding this recall call 714-247-8656. |  | 96 | 96 | Worldwide Distribution including China, India, Indonesia, Korea, Malasyia, Taiwan, Thailand, Austria, Finland, France, Germany, Great Britain, Hungary, Italy, Netherlands, Portugal, and Saudy Arabia | Terminated | January 08, 2014 | 841 | 179
Z-0280-2012 | 60219 | No | DHN | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Inova Diagnostic Inc., NOVA Lite ANA KSL Kit |  |  | September 26, 2011 | 2011 | Inova Diagnostics Incorporated | The NOVA Lite ANA KSL kit label has the wrong fill volume for the IFA System Negative Control (P/N 508186) with 1mL. The component in the kit is correctly labeled and filled with the 0.5mL fill volume. | Process change control | On 9/26/11, Inova Diagnostics forwarded an Urgent Advisory Notice and Product Advisory Notice Fax Back Form to all their customers who purchased the Nova Lite Ana KSL, Product Code: 708390. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to complete the attached fax back form indicating how many of these kits they have in their inventory, quantity destroyed, method of destruction and the quantity of kits that need to be replaced. Customers were also instructed to notify their customers of the advisory notice. Customers were instructed to contact Tana Keivens with any questions or comments at (858) 586-9900 ext 351. |  | 500 | 500 kits | Worldwide distribution to: USA, Mexico, Kuwait, Uraguay, Romania, Italy, Philippines, India, Malaysia, Algeria, Czechoslovakia, Brazil, Colombia, Ireland, Korea, and Saudi Arabia. | Terminated | April 04, 2012 | 191 | 500
Z-0284-2012 | 60142 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions USA, Inc., Symbia T Series SPECT-CT (Single-Photon Emission Computed Tomography and Computed Tomography) System. |  |  | October 14, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Errors could potentially impact image interpretation. The acquisition software is exhibiting a defect in the energy window position when an isotope is peaked. There is a potential patient safety issue when running syngo MI Applications version VA60A with Foresight detectors firmware version v1.5.9.12 on the Symbia systems. | Component change control | Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated October 14, 2011 to all affected customers. The letters informed customers of the acquisition software defect in the energy window position when an isotope is peaked, which could potentially impact image interpretation, and that the AEB firmware version on their imaging system has been changed to a previously released version until acquisition software service pack is available. For questions or assistance they were instructed to contact Siemens Medical Solutions USA, Inc. at (800) 888-7436. |  | 3 | 3 units | Florida, Illinois and Texas | Terminated | October 16, 2012 | 368 | 5
Z-0274-2012 | 60328 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSDV), sterile EO, Rx Only, Synovis Surgical Innovations, A Division of Synovis Life Technologies, Inc, 2575 University Ave W., St Paul, MN 55114-1024. |  |  | October 26, 2011 | 2011 | Synovis Surgical Inovation Div. of | Synovis Surgical Innovations is voluntarily recalling one lot of Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSD-V) product because some of the devices were mislabeled. | Labeling mix-ups | Synovis Surgical Inovation sent a "Recall Notification" letter dated October 26, 2011. The letter described the problem and product affected. Customers were instructed to identify and quarantine the affected product and to complete and fax the Product Return/Replacement Form. |  | 132 | 132 units | Nationwide Distribution including CT, GA, NJ, NY, OH, TX, VA. | Terminated | April 24, 2012 | 181 | 132
Z-0277-2012 | 60436 | No | LNX |  |  | General Hospital | 510(k) | SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. |  |  | May 03, 2011 | 2011 | SCC Soft Computer | On 05/03/2011 SCC Soft Computer, Clearwater, FL initiated a correction on the following SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. A client reported the PT and INR results did not get flagged. The PROT should have been flagged as HIGH and the INR should have been flagged as PANIC. Neither test was flagged. All affected clients were notified of the issue on 05/03 | Software design | SCC Soft Computer initiated a Field correction on May 3, 2011. All customers with affected software have been notified of this issue. Clients were instructed to have a "valcheck" utility run to determine if there are any tests in their system affected by the issue. They were supposed to indicate whether they wanted to schedule to receive the patch or hot fix.If you have any further questions please call ( 727 ) 789-0100. |  | 69 | 69 | Worldwide Distribution-- USA ( nationwide) and Canada. | Terminated | August 07, 2013 | 827 | 69
Z-0285-2012 | 60166 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Curix Opthos H, size 35x43 cm |  |  | October 04, 2011 | 2011 | AGFA Corp. | The film notch was located in the wrong position. | Process control | An "URGENT SAFETY NOTICE" letter was sent via email to the dealers on October 7, 2011. The letter describes the problem and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood has been requested from the dealer sites. Included in each dealer letter, Agfa also provided a customer notification letter to be sent to their end-use customers. Agfa has requested dealers and end-use customers to file a complaint at 1-877-777-2432 for any identified product still in their inventories. Agfa will replace product and arrange for return shipment. |  | 125 | 125 | Worldwide Distribution -- USA, including the states of IN, KS, MI, MN, MO, NC, NE, NY, OH, and TN and the country of Canada. | Terminated | December 22, 2011 | 79 | 125
Z-0292-2012 | 60413 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Hospira LifeShield Latex-Free 150 mL Burette Set, with Calibrated Burette and 2 Clave Ports, Convertible Pin, 77 inch , Soluset Microdrip with Option-Lok; an Rx sterile medical device used to administer fluids; 20 individually wrapped sets per case; Made in Costa Rica, Hospira, Inc., Lake Forest, IL 60045 USA; list 19208-01 |  |  | November 10, 2011 | 2011 | Hospira Inc. | The float valve in the burette sticks to the burette wall and does not open or close properly. | Employee error | The firm, Hospira, sent an "URGENT DEVICE RECALL" letter dated November 10, 2011 to the customers of record via UPS on the same date. The customers were instructed to check their inventory and immediately quarantine any affected product; complete and return Reply Form via fax to 1-866-913-0667 or email to: hospira2530@stericycle.com; return affected product to Stericycle using the label provided with this letter; and if they have distributed the product further, notify their accounts that received the product . Please contact Hospira Customer Care at 1-877-946-7747 or your Hospira representative for information regarding product availability. |  | 12580 | 12,580 units | Worldwide distribution: USA (nationwide) including states of: Alabama, Alaska, California, Colorado, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kentucky, Massachusetts, Michigan, Minnesota, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New York, North Carolina, North Dakota, Ohio, Oregon, Pennsylvania, Rhode Island, Tennessee, Texas, Vermont, Virginia, Washington and Wisconsin, and internationally to Canada | Open, Classified |  |  | 12580
Z-0288-2012 | 60355 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | ImmuneTech brand MyAllergyTest, Specimen Collection Kit, Part No. |  |  | November 04, 2011 | 2011 | Immunetech Inc | ImmuneTech is recalling certain Specimen Collection Kits that contain Triad brand non-sterile alcohol prep pads which have been recalled by Triad Group, Hartland, WI due to Bacillus cereus, microbial contamination. | Material/Component Contamination | The firm, IMMUNETECH, sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated November 4, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their inventory and quarantine all affected kits; if they have further distributed the affected kits, contact Immunetech to accurately identify and immediately notify their customers and facilitate the return and exchange of all remaining kits. If you have any questions, call (650) 312-1066, Monday-Friday, 9:00am -5:00pm (pacific) or email: customersupport@immunetech.com. |  | 310 | 310 | Worldwide distribution. | Terminated | June 14, 2012 | 223 | 3340
Z-0294-2012 | 60031 | scale, stand-on, patient | FRI | Stand-on patient scale. | General Hospital | General Hospital | 510(K) Exempt | Health o Meter Professional Remote Display Digital Physician Scale, Model 498KL; digital patient weighing scale; Pelstar LLC, 11800 South Austin Avenue, Unit B, Alsip, IL 60803, U.S.A., Made n China; Model 498KL |  |  | October 07, 2011 | 2011 | Pelstar, Llc | There is a potential for an electrical short when installing or replacing the batteries in the remote display head of the scales. This short may cause the batteries to smoke, burn or burst. | Device Design | Pelstar LLC. sent a Urgent Device Recall letter dated October 7, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informed them of the potential of an electrical short caused when installing or replacing the batteries which may result in minor injuries to the person installing the batteries. The customers were requested to immediately examine their inventory and quarantine all product subject to the recall. In addition, if they distributed this product to end users, they were requested to identify all customers and notify them of this product recall at once in writing, using the enclosed recall notification letter. The customers were also requested to complete and return the enclosed mandatory response form as soon as possible. Any questions were directed to Health O Meter Professional at 800-815-6615 or 708-377-0600. |  | 2498 | 2,498 units | Worldwide Distribution -- USA (nationwide) including the countries of Canada, New Zealand, Israel and Lebanon. | Open, Classified |  |  | 2498
Z-0293-2012 | 60304 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia 1.5T and 3.0T MRI Systems, Model Numbers: 781396 and 781377, 510(k) #K110151. System units equipped with software version: 4.1.1 are subject to recall/field correction. |  |  | August 03, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The firm determined that during scanning procedures, for some larger patients, some reddening of the skin of can occur in certain instances due to RF heating when the torso of the patient is positioned within the body coil. | Software design | Philips Medical Systems ( Cleveland) Inc. sent a Urgent Field Safety Notice dated August 3, 2011, to all affected customers. Philips Field Service Engineers were dispatched to the affected customer sites to deliver the recall notification and to perform the necessary software corrections. The letter lists the affected MRI systems; describes the problem; the actions to be taken by the customer/user; it also informs the customer of the upcoming software correction which will be applied to each unit by the Philips Field Service Engineers in order to address this issue. If any additional information is required, the customer is instructed to contact their local Philips Field Service Engineer. For further questions please call Customer Care Solutions Center (CCSC) at 800-722-9377. |  | 54 | 26 units - 12 units distributed in U.S. &amp; 16 distributed to foreign customers | Worldwide Distribution -- USA (nationwide) including the states of AL, AZ, CA, GA, OR, MI, NJ and TX., and the countries of AU, BE, CH, DE, FR, JP, NL and SE. | Terminated | February 22, 2013 | 569 | 54
Z-0297-2012 | 60233 | No | DLJ | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | Multigent Vancomycin Calibrators; |  |  | June 08, 2011 | 2011 | Microgenics Corp | Complaint from the sole distributor that the third party control material is recovering high, and outside the published range. | Nonconforming Material/Component | Microgenics sent a Product Removal notification letter dated June 15, 2011 to all affected customers. The letter identified the affected product, the problem, and the necessary actions to be taken. Customers were instructed to discontinue use and dispose of the affected product per local waste ordinances, review all patient results reported with the use of this lot, and retain a copy of this notification for their records. The letter states that if the affected product has been forwarded to another laboratory, a copy of the letter is to be provided to them. An alternative calibrator is available to customers impacted. It can be ordered by calling Thermo Fischer Scientific Customer Service at 1-800-232-3342.Questions regarding the information should be forwarded to Abbott Customer Service at 1-877-4ABBOTT. |  | 118839963 | 809 kits lot 59446762; 651 kits lot 59391741 | Nationwide Distribution (USA) in the state of ILLINOIS | Terminated | February 17, 2012 | 254 | 118839963
Z-0303-2012 | 59479 | Yes | DWB | Roller-type cardiopulmonary bypass blood pump. | Cardiovascular | Cardiovascular | 510(k) | Battery Module for Sarns Modular Perfusion System 8000 |  |  | October 21, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems is aware on one instance in which an incorrectly labeled printed circuit board for the Sarns Modular Perfusion System 8000 was shipped to a single customer. The affected board mislabeled two connection between the board and the transformer. | Employee error | Terumo conducted a Correction/Removal - Customer Communicated Script on November 7, 2011 to the affected customer. The phone script communicated the correction and reason for removal, potential hazards, affected population, and confirmation of possession of affected module. A replacement date was set for November 7, 2011. For questions or concerns contact Terumo CVS Customer Service at 1-800-521-2818, hours are Monday to Friday. 8 AM -6 PM EST. |  | 10 | 10 units | Nationwide Distribution (USA) including the state of San Antonio, TX | Terminated | February 17, 2012 | 119 | 10
Z-0305-2012 | 60121 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI 4.5 |  |  | March 09, 2011 | 2011 | Elekta, Inc. | If "Confirm" or "Unconfirm" settings are clicked during the "Terminated Checking" phase of treatment delivery, Desktop Pro cancels the loaded beam and sets the Beam MUs to zero. Desktop Pro then sends the delivered MUs to MOSAIQ a a zero value. | Software design | Elekta sent an IMPORTANT NOTICE letter dated March 9, 2011 to affected customers. The notice identified the product, the problem, and the corrective action to be taken by the customer. The notice advise all customers to follow any instructions or recommendations covered in the Notice. For questions call +44(0)1293 654200. |  | 56 | 56 | Nationwide Distribution (USA) including the states of: AZ, CA, GA, IL, LA, MD, MA, MI, MN, NJ, NY, NC, OH, PA, TN, VA, WI and Puerto Rico | Terminated | August 09, 2012 | 519 | 56
Z-0306-2012 | 60192 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | ON-Q Silver Dressing is a Antimicrobial wound dressing containing silver. It is individually wrapped in a foil pouch. Each pouch contains 1 single use 4" x 4.5" pad. There are 5 pouches per box. Ref: SD4X4.5; Part Number: IF-151023 and IF-151026. |  |  | June 30, 2008 | 2008 | Acrymed Incorporated | The Silver Dressing, lot H0705J and H0706J may contain dressings that are not properly sealed in the foil pouch, which could affect sterility of the dressing. | Process control | On June 30 2008, the firm e-mailed all affected customers with an Important Product Recall Notification letter.The customers were notified that two lots of SilverDressing may contain dressings that are not properly seated in the foil pouch.The customers were instructed to inspect the inventory, ceased distribution of the affected lots and quaratine the remaining inventory and notify AcryMed.Customers can call AcryMed at 503-624-9830 for any questions about this recall. |  | 41061 | 14,125 units from lot # H0706J and 25,525 units from lot # H0705J | Nationwide distribution -- DE, FL, KY, OR, IN, TX, NV, MT, CA and NY. | Terminated | December 20, 2011 | 1268 | 41061
Z-0299-2012 | 59581 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | The Terumo Advanced Perfusion System 1, Catalog Number 801764 System 1 Base, 220 V |  |  | November 08, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | 1 US customer was operating the Advanced Perfusion System 1 with an unapproved, modified cable, which was in product development, to connect the system an an occluder for the Sarns Modular Perfusion System 8000 . Five(5)of twelve(12) cables from Terumo Japan, which were in product development, were given to a US customer. | Component design/selection | Terumo sent an URGENT MEDICAL DEVICE CORRECTION: Recall Notification dated November 8, 2011 to all affected customers. The letter described the reason for the recall, known/potential hazards and request immediate location and correction of product. The letter instructs customers to stop using the modified cables and return to Terumo for replacement and destruction. Customers were asked to complete the attached response form and fax to the number indicated on the form. All affected customers will be receiving a safety advisory to advise them that Terumo do not recommend the alternative configuration of a System 1 Base with an 8K Occluder. For questions or concerns contact Terumo CVS Customer Service at 1-800-521-2818 or 1-800-292-6551 (fax). |  | 12 | 12 units | Worldwide Distribution - (USA) Nationwide Distribution | Terminated | August 22, 2012 | 288 | 12
Z-0302-2012 | 60289 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1- 4" Roller Pump for Terumo System 1 |  |  | November 10, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo CVS discovered that our In house service center has incorrectly installed a felt washer into the main bearing of nine roller pumps that had been returned for repair or preventative maintenance unrelated to the felt washer. The felt washer was added as a validated running change to production to all roller pumps manufactured since February 2009. The felt washer should not have been installed | Nonconforming Material/Component | Terumo sent an URGENT MEDICAL DEVICE REMOVAL notification letter dated November 10, 2011 to all affected customers. The letter described the affected products, problem, potential hazards, and important clinical information. Replacement roller pumps will be configured at the customer site by a Terumo Field representative with the affected roller pumps being returned via normal returned goods process. Customers were asked to complete the attached response form and fax to the number indicated on the form. For questions or concerns contact Terumo CVS Customer Service at 1-800-521-2818. |  | 3 | 3 units | Worldwide Distribution - USA (nationwide) including the states of: CA, KY, MO, NE, PA and TX and the countries of Canada and Europe. | Terminated | September 06, 2012 | 301 | 9
Z-0304-2012 | 60204 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Ambulatory Blood Pressure (ABP) Monitor, Model 90217A. |  |  | October 12, 2011 | 2011 | Del Mar Reynolds Medical, Ltd. | If external pressure is applied to the cuff of the Ambulatory Blood Pressure, Model 90217A, during deflation, the monitor will stop working and display an EC75 error code. | Device Design | On October 21, 2011, a Medial Device Correction customer letter was mailed to all US customers. A return receipt was requested.The letter identified the affected devices and described the problem. If the device displays an EC75 error code, customers can cycle the power off and then on to restore normal functionality. Please advise patients of this mitigation to avoid a delay in obtaining results. The firm is in the process of upgrading their software to avoid the issue and will contact customers to arrange for an update to their monitors.Customers can call Spacelabs Healthcare at 1-800-522-7025, select 2 for Technical support for any question about this recall. |  | 785 | 785 units worldwide (83 in US and 702 outside US) | Worldwide Distribution | Terminated | May 30, 2012 | 231 | 785
Z-0311-2012 | 60146 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche Cobas c 111 Analyzer |  |  | October 18, 2011 | 2011 | Roche Diagnostics Operations, Inc. | Internal investigations have revealed a potential printer malfunction on the Cobas c 111 analyzer, where one or more numbers or characters can be missing on the hard copy result printout, or the printer may skip leaving blanks which may result in an incorrect print layout. The missing number of characters can be in any position, therefore it could be in the middle of a result, (e.g. 48.53 mmol in | Software design | Roche Diagnostics Operations, Inc sent a Urgent Medical Device Correction dated October 18, 2011, via UPS Ground (receipt required) to all affected customers and one distribution center, with instructions to disable the Auto Print Result feature until new software is available. All printed results will need to be confirmed against the results indicated on the display screen. Questions are to be directed to Roche Diagnostics Technical Support at 1-800 428-2336, |  | 135 | 135 | Worldwide Distribution - USA including West Virginia and Texas. | Terminated | February 21, 2013 | 492 | 135
Z-0309-2012 | 60314 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Arrow International StimuCath(R) Continuous Nerve Block Kit, ASK-19608-SFH, Lot # MF0127039 |  |  | June 20, 2011 | 2011 | Arrow International Inc | Some of the outer trays used in the packaging of the ASK-19608-SFH Continuous Nerve Block Kits, lot # MF0127039 were found to have an insufficient seal. Therefore, the packaging may have been compromised, and sterility of the product cannot be guaranteed. | Packaging | The firm, Arrow International, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 20, 2011, to its customer. The letter included a stock response form to be completed and returned to the firm. The letter also described the product, problem and action to be taken. The customer was instructed to check their stock for the products included in the scope of the recall; cease use and distribution; quarantine all affected product immediately; complete and return the enclosed Recall Acknowledgement &amp; Stock Status Form to their sales representative, and return any affected product freight collect to: Arrow International, Inc., 312 Commerce Place, Asheboro, NC 27203, attn: Asheboro Quality Assurance Manager. Arrow International can provide replacement with a similar product or credit their account.If you have questions or concerns, please call this toll free number 800-233-3187. |  | 17 | 17 kits | Nationwide distribution: CT | Terminated | October 01, 2012 | 469 | 17
Z-0313-2012 | 60323 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Rel 3, a radiographic system, digital |  |  | November 01, 2011 | 2011 | Philips Healthcare Inc. | Missing a small, metal protective cover on the bottom side of the control handle for the ceiling suspension. exposes printed circuit | Process control | The firm, Philips Healthcare, sent a "URGENT - Field Safety Notice" letter dated October 25, 2011, to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed: Do not place your hand in the open slot below the Control Handle and ensure that no patient places his/her hands the opening; review the following information with all members of their staff who need to be aware of the contents of this communication, and retain a copy with the equipment instruction for use. Philips has issued a Field Change order instructing Service Engineers to check affected systems and install the cover where it is missing free of charge.If you need any further information or support concerning this issue call 978-659-4519. |  | 28 | 28 (7 US; 2 Canada; 19 OUS) | Worldwide distribution: USA (nationwide) including states of: NC, SC, TX, and WV; and countries of: Australia, Austria,Canada, Germany, Italy, Netherlands, Sweden, and Switzerland. | Open, Classified |  |  | 28
Z-0319-2012 | 60150 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stereotactic Circular Collimator |  |  | March 25, 2011 | 2011 | Elekta, Inc. | Recent newspaper articles outlined improper use of SRS Cone Collimator accessories that injured patients on Brainlab and Varian systems. | Under Investigation by firm | Elekta sent an IMPORTANT NOTICE letter dated March 25, 2011 to all affected customers. The letter identified the affected product, the problem, key safety procedures and information to users for safe operation of the Stereotactic Circular Collimator. Elekta will issue a Mandatory Action Field Change Order which will include an updated Circular Collimator Instructions for Use - 102022201 and new labels. The notice instructs Users to file this document in the Important Notice section of the appropriate User Manual. Also, Elekta is offering some free WebEx sessions which will include the procedures and recommended operation of these types of accessories. Customers are instructed to register at www.elekta.com/SafeAccessories. |  | 89 | 89 devices | Nationwide Distribution (USA) - including the states of: AZ, CA, FL, GA, MO, NJ, NM, NC, OR, TN, and WV | Terminated | October 02, 2012 | 557 | 89
Z-0310-2012 | 60282 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas c 311 Analyzer, Part Number: 04826876001 |  |  | October 26, 2011 | 2011 | Roche Diagnostics Operations, Inc. | The cobas c 311 analyzer software may incorrectly calculate the volume of the ISE Reference Electrolyte solution remaining on the analyzer. In some cases the incorrect results obtained will not be flagged with any data alarms. | Software design | October 26, 2011 Urgent Medical Device Corrections were sent to all direct accounts in the USA that have installed the cobas c 311 analyzer and included a faxback form to be completed and faxed to 1-877-766-7452. Consignees were instructed to temporarily replace the ISE Reference Electrolyte bottle based upon the frequency calculated from the maximum number of ISE samples run per day and activate the yellow threshold alarm on the system. Customers with question and concerns are directed to call Roche Diagnostics Technical Support at 1-800-428-2366, available 24-hours, 7 days a week. |  | 940 | 940 worldwide | Worldwide Distribution -- USA including the states of Maryland and Texas. | Terminated | July 09, 2012 | 257 | 940
Z-0314-2012 | 60287 | No | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | DePuy Mitek Panalock Loop, Lupine Anchor |  |  | October 31, 2011 | 2011 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Incorrect suture configuration | Employee error | The firm, DePuy Mitek Inc., sent an "URGENT VOLUNTARY PRODUCT RECALL"letter dated October 31, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to not use or sell any of the identified problem products included in this recall; complete the enclosed business reply form and fax to: 1-508-828-3750 or email to: mitekcomplaints@its.jnj.com (note: It is important that we receive this business reply form, even if you have no product remaining in your inventory from the affected lots); isolate all inventory and call 1-877-379-4871 for return goods authorization and replacement product; and return all identified affected product to: (USA)- DePuy Mitek, ATTN: Recall Coordinator, 50 Scotland Boulevard, Bridgewater, MA 02324 or (foreign)-GMED Healthcare EDC Quality Dept Rue de Luxembourg 5, ZI Trazegnies, BE-6180 Courcelles, Belgium. If you have any questions or concerns with regard to this program, please contact your local DePuy Mitek Representative or DePuy Mitek Customer Service directly at 1-877-379-4871. |  | 298 | 298 units | Worldwide distribution: USA (nationwide) and countries of: Austria, Australia, Brazil, Canada, Estonia,Switzerland, China, Czechoslovakia, Denmark, Spain, Finland, France, Great Britain Israel, India, Japan, Korea, Mexico, Martinique, Norway, Portugal, Sweden, Singapore, Solvonia, and Slovakia. | Terminated | August 05, 2015 | 1374 | 1768
Z-0307-2012 | 60125 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Leksell Gamma Plan |  |  | June 23, 2011 | 2011 | Elekta, Inc. | There is a risk that the user may interpret the instructions in the manual too literally, and adjust the TMR 10 dose description to obtain an ETCPD that matches previous clinical prescriptions, which in turn could lead to a prescription dose that deviates a few percent from the intended dose. | Labeling False and Misleading | Elekta sent an IMPORTANT NOTICE letter dated June 23, 2011 to affected customers. The notice identified the product, problem, and corrective actions to be taken by the user. The letter advise all customers to follow any instructions or recommendations covered in the Notice. The letter states that Elekta is working on a new software version and customers will receive the future version when released. For questions contact your local Elekta representative. |  | 2 | 2 | Nationwide Distribution (USA) - including the states of: NM and WA | Terminated | January 31, 2012 | 222 | 2
Z-0295-2012 | 59856 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Flexible Applicator Probe, Part GM11002560 - 3.2 mm diameter braced PVDF. Varian Medical Systems, Palo Alto, CA, Mfg. by Varian Medical Systems, Haan GmbH. |  |  | August 05, 2008 | 2008 | Varian Medical Systems, Inc. | GammaMed Flexible Applicator Probes may not have been heat annealed during manufacturing to prevent shrinkage after autoclave. | Process control | Varian Medical Systems, Inc. sent a Customer Technical Bulletin CTB-GM-603A on August 5, 2008 to all affected customers. The letter included description of device and problem. The bulletin also advised that if they find that the applicator and source guide tube combination is not within the indicated range on the length gauge, to call VBT technical support using the contact information on page 5 of the bulletin to obtain a replacement applicator probe. Customers are to complete and return the attached verification card. If support is needed the letter provides customer support contact information. |  | 653 | 653 parts | Worldwide Distribution -- USA (nationwide): Arkansas,California, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Kentucky, Maine, Maryland, Massachusetts, Minnesota, Montana, Michigan, Missouri, Nebraska, Mississippi, New Jersey, Nevada, New York, North Carolina, Ohio, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Washington, West Virginia, and Wisconsin, and the countries of Austria, Azerbaijan, Brazil, British Columbia, Canada, Chile, Germany, Greece, India, India, Ireland, Israel, Mexico, Morocco, Pakistan, Poland, Republic of Macedonia, Russia, Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, United Kingdom, Ukraine, and Venezuela. | Terminated | December 07, 2011 | 1219 | 653
Z-0312-2012 | 60128 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Leksell Gamma Knife C 1.2, 4 and 4C |  |  | May 26, 2011 | 2011 | Elekta, Inc. | Several of the LMR03 actuators with bronze drive nut have failed unexpectedly, creating a potential safety hazard for the operator and patient. | Pending | Elekta sent an IMPORTANT NOTICE: Urgent Field Safety Notice letter dated May 12, 2011 to affected customers. The notice identified the affected product, problem, and actions to be taken by the user. The letter describes special care to be used in operating the helmet changer actuator LMR03. The letter states that Elekta will replace all actuators LMRO3. For more details contact your local Elekta representative.. |  | 44 | 44 | Nationwide Distribution (USA) - including the states of: CA, CO, DAL, FL, GA, IL, MD, MA, MI, MN, MO, NE, NV, NJ, NY, NC, OH, OK, OR, PA, TN, TX, VA, WA, WI and Puerto Rico | Terminated | October 06, 2014 | 1229 | 44
Z-0326-2012 | 59779 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Plastic Needle with Mandrin; 2.0 mm diameter, length 113 mm. Remote controlled radionuclide applicator system. Varian Medical Systems, Inc. Mfg by Varian Medical Systems Haan GmbH, Haan, Germany. |  |  | September 30, 2008 | 2008 | Varian Medical Systems, Inc. | Plastic needle tip could possibly crack or separate. | Nonconforming Material/Component | Varian Medical Systems sent a "CUSTOMER TECHNICAL BULLETIN" letter dated September 30, 2008 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructed users to inspect needle tips before and after every clinical application to prevent needle tip separation or cracking. |  | 1716 | 1716 parts | Worldwide Distribution - USA (nationwide) including the states of Arizonia, California, Florida, Illinois, Indiana, Michigan, North Carolina, and Vermont, and the countries of Argentina, Austria, Belgium, Canada, Estonia, France, Germany, India, Norway, Pakistan, Russia, Slovenia, So. Africia, Spain, Switzerland, United Kingdom, and Ukraine. | Terminated | December 07, 2011 | 1163 | 1716
Z-0331-2012 | 60229 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Model CAT-850B Catheterization Table, a component of the Toshiba INFX-8000 series X-Ray Interventional System. |  |  | October 03, 2011 | 2011 | Toshiba American Medical Systems Inc | The recall was initiated because Toshiba has confirmed that during patient transfer from a gurney to the catherization table, the tabletop may rotate unexpectedly when sufficient force is applied to the tabletop. The result of the tabletop rotating unexpectedly is the patient being tranferred may fall from the tabletop. | Device Design | Toshiba America Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 3, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Customers were instructed to follow the user instructions by positioning the tabletop at the foot-end position, and ensuringsufficient assistance is available, when transferring a patient from the gurney to the CAT-850B catheterization table. Customers were also instructed to follow the instructions for transferring a patient to the CAT-850B that are included in the operation manual. A Customer Reply Form was included with the letter for customers to complete and return via fax to (877) 349-3054. The customer reply form can also be sent via e-mail to raffairs@tams.com. Questions regarding the recall letter can contact the Director of Regulatory Affairs at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 300 | 300 units | Nationwide Distribution. | Terminated | March 18, 2013 | 532 | 300
Z-0337-2012 | 59767 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Dual Incu i (Atom Infant Incubator Model 100) |  |  | August 23, 2011 | 2011 | Atom Medical Corporation | The caster with lock on the incubator or infant warmer may become loose or break. | Device Design | The firm, Atom Medical, sent an "Urgent: Field Safety Notice" letter dated August 18, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify all devices in their inventory and quarantine them until the rework is complete; identify any customers to whom they have further distributed the product; contact their customers at once and arrange a time in which the field service technician can perform the rework (the parts needed to complete the rework will be sent from Atom Medical's export department ); complete rework of all affected devices and document rework on the Distributor FSCA/Recall Response Form and return via Fax to: +81 3 3812-3199 or email to kira-utc@atomed.co.jp; t-sugino@atomed.co.jp or kenichi.shimada@atomed.co.jp (all forms should be completed and returned by January 31, 2012.) If you have any questions, contact the Manager Quality Control or Manager, Regulatory Affairs at email addresses above, phone: +81-3-3815-3632 or fax +81-3-3812-3199. |  | 105 | 105 units | Worldwide distribution: USA (nationwide) including state of: PA and countries of: Algeria, Austria, Australia, China, Egypt, Germany, Italy, Japan, Korea, Kuwait, Malaysia, Portugal, Russia, Saudi Arabia, Spain, Switzerland, Thailand, The Netherlands, and UAE. | Terminated | September 09, 2013 | 748 | 540
Z-0345-2012 | 59934 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Odyssey |  |  | November 04, 2010 | 2010 | PerMedics, Inc. | Anomaly 1 - Invalid Export of Beam Table and or Collimator Angle for DICOM RT Ion, DICOM Image and RTP Link export files:During PerMedics' internal code reviews for lEC 61217 compliance verification, an issue was identified by PerMedics engineers in three export files: D1COM RT Ion Plan, D1COM RT Image and RTP Link. In some cases, this issue can result in table and/or collimator angle values t | Device Design | PerMedics sent an Odyssey TPS URGENT NOTICE OF CORRECTION letter dated November 4, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Anomaly 1 User Corrective Action:Install the mandatory new version correcting the software as soon as it is available (see PerMedics Corrective Action below). In the meantime, the following recommendations were provided if they planned to use Odyssey RT Ion Plan, DICOM RT Image, and/or RTP Link export:Workaround for the invalid exported table and/or collimator angle errors: If, prior to the software correction, any of these affected file types were exported from Odyssey, manually verify that the table and collimator angles are correct. Additionally, apply appropriate changes to any angles deemed incorrect that were obtained from Odyssey's DICOM RT Ion Plan, DICOM RT Image, and/or RTP Link files.Only those machines with table and/or collimator angles that increase in the clockwise direction will have the incorrect angles.Anomaly 2 User Corrective Action:Install the mandatory new version correcting the software as soon as it is available (see PerMedics Corrective Action below). In the meantime, the following recommendations are provided if the user- Changes the patients scheduled start date prior to the software correction:Workaround for Incorrect Monitor Unit Values when Changing Treatment Start Date: If, prior to the software correction, the user changes the patient's treatment start date on any fully-computed IMRT plan, IMAT plan or plan containing a "Region DVH" prescription, the user should force a full recommendation oft hose patient plans prior to fixing and exporting the plan.The following procedure can be used to force a full re-computation of the plan:- Open the patient's plan worksheet- Change the algorithm type of the plan- Change the algorithm type back to its original value- Close t |  | 15 | 15 | Worldwide Distribution - USA including NY, GA, CA and the countries of Hungary, Brazil, China, India, Italy, and Mexico | Terminated | March 02, 2012 | 484 | 15
Z-0325-2012 | 59885 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | VariSource iX, Software Version 1.1, Part Number AL20020000 VariSource iX Afterloader. |  |  | March 01, 2010 | 2010 | Varian Medical Systems, Inc. | VariSource iX console software may cause incorrect dwell times to be delivered to patient and may result in a mistreatment. | Software design | Varian Medical Systems sent a "MEDICAL DEVICE RECALL" letter dated April 1, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were asked to conduct a simple test to establish if their machine is affected by the problem. Device users were requested to complete an attached receipt verification form and return it to Varian BrachyTherapy via fax or e-mail. Call the Contact Service Dispatch number at 1-800-864-1672 for questions regarding this detail. |  | 3600 | 72 worldwide/50 in U.S. | Worldwide Distribution-USA (nationwide) | Terminated | May 30, 2013 | 1186 | 3600
Z-0336-2012 | 60286 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 10mm x 2.7mm RibLoc Screw, a component of the RibLoc Rib Fracture Plating System, Model RBP1202. |  |  | February 28, 2011 | 2011 | Acute Innovations Llc | RibLoc screw from lot L1008007 was out of specification. It would not engage with the hex tip properly. | Process control | The firm, ACUTE Innovation, sent an "URGENT NOTICE: DEVICE RECALL" letter dated February 21, 2011, to all consignees/customers via FED-EX on February 28, 2011. The letter describes the product, problem and actions to be taken. The customers were instructed to identify product having the lot code(s) and quarantine them; return all product from this lot to Acute Innovations immediately using shipping labels provided or return to: Acute Innovations, LLC, 21421 NW Jacobson Rd, Suite 700, Hillsboro, OR, 97124, and if they further distributed this product, notify those customers of this recall and request return of product(s).Customers can call ACUTE Innovation at 1-866-623-4137 for any questions about this recall. |  | 299 | 299 units (267 in US and 32 to South Africa) | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, GA, IL, MN, NC, NY, UT, and WA; and countries of: South Africa. | Terminated | December 06, 2011 | 281 | 299
Z-0330-2012 | 60231 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Varian brand GammaMed 12i Accessory: Clamping adapters for bronchial catheters, 1.8mm diameter, GammaMed 12i; Ref: BT-00264, Model 2-H63; Product is distributed by Varian Medical Systems Inc., Palo Alto, CA and Manufactured by Varian Medical Systems Haan GmbH, |  |  | October 06, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The Product Models 1 &amp; 2 have a design flaw problem where the clamping adapter canobstruct movement of the source cable in the event the bronchial catheter is dislodgedfrom the connector. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION LETTER/FIELD SAFETY NOTICE" letter dated October 6, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Contact the Varian Help Desk at 1-888-827-4265 for questions regarding this notice. |  |  |  | Worldwide Distribution--USA (nationwide). | Terminated | October 22, 2012 | 382
Z-0344-2012 | 60302 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips GEMINI TF 64 Diagnostic CT X-ray/PET Imaging System, Catalog/Model #882471, 510(k) #K052640. System units equipped with software version: 3.5.2 are subject to recall/field correction. |  |  | April 25, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | In response to complaints received by Philips Medical concerning certain identified software errors reported by their customers, the firm is conducting a recall/software correction to several of their GEMINI PET/CT X-ray system units in the field. The firm believes that if these software defects were to recur, it could impact the performance of reliability of the these PET/CT system units. | Software design | Philips Healthcare sent a "CPE CUSTOMER LETTER" dated October 28, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Additionally, device operation information is provided to customers. A Philips Field Service Engineer will be visiting each customer site to install the necessary software upgrades. Contact the Customer Care Solutions Center at 1-800-722-9377 for questions concerning this notice. |  | 10 | 10 system units were distributed. | Nationwide Distribution including the states of CO, LA, MD, MI, MN, PA, TX, VA, VT, and WI. | Terminated | February 22, 2013 | 669 | 38
Z-0322-2012 | 59814 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ZMR Porous Revision Hip Prosthesis Surgical Technique, utilized for the following ZMR devices: Porous Stem,Nitrited Porous Revision, Slotted Porous Bowed Stems with and without Compression Nut, Splined Straight Stem, Taper Stem and Bodies, Calcar Bodies, and Spout Bodies. |  |  | October 21, 2011 | 2011 | Zimmer Inc. | This recall was initiated because it may not be clear to the surgeon whether the device can be used based on the instructions for use. | Error in labeling | The firm Zimmer, sent an "URGENT DEVICE CORRECTION" letter dated August 26, 2011 to Marketing Personnel, Distributors, Sales Representatives, Distribution Operation Managers, and Surgeons. The letters identified the reason for the recall, the product, and required actions to be taken. The letter also included Medwatch reporting information and a Certificate of Acknowledgement form to fax back to the recalling firm to confirm receipt of the recall notice. For questions contact Zimmer at 1-877-946-2761. |  | 119284 | 119,284 total devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Dubai/UAE, Eschbach, Finland, France, Germany, Greece, Hong Kong, Hungary, Italy, Japan, Jordan, Israel, Korea, Lebanon, Mexico, Netherlands, New Zealand, Peru, Poland, Portugal, Russia, S. Africa, Saudi Arabia, Singapore, Slovakia, Spain, Switzerland, Taiwan, Thailand, Turkey, UK, Ukraine, and Venezuela, | Terminated | August 29, 2013 | 678 | 119284
Z-0320-2012 | 60250 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI R.4.5 and R4.6 |  |  | October 14, 2011 | 2011 | Elekta, Inc. | The software did not operate as expected. | Software design | Elekta sent an IMPORTANT NOTICE letter dated October 14, 2011 to all affected consignees. The letter identified the affected product, the problem, and the actions to be taken. This notice contained precautionary measures that are required for the operation of the equipment. The notice advise all customers to follow any instructions or recommendations covered in the Notice. The notice instructs Users to file this document in the Important Notice section of the appropriate User Manual. For questions contact your local Elekta representative. |  | 77 | 77 | Nationwide Distribution (USA) - including the states of: AZ, CA, CT, GA, ID, IL, LA, MD, MA, MI, MN, MO, NJ, NY, NC, ND, OH, OK, OR, PA, RI, TN, VA, WA, WI and Puerto Rico | Terminated | November 03, 2014 | 1116 | 77
Z-0333-2012 | 60244 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Patient Management System, Models 6441*, 6442*, 6465*, 6488 (USA). |  |  | October 12, 2011 | 2011 | Boston Scientific CRM Corp | Current Boston Scientific defibrillator systems measure shock lead impedances up to 200 ohms, and export results to the programmer and LATITUDE monitoring system via the Daily Measurements feature. A measured shock lead impedance that is greater than 125 ohms or less than 20 ohms will generate an alert message on programmer screens and a Red Alert within LATITUDE (if activated). However, the mea | Software design | Boston Scientific sent an "IMPORTANT DIAGNOSTIC INFORMATION" letter dated October 2011 to all affected customers. The letter described the product, problem and actions to be taken by the customers. |  | 25300 | 25,300 (14,700 US; 10,600 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the countries of AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NORWAY, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, CANADA, AUSTRALIA, NEW ZEALAND, CYPRUS, CZECH REPUBLIC, GREECE, GUADELOUPE, HUNGARY, BANGLADESH, CHINA, COCOS ISLAND, HONG KONG, INDIA, INDONESIA, KOREA, MALAYSIA, PHILIPPINES, SINGAPORE, SOUTH KOREA, SRI LANKA, THAILAND, RUSSIA, SAUDI ARABIA, SLOVAKIA, SLOVENIA, SOUTH AFRICA, NEW CALEDONIA, POLAND, PORTUGAL, QATAR, REUNION,IRAN, ISRAEL, JORDAN, KUWAIT, LATVIA, LEBANON, LUXEMBOURG, MONACO, MOROCCO, JAPAN, ARGENTINA, ARUBA, BAHAMAS, BARBADOS, BELIZE, BERMUDA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, ECUADOR, EL SALVADOR, FALKLAND ISLANDS, HAITI, HONDURAS, JAMAICA, MEXICO, PANAMA, PERU, TRINIDAD &amp; Tobago, URUGUAY, VENEZUELA, AND NETHERLANDS ANTILLES. | Terminated | December 04, 2012 | 419 | 25300
Z-0332-2012 | 60238 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Head Ring Posts with part number 970.280 - re-usable components of the Frame Array Module (of the Optical Guidance Platform and Floorstand devices. |  |  | October 04, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the Head Ring posts used by both the Optical Guidance Platform FrameArray module and the Floorstand where the head ring posts may be damaged due to excessive mechanical stress resulting in possible failure during usage. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated October 4, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Customers are advised to verify that they are using the current design of the head ring posts (part # 970.284) and to contact Varian if replacements are needed. Customers may contact the Varian Help Desk for questions regarding this notice at 1-888-827-4265. |  | 820 | 205 kits ( Each kit contains 4 Head Post Rings) 820 HPR's | Worldwide Distribution. | Terminated | July 25, 2012 | 295 | 820
Z-0335-2012 | 60228 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography X-Ray System, 510(k) #K033326, Model Number: 728231, Serial Numbers: 9745, 90169, and 95414. The recalled device systems are equipped with Version 3.5 Beta 1 software. |  |  | September 22, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were found for the Brilliance Big Bore and Brilliance 64 Computed Tomography X-ray Systems equipped with software version; 3.5 Beta1. The non-conformances stem from software errors identified in the operating software. | Software design | Philips Healthcare released a 'Field Change Order' (FCO #72800531) on September 12, 2011. Philips sent an - URGENT Medical Device Correction letter dated October 17, 2011, to all affected customers. The letter listed the affected device systems; described the nature of the software problem; and any possible hazard inherent to this situation. The letter informed customers on how to identify the affected products and what action to take. In this instance, the customer was asked to review the Customer Release Notes which were provided to each customer location once the software update is installed. The letter notified the customer that a Philips representative would be installing a software upgrade in order to address the referenced software issues free-of-charge, through a Field Change Order. Customers were instructed to contact their Philips representative or local Philips Healthcare Office for further information or support for this issue. For North America and Canada customers were instructed to contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 3 | three (3) units were distributed. | Worldwide Distribution - USA including Texas and the countries of Italy, India and China. | Terminated | February 22, 2013 | 519 | 10
Z-0342-2012 | 60497 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI X-ray Volume Imaging System |  |  | October 20, 2011 | 2011 | Elekta, Inc. | A small number of sites have experienced situations with kV generator transformers overheating due to a diode failure on the interface board. | Nonconforming Material/Component | The firm, Elekta, issued a "FIELD CHANGE ORDER MANDATORY ACTION" notice dated October 20, 2011 to its customers to add fuse protection for the kV transformers. The notice describes the product, problem, provides instructions and procedures and actions to be taken. The customers were instructed to follow any instructions or recommendations covered in the notice and to file the document in the Important Notice section of the appropriate User Manual. Elekta will contact the customer if the FCO is applicable to their kV generator.If you have any questions, call 770-300-9725. |  | 123 | 123 | Nationwide distribution: USA including states of: AL, AK, AZ, CA, CT, FL, GA, IL, IA, KY, MD, MS, MI, MN, MS, MO, MT, NE, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TX, VT, VA, WA, WV, WI, Puerto Rico, &amp; DC. | Terminated | August 09, 2012 | 294 | 123
Z-0341-2012 | 60153 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | TC pCO2 Electrodes, Model/Part #945-655-E5260, 510(k) #K043003. |  |  | September 27, 2011 | 2011 | Radiometer America Inc | Transcutaneous pCO2-electrodes were wrongly coded as: pO2 electrode on the hardware plug for the firm's TCM series monitors. This causes the monitor to interpret the electrode wrongly, causing it to fail all calibrations and never get ready for use. | Employee error | Radiometer contacted the affected U.S. customer and informed them of the recall via telephone on September 27, 2011. Radiometer requested the return of the affected product received from the firm. For questions call (440) 871-0463. |  | 4 | 4 units were distributed. | Worldwide Distribution - USA (nationwide) in the state of San Diego, CA. | Terminated | February 19, 2012 | 145 | 4
Z-0347-2012 | 60335 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Product is labeled in part - Pouch label: "***NanoFUSE***DBM***Manufactured and Distributed By: Nanotherapeutics, Inc. 13859 Progress Blvd, Suite 300 Alachua, FL 32615 (386) 462-9663***Description: nanoFUSE DBM***Size***Order No.:***ID***Expiration***STERILE R***Donated Human Tissue Allograft***SINGLE PATIENT USE ONLY***Rx Only***nanotherapeutics***" |  |  | March 10, 2011 | 2011 | Nanotherapeutics, Inc | On 03/10/2011 Nanotherapeutics Inc, Alachua, FL initiated a recall of their Origen DBM with Bioactive Glass Catalog # 22-2002 (2cc), 22-2005 (5cc) and 22-2010 (10cc). The same product is also distributed as NanoFUSE DBM Catalog # NAN109-02 (2cc) NAN109-05 (5cc) and NAN109-10(10cc). The products were manufactured without an approved 510k. Direct accounts were notified of the recall on 04/08/2011 | No Marketing Application | Nanotherapeutics sent an Urgent: Medical Device Recall letter dated April 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove any unimplanted/unsold product shipped to them. Customers were asked to notify their customers and instruct them to make notification until the end user (hospital/physician) is reached. Recall Instructions:Immediately examine inventory and quarantine product subject to recall.Return all quarantined product subject to recall to:Nanotherapeutics, Inc.Attn: Receiving Department13859 Progress Blvd, Suite 300Alachua, Florida 32615For questions call 386-462-9663 |  | 131 | 131 units | Nationwide Distribuiton including TX, FL, AZ, CA, NY, and MA. | Terminated | April 03, 2012 | 390 | 12814
Z-0363-2012 | 60458 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath GUI versions: 4.2.3.3-4.2.3.6 and 4.3.0.5-4.3.0.12. |  |  | July 20, 2011 | 2011 | SCC Soft Computer | On 07/20/2011 SCC Soft Computer, Clearwater, FL, initiated a correction on the following SoftPath GUI versions: 4.2.3.3-4.2.3.6 and 4.3.0.5-4.3.0.12. A client reported that the HPV results on a patient report were incorrect. The HPV results were from the previous order on the patient and not the current order. All affected clients were notified about the issue on 07/20/2011. | Software design | SCC Soft Computer sent a "CORRECTION COMMUNICATION" dated July 20, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. SCC personnel will review the customized Lab Results AP macro and/or the Micro Results AP macro on customer systems and will either discuss with them the irregularities requiring change or reconfigure macros as necessary. |  | 97 | 97 | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | May 21, 2013 | 671 | 97
Z-0355-2012 | 60451 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI Versions: 4.3.0.10, 4.3.0.11, and 4.3.0.12. |  |  | June 13, 2011 | 2011 | SCC Soft Computer | On 06/13/2011, SCC Soft Computer initiated a correction on the following SoftPath GUI Versions: 4.3.0.10, 4.3.0.11, and 4.3.0.12. A client reported when they verified the contents of the fax reports, they found one of the specimen diagnoses to be missing. Parts A through D were present as well as Part F, however, Part E diagnosis was missing for one of the patient reports. All affected clients wer | Device Design | SCC Soft Computer initiated a Field Correction on June 13, 2011. All customers with affected software have been notified of this issue.SCC recommends until the corrected patch is loaded to temporarily use one of the following alternative working solutions: 1. Use MS Word Editor functionality within SoftPath application for creating patient reports. In order to ensure Text Control editor is not inadvertently activated, the templates for Text Control editor will be deleted and reinstalled at the clients request by their SCC representative after receipt of the correction. 2. In Report Template set-up, the bottom margin should be made slightly larger than the size of the footer plus its distance from the bottom page border. For further explanation, please consult your SCC representative for further details, or call (727) 789-0100. |  | 14 | 14 | Class II Recall - Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | May 31, 2012 | 353 | 14
Z-0359-2012 | 60409 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Version 2.0.0.15 Consult instruction for use. RaySearch Laboratories AB Sveavagen 25, SE-111 34 Stockholm Sweden. support@raysearchlabs.com |  |  | October 19, 2011 | 2011 | RAYSEARCH LABORATORIES AB | The firm has found an issue with DICOM RT Image representation of DRR images from RayStation 2.0 which needs to be clarified in the user manual. The symptoms involve the Pixel Spacing and Image Position attributes in the RT images being incorrect, thus enabling the user to make measurements that are erroneous and possibly misalign the patient. | Device Design | The firm, RaySearch Laboratories, sent an email with an "Urgent Field Safety Notice Medical Device Correction" notice dated October 19, 2011 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to pass on the information in the field safety notice to everyone involved and create routines to compensate for the incorrect DRR images; if the DICOM DRR images are not used to make any form of absolute measurements, no action needs to be taken. Raysearch laboratories will provide a script to update these DICOM files to have the correct values upon request.For further information please contact RaySearch Laboratories AB Quality and Regulatory Affairs Manager at +46 722 366 110 or david.hedfors@raysearchlabs.com; RaySearch Americas Director of Customer Support at +1 877 778 3489 or freddie.cardell@raysearchlabs.com. |  | 19 | 19 units | Worldwide distribution: USA (nationwide) including states of: MA, NJ, and NY; and country of: Netherlands | Terminated | May 11, 2012 | 205 | 19
Z-0348-2012 | 60237 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Oscillating Saw Attachment/Large w/Key for Small Battery Drive. |  |  | October 17, 2011 | 2011 | Synthes USA (HQ), Inc. | The potential exists for the Large Oscillating Saw Attachment to disengage from the Small Battery Drive handpiece. If the attachment disengages from the handpiece, there is a potential for injury to the patient and/or user. | Device Design | Synthes sent a "NOTICE: MEDICAL DEVICE RECALL" letter dated October 17, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The firm requests that customers return their affected device(s) for servicing. Instructions are provided for customers to follow. Customers having questions regarding this notice should contact the firm at 610-719-5450. |  | 5 | 5 | Nationwide Distribution including the states of CO, LA, NY, and OK. | Terminated | December 18, 2014 | 1158 | 5
Z-0349-2012 | 60245 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Foundation Tibial Broach P/N 801-01-013 |  |  | October 26, 2011 | 2011 | Encore Medical, Lp | The strike plate may dislodge from the broach stem handle during impaction of the device. | Nonconforming Material/Component | DJO Global sent an URGENT FIELD SAFETY NOTICE letter dated October 26, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the information to all those who need to be aware within their organization or to any organization where the affected devices have been transferred. Customers should contact Customer Service to place a replacement order and receive an RMA number and return all affected devices using the RMA number. For any questions call (512) 834-6255. |  | 30 | 30 units | Worldwide Distribution - USA including AR, AZ, CA, LA, MD, NY, TX, UT and Puerto Rico and the countries of United Kingdom and Italy | Terminated | September 05, 2012 | 315 | 30
Z-0351-2012 | 60347 | No | FCK | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The EnCorÂż biopsy probe, designed for use with EnCorÂż and EnCorÂż Enspire Vacuum Assisted Biopsy (VAB) driver systems. |  |  | November 03, 2011 | 2011 | Bard Peripheral Vascular Inc | This recall was voluntarily initiated by the firm due to the product 's damaged packaging trays; cracked EnCor (Trademark) packaging trays, plastic PETG tray the probe is packaged in was cracked. | Packaging | The firm sent the customer notification letter on 11/07/11, via FedEx with proof of delivery notice. The letter notified customers of the reason for recall that units from the affected lot may have packaging damage resulting from transportation. The letter instructed customers to check all inventory locations , remove any identified product, to not use the affected product, and return the affected product to Bard Peripheral Vascular, Inc., 1415 W. 3rd Street, Tempe, Arizona 85281. The letter also instructed customers to fill out the recall and effectiveness check form and fax it to 1-800-894-6772. Customers were instructed to call 1-800-321-4254, Option #2 Ext 2727 for a return authorization number to facilitate the expedient return of the product.For questions regarding this recall call 480-894-9515 |  | 64570 | 64570 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Europe, LAPAC, and Australia | Terminated | March 29, 2012 | 147 | 64570
Z-0362-2012 | 60416 | No | LNX |  |  | General Hospital | 510(k) | SoftReports version 1.1.8.2.4 with Hot Fixes 1.7199.7 or 1.7200.7. |  |  | October 25, 2011 | 2011 | SCC Soft Computer | On 10/25/2011 SCC Soft Computer, Clearwater, FL initiated a correction on the SoftReports version 1.1.8.2.4 with Hot Fixes 1.7199.7 or 1.7200.7. When SoftReports Designer opened a layout, the system did not load the layout with the previously saved value in the NewPage property. This caused multiple patient records to be printed on the same page. All affected clients were notified on the issue on | Software design | SCC Soft Computer sent a "Correction Notification" letter dated October 25, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The hot fixes that were previously loaded which introduced this issue were corrected and delivered to all affected customers as a mandatory correction in hot fixes 1.7199.8 and 1.7200.8 and higher. |  | 5 | 5 | Nationwide Distribution. | Terminated | March 16, 2012 | 143 | 5
Z-0364-2012 | 60411 | [6 empty] | TrueBeam and TrueBeam STx, Model Number: H19, Ref/FSCA identifier: CP-06381 are intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body where radiation treatment is indicated. |  |  | October 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the TrueBeam and TrueBeam STx systems where, under certain tuning conditions, the electron beam emerging from the bend magnet may have an elongated spot shape. | Device Design | On 11/04/2011, an Urgent Medical Device Correction Letter was distributed to all affected users, with a description of the problem and user corrective action steps. Varian is developing a corrective action for this issue. A customer service representative will contact you when the correction is available to schedule installation on your system. Questions regarding the letter should be directed to the Varian Oncology Help Desk Contact Information at USA and Canada: 1-888-827- 4265 or Europe +41-41-749-8844. |  |  |  | Worldwide distribution including the US and Internationally. | Terminated | July 25, 2012 | 268
Z-0360-2012 | 59886 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography System, Model #728231, 510(k) #K 033326, Serial Numbers: 95504, 91003, 300030, 95130, and 95157. |  |  | September 06, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were logged during in-house testing and external validation testing of Brilliance 64 and Ingenuity CT software version 4.0 Beta 2. The non-conformances are due to software errors identified in these device systems. | Software design | Philips Healthcare released a 'Field Change Order' (FCO #72800542) on September 15, 2011. Philips sent an URGENT - Medical Device Correction letter dated October 10, 2011, to all affected customers. The letter listed the affected device systems; described the nature of the software problem; and any possible hazard inherent to the situation. The letter was mailed by Philips to their customers on October 27, 2011. The letter provided information telling the customer how to identify the affected products and what action to take. In this instance, the customer was asked to review the Customer Release Notes which were provided to each customer location once the software update was installed. The letter notified the customer that a Philips representative would be installing a software upgrade in order to address the referenced software issues free-of-charge, through a Field Change Order. The letter closes by providing information which all customers in Canada and North America can use to contact the firm through a toll-free telephone number 1-800-722-9377, option 5; enter side ID or follow the prompts.For questions regarding this recall call 440-483-4918. |  | 5 | 5 units | Worldwide Distribution - USA including Ohio and the countries of Australia, China, Belgium, Israel, and Sweden | Terminated | March 15, 2013 | 556 | 7
Z-0353-2012 | 60338 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Maquet Servo-S ventilator system |  |  | September 19, 2011 | 2011 | Maquet Inc. | Servo-I and Servo--S ventilator systems are equipped with a specific revision of a printed circuit board (PC1772 revision 22) that includes a component that has generated a number of nuisance alarms due to technical error message 24 (TE 24). | Vendor change control | The firm, Maquet Getinge Group, sent an " PRODUCT RECALL URGENT - MEDICAL DEVICE CORRECTIVE ACTION" letter dated September 19, 2011, to its customers. The letter describes the product, problem and actions to be taken. Maquet service team will contact the customers within two weeks of receiving this letter to schedule an on-site service of the affected equipment. If you have any questions, call technical support at our toll free number at 1-888-627-8383. |  | 26 | 26 units | Nationwide distribution. | Terminated | July 22, 2015 | 1402 | 202
Z-0356-2012 | 60300 | No | FCW | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Integra Surgical Luxtec MLX Light Source |  |  | October 14, 2011 | 2011 | Integra LifeSciences Corp. | A loose screw may result in shifting of components and improper function of the turret or actuator assembly. | Process control | Integra sent an URGENT: Field Corrective Action letter dated October 14, 2011 to all affected customers via FedEx. The letter identified the affected product, the problem, and the actions to be taken with instructions to immediately inspect their inventory for potentially loose screws within the affected unit(s). If the unit(s) were observed with loose screws, the unit(s) were to be returned to Integra Surgical for repair and servicing. Customers were instructed to complete the Acknowledgement and Return Form per the instructions on the form. For questions call Integra Surgical Customer Service at 1-800-431-1123. |  | 66 | 66 units | Worldwide Distribution - USA (nationwide) including the states of: Arizona, California, Colorado, Florida, Indiana, Mississippi, New Jersey, New York, North Carolina, Ohio, Pennsylvania, Montana, Washington, Wyoming, and Puerto Rico, and the countries of: Australia, Germany, Italy, Japan, Netherlands, New Zealand, Pakistan, Sweden, Switzerland, and Turkey. | Terminated | March 08, 2012 | 146 | 97
Z-0383-2012 | 34843 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Product is 40mm Flexible Tube. ***NON-STERILE PRODUCT***CAUTION: FEDERAL LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN." |  |  | February 17, 2006 | 2006 | Walter Lorenz Surgical, Inc. | Lorenz Engineer discovered flex tube could break if bent too much. | Device Design | W. Lorenz Surgical notified consignees of recall by an Urgent Advisory Notice letter. Consignees were requested to acknowledge receipt of recall and provide product disposition. Initial letter was sent on March 2, 2006, March 15, 2006, follow-up letters sent on April 4, 2006 and May 15, 2006. For additional information call (904) 741-4400 or fax (904) 741-4500. |  | 76 | 76 pieces | Worldwide Distribution - USA (nationwide) including the states of: GA and TX and the country of Japan. | Terminated | December 12, 2011 | 2124 | 182
Z-0468-2012 | 51146 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Sterilizer Hydrogen Peroxide Gas Plasma Sterilizer Product Code 10133. |  |  | December 10, 2008 | 2008 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) has discovered a defect in the barcode labeling of the STERRAD NX System Cassettes 10133, Lot #08I032. This is the only lot number affected by this action. The barcodes for Lot #08I032 were printed improperly and are not capable of being read by the STERRAD unit. | Error in labeling | Advance Sterilization Products sent an "URGENT: PRODUCT RECALL" letter dated December 10, 2008 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The letter instructs customers to examine inventory and return all sealed cassettes, using an enclosed prepaid UPS return label to: Stericycle, 2670 Executive Drive Suite A, Indianapolis, IN 46241. Additionally, a Business Reply Card and packing slip were included with the letter for customers to complete and return. Questions or further assistance was directed to ASP Customer Care Center at 888-783-7723, option 2. |  | 776 | 776 cases | Nationwide Distribution. | Terminated | December 16, 2011 | 1101 | 776
Z-0487-2012 | 56540 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Harmonie software in use with: |  |  | April 16, 2009 | 2009 | Stellate Systems | The following implied or stated indications in Stellate labeling with Harmoniae software have not yet been cleared by the Food and Drug Administration (FDA) for sale in the United States of America. | No Marketing Application | Stellate Systems sent a "DEVICE CORRECTION" letter dated April 6, 2009 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The firm provided information on how to verify if customers had the affected product. Software updates will be available to all customers in possession of the affected product. Customers were instructed to call the Customer Support line at 1-888-742-1306 with any questions or concerns they may have regarding the notification. |  | 1041 | 1,041 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Canada, Denmark, Greece, Hong Kong, Hungary, Italy, Japan, Mexico, Norway, Rep. DE Panama, Spain, Switzerland, Taiwan ROC, and Turkey. | Terminated | December 27, 2011 | 985 | 1041
Z-0569-2012 | 57863 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Appraise/Heritage Combo A1C &amp; Blood Test Kit, Product K8300, each kit packed in a bag or envelope and contains 1 Triad Medium Alcohol Prep Pad, Cat. No. 10-3001, NDC 50730-3001-0, and blood specimen collection supplies, forms, and instructions. |  |  | January 09, 2011 | 2011 | Heritage Labs International LLC | Kits contain recalled Triad alcohol pads | Material/Component Contamination | Heritage Labs sent a Device Notification and Safety Alert letter dated January 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Heritage Labs sent replacement alcohol prep pads from a different supplier not affected by the recall to user of Heritage Labs kits. Heritage Labs recommended the replacements be on "as needed" basis. Upon opening kit, users should immediately dispose of sealed TRIAD brand alcohol pads. Users are to use the replacement wipes instead of the wipes contained in the kits. For any questions call (888)764-4120. |  | 4114 | 4,114 kits | Worldwide Distribution - USA (nationwide) and Germany | Terminated | December 26, 2012 | 717 | 683376
Z-0393-2012 | 60386 | No | HEB | Unscented menstrual tampon. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Kotex Natural Balance Security tampons, regular absorbency, unscented. packaged in 18 count or 36 count retail containers. Consumer Services, Kimberly-Clark Corp. Dept. KNBTR-18, P.O. Box 2020, Neenah, WI 54957-2020 USA. |  |  | November 03, 2011 | 2011 | Kimberly-Clark Corporation | A limited number of Kotex Natural Balance Security Unscented Tampons, regular absorbency, are being recalled after a raw material contaminated with a bacterium, Enterobacter sakazakii, was found when testing the plunger portion of the applicator. | Material/Component Contamination | Kimberly-Clark sent an Urgent Product Recall letter, dated 11/09/2011, via FedEx to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify, quarantine, and discontinue use immediately. Customers were asked to complete the response form and fax to +1 (920) 380-6467. A Kimberly-Clark representative would contact them regarding return of the product, replenishement and/or account credit. For questions regarding this recall call +1 (920) 216-8152. |  | 1400 | 1400 Cases (216 tampons per case) | Nationwide Distribution | Terminated | December 13, 2012 | 406 | 1400
Z-0388-2012 | 60418 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea (available in various configuration as Vitrea Enterprise Suite). Vitrea Enterprise Suite 1.2, Vitrea Enterprise Suite 1.3, Vitrea enterprise Suite 6.0 and Vitrea Enterprise Suite 6.1. |  |  | November 11, 2011 | 2011 | Vital Images, Inc. | Vital Images has found two issues that may display incorrect Standard Update Values (SUV) when viewing certain PET/CT images. This issue affects the VitreaCore PET/CT feature for Vitrea Enterprise Suite 1.2, 1.3, 6.0, 6.1, including all updates to those versions. | Software design | Vital Images, Inc. sent an URGENT DEVICE RECALL letter dated November 11, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to inform all potential users of the issues. Customers were to be contacted by their Service Account Manager to discuss and schedule resolutions.For questions customers were to contact Vital Images Customer Support with any questions toll free at 1-800-208-3005, in Europe +31 (0) 70 4135801, in Asia t +86 10 8588 1800, or online at support@vitalimages.com. |  | 338 | 338 (233 US) | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, ND, OH, OK OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of IRELAND, FRANCE, UNITED KINGDOM, SWITZERLAND, AUSTRALIA, SLOVENIA, GERMANY, NETHERLANDS, BELGIUM, PAKISTAN, BRAZIL, ARGENTINA, SWEDEN, QATAR, CHINA, FINLAND, ITALY, DENMARK, PORTUGAL, REPUBLIC OF KOREA, SPAIN, and CANADA | Terminated | December 03, 2013 | 753 | 338
Z-0523-2012 | 60423 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | MAMMOMAT Inspiration |  |  | September 15, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Lowering the swivel arm with "object table" while an operator's or patient's body parts are underneath the "object table" may cause the "object table" to collide with and injure the operator or patient. Therefore attention should be paid when adjusting "object table" height to avoid a possible risk of injury to the operator or the patient.The movement of the swivel arm may also cause an injur | Use error | Siemens Medical Solutions USA, Inc sent a Safety Advisory Notice dated January 8, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.The notification included an Addendum to the User Manual.We appreciate your understanding and cooperation with this safety advisory and ask you to immediately instruct your operating personnel accordingly. Please ensure that this safety advisory is placed in the System's Operator Manual. Your operating personnel should maintain awareness over an appropriate defined period. Further questions please call (610) 219-4834. |  | 12 | 12 | Nationwide Distribution including the states of FL, MD, MI, NE, NY, NC, ND, OH, and SC. | Terminated | March 05, 2014 | 902 | 12
Z-0522-2012 | 60426 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo.plaza |  |  | September 21, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Potential malfunction when using syngo.plaza. On syngo.plaza workplaces large studies can be loaded into the viewer if transferred from a DICOM node or imported even while image transfer is not yet completed. In these situations more images of that study may exist which are not shown in the syngo.plaza viewer. | Software design | Siemens Medical Solutions USA sent a Customer Safety Advisory Notice dated March 10, 2011 to the affected customer. The letter identified the product the problem and the action needed to be taken by the customer.The letter informed the customer of the potential issues and provides additional information on the implemented resolution.For further questions please call (610) 219-4834. |  | 1 | 1 | Nationwide Distribution New York | Terminated | January 31, 2014 | 863 | 1
Z-0508-2012 | 60325 | No | HXK | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Offset Wire Fixation Bolt, catalog number 8180-50-006. The bolts are used with the DePuy Ace-Fischer External Fixation System and are packaged in a flexible polyethylene bag. |  |  | November 08, 2011 | 2011 | DePuy Orthopaedics, Inc. | DePuy has received three complaints concerning oversized bolt threads in the Offset Wire Fixation Bolts used in the Ace- Fischer External Fixation System. This could lead to the inability to properly thread the nut onto the bolt. | Component change control | On 11/9/2011 DePuy sent "Medical Device Recall Notice" letters to Distributors and to direct account hospitals. Recall notices included instructions to cease use or distribution of recalled products and to return them to DePuy in Warsaw, IN via distributors and sales representatives. The product is going to be discontinued and no reconditioning or reworking for the product is planned. For questions about recall information provided, please contact Katie Seppa, Manager of Customer Quality, 574-372-7333 (M-F; 8 am - pm EST.) |  | 890 | 890 | Nationwide Distribution | Terminated | March 07, 2013 | 485 | 890
Z-0350-2012 | 60434 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | AVAIRA AQUAFORM Comfort Science enfilcon A Daily Wear SPHERE Soft Contact Lenses (disposable), Rx only, STERILE --- CooperVision: Hamble, SO31 4RF, UK -- Scottsville, NY 14546 USA -- Juana Diaz, PR 00795 USA --- Device Listing # D033736. Packaged: 6 Soft Contact Lenses per unit carton or trial product 3 pack. |  |  | November 15, 2011 | 2011 | Coopervision Inc. | Presence of silicone oil residue on Avaira Sphere Soft Contact Lenses. | Process control | CooperVision Inc. issued a press release on November 15, 2011. An "URGENT: MEDICAL DEVICE RECALL- AVAIRA SPHERE CONTACT LENSES" letter dated November 16, 2011 was sent to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their inventory, stop further distribution of the affected product, contact their patients and to advise their patients to check the firm's website to confirm they have the affected product. Additionally, a response form was included with the letter for customers to complete and return via fax at 866-484-1048. Contact Coopervision at 800-341-2020 for questions regarding this recall |  | 4877897 | Domestic: approx. 3,155,781 lenses; Foreign: estimated 1,722,116 lenses | Worldwide Distribution--USA (nationwide) including Puerto Rico. | Open, Classified |  |  | 4877897
Z-0444-2012 | 60455 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab used with SA INST 4.0.2.0 to 4.0.2.51, 4.0.3.0 to 4.0.3.13, 4.5.0.0 to 4.5.0.9, 4.5.2.0 to 4.5.2.8 and SA INST 3.1.3.0 to 3.1.3.29, 3.1.6.0 to 3.1.6.17, 4.0.0.0, 4.0.1.0 to 4.0.1.32. |  |  | July 26, 2011 | 2011 | SCC Soft Computer | SCC Soft Computer, Clearwater, FL, initiated a correction for the SoftLab used with SA INST. A test performed on the Roche Cobas Integra series of instruments when used with the GenInst/Dispatch listening program could cause incorrect flagging for a test result posted to the lab database. | Software design | The firm, SCC Soft Computer, sent a "Correction Communication" letter on July 26, 2011 to customers/consignees. The letter describes the product problem and actions to be taken. The firm will correct the software to recognize and support all known formats of the resulting date and time sent by an instrument. The software is available via a recommended patch or hot fix to the client's current release. The customers were asked to acknowledge receipt of the task and grant permission for SCC to correct the affected configuration files. If you have any questions, call (727) 789-0100. |  | 265 | 265 | Worldwide Distribution - USA (Nationwide) and Country of: Canada. | Terminated | May 21, 2013 | 665 | 265
Z-0389-2012 | 60344 | [6 empty] | Magnus Hybrid operating table columns 1180.01A1 and 1180.01B1 |  |  | September 19, 2011 | 2011 | Maquet Inc. | A component of the Magnus hybrid OR system fails to respond to a position command. | Device Design | Maquet sent a PRODUCT RECALL URGENT MEDICAL DEVICE CORRECTIVE ACTION letter dated September 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. To avoid the potential of a malfunction occurring prior to their scheduled Service procedure, customers were instructed to ensure that the MAGNUS table-top is positioned at the "0" pivot position and that the MAGNUS column pivot brake is secured. If it is necessary to pivot the MAGNUS table-top prior to procedure start, ensure that the column pivot brake is completely released and then completely secured when returned to "0" pivot position. Customers were instructed to ensure that all user of the affected product and other necessary persons are informed of the product correction notice. For technical support or questions call 1-888-627-8383 |  | 45 | 45 UNITS | Worldwide Distribution - USA including VA, LA, MO, NC, MD, and AZ and the countries of AE, AU, CA, CH, CN, DE, DK, ES, IE, IL, SP, PL, RU, and TH. | Terminated | December 08, 2011 | 80 | 45
Z-0601-2012 | 60367 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter Âż HmX Hematology Analyzer with Autoloader, |  |  | October 13, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman has identified a problem with Integrated Circuit chips used in different boards throughout the analyzer. | Nonconforming Material/Component | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated October 13, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were instructed to contact Beckman Coulter Service at (800) 526-7694 if they were experiencing failures. Customers not experiencing failures, but their system may have the suspect chips, were told they would be contacted by Beckman Coulter Service to correct their problem. Customers were also instructed to complete and return an enclosed Response Form. |  | 98 | 98 units total | Worldwide Distribution. | Terminated | June 01, 2012 | 232 | 98
Z-0394-2012 | 60440 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OR table TITAN, Surgical Table. |  |  | January 21, 2011 | 2011 | Trumpf Medical Systems, Inc. | Product is being recalled due to unintended movement on the Titan Tables. | Device Design | Trumpf Medical Systems sent an initial notification to customers dated January 24, 2011. A subsequent "URGENT SAFETY NOTICE-UPDATE" letter dated May 26, 2011 was sent to all affected customers. The letter described the product, problem, and actions to be taken by the customers. A Corrective Action Plan includes a software update to be installed by a Trumpf Technical Service Representative. Contact Technical Service at 888-474-9359 for questions regarding this notice. |  | 94 | 94 beds total for all types. | Worldwide Distribution. | Terminated | August 13, 2012 | 570 | 98
Z-0417-2012 | 60447 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OSS INTERLOK BOWED IM STEM W/SCREW 14 MM X 150 MM TI 6AL 4V ALLOY. |  |  | November 01, 2011 | 2011 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that Part 150367, (OSSCemented 1M Stem 13x150) was incorrectly labeled as part 150368 (OSS Cemented 1M Stem14x150). | Mixed-up of materials/components | Biomet, Inc. sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated November 1, 2011 to the one distributor with the affected product. The letter described the product, problem, and actions to be taken by the customer. The letter instructs the customer to immediately locate and remove the affected product from inventory. Additionally, the customer was asked to complete and return a Response Form via fax to 574-372-1683. Call 574-372-3983 for questions related to this notice. |  | 4 | 3 units &amp; 1 pending release | Distributed only in The Netherlands. | Terminated | September 10, 2012 | 314 | 4
Z-0454-2012 | 60359 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Dental cement under the brand name Cement-It and Natural Elegance (private label of Cement-It). |  |  | March 17, 2011 | 2011 | Kerr/pentron Dba Kerr Corporation And Pentron Clinical | The firm initiated a recall because the product sets faster than specified in the "Directions for Use" and in release specifications. | Other | The firm, Pentron Clinical, sent an " URGENT: MEDICAL DEVICE RECALL" letters dated April 4, 2011 and April 20, 2011 via USPS 1st class mail to its customers. The letter describes the product, problem and actions to be taken. The letter instructed customers to review the table with affected product lots and determine if the affected products was in their inventory; return the affected product for a replacement at no charge; contact Pentron Clinical Customer Service at (800) 551-0283; 1-800-243-3969 ext. 677 or email: aestridge@pentron.com directly to handle the arrangements of a quick return and replacement; and complete and fax back the enclosed Recall Return Form at 1-877-677-8844. The letter also request the customer to identify and recover the affected product lots that may have been shipped to thier customers.If you have any questions, contact Pentron Clinical Customer Service at (800) 551-0283. |  | 157 | 157 | Worldwide distribution: USA (nationwide) including states of: PA, IL, NV, WA, OH, TN, WI, NY, NJ, CT, IN FL, MI, OK, NE, HI, VA, CA, and TX; and countries of : Canada, Egypt, Isreal, Netherlands, Sweden and UK. | Terminated | January 22, 2014 | 1042 | 157
Z-0370-2012 | 60441 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Eclipse Release 10 with RT Chart Release 10 (Client Build 10.0.28 or Build 10.0.34), An |  |  | October 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with Varian Eclipse and RT Chart where adding a reference pointin RT Chart and editing the field dose contributions may change the Prescribed Dose Per Fractionin Eclipse. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" dated October 31, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Varian has corrected the issue and will contact customers to schedule installation of the corrected version of the system. Contact the Varian Oncology Help Desk at 1-888-827-4265 for questions regarding this notice.. |  |  |  | Worldwide Distribution. | Terminated | June 06, 2014 | 949
Z-0596-2012 | 60504 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, pre-filled flush solution, for IV Flush only, sterile fluid path; 10 mL Single-Use Syringe with Male Luer Lock, 100 syringes per pack, 4 packs per case; Made in Costa Rica, Hospira, Inc., Lake Forest, IL 60045 USA; list no. 1978-20 |  |  | November 18, 2011 | 2011 | Hospira Inc. | There is particulate matter in the fluid pathway, which has been identified as the same material as the rubber tip of the syringe plunger. | Material/Component Contamination | Hospira sent a "URGENT DEVICE RECALL" letter dated November 18, 2011, to all affected customers via UPS. The letter identified the product the problem and the action needed to be taken by the customer.The customers were instructed to check their inventory and immediately quarantine any affected product complete and return Reply Form via fax to 1-866-912-2512. Return affected product to Stericycle using the label provided with the letter and if they have distributed the product further, notify their accounts that received the product . Please contact Hospira Customer Care at 1-877-946-7747 or your Hospira representative for information regarding product availability.Hospira expanded the recall to include thirteen additional affected lots via letter dated January 31, 2012. The recall instructions remain the same as in the November 18,2011 letter. |  | 5817600 | 5,817,600 syringes | Nationwide distribution including states of: Alabama, Alaska, Arizona, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Mississippi, Missouri, Nebraska, New Hampshire, New Jersey, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, West Virginia and Wisconsin. | Open, Classified |  |  | 5817600
Z-0488-2012 | 60452 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab GUI versions: 3.18.1.0-3.18.1.21; 3.19.1.0-3.19.1.25; 3.19.2.0-3.19.2.21; 4.0.0.2; 4.0.1.0-4.0.1.16; 4.0.2.0-4.0.2.10; 4.0.3.0-4.0.3.16; 4.0.4.0-4.0.4.6; and 4.5.0.0-4.5.0.8. |  |  | May 03, 2011 | 2011 | SCC Soft Computer | On 05/03/2011, SCC Soft Computer, Clearwater, FL, initiated a correction on the following SoftLab GUI versions: 3.18.1.0-3.18.1.21; 3.19.1.0-3.19.1.25; 3.19.2.0-3.19.2.21; 4.0.0.2; 4.0.1.0-4.0.1.16; 4.0.2.0-4.0.2.10; 4.0.3.0-4.0.3.16; 4.0.4.0-4.0.4.6; and 4.5.0.0-4.5.0.8. A client reported that results were not flagged correctly. This can occur when a test is ordered in Order Entry, then collected | Software design | SCC Soft Computer sent a "CORRECTION COMMUNICATION" notification letter to all affected customers. The letter describes the product, Risk-to-Health notification task, potential hazard, and corrective action. The customers were asked to acknowledge receipt of the task and grant permission for SCC to make the required and recommended changes. For any questions, call (727) 789-0100. |  | 220 | 220 | Worldwide Distribution - USA (nationwide) and the country of: Canada. | Terminated | August 07, 2013 | 827 | 220
Z-0447-2012 | 60421 | [6 empty] | Box labeled in part, "REF: 72-1000***IQ SYSTEM IQ INTELLIGENT DRIVER 16.51CM (6.5")(L) RABALON, STAINLESS STEEL, GOLD PLATED***NON-STERILE PRODUCT*** MADE IN THE PHILIPPINES*** BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***" |  |  | July 08, 2011 | 2011 | Biomet Microfixation, Inc. | Biomet Microfixation, Inc., Jacksonville, FL, recalled their IQ Intelligent Device, 72-1000 due to failure of the front bit sensor to detect the correct bit that is installed. When a screw is inserted with the device in drill mode, it is possible for the screw or bone to strip. | Component design/selection | Biomet Microfixation called, emailed, or sent an Urgent Medical Device Recall letter dated July 8, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the remaining device(s) at their facility and return immediately to PIVOT for the required software upgrade. Customers were also instructed to return the attached reconciliation form to:PIVOT International10916 Strang Line RoadLenexa, KS 66215Attn: Julie LawrenceFor any questions regarding this recall call 1-800-874-7711 or 904-741-4400, ext 9468. |  | 161 | 161 devices were distributed. | Worldwide Distribution - USA AZ, CA, CO, DC, FL, GA, HI, ID, IL, IN, KS, LA, NC, NH, NY, OK, OR, PA, TN, TX, and WI and the countries of Argentina, Brazil, France, Italy, Japan, South Africa, and Thailand. | Terminated | March 09, 2012 | 245 | 161
Z-0462-2012 | 60461 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | syngo RT Therapist and syngo RT Therapist Connect/MOSAIQ OIS. |  |  | November 02, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The software release will bring current Installed Base system to the forward production version. It is a result of the final system testing addressing several safety issues that are not critical for the safe use of the device and multiple performance defects. | Software design | Siemens sent Customer Information letters and Update Instructions TH024/11/S for the installation of software release 4.2.108 beginning on November 2, 2011 to all affected consignees. Sites with active devices for the Oncology system were sent the current Update Instructions TH014/11/S for the installation of software release 4.2.108 and Customer Information letter beginning November 22, 2011.Consignees were asked to include the Customer Information in their syngo RT Oncologist System Owner Manual chapter "Safety Advisory Letters" where it should remain. For questions regarding this recall call 925-602-8157. |  | 147 | 85 affected customers RT Therapa[ist, 62 customers Oncology | Worlwide Distribution - USA (nationwide) including UT, OH, TN, FL, WI, IN, KY, CA, PA, SD, NY and the countries of India, Germany, United Kingdom, Australia, Poland, Japan, Croatia, India, Columbia, Belgium, Puerto Rico, China, Sweden, South Africa and Czech Republic. | Terminated | March 21, 2012 | 140 | 147
Z-0446-2012 | 59231 | No | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Osteopore PCL Scaffold Bone Filler: Osteopore Osteoplug-C, Osteoplug, and Osteomesh. |  |  | April 27, 2011 | 2011 | OSTEOPORE INTERNATIONAL PTE LTD | Product sterility may be compromised. | Nonconforming Material/Component | Osteopore sent an Urgent Voluntary Product Recall notice to their consignees by letter dated April 27, 2011. |  | 330 | 330 | Nationwide distribution - one distributor in TX. | Terminated | June 06, 2012 | 406 | 330
Z-0461-2012 | 60542 | No | NYQ |  |  | Immunology | PMA | SPOT-Light" HER2 CISHÂż Kit, catalog #84-0150, 20 tests/kit, labeled in part ***Invitrogen Corporation 7335 Executive Way Frederick MD 21704 www.invitrogen.com*** |  |  | November 14, 2011 | 2011 | Life Technologies Corporation | In vitro diagnostic reagent may be contaminated with a fungal contaminant. | Material/Component Contamination | Life Technologies sent an URGENT DEVICE RECALL letter dated November 14, 2011 by mail to all affected customers. The letter identified the product, description of problem, and actions to be taken by the user. Customers were instructed to examine inventory immediately, discontinue use, quarantine the material and return remaining inventory. Customers were advised to complete the Customer Response Sheet included for recovering inventory data and fax back to (+1) 716-774-6727 Technical Support or e-mail to techsupport@lifetech.com. Customers who had distributed the affected kits outside of their facility were instructed to immediately notify their customers of the recall. For questions call Technical Support, at (+1) 800-955-6288 opt. 2 then opt. 3. |  | 29 | 29 kits | Worldwide Distribution - USA (nationwide) in the states of CA and NV and the countries of: Canada, Venezuela, and the UK. | Terminated | February 10, 2012 | 88 | 29
Z-0466-2012 | 60459 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftMic GUI versions: 4.0.3.0-4.0.3.15, 4.0.4.1-4.0.4.7. |  |  | August 25, 2011 | 2011 | SCC Soft Computer | On 08/25/2011, SCC Soft Computer initiated a correction on the following SoftMic GUI versions: 4.0.3.0-4.0.3.15, 4.0.4.1-4.0.4.7. For clients using SoftMic Result Entry, when a certain series of keystrokes on the test grid are performed while entering results and statusing a test, the test status may not be maintained upon saving. The status may be removed or copied from another test on the order | Software design | SCC Soft Computer initiated a Field Correction dated August 25, 2011. to all affected customers.Customers with the affected software have been notified and are scheduling receipt of the hot fix..Further questions call (727) 789-0100. |  | 144 | 144 | Worldwide Distribution -- USA (nationwide) and Canada. | Terminated | October 03, 2014 | 1135 | 144
Z-0445-2012 | 60565 | No | NAA | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | CrystalensÂż Accommodating Posterior Chamber Intraocular Lens, AT-50AO. |  |  | November 14, 2011 | 2011 | Bausch and Lomb, Incorporated | Bausch + Lomb voluntary initiated the recall due to a physician complaint of haze on the Crystalens AOwhich was visualized using an angled beam from a slit lamp during the post-operativeexamination. | Nonconforming Material/Component | Bausch + Lomb sent an "URGENT- MEDICAL DEVICE RECALL" letter dated November 14, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to remove the affected products from inventory and return the lenses to the firm. Additionally, a Recall Acknowledgement Form was attached to the letter for customers to complete and return with the affected products. Contact Bausch + Lomb at 949-521-7895 for questions regarding this recall.. |  | 62 | 62 units | Nationwide Distribution | Terminated | May 23, 2012 | 191 | 62
Z-0457-2012 | 60571 | Yes | LWS |  |  | Cardiovascular | PMA | St. Jude Medical Riata (8Fr), Models: 1560, 1561, 1562, 1570, 1571, 1572, 1580, 1581, 1582, 1590, 1591, 1592 |  |  | November 28, 2011 | 2011 | St Jude Medical CRMD | St. Jude Medical has confirmed the failures associated with all cause insulation failure on the St. Jude Medical Riata and Riata ST Silicone Endocardial Defibrillation Leads with specific emphasis on externalized conductors. | Device Design | St. Jude Medical sent a Medical Device Advisory Important Product Information Update letter dated November 28, 2011, Physician Communication dated December 15, 2010, Product Performance Report Extracts, and a Physician Device Advisory Notification to physicians by FedEx priority overnight service for receipt on November 29, 2011. The letters provides the customers with an explanation of the problem, the product, and the actions to be taken. The issuance of the recall letter is in conjunction with St. Jude's recent release November 2011 Product Performance Report (available online at http://sjmprofessional.com). The letter also explains about the Lead Peformance Summary, Root Cause, Clinical Implications, Rate of Occurrence from Complaints and Returns, New Peer Reviewed Publications, Data on Optim Insulated Leads (Durata and Riata ST Optim), and Recommendations and Mitigations. Physicians with any questions or concerns can contact their local St. Jude Medical representative or the Technical Services Department at (800) 722-3774. |  | 128000 | 128,000 units total | USA (Nationwide) and worldwide to: Austria, Belgium, Bulgaria, Canada, Czech Republic, Denmark, Estonia, Finland, France, Germany, Hungary, Iceland, Italy, Romania, Russia, Netherlands, Norway, Poland, Portugal, Russia, Serbia, Slovakia, Spain, Sweden, Switzerland, The Netherlands, UK, Australia, China, Hong Kong, Indonesia, Japan, Korea, Malaysia, Philippines, Singapore, Taiwan, Thailand, Vietnam, New Zealand, Argentina, Brazil, Colombia, Costa Rica, Ecuador, Mexico, Puerto Rico, Uruguay, Egypt, Greece, India, Iran, Israel, Joran, Kuwait, Lebanon, Oman, Qatar, Saudi Arabia, South Africa, Turkey, Malta, Cyprus, Iraq, Egypt, Lybia, Yemen, and U.A.E. | Terminated | May 09, 2012 | 163 | 128000
Z-0520-2012 | 60592 | No | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | Lifescreen Apnea is Lifescreen ECG Screening System, Version 3.31. |  |  | November 11, 2011 | 2011 | Del Mar Reynolds Medical, Ltd. | When Lifescreen is left open and a new recording is loaded from Sentinel Cardiology InformationManagement System, customers can get corruption in the data or a wrong ECG after the first32 hours. Both recordings have to be longer than 32 hours for this to happen. | Software design | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated November 11, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to close Lifescreen every time before loading another recording from the Sentinel database. Spacelabs is in the process of upgrading software and will contact customers to arrange for an update to affected systems. If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 523 | total 523 units and 171 kits (34 units and 7 kits in US; 489 units and 164 kits outside US) | Worldwide Distribution-USA (nationwide) including the states of AR, CA, CO, CT, FL, GA, KY, LA, MA, MD, MI, MN, NC, NY, OH, TN, TX, UT, WV, and PUERTO RICO and the countries of AUSTRIA, BELGIUM, CHILE, CHINA, COLOMBIA, COSTA RICA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, KUWAIT, MALTA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, Slovakia, SPAIN, SWEDEN, SWITZERLAND, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM | Terminated | June 05, 2012 | 207 | 523
Z-0489-2012 | 60502 | No | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur CA19-9 Assay |  |  | November 11, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc | Positive bias of 14% and higher for some individual points for CA 19-9 results generated on the ADVIA Centaur CP system when they are compared to CA 19-9 assay results generated on either the ADVIA Centaur or ADVIA Centaur XP systems | Nonconforming Material/Component | Siemens Healthcare Diagnostics issued a Urgent Device Notification email dated November 11, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed the customers conduct a look back if they are using the ADVIA Centaur CA 19-9 interchangeably with the ADVIA Centaur or ADVIA Centaur XP systems and the ADVIA Centaur CP system and to review the contents of the recall notice with their Laboratory Director. A confirmation fax-back form is included to ensure that all customers worldwide have been notified and have complied with the mandatory actionFor further questions please call (508) 359.3825. |  | 21593 | 50 Test Kits =21,160; 250 Test Kits = 21,593 | Worldwide Distribution -- USA (nationwide) and the countries of Australia, Malyasia, Korea, Hong Kong, Singapore, Myanmar, Mexico, Egypt, Taiwan, South Africa, Brazil ,India, Japan, Argentina, Chile, Peru, Germany, Malaysia, Saudi Arabia, and Canada. | Terminated | April 12, 2013 | 518 | 21593
Z-0590-2012 | 59833 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DURAMAX 15.5 F x 36 cm Chronic Hemodialysis Catheter Set (With Cuff 31 cm from Tip), DuraMax Stacked Tip 36 cm Str. Basic Kit, Catalog No./REF 10302806, STERILE --- AngioDynamics, Inc. 603 Queensbury, NY 12804 USA |  |  | January 24, 2011 | 2011 | Angiodynamics, Inc. | AngioDynamics received customer complaints for breakage of the tunneler sleeve during product use. | Nonconforming Material/Component | The firm, AngioDynamics, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 28, 2011 and response form via Certified Mail, Return Receipt Requested, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to identify and segregate the recalled lots; complete and return the enclosed DuraMax HemoDialysis Catheter Recall Reply Form via fax to: Attn: DuraMax HemoDialysis Catheter Recall Coordinator at 518-798-1360 (even if they don't have any product); and do not use the affected product, return and ship the recalled product to ANGIODYNAMICS (replacement product will be shipped upon receipt and confirmation of the returned product). If the customers are in immediate need for replacement product, they should fax the completed Recall Reply Form, along with a copy of the product label, to Customer Service at (518) 798-1360 or call us at (800) 772-6446, x1358 or x1363.If you have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please do not hesitate to contact your local sales representative or call ANGIODYNAMICS customer service at 1-800-772-6446 or email to: customerservice@angiodynamics.com. |  | 39 | 39 units | Nationwide distribution. | Terminated | February 02, 2016 | 1835 | 1124
Z-0519-2012 | 60518 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Smith &amp; Nephew, SPIDER2 Limb Positioner, including accessories Battery, foot pedal and charger |  |  | July 26, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | SPIDER2 may unlock (loss of traction) resulting inpotential movement of the patient when the device is within close proximity to an RF source. Lossof traction can potentially result in the unintended repositioning of the extremity and potential for tissue or cartilage damage | Device Design | Smith &amp; Nephew Advanced Surgical Devices Division sent an "URGENT- PRODUCT RECALL NOTIFICATION-URGENT" letter dated July 26, 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. Customers were instructed to complete and return a form via fax at 405-917-8650. Contact the firm at 405-917-8556 for questions regarding this recall. |  | 52 | 52 units | Worldwide Distribution-USA (nationwide) and the countries of Canada and France. | Completed |  |  | 52
Z-0459-2012 | 60525 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE and ONCOR linear accelerators with G31 mother board. |  |  | November 09, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Reports of gantry rotation without receiving a movement command caused by an overload within the drive signal power line. | Component design/selection | Siemens Medical Solutions USA, Inc. sent a Customer Information letter dated November 9, 2011, by federal express. The letter identified the product the problem and the action needed to be taken by the customer.The HW Update included information regarding the potential problem and provide instructions in the Safety update.Further questions call (925) 602-8157. |  | 13 | 13 active devices | Worldwide Distribution-- USA ( nationwide ) including the states of : MO, NC, OH, and OK and the countries of Turkey, Japan, Germany and Portugal | Terminated | April 15, 2013 | 523 | 13
Z-0606-2012 | 60626 | No | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Eius Unicompartmental Knee System |  |  | August 17, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Two National Joint Registries suggest that the EIUS Unicompartmental System is associated with a higher revision rate then the unicompartmental devices. | Other | Stryker Orthopaedics sent "URGENT PRODUCT RECALL" letters dated August 17, 18 and 25, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and return an attached Product Recall Acknowledgement Form via fax to 201-831-6069. Contact the firm at 201-972-2100 for questions concerning this recall. |  | 26648 | 26,648 units | Nationwide Distribution. | Terminated | March 26, 2014 | 952 | 26648
Z-0595-2012 | 60628 | No | LNR |  |  | Gastroenterology/Urology | PMA | Therakos Cellex Photopheresis System |  |  | September 01, 2011 | 2011 | Therakos, Inc. | There have been three confirmed reports of post reinfusion hemolysis or hemoglobinuria in patients treated on the Therakos Cellex Photopheresis System. | Other | Therakos sent an "URGENT PRODUCT CORRECTION NOTIFICATION" dated September 1, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Additionally, a Confirmation of Receipt was attached to the letter for customers to complete and return. Contact Therakos Technical Support at 1-877-865-6850 for questions regarding this notice. |  | 114 | 49 USA, 65 Foreign | Worldwide Distribution. | Terminated | March 27, 2014 | 938 | 114
Z-0456-2012 | 60636 | No | NZE | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | ThermoSuit Manufactured for Life Recovery Systems. HD, LLC, 170 Kinnelon Rd. Kinnelon, NJ 07405 USA |  |  | October 21, 2011 | 2011 | Life Recovery Systems HD, LLC | ThermoSuit System Model TSP-50 Pump/Controller may have corrosion on the temperature circuit board which may result in incorrect patient temperature reading. | Process control | Life Recovery Systems sent recall letters/return response forms dated 10/24/2011 to all consignees. Life Recovery Systems has expanded their recall to include additional units and to provide additional service to products out in the field. A new recall letter was sent to consignees on 4/27/2012. |  | 28 | 28 units | Nationwide Distribution including NY, NC, LA, NJ, AL. PA TX, KY, MO, MA, and TN | Terminated | January 24, 2013 | 461 | 28
Z-0509-2012 | 60587 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Benephit XT Infusion System, Targeted Renal Therapy Infusion System, Catalog/REF 70035, STERILE -- AngioDynamics Inc. 603 Queensbury, NY 12804 USA --- COMMON/USUAL NAME - Continuous flush catheter; CLASSIFICATION NAME: Continuous flush catheter; Device Listing # D069932 |  |  | December 05, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The Benephit XT Infusion System may contain hairline cracks in the sheath hub. | Nonconforming Material/Component | Angiodynamics Worldwide Headquarters sent a Urgent Medical Device Recall Letter and Response Forms dated December 2, 2011, to all affected customers via Certified Mail on December 5, 2011. The recall is being extended to the medical user level. Customers will be instructed not to use the product and to return it to AngioDynamics, Inc. in Queensbury, NY. A Return Authorization number (RA #) will be assigned to each customer in the recall notification letter and a UPS account will be provided for return freight. A reply form was provided with the recall notification letter to be completed by each consignee. ACTIONS TO BE TAKEN BY CUSTOMER / USER: (1.) Identify and segregate the recalled lots that are in your possession. (2.) Complete the enclosed Benephit XT Infusion System Recall Reply Form and fax it to the attention of the Benephit XT Infusion System Recall Coordinator at 518-798-1360. The form lists the catalog number, lot numbers, and quantity our records indicate your facility has received. It is important that even if you do not have any product remaining in your possession that you fill out the attached form noting zero quantity to be returned and fax the form to ANGIODYNAMICS. (3.) Ship the recalled product to ANGIODYNAMICS. Replacement product will be shipped upon receipt and confirmation of the returned product. Reference Return Authorization Number on the outside of the shipping box and include a copy of the Benephit Infusion System Recall Reply Form with your shipment. (4.) If you have distributed the affected product/lot, please inform your customers of this recall action immediately, and have them return the recalled units to you. --- Certified Mail will verify delivery to each domestic consignee. Non-respondents will be contacted by telephone or via e-mail. --- Corrective Action: Returned product will be quarantined and final disposition has not been determined at this time. A corrective action has been initiated by AngioDynam |  | 55 | 55 units domestically | Nationwide Distribution including the states of New York, Pennsylvania, Virginia, Kentucky, Iowa, Illinois, Missouri, Louisiana, Texas, and Utah. | Terminated | January 29, 2016 | 1516 | 55
Z-0597-2012 | 60456 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | Midmark M11 Ultra Steam Sterilizer, Distributed by: Midmark Corporation, 60 Vista Drive, Versailles, Ohio 45380. |  |  | October 24, 2011 | 2011 | Midmark Corportation | During and FDA inspection on 8/23/2011, it was discovered that the firm received numerous reports of the steam sterilizer malfunctioning. | Employee error | Midmark Corp. will begin sending customer communications out the week of January 9, 2012 to all affected customers. The letter will describe the product, problem, and actions to be taken by the customers. Additionally, a warning label and an illustration of where to place the label will be included with the letter. Customers may contact the firm at 1-800-643-6275. |  | 30255 | 30,255 units | Worldwide Distribution- USA (nationwide) including Puerto Rico and the countries of AUSTRALIA, BELGIUM, BRAZIL, CHILE, CHINA, EGYPT, FRANCE, GUAM, HONG KONG, HUNGARY, ICELAND, INDIA, KOREA, KUWAIT, MALAYSIA, MEXICO, NEW ZEALAND, POLAND, RUSSIA, SAIPAN MP, SAUDI ARABIA, SEOUL, SINGAPORE, TAIWAN, THAILAND, TURKEY, U.A.E., UNITED KINGDOM and WEST INDIES. | Terminated | June 20, 2014 | 970 | 30255
Z-0455-2012 | 60551 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Artiste Maverick Tint, Part Number: N36TB, Lot/Serial Numbers: 3416180, Product Description and Intended Use: The intended use of this device is for internal color characterization and includes 3 opaques for blocking out deep staining or metal substrates. |  |  | March 03, 2011 | 2011 | Kerr Corporation | The resin material separates when the product is extruded from the syringe. Although there are no safety issues related with this situation, please refrain from using the affected product. | Nonconforming Material/Component | Pentron Clinical sent an URGENT: Medical Device Recall letter dated April 20, 2011 to all affected customers. The letter described the product, problem and actions to be taken. Customers were instructed to examine their inventory and quarantine all product(s) subject to recall for return and exchange at no charge. In addition, the letter states that if the product was further distributed to contact their customers and notify them of the recall. Customers were also asked to complete and return the enclosed Recall Return Form. Contact Pentron Clinical Customer Service at (800) 551-0283, (option 1) directly to handle the arrangement of a quick return and replacement. |  | 100 | 100 Units | Worldwide Distribution - USA (nationwide) and the countries of: Armenia, Russia and Ukraine. | Terminated | January 09, 2012 | 312 | 100
Z-0308-2012 | 55294 | No | CAH | Breathing circuit bacterial filter. | Anesthesiology | General Hospital | 510(k) | ViroMax viral/bacterial filter. The filter reduces the risk of infection due to cross-contamination between patients |  |  | March 31, 2010 | 2010 | A M Systems Inc | Some ViroMax viral/bacterial filters are leaking at the point where the two halves of the filter are ultrasonically bonded together. | Nonconforming Material/Component | A - M Systems sent a Urgent Notification of Product Recall/Correction dated March 31, 2010, to all affected customers. The letter identificed the product the problem and the action needed to be taken by the customer.Customers were instructed to cease use and/or distribution of the affected product and to immediately notify any of their sub-accounts with the same information. Customers are asked to provide A - M Systems with the amount distributed and amount remaining in inventory and to then destroy or discard any remaining product in their possession.Please call 360.683.8300, if you have any questions or concerns. |  | 37800 | 37800 | Worldwide Distribution -- USA ( nationwide) including the states of AZ, CA, FL, KY, MA, MI, NC, OH, OR, and WI., and the countries Canada, Mexico, and Taiwan. | Terminated | January 05, 2012 | 645 | 37800
Z-0582-2012 | 48745 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Artificial Ligament Fixation Device 15MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 25MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 20MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 30MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 35MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device #7 PE Ziploop (TM) ToggleLoc (TM). |  |  | April 24, 2008 | 2008 | Biomet Sports Medicine | Six sizes of ToggleLoc, Artificial Ligament Fixation Device, have been manufactured with the suture pull hole drilled in the wrong location on the device.The ToggleLoc implant is deployed with a suture that pulls the device through the tunnel. The hole being drilled in the wrong location will orient the device to pass through the tunnel with the angled end of the implant first. This may caus | Employee error | Biomet Sports Medicine sent an "URGENT MEDICAL DEVICE REMOVAL" notice dated May 15, 2008 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The letter instructs customers to locate and discontinue use of the affected product. Customers should contact the firm at 800-535-8692 for questions related to this notice. |  | 203 | 203 | Worldwide Distribution-USA (nationwide) including the states of CA, FL, GA, ID, IL, IN, KS, TX, and WA, and the countries of Australia, Chile, France, Poland, The Netherlands and Turkey. | Terminated | December 27, 2011 | 1342 | 203
Z-0604-2012 | 60479 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI 101 Hematocrit/Oxygen Saturation Monitoring System. The CDI 101 Hematocrit/Oxygen Saturation Monitoring System is intended for use during cardiopulmonary bypass procedures where continuous monitoring of the blood's oxygen saturation, hematocrit and hemoglobin levels is needed. |  |  | December 07, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo is recalling the CDI 101 Hematocrit/Oxygen Saturation Monitoring System because they have not yet submitted a 510(k) to FDA to market the device. | No Marketing Application | 12/7/11 &amp; 12/8/11 All consignees received a "Transition and Notification Guide" dated 12/6/11 and a "URGENT MEDICAL DEVICE REMOVAL" letter dated 12/4/11, which will inform them of the removal of the CD l 101 and provide the user with various options. Terumo will remove CD l 101 units from the US market over the next 18 months. |  | 366 | 366 | Nationwide Distribution - Including the states of: NV, MO, SC, WA, MI, FL, GA, NY, CT, IL, VA, NY, TN, IN, HI, TX, AL, OH, MD, AZ, CA, KS, MD, PA, and NM. | Terminated | August 30, 2013 | 632 | 366
Z-0420-2012 | 60096 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Single Use PERPOS PLS System, Catalog #9045-01 |  |  | February 22, 2008 | 2008 | Interventional Spine Inc | The recall was intiated because Interventional Spine has confirmed that the PERPOS 2-in-1 Drill (6070) manufactured by a contract manufacturer for Interventional Spine, Inc., may break during use. | Pending | The recall communication was initiated on February 22, 2008 with Interventional Spine forwarding an Urgent Recall Notification to all their customers (via email). The notification informed the customers of the problem identified and instructions for the recall.If you have any questions, contact the Vice President, Quality, Regulatory and Clinical Affairs at (949) 525-1493 or email: jmetcalf@i-spineinc.com. |  | 65 | 65 | Worldwide distribution: USA (nationwide) including states of: CA, FL, MO, NY, OH, PA, TN, WA, and WI; and country of: Europe. | Terminated | February 29, 2012 | 1468 | 69
Z-0453-2012 | 60481 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Cepheid brand Xpert MRSA/SA-BC US-IVD, an In-vitro diagnostic test kit, Catalog Numbers: GXMRSA/SA-BC-10, Lots Affected: 2803 and 3001; |  |  | November 17, 2011 | 2011 | Cepheid | Cepheid has received reports from some of its customers regarding higher than desired (&gt;5%) rates of invalid results when using lot numbers 2803 and 3001 of the Xpert MRSA/SA BC (blood culture) product. | Nonconforming Material/Component | On 11/17/11, Cepheid issued letters via mail to all its direct consignees, informing them of the affected products and providing instructions on the recall. |  | 1289 | 1289 | Product was released for distribution in US. | Terminated | January 09, 2012 | 53 | 1289
Z-0515-2012 | 60424 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS, versions VA70A and VA70B Client/Server Installations |  |  | September 19, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Unintended behavior when using syngo Imaging XS, versions VA70A and VA70B Client/Serve installations. Loading additional images/series from Patient Browser using the thumbnails can cause an unintended deletion of images/series. | Software design | Siemens Medical Solutions USA, Inc sent a Customer Safety Advisory letter dated August 24, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you detect patient data inconsistenceis in your system caused by this problem, please contact Siemens Service for support.For further questions please call (610) 219-4834. |  | 12 | 12 | Nationwide Distribution including the states of CA, IN, IA, NV, NE, NY, OH, TN and TX. | Terminated | January 24, 2014 | 858 | 12
Z-0390-2012 | 60407 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Maxcem Elite(TM) self-etch/self-adhesive resin cement, Refill Kit, net wt. 10 g, REF/Part # 33872, CLEAR --- Refill Kit Contents: 2 Syringes (5 g ea), 10 Automix tips (regular) 4:1, 10 Automix tips (wide) 4:1, 10 Intra-oral tips |  |  | November 10, 2011 | 2011 | Henry Schein, Inc. | The specific lots of Maxcem Elite Refill Kit appear not to be authentic. | Pending | The firm, Henry Schein, Inc., sent an "URGENT: Product Recall (Dental Practitioners/User Level) letter and response form dated November 10, 2011, via first class mail to its customers. The letter instructed the customers to examine their inventory to verify if they have any of the specified lots on hand. If so, remove it from their shelves. The product may be returned for credit within 30 days, to the following address: Henry Schein, Inc., 41 Weaver Road, Denver, PA 17517. Additionally, the letter instructed the customers to complete, sign, date, and return the enclosed Henry Schein Response Form via fax to: (631) 843-5557 or mail to: Regulatory Affairs Department (E-355) 135 Duryea Road, Melville, NY 11747, Attn: Regulatory-RR. The letter stated "Only returns of the above noted recalled item/lot numbers purchased from Henry Schein, Inc. will be credited to your account. In order to expedite your credit, please include a copy of your response form and invoice, if available, with your return".If you have any questions regarding this matter, please contact Recall Coordinator at 631-390-8000 ext. 8476. |  | 18855 | 18,855 units total: Domestic - 18,820 units; Foreign - 35 units | Worldwide distribution: USA (nationwide) and countries of: Bermuda, Martinique FWI, Nassau Bahamas, St Maarten NA, Australia, and South Africa. | Open, Classified |  |  | 18855
Z-0507-2012 | 56545 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga Unified Intelligence System (UIS), This is a software-only device, WHIE Implementation v1.1i. A calculator/data processing module, for clinical use. |  |  | July 27, 2010 | 2010 | Microsoft Corporation Health Solutions Group | The database may incorrectly match patient data resulting in the incorrect patient data being displayed. This could potentially present a risk to patient safety. | Software design | Microsoft notified the sole consignee by FAX on 7/22/2010 with an Urgent Medical Device Safety Information. The letter identified the affected product, the description of the problem, and actions to be taken to correct the problem, along with some additional information for the consignee. In the interim, the firm proposes to block access to patient records identified as mismatched. The firm also stated that the issue has been addressed in version 2.0. The customer should report any incidents to Michael Stokes, Director of Policy Compliance at 208-938-5397 or via e-mail at Michael.Stokes@microsoft.com. |  | 1 | 1 | Nationwide Distribution -- Wisconsin. | Terminated | December 20, 2011 | 511 | 1
Z-0505-2012 | 60412 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | REF UT8150 Urethal Insertion Tray 14Fr Red Rubber Catheter with PVP Swabs Qty: 20 For Rx Only Sterile unless package is damaged or opened. Latex Free Manufactured for: Coloplast A/S 3050 Denmark STERILE EO Distributed by: Coloplast Corp. Minneapolis, MN 55411 USA 1-800-328-3863 Distribue par: Coloplast Canada Corporation 3300 Ridgeway Drive Unit 12 Mississauga, Ontario Canada L5L 5Z9 1-905-820-7588 |  |  | September 02, 2011 | 2011 | Nurse Assist, Inc | Nurse Assist, Inc. received a recall notice from the Triad Group. This recall has been initiated due to a recall of the PVP Solution Packs manufactured by the Triad Group that are contained in certain procedure trays. | Component design/selection | Nurse Assist sent a recall notice dated September 2, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately and quarantine the affected product. Customers were also instructed to identify their customers and notify them at once of the product recall. Customers were instructed to complete and return the enclosed response form as soon as possible. For those distributors that choose to destroy the affected product they must use the attached Notice of Destruction to report the specifics of the destroyed product. and contact Nurse Assist Customer Service for a RGA # prior to destruction. For any questions call Nurse Assist Customer Service at 1-800-649-6800. |  | 968 | 968 cases | Worldwide Distribution - USA (nationwide) and the countries of Canada and Saudi Arabia. | Terminated | October 26, 2012 | 420 | 55118
Z-0296-2012 | 60099 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Trilogy 100, 200 and 202 Ventilators |  |  | October 11, 2011 | 2011 | Respironics, Inc. | Trilogy 100, 200 and 202 units are being recalled due to low impeller press-force used during the manufacturing process. If the press force were low enough, theoretically the impeller could move up the shaft and come into contact with the inside of the blower housing/impeller enclosure. This may impact therapy. | Process design | Philips Healthcare, began notifying customers via telephone on October 11, 2011. The telephone notification discussed the product, problem and actions to be taken. The customers were instructed to not remove a ventilator from a patient without providing a replacement unit; provide a fax number or email address Philips can use to send the Business Reply Form to complete and return via fax to: 800-733-9962 or email: report.request@philips.com; locate their affected devices (if the device is with a patient, replace it with a non-affected unit-Philips will provide a loaner overnight if needed)-(if the device is in their inventory, quarantine it and return it using the RA number provided in the email/fax that Philips sends them); and If loaner units are needed, they are to call 1-877-387-3311 to arrange return of your affected device and to obtain a loaner unit.If you have any questions or concerns, contact Philips Respironics at 877-387-3311. |  | 127 | 127 | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CO, GA, ID, IL, KS, MI, MO, NE, NV, NJ, OH, PA, TN, and TX; and countries of: Australia, Austria, Canada, China, Denmark, Germany, Great Britain, India, Italy, Japan, Korea (Republic of [South] Korea), Netherlands, Philippines, Saudi Arabia, Spain (Espana), and Turkey. | Terminated | November 06, 2012 | 392 | 127
Z-0518-2012 | 60381 | No | BSF | Carbon dioxide absorber. | Anesthesiology | Anesthesiology | 510(K) Exempt | The IS Can Disposable CO2 absorber x 6 --- |  |  | February 09, 2011 | 2011 | Intersurgical Inc | Intersurgical Incorporated has issued a voluntary recall of the IS Cans Carbon Dioxide Absorber due to a possibility that loose granules may enter the inspiratory path of the breathing system. | Device Design | Intersurgical Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 8, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately discontinue use and quarantine any affected product found. A Recall Response Form was attached to the letter for customers to complete and return via fax at 315-451-3696, Attn: Customer Service. Contact Customer Service at 1-800-828-9633 if you have questions concerning this recall. |  | 110 | 110 units | Worldwide Distribution- USA (nationwide) including the states of Florida, North Carolina, New York, Ohio, and the country of Canada. | Terminated | January 15, 2013 | 706 | 110
Z-0428-2012 | 60516 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Level 1Âż Normothermic I.V. Fluid Administration Sets, |  |  | November 21, 2011 | 2011 | Smiths Medical ASD, Inc. | Disconnected/ loose Luer connectors at the patient end of the triple lumen tubing of the HOTLlNE Disposable Administration Sets and Level 1 D/DI-60HL Normothermic IV Administration Sets | Device Design | Smiths Medical notified consignees by letter dated November 28, 2011 titled: Urgent Field Safety Notice. The letter identified the affected product, the problem, and the actions to be taken. Customers were instructed to inspect inventory and segregate unused affected product. Distributors were instructed to immediately stop distributing and quarantine all inventory. The letters included a Confirmation Form to be complete and return, by fax to 781-610-9859 or by email to hotlineconnectorrecall@smiths-medical.com. For questions consignees were instructed to contact Smiths Medical at 1781-763-9330. |  | 7631 | 7631 units | Worldwide Distribution  (USA) (nationwide) - Argentina, Austria, Australia, Belgium, Bahrain, Bermuda, Canada, Chile, China, Columbia, Czech Republic, Germany, Denmark, Estonia, Egypt, Finland, France, Great Britain, Guernsey, Greece, Guatemala, Hong Kong, Hungary, Indonesia, Ireland, Israel India, Iceland, Italy, Japan, South Korea, Kazakhstan, Lebanon, Luxembourg, Mexico, Malaysia, New Caledonia, Netherlands, Norway, New Zealand, Panama, Poland, Portugal, Qatar, Singapore, Sweden, Slovak Republic, South Africa, Spain, Sri Lanka, Switzerland, Taiwan, and Vatican City State. | Terminated | August 07, 2014 | 990 | 805506
Z-0168-2012 | 60236 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | SafeSheath(R) CSG(R) Coronary Sinus guide Braided Core Hemostatic Introducer, SafeSheath(R) Worley Telescopic Braided Series, and Situs LDS2 Lead Delivery System - Catalog numbers: CSG/WORLEY/BCor-1-09; CSG/WORLEY/BCor-2-09; CSG/WROLEY/L/BCor-1-09; CSG/Worley/R-1-09; CSGWORBC10M; CSGWORBC29M; CSGWORLBC19M; LVI/66-5.5-CR; LVI/66-5.5-RE; LVI/75-5-62-07-HO; LVI/75-5-62-07-HS; LVI/75-5-62-07-MP; LVI/75-5-62-07-RE; LVI/75-5-66-5.5-RE; LVI/Tele/B-60-07-HO; LVI/Tele/B-60-07-HS; LVI/Tele/B-60-07-MP; LVI/Tele/B-60-07-RE; RL065; RL066; RL067; and RL068 |  |  | October 07, 2011 | 2011 | Thomas Medical Products Inc | Radiopaque (RO)/soft tip of the braided core component in the CSG kits may detach from the braided core tube. | Nonconforming Material/Component | The firm,GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 7, 2011, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately segregate all affected product in a manner that ensures it will not be used; return any product with reference to the return authorization (RA) number or contact customer service representative to facilitate return of product; if any affected product has been forwarded to another facility, contact that facility to arrange return, and complete and return the attached product reconciliation form.If you have any questions, you may contact the Customer Service number at 1-866-446-3003 Monday through Friday from 8:30am to 5:00pm EDT. |  | 33134 | 33,134 | Nationwide distribution: CA only. | Terminated | August 29, 2012 | 327 | 33134
Z-0381-2012 | 60214 | No | FST | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | UltraLite Headlight module-grey |  |  | October 10, 2011 | 2011 | Integra LifeSciences Corp. | Headlight assembly (module) may become loose and separate from the headband assembly | Device Design | Integra Life Sciences notified consignees by URGENT DEVICE RECALL letter dated October 10, 2011, via FedEx, email or telephone conference calls. The letter identified the affected product, the problem, and the actions to be taken. Customers were instructed to identify and return any affected product. Customers were requested to complete and return the attached Field Corrective Action Acknowledgement Form to Integra Surgical indicating receipt and review of this notification. For assistance and additional information contact your local and Integra Sales Representative or call Integra Surgical Customer Service at 1-800-431-1123. |  | 60 | 60 units | Worldwide Distribution - USA (nationwide) including Puerto Rico, and the countries of: Africa, Algeria, Australia, Austria, Brazil, Canada, Check Republic, China, Chile, Costa Rica, Denmark, Dominican Republic, Ecuador, Ethiopia, Finland, France, Georgia, Germany, Greece Hungary, Guatemala, India, Ireland, Israel, Italy, Iraq, Japan, Jordan, Kenya, Kuwait, Lebanon, Lithuania, Macedonia, Mexico, Netherlands, New Zealand! Norway, Oman, Pakistan, Peru Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, Qatar, United Kingdom, Yemeni United Arab Emirates, Uruguay, and Venezuela. | Terminated | October 01, 2012 | 357 | 7216
Z-0486-2012 | 60357 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare AW Server version 2.0 with release 2.0 - 4.0 |  |  | November 17, 2011 | 2011 | GE Healthcare, LLC | 2 Unrelated Software Issues: An incorrect patient name maybe displayed when switching between two Volume Viewer tabs with different patients. Also, if a second patient exam is selected while the application is starting, the second patient's data will be displayed with the first patient's name display on the applications tab. | Software design | Consignees were sent on 11/17/11 a GE "Urgent Medical Device Correction" letters dated November, 16 2011 The letter was addressed to Hospital Administrators, Radiology Mangers and Risk Managers. The letter described the Safety Issue, Safety Instructions Affected Product Details, Product Correction and Contact Information. For the first issue, customers were asked to contact their field engineer to install the correct configuration file. Also, customers should not apply the recommended changes from the Customer Release Note as those changes cause the issue to occur. With the second issue, customers are to wait until the application has fully initialized before making another patient exam selection. GE Healthcare will correct all affected AW server installations via a software upgrade at no cost to the customer. A representative will contact customers to arrange for the correction. Questions should be directed to customers' local GE Healthcare Service Representative or at the numbers provided in the letter. |  | 308 | 308 | Worldwide Distribution -- USA, including states of AZ, CA, CO, CT, DE, DC, FL, HI, IL, IN, LA, MD, MI, MN, MS,MO, NJ, NM, NY, NC, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, and WI and countries of UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, REUNION, QATAR, POLAND, NORWAY, NEW ZEALAND, LEBANON, SOUTH KOREA, JAPAN, ITALY, ISRAEL, HONG KONG, GUADELOUPE, GREAT BRITAIN, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CANADA, BRAZIL, BOSNIA AND HERZEGOVINA, BELGIUM, and AUSTRALIA. | Terminated | July 03, 2012 | 229 | 308
Z-0365-2012 | 60117 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000sp Instrument; an automated system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list 09K14 |  |  | October 11, 2011 | 2011 | Abbott Molecular | Abbott Molecular has received reports in which the liquid waste sensor and components in the m2000sp lower instrument cabinet overheated and/or caught fire. | Component design/selection | Abbott Molecular sent a Field Correction Recall letter dated October 10, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until further notice, and to complete the enclosed reply sheet and fax it to (847) 775-6728 at their earliest convenience to acknowledge receipt and understanding of the letter.For questions regarding this recall call 224-361-7133. |  | 432 | 432 units | Worldwide Distribution -- USA (nationwide) including the countries of Algeria, Australia, Austria, Azerbaijan, Belgium, Brazil, Burundi, Cameroon, China, Croatia, Estonia, Ethiopia, Finland, France, Gaza &amp; Jericho, Germany, India, Ireland, Israel, Italy, Kazakhstan, Kenya, Libya, Malawi, Malaysia, Mali, Mozambique, Netherlands, New Zealand, Pakistan, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, United Arab Emirates, Uganda, Ukraine, United Kingdom, Vietnam and Zambia | Terminated | February 23, 2016 | 1596 | 432
Z-0452-2012 | 59562 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC and Synchron LX Clinical Systems 100 ÂżL Sample Syringe, 100UL-PKGD, Part Number 474171 |  |  | September 02, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Modular Chemistry (MC) and Cartridge Chemistry (CC) 100 uL sample syringes may experience premature wear as indicated by brown discoloration on the plunger and/or inside the glass syringe barrel. Excessive wear of the 100 uL sample syringe barrel and plunger could lead to leaking and may affect results. | Component design/selection | Beckman Coulter sent an URGENT: Product Corrective Action letter to all affected customers. The letter provided customers with an explanation of the problem identified and actions to be taken. The letter instructs customers to share this information with their laboratory staff and retain a copy as part of their laboratory Quality System documentation. In addition, the letter states that if the affected product was further distributed to provide a copy of the letter to them. Customers were also asked to complete and return the enclosed Response Form. For any questions contact Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. |  | 3278 | 3278 total (3083 in US, 195 in Canada) | Worldwide Distribution - USA (nationwide) and the country of Canada. | Terminated | June 30, 2014 | 1397 | 3278
Z-0418-2012 | 59739 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | FOUNDATION HIP SYSTEM Series 460 Hip Stem STERILE R CEMENTED Sz 16.5mm |  |  | August 17, 2011 | 2011 | Encore Medical, Lp | The sterility has the potential to be compromised. | Packaging | Encore Medical Lp sent a Urgent Field Safety Notice dated August 17, 2011. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately quarantine these devices if you have not already done so. Contact Customer Service to obtain replacement devices as well as an RMA. Return all affected devices by October 31, 2011, using the RMA Customer Service provided in step #3.If you have any further questions you may call (512 ) 834-6330. For International Customer Service please call (512) 834-6275. |  | 169 | 169 units | Worldwide Distribution -- USA ( nationwide ) including the states of WI, NY, MO, TX, FL, CA, UT, AZ, Puerto Rico, and the country Japan. | Terminated | May 11, 2012 | 268 | 169
Z-0372-2012 | 60307 | No | BZT | Esophageal stethoscope with electrical conductors. | Anesthesiology | Anesthesiology | 510(k) | Cincinnati Sub Zero, Steri-Probe Temperature Probe, Esophageal/Rectal Temperature Probe, Part Number: 39622, Catalog #491B. TA disposable temperature probe used in conjunction with a measuring/monitoring device to continuously measure/monitor a patient's core body temperature. |  |  | September 26, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | Reports that the Rectal/ Esophageal Temperature Probes are not reading correctly. This situation causes the Blanketrol device, which the probes connect to, to shut down operation and display a 'check probe' error. | Nonconforming Material/Component | Cincinnati Sub-Zero Products, Inc., issued an "Urgent Medical Device Field Action" letter on November 1, 2011 to all affected customers. The letter includes; description of product and problem, and a list of the affected lot numbers. The letter instructs the customers to discontinue use of the device. All unused product is to be returned to the Biomedical/Clinical Engineering and/or Service Department for an replacement and all used probes are to be disposed of in accordance with hospital policy. The customers are also instructed to complete and return an attached Recall Response Form ASAP. For additional information, customers may contact the Beatrice Washington at 1-800-989-7373 or (513) 772-8810. |  | 369680 | 369,680 | Nationwide distribution and Canada. | Terminated | August 13, 2013 | 687 | 369680
Z-0506-2012 | 60581 | No | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Walgreens Quick Read Digital Thermometer; an electronic digital thermometer; Model WIC551226 |  |  | November 23, 2011 | 2011 | Walgreen Company | The thermometers have been reported to have inaccurate readings | Labeling design | On November 23, 2011, Walgreens issued a recall notice electronically to their stores, instructing them to pull and quarantine the Thermometers and placing a Stop Scan on the products in the computer cash register system on the same date. Walgreens also removed the thermometer from their internet sales site. |  | 92000 | estimated 92,000 units | Nationwide through Walgreens retail stores and their internet site. | Terminated | August 14, 2017 | 2091 | 92000
Z-0391-2012 | 60371 | stocking, medical support (for general medical purposes) | FQL | Medical support stocking. | General Hospital | General Hospital | 510(K) Exempt | Truform Therapeutic Classic Medical Style Compression Stockings 20-30 mmHG, Truform Therapeutic Classic Medical Style Compression Stockings 30-40 mmHG, A Div of SAI, Cincinnati, OH 45209 |  |  | October 07, 2011 | 2011 | Surgical Appliance Industries | During an audit with the FDA, it was discovered that a new compression testing device used for testing hosiery had been implemented and that the operators were not properly trained. Additionally, the acceptance criteria for the new device were not changed from the old testing device to reflect additional data (compression profile percentages). This resulted in the device history files containing s | Employee error | The firm, Surgical Appliance Industries sent an URGENT: MEDICAL DEVICE RECALL letters dated November 7, 2011 to all affected customers. The letter described the products, problem and actions to be taken. Customers were instructed to immediately examine their inventory and quarantine all product(s) subject to recall for return and exchange. In addition, the letter states that if the product was further distributed to contact their customers and notify them of the recall. Customers were asked to complete and return the enclosed Recall Response Form. For any questions contact Quality Systems Manager, at 513-271-4594 ext. 282. |  | 351896 | 351896 pairs | Worldwide Distribution - USA (nationwide) and the countries of: Australia, Canada, Colombia, Hong Kong, Peru, Singapore, Spain, and Venezuela. | Terminated | March 05, 2013 | 515 | 363715
Z-0463-2012 | 60375 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew 5.5mm DYONICS BONECUTTER ELECTROBLADE Resector |  |  | November 07, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | Manufacturing issue: The distal tip of the outer sheath (return electrode) may detach and become loose in the joint during use. | Process control | Smith &amp; Nephew Inc. Endoscopy Division notified facilities on 11/7/11 via an "Urgent - Product Recall 1st Notification - Urgent" letter and/or phone. The letter identified the affected product and the reason for the recall. It also discussed the potential risks, customer actions, and instructions for the return of the product. Customers were to check their inventory and locate any of the affected product. They were to also complete the requested contact information and return the notification letter. In addition, customers were to contact Smith &amp; Nephew Endoscopy Division Returns Group at endo.andreturns@smith-nephew.com for information on how to return the affected product and receive replacement product. Questions, contact Hoangthi Le directly by phone at 508-337-3731. |  | 1743 | 1743 units | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Dubai, Finland, France, Germany, Norway, South Africa, Switzerland, United Kingdom, and Venezuela. | Terminated | February 22, 2016 | 1568 | 1743
Z-0603-2012 | 60390 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The GSP 2021-0010 (Genetic Screening Processor) instrument is a fully automated, high through put batch analyzer for time- resolved and prompt fluorescence analysis of samples in microtitration plates. It is intended for in vitro quantitative / qualitative determination of analytes in body fluids. The chemistries running on the GSP instrument are used to test blood specimens dried on filter paper as an aid in screening newborns for congenital disorders. |  |  | October 24, 2011 | 2011 | PerkinElmer LAS, Inc | Malfunction of the GSP disk remover unit may cause a punched disc to remain in the well during the measurement of the plate. This may cause a false negative screening result being reported for 17-OHP, resulting in failure to identify cases of congenital adrenal hyperplasia (CAH). | Device Design | Recall (correction) letters were sent to all customers on October 24, 2011. An improved disk remover design is under development to eliminate the reported problem. Until the final field correction is available, customers should follow the precautions provided. The needed protocol update will be performed by the local service engineer. If customers observe an increase in the occurrence of the problem, the laboratory should contact their local PerkinElmer representative. |  | 52 | 52 | Worldwide Distribution - USA, including the states of: Missouri, Florida, Texas, Oregon, Alabama, Louisiana, New Jersey, Georgia, Virginia, Kentucky and the countries of: Australia, Belgium, Brazil, Egypt, France, Germany, Italy, Netherlands, Norway, Qatar, Saudi MOH, Switzerland, and the UK. | Terminated | July 20, 2012 | 270 | 52
Z-0591-2012 | 60625 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Agfa Computed Radiography Systems with DX-G Digitizer. |  |  | December 06, 2011 | 2011 | AGFA Corp. | Occasionally images can be lost or sent to an incorrect folder or the digitizer may stop either during startup or during a cassette scan cycle. | Software design | Agfa HealthCare sent a 'URGENT SAFETY NOTICE" letter dated December 6, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Agfa Service will contact your facility and schedule a software upgrade if this has not already been performed. Please complete the attached feedback form as soon as possible and fax it to 864-421-1662.If you have any questions about this matter, please contact your local Agfa HealthCare Service at 877-777-2432. |  | 125 | 125 | Worldwide Distribution-- USA (nationwide) including the states of AZ, CA, CO, FL, GA, HI, IL, IN, KS, LA, ME, MI, MN, MO, NE, NH, NJ, NY, OH, OK, PA, SD, TX, VA, WA, WI and WV, and Canada. | Terminated | April 11, 2012 | 127 | 125
Z-0382-2012 | 60437 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System: Lung Optimized Treatment option. |  |  | November 07, 2011 | 2011 | Accuray Incorporated | User Facility reported an anomaly where inhale and exhale CT pairs used for treatment planning did not represent the same magnitude of respiration that was displayed during treatment delivery. | Labeling design | Accuray sent a Urgent Medical Device Correction letter dated November 7, 2011, via Federal Express to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The information provided in this letter will be incorporated in user manuals shipped with the CyberKnife system.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at 1. (877) 668-8667 ( USA) 1. (408) 716-4700 ( non USA ) |  | 30 | 30 | Worldwide Distribution -- USA ( nationwide ) including the states of FL, TN, IL, NJ, CO, PA, MD, CA, MS, MN, SC, WA, OK, GA, AZ, and TX., and the countries of France and United Kingdom. | Terminated | February 27, 2012 | 112 | 30
Z-0387-2012 | 59812 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | ActiveCare DVT |  |  | August 28, 2011 | 2011 | Medical Compression Systems | This voluntary field action is being undertaken because MCS has received a limited number of reports involving the use of a non-MCS AC/DC adaptor with the ActiveCare DVT and ActiveCare+S.F.T. devices. The firm found that an ActiveCare DVT device may work improperly if connected to a non-MCS AC/DC adaptor with a higher nominal voltage. As specified in each product's User Manual, the ActiveCare DVT | Device Design | Medical Compression Systems (DBN) Ltd, sent out "VOLUNTARY FIELD ACTION" letters dated September 15, 2011 to all affected customers. The letter included description of the device and problem. Customers were instructed to check their inventory and make a list of serial numbers of all applicable ActiveCare+S.F.T. and ActiveCare DVT devices in their inventory, include the list on the attached excel file (Field Action Verification Form) and email to maya@mcsmed.com; even if they do not have any such devices. The letter also states that within 15 days from the email, they will receive instructions for the next steps. In addition, the letter stated -YOU MAY CONTINUE TO USE THE ActiveCare+S.F.T.Âż DEVICES WHILE AWAITINGTHE INSTALLATION OF THE OVP.In case you have questions or you need any assistance, you can contact the firm at (317) 485-7887 or e-mail to: maya@mcsmed.com or mahlet@mcsmed.com. |  | 2217 | 2217 units | Nationwide distribution. | Terminated | July 11, 2012 | 318 | 2768
Z-0513-2012 | 60643 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Dimension Vista(R) Cardiac Troponin I Calibrator (KC678) lot 1DD085, in vitro diagnostic |  |  | November 17, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed positive bias on QC and Patients. | Nonconforming Material/Component | Siemens healthcare Diagnostics initiated their recall of lot 1DD085 on November 17, 2011 by sending an Urgent Field Safety Notice letter to all affected customers. Customers were requested to immediately discontinue use and discard any remaining inventory of the CTNI CAL lot 1DD085. |  | 2281 | 2281 cartons | Worldwide distribution, including USA, Austria, Belgium, Canada, Denmark, Finland, France, Germany, Italy, Japan, Netherlands, New Zealand, Portugal, Puerto Rico, Slovenia, South Korea, Spain and Switzerland. | Terminated | August 26, 2014 | 1013 | 2281
Z-0473-2012 | 60608 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Y-Type Solution Set, 78" (2.0 m), 2 Injection Sites, Male Luer Lock Adapter; an Rx sterile, nonpyrogenic fluid pathway for the administration of sterile I.V. fluids to the patient; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 1C8581 |  |  | December 06, 2011 | 2011 | Baxter Healthcare Corp. | During production of the IV sets, solvent was inadvertently applied to the locking collar of the male luer connector, which may have damaged the component, preventing proper engagement of the locking collar to secure a connection. | Employee error | The firm, Baxter, sent an "Urgent Product Recall" letter dated December 6, 2011, via first class mail to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to inspect all locations where their facility's inventory is located for products with the affected lot numbers; return the affected product; request a credit or replacement by calling Baxter Healthcare Center for Service at 1-888-229-0001 between 7:00am to 6:00pm Central time, Monday through Friday; complete and return the attached Customer Reply Form via fax to: 1-847-270-5457, or scan and email to fca@baxter.com; and if they are a dealer, wholesaler, or distributor/reseller that distributed any affected product to other facilities, please notify their customers of this recall.If you have any questions regarding the communication, please call The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 2352 | 2,352 units | Worldwide distribution: USA (nationwide) including states of: Alaska, California, Colorado, Florida, Georgia, Hawaii, Illinois, Kansas, Kentucky, Massachusetts, Maine, Michigan, Minnesota, Missouri, North Carolina, New Hampshire, New Jersey, New York, Ohio, Pennsylvania, Rhode Island, South Carolina, Tennessee, Virginia, Washington, Wisconsin and West Virginia; and country of: New Zealand. | Terminated | November 01, 2012 | 331 | 48480
Z-0464-2012 | 60568 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | CD HORIZON (R) Spinal System, MULTI-AXIAL SCREW FOR 5.5mm ROD W/DUAL LEAD THREAD, Size 6.5mm X 40mm, Mat'l TITANIUM ALLOY, REF X0809439, Rx only, Medtronic Sofamor Danek USA, 1800 Pyramid Place, Memphis, TN 38132 |  |  | November 30, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | One screw may be longer then size etched on the screw. | Mixed-up of materials/components | Medtronic Sofamor Danek USA, Inc sent a Voluntary Field Correction letter dated November 30, 2011, via FedEx to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The firm will visit each consignee to perform a visual check of inventory.Please utilize the attached questionnare to account for the affected product and to record any patient injuries that may have occurred specifically related to the use of these items. If you need assistance, or if you have any related questions or concerns please call (800) 876-3133 extension 3173 |  | 252 | 252 units | Nationwide Distribution including the states of AL, PA, IN, CO, MI and IA. | Terminated | April 05, 2012 | 127 | 282
Z-0583-2012 | 60258 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | INOmax DS (Delivery System) Model 10003 &amp; 10007. |  |  | November 30, 2010 | 2010 | Ino Therapeutics | Through the review and monitoring of its complaint data, Ikaria identified a product anomaly with the INOmax DS Model 1003. Ikaria concluded the anomaly risk to the patient safety was low and non reportable (4/23/10). During a subsequent Level 3 inspection by FDA, it was concluded that the action should have been reported under 21 CFR &amp; 806.10. | Nonconforming Material/Component | Ikaria sent a "RECALL- CUSTOMER PRODUCT REMEDIATION" notice dated December 22, 2011 to all affected customers. The letter describes the product, problem, actions taken by the firm. No action was required by customers. The notice served as correspondence to inform customers of the issue. Customers may direct their questions to Ikaria Customer Care at 1-877-566-9466. |  | 2446 | 2446 | Worldwide Distribution-USA (nationwide) and the country of MALAYSIA. | Terminated | May 29, 2012 | 546 | 2446
Z-0460-2012 | 60051 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | SRS FAST SET PUTTY 3 CC BONE VOID FILLER-STERILE, Part Number SRS-005-FSP, Lot number N001101. |  |  | September 29, 2011 | 2011 | Synthes USA | A recent shipment carrying Norian SRS was lost by the carrier. Since the product was lost before final product release testing, Synthes is unable to confirm that Part Number SRS-005-FSP- Lot Number N001101 meets final product release requirements. The product contains all final labeling and cannot be visually distinguished from the product that has met final product release requirements. | Labeling Change Control | All Synthes accounts with a history of purchasing the product in the past two years will be notified by letter of the incident. All firm sales consultants will also receive written notification of the incident. |  | 131 | 131 units | Synthes USA sent a "NOTICE: MISSING SHIPMENT-NORIAN SRS" letter dated September 29, 2011 to all accounts with a history of purchasing the product in the past two years. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory to verify they have not received the affected product. Call 1-800-620-7025 ext. 5450 for questions concerning this notice. | Terminated | December 18, 2014 | 1176 | 131
Z-0598-2012 | 60521 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Product Code: 2M8161. |  |  | May 08, 2009 | 2009 | Baxter Healthcare Corp. | There was no documentation to show the required tests were performed following the replacement of the battery and battery harness. | Process control | Baxter sent a service technician to the consignee location on 05/08/2009 to perform the battery check and discharge test. The pump was found to be within specification. |  | 1 | one pump | Puerto Rico | Terminated | January 03, 2012 | 970 | 1
Z-0421-2012 | 60454 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | MAXLOCK EXTREME EXTREMITY PLATING SYSTEM WITH VARIABLE ANGLE TECHNOLOGY. |  |  | August 22, 2011 | 2011 | Orthohelix Surgical Designs Inc | The firm discovered a testing protocol error, in both dynamic and static cantilever comparative testing, during the initial comparative testing. | Nonconforming Material/Component | Orthohelix Surgical Designs Inc. sent an URGENT RECALL NOTIFICATION letter dated August 22, 2011 to all affected customers. The letter identified the affected products, the problem, and the actions to be taken. Customers were instructed to immediately inventory, remove and return the recalled product to the firm via FedEx or any shipping method with tracking capability. Customers were also requested to notify any affected user facilities of the recall. For any questions call 330-247-1444 or 330-247-1445. |  | 2722 | 2,722 units | Nationwide Distribution (USA) - including the states of: CA, CO, CT, FL, GA, IA, IL, KY, LA, MD, MI, MO, NC, NJ, NM, NY, OH, OK, PA, SD, TN, TX, UT, VA, WA, WI and WY. | Terminated | February 19, 2012 | 181 | 2722
Z-0517-2012 | 60583 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | Rusch Flexi-slip Endotracheal Tube Stylet with Soft Distal Tip, 6FR, Catalog numbers: 502501, Teleflex Medical Sdn Bhd, P.O. Box, Industrial Estate, 24600 Kamunting, Perak, Malaysia. |  |  | December 02, 2011 | 2011 | Teleflex Medical | Complaints of difficulty removing stylet from the endotracheal tube and possible PVC sheath separation were received. | Pending | The firm, Teleflex Medical, sent an "Urgent Medical Device Recall Notification" letter dated December 2, 2011 to its Consignees/Customers. The letter described the product, problem and actions to be take. The customers were instructed to immediately discontinue use and quarantine any products with the listed catalog and lot numbers; return the product; complete and return the enclosed Immediate Attention Requested Recall Acknowledgement Form via fax to 1-866-804-9881, Attn: Customer Service (a customer service representative will contact the customer with a returns good authorization number and provide instructions for the return of product to Teleflex medical), even if they have no affected product.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 73281 | 73,281 units | Worldwide distribution: USA (nationwide) and countries of: Canada and Europe. | Terminated | April 29, 2014 | 879 | 73281
Z-0485-2012 | 60528 | No | RER |  |  |  |  | Raymond Laser alignment guide - Model #1096597/002 Red Laser. |  |  | November 28, 2011 | 2011 | The Raymond Corporation | Raymond shipped a laser alignment guide as an option on its powered industrial trucks from February 2, 2010 through April 21, 2011, but was not aware of its obligation under 21 C.F.R. 1010.2 and 1010.3 to certify and label. | Radiation Control for Health and Safety Act | Raymond will be initiating a Field Campaign to ensure that all customers with Alignment Guides have the confirmatory laser caution decals and manufacturer label, an updated installation and service manual, the operator instructions, and informational brochure. This campaign will be performed at no cost to Raymond's dealers and customers, and will include documenting the laser serial number for each truck configured with an Alignment Guide for product traceability. |  | 493 | 493 total | Nationwide Distribution | Open, Classified |  |  | 493
Z-0521-2012 | 60519 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Triage Drugs of Abuse Panel Plus TCA 25 Test Kit, PN (Model Number) 92000 |  |  | July 01, 2011 | 2011 | Alere San Diego | The recall was initiated because Alere San Diego has confirmed false positive THC results obtained on Triage Drugs of Abuse Plus TCA 25 devices from lots 235991 and 238928. | Nonconforming Material/Component | A recall communication was initiated on July 1, 2011 with Alere San Diego forwarding a Product Notification and a Verification Form (via fax, e-mail or direct mail) to all their customers who purchased the Triage Drugs of Abuse Plus TCA 25 Test Kit, PN 9200, Lots 235991 and 238928. The letter provides the customers with an explanation of the problem identified and an action to be taken. The firm reminds customers that clinical consideration and professional judgment must be applied to any test result. Customers were instructed to retain the Product Notification letter for their records as documentation that they have received the information and to ensure that all users have a copy. Alere will replace the lots at the customer's request. As the customers complete the fax back form, they were instructed to indicate whether they wish to replace the affected product. Customers were instructed to complete and fax the enclosed Verification Form within 10 days to confirm their receipt of the notice. If customers have any further questions or need to report any concerns they have about the Triage Drugs of Abuse Plus TCA performance, they were instructed to contact Technical Services at (877) 308-8287 or by e-mail at techservice@alere.com. |  | 997 | 997 total kits | Worldwide Distribution -- US and to Canada. | Terminated | April 04, 2012 | 278 | 997
Z-0594-2012 | 60549 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 Radiation Therapy Planning System, software version 9.0. |  |  | November 18, 2011 | 2011 | Philips Medical Systems (Cleveland), Inc. | Phillips Medical Systems have recently determined that a software (Pinnacle3 Software version 9.0) nonconformance can cause incorrect beam geometry. | Software design | Consignees were sent on 11/18/11 a Philips "Urgent Medical Device Correction" letter dated 2011 Nov 16. The letter identified the affected product and described the Problem, Hazard Involved, How to Identify Affected Products, Action to be Taken by Customer/User, Actions Planned by Philips and Further information and Support. If customers observed the issue, they were to manually make the suggested change. In addition, customers were to place the letter in the Instructions for Use until otherwise notified. The firm will deliver 9.2 software release free of charge. Customers in North America are to contact Customer Care Solutions Center at 1-800-722-9377, if they have questions or need support concerning the recall. |  | 715 | 715 (463 US) | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and countries of AUSTRALIA, AUSTRIA, BELGIUM, CHINA, DENMARK, FRANCE, GERMANY, HONG KONG, ITALY, JAPAN, MALAYSIA, MONTENEGRO, NETHERLANDS, NEW ZEALAND, OMAN, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM, and CANADA. | Terminated | February 28, 2014 | 833 | 715
Z-0511-2012 | 60476 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Urine Reagent Strips (URS) 10 Parameters |  |  | October 14, 2011 | 2011 | Teco Diagnostics | This recall was initiated in response to findings during a FDA inspection of the firm. The firm's recall has been initiated due to improper manufacture of sub-component nitrite paper Lot 35134, Expiration Date: 01/2014. Discrepancies were found on the information recorded on formulation sheet and manufacturing log sheets (Refer to next paragraph). Sampling size submitted for QC testing was insuff | Nonconforming Material/Component | Teco Diagnostics sent an "URGENT: MEDICAL DEVICE RECALL" letter dated October 14, 2011 to affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructed customers to examine their inventory for the affected product and to quarantine any product found. A Recall Return Response Form was included with the letter for customers to complete and return. Contact the firm's Technical Support Department at (800) 222-9880 for questions regarding this recall. |  | 550 | 550 kits | Distributed in CA and Poland. | Terminated | April 09, 2013 | 543 | 550
Z-0416-2012 | 60263 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Grafton DBM A-Flex; Unit Size 60 mm diam; 1ea |  |  | August 08, 2011 | 2011 | Osteotech Inc | Aseptic conditions during the production of these products may have been compromised. | Material/Component Contamination | Osteotech sent a recall letter dated August 8, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all units of the affected product. Customers were asked to contact Osteotech Customer Service at 732-578-6699 within two business days of receiving the letter to acknowledge receipt of the letter and to arrange for a return and exchange or credit for the product.For questions regarding this recall call 732-542-2800. |  | 18 | 18 unts | Worldwide Distribution | Terminated | December 08, 2011 | 122 | 33
Z-0516-2012 | 60684 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Portal Vision, R-Arm with 4D Integrated Treatment Console; |  |  | November 23, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly was discovered when using the 4DITC in combination with an R-Arm to acquire double exposure MV images. | Software design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" dated December 1, 2011 to all affected sites via Certified mail, return receipt. Canadian accounts were sent by Fed Ex and internationally to Varian International Customer Support Managers for translation and notification of customers. The notice described the product, problem and actions to be taken. The customers were instructed to always acquire both the treatment port and the open port images BEFORE displaying the digital graticule, when using a standard double exposure template; and to advise the appropriate personnel working in their radiotherapy department of the content of this letter.Special instructions for customers outside the USA and Canada: these customers were instructed to complete and return the attached Proof of Notification or Receipt Verification Card to Varian Medical Systems, once they have read this document. Varian is developing a correction for this issue and a Customer Service rep will contact the customers when the correction is available to schedule its installation on their systems.If you have any questions, contact USA and Canada at 1-888-827-4265; Europe: +41 41 749 8844; Email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: China.apps.helpdesk@varian.com; Japan: Japan-Apps.Helpdesk@varian.com and Latin America: soporte.al@varian.com or Internet: Oncology systems customer site: www.myvarian.com and Varian Medical Systems public site: www.varian.com. |  | 1486 | 1486 devices | Worldwide distribution: USA (nationwide) and countries of: Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Belarus, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France,, Georgia, Germany,Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Lativia, Macedonia, Malaysia, Martinique, Mexico, Montenegro, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Panama, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom and Venezuela. | Terminated | August 06, 2014 | 987 | 1486
Z-0510-2012 | 60148 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | AVEA Ventilator (Catalog Code 17310, 17311, 17312). |  |  | September 08, 2011 | 2011 | Carefusion 211 Inc | The recall was initiated because Carefusion has identified a potential risk associated with certain AVEA ventilators and affected replacement parts manufactured between March 1, 2009 and June 30, 2011. Carefusion is voluntarily initiating a field correction of the affected devices to preclude the possibility of this risk. | Nonconforming Material/Component | CareFusion sent an "URGENT PRODUCT RECALL" letter dated September 25, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers were instructed to contact the firm by telephone to coordinate implementation of the corrective action at their site. Customers were instructed to return an enclosed response card to expedite the correction process and acknowledge receipt of the recall notification. Carefusion does not require that the customers return their devices. Customers will be contacted by a member of the CareFusion Technical Support Department to arrange for onsite remediation of the affected devices. In the interim, if any affected AVEA ventilator units in their facility exhibits a sustained Extended High Ppeak alarm followed by the opening of the Safety Valve, customers were instructed to remove the ventilator from service, provide alternate ventilation and contact Carefusion Technical Support (800) 213-2466 to report the issue. All ventilator dependent patients should be constantly monitored by qualified personnel to ensure that if a malfunction were to occur, alternate ventilation can be provided. |  | 6797 | 6797 total, 1891 in US, 4906 worldwide | Worldwide Distribution-USA (nationwide) and the countries of Cyprus, Estonia, Brazil, Kuwait, United Arab Emirates, Slovenia, Lithuania, Latvia, Belgium, Bermuda, Ecuador, Guatemala, Italy, Canada, Australia, Netherlands, Germany, United Kingdom, Spain, Czech (Rep), Nicaragua, China, Bangladesh, Morocco, Bolivia, Vietnam, Slovakia, Paraguay, Argentina, Turkey, Chile, Singapore, Japan, Colombia, Egypt, Hong Kong, Honduras, India, Croatia, Venezuela, Saudi Arabia, Greece, Hungary, Mexico, Austria, Ireland,Indonesia, Dominican Republic, Portugal, South Africa, Panama, Philippines, Libyan Jamahiriya, Jordan, Malaysia, France, Thailand, Peru, Costa Rica, Korea, and Taiwan. | Terminated | October 29, 2013 | 782 | 6797
Z-0592-2012 | 60524 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Orbit and Orbit Micro |  |  | March 01, 2011 | 2011 | ICU Medical, Inc. | The recall was initiated because ICU Medical Inc. has become aware of the market recall of the Triad Group's alcohol prep pads, alcohol swabs, and alcohol swabsticks manufactured by Triad in the United States and marketed under various brand names. The Triad alcohol prep pads have the potential contamination of the bacteria, Bacillus cereus. The Triad Group alcohol prep pads have been copackaged | Pending | The firm, ICU Medical, Inc., sent an "IMPORTANT-Important Information About Triad Group's Alcohol Prep Pads" letter dated March 2, 2011 (via email) to their distributors. Hard copy notices were shipped (via next day express delivery with receipt confirmation) to all their customers. The Important letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were informed that the Orbit and Orbit Micro subcutaneous infusion sets as well as the IV start kits are not contaminated and may continue to be used according to the directions for use. The issue is confined to the actual Triad alcohol prep pad products. Patients and healthcare providers should not use the alcohol prep pads packaged with these devices and should instead use an alternate alcohol prep pad that is not involved with the Triad Group recall. ICU Medical has halted all shipments of product that contain the Triad alcohol prep products to its distribution network. An alternate wipe has been qualified and is now being shipped in new orders. Customers were instructed to forward the Important letter to their customers and on to any end users of the devices. For Orbit and Orbit Micro devices, patients are instructed to immediately discontinue using the Triad alcohol prep pads included in the Orbit packaging and dispose of those pads in the trash. Patients should use an alternative prep pad that is not subject to the Triad recall. For IV Start Kits, distributors and users are instructed to place the enclosed label on the device prior to use. Additional labels can be obtained by contacting ICU Customer Service Monday through Friday between the hours of 8:30AM and 4:00PM Pacific Standard time: (800) 866-9025 and select option 8, or email the following address: customerservice@icumed.com. Customers may also contact regulatory@icumed.com for MSWord template file for Avery 5160 labels and print additional labels. Customers were instructed to complete |  | 376290 | 376,290 units | Worldwide distribution: USA (Nationwide) and countries of: Austria, Belgium, Chile, Cyprus, CZ, France, Germany, Great Britain, Greece, Israel, Italy, Lithuania, Netherland, Poland, Russia, Saudi Arabia, Sweden, andSwitzerland. | Terminated | February 22, 2013 | 724 | 957990
Z-0585-2012 | 60132 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery XR650, Digital Radiographic Systems. |  |  | November 22, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue associated with the patient barrier rotating arm that is a part of the Definium 8000 and Discovery XR650 image pasting systems that may impact patient and user safety. The patient barrier rotating arms that are used to position and stabilize the patient during image pasting procedures may fail to retain residual friction at the arm pivot. Thi | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 17, 2011 to its customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The customers were instructed to ensure that all potential users in their facility are made aware of this safety notification and the recommended actions; discontinue use of the patient barrier and remove from service (away from patients) if either one the rotating arm locking levers is damaged (does not hold the rotating arm in place when engaged); take extra precautions to hold the patients barrier rotating arm during release of the locking lever and repositioning of the arm and store the rotating arms in the lowest pivot position (down) when not in use. The letter also stated that the patient barrier may continue to be used if the rotating arm locking lever securely holds the rotating arm in place when it is engaged.Additionally, GE Healthcare will modify the rotating arms to restore the residual friction so that the arm will retain position (rather than swing freely) while the locking levers are disengaged. Also, GE will update the product labeling to improve awareness and identify the hazard relative to proper storage.If you have any questions, call Center in US: 800-437-1171; Japan: 0120-055-919 and for other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. |  | 229 | 229 | Worldwide distribution: USA (nationwide) including states of: AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID,IL, IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NH, NJ, NM, NY,NC, OH, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY; and countries of: VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, NORWAY, NEW ZEALAND, NETHERLANDS, MARTINIQUE, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, HONG KONG, GREECE, GERMANY, FRENCH POLYNESIA, FRANCE, FINLAND, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BOSNIA AND Herzegovina, BELGIUM, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | October 16, 2012 | 329 | 860
Z-0430-2012 | 60134 | No | DXY | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular | PMA | Anthem CRT-P models PM3110 and PM3210 |  |  | September 23, 2011 | 2011 | St Jude Medical CRMD | St. Jude Medical has received reports of out of range pacing lead impedance (PLI) measurements. These measurements are low and out of range (&lt;200 ohms) but measure in range during visits to the clinic. The out of range readings are concluded as false. | Pending | The firm, St. Jude Medical, sent an "IMPORTANT PRODUCT FIELD CORRECTION" letter dated September 22, 2011 to its customers by FedEx priority overnight service for planned receipt on September 23, 2011. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed that in order to prevent a false reading, a new programmer software version (12.1.0.1) is available. This software when used to interrogate an Accent DR or Anthem CRT-P pacemaker will eliminate the potential for this anomaly to occur. The onetime upgrade is performed automatically on affected devices and will not change the operation of the implanted device. A St. Jude Medical Sales Representative will assist the customers in loading the new programmer software onto their Merlin programmer. Customers with questions or concerns, can contact their local St. Jude Medical representative or their Technical services department at (800) 722-3774. |  | 14000 | 14,000 units | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Australia, Asia Pacific, Belgium, Bulgaria, Canada, China, Colombia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, Hungary, India, Iran, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latin America, Lebanon, Malaysia, Netherlands, Norway, Oman, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, UK, and U.A.E. | Terminated | April 10, 2012 | 200 | 166000
Z-0607-2012 | 60660 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | SKYLight Imaging Systems, Model 882050; |  |  | December 08, 2011 | 2011 | Philips Medical Systems | During preventative maintenance, a crack was discovered in the lower portion of the old style, non-common arm for the detector assembly, which could result in the detector dropping in an uncontrolled manner. | Device Design | Urgent Medical Device Correction/Field Safety Notices, dated Dec 08, 2011, were mailed to affected customer sites on December 9th and 12th, 2011 by Return Receipt certified mail. The letter identified the affected product, described the issue and when it could occur, along with the hazards involved. Customers were advised to stop using the system and contact a Philips service representative for an immediate inspection. Customers were also to follow the Instructions For Use to avoid the issue. Philips service will contact customers to arrange a time for Philips to inspect the system. The letter should be placed in the Instructions for Use. The Field Change Order will be released on December 16, 2011. If further information or support is needed, customers should contact Customer Care Solutions Center at 1-800-722-9377. |  | 129 | 129 | Worldwide Distribution -- USA, Canada, Italy, Spain, France, United Kingdom, Sweden, Switzerland, Germany, Japan, Denmark, and the Netherlands. | Terminated | August 13, 2012 | 249 | 129
Z-0399-2012 | 60230 | No | KRZ | Calibrator for hemoglobin or hematocrit measurement. | Hematology | Hematology | 510(K) Exempt | VARIANT II Link Hemoglobin A1c Program Reorder Pack, |  |  | October 25, 2011 | 2011 | Bio-Rad Laboratories, Inc. | It was discovered that the Variant II Turbo Link Calibrator Values card contains incorrect values for IFCC units. The numbers are missing a decimal point. | Process control | Bio-Rad Laboratories, Inc. sent a Medical Device Correction Notification letter dated October 25, 2011, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The correction was initiated by the firm, customers were instructed to contact their regional Bio-Rad office for any questions.For questions regarding this recall call 510-741-6114. |  | 63 | 63 | Nationwide Distribution including MI, SC, OH, IL, MN, MD, VA, NY and WI. | Terminated | February 02, 2012 | 100 | 63
Z-0613-2012 | 60049 | [6 empty] | OPUS SmartStitch M-Connector Product Number OM-8007 is a suture passer |  |  | August 08, 2011 | 2011 | Innovatech Medical Resources L.P. | Product is being removed from the field due to an over-label issue. | Reprocessing Controls | Innovatech Medical Resources L.P. decided to recall all products resterilized and overlabeled at their facility. A firm representative would visit each customer and request that they remove the overlabeled product from their shelves and secure it. Customers could call the MSI office to get a RMA (Return Materials Authorization) for a refund on this product.For questions regarding this recall call 817-714-9066. |  | 113086 | Lot Number (QTY): 113086 (1 unit), 107304 (1 unit), 110511 (1 unit), and 112473 (2 units) | Nationwide Distribution including Texas | Terminated | September 27, 2012 | 416 | 99641488
Z-0467-2012 | 60590 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | MORPHEUS CT PICC with STYLET, 5F Dual Lumen 65 cm Insertion Kit, Catalog No./REF 1200923, STERILE, AngioDynamics Queensbury, NY 12804 USA --- COMMON/USUAL NAME: Percutaneous, implanted, long-term intravascular catheter; CLASSIFICATION NAME: Percutaneous, implanted, long-term intravascular catheter; Device Listing # D095274. |  |  | December 05, 2011 | 2011 | Angiodynamics Worldwide Headquarters | AngioDynamics distributed the affected PICC catheters that were beyond their expiration date. | Incorrect or no expiration date | AngioDynamics sent an "URGENT-MEDICAL DEVICE RECALL" letter dated December 2, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. A Recall Reply Form was attached to the letter for customers to complete and return to the firm via fax at 518-798-1360. Contact your local sales representative or AngioDynamics Customer Service at 800-772-6446 for questions regarding this notice. |  | 16 | 16 units domestically | Nationwide Distribution-including the states of Florida, Georgia, Montana, and Tennessee. | Terminated | September 27, 2012 | 297 | 16
Z-0438-2012 | 60138 | No | NRA |  |  | Orthopedic | PMA | "***BUECHEL-PAPPAS CRUCIATE SACRIFICING BEARING ***MATERIAL UHMWPe***STERILE EO***ENDOTEC***Endotec, Inc. 300 Sunport Lane, Suite 500 Orlando, FL 32809 USA***CAT.NO.***SIZE***" |  |  | September 21, 2011 | 2011 | Endotec, Inc. | On 09/21/2011 Endotec, Orlando, FL initiated a recall for the Beuchel-Pappas Mobile Bearing Knee System (BP Knee). The BP Mobile Bearing Knee System consists of unapproved mobile (rotating) bearing tibial and patellar components. | PMA | Endotec sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 21, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to examine and quarantine product subject to the recall. Additionally, a Response Form was included with the letter for customers to complete and return to the firm. Call 862-703-6730 for questions regarding this recall. |  | 60 | 60 | USA (nationwide) including the states of AZ, CA, CT, FL, NE, NJ, OH, and OK. | Terminated | September 27, 2013 | 737 | 2400
Z-0514-2012 | 60425 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS, all versions |  |  | September 20, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Unintended behavior when using syngo Imaging XS, all versions. The "Rename" functionality can cause an unintended patient merge if the renamed patient is sent to another system. | Labeling False and Misleading | Siemens Medical Solutions USA, Inc sent a Customer Safety Advisory Notice letter dated August 3, 2011 to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you detect patient data inconsistencies in your system caused by this problem, please contact Siemens Service for support.For further questions please call (610) 219-4834. |  | 38 | 38 | Nationwide Distribution including the states of CA, FL, GA, IN, IA, KS, MN, MO, NE, NV, NJ, NY, NC, OH, PA, TN, TX, and WI. | Terminated | March 05, 2014 | 897 | 38
Z-0448-2012 | 60387 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Microcyn Solution with Preservatives (Rx), 16.9 fl oz (500 mL) and 33.5 fl oz. (990 mL) |  |  | November 09, 2011 | 2011 | Oculus Innovative Sciences Inc | Unapproved drug claims and those not approved under 510K clearance. | Process control | Oculus Innovative Sciences sent an"URGENT MEDICAL DEVICE RECALL" letter dated November 15, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to identify and remove all affected product from inventory. A Business Reply Card was attached to the letter for customers to complete and return via fax to 1-415-462-5181. Call Oculus Innovative Sciences Customer Service at 1-800-759-9305 for questions regarding this recall. |  | 4890 | 1363 x 500 mL; 2037 x 990 mL. | Nationwide Distribution. | Terminated | December 07, 2012 | 394 | 19918
Z-0414-2012 | 60467 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VenaCure EVLTÂż NeverTouch-FRS PROCEDURE KIT Featuring Gold-Tipped Fiber, 65 cm, VC NT FRS 65cm Clear Non-Echo Needle, Catalog No./REF 51403005, STERILE -- AngioDynamics, Inc. 603 Queensbury, NY USA 12804. This is a Special Order. |  |  | November 10, 2011 | 2011 | Angiodynamics Worldwide Headquarters | During an in-process inspection the hemostasis sidearm assemblies were identified to have the potential for flash to exist on the housing at the base of the threads. This flash has the potential to break off and relocate within the fluid path. | Process control | AngioDynamics sent an URGENT: Medical Device Recall letters and response form (dated November 10, 2011) to all affected customers. The letter identified the affected products, the problem, and the actions to be taken by customer/user. The customers were instructed to do the following: 1) Identify, segregate and return the recalled products to ANGIODYNAMICS and 2) Complete the enclosed Recall Reply Form even if they do not have any products remaining in their possession, noting zero quantity to be returned and fax to 518-798-1360. The letter states that a returned authorization number (RA#) will be assigned to each customer in the recall notification letter, an UPS account will be provided for return freight and a replacement product will be shipped upon receipt and confirmation of the returned product. If you have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please do not hesitate to contact your local sales representative or call the Customer Service Manager, at ANGIODYNAMICS Customer Service at 1-800-772-6446. You may also e-mail your questions to customerservice@angiodynamics.com. |  | 105 | 105 units | Worldwide Distribution - USA (nationwide) and the countries of: Canada, Great Britain, France, India, New Zealand, Netherlands, Saudi Arabia, Turkey, and Singapore | Terminated | February 01, 2016 | 1544 | 12205
Z-0605-2012 | 60658 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | ETCO2 Divided Sampling Cannula with Simultaneous Oxygen Delivery, Model number 4707 7-7-24. |  |  | December 07, 2011 | 2011 | SALTER LABS | Mis-labeling: report of cases labeled as 4707 (cannulas with male luer lock connector) actually contained 4707F (with female luer lock connectors). | Process control | The firm notified the three affected customers directly by phone, followed by a recall form for each direct customer indicating the product, part number, lot number and shipment date and customer order specific for each firm. Customers were to complete and return the form to the fax number listed or call a Customer Service Representative at 1-800-628-4690. |  | 240 | 240 cases | Nationwide Distribution -- TX and VA. | Terminated | January 09, 2012 | 33 | 240
Z-0366-2012 | 60419 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer ZMR Femoral Spout &amp; Cone Bodies and Trabecular Metal Primary Hip Prosthesis Femoral Stem |  |  | November 16, 2011 | 2011 | Zimmer Inc. | Lack of assurance of proper metal fatigue strength due to metal grain structure anomaly. In 2006 a recall was initiated as a result of the grain anomaly; however, these 4 lots had been split off during normal production and had not been included in the 2006 recall. | Nonconforming Material/Component | On 11/16/11, Zimmer sent letters titled "URGENT MEDICAL DEVICE RECALL" to distributors via electronic mail and to risk managers and surgeons via FedEx indicating that the product must be immediately removed from use and quarantined." Stop using the affected lots and place affected product in quarantine." Review available product on the shelf" Identify those part/lot combinations listed in the recall communication" Immediately remove any implants found that match part/lot combinations listed in the recall communication and quarantine" Return affected lots and signature of acknowledgement to Zimmer |  | 2 | 2 (1 device is domestic) | Worldwide distribution: USA (NY), AUSTRALIA, SWITZERLAND, and DENMARK | Terminated | August 28, 2013 | 651 | 4