## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities
Z-0832-2012 | 60615 | No | NZG | Pediatric medical crib. | General Hospital | General Hospital | 510(K) Exempt | Bassinet Models: BSNT/01; BSNT/02, BSNT/03, BSNT/04 |  |  | October 14, 2011 | 2011 | Nemschoff Chairs, Inc. | To correct issues with casters, door hinges, drawer slides, and the bassinet surround. Also, to add labeling to the product regarding specific instructions for cleaning the bassinet tub and drawer weight limits. | Device Design | The firm, Nemschoff, sent an "Important Medical Device Correction" letter and email dated October 14, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to check inventory, quarantine and hold product until they receive instructions for servicing the units; once corrections are complete on the units at their locations, document all corrections on the Correction Acknowledgment Letter and return to Nemschoff, Inc. as instructed. The customers were also instructed to complete and return the enclosed Receipt Acknowledgment Form by October 21, 2011.Nemschoff will provide instructions and assistance with locating and repairing the units.If you have any questions, call Nemschoff, Inc at (800) 203-8916 and ask for our Quality Representative or your Customer Service Representative. |  | 1214 | 1214 units. | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, FL, GA, IA, IL, IN, KS, LA,MA, MD, MI, MO, MS, NC, ND, NJ, NY, OH, OK, SC, TX, VA, WA, WI, and WY; and country of: Canada. | Terminated | November 27, 2012 | 410 | 1214
Z-0935-2012 | 60842 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | FloWire Doppler Guide wires. |  |  | October 06, 2009 | 2009 | Volcano Corporation | Reports of no signal displayed on the monitor when FloWire Doppler Guide wires were connected to the Combomap. | Employee error | Volcano Corporation sent recall letters dated October 6, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and fax the attached Customer Reply Letter so that an RMA could be issued for the return of the unused affected product, and any replacement or credit could be arranged. Customers were instructed to return the Customer Reply letter even if they no longer have the affected product in their inventory. For questions regarding the use of the product. customers were instructed to contact their Sales Representative.For questions regarding this recall call 916-281-2790. |  | 71 | 14 in US; 42 to EU and 15 to Japan | Worldwide Distribution - USA including MI, AZ and FL and the countries of Japan and Europe | Terminated | February 01, 2012 | 848 | 71
Z-0934-2012 | 60911 | No | NQQ | Ultrasonic pulsed echo imaging system. | Radiology | General &amp; Plastic Surgery | 510(k) | Imalux Niris OCT Imaging System Probe, Model Number 1300, Serial Number F464019, IMALUX Corporation, Cleveland, OH |  |  | December 22, 2011 | 2011 | Imalux Corporation | During an FDA inspection, the firm was notified that the wrong tubing was used during the manufacture of the probe for the unit. | Process control | Imalux Corp contacted their one customer via telephone on December 14, 2011. They informed the customer of the problem and requested the probe returned. The probe involved was returned and received by Imalux on December 22, 2011, and will be stored offsite in quarantine.Any further questions please call (216) 502-0755 ext. 301. |  | 1 | One Unit | Nationwide Distribution only to Texas. | Terminated | January 09, 2014 | 749 | 1
Z-0942-2012 | 52807 | No | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Vigilance II Monitor Model: VIG2/VIG2E Software: versions 00.51, 00.60, 00.61, 00.62, 00.63, 00.64, 00.67, 00.69, |  |  | September 12, 2008 | 2008 | Edwards Lifesciences, Llc | Edwards Lifesciences has made the decision to initiate a field corrective action for the Edwards Vigilance II Model: Software: versions 00.51, 00.60, 00.61, 00.62, 00.63, 00.64, 00.67, 00.69, 00.71, 00.73 due to the potential of inaccurate Injectate Cardiac Output (ICO) measurements. | Device Design | Edwards initiated two Urgent Product Notification - Action Required letters to all affected customers. The letters identified the affected product, problem, and actions to be taken. The letters instructed users to immediately discontinue use of in-line temperature probes for Intermittent Cardiac Output (ICO) measurements on the Vigilance II instrument and recommends use of the bath probe only (model 9850A) to measure the temperature of the injectate solution for ICO calculations, until software update have been completed. Customers were instructed to complete and return a software upgrade form to Edwards as record of completion. For questions call Edwards Technical Support department at 800-822-9837. |  | 6934 | 6934 instruments | Worldwide Distribution - Nationwide Distribution (USA) - and the countries of: Argentina, Australia/ New Zealand, Brazil, Canada, China, Columbia, Ecuador, Europe, Indonesia, Japan, Korea, Macao, Malaysia, Mexico, Middle East, Peru, Puerto Rico, Singapore, Thailand, Taiwan and Venezuela | Terminated | February 08, 2012 | 1244 | 6934
Z-0936-2012 | 60936 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva FD Software release: PBL 6.1.1 |  |  | January 09, 2012 | 2012 | Philips Healthcare Inc. | The skin dose and skin dose rate values are calculated using an incorrect point of reference. Therefore, the AKR as displayed and stored in patient records is too low, although the dose as received by the patient is not affected. | Software design | Philips Healthcare issued on January 9, 2012 a mandatory Field Safety Notice to all affected customers. The letter identified the affected product, problem description, hazard involved and actions that should be taken by the customer/user. The letter contains important information for the continued safe and proper use of the equipment. Customers were instructed to review the following information with all members of there staff and retain a copy with the equipment "Instruction for Use". For information or support concerning this notice contact your local Philips representative 866-767-2822. |  | 28 | 28 units (US) | Worldwide Distribution - Nationwide Distribution (USA) including the states of AR, AZ, CA, CO, FL, IA, KS, LA, MI, MO, NC, NY, OH, PA, TX, VA, VT, WA and WV | Open, Classified |  |  | 28
Z-0941-2012 | 61010 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DRAD-3000E: Radrex-i X-ray is a general purpose X-ray System. Labeling reads in part "TOSHIBA***DIGITAL RADIOGRAPHY SYSTEM RADREX-i DRAD-3000E Systems Data No. MSDXR0144EA***" |  |  | December 29, 2011 | 2011 | Toshiba American Medical Systems Inc | It has been found that some of the starter cables used for x-ray tube high speed rotation are not sufficiently durable. | Nonconforming Material/Component | The recall notification titled "URGENT: MEDICAL DEVICE CORRECTION", was sent on 12/29/11 to customers via certified mail. The letter informed the customer of an issue with the starter cable used for high speed rotation on the Radrex X-Ray System. The letter also informs customers of the corrective measures: a replacement cable will be installed on your system to prevent this issue from occurring. An 800 number was provided for any questions regarding the letter. |  | 39 | 39 | Class II Recall - Nationwide Distribution (USA) - Including the states of VA, IL, CA, PA, TN, OH, LA, TX, OK, AR, AZ, MS, WV, TN, WI, NJ, MN, NY, IA, FL, and Washington D.C. | Terminated | March 24, 2014 | 816 | 39
Z-0938-2012 | 60900 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | KIMBERLY CLARK MIC GASTRO-ENTERIC FEEDING TUBE, 22F. |  |  | December 21, 2011 | 2011 | Kimberly-Clark Corporation | The product is misbranded, i.e. the Gastric (center) port may be mislabeled as the Jejunal port, and the Jejunal (side) port mislabeled as the Gastric port. | Labeling False and Misleading | Kimberly-Clark Health Care sent an "URGENT: VOLUNTARY MEDICAL DEVICE RECALL NOTIFICATION" letter and e-mail to all affected customers. The communications described the product, problem, and actions to be taken by the customers. Customers were instructed to locate and quarantine any affected product. A Product Recall Response Sheet was enclosed for customers to complete and return to the firm via fax at 920-380-6682. Contact the firm at 770-587-7400 for questions regarding this recall. |  | 50 | 50 devices | Worldwide Distribution-USA (nationwide) including the states of CA, CO, ID, IL, MD, MO, NC, NJ, NY, PA, SC, UT, VA, WA, WI, and WV and the country of Canada. | Terminated | July 02, 2012 | 194 | 50
Z-0939-2012 | 60016 | No | OGL | Portable liquid oxygen unit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Broselow/Hinkle Pediatric Emergency System |  |  | October 17, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | Povidone Iodine Pads included in the Vital Signs Broselow Pediatric Emergency Kit may be contaminated with Elizabethkinga meningoseptica. | Material/Component Contamination | GE Healthcare sent a recall letter/return response form dated October 17, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to not use the affected product, isolate and discard all affected product.Fill out the attached Confirmation form and fax back to 800-535-7923 per the instructions to obtain replacement Busse IV Start Kits. Also, consignees were instructed to return the attached response form even if no recalled product was in their inventory.If they forwarded any affected product to any other healthcare institutions, they should forward a copy of the letter to those institutions.For any questions regarding this recall call 1-800-932-0760. |  | 6675 | 6675 kits | Nationwide and Foreign | Terminated | January 15, 2013 | 456 | 6675
Z-0940-2012 | 61030 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Aera/Skyra, model numbers 10432914 and 10432915. |  |  | January 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During product monitoring firm has discovered that on a few systems the gradient cable connections did not meet their specifications after installation. Locally, the cable connections overheated, making servicing necessary. The potential existed for a cable connection to become overheated, leading to the emission of smoke and possible material damage. Firm preventively checked the gradient cabl | Nonconforming Material/Component | Siemens Medical Solutions sent a "CUSTOMER SAFETY ADVISORY NOTICE" letter dated November 24, 2010 to all affected customers. The letter described the product, problem, and actions to be taken. Contact the firm at 610-219-4834 for questions regarding this recall. |  | 128 | 128 | Nationwide Distribution-including the states of AK, AZ, CA, CT, DE, FL, GA, IL, IN, KS, MD, MA, MI, MN, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, and WI. | Terminated | January 31, 2014 | 736 | 128
Z-0937-2012 | 60969 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF 16 PET/CT System &amp; GEMINI TF 64 PET/CT System, Philips Medical Systems, Cleveland, OH |  |  | October 24, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Philips became aware of an issue relating to a UPS cabinet that contained a melted battery which caused damage to other batteries lower in the cabinet, leading to battery acid leaking on the floor. | Material/Component Contamination | Philips Medical Systems (Cleveland) Inc communitcated the problem to the facilities via telephone. No customer letter was sent since the correction was already performed by the supplier. |  | 2 | 2 | Nationwide Distribution including Maryland and Colorado | Terminated | June 18, 2013 | 603 | 2
Z-0947-2012 | 58937 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | ***CAT 7205305***DYONICS***Arthroscopic Blade***3.5mm FULL RADIUS, 13cm long***Fits Power/EP-1 Motor Drives***Beige***SAMPLE LATEX-FREE***Made in the USA***Manufacturer: ASCENT HEALTHCARE SOLUTIONS***Lakeland, FL*** Lot number: 199468***Expiration Date: 12/2011. |  |  | January 26, 2009 | 2009 | Ascent Healthcare Solutions | On 01/26/2009, Ascent Healthcare Solutions, Lakeland, FL initiated a recall of various models of Dyonics, Stryker, and Linvatech Reprocessed Arthroscopic Shaver Blades. The sterility of the devices may be compromised due to a potential breach in the seal of the Tyvek cover to the plastic tray. | Packaging process control | Ascent Healthcare Solutions sent an "URGENT MEDICAL DEVICE VOLUNTARY RECALL" letter dated January 26, 2009 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were instructed to examine their inventory, remove and quarantine all affected devices. A Recall Effectiveness Check Form was attached for customers to complete and return via fax to 1-866-325-4760. Contact the firm at (888) 888-3433 ext. 5519 for questions regarding this recall. |  | 330 | 330 devices | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NH, NV, NY, OH, OK, OR, PA, RI, SD, TN, TX, UT, VA, WA, WI, and WY, and the country of Canada. | Terminated | March 09, 2012 | 1138 | 3702
Z-0944-2012 | 60836 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra PAC1 Mod 2 Cables used with the Integra MPM, CAM01 Monitors |  |  | October 20, 2011 | 2011 | Integra LifeSciences Corp | Some of the connector cables associated with Integra's intracerebral pressure (ICP) monitors may cause temporary interference with the ICP waveform on the monitor display. | Under Investigation by firm | Integra sent a Urgent Product Recall Notification letter dated October 20, 2011, via FedEx, email, and fax to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed to examine their inventory at your earliest convenience. If you have an effected PAC1 cable, please complete the attached Recall Acknowledgement Form and return it to Integra Customer Service per the instructions on the form. An Integra Sales Representative will contact you to assist in the return and replacement of new cables at no cost. In the meanwhile, you may continue to use the cables until you have been contacted and replacement product has been provided. Distortion of the ICP wave form can be stopped or prevented by keeping the PACI catheter receptacle out of human contact.Please contact Integra Customer Service at 1-800-654-2873. |  | 221 | 221 Units | Worldwide Distribution-- USA (nationwide ) . | Terminated | June 18, 2013 | 607 | 221
Z-0946-2012 | 60977 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance Big Bore Model #728243 and #728244, Philips Medical Systems. |  |  | December 13, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The Block and Multi Leaf Collimator (MLC) shape is not propagated to the beam that is modified by the Oppose operation. The contours can appear with the incorrect color in the Beam's Eye View Viewer, resulting in organs being incorrectly placed or incorrectly shaped. | Software design | On 12/13/2011, the firm sent Field Safety Notices, entitled URGENT - Medical Device Correction, to their customers. The letter identified the affected product, the problem, the hazard involved, and the actions to be taken by customers and the firm. Customers were to follow the instructions provided in order to address or avoid the issues. Philips will be installing a software update to address the identified issues free of charge. If customers need further information or support in regards to these issues, they are to contact their local Philips representative or local Philips Healthcare office. In North America and Canada, customers can contact the Customer Care Solutions Center at 1-800-722-9377. |  | 142 | 142 Units | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, CT, FL, IA, IL, IN, KY, MA, MD, MI, MN, MO, NC, NH, NM, NY, OH, OR, PA, PR, TN, TX, VA, &amp; WA and countries of AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, DENMARK, DOMINICAN REPUBLIC, FRANCE, GABON, GERMANY, HUNGARY, ISRAEL, ITALY, JORDAN, LUXEMBOURG, MALI, MEXICO, NETHERLANDS, OMAN, SLOVAKIA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UKRAINE, UNITED ARAB EMIRATES, &amp; UNITED KINGDOM. | Terminated | February 22, 2013 | 437 | 142
Z-0945-2012 | 61012 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian Clinac Accelerators with Mark, Millennium MLC 120 (Multileaf Collimator) and BrainLAB m3 MCL, Reference/FSCA CP-05591, Model Number: H53, Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | November 10, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Failure to retract or park the Varian MLC leaves prior to use of the demountable BrainLAB m3 could result in the Varian MLC leaves interfering with the treatment field defined by the BrainLAB m3 without an interlock being asserted. | Device Design | The firm, Varian Medical Systems, Inc., sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" letter dated November 10, 2011 to all affected users. The letter describes the product, description of problem, details, recommended actions and Varian's actions. Varian Service will contact customers to arrange installation, once modifications are available.***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter also included special instructions for customers outside the USA and Canada. The customers were instructed to complete and return the attached Proof of Notification or Receipt Verification Card to Varian Medical Systems once they have read the document.If you require further clarification, feel fee to contact your local Varian Customer Support District or Regional Manager at: USA and Canada: 1-888-VARIANS (1-888-827-4265); Europe: +41 41 749 8844; or Email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: support-sea@varian.com; China/Asia: support-china@varian.com; Japan: support-japan@varian.com; Latin America: soporte.al@varian.com or Internet: Oncology Systems customer site-www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 418 | 205 units***12-12-12 AMENDED TO: 213 units*** | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Algeria, Austria, Australia, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Malaysia, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Saudi Arabia, Slovenia, South Korea, Spain, Sri Lanka, Sweden, Taiwan, Turkey, United Kingdom and Venezuela. | Terminated | February 19, 2015 | 1197 | 418
Z-0963-2012 | 60852 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sterile surgical kits: |  |  | June 14, 2011 | 2011 | DeRoyal Industries Inc | Custom surgical kits contain recalled Kimberly Clark surgical drapes which were found to contain metal shavings in the folds of the drapes. | Nonconforming Material/Component | DeRoyal Industries Inc. sent an "URGENT VOLUNTARY PRODUCT RECALL OF KIMBERLY-CLARK THREE-QUARTER DRAPE 79147" letter dated June 14, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to identify and quarantine all affected product. Further instructions by the firm asks customers to complete and return a Recall Response Spreadsheet via fax to 865-362-3744. Customers should contact their Deroyal Sales Representative for questions or assistance regarding this recall. |  | 3881 | 3881 kits | Nationwide Distribution. | Terminated | June 27, 2013 | 744 | 3881
Z-0964-2012 | 60391 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX Analyzer, Model Number: 10-004, 14-874, 14-875, in vitro diagnostic. |  |  | November 16, 2010 | 2010 | Alere San Diego | Alere has identified that a bias on results for certain analytes tested on the Alere Cholestech LDX System may exist due to humidity variability. | Device Design | Alere sent an Updated Urgent Recall Notification Letter in March 2012 to all their customers who are in possession of one or more Cholestech LDX Systems. The updated letter provides the customers with important information about the problem identified and the actions to be taken. Customers are instructed to operate their device in a controlled environment that maintains humidity within 40-60% relative humidity if the device has not been upgraded with ROM pack v3.40. If humidity cannot be monitored, the device is no longer CLIA waived. Customers are instructed to complete the customer verification form. A previous notice from Alere sent in November 2010 instructs users to contact Alere to receive the ROM pack upgrade. Customers may email LDXROM@alere.com to request an electronic form to complete and return via email. Customers can expect their ROM packs approximately 10 to 14 business days after Alere receives their request. If customers have transferred ownership to another party, they can email Alere at LDXROM@alere.com with the updated contact information. If customers are not currently using their Cholestech LDX System, it is recommended that the ROM Pack upgrade be installed prior to restarting use of any Cholestech LDX analyzer. Customers with questions regarding this communication can call Alere at (877) 441-7440, option 1 or email at LDXROM@alere.com. |  | 22000 | 22,000 meters | Worldwide distribution including USA | Terminated | August 22, 2013 | 1010 | 22000
Z-0943-2012 | 60422 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Foot Rest Kit (pN 6252700-003) for Models 6251 and 6252 Stryker (EMS Stair -PRO), OE 2065939, 1739631, 2066066, 6066803 and 20666941 with optional , Stryker Medical, Portage, MI 49002. |  |  | October 31, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | The Foot Rest Kit may have the wrong pivot spacers making the joint more susceptible to relaxation and prone to fatigue. There is a moderate risk of injury to the patient should the handles become detached during use. | Mixed-up of materials/components | Stryker Medical spoke with all customers who received these shipments prior to 10/31/2011 requesting information on kits affected and recalled foot rest units installed. December 5, 2011 Urgent Medical Device Recall Letters, with instructions on how to to either install a new Foot Rest Kit ( enclosed ) or dispose of any kits that were not installed were sent to both customers who could have installed recalled foot rest units. Customers needing help or assistance with installation are asked to call Stryker Technical Support a 1-800-327-0770. Customers were to complete and return a postage paid postcard confirming receipt. Customers with questions or concerns are asked to contact Kristin Dudek at 269-389-6927 M-F m 8 a.m. - 5 p.m. (EST) |  | 3 | 3 units | Nationwide Distribution -- MI &amp; MA. | Terminated | August 28, 2013 | 667 | 3
Z-0966-2012 | 61018 | No | LLT |  |  | Obstetrics/Gynecology | Contact ODE | Bravo pH Monitoring capsule, UPC 7290101361695 (FGS# 0312) and UPC 7290101361688 (FGS# 0313). |  |  | January 11, 2012 | 2012 | Given Imaging Ltd. | Failure of the Bravo capsule to attach to the esophagus or alternatively, failure of the capsule to detach from the placement device. | Nonconforming Material/Component | Given Imaging sent a Urgent Product Recall letter to all affected consignees. The letter identified the product, description of problem, actions to be taken and advised customers to read this important information packet in its entirety. Customers were advised to complete the Customer Response Sheet included and return per the instructions provided. Also attached to the letter was a Customer Bulletin: Update to Vacuum Settings. For questions call 1-800-448-3644 option 3. |  | 21885 | 21,885 units | Worldwide Distribution - USA (nationwide) and the countries of Austria, Belgium, Brazil, Denmark, Canada, Cyprus, Finland, Germany, Hong Kong, India, Indonesia, Israel, Italy, Japan, Mexico, Netherlands, Norway, New Zealand, Panama, Portugal, Singapore, Sri Lanka, South Africa, Spain Thailand, Turkey, United Kingdom and Venezuela. | Terminated | September 14, 2012 | 247 | 21885
Z-0967-2012 | 60987 | No | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | ACROBAT Stabilizers- |  |  | November 16, 2011 | 2011 | Maquet Cardiovascular, LLC | There may be a small perforation in the sterile pouch of select lots of the ACROBAT V, ACROBAT SUV, and ACROBAT Mechanical Stabilizers and the XPOSE 3 and XPOSE 4 Positioners. | Employee error | Maquet sent an Urgent Device Removal Immediate Action Required letter dated November 16, 2011 and a Field Action Response forms via Federal Express to to affected customers. The letter identifies the problem, product. risk factors and actions to be taken. Customers were instructed to examine their stocks immediately to determine if they have any of the recalled product. If so, discontinue dispensing (distributing) the lot and complete the Field Action Response form attached to complete and fax the attached Field Action Response form. In addition, it provides instructions for how to return the recalled product to MAQUET. Distributors are instructed to provide the same instructions to their customers supplied with the recalled products. If you have any questions contact the Recall Hotline Center at 1-888-613-0271. If you have any additional questions, contact your local MAQUET Cardiovascul'ar sales representative or Customer Service at 1-888-880-2874 (6AM -5PM PST). |  | 9436 | 9436 unit total (US AND WORLDWIDE) | Worldwide Distribution - USA Nationwide | Terminated | January 10, 2014 | 786 | 9436
Z-0965-2012 | 53127 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Total Knee Junior Model Number 1100 |  |  | August 10, 2009 | 2009 | Ossur H / F | Some units of the Total Knee Junior produced between November 25th, 2008 and through April, 2009 contain balancing units that may fail under extreme use over a long period. The fault is noncritical and no incidents or injuries resulting from this situation have been reported to Ossur at this stage. Over long term use (estimated to be in the four to six year range) the prosthetic device may fail a | Nonconforming Material/Component | Ossur sent a Safety Notice dated August 10, 2009 by mailed to all affected consignees. Ossur informed consignees of the above mentioned public reason for recall and asked consignees to take the following action: Identify the users of each of the Total Knee Junior model devices listed above and check inventory and isolate any recalled devices on hand but have not fitted on users. The letter states that Sue Fitzgerald from customer service will be contacting consignees in the next few days to coordinate replacement of the affected devices. For question call 1-(800) 219-2150, Ext: 4109. In addition, Ossur website: www.ossur.com we will be posting updates as appropriate. |  | 141 | 141 units | Worldwide Distribution - USA (nationwide) in the state of: AL, AR, AZ, CA, CO, DE, FL, GA, IL, IN, KS, KY, LA, MA, MI, MN, MO, NC, NJ, NV, NY, OH, OK, OR, PA, SC, TX, and WV and the countries of: Australia, Canada, China, France, Germany, Hong Kong, Israel, Japan, Malaysia, Netherlands, New Zealand, Portugal, Puerto Rico, Saudi Arabia, South Africa, Spain, Sweden, Taiwan, United Kingdom and United States. | Terminated | February 06, 2012 | 910 | 141
Z-0973-2012 | 60867 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano pcFM Software Kit ; |  |  | June 11, 2010 | 2010 | Volcano Corporation | Internal review of validation of pcMF Software kit found that instructions directs user to remove the rear panel, causing the potential for electric shock. | Software design (manufacturing process) | The firm, Volcano Corporation, sent an "IMPORTANT PRODUCT INFORMATION" letter dated June 17, 2010 to all customers. The letter describes the product, problem and actions to be taken. The customers were instructed to return the pcFM Software kit to: Volcano Corporation at 2870 Kilgore Road, Rancho Cordova, CA 95670 and contact (800) 228-4728 for instructions on returning the product. The customers were also instructed to complete and return the attached Customer Reply Letter form to Volcano even if they no longer use the pcFM Software Kit via fax to Volcano Customer Support at (916) 638-8812.If you have any questions, contact your Sales Representative or Customer Service at (800) 228-4728. |  | 8 | 8 units | Worldwide distribution: USA (nationwide) including states of: CA, MI, TX, and WI; and countries of: The Netherlands, Sweden and United Kingdom. | Terminated | February 07, 2012 | 606 | 8
Z-0971-2012 | 60794 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂż Chemistry Products Calibrator Kit 24, Catalog No. 680 1708. Common/Usual Name: VITROS Calibrator Kit 24. For in vitro diagnostics use only. VITROS Chemistry Products Calibrator Kit 24 is used to calibrate the VITROS 5,1 FS Chemistry System, the VITROS 4600 Chemistry System and the VITROS 5600 Integrated System for the quantitative measurement of albumin in urine. Ortho-Clinical Diagnostics Inc. |  |  | October 27, 2011 | 2011 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics received numerous complaints of lower than expected results for control fluids and patient samples upon initial calibration or following a recalibration using VITROSÂż Chemistry Products Calibrator Kit 24 and VITROSÂż Chemistry Products mALB Reagent. | Nonconforming Material/Component | All consignees were sent "Important Product Correction Notification" communications (dated 10/27/11) beginning on 10/27/11 informing them of the issue, and instructing them to discontinue use of their remaining inventory of VITROSÂż Chemistry Products Calibrator Kit 24, VITROSÂż Chemistry Products mALB reagent, and VITROSÂż Chemistry Products mALB Performance Verifiers. For questions, please contact OCD's Customer Technical Services at 1-800-421-3311. |  | 4436 | 4436 units worldwide (domestic - 2201; foreign - 2235) | Worldwide Distribution - USA, including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Terminated | February 12, 2016 | 1569 | 4436
Z-0979-2012 | 60577 | Yes | OAE |  |  | Cardiovascular | PMA | EZ Steer ThermoCool Nav Part Number D-1292-01-S Diagnostic/Ablation Catheter instructions for use. |  |  | November 22, 2010 | 2010 | Biosense Webster, Inc. | Biosense Webster, Inc. has recently become aware of the omission of the following two Warning and Precaution statements on the Instructions For Use (IFU) for the above catheter:1. If phrenic nerve location is a concern, precautionary measures are recommended to evaluate the proximity of the nerve to the ablation electrode, such as pacing maneuvers.2. To minimize risk of atrio-esophageal fist | Labeling design | Biosense Webster, Inc. sent an Urgent Field Safety Notice letter dated Novembe 22, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were advised to continue use of the affected product and that no further action is required on their part. Consignees were advised that the updated IFU can be found at www.jnjgatewayifu.com. Consignees were instructed to contact their local Biosense Webster representative if they have any questions or require additional information.For questions regarding this recall call 909-839-8604. |  | 24713 | 24,713 units in total | Nationwide Distribution | Terminated | February 08, 2012 | 443 | 24713
Z-0970-2012 | 60721 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Hemotherm 48ÂżC Dual Reservoir Heater/Cooler System, Model Number: 400 and Model Number: 400MR. The 48 o ( Degree) C Hemotherm Dual Reservoir Heater/Cooler I provides a flow of temperature-controlled water to a heat exchanger. The device, in conjunction with a heat exchanger, is used to warm/cool the fluid that enters and exits a heat exchanger. The conductive heat/cool transfer is accomplished through any type of blood oxygenator/heat exchanger unit. The operator of the unit regulates the desired fluid temperature. Frequent observation of the actual blood/fluid temperature and body temperature is required. The Hemotherm does not contact the patient of the fluids; the water being circulated by the Hemotherm is separated from the fluid by the heat exchanger. To active temperatures above 42 Degrees C, the user must follow explicit instructions in to the owner's manual in order to activate the higher setting. Without following the explicit instructions, the device can only operate within the normal maximum temperature range of: 42 degrees C. The Hemotherm can be used with adult and pediatric patients for extracorporeal circulation of the patient's limb, entire body, or a chemo medication lavage of the abdominal area. |  |  | October 27, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | The suspect hyper-hypothermia device systems sold to physicians for research and/or the treatment of AIDS and Cancer patients, were not appropriately cleared by the FDA prior to distribution. | Software design | CSZ sent an "URGENT: FIELD ACTION NOTICE" letter dated January 20, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to become familiar with the amended manual page and to get approvals from the Institutional Review Board before further use of the device. Additionally, a response form was enclosed with the letter for customers to complete and return via fax to (513)772-9119. Contact CSZ Service Department at 1-800-989-7373 for questions regarding this notice. |  | 60 | A total of 20 devices were distributed in the U.S. &amp; 3 units were distributed internationally. | Nationwide Distribution-including the District of Columbia and states of CA, GA, MA, MD, MI, NJ, NY, OH, and UT. | Terminated | February 26, 2013 | 488 | 76
Z-0975-2012 | 60919 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU Clinical Chemistry Analyzers: Model Number: |  |  | July 12, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman has confirmed that the AU Clinical Chemistry Analyzers may experience cuvette overflow. | Under Investigation by firm | Beckman Coulter sent an "UPDATE TO PCA -14563: URGENT PRODUCT CORRECTION" letter dated April 27, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter is a communication which describes an alternative approach to detecting cuvette overflow. Additionally, a response form was enclosed for customers to complete and return to the firm. Contact the Customer Support Center at (800) 854-3633 or on the Beckman Coulter website at www.beckmancoulter.com/customersupport/support/default.asp for questions regarding this notification. |  | 8225 | 8225 units | Worldwide Distribution-including USA (nationwide). | Terminated | August 04, 2014 | 1484 | 8225
Z-0978-2012 | 60011 | No | DWH | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Terumo, Sarns Sternal Saw I System, Catalog 6090. The Sarns Sternal Saw I is indicated for use in splitting the sternum during thoracic surgery. |  |  | November 17, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received one report of minute metal shavings leaking from the Sarns Sternal Saw I System at the point where the saw blade is inserted into the saw handpiece. | Equipment maintenance | On 11/16/2011, 11/23/11, 02/3/12, 03/22/12, and 4/24/12 the firm sent a letter entitled "URGENT MEDICAL DEVICE RECALL" to it's domestic direct accounts that included the name of the device, catalog and lot numbers, dates of distribution, and instructions to discontinue use and return the device to the firm for destruction, with a Fax-back response form to be submitted. Medwatch reporting information was also included in the letter. Questions and comments are to be addressed by Customer Service @ 1-800-521-2818, M-F, 8-6 EST |  | 10186 | approx. 10186 | Worldwide Distribution - USA, and the countries of Colombia, Argentina, Lebanon, Belgium, Chile, India, Sri Lanka, Jordan, Egypt, Mexico, Brazil, Taiwan, Honduras, Canada, Germany, Singapore, Republic of Georgia, Hong Kong, Japan, UAE, Philippines, Thailand and South Africa. | Terminated | December 06, 2012 | 385 | 10186
Z-0972-2012 | 60662 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens CLINITEST hCG -Pregnancy Test (qualitative), Radio immunoassay,Human Chorionic Gonadotropin (hCG) test system. Package of 25. Product Code 1760/ Catalog# 06484105). |  |  | December 15, 2011 | 2011 | Siemens Healthcare Diagnostics | A potential for false positive or increase in borderline results with patients and commercially available negative control material. | Process control | Siemens issued Overnight Mailings of Urgent: Recall Notices on December 15, 2011 for delivery beginning December 16, 2011 to Distributors and end users. Customer bulletin has been issued instructing customers to discontinue use of product and discard impacted inventory. Customers are to contact their distributor regarding credit for unused product. For technical questions contact Siemens Technical Solutions Center at 877-229-3711. |  | 48578 | 48,578 Kits (25 tests / kit) | Worldwide Distribution -- US, Austria, Germany, Great Britain, Denmark, Norway, Sweden, Poland, Netherlands, Luxembourg, Belgium, Finland, France, Greece, Ireland, Italy, Portugal, Qatar, Saudi Arabia, Slovakia, Slovenia, and Switzerland. | Terminated | March 30, 2015 | 1201 | 48578
Z-0984-2012 | 60800 | No | LEG | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | VITROSÂż Valproic Acid (VALP) Reagent, REF/Catalog No. 6801710 --- Ortho- Clinical Diagnostics Rochester, NY USA -- Ortho- Clinical Diagnostics High Wycombe, UK --- An in vitro diagnostic. |  |  | May 10, 2011 | 2011 | Ortho-Clinical Diagnostics | Results using VITROSÂż Chemistry Products VALP Reagent were negatively biased compared to other manufacturers. OCD confirmed that patient results were also affected by this issue. | Manufacturing material removal | All domestic end-user consignees were initially sent "Important Product Correction Notification" recall communications (dated 5/10/11) by Federal Express overnight mail on 5/10/11. The letters informed users of the issue, and advised them to discontinue using their VITROS VALP reagent until adjusted calibrator values are provided for GEN 14 of VITROS VALP Reagent. The adjusted calibrator values for VITROS VALP Reagent were sent to customers via Assay Data Disk (ADD) Data Release Version (DRV) 5632. Customers were to refer to the Technical Support Communication for details. Users were to confirm receipt of the notice by completing and returning the Fax Reply Form - Confirmation; confirm that all customers or distributors were informed of the issue and instructed to discontinue use of affected product; and lastly confirm that all customers or distributors were sent adjusted calibrator values. |  | 5800 | Approximately 5800 units worldwide (domestic - 4549, foreign - 1287) | Worldwide Distribution -- US, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Terminated | February 12, 2016 | 1739 | 5800
Z-0985-2012 | 60898 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | MultiDiagnost Eleva IITV (MD-EII) with Software Release: PBL 6.1.1 |  |  | January 09, 2012 | 2012 | Philips Healthcare Inc. | Incorrect Skin Dose calculation. AKR (Air Kerma Rate) displayed and stored in patient records is too low | Software design | Philips Healthcare issued a mandatory Urgent Field Safety Noticedated October 3, 2011 to all affected customers. The letter identifies the affected product, problem description, hazard involved, action to be taken by customer/user and actions planned by Philips to correct the problem. The letter provide important information for the continued safe and proper use of the equipment. The letter instructs customers to review the following information with all members of their staff who need to be aware of the contents of this communication and retain a copy with the equipment Instruction for Use. For further information or support concerning this issue contact your local Philips representative at 866-767-2822. |  | 10 | 10 units (US) | Worldwide Distribution - USA (nationwide) including the states of CA, IL, OH, SD, TX, WY and the countries of: Australia, Austria, Canada, Egypt, Germany, mexico, Turkey and UK | Open, Classified |  |  | 10
Z-0148-2012 | 59982 | No | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | GE Healthcare Signa, Discovery and Optima. |  |  | September 15, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of missing labels on the alignment lasers. | Labeling design | GE Healthcare will execute a field action in which GE Healthcare Field Engineers will perform an inspection of all suspect MR systems. Any systems lacking any of the 3 laser warning labels will be remediated by applying the missing labels). We expect this field action will be completed by June 30, 2012.1. Notification of all distributors and purchasers be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. The FDA district office coordinator noted below is to be included in the notifications.2. Corrections be made at no cost to the purchasers. |  | 4614 | 4614 | Nationwide Distribution | Terminated | June 29, 2012 | 288 | 4614
Z-0986-2012 | 60938 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Integra OS Osteoconductive Scaffold Putty (2.5 cc and 5 cc). Integra OS Putty, combined with bone marrow aspirate, is intended for use as a bone void filler to fill voids or gaps of the skeletal system. Following placement in the bony void or gap (defect), Integra OS Putty is resorbed and replaced with bone during the healing process. |  |  | October 07, 2011 | 2011 | Integra LifeSciences Corp | Integra OS Osteoconductive Putty is labeled with incorrect storage conditions. | Error in labeling | Integra issued an "Urgent: Product Recall Notification" to customers and sales specialists on 10/7/2011. They were advised of the problem and provided instructions on how to handle the affected product. |  | 45 | 45 units | Nationwide Distribution - Including the states of CA, OK, MD and VA. | Terminated | May 21, 2013 | 592 | 45
Z-0987-2012 | 59092 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ***REF C9944A***"FULL RADIUS RESECTOR"***2.9mm***Qty 6***STERLING***STERILE***Rx Only.*** Made in the USA***Manufacturer: CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA***ConMed Linvatec Biomaterials, Ltd.Hermiankatu 6-8L 33720 Tampere, FINLAND.***Lot number # 95987. |  |  | November 12, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Largo, FL recalled catalog number, C9944A, lot # 95987, 2.9MM Full Radius Resector Blade in November 2009 because there was a possibility that the device was assembled with an incorrect outer tube. | Employee error | A Linvatec Corporation dba ConMed Linvatec sales representative contacted the affected customers by telephone on November 12, 2009. The customers were also contacted by the Recall Coordinator. The product, problem, and action to be taken by the customers were discussed. Customers with the product in their inventory were instructed to return the product to the firm. Contact the firm at 727-399-5205 for questions regarding this recall. |  | 30 | 30 Units | Product Distributed in the states of Georgia and New York. | Terminated | February 13, 2012 | 823 | 30
Z-0988-2012 | 60902 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Gamma camera systems: AXIS SPECT Systems &amp; IRIX SPECT Systems, Model #: AXIS- 882130, IRIX- 882140. Philips Healthcare. |  |  | November 18, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Philips was informed that the detector buckets may experience weakness over time. As a result, this could cause the detector bucket to separate from the mounting plate of the gantry and eventually result in the detector falling. | Nonconforming Material/Component | On 11/15/2011 Philips sent Urgent - Medical Device Correction letters to their customers. The letters identified the affected product, the problem, the hazard involved, and the actions to be taken by users and the Philips. Customers are to place the letter in their User Documentation. Philips service will arrange a time to inspect customers' systems and conduct the appropriate field safety correction. Customers are to contact their local Philips representative if they need further support or information. In the US and Canada, customers can contact Philips Healthcare Customer Care Solutions Center at 1-800-722-9377. |  | 795 | 795 units | Worldwide Distribution -- USA, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PRO, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY and the territory of Puerto Rico; and countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Egypt, Egypt, Finland, France, Germany, Greece, Iceland, India, Ireland, Israel, Italy, Japan, Lebanon, Luxembourg, Malaysia, Mexico, Netherlands, Peru, Philippines, Poland, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, &amp; United Kingdom. | Terminated | February 22, 2013 | 462 | 795
Z-0990-2012 | 60693 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Flow Tec Nano-Hybrid Flowable Affected Product |  |  | March 03, 2011 | 2011 | Sybron Dental Specialties | The recall was initiated because Pentron Clinical has confirmed that the Flow Tec Nano-Hybrid Flowable Composite material has been found to be difficult to extrude. | Nonconforming Material/Component | Pentron Clinical sent an initial "URGENT: MEDICAL DEVICE RECALL" letter dated April 20, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and fax back an enclosed recall return form in order to confirm their receipt of the recall notification. A second "URGENT: MEDICAL DEVICE RECALL" letter dated November 10, 2011 was sent to all affected customers because Pentron Clinical expanded the recall. On December 8, 2011 another recall communication was forwarded to all affected customers to change the wording on the letter to make it more clear to the customers about the recall expansion. Contact Pentron Clinical Customer Service at 1-800-551-0283 for questions regarding this recall. |  | 77289 | 77,289 Units total | Worldwide Distribution: USA (nationwide) and the countries of Canada, Egypt, France, Germany, Italy, and Spain. | Terminated | October 05, 2012 | 582 | 77289
Z-1008-2012 | 60929 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | S1827 Crea Calibration solution 1, 944-135, Lot DV-02 Radiometer Medical, Bronshoj Denmark |  |  | December 16, 2011 | 2011 | Radiometer America Inc | When the user activates the dosicap (containing the crea additive) and mounts it onto the bottle, the threads will not seal correctly causing approximately 10% of the solution to leak out when the bottle is placed horizontally for dissolving the additive. This will cause the concentration of creatinin and creatin in the calibration solution to be wrong, and results in a bias on the Crea results. | Device Design | Radiometer sent a Medical Device Recall letter dated December 28, 2011 to their customers. The letter provided details on the affected devices, description of the problem and actions to be taken. The firm instructed their customers to immediately examine their stock and return any remaining product to Raiometer America Inc. for disposal. The letter states that Radiometer will replace all bottles of the affected lots, at no charge. Customers were asked to complete the attached Recall Response Fax Form and fax to 800-736-0601. For questions regarding the Recall contact Radiometer America Technical support at 1-800-736-0600 opt. 4. |  | 42 | 7 boxes/6/175mL bottles | Nationwide Distribution (USA) including the states of: NY, PA and TN. | Terminated | January 17, 2014 | 763 | 42
Z-1000-2012 | 59936 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Artiste Flowable Value Pack Part Number N28VA, Syringe Lot Number 165741 |  |  | October 22, 2010 | 2010 | Sybron Dental Specialties | The recall was initiated because Pentron Clinical has confirmed the Flow Tec Nano Hybrid Flowable Composite material has been found to be difficult to extrude. | Nonconforming Material/Component | Sybron Dental sent a Recall Letter dated November 15, 2010 to consignees informing them of the above mentioned public reason for recall and asking them to determine if they have any affected products in their inventory and requesting their cooperation in returning any affected product which they still have in their stock. Product was to be returned and replaced at no charge.They also requested that consignees complete and fax back an enclosed return from , whether or not they had any product to return.Consignees were told to call Pentron Customer Service at (800) 551-0283 directly to handle the arrangements of a quick return and replacement.Authorized Petron Clinical distributors were asked to identify and recover the affected product lots that may have been shipped to their customers. |  | 771 | 771 units in total | Worldwide Distribution - USA (nationwide) and the countries of Cyprus, Germany, Italy, Poland, Canada, Ecuador, and Kuwait | Terminated | February 13, 2012 | 479 | 771
Z-1011-2012 | 60835 | [6 empty] | Zoll pedi.padz Reduced Energy Multi-Function Electrodes (MFE) |  |  | December 30, 2011 | 2011 | Bio-Detek, Inc. | Lack of an adequate silicone seal around the periphery of the attenuator enclosure may allow fluid to enter. | Process control | Zoll Medical issued Urgent Device Correction notices dated 12/30/11 via UPS/Registered Mail. The letter described the issue and identified the affected product. ZOLL recommended that customers alert their staff of the issue, that the electrodes should be kept away from a fluid environment, to use caution in handling the electrodes if they are subjected to a fluid environment and maintain safe defibrillation practice, and seek a replacement if available. Also customers should contact Zoll 1-800-348-9011 for product replacement. The firm plans to replace the affected lots free of charge. |  | 3080 | 3,080 units | Worldwide Distribution -- US and Canada. | Terminated | December 11, 2012 | 347 | 3080
Z-1003-2012 | 61038 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Custom surgical kits: |  |  | January 10, 2012 | 2012 | DeRoyal Processing Center | The firm distributed surgical kits which contained DermaHook Neuro Elastic Retractors which were subsequently recalled by Teleflex. | Nonconforming Material/Component | DeRoyal Industries sent a Customer letter dated January 6, 2012, to the affected customer. The letter identified the product the problem and the action needed to be taken by the customer. An email was also sent on January 10, 2012, to the affected customer. The customer was requested to follow these instructions. 1) Identify any affected inventory using the product and lot numbers on Attachment 1.2) Attach the provided warning label to each affected kit and tray.3) Forward this recall to any of your end users that may have affected product.4) Complete and return Attachment 1 by fax (865-362-3716) or email (recalls@deroyal.com) within two weeks of the date of this letter, even if you no longer have affected inventory.If you have questions or need assistance with the recall, including additional labels, please call (865) 362-1037 . We apologize for any inconvenience this may caused you. |  | 214 | 214 units | (USA) including Indiana | Terminated | November 05, 2012 | 300 | 214
Z-1012-2012 | 61044 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Model #03001-02: Periophil Biphasic, Synthetic Biphasic Dental Bone Graft Granules, 1.0 cc. |  |  | January 11, 2012 | 2012 | Cytophil, Inc | Cytophil Corporation is recalling Periophil Biphasic Periophil Biphasic, Periobase Biphasic, and Periobone Biphasic because devices were inadvertently shipped after receiving an inadequate dose of radiation and this product may not be sterile. | Process control | Consignees were sent an "Urgent Medical Device Recall" letter dated January 11, 2012 by Cytophil. The letter described the problem and the product involved in the recall. The firm advised consignees to immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, they were requested to identify their customers and notify them at once of this product recall. For questions, call 262-642-2765. |  | 547 | 547 | Nationwide Distribution - Including the states of WI, GA, IL, SC, MO, and TX. | Terminated | May 01, 2012 | 111 | 547
Z-1010-2012 | 61004 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical lighting systems iLED and TruLight in use with the ALC+ function. |  |  | January 09, 2012 | 2012 | Trumpf Medical Systems, Inc. | The ALC+ function may not be properly locked during use resulting in the possibility of the handle falling off into the surgical field causing injury to the patient. | Device Design | The firm, TRUMPF Medical Systems, Inc., sent an "Urgent Safety Information" notice/package dated January 9, 2012 to its customers. (Note: Some were delivered by TRUMPF representatives in person and some via mail with Urgent Safety identification on the envelope). The notice describes the product, problem and actions to be taken. TRUMPF will replace all sterilizable handles of this type in the field with new handles, in addition, TRUMPF will send new handles and verify disposal of the old version at the customers location. Distribution of new handles will start January 2012. The customers were instructed to ensure that all users and other relevent persons are informed of the Urgent Safety Information; store the information until the measure has been completed and complete and return the Confirmation of User Receipt form via mail to TRUMPF Medical Systems, Inc. 1046 LeGrand Blvd., Charleston, SC 29063 or email: Lindsey.ronnenberg@us.trumpf-med.com.If you have any questions, call 843-822-6939. |  | 8733 | 8,733 units | Worldwide distribution: USA (nationwide) including states of: AL, CA, CO, CT, DE, FL, GA, IL, IN, MA, MD, MO, NC, NM, NJ, NY, OH, PA, TX, VA, VT, WA, and WI; and countries of: Australia; Austria, Belguim, China, Cyprus, Denmark, Estonia, Finland, France, Germany, Great Britain, Hong Kong, Indonesia, India, Ireland, Italy, Japan, Kazakhstan, Kenya, Korea, Latvia, Luxembourg, Mexico, Netherlands, Norway, Pakistan, Peru, Philippines, Poland, Romania, Russian Federation, Saudi Arabia, Seychelles, Singapore, South Africa (Zuid Afirka), Spain, Sweden, Switzerland (Confederation of Helvetia), Taiwan, Thailand, Tunisia, Turkey, Uruguay, Venezuela, and Viet Nam. | Terminated | October 09, 2012 | 274 | 8733
Z-1009-2012 | 60988 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD 30G1/2" Precision Glide" Needle |  |  | October 08, 2011 | 2011 | Becton Dickinson &amp; Company | Becton Dickinson has identified clogged needles. | Device Design | Becton Dickinson sent Urgent Product Recall letters dated October 18, 2011 and October 17, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory to deterine if they have any of the affected product. If they have affected product customers should return the product to Beckton Dickinson. For any questions regarding this recall call 201-847-4344. |  | 480000 | 480,000 units | Nationwide Distribution including IL and NJ | Terminated | December 11, 2012 | 430 | 480000
Z-1004-2012 | 60943 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CellTracks Analyzer II. |  |  | August 17, 2011 | 2011 | Veridex, LLC | Control cell images may i intermix with Patient cell images while changing between the Control review screen and Patient review screen on the Cell tracks Analyzer II. | Software design | Veridex, LLC sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated August 17, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter provides procedure sheets with step by step directions for customers to follow. Additionally, a Confirmation of Receipt form was enclosed for customers to complete and return to the firm. Contact Customer Technical Services at 1-877-837-4339 for questions regarding this notice. |  | 37 | 37 Units | Worldwide Distribution. | Terminated | June 17, 2015 | 1400 | 37
Z-0997-2012 | 60901 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DUET TRS 60 4.8MM STRAIGHT SULU |  |  | January 12, 2012 | 2012 | Covidien LP | Duet TRS Straight and Articulating Single Use (SULU)Loading Staplers used in thoracic surgery may result in serious injury or death | Under Investigation by firm | Covidien issued an Urgent Medical Device Recall letter dated January 12, 2012 to all affected customers. The letter identified the affected products, problem, and actions to be taken. Customers were instructed to immediately discontinue use of the device in thoracic surgery, Immediately advise all surgeons/surgery personnel of this recall and return products intended for thoracic use to Covidien by contacting Customer Service at SDFeedback@Covidien.com or 1-800-722-8772, option 1, to obtain a Return Goods Authorization prior to returning the affected units. Customers were advised to complete the Recall product return form in its entirety and return it with the returned units. For questions contact your Covidien representative or Quality Assurance at 203-492-5232. |  | 9687 | 9,687 units | Worldwide Distribution - USA (nationwide) and the countries of: ÂżeskÂż republika, ASSAGAY, Australia, Austria, Belgium, Bosnia, Herzegovina, Bulgaria Cairo,Egypt, Canada, Canary Islands, China, Croatia,Cyprus, Czech Republic, Denmark, Egypt, England, Finland, France,Germany, Guadeloupe, Hong Kong, Hungary, Iceland, India, Iran, Ireland,Israel, Italy, Japan, Jordan, Kingdom of Bahrain, Korea, Republic of Kuwait,Latvia, Lebanon, Libya, Luxembourg, Nederland, Nelspruit, New Redruth,New Zealand, Norway, Nouvelle Caledonie, Philippines, Poland, Portugal, Qatar, Reunion, Russia, Romania, Rustenburg, Saudi Arabia, Serbia,Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Turkey, Taiwan, Thailand, Tunisia, Uae, Uae - Dubai, Umhlanga Rocks,United Kingdom, Vietnam and Wales. | Open, Classified |  |  | 666880
Z-1014-2012 | 50701 | No | NOP | West Nile virus serological reagents. | Microbiology | Microbiology | 510(k) | West Nile Virus IgM Capture ELISA |  |  | October 08, 2008 | 2008 | Inverness Medical | It has come to our attention that users of these product lots may experience an elevation in the rate of false positivity when compared with previous product lots. Internal quality assurance testing has shown that the products are performing within the release specifications, although there has been a slight shift in the cut-off which may lead to an increase in the number of positive results obta | Nonconforming Material/Component | Inverness Medical sent an Urgent In Vitro Diagnostic Recall letter dated October 6, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use, destroy any unused or partially used test kits of the affected lot(s), and complete the table and certification of compliance. Customers were asked to fax back the completed form to inverness Medical at 1-609-527-8054. Replacement stock would be forwarded to them at no cost when the new lot becomes available.For questions regarding this recall call 609-627-8038. |  | 289 | 289 Kits | Worldwide Distribution including ND, PA, TX, MN; and the country of Canada | Terminated | February 15, 2012 | 1225 | 289
Z-1013-2012 | 59094 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | CI healthcare***Model: CI8092***S/N: 5-60-85-044 MFD: JUN, 2010 MADE IN CHINA***cihealthcare.com |  |  | February 17, 2011 | 2011 | Cardiopulmonary Instrumentation, Inc | Cardiopulmonary Instrumentation, Inc. dba CI Healthcare, Miami, FL, is recalling their Stainless Steel Patient Transport Stretcher, model # CI-8092, Lot 002, because of a malfunction that occurs when weight is applied to the backrest. This malfunction causes the backrest support bar, when secured in the first hook, to fall to a flat position. | Device Design | CI Healthcare sent an "URGENT RECALL NOTIFICATION" letter dated February 17, 2011, to its customer. The letter describes the product, problem and actions to be taken. The customer were instructed to do the following: 1) Examine their inventory and isolate the product(s). 2) If they have distributed any of the products to other services or facilities, please forward this information as appropriate. 3) Discontinue use of the product until a correction has been properly executed and 4) Complete and return the "Recall Response Form" as soon as possible. Should you have any questions, please call Robert Nunez at (305) 592-8196. |  | 72 | 72 Units | Nationwide Distribution - to one distributor in Puerto Rico | Terminated | April 19, 2012 | 427 | 72
Z-1032-2012 | 60996 | Yes | OAD |  |  | Cardiovascular | PMA | The 1500T9 Cardiac Ablation Generator, RO#lBl-89000 is part of the Therapy Cardiac Ablation System, along with the Cool Point irrigation Pump, the Therapy 1300 Series Steerable Ablation Catheters, the Therapy Cool Path Ablation Catheters and accessories. The generator is a microprocessor-controlled device that produces a continuous unmodulated radiofrequency (RF) output delivered from the generator between its tip electrode and a dispersive pad for the treatment of atrial flutter. |  |  | January 05, 2012 | 2012 | St Jude Medical Inc | Two St Jude Medical 1500T9 RF Generator units were distributed to customers while still under FDA review as part of a PMA supplement to the approved PMA applications. | Vendor change control | The notification letter was dated January 5th, 2012 and titled "Product Recall RF Generator 1500T9-CP". The letter informed customers that a St Jude Medical 1500T9-CP RF Generator unit had been distributed to them while currently under FDA review as part of a PMA supplement application. The letter requested that the use of the impacted generator be discontinued. Customers were requested to return the field action form and return the unit identified in the letter. The contact number provided in the letter is: 651-756-2000. |  | 2 | 2 units | Nationwide Distribution, including the states of Kansas and Maryland. | Terminated | September 04, 2014 | 973 | 2
Z-1016-2012 | 60947 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | DYNASTY(R) BIOFOAM(R) SHELL, SIZE: 46mm GROUP B, REF DSBFGB46, Rx ONLY, STERILE R, Wright Medical Technology, Inc., hip prosthesis acetabular component. |  |  | January 03, 2012 | 2012 | Wright Medical Technology Inc | The possibility exists that particulate debris from the packaging spacer could be unintentionally implanted into a patient and could potentially lead to complications that might necessitate a revision surgery. | Packaging | Wright Medical sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated December 27, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product and to return any affected product to the firm. A Product Recall Verification Form was enclosed for customers to complete and return via fax to 901-867-7401. Contact Customer Service at 1-800-238-7117 for questions regarding this recall. |  | 160 | 160 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Brazil, Canada, Denmark, and Italy. | Terminated | March 13, 2013 | 435 | 5342
Z-1036-2012 | 60946 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Viterion 100-BGM TeleHealth Monitor, Model No. 09377512, Product Code/Part No. 81842478, Bayer HealthCare LLC 555 White Plains Road, Tarrytown, NY 10591-5097 USA --- Manufactured in Japan for Bayer HealthCare LLC -- Common Name: Viterion100 TeleHealth Monitor -- Classification Name: Monitor, Physiological, Patient (Without Arrhythmia Detection or Alarms) [ref. 21CFR 870.2300] |  |  | May 25, 2011 | 2011 | Viterion TeleHealthcare Llc | Software defect; Specifically, blood glucose (BG) results from the Roche Aviva meter that are tagged by the user with either the meter hypoglycemic or general "asterisk" flags (*) will not be automatically transmitted through the Viterion 100-BGM monitor containing software version 3.2.0 or displayed to the health care provider through the data server. | Software design | On May 12, 2011, Bayer Healthcare (aka Viterion) telephoned their consignees regarding the issue with the Viterion 100-BGM and Roche ACCU-CHEK Aviva Flagged Data and followed-up with a letter (dated May 27, 2011) via e-mail. The letter notified customers of the issues and provided information on how to work around the problems. The letter provided the following recommendations: (1) The information concerning the general asterisk (*) flag be communicated within their organization as an amendment to the current patient Safety Alert Notice that was distributed and posted in April 2011. (2) For patients who need to use the Aviva Hypo and/or the asterisk flag, please assign Manual Data Entry for blood glucose on the V100-BGM. The letter informs customers that Bayer has completed the software development to rectify the aforementioned data transmission issue and will upgrade all affected products. Bayer request that customers confirm their receipt of this memo in writing and for any questions concerning this situation, contact Thuy Nguyen, Manager Quality Assurance and Regulatory Affairs at 914-333-6821. |  | 3905 | 3905 units domestically | (USA) Nationwide Distribution | Open, Classified |  |  | 3905
Z-1040-2012 | 61098 | Yes | NIO |  |  | Cardiovascular | PMA | S.M.A.R.T. CONTROL NITINOL STENT SYSTEM ILIAC: Self-Expanding Vascular Stents delivered by catheter for the indication of improving luminal diameter in patients with symptomatic atherosclerotic disease of the iliac arteries. catalog numbers are CxxyyyMB, CxxyyySB, CxxyyyML, and CxxyyySL, where xx is the stent diameter in mm and yyy is the stent length in mm; but not including catalogs where yyy is 120 or 150 . The Removal only includes Catalog Numbers in the above sequence with S.M.A.R.T.Âż CONTROLÂż in the logo/description. All unexpired lots from 15112447 through 15525516. |  |  | January 31, 2012 | 2012 | Cordis Corporation | Cordis Corporation initiated a recall of S.M.A.R.T. CONTROL Nitinol Iliac Stent System due to a potential sterility breach in the pouches. | Packaging | All affected consignees were notified by an "Urgent Medical Device Recall (Removal) Letter" dated January 31, 2012 and mailed by overnight delivery on February 1st and 2nd, 2012. Consignees were instructed to identify and quarantine affected product; review, complete, sign, and return an Acknowledgement Form; return the affected product per instructions. Questions should be directed to Customer Service at (800) 551-7683. |  | 132000 | 132,000 units | Worldwide Distribution - USA, and the countries of Japan, Canada, Mexico, South Korea, and Taiwan. | Terminated | August 08, 2014 | 920 | 150000
Z-1033-2012 | 61129 | No | OHP | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | Dental | Dental | 510(k) | Provent -SR - Sleep Apnea Therapy-SR (standard resistance). |  |  | February 03, 2012 | 2012 | Ventus Medical, Inc. | A discrepancy was found during inspection of a returned unit that the lot number and SKU on cartoon did not match. | Labeling mix-ups | Ventus Medical sent a Urgent Medical Device Correction Letter dated February 3, 2012, to all affected customers. The letter identified the problem, the product and the action needed to be taken by the customer.There is no health hazard associated with the affected devices or with the continued use of the affected devices.Ventus Medical wanted Customers to be aware of the problem so they can respond to patients if inquiries are received. A sample letter to patients was also provided, should physicians choose to provide it to patients. Questions may be directed to Ventus Medical Territory Business Manager at 650-632-4199. |  | 288 | 288 units | Nationwide Distribution including the states of OH, TX, IL, OR, CA, FL and NY. | Terminated | February 15, 2012 | 12 | 288
Z-1038-2012 | 60942 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Temporary Pacing Wire, BIPOLAR, Models BM, V, VE and VF, Rx Only, STERILE/EO, Single use only -- Mce (MEDICAL CONCEPTS EUROPE) --- Manufacturer ECM B.V. Oost-Om 54, P.O. Box 53, NL 5420 AB GEMERT --- Shipping or Unit Package: 20 individual units packaged per box with 6 boxes per case. --- Device Classification Name: Electrode, pacemaker, temporary. ---The device is supplied sterile and intended for single use. |  |  | April 18, 2011 | 2011 | European Custom Manufacturing B.V. | The crimping operation during product production was not performed properly. As such, the pacing wire may separate from the connection pin, which causes non-functionality. | Process design | ECM sent an "ECMbv RECALL OF MCE TEMPORARY PACING WIRES" letter dated May 2, 2011 to all affected customers. Customers were also informed of a potential problem via e-mail and/or phone between April 18 and 21, 2011. The letter describes the product, problem, and actions to be taken by the customers. Customers were instructed to place the affected product on hold and to return the product to the firm. For questions on this matter, contact ECM or MCE Customer Service at 1-877-886-1985. |  | 2120 | 2,120 individual units (106 boxes) | Nationwide Distribution-including the states of AL, AR, CA, FL, GA, KY, NC, NV, NY, TX, VA, and WI. | Terminated | February 15, 2012 | 303 | 2120
Z-1037-2012 | 61117 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Orthofix rotating cutter used in spinal fusion procedures with the FORZA Spacer System. Catalog Number 89-0452 (8mm Rotating Cutter), 89-0454 (10mm Rotating Cutter), and 89-0030 (FORZA Discectomy Case). Manual surgical instrument comprised of stainless steel and titanium. |  |  | February 06, 2012 | 2012 | Orthofix, Inc. | The 8mm and 10mm Rotating Cutters may fracture during use, specifically during disc space preparation. If an 8mm or 10mm Rotating Cutter does Fracture during use, it may result in a surgical delay estimated to be 30 minutes or greater due to preparing the disc space using other instruments or methods and/or retrieving the instrument fragment. | Device Design | Orthofix mailed recall letters to the distributors on 02/06/2012. Orthofix does not sell or consign these instruments to hospitals. |  | 52 | 52 | Nationwide and Puerto Rico | Terminated | June 11, 2012 | 126 | 52
Z-1034-2012 | 61111 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301A: AQUILION ONE; System, X-Ray, Tomography, Computed |  |  | January 03, 2012 | 2012 | Toshiba American Medical Systems Inc | The recall was initiated because Toshiba has confirmed that their is a potential software issue for the TSX-301A Aquilon One and the TSX-301B Aquilon One. | Software design | Toshiba sent an Urgent Medical Device Correction letter to all affected customers. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to complete and fax the attached form to (877) 349-3054. The form can also be sent via email to raffairs@tams.com.Customers with any questions regarding this letter were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 82 | 82 units total | Nationwide Distribution - USA (nationwide) and Puerto Rico | Terminated | April 08, 2013 | 461 | 82
Z-1042-2012 | 60866 | No | ETP | Air or water caloric stimulator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Leica M822 Surgical Microscope Device Controller Software (MDC version 3.6, pack F, affecting the control unit component, Article #10448446, in combination with a Leica M822 Optics Carrier-XY-Focus-Tilt-Unit, article #10448587; Leica Microsystems (Schweiz) AG, Max Schmidheiny-Strasse 201, CH-9425, Heerbrugg, Switzerland |  |  | December 21, 2011 | 2011 | Leica Microsystems, Inc. | Leica Microsystems received complaints stating that the zoom function of the Leica M822 surgical microscope locked up during use. | Device Design | The firm, Leica Microsystems, Inc., sent a "Medical Device Correction" notice dated December 21, 2011 to all affected customers by e-mail. The product, problem, and actions to be taken by the customers were addressed. Customers were instructed- DO NOT USE use the affected product until the software could be upgraded and complete and return the Acknowledgement Form via fax to: 1-847-607-3147 or Scan and send to elizabeth.culotta@leica-microsystems.com. Leica Microsystems representatives will contact the customers to make necessary arrangements to complete the software upgrade.If you have any questions, contact Director, Global Product Quality Management at 1-847-317-7209. |  | 29 | 29 units (10 units-USA) | Worldwide distribution: USA (nationwide) including states of: California, Colorado, Illinois, New York, Oregon, South Carolina; and countries of: Australia, Japan, Korea, Pakistan, Romania, South Africa, Switzerland and Turkey. | Terminated | February 16, 2012 | 57 | 29
Z-1045-2012 | 60990 | No | JWX | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | SLIDEX Staph kit (250 Tests). |  |  | December 21, 2011 | 2011 | bioMerieux, Inc. | Reports of false negative results in organism identification. | Nonconforming Material/Component | Biomerieux sent an "URGENT PRODUCT REMOVAL NOTICE" dated December 21, 2011 to the affected customer. The letter describes the product, problem, and actions to be taken by the customer. The letter instructs the customer to stop using and to destroy the affected products. The firm recommends steps to follow when interpreting results. An acknowledgement and Product Replacement Form was enclosed for the customer to complete and return to the firm. Contact Biomerieux Customer Service representative at (800) 682-2666 for information regarding this notice. |  | 100 | 100 kits | Distributed to one customer in California. | Terminated | September 17, 2013 | 636 | 100
Z-1050-2012 | 60599 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | "***REF NL82059-03***INTEGRA NEUROSCIENCES***Touhy Needle***Non Sterile for Single Use Only***14g, 6"***Plainsboro, NJ***" |  |  | November 28, 2011 | 2011 | Integra Neurosciences PR, Inc. | The current packaging configuration for the non-sterile shunt accessories could lead to the potential for confusion regarding product sterility. | Labeling mix-ups | Integra sent an Urgent Product Recall Notification letter dated November 28, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to fill out and sign the attached Recall Acknowledgement Form noting the part numbers, lot numbers and quantity of the affected product in their inventory. Customers were asked to contact Integra LifeSciences Corporation International Customers Service at 609-936-5400 to arrange for Returns. If customers do not have any of the affected product in their inventory they are required to complete and sign the form and indicate they have NO affected product in their inventory. Customers should fax or email their Recall Acknowledgement Form to:609-750-4259 or LAPAC_CS@Integralife.comFor questions call 609-936-2495. |  | 357 | 357 units | Worldwide Distribution - USA (nationwide) and the countries of Austria, Canada, Italy, Australia | Terminated | October 29, 2013 | 701 | 1303
Z-1044-2012 | 61019 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | STA - System Control N + P |  |  | August 23, 2011 | 2011 | Diagnostica Stago, Inc. | Factor VIII quality control results are being recovered below the assigned ranges for STA System Control N+P. | Nonconforming Material/Component | Stago sent a recall letters via email dated August 23, 2011 to all affected customers . The letter identified the product, description of problem, and corrective actions to be taken. Customers were asked to immediately examine their inventory and follow the corrective actions provided. Customers were instructed to complete the response form attached to the letter. Because of incorrect manufacturer's instructions included in the initial letter another letter was emailed by the firm on August 30, 2011. For questions or technical assistance for updating the parameters please call the Diagnostica Stago 24 hour hotline at 1-800-725-0607. |  | 1259 | 1259 units | Worldwide Distribution and USA (Nationwide) and including Canada | Terminated | September 04, 2013 | 743 | 1259
Z-1043-2012 | 60982 | No | GGW | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Dade(R) Actin(R) FSL Activated PTT Reagent |  |  | December 19, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | The listed lots of Dade(R) Actin(R) FSL Reagent show an increase in Heparin sensitivity over the shelf life which is demonstrated by testing with Dade(R) Citrol(R) Heparin Control, low and high. | Unknown/Undetermined by firm | Siemens sent an Urgent Field Safety Notice letter dated December 2011 to all customers of the affected lots. The letter identified the product, description of problem, and actions to be taken by the user. DadeÂż ActinÂż FSL customers were instructed to exercise one of two options: 1) If customers would like to continue working with their Dade(R) Actin FSL reagent lot (if one of the impacted lots) they must verify on a regular basis (monthly) that their therapeutic range is still valid, e.g. by using a Heparin containing control. In the case that it is not valid, they are to adjust their existing therapeutic range or establish a new one. 2) If customers decide not to continue working with the Dade(R) Actin(R) reagent lot (if one of the impacted lots) they are to contact their Siemens representative at 1-800-242-3233, option 5, for a replacement lot. Once they have established a new normal, therapeutic, and quality control ranges, they are to discard the remainder of the old lot. Customers were advised to complete the response form included to confirm receipt of Safety Notice and fax back to (302) 631-8467. If you have any questions or concerns, please contact the Technical Solutions Center at 1-800-242-3233, option 1. |  | 24672 | 24,672 Kits | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY and the countries of Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Columbia, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Guatemala, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, Mongolia, Netherlands, Oman, Philippines, Poland, Portugal, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, United Kingdom and Vietnam. | Terminated | April 03, 2014 | 836 | 24672
Z-1046-2012 | 61102 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity CT (Computed Tomography) System, 510 (k) #K033326, Model #728326, Serial Numbers: 300003, 300004, 300005, and 300010. |  |  | December 08, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Design defect. Philips learned of this device issue following their receipt of a customer complaint reporting that a 'gap' is created in the side covers when the gantry of the device is tilted to +30 degrees (forward tilt). A hazard to the operator/user presents itself if the operator puts his or her fingers or hand inside the gap and the gantry tilt is returned in position while the fingers or | Device Design | The firm, Philips, sent an 'URGENT - Medical Device Correction Ingenuity CT***' letter dated December 08, 2011. The letter describes the affected product, the nature of the problem along witth the circumstances under which this problem can occur; the hazard involved and actions to be taken by the customer/user. The customers (operator/users) were instructed not to insert hands in the gap which will be located between the gantry and stationary support column located on either side of the CT System (illustrations are provided). The letter notifies the customers that a Philips representative will be visiting each customer site and installing a hardware upgrade through a Field Change Order. For further information or support concerning this issue contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center at (1-800-722-9377, option 5; enter site 10 or follow prompts). |  | 8 | Four (4) device system units were distributed. Two (2) in U.S. and two (2) overseas | Worldwide distribution: USA (nationwide) including states of: IN and OR; and countries of: Australia and Germany. | Terminated | March 01, 2013 | 449 | 8
Z-0833-2012 | 57842 | No | RCE |  |  |  |  | GE eXplore Locus Micro CT cabinet x-ray system model RS9-80. |  |  | January 31, 2011 | 2011 | GE Healthcare, LLC | Product fails to comply with the federal performance standard for cabinet x-ray systems (Title 21 CFR 1020.40(c)(1)(i)). Specifically, the x-ray emission limit of 0.5 milliroentgen in one hour could be exceed. Highest emission measured was 3 mR/hr. | Radiation Control for Health and Safety Act | GE Healthcare will notify purchasers of the problem, the affected product, instructions for users, and specify the maximum emission rate measured. GE Healthcare will bring defective systems into compliance free of charge. This will be done via Field Modification Instructions IFMI) 11101 which is scheduled for deployment on March 1, 2011 with an estimated completion date of July 1, 2011. |  | 21 | 21 | Worldwide Distribution | Terminated | September 06, 2012 | 584 | 21
Z-1059-2012 | 60972 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ENDURAMESH CORPECTOMY SPACER SYSTEM 013MM, 017MM, 023MM by varying heights. Dist. By: Lucero Medical, PO Box 67, Richfield, Ohio 44826 |  |  | January 11, 2012 | 2012 | Lucero Medical LLC | During an FDA inspection the firm was informed that the device marketed is not the device cleared for the 510K associated with the device. The product is sold unlabeled. | Device Design | Lucero Medical LLC made a personal visit to facilities in regards to an Urgent Medical Device Recall as of January 12, 2012. The firm instructed the customer that they were removing and quarantining all product from the facility. If you have any further questions, please contact Lucero Medical, PO Box 69, Richfield, Ohio 44826, Director of Quality Assurance/Natalie Williams/ (440) 670-9583 |  | 190 | 190 units | USA Nationwide Distribution - in the state of OH. | Terminated | April 11, 2014 | 821 | 190
Z-1060-2012 | 60855 | No | KXE | Wrist joint ulnar (hemi-wrist) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Ascension Orthopedics First Choice Partial Ulnar Head System ("DRUJ System"). Head sizes 14.5 mm, 16.0 mm, 17.5 mm, and 19.0 mm each with stem sizes 4.5 mm, 5.5 mm, or 6.5 mm. |  |  | December 03, 2011 | 2011 | Ascension Orthopedics, Inc | A recent review conducted by Ascension Orthopedics revealed that certain modifications were made to the First Choice System and should have been preceded by a submission of a filing with the FDA. | Device Design | The Firm, Ascension Orthopedics, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated December 3, 2011 via FedEx, email or fax to all customers. The letter described the product, problem, and actions to be taken. The customers were instructed to immediately cease any further use of any affected product; remove from their stockrooms until it is shipped back to Ascension Orthopedics; promptly return any affected product to: Ascension Orthopedics, 8700 Cameron Road, Austin, TX 78754, Attn: First Choice DRUJ Recall Administrator, and complete and return the reply form even if you do not have affected product in stock, to Regulatory Affairs Dept. via fax to 512-836-6933.If you have any questions, please do not hesitate to contact customer service at 1-512-836-5001. |  | 1303 | 1303 units released to customers | Worldwide distribution. | Terminated | January 14, 2013 | 408 | 1303
Z-1056-2012 | 61110 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | TruFlight Select PET/CT System equipped with software version 3.5.3.10, Model #459800111661. |  |  | December 13, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances have been noted in the TruFlight Select PET/CT System. A number of these non-conformances have been corrected to improve the overall quality and functionality of the suspect system units. | Software design | Philips sent an 'URGENT - Medical Device Correction TruFlight Select PET/CT Serial Numbers: 17002, 17003 letters. The letter was dated December 13, 2011. The letter describes the affected product, the nature of the problem along with the circumstances under which this problem can occur; the hazard involved and actions to be taken by the customer/user. In this case, the customers (operator/users) are provided with separate recommended actions. The letter goes on to notify the customers that a Philips representative will be visiting each customer site and installing a software update (3.5.3) in order to address several of the non-conformance issues through a Field Change Order. For further information or support concerning this issue contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center at (1-800-722-9377, Customers in all other countries are advised to dial their local Philips Healthcare office. |  | 2 | two (2) units were distributed | Nationwide Distribution -- OH and MD. | Terminated | February 22, 2013 | 437 | 2
Z-1055-2012 | 60513 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | ImmunoCAP Allergen Ri206, Cockroach American, Periplaneta americana |  |  | January 03, 2012 | 2012 | Phadia US Inc | The firm found this lot of ImmunoCAP Allergen Ri206, American Cockroach does not contain sufficient levels of some cockroach allergens e.g. tropomyosin. Approximately 2% of the positive samples have a risk to fall below the Limit of Quantization (o .1 kUA/L). | Release of Material/Component prior to receiving test results | ThermoFisher sent an URGENT RECALL NOTIFICATION letter dated December 22, 2011, to all affected consignees. The letter identifed the product, the problem, and the action to be taken by the conignee. Consignees were recommended to cease testing on the affected lot of product and, if so desired, substitute with ImmunoCAP Allergen i6, German Cockroach, until ImmunoCAP Allergen Ri206 was in stock. For any questions regarding this notification, contact Technical Support at 800.346.4364, option 2. |  | 10 | 780; all sent to US, with 10 assays per sleeve | Nationwide to include: MN, VA, NC, MO, SC, TX, MA, MI, FL, IL, PA, MD, UT, CO, GA, CA, | Terminated | February 04, 2013 | 398 | 10
Z-1054-2012 | 60974 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Liquid Concentrated Mouse Monoclonal Antibody CD7 (LP15) Reagent; an in vitro diagnostic primary antibody for use as part of an antibody panel for the classification of tumors of T-cell origin; Catalog No. NCL-L-CD7-580; packaged in 1 mL vials and 0.1 mL vials; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom. |  |  | January 10, 2012 | 2012 | Leica Microsystems, Inc. | The primary antibodies to CD2 (clone 11F11) and CD7 (clone LP15) have been adulterated. | Labeling mix-ups | Leica Microsystems, Inc. sent a Medical Device Field Notification letter dated January 4, 2012, via first class mail on January 10, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to discontinue use of the product and to return the product to the firm. Customers who have received the affected lot numbers of the listed products were requested to return any unused/part used reagents to the Buffalo Grove, IL, location for replacement. The accounts were asked to complete the attached Acknowledgement Form, acknowledging receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the account's Leica Representative. For further questions please call (847) 405-0123 |  | 72.1 | 60 - 1 mL and 11 - 0.1 mL vials | Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | April 18, 2012 | 99 | 339
Z-1062-2012 | 61052 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Agfa DX-D Imaging Package - (The detector power cables described in the Report of Correction and Removal are connected to portable detector panels used in the DX-D systems) |  |  | January 16, 2012 | 2012 | AGFA Corp. | The firm received a complaint that when using the DX-D Mobile DR X-Ray unit they noticed the detector cable was hot. | Device Design | AGFA Corp. sent an "URGENT FIELD SAFETY NOTICE" letter dated January 16, 2012 to affected consignees. The letter described the affected product, concerns, issues and actions to be taken. Customers have been requested to place the Varian warning label included in a highly visible and suitable position on the cable of the portable detector. Customers were instructed to complete and return the feedback form by fax to 864-421-1664, as soon as possible. For questions about this matter contact your local AGFA HealthCare Technical Support at 877-777-2432. |  | 42 | 42 units | Nationwide Distribution (USA) including the states of: FL, IL, IN, KS, MI, MN, MS, NE,NY, OH, SC and WV. | Terminated | May 10, 2012 | 115 | 42
Z-1061-2012 | 60863 | No | GWG | Neurological endoscope. | Neurology | Neurology | 510(k) | ASSISTO Arm System Table-mounted self retaining instrument holder system consisting of stainless steel tubular, articulated arms. |  |  | December 16, 2011 | 2011 | Geomed Medizin-Technik Gmbh &amp; Co. | The product may fracture during surgical procedures. | Nonconforming Material/Component | Geomed Medizin-Technik Product sent an Urgent Voluntary Medical Device Recall notification on December 9, 2012 and December 16, 2012 to affected customers. The letter identified the affected products, reason for recall and actions to be taken. Customers were instructed to (1) fill out the enclosed return form and fax back immediately at +49 (0)7461-935593, (2) return the affected products to the address provided, including the completed return form, (3) use the UPS account number 03710R provided to return the products free of charge and (4) immediately forward a copy of this recall notification to third parties to which you have distributed affected products. For questions and support, please contact Mr. Hanno Haug, VP -AMB at +49 (0)7461 9355-30 or by email: hhaug@geomed.de. |  | 1205 | 1,205 units | Worldwide Distribution - USA (nationwide) including the countries of: Australia, Belgium, China, Denmark, Finland, France, Germany, Great Britain, India, Italy, Japan, Lebanon, Libya, Lithuania, Luxemburg, Netherlands, Russia, Sweden, Spain and Turkey. | Terminated | April 15, 2014 | 851 | 1205
Z-1063-2012 | 60058 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE 2000D FULL FIELD DIGITAL MAMMOGRAPHY SYSTEM ( SENO 2000D). |  |  | October 17, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of inaccurate measurements on GE magnified images provided by Senographe 2000D Mammography systems when reviewed on IHE compliant workstations. This issue may impact patient safety. | Device Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 17, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Contact the Call Center at 800-437-1171 for questions regarding this notice. |  | 1060 | 1060 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, AR, CA,CT, DC,FL, GA, HI, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, VA, WV,and WI and countries of VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, TAIWAN, SWITZERLAND, SPAIN, SINGAPORE, SAUDI ARABIA, RUSSIA, QATAR, PORTUGAL, POLAND, NETHERLANDS, MOROCCO, MEXICO, MALAYSIA, LEBANON, KUWAIT, KOREA, JAPAN, JAM,AICA, ITALY, IRELAND, HUNGARY, HONG KONG, HONDURAS, GREECE, GERMANY, GEORGIA, FRANCE, EGYPT, DOMINICAN REPUBLIC, COLUMBIA, CHINA, CHILE, CANADA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | March 12, 2014 | 877 | 1060
Z-1064-2012 | 61071 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Imaging XS. |  |  | January 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | In September 2011 firm issued two Customer Safety Advisory Notices for syngo Imaging XS. Firm became aware of an unintended behavior when using syngo Imaging XS, version VA70A and VA70B with Client/Server Installations. Loading additional images/series from the Patient Browser using the thumbnails may cause unintended deletion of images/series. In response, firm issued a Customer Safety Advisor | Software design | Siemens Medical issued an "UPDATE INSTRUCTIONS SY055/11/S/UPDATE PACKAGE VA70B_0211" to all affected customers. The package described the product, problem, and instructions for customers to follow. |  | 128 | 128 | Nationwide Distribution-including the states of CA, CO, CT, FL, GA, IL, IN, IA, KS, MD, MN, MO, NE, NV, NJ, NY, NC, OH, PA, TN, TX, and WI. | Terminated | September 29, 2012 | 247 | 128
Z-1065-2012 | 60126 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | REVOLUTION XR/D DIGITAL RADIOGRAPHIC IMAGING SYSTEM WITH TOMOSYNTHESIS |  |  | October 26, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue associated with the hand grip that is clamped to the table top of your Revolution XRD X-ray systems that may impact patient safety. The hand grip accessory that is clamped to the table top may not be properly adjusted and can dislodge during use. Should this happen during patient loading or unloading, patient injury could occur. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to remove the hand grip accessory from use until the product is corrected and to take extra precautions when loading or unloading patients from the table while the hand grips are not in use. A GE Healthcare service representative would contact the consignee to arrange for the correction. For questions regarding this recall call 262-513-4122. |  | 772 | 772 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WY, WI and the countries of VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TAIWAN, SWITZERLAND, SPAIN, SLOVENIA, SAUDI ARABIA, REPUBLIC OF KOREA, NORWAY, NEW ZEALAND, NETHERLANDS, MALTA, KUWAIT, JORDAN, JAPAN, ITALY, ISRAEL, HONG KONG, GERMANY, FRANCE, ECUADOR, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, and AUSTRALIA. | Terminated | August 02, 2012 | 281 | 772
Z-1067-2012 | 60751 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE SIGNA EXCITE MR SURGICAL OPTION (MR450w Surgical Suite). |  |  | December 15, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the failure of the transfer board release lever associated with the OR Compatible patient transfer table of your Optima MR450w MR Surgical Suite. | Nonconforming Material/Component | GE Healthcare, sent a Urgent Medical Device Correction letter dated December 15, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed discontinue use of OR Compatible table until issue is corrected. GE Healthcare will correct all affected systems by providing a replacement table cradle incorporating design changes to the cradle release lever mechanism at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.Contact your local service representative if you have any questions concerning this notification. For further questions please call (262) 513-4122. |  | 2 | 2 | Worldwide Distribution -- USA including Wisconsin and the country of Austria. | Terminated | May 29, 2012 | 166 | 2
Z-1070-2012 | 60828 | dispenser, liquid medication | KYX | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | Children's Medical Ventures 1 ml Amber Oral Syringe. The product is a Liquid Medication Dispenser. |  |  | December 12, 2011 | 2011 | Benlan, Inc. | The tip caps may come loose from the syringe, potentially resulting in leakage of medication from the syringe. | Process control | The firm initiated their recall of the specified lots of product in November, 2011 by sending a "Customer/Distributor Medical Device Correction" letter to their consignee. The letter requested they examine their inventory, quarantine any affected product and complete a verification and response form to be returned to the firm. Questions were directed to (905) 829-5004 ext 233. |  | 143250 | 143,250 | Nationwide Distribution, USA - including only the state of Pennsylvania | Terminated | March 19, 2013 | 463 | 143250
Z-1041-2012 | 59827 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | V24-00400-29 cable, an accessory to the Newport HT50 ventilator. |  |  | August 10, 2011 | 2011 | Flight Medical Ltd. | The recall was initiated because Flight Medical Innovations Ltd. has identified a problem may occur when contact between an exposed unused wire and the grounding wire occurs. If this contact occurs the ventilator may reset and switch off. | Device Design | Newport Medical Instruments, Inc., the U.S. distributor for Flight Medical Ltd., sent a letter dated August 23, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately discontinue use of the affected product. An Acknowledgement Form was attached to the letter for customers to complete and return via fax at (714) 427-0489 or e-mail to Regulatory@NewportMedical.com. Replacement product will be provided once the form is received. Contact Newport Medical's Technical Service Group at 800-451-3111 ext. 500 for questions regarding this notice.Updated: The firm posted a press release on 02/12/12 Isreal time, notifying customers of update regarding its voluntary recall of remote alarm cable, part number V24-00400-29, used to connect a Newport HT50 mechanical ventilator to a nurse station. The press release notified: The recalled cable may negatively affect Newport HT50 ventilators manufactured prior to 2007. Affected units may shutdown unexpectedly. The FDA has classified this action as a Class I recall.On August 23, 2011, Flight Medical sent an urgent Recall Notice to affected customers through its distributor, Newport Medical Instruments, Inc. (Newport Medical), to stop using the alarm cable immediately and to contact Newport Medical in order to receive an upgraded version of the cable. In the notification letter, the customers were notified of the affected cables and that an upgraded version of the cable would be provided to them free of charge. Customers, including Newport Medical are requested to dispose of the affected cables properly.Instructions to customers: Customer inquiries related to this action should be addressed to Newports Technical Service Group at (800) 4513111, extension 500. |  | 363 | 363 total units, all in the US | Nationwide Distribution. | Terminated | May 21, 2012 | 285 | 363
Z-1072-2012 | 60881 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂż 3600 Immunodiagnostics Systems, Catalog # 6802783; COMMON/USUAL NAME: VITROSÂż 3600 --- NOTE: VITROSÂż 3600 that are e-Connected are not affected by this issue. |  |  | April 13, 2011 | 2011 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) received customer complaints regarding the loss of calibration data following the installation of a development Assay Data Disk (ADD) that was available in the on-screen mailbox on the Main Menu System Status screen of their VITROSÂż 5600 Integrated System. OCD's investigation determined that ADD, Data Release Version (DRV) 9241 or 9251 was electronically transferr | Software design | Ortho-Clinical Diagnostics (OCD) sent an Important Product Correction Notification letters dated April 13, 2011 via FedEx overnight to customers with an affected VITROSÂż 3600 or 5600 System (those not e-Connected). Foreign affiliate consignees were informed of this action via e-mail on April 13, 2011. The letters described the affected products, problem and actions to be followed. The letters states that an OCD Field Engineer will contact customers who has the affected products and perform a Modification (MOD) procedure that will remove ADD DRV 9241 or 9251 from the on-screen mailbox of their system(s). Customers were instructed to complete and return the accompanying Confirmation of Receipt Form and fax to1-888-557-3759 or 1-585-453-4110. The letter asks that distributors forward a copy of the enclosed customer communication to all customers who were shipped an affected VITROSÂż 3600 or 5600 System. For questions, please contact OCD's Customer Technical Services at 1-800-421-3311. |  | 96 | Domestic - 15 units, foreign - 81 units | Worldwide Distribution - USA (nationwide) including the countries of: Brazil, Canada, Chile, China, England, France, Germany, India, Italy, Japan, Mexico, Singapore, Spain and Venezuela. | Open, Classified |  |  | 228
Z-1071-2012 | 60971 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Discovery Elbow System Humeral Condyle Set- Hexalobular, REF 114700, Lot 797550 Alloy, Sterile, Biomet Orthopedics Warsaw IN 46581. |  |  | December 14, 2011 | 2011 | Biomet, Inc. | A Discovery Condyle Kit w/ Hexalobula was opened during an elbow joint replacement surgery, and was found to contain two male condyles. | Process change control | December 14, 2011 Urgent Medical Device Recall Notices were sent to Distributors and hospital OR managers describing the product, problem, risks associated with the recall, and actions to be taken. The notice included Instructions to locate and remove the product. Customers were asked to complete a response form and fax to 574-372-1683, prior to returning product by priority mail to the attention of Angi Dickson, Biomet , 56 East Bell Drive Warsaw, IN, USA 46580 with a copy of the response form. Receipt of the notice should also be confirmed by calling 800-348-9500, extension 3983 or 3009. Questions related to this notice should be directed to 574-371-3983 or 574-371-3009 Monday through Friday 8 AM to 5 PM. |  | 10 | 10 units | International Only Distribution -- Netherlands &amp; UK. | Terminated | December 12, 2012 | 364 | 10
Z-1066-2012 | 61050 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI 4.5 and R4.6 |  |  | January 26, 2012 | 2012 | Elekta, Inc. | If a user, using an on-line correction protocol, clicks the accept correction button, an approval dialog box appears. If cancel is selected, the table correction values return to their initial values, but the image panes continue to display the changed position. | Software design | ELEKTA sent an Important Field Safety Notice (20001507042) dated January 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.For questions regarding this recall call 770-670-2548 |  | 96 | 96 units | Nationwide Distribution including AR, AZ, CA, CT, FL, GA, ID, IL, LA MD, MA, MI, MN, MO, NJ, NY, NC, ND, OH, OK, OR, PA, RI, TN, VA, WA, WI, and Puerto Rico | Terminated | April 17, 2012 | 82 | 96
Z-1069-2012 | 61124 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | AngioDynamics SPECIAL ORDER Morpheus Smart PICC CT --- 4F Single Lumen 65cm --- Max CT Flow Rate: 4mL/sec --- Kit Contains: --- 1 each Nitinol Guidewire, 0.46mm (0.018in.) x 130cm --- Catalog No. [REF] 52110902 --- [Mfg symbol] AngioDynamics, Inc., 603 Queensbury, NY USA 12804. |  |  | February 06, 2012 | 2012 | Angiodynamics Worldwide Headquarters | The Morpheus Smart PICC label states the kit contains a Nitinol guidewire; however the kit contains a Platinum guidewire. | Labeling mix-ups | AngioDynamics sent an "URGENT-MEDICAL DEVICE RECALL" letter dated February 2, 2012 to the single customer. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs the customer to identify and segregate the affected product. A Morpheus Smart PICC Recall Reply Form was enclosed for the customer to complete and return via fax to 518-798-1360. Contact AngioDynamics Customer Service at 1-800-772-6446 for questions regarding this notice. |  | 25 | 25 units | Distributed to one customer in Wisconsin. | Terminated | February 10, 2016 | 1465 | 25
Z-1081-2012 | 60748 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | LH Lithium Heparin tubes, Greiner bio-one, North America, 4238 Capital Drive, Monroe, NC 28110. |  |  | December 14, 2011 | 2011 | Greiner Bio-One North America, Inc. | The lithium heparin tubes may have a contamination with EDTA potassium. | Material/Component Contamination | Greiner Bio-One sent a "RECALL OF LITHIUM HEPARIN TUBES WITH EDTA CONTAMINATION" letter dated December 13, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete an attached Product Disposition Site Confirmation Form and to fax the form to 800-726-0052. This form will confirm that customers have discarded/destroyed all affected product. The firm will replace the product after receipt of the form. Contact the firm at 800-515-8112 for questions regarding this recall. |  | 2082000 | 2,082,000 pieces | Nationwide Distribution. | Terminated | June 21, 2012 | 190 | 2082000
Z-1080-2012 | 61036 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CORFLO-Ultra Jejunal Tube with CORLOCK-Twoomey Y-Adaptor for use with the CORFLO-Max PEG; enteral feeding tube,10 Fr, 36" (91 cm), 3 Gram weighted bolus; individually packaged, 5 tubes per case; CORPAK MedSystems, Wheeling, IL 60090; catalog/reorder #30-7361; For delivery of enteral nutrition |  |  | January 26, 2012 | 2012 | Corpak Med Systems | The wrong size Twoomey adapter was packaged with the Jejunal tube. | Mixed-up of materials/components | Corpak MedSystems telephoned customers on January 26, 2012 and sent a follow-up Recall Notification letter dated January 27, 2012 via e-mail reiterating the telephone conversations. Customers were informed of the urgent product recall, the problem and the actions to be taken. Customers were instructed to inspect inventory, segregate any affected product and contact Corpak at 1-800-323-6305 to arrange for the return and replacement of the units found. Customers were advised to complete and return the enclosed acknowledgment form, indicating the number of affected units found in their inventory. Distributors were requested to notify their customers. For questions contact Corpak MedSytems at 1-800-403-3400. |  | 145 | 145 units | Worldwide Distribution - USA (nationwide) including the states of Georgia, Indiana, Kentucky, Maryland, Michigan, Minnesota, Tennessee, Utah and the countries of to Sweden and New Zealand. | Terminated | March 21, 2012 | 55 | 145
Z-1082-2012 | 60890 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga - previously known as Microsoft Amalga Unified Intelligence System (UIS) 2009 |  |  | January 04, 2012 | 2012 | Microsoft Corporation Health Solutions Group | Healthcare providers reported that when sharing a single browser session on Amalga Web, patient A's data can be displayed under patient B's name. | Software design | Microsoft Corporation sent an "Urgent Field Safety Notice" letter dated January 5, 2012 by email to all affected customers. The letter described the product, problem, and actions to be taken by the customers. If you have any questions about this corrective action program, please contact Microsoft Corporation at 425-882-8080. |  | 9 | 9 units | Worldwide Distribution - USA (nationwide) including the states of California, China, Kentucky, North Carolina, Wisconsin and country of Switzerland. | Terminated | December 12, 2012 | 343 | 9
Z-1074-2012 | 60959 | No | GTY | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Strep A Rapid Test Cassette (K031784), IST-502 Cassettes. |  |  | September 01, 2011 | 2011 | Alere San Diego | Alere San Diego is expanding the scope of the voluntary recall of the Strep A Rapid Test Cassette and Twist devices from the single lot that was recalled in January 12, 2011 to all lots of product that were on the market within the expiry period. | Process change control | The firm, Alere, sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 1, 2011 to its customers. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to retain the recall letter for their records as documentation that they have received the recall information. Customers were instructed to follow the directions and return the attached Verification Form. Customers were instructed to use the table in the Verification Form to document all inventory at any of their inventory locations. Customers may use multiple forms to document all sites, products, and lots. Alere will issue credit against PO#ASD811STA for any unused portion of the product based on completion and return of the attached Verification Form. Customers with any technical questions about the information contained in the notification were instructed to contact Alere by phone (877) 866-5309, fax (877) 866-9304 or email: ca2services@Alere.com. |  | 1859430 | 1,859,430 devices (228 lots) | Worldwide distribution: USA (nationwide) and countries of : Austria, Canada, Cyprus, Denmark, Finland, Germany, Greece, Hungary, Iceland, Italy, Poland, Portugal, Romania, Slovenia, Spain, Switzerland, and United Kingdom. | Terminated | June 28, 2012 | 301 | 4657030
Z-1077-2012 | 60962 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Precision Aspheric Lens (PAL) Intraocular Lens |  |  | August 02, 2011 | 2011 | Aaren Scientific, Incorporated | The recall was initiated because Aaren Scientific has confirmed that a specific group of lenses were shipped without a complete expiry date showing on the label. | Incorrect or no expiration date | Aaren Scientific sent a recall communication letter dated August 2, 2011, August 3, 2011, and August 30, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Aaren Scientific is requesting that you return the lenses and we will rework the product and return them back to you. Please let me know and I will issue an RGA number and inform you of our FedEx account number to reference.In the meantime, we are conducting an internal investigation regarding this issue. Once this has been closed out then we will notify you of our findings.For further questions please call (909) 937-1033 ext. 108. |  | 3 | 3 units | Worldwide Distribution - USA including CA and TN, and the countries of India and Lebanon. | Terminated | August 10, 2012 | 374 | 153
Z-1083-2012 | 60967 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | UltraVision Detection System; |  |  | January 11, 2012 | 2012 | Lab Vision Corporation | Firm was notified by vendor that a component in kits was not meeting internal specifications. Declination in product performance confirmed by Thermo Fisher. | Nonconforming Material/Component | Lab Vision Corporation of Thermo Fisher Scientific sent a Product Recall Notification letters dated January 11, 2012 to all affected customers. The letter identified the affected products, reason for recall and necessary actions to be taken. Customers were instructed to examine inventory, discontinue use, destroy any remaining inventory, follow normal procedure to order an alternate lot of the affected products, retain a copy of this letter for their laboratory records, if the affected products were further distributed provide a copy of this letter to them and complete the attached Customer Recall Response Form and return within 5 days. For questions regarding this information, U.S. customers should call Technical Support at 510-771-1595. Customers outside the U.S. should contact their local area Customer Service. |  | 31 | 31 kits in total | Worldwide Distribution - USA (nationwide) FL, IL, MA, TX and countries of Denmark, Turkey, Italy, Poland, Canada, Japan, Germany, Egypt and Greece. | Terminated | July 20, 2012 | 191 | 31
Z-1093-2012 | 60932 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ S8 DIAGNOSTIC ULTRASOUND SYSTEM with R1.1.0 Software. |  |  | February 07, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue where an image from a different patient, or an image from the current patient with title information from a different patient maybe noticed. | Software design | Consignees were sent on 2/7/12 a GE Healthcare "Urgent Medical Device Correction" letter dated February 6, 2012. The letter was addressed to Hospital Administrators / Risk Manager, Biomedical Engineering and Head of General Imagining Ultrasound Department. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Users were instructed not to press the PRINT key during the specified view in order to avoid seeing the issue. GE will provide a software update at not cost to customers in order to remove the malfunction. Customers are to contact their local GEHC service representative or call 1-800-321-7937 if they have questions or concerns. |  | 161 | 161 | Worldwide Distribution -- US, including the states of NJ, PA, and WI and countries of UNITED KINGDOM, THAILAND, SOUTH KOREA, SLOVENIA, SINGAPORE, ROMANIA, REPUBLIC OF MOLDOVA, POLAND, MALAYSIA, JAPAN, ITALY, INDIA, GERMANY, FRANCE, EUROPE, CZECH REPUBLIC, AUSTRIA, AUSTRALIA, and ARAB EMIRATES. | Terminated | July 06, 2012 | 150 | 161
Z-1098-2012 | 61045 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System with 4D Planning procedure option of MultiPlan Treatment Planning System, versions 3.0, 3.1, 3.5.1, 3.5.2 and 3.5.3 with Ray-Tracing dose calculation. |  |  | January 16, 2012 | 2012 | Accuray Incorporated | An anomaly was discovered during internal regression testing, where the dose information is displayed incorrectly during treatment planning during a specific workflow using the optional 4D Ray Tracing dose calculation algorithm. As a result, it is possible that the dose calculation will display a lower dose than the intended dose prescribed for treatment delivery. | Software design | Accuray sent an "URGENT DEVICE CORRECTION" letter dated January 16, 2012 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter provides a work around to prevent the issue from happening. Contact Accuray Customer Support at 1-877-668-8667for questions regarding this notice. |  | 207 | 207 | Worldwide Distribution-USA (nationwide) and the countries of Korea, Hong Kong, China, Canada, Turkey, India, Russia, Italy, France, Belgium, Poland, Germany, Netherlands, United Kingdom, Spain, Greece, Saudi Arabia, Switzerland, Ukraine, Czech Republic, and Japan. | Terminated | September 28, 2012 | 256 | 207
Z-1094-2012 | 60824 | No | MQK | Assisted reproduction labware. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | ORIGIO MIDATLANTIC DEVICES CRYOPETTE (r) Rx only Single use only CRY-PETTE-CL LOT 123456 2013-10-10 Exp. date *** MADE IN THE USA STERILE R ***Origio, Inc. 3400 Hunter's Way, Charlottesville, VA 22911 www.origio.com |  |  | November 28, 2011 | 2011 | ORIGIO HUMAGEN PIPETS INC. | Cryopettes were rupturing during the warming process following vitrification. | Device Design | ORIGIO sent a Field Safety Corrective Action letter dated November 21, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ORIGIO feels it is important to eliminate this risk completely. As soon as we are confident that this danger of rupture has been eliminated, we will make Cryopettes available again. Please use the attached UPS call tag to return any unused Cryopettes to ORIGIO Inc. Upon receipt, we will issue a credit to your account.We sincerely apologize for the problems this issue has caused in your laboratories. For further questions please call (434) 979-4000. |  | 73850 | 73850 Cryopettes | Worldwide Distribution - U.S. included states of AL, AR, DE, PR, NY, MA, NH, VT, IL, CT, NJ, PA, VA, WV, NV, NC, SC, GA, FL, OH, MD, MS, IN, MI, PR, WI, MN, SD, MT, MO, TX, CO, UT, AZ, CA, HI, OR, AK, WA and in countries of Turkey, Saudi Arabia, Kuwait, Jordon, So. Korea, Taiwan, Brazil, So. Africa, Oman, Kazakhstan, Honduras, Slovenia, Serbia, Poland, Czech Republic, Romania, Croatia, Ukraine, Netherlands, Bulgaria, Hungary, Russia, Lebanon, Singapore, Malaysia, Vietnam, China, Thailand, Nigeria, Japan, Spain, Greece, Canada, Denmark, France, United Kingdom, Italy, Germany, and Australia, | Terminated | February 04, 2013 | 434 | 73850
Z-1097-2012 | 61206 | dispenser, liquid medication | KYX | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | BD Oral Dispensing Syringe Oral Syringe Tip Caps |  |  | October 05, 2011 | 2011 | Becton Dickinson &amp; Company | There are complaint reports for loose tip caps in various types of oral syringes | Device Design | Becton Dickinson (BD) sent an Urgent: Product Recall letters dated October 5, 2011 to all affected consignees. The letter identifies the affected product, problem and actions to be taken. Customers are instructed to follow the instructions on how to handle the affected product: 1) Check their inventory to determine if they have any Oral Syringe Tip Caps from the impacted lot numbers. 2) If they have any of the impacted product return the Oral Syringe Tip Caps to BD for replacements as directed on the attached recall packing list. 3) If they use the Oral Syringe Tip Caps and require replacement, follow the instructions on the attached form. 4) If they have no inventory of the affected lots, indicate this on the website. |  | 22827500 | 22,827,500 units | (USA) Nationwide Distribution | Terminated | February 01, 2013 | 485 | 22827500
Z-1092-2012 | 61121 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE, ONCOR and PRIMUS Linear Accelerators with RT Therapist Workspace, version 4.1. Manufactured by Siemens AG, Kemnath, Germany. The indication is to deliver X-ray, photon and electron radiation for the therapeutic treatment of cancer. |  |  | January 22, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Software upgrade to installed base to the latest release of the RT Therapist version 4.1.144 which corrects multiple safety issues that could lead to patient mistreatment. | Software design | Siemens customers with active devices will receive the software update through the distribution of the "Update Instructions" with a release date of January 22, 2012. The software update will be delivered and installed by a qualified Field Service Engineer utilizing a SW kit and instruction distributed by the Update instruction. Consignee notification of the Update Instructions will be delivered by Customer Service or by certified mail. |  | 182 | 182 active devices | Worldwide Distribution - USA, including the states of WI, UT, PA, TN, OH, LA, DE, MI, NJ, FL, NE, WV, MO and the countries of Germany, Spain, Poland, Australia, Malaysia, New Zealand, India, Republic of Korea, Canada, Italy, Egypt, Philippines, France, Belgium, Saudi Arabia, South Africa, Norway, Turkey , Hungary, Russian Federation, P.R. China, Lebanon, Iran, and the United Kingdom. | Terminated | November 07, 2012 | 290 | 182
Z-1090-2012 | 61074 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian 52-leaf Standard MLC multileaf collimator; |  |  | January 16, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | When using the Standard and Millennium 52-leaf MLC multi-leaf collimators for beam shaping, under specific treatment conditions, dose may be delivered outside of the intended treatment field that exceed the expected transmission dose. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 16, 2012 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter provides recommended actions when planning and treating dynamic fields on the 52 leaf MLCs. Contact the Help Desk at 1-888-827-4265 for questions regarding this recall. |  |  |  | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Canada, China, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Hungary, Ireland, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Slovakia, Spain, Sweden, Switzerland, United Kingdom, Singapore, Bahrain, Egypt, Hong Kong, India, Indonesia, Israel, Malaysia, Morocco, Nepal, South Africa, South Korea, Taiwan, Turkey, Ukraine, United Arab Emirates, Japan, Brazil, Chile, Colombia, Mexico, Panama, Venezuela, Saudi Arabia, and Australia. | Terminated | July 26, 2012 | 192
Z-1096-2012 | 60945 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDE Patient management System, Model 6488, software version 7.2. It is indicated for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. |  |  | December 19, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific is conducting a recall involving the LATITUDE Patient Management System - Model 6488, software version 7.2 because patients are not able to perform setup. There is no risk to health associated with this issue. | Software design | Beginning on December 19, 2011, for the patients that have attempted to set-up their communicators, Boston Scientific called the patients and their associated LATITUDE clinics to let them know that they can now complete the set-up. |  | 1 | 1 | Nationwide Distribution, USA - including the states CA, MA, MI, NC, OK, PA, RI, TX, and WA. | Terminated | October 25, 2012 | 311 | 1
Z-1104-2012 | 61048 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Presource Custom Sterile Surgical and Procedure Kits containing Ethicon Ethilon Black Nylon Nonabsorbable Surgical Sutures in new Tyvek packaging; surgical trays; Cardinal Health. The kits are custom packed with all of the specific disposable class 1 and 2 devices required by the customer for specific types of surgery and medical procedures. These packs are then ethylene oxide sterilized and are ready to be used by the customer for the specific procedure. |  |  | January 27, 2012 | 2012 | Cardinal Health | The Presource kits were assembled with Ethicon Sutures in new Tyvek packaging which could no longer be placed in kits that will undergo ethylene oxide sterilization. | Packaging change control | The firm, Cardinal Health, sent an "URGENT: PRODUCT RECALL" letter dated January 26, 2012 to the affected consignees via UPS next day air. The letter identified the product, problem, and actions to be taken. Enclosed with the letter were recall labels, instructing the customer not to use the affected sutures and to remove and destroy the sutures upon opening the kit, to affix recall label to the front side of each affected Presource kit and shipping carton currently in their inventory. The customers were also requested to complete and fax back to Cardinal at 847-689-9101 or 614-652-9648 regardless of whether or not you have inventory. For component credit/replacement the consignees were instructed to contact the following: Presource Sales Operations at 800-766-0706; Government Customer Service at 800-444-1166; and Distributor Management Group at 800-635-6021. For any questions, require additional labels or desire special assistance relating to this recall contact Cardinal Health Quality Systems, 800-292-9332. |  | 18300 | 18,300 kits | Worldwide distribution: USA (Nationwide) including the states of: Alabama, Arizona, Arkansas, California, Colorado, Connecticut, Delaware, Florida, Georgia, Illinois, Indiana, Kentucky, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Jersey, New Mexico, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, West Virginia, Wisconsin and Washington; and country of Canada. | Open, Classified |  |  | 18300
Z-1099-2012 | 61180 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Retcam3 &amp; Retcam Shuttle with version 6.0 software. |  |  | February 02, 2012 | 2012 | Clarity Medical Systems Inc | Still images extracted from video recordings may be mislabeled as to right and left eye. | Software design | Clarity Medical Systems, Inc. sent a Safety Alert letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed the following :1. Do not use the video mode until the software has been updated.2. Install the software update using the enclosed CD. The instructions for installation are also enclosed. This software update will correct the anomaly.3. If you have extracted still images from video recordings, review all potentiallyaffected patient files to determine if any right eye images are mislabeled as left eye images and/or left eye images are mislabeled as right eye images. If any images are mislabeled, make the appropriate corrections in the annotation section of the affected file. Consider whether there is any diagnostic impact or need to re-irnage the patient.We apologize for the inconvenience. If you have questions or concerns, please contact Clarity at 800-215-6005. |  | 33 | 33 RetCam Shuttles and 20 RetCam 3 - total 53. | Worldwide Distribution -- USA (nationwide) including the states of TN, MO, TX, FL, NJ and Puerto Rico and countries of Italy, Portugal, Czech Republic, Thailand, Chile, Oman, Hong Kong, Russia, Philippines, Sweden, Finland, Ecuador, Australia and Canada. | Terminated | June 07, 2012 | 126 | 33
Z-1103-2012 | 61082 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Sunrise Medical Zippie TS wheelchair |  |  | January 13, 2012 | 2012 | Sunrise Medical (US) LLC | One Zippie TS Wheelchair was made per customer order was shipped with transit brackets for use a seat in a motor vehicle, however the chair was not configured with a back tested seat for motor vehicle use. | Employee error | Sunrise Medical contacted customer by telephone on January 17, 2012 to arrange for removal of transit brackets. The Sunrise Sales rep then visited the customers on January 18, 2012 and removed the bracket from the chair making it appropriate for intended use.If you need any additional information or have any questions, call 559-348-2572 or email: laurie.roberts@sunmed.com. |  | 1 | 1 chair | Nationwide distribution: OK only. | Terminated | February 27, 2012 | 45 | 1
Z-1101-2012 | 60981 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Oral-Eze (R) Oral Fluid Collection System, manufactured by Capitol Vial, Inc., Auburn, AL 36832 |  |  | October 07, 2011 | 2011 | Capitol Vial Inc | Low buffer fill volumes on Oral-Eze, an oral fluid collection system. | Device Design | Thermo-Fisher Scientific/Capital Vial sent a Recall Notification letter dated October 7, 2011 to the sole consignee. The letter identified the affected product, problem, result impact and necessary actions to be taken. The consignee was instructed to discontinue use of the affected product and follow the necessary actions provided. If you have any questions regarding this information, call Michael Hect, Manufacturing Manager through Customer Service at (800) 772-8871. |  | 135000 | 135,000 units | (USA) Nationwide Distribution - in the state of KS | Terminated | August 24, 2012 | 322 | 135000
Z-1102-2012 | 60976 | No | DRT | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | V Series Monitor |  |  | October 14, 2011 | 2011 | Mindray DS USA, Inc. d.b.a. Mindray North America | A system database corruption may occur that causes the V Series Monitor to reset. | Process control | The firm, Mindray, sent an "URGENT PRODUCT CORRECTIVE ACTION" letter dated October 14, 2011 on October 17, 2011 via certified mail, return receipt requested, signature required to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to contact their Mindray Service Representative at 1-800-288-2121, to arrange for software upgrade of the V Series Monitor(s) in their facility.If you have any questions, contact Manager, Regulatory Affairs and Clinical Affairs at 201-995-8407 or email: d.arpino@mindray.com. |  | 695 | 542 units US ; 153 units International | Worldwide distribution: USA (nationwide) including states of: AL, CA, CO, IN, MO, NE, PR, and SD; and countries of: Australia, Chile, Columbia, United Kingdom, and Venezula. | Terminated | June 11, 2013 | 606 | 695
Z-1100-2012 | 61235 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Breeze Self-Adhesive Resin Cement, part number: N97F. The intended use is for cementation of porcelain, resin, and metal-based inlays, onlays, crowns, bridges, and posts. |  |  | December 02, 2010 | 2010 | Kerr Corporation | The recall was initiated because Pentron Clinical has confirmed that the Breeze Self-Adhesive Resin Cement exhibits faster than expected gel and set times. If the work time is exceeded when using this material, it could possibly result in inadequate bond strength. | Nonconforming Material/Component | An "Urgent Medical Device Recall" letter was sent on 12/17/2010 via USPS 1st class mail to all their customers. The letter provides the customers with an explanation of the problem identified and the action to be taken. Customers were instructed to complete and fax back the enclosed recall return form. Customers with questions were instructed to call (800) 551-0283. |  | 622 | 622 units total (12 units in US) | Worldwide Distribution - Nationwide (USA) and the countries of Vietnam, Jordan,Turkey, and Italy. | Terminated | February 24, 2012 | 449 | 622
Z-1106-2012 | 61049 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Premier Care Recliner Model 5570. |  |  | October 17, 2011 | 2011 | Winco Mfg., LLC | On 10/17/2011 Winco Mfg. LLC, Ocala, FL initiated a recall of The Premier LifeCare Recliner and Premier Care Recliner Models 5400 and 5570 distributed between January 4, 2011 and October 12, 2011. The Position Lock Pin(s) could potentially fail if exposed to excessive force when the chair is in the Locked Recline position. This failure will result in an inoperable Position Lock Bar feature. | Device Design | Winco Mfg. sent a "POSITION LOCK PIN RECALL NOTIFICATION" letter dated October 24 , 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and return an enclosed response form and provide product disposition. Winco will ship retro-fit parts to install on the affected product. Instructions will explain step by step, how to implement the fix. Contact Winco at 800-237-3377 for questions regarding this recall. |  | 79 | 79 chairs | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, FL, GA, IL, KS, MA, MI, MO, MS, NC, NM, NY, OK, PA, SC, SD, TN, TX, VA, WI, and the country of Canada. | Terminated | May 04, 2012 | 200 | 133
Z-1111-2012 | 60790 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Discovery NM 630 |  |  | January 11, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue associated with the collimator cart and lead cover of your Discovery NM/CT 670 and/or Discovery NM 630 system that may impact operator safety. 1.The cart locking pin release and cart disengagement could result in the collimator falling down and could lead to operator or service personnel injury.2.Another issue is that the | Nonconforming Material/Component | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated January 13, 2012 to its Consignees/Customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The customers were instructed to continue to use product, ensure that the collimator exchange monitor the collimator cart are properly locked in place; pay attention to any unexpected cart motion during the exchange procedure; cease the process immediately if the cart has moved from the docking spot, dock cart in place again, and complete the exchange; and if a 3rd party service engineer is servicing rather than the firm, GE, advise according to the instructions in letter.A GE Healthcare Service representative will perform the required hardware update on each affected system.If you have any questions or concerns regarding this notification, contact your local GE Healthcare Service Representative. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other countries: 262-896-2890. |  | 10 | 10 | Worldwide distribution: USA (nationwide) including states of: CA, FL, GA, IL, NJ, NY, NC, TX, and WI; and countries of: AUSTRALIA, BELGIUM, CANADA, CHINA, DENMARK, FRANCE, GERMANY, INDIA, ISRAEL, ITALY, KOREA, JAPAN, NETHERLANDS, NORWAY, POLAND, RUSSIAN FEDERATION, SPAIN, SWITZERLAND, UNITED KINGDOM. | Terminated | December 06, 2012 | 330 | 89
Z-1109-2012 | 61006 | No | LKA | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | BBL[TM] Lysine Iron Agar, catalog #211363,500 g bottle 2014-03-31. Lysine Iron Agar was developed for use in the differentiation of enteric organisms based on their ability to decarboxylate or dominate lysine and to form hydrogen sulfide. Becton, Dickinson and Company, Sparks, MD 21152 USA 38800 Le Pont de Claix, France. |  |  | January 11, 2012 | 2012 | Becton Dickinson &amp; Co. | Lysine Iron Agar may not produce the correct biochemical reactions with Quality Control organisms over the shelf life of the product.. | Process design | BD Diagnostics Systems sent an "URGENT PRODUCT RECALL" letter dated January 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard any remaining packages. A form was attached for customers to complete and return via fax to: Attention: Regulatory Compliance at 410-316-4258. Contact the BD Technical Services Department at 1-800-638-8663 for questions regarding this recall. |  | 33 | 33 bottles | Worldwide Distribution-USA (nationwide) including the states of AL, AR, CA, CO, CT, DC, FL, GA, IL, IN, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OR, PA, SD, TX, VA, and WI, and the countries of Belgium, Canada, Chile, China, Columbia, Costa Rica, Mexico, El Salvador, Ecuador, Japan, Manila, Singapore, Peru, Taiwan, and Venezuela. | Terminated | March 30, 2012 | 79 | 297
Z-1057-2012 | 60992 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Powerheart G3 Automated External Defibrillators (Non-Wearable) and include the following models: Powerheart 9300A, 9300E, 9300P, 9390A, 9390E; CardioVive 92532, 92533; Nihon-Kohden 9200G, 9231; and GE Responder 2019198, 2023440 AEDs. |  |  | January 13, 2012 | 2012 | Cardiac Science Corporation | The affected Automated External Defibrillators (AEDs) contains a circuit board manufactured with a component that may fail unexpectedly due to a supplier manufacturing defect. If the component were to fail during a rescue attempt, the AED may not deliver defibrillation therapy. | Process control | Cardiac Science sent an "URGENT-Voluntary Medical Device Recall" letter dated January 13, 2012 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to immediately return the device to the manufacturing location for repair. The affected devices will receive a hardware correction, and the same serial number device returned to the consignee in most cases.Customers in the US can call 1-888-402-2484 and customers outside the US can call 1-425-402-2482 or email at aed210@cardiacscience.com to arrange delivery of the shipping materials.Customers can also visit www.cardiacscience.com/aed210 for more information about AEDs affected by this recall. |  | 11019 | 11,019 units (6142 units in the US and 4877 units international) | Worldwide distribution: USA (nationwide) and Puerto Rico; and to the following countries: Australia, Belgium, Brazil, Canada, China, Colombia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Malaysia, Maldives, Malta, Netherlands, New Zealand, Norway, Panama, Peru, Poland, Portugal, Qatar, Romania, Russia, Saint Lucia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates, Ukraine, and United Kingdom. | Terminated | September 27, 2013 | 623 | 11019
Z-1117-2012 | 60831 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | VITROSÂż Chemistry Products PHBR Slides, REF 822 1384 |  |  | June 23, 2011 | 2011 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) received customer complaints of intermittent imprecision, outliers, or shifts in control fluid values when using various lots of VITROSÂż Chemistry Products PHBR Slides (Phenobarbital) from Coating 0053. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated June 23, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue using and discard all remaining lots of the affected product. A Confirmation of Receipt Form was attached so customers may complete and return to the firm. The firm recommends that customers consult with their Laboratory Medical Director and requesting physician to resolve any concerns they may have regarding previously reported results. Contact Customer Technical Services at 1-800-421-3311 for questions regarding this notification. |  | 2977 | Domestic - 2345 sale units, Foreign - 632 sales units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 2977
Z-1114-2012 | 61123 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics BENEPHIT CV Infusion System Targeted Renal Therapy Infusion System --- Catalog No. [REF] 60040 --- Manufactured in USA --- [Mfg symbol] AngioDyanamics, 603 Queensbury Avenue, Queensbury, NY USA 12804. |  |  | February 06, 2012 | 2012 | Angiodynamics Worldwide Headquarters | The Benephit CV Infusion System contains a Sheath Y Connector that has the potential to develop cracks. | Nonconforming Material/Component | Field Safety Notices (FSNs), entitled Urgent - Medical Device Recall, were sent to consignees on 02/06/2012 via USPS certified mail, return receipt requested. The letters identified the affected devices, the issue, hazard involved, and actions to be taken by the customer and AngioDynamics. The FSN directs consignees to identify and segregate the recalled lots, returning any affected product to AngioDynamics, Inc., Queensbury, NY 12804. The Recall Reply Form should be completed and faxed to 518-798-1360. --- If customers have distributed the affected product/lot, they are instructed to inform their customers of this recall action immediately, and then return the recalled units. --- If customers have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please contact the local sales representative or call Julie Blair, Customer Service Manager, at ANGIODYNAMICS Customer Service at 1-800-772-6446. Customers may also e-mail questions to customerservice@angiodynamics.com |  | 4 | 4 units | Nationwide Distribution -- Florida, Iowa, Maryland, New York, and Texas. | Terminated | February 10, 2016 | 1465 | 15
Z-1116-2012 | 61031 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Perkin Elmer Genetic Screening Processor (GSP) Product code: 2021-0010. It is intended for in vitro quantitative / qualitative determination of analytes in body fluids. |  |  | January 09, 2012 | 2012 | PerkinElmer Life and Analytical Sciences, Wallac, OY | GSP instruments with defective washer modules may cause a false negative screening result being reported for T4, TSH and IRT. For 17-OHP, a false positive result may be generated if the failure occurs for a normal sample. | Nonconforming Material/Component | Perkin Elmer issued an "Urgent Medical Device Recall" letter dated 1/25/12 to all customers and informed them about the possible issue and requested permission for service personnel to check their instrument. The customers were instructed that the instrument could not be used and the washer modulerequired replacement. Customers were told to expect to be contacted by the firm's local representative for further details. |  | 10 | 10 units | Worldwide Distribution - USA, including the states of Florida, Ohio, Oregon, Texas and the countries of Brazil, Belgium, France, Germany, and the UK. | Terminated | March 06, 2013 | 422 | 10
Z-1115-2012 | 61239 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Brand Name: Fusio Liquid Dentin, part number N21SB, N21SC. |  |  | August 19, 2010 | 2010 | Kerr Corporation | A recall was initiated because Pentron Clinical has confirmed that the Fusio Liquid Dentin is not as flowable as ideally expected. | Nonconforming Material/Component | Pentron Clinical sent a Medical Device Recall letter dated September 8, 2010 (via USPS 1st class mail) to all affected customers. The letter identified the affected product, the reason for the recall and the actions to be taken. Customers were instructed to examine their inventory for the affected product, return any remaining in stock and complete and fax back the enclosed recall return form. Customers were advised to contact Pentron Clinical Customer Service at (800) 551-0283, (option 1) directly to handle the arrangements of a quick return and replacement. |  | 382 | 382 units total (207 units in US) | Worldwide Distribution - USA (nationwide) in the countries of Austria, Bulgaria, France, Great Britain, Netherlands, Japan, Lebanon, Malaysia and South Africa. | Terminated | February 28, 2012 | 558 | 382
Z-1126-2012 | 61184 | No | FIE | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medisystems (a Nxstage Company) Hemodialysis Fistula Needle Set with MasterGuard Ant-Stick Needle Protector 15 GA x 1" |  |  | January 31, 2012 | 2012 | Medisystems Corp | Mislabeled outer carton: Outside shipping cartons is stamped 15Ga x 1", product contained within properly labeled as 17Ga x 1" needle | Employee error | Medisystems Corporation issued notice of recall to distributors on January 31, 2012. Distributors to notify affected end users by letter. Each letter was accompanied by a response form and instructions on how to return the response form. End users were given the option of returning involved product for credit or replacement. Contact Medisystems Quality Assurance Manager at (978) 332-5973 for information. |  | 20750 | 20,750 units (83 cases) | USA (nationwide) including the states of AZ, CA, CO, OR, NV, NY, TN, and WA. | Terminated | September 12, 2012 | 225 | 20750
Z-1015-2012 | 60995 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | B. Braun Infusomat Space Pump - US Version Infusion pump. Model number 8713050U. |  |  | July 12, 2011 | 2011 | B. Braun Medical, Inc. | If the IV set anti free flow clip is incorrectly inserted into the pump and the pump door forced closed, the clip catch on the inside of the pump door may break. If the clip catch is broken and the door opened, free flow protection is ensured. If IV set is then removed from pump without closing the roller clamp, the clip may not re-engage the tubing, leaving an open fluid path with free flow pote | Device Design | B Braun sent a letter on August 16, 2011, to their customers, that detailed the proposed G03 implementation in response to reported customer complaints. The software upgrade with the door jam was agreed upon and the eight hospitals were upgraded in November 2011. |  | 5769 | 5769 | Product was distributed to hospitals in CA and MO. | Terminated | November 03, 2014 | 1210 | 5769
Z-1127-2012 | 61141 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza with software version VA20B_HF04 |  |  | January 20, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction when using syngo.plaza VA20B_HF04 with USB dongle and archive license. In syngo.plaza VA20B_HF04, there is a scenario where data in Short-Term Storage (STS) could reaming unprotected and not archived. this can happen when an archive server with a USB dongle-based license is enabled with auto archiving. If the disk fill level is reached which triggers auto deleti | Software design | Siemens sent a Customer Safety Advisory Notice letter dated January 20, 2012 to all affected consignees. The letter identified the affected product, problem, steps to take to avoid this issue, and actions to be taken. Customers were instructed to contact their local Siemens Service to plan for their software update and ensure that this safety advisory is placed in the system's instruction for use. |  | 7 | 7 | Nationwide Distribution (USA) - CA, FL, GA, MO, NY and WA. | Terminated | September 29, 2012 | 253 | 7
Z-1129-2012 | 59285 | No | REK |  |  |  |  | DirectFRAP (Direct Flourescence Recovery After Photobleaching) Imaging System is an extension unit for Axio Observer that is used for microscopic analyses of dynamic processes within living cells. The DirectFRAP module is a component used in conjunction with Laser Scanning Microscopes, Non-Medical Devices for Research Use. The DirectFRAP Imaging System is a Class 3B laser system. |  |  | April 28, 2011 | 2011 | Zeiss, Carl Inc | The manufacturer confirmed that the opto-semiconductor sensors can be triggered by exposure to incandescent lighting. The safety interlock is in place to prevent unintended exposure to Class 3B laser radiation. | Nonconforming Material/Component | Carl Zeiss MicroImaging LLC, through trained personnel, first visited each user site, provided additional user instructions, re-labeled the components to provide adequate instruction and warning and further placed a label over the access to the locking screw which secures the component to the microscope. Carl Zeiss Inc. issued "IMPORTANT LASER SAFETY NOTICE: URGENT: FIELD CORRECTIVE ACTION" notices to each user by express courier in late April 2011, Carl Zeiss personnel will confirm that the relabeling has been performed and updated instructions for use have been provided. For questions, please contact the company.The firm will implement 2 actions. Service Representatives will visit each user site, provide additional user instructions, relabel the components to provide adequate instruction and warning, and add a label over the access to the locking screw which secures the component to the microscope. Secondly, when new replacement sensors are available, Carl Zeiss field service technicians will perform the component replacements. Modified componentsare anticipated to be available within 6 months and replacement completed by April 2012. |  | 11 | 11 units | Worldwide Distribution -- USA (nationwide) | Terminated | December 11, 2012 | 593 | 111
Z-1125-2012 | 61043 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Philips Intellispace Portal, Release 4.0.0. |  |  | January 27, 2012 | 2012 | Philips Medical Systems Nederlands | There maybe incorrect values in the diffusion/perfusion maps. | Software design | Philips issued URGENT - Medical Device Correction Notifications, dated January 20, 2012. The letter identified the affected product, all Intellispace Portal with software release 4.0.0 Installed, and gave a description of the problem. Users were provided with the actions to be taken prior to the software upgrade( to version 4.0.1). Users are to follow the instructions provided in order to generate maps with the correct values and to make sure that the setting is correct before using the generated data. Philips will contact customers in order to install the update free of charge. Customers should contact their local Philips representative at 866-767-2822 |  | 25 | 25 units | Worldwide Distribution -- US, including states of IN and TX, and countries of Australia, Austria, Czech, Denmark, France, Germany, Indonesia, Netherlands, Poland, Spain, and Switzerland. | Open, Classified |  |  | 25
Z-1118-2012 | 61187 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Liebel-Flarsheim Optistar MR 60 mL Syringe w/lateral flow needle, Mallinckrodt, Cincinnati, OH 45237. Made in Mexico. Devices packaged in Plastic and Tyvec pouch, with 50 pouches per case. |  |  | February 10, 2012 | 2012 | Mallinckrodt Inc | The machine used to package the product was missing a tool, which could prevent the packaging from sealing properly, causing the sterility of the product to be compromised. | Packaging process control | Covidien sent an Urgent Device Recall letter dated February 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product immediately. Customers were to use the attached business response form to report the quantity of each lot that is currently in their stock. If customers did not have any of these lots in stock, they should indicate by checking the appropriate box on the business report form. When completed, customers should email the completed form to covidien2600@stericycle.com or fax to 888-257-7910. For any questions customers should call 866-761-9571. |  | 665 | 665 Cases | Worldwide Distribution - USA including Distributed worldwide: USA including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, Puerto Rico and the countries of Canada, Australia, Belgium, Brazil, Chile, Ecuador, Guatemala, Mexico, Panama, Singapore, and Uruguay. | Terminated | December 24, 2013 | 683 | 665
Z-1133-2012 | 61185 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | BardPort M.R.I. Implanted Port with Attachable 6 Fr. ChronoFlex Open-Ended Single-Lumen Venous Catheter and Peel-Apart Introducer Kit. Subcutaneous implanted intravascular infusion port and catheter is used to provide repeated access to the vascular system for the delivery of medications, intravenous fluids, parenteral nutrition solutions, and blood products. They are also indicated for the withdrawal of blood samples. |  |  | February 16, 2012 | 2012 | Bard Access Systems | The end flap label on the kit incorrectly describes the product as "Bard Port" MRI DUAL Implanted Port instead of a "Bard Port MRI Implanted Port" as correctly indicated on the main label of the kit. | Labeling Change Control | A "Labeling Discrepancy Notification" letter was sent to US consignees via FED EX on February 16, 2012 and on February 20, 2012 for the foreign consignee explaining the labeling discrepancy and providing a visual example of the incorrect label. The customers were given the option of continuing to use the product with the understanding the end flap label was incorrect or they could return the affected products for replacement (Call 1-800-290-1689 to arrange a replacement). |  | 545 | 545 units (Lot RESF0856 - 150 units; Lot RESH0067 - 295 units; Lot RESH1150 - 100 units) | Worldwide Distribution -- USA, including the states of CA, CT, FL, KS, IL, MI, MN, OH, OK, SD, TN, TX, VA, WA and the country of Belgium. | Terminated | March 30, 2012 | 43 | 545
Z-1132-2012 | 60888 | No | INX | Air-fluidized bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | Clinitron Rite Hite Bed headboard and foot side rails. |  |  | February 03, 2012 | 2012 | Hill-Rom, Inc. | The firm identified in October 2011 that the Clinitron Rite Hite Bed and Clinitron At Home Bed headboards failed to meet the safety standard force requirements. The firm identified in October 2011 that the Clinitron Rite Hite Bed and Clinitron At Home Bed outer rail, which is extruded Aluminum tubing, is intended to be one piece joined together by a weld at the lower center. The parts receiv | Nonconforming Material/Component | Hill-Rom, Inc. sent a Modification Notice Urgent Medical Device Correction Notification dated February 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The modification notice instructed all customers to find and replace all headboards and food side rails that do not meet design intent. All technicians will inspect all affected product in the service centers and all at home and rite hite beds for an "x" or "L" stamp. The technicians will inspect the beds that are still in use during scheduled maintenance or before each rental. All will be completed before 01-JUL-12.For support call 1-800-445-3720 |  | 103 | Foot Siderail = 70; Headboards = 103 | Worldwide Distribution -- USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, WV, and WY. and the country of Canada. | Terminated | December 05, 2012 | 306 | 103
Z-1130-2012 | 60980 | No | DIS | Barbiturate test system. | Toxicology | Toxicology | 510(k) | Product Brand Names: TriageÂż Drugs of Abuse Panel 8 Test Kit (K924476), TriageÂż Drugs of Abuse |  |  | January 04, 2012 | 2012 | Alere San Diego | Alere San Diego received a rise in customer complaints regarding false positive Barbiturates (BAR) results on the Triage Drugs of Abuse Panel Plus TCA. | Nonconforming Material/Component | Alere sent a "PRODUCT NOTIFICATION" letter dated January 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Verification Form was enclosed for customers to complete and return to the firm. Contact Technical Services at 1-877-308-8287. |  | 693054 | 693054 kits | Worldwide Distribution-USA (nationwide) and the countries of Australia, Brazil, Canada, EU, Guatemala, Japan, Lebanon, and Saudi Arabia. | Terminated | July 05, 2013 | 548 | 693054
Z-1095-2012 | 61070 | No | KTI | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Inspira AIR Balloon Dilation System, |  |  | January 30, 2012 | 2012 | Acclarent, Inc. | Acclarent received reports of difficulty deflating the balloon during the procedure, which could potentially result in airway obstruction. | Component design/selection | Acclarent sent an "Urgent Voluntary Product Recall" letter dated January 31, 2012 by US mail, return receipt requested to all affected customers. Additionally, customers known to have scheduled procedures using the affected product were notified by phone to advise them of the recall. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately, discontinue use and return any remaining in stock per the instructions provided. The letter provides additional information and photographs illustrating how to identify the recalled product. Customers were advised to complete the enclosed Business Reply Card regardless of whether they have or do not have product subject to this recall in their possession and mail to Stericycle. Customers were instructed to distribute this information to all staff within their department who uses the affected product. For further questions call 1-866-781-1173 or contact your Acclarent Sales Representative. |  | 633 | 633 units | Nationwide Distribution (USA) only sixe 18 x 40Worldwide distribution other sizes.Distirbution to France, Canada, Saudi Arabia | Terminated | August 29, 2013 | 577 | 633
Z-1137-2012 | 61073 | No | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | Perfector |  |  | January 06, 2012 | 2012 | Ion Genius, Inc. | Wires do not meet performance standards under 21 CFR 898 and have unprotected electrode configurations. Additionally, devices were sold/distributed without 510(k) or PMA | Nonconforming Material/Component | Ion Genius sent a Medical Device Product Recall letters to all affected customers on January 6, 2012. The letter identified the affected products, problem and actions to be taken. The letter instructed customers to check their inventory for the affected products and return immediately to Ion Genius for replacements. Customers returning affected products should call the firm so that a shipping label can be provided.Customers were instructed to be completed the Response Form and returned to Ion Genius by fax at 808-395-0787. For questions and concerns contact Xanya Sofra-Weiss at 808-222-2664 or email x@arasysperfector.com. |  | 72 | 72 cables in total, amounts of other devices not reported | Worldwide Distribution - (USA) Nationwide Distribution | Terminated | May 21, 2015 | 1231 | 72
Z-1135-2012 | 61107 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific, LATITUDEÂż Patient Management System Communicator, Models 6476. |  |  | January 19, 2012 | 2012 | Boston Scientific CRM Corp | One (1) Model 6476 LATITUDE Communicator is unable to interrogate the patient's pulse generator. Boston Scientific has determined that this is due to a patient receiving an incorrect communicator. Therefore the communicator did not find and was unable to interrogate the assigned patients PG. Additionally when the patient attempted to set-up this incorrect communicator, the system reflected | Packaging process control | On January 19, 2012 Boston Scientific telephoned the affected patient and their clinic. The patient was notified that the communicator they received will not be able to activate successfully and that their communicator and sensors would be replaced. The clinician was informed of the patient's communicator was sent incorrectly to the other patient who attempted to activate it, the monitored status of the website was incorrect and that a new communicator would be shipped to the patient that will work with their device. |  | 1 | 1 | Nationwide distribution: UT only. | Terminated | July 05, 2012 | 168 | 1
Z-1136-2012 | 61221 | No | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Brand name: Orthos CM Bracket, Orthodontic Metal Bracket, |  |  | February 13, 2009 | 2009 | Ormco Corporation | The recall was initiated because Ormco has confirmed that manufacturing error resulted in the Orthos CM Brackets being incorrectly manufactured with a torque of -9Âż. The packaging was labeled with +9Âż torque; however the brackets contained inside the packaging are actually Orthos CM Brackets with a torque of -9Âż. | Process control | Ormco Corporation sent an Urgent Field Safety Notice dated February 13, 2009, to all affected customers. The recall communication for consignees in the United States, Canada, Colombia, HongKong, Kuwait, United Arab Emirates, Mexico, Russia, South Africa and Croatia was sent on February 13, 2009 via USPS First Class Mail. The recall communication for consignees in Japan was sent on February 17, 2009. European consignees were sent recall notification on February 23, 2009. The Australian consignees and New Zealand consignees were the sentrecall communication on February 25, 2009. Customers were instructed to contact Ormco Customer Care at (800) 854-1741 to receive an RMA number. The affected product should be returned to Ormco Corporation at the following address: Ormco Corporation, 1332 South Lone Hill Avenue, Glendora, CA 91740.Customers were instructed to label their return product "RECALLED PRODUCT Attention: Customer Returns". Customers were instructed to complete the enclosed Return Form and return it by fax to (909) 962-5605.For questions regarding this recall call 909-962-5600. |  | 304 | 304 units | Worldwide Distribution - USA (nationwide) and worldwide to Canada, Colombia, Hong Kong, Kuwait, United Arab Emirates, Mexico, Russia, South Africa, Croatia, Japan, France, Bulgaria, Germany, Estonia, United Kingdom, Hungary, Ireland, Italy, Netherlands, Australia, and New Zealand. | Terminated | March 02, 2012 | 1113 | 304
Z-1134-2012 | 61127 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Patient Monitors: |  |  | February 09, 2012 | 2012 | Philips Healthcare Inc. | Central station monitors may not alarm (either visual or audible) when an alarm is announced at the affected patient monitor | Software change control | Philips Healthcare sent a Urgent Medical Device Correction letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips has recently received reports that, under certain circumstances, alarms announced at the patient monitor are not announced (either visual or audible) at the central station. This issue affects the following IntelliVue patient monitor models with SW rev. H.xx.xxMP20 (M8001A), MP30 (M8002A), MP40 (M8003A), MP50 (M8004A), MP60 (M8005A), MP70 (M8007A), MP80 (M8008A), MP90 (M8010A), Intelligent Display D80 (M8016A), MX600 (865242), MX700 (865241) and MX800 (865240)If this issue occurs, the primary alarm function at the bedside monitor is not affected.Philips is conducting this voluntary correction to upgrade software on affected devices. Please refer to the following pages, which provide instructions for actions to be taken. Follow the Action to be taken by Customer/User section of the instructions. This issue has been reported to the appropriate regulatory agencies.Ensuring that you have the highest quality medical devices, accessories and supporting documentation is our top priority. Your satisfaction with Philips products is very important to us.Review this information with all staff members who interface with the central station and need to be aware of the contents of this communication.The Product Number and Serial Number is contained on the devices product label, located on the front of the device. The SW revision can be accessed via the Revision Screen at the bedside monitor. charge.A Philips Healthcare representative will contact customers with affected devices contact Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this correction to arrange an upgrade of the Intellivue software. |  | 83534 | 83,534 units | Worldwide Distribution -- USA (nationwide) and the countries of ALBANIA, ANGOLA, ANTILLES, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL, BYELORUSSIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GABON, GERMANY, GHANA, GREECE, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KUWAIT, LEBANON, LIBYAN, ARAB, JAMAHIRI, LITHUANIA, LUXEMBOURG,MACAU, MALAYSIA, MOROCCO, MAURITIUS, MEXICO, MOROCCO, NAMIBIA, NEPAL, NETHERLANDS, ANTILLES, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PALESTINE, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB, EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIET NAM, YEMEN, ARAB REP and ZAMBIA | Completed |  |  | 83534
Z-1143-2012 | 61100 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OSS AVL Poly Tibial Bushing Set |  |  | February 07, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action after complaints were received where OOS AVL Bushing Sets were opened during OSS knee replacement procedures, and were found to contain two tabbed bushings. | Employee error | BIOMET Orthopedics sent an URGENT MEDICAL DEVICE recall notice dated February 7, 2012, to all affected consignees. The recall notice instructed consignees to immediate locate and remove the devices from circulation and to fax back the response form to 574-372-1683. If product was further distributed, consignees must notify hospital personnel of this action via the enclosed notice. Consignees were charged with the location and return of the affected product. For questions regarding this recall call 574-372-3983 or 574-371-3009. |  | 27 | 27 units distributed US and foreign | Worldwide Distribution - USA including IN. KS, NJ, CA, and OHInternationally: The Netherlands | Terminated | February 06, 2013 | 365 | 27
Z-1147-2012 | 61125 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Titanium SURFIX ALPHA SCREWS, a component of the Newdeal Hallu Lock Plate System. Surfix Alpha Variable angle locking system exists in different diameters and lengths. The product can be sold as a replenishment part individually packaged sterile, or as one (1) unit which is kitted within an instrumentation tray at which point, the product is non-sterile until the tray in which it is kitted is sterilized a the designated facility. The product is indicated for use in fixation of fractures, osteotomies or arthrodesis of the first metatarso-phalangeal joint. |  |  | December 21, 2011 | 2011 | Integra LifeSciences Corp. | Manufacturing error. The part number which corresponds to the length dimension of : 14 mm. was incorrectly etched on to screws with the actual measurement of: 12 mm. | Error in labeling | Integra contacted the Sales Representatives and All USERS on December 21, 2011, with a 'SUBJECT: MEDICAL DEVICE RECALL' letter. The letter was sent via FedEx and/or E-mail. The letter describes the product, problem, and the actions to be taken by the customers and the firm. The customers are instructed to immediately examine their inventories for the presence of the recalled product, and complete and return the attached, MEDICAL DEVICE RECALL ACKNOWLEDGMENT AND RETURN FORM via E-mail or FAX #1-609-0750-7999. Any questions concerning the recall document are to be addressed to the customers' regional managers or to the E-mail address and/or telephone number provided on the recall notification. |  | 105 | 105 units | Nationwide Distribution, USA - including the states of OH, FL, WI, CT, OK, TN, OR, ID, CA, PA, MD, VA, FL, MS, NJ, LA, IA, GA, AL, MA, NY, MN, TN, IL, AR, NE, AZ, MI, WA, OR, NC, KY, MO, UT, NV, IN, SC, CO and the countries of the United Kingdom, France, Switzerland, and Austria. | Terminated | February 19, 2013 | 426 | 105
Z-1146-2012 | 59959 | No | ELW | Impression material. | Dental | Dental | 510(k) | Correct VPS Vinyl Polysilocane Impression Material, Product Number Q01B. The intended use of this device is as a dental impression material. |  |  | November 04, 2010 | 2010 | Sybron Dental Specialties | Pentron Clinical is voluntarily recalling the following lot of mislabeled Correct VPS Medium Body Tubes. It was discovered that the individual tubes inside the Correct VPS Medium Body Tube boxes are missing the lot/expiration date sticker. | Packaging process control | Pentron Clinical sent "Urgent Medical Device Recall" letters dated November 17, 2010 to consignees explaining the above mentioned public reason for recalls and asking them to determine if they have any of the affected products in their in stock. Product that is returned will be replaced at no charge. They also asked that they complete and fax back an enclosed recall return form.Questions can be directed to Customer Service at (800) 551-0283. |  | 176 | 176 units | Worldwide Distribution, USA including the states of FL, NY, WA, CA, IL, MA, NC, NJ, AZ, FL, KY, CT, PA, IN, NJ, PA, TN, NV, TX and the countries of Canada and Pakistan. | Terminated | March 14, 2012 | 496 | 176
Z-1142-2012 | 61265 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Brand name: Bond-1 Primer/Adhesive, a resin tooth bonding |  |  | May 27, 2010 | 2010 | Kerr Corporation | The recall was initiated because Pentron Clinical has confirmed that various lots of Bond-1 Primer/Adhesive have been found to have gelled and therefore becomes unusable. | Nonconforming Material/Component | Pentron Clinical sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 7, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were instructed to complete and return an enclosed Recall Return Form via fax at 1-877-677-8844. Contact Pentron Clinical Customer Service at (800) 551-0283, (option 1) directly to handle the arrangements of a quick return and replacement. |  | 30276 | 30,276 units total | Worldwide Distribution-USA (nationwide) and the countries of Germany, Canada, Ecuador, Egypt, Hong Kong, India, Israel, Japan, Kenya, Lebanon, Morocco, Saudi Arabia, Singapore, El Salvador, Thailand, Turkey, Taiwan, UAE, Vietnam, Armenia, Austria, Belgium, Bulgaria, Byelorussia, Cyprus, Germany, Spain, Russia, Finland, England, Greece, Hungary, Ireland, Italy, Latvia, Netherlands, Poland, Romania, Sweden, and Serbia. | Terminated | March 02, 2012 | 645 | 30276
Z-1144-2012 | 61041 | No | OYK | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TibiaxysÂż Ankle Arthrodesis Plate Compression Guide, Model #159635, part # 159 635, Batch #: EAHZ, EAJ1, ELGV, ENQE &amp; ENQH. |  |  | November 17, 2011 | 2011 | Integra LifeSciences Corp. | During 6/30/2010 to 8/17/2011 the firm received five (5) incident complaints of breakage of the threaded part of the compression guide which potentially occurred when: 1) a compression is applied between talus and tibia. (Step 3 of the product description) or when 2) the surgeon screws the compression guide into the plate which most likely was attributed to improper surgical techniques used durin | Nonconforming Material/Component | Integra sent a "Field Safety Notice" letter dated November 17, 2011 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review the Surgical Techinique Manual and follow the recommended guiidelines to avoid breakage. ensure that this information is distributed to their surgical teams, and fill out the enclosed "Field Safety Notice Acknowledgement and Return Form" and return to bonnie.kuzbyt@integralife.com or fax to 1-609-750-7999. For questions call 609-936-6932 or email at sean.luland@integralife.com. |  | 168 | 168 units | Worldwide Distribution - USA (nationwide) including the states of; AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KY, MA, MD, MI, MO, MN, NC, NJ, NM, NV, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA &amp; WI. Product was also shipped to the following countries: Austria, Belgium, Canada, Czech Republic, France, Germany, Ireland, Israel, Italy, Lebanon, Saudi Arabia, South Africa, Spain, Sweden, Switzerland &amp; United Kingdom. | Terminated | February 19, 2013 | 460 | 168
Z-1145-2012 | 59740 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 42-422561 SIGNATURE QTY, 1 TKA GDE/MDL (Prosthesis, knee) SET 04-05 DGHUMPHR-NOV52LKA-J_K-00432022 SLIDEX DISTAL-MP/PREMIER TIBIAL CUSTOM-MADE DEVICE |  |  | August 04, 2011 | 2011 | Biomet, Inc. | Biomet has initiated this action following investigation that identified two (2) specific lots of part number 42-422561 were incorrectly labeled. One (1) unit (lot 025640) was received back from distributor's office as a product complaint. | Labeling mix-ups | The firm, BIOMET Orthopedics, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated January 23, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:-Immediately locate and remove the above identified Signature " Guide and Model Set.-Carefully follow the instructions on the enclosed FAX Back Response Form .- Fax a copy of the Response Form to 574-372-1683 prior to return of product.- Use priority carrier for your shipment.- If you have further distributed this product, you MUST notify hospital personnel of this action, via the enclosed recall letter. This letter MUST be given to hospital personnel responsible for receiving recall notices. Confirm receipt of this notice by calling, 800-348-9500, extension 3983. Questions related to this notice should be directed to 574-372-3983, Monday through Friday, 8AM to 5PM. |  | 2 | 2 units | Nationwide distribution:Virginia and Ohio only. | Terminated | September 06, 2012 | 399 | 2
Z-1148-2012 | 60657 | No | DYJ | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | AdvanSync is an orthodontic device intended for the orthodontic treatment of Class II malocclusions. The AdvanSync part numbers and lot numbers affected by this recall are as follows: |  |  | October 02, 2009 | 2009 | Ormco Corporation | Manufacturing Error: There is a potential for a braze joint failure in the upper telescopic rod assembly. Although an incident could occur in which the eyelet loses connection from the rod, there is no risk of swallowing or aspirating the appliance because it will still be attached to the patient's crown. | Process control | Ormco sent a Urgent Medical Device Recall letter dated October 9, 2009,( via USPS 1st class mail) to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are instructed to return the affected lot numbers of the 5-Pack Left and Right Telescopic Rod Assemblies (600-4040 and 600-4041) and the Left and Right Telescopic Rod Assemblies contained within the AdvanSync Kit (600-4000). The affected parts are identified by the items contained inside the wells titled, "Left Rods" and "Right Rods." The remaining items within the kit are not affected by this recall notification; therefore, there is no need to return the entire kit.Customers were instructed that any product which they have in stock will be replaced at no charge or credited to their account. Customers were instructed to contact Ormco Customer Care at (800) 854-1741 (ext. 1) directly to handle the arrangements of a quick return and replacement. The affected product should be returned to Ormco Corporation at the following address: Ormco Corporation, 1332 South Lone Hill Ave., Glendora, CA 91740. Customers were instructed to label their returned product "RECALLED PRODUCT Attention: Customer Returns". In addition, customers were asked to complete the enclosed Return Form and return it by fax at (909) 962-5605. |  | 3635 | 3,635 Units (237 kits, 315-pack left assemblies and 305-pack right assemblies) | Worldwide- USA (nationwide) and the countries of Europe, Australia, Canada and New Zealand. | Terminated | March 05, 2012 | 885 | 3635
Z-1153-2012 | 61208 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5/C5 Control panel for 2 mast roller pumps 85, Part No.: 28-95-85, Sorin Group Deutschland GMBH, Lindberghstrasse 25, 80939 Munchen, Germany, Sorin Group USA, Inc., 14401 W. 65th Way, Arvada, CO 80004, USA |  |  | February 16, 2012 | 2012 | Sorin Group USA, Inc. | Touch screen may be unresponsive, inhibiting user input. | Mixed-up of materials/components | All affected US customers and distributors were notified of the Sorin S5 HLM Touch Screen Field Correction by certified mail on 02/15/12. The letter informed users of the potential for unresponsive S5 HLM touch screens and notified them that they would be contacted by the Sorin Service Group to arrange an appointment to replace affected touch screens on site.Customers and distributors were also provided with a Response Form to confirm they had received, read and understood the Field Correction Notice. They were instructed to fax the completed form to (303) 467-6502.For questions regarding this recall call 303-467-6306. |  | 2 | 2 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Belgium, Bulgaria, China, Denmark, Germany, Ecuador, Egypt, Finland, France, Georgia, Greece, United Kingdom, Hong Kong, Indonesia, Iraq, Iran, Israel, Italy, Japan, Jordan, Canada, British Columbia, South Korea, Kuwait, La Reunion, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Morocco, Mexico, Monaco, New Zealand, Netherlands, Norway, Oman, Austria, Phillippines, Poland, Portugal, Puerto Rico, Qatar, Serbia, Russia, Saudi Arabia, Sweden, Switzerland, Singapore, Spain, South Africa, Syria, Thailand, Czechoslavakia, Tunisia, Turkey, Hungary, Vietnam, and the United Arab Ermirates | Terminated | November 09, 2012 | 267 | 187
Z-1159-2012 | 61251 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Ysio, AXIOM Luminos dRF, AXIOM Luminos Agile/TF, Uroskop Omnia |  |  | February 15, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Firm has become aware of unintended behavior when using Ysio, AXIOM dRF, AXIOM Luminos Agile/TF and Uroskop Omnia with RAD Fluoro Uro function using software version VC10A. Due to an error in the image pipeline, images may display dark and not suitable for diagnostic. Due to this software error customers have to repeat x-ray exposure to receive diagnostic images. | Software design | Siemens sent Update Instruction XP057/11/S by mail on February 15, 2012, which releases software update package VC10D to remedy the described issues.For questions regarding this recall call 610-219-4834. |  | 58 | 58 | Nationwide Distribution including AK, CA, CO, FL, GA, IL, IN, IA, KY, MD, MA, MO, NE, NY, NC, OH, OR, PA, SC, TX, and UT. | Terminated | January 14, 2014 | 699 | 58
Z-1160-2012 | 60285 | No | FEM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Olympus Mobile Workstations (Endoscopy Carts), Models WM-DP1, WM-NP1 and WM-WP1 |  |  | October 28, 2011 | 2011 | Olympus America Inc. | Replacement wheels may break off from the base of the cart | Equipment maintenance | OLYMPUS sent an Urgent - Device Correction letter dated October 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. OLYMPUS required customers indicate on the attached questionnaire the contact information for their facility so that they can schedule a convenient time for an Olympus Field Service Representative to visit. Service personnel would inspect, and replace, as necessary any user installed casters on the WM-P1 cart. For questions regarding this recall call (483) 896-5688. |  | 5 | 5 | Nationwide Distribution including AL, CA, NY and WV. | Terminated | January 14, 2014 | 809 | 5
Z-1176-2012 | 61055 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Cell Marque CD7 (MRQ-12); |  |  | January 23, 2012 | 2012 | Cell Marque Corporation | Letter received from supplier that material received in October 2009 was labeled as CD7, but was actually CD 2. | Mixed-up of materials/components | Cell Marque sent a Field Notification letter dated January 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete and return the attached form indicating they have received the notification. Customers were also asked to destroy any existing inventory of the affected product. If the product was shipped to their customers, customers should confirm that they have contacted those customers. Customers were asked to complete, sign and date and return the attached Acknowledgement Form by faxing it to 1-916-746-8989 or by scanning it and emailing it to bdevos@cellmarque.com.For questions regarding this recall call 1-800-665-7284 or 1-916-746-8900. . |  | 110 | 110 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Greece, Turkey, Italy, The Netherlands, Spain, Portugal, France, Israel, S. Korea, Pakistan, Morocco, Russia and Brazil. | Terminated | October 19, 2012 | 270 | 110
Z-1174-2012 | 60937 | No | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | Synchron Phenobarbital (PHE) Reagent, P/N 469785, |  |  | August 11, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed reports of variability in results with the lots of Phenobarbital (PHE) reagent and calibrator. | Nonconforming Material/Component | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 26, 2009 to all customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were instructed to share the recall information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. Customers were asked to complete and return a response form. Contact Beckman Coulter Customer Support Center (Hotline) at (800) 854-3633 for questions regarding this recall. |  | 31559 | 31,559 units total (30,617 units in the US) | Worldwide Distribution--USA (nationwide) and the country of Canada. | Terminated | December 05, 2013 | 1577 | 31559
Z-1165-2012 | 61088 | No | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Depuy Mitek Lupine Loop Plus Anchor w/Orthocord |  |  | January 27, 2012 | 2012 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Lupine BR Anchor w/ Orthocord may be assembled with an incorrect suture configuration and result in a loose repair | Process control | DePuy Mitek sent an Urgent Voluntary Product Recall letter dated January 27, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to immediately check all inventories to determine if they have any affected product. Customers were asked to refer to the instructions to report their inventory status, return and obtain replacement or credit for the product. For questions regarding this recall call 1-877-379-4871. |  | 32408 | 32408 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, AE, Argentina, Australia, Austria,Belgium,Brazil, Bahrain, Chile, China, Columbia,Cyprus, Check Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, France, Finland, Great Britain, Guadeloupe, Hong Kong, Hungary,Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Luxembourg,Latvia, Martinique, Montenegro, Mexico, Maylasia, Netherlands,Norway, New Zealand, Oman, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Sweden, Singapore, Slovakia, Thailand, Tunisia, Turkey,Uruguay, Venezuela, and South Africa | Terminated | August 05, 2015 | 1286 | 190853
Z-1175-2012 | 61196 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam and True Beam STx versions 1.0 through 1.5, Model number H19; Varian Medical Systems, Palo, Alto, CA 94304. TrueBeam Radiotherapy Delivery System is intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation treatment is indicated. |  |  | January 30, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the respiratory gating software of TrueBeam. When importing breath-hold gating protocols, the gating thresholds can be reset to default values, rather than retaining the thresholds established during planning. | Device Design | Urgent Medical Device Correction/Urgent Field Safety Notices were sent on January 30, 2012 by certified mail to US and Canadian consignees.***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***Varian Sales, Marketing and Service Organizations also were informed via letter the same day. The Field Safety Notice with customer distribution list was e-mailed to Varian International Customer Support Managers for translation (if necessary) to be sent to affected sites with return proof of notification to be sent to Varian. The notice informed users of the recommended actions. Questions were directed to the Varian Help Desk 1-888-827-4265 (US and Canada). |  | 382 | 182 units ***12-12-12 AMENDED TO: 200 units*** | Nationwide Distribution (USA) and the countries of Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Italy, Netherlands, New Zealand, Norway, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Thailand, Turkey and the United Kingdom. | Terminated | April 16, 2014 | 807 | 382
Z-1172-2012 | 61224 | No | EKS | Dental hand instrument. | Dental | Dental | 510(K) Exempt | K3 Engine File G-Pack, |  |  | August 01, 2008 | 2008 | Ormco Corporation | A recall was initiated because SybronEndo has confirmed that the K3 Engine File G-Pack has label packaging and file mis-marking error. | Process control | A recall communication was initiated on 07/31/2008 with SybronEndo forwarding an Urgent Medical Device Recall letter (via USPS 1st class mail) to all their customers who purchased the K3 Engine File G-Pack (Part No. 830-9925). The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to contact SybronEndo Customer Care at (800) 346-3636 directly to handle the arrangements of a quick return and replacement. Customers were instructed to complete the Return Form and return any affected product in their inventory. |  | 199 | 199 units | Worldwide Distribution -- US, including the territory of St. Thomas, and countries of Australia, Canada, New Zealand, China, Singapore, India, and Chile. | Terminated | March 08, 2012 | 1315 | 199
Z-1177-2012 | 61262 | No | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | Pelton &amp; Crane Spirit 3000 dental operative unit with the TLC SAFEdrive option, 1727 Fuehauf Dr. Charlotte, NC 28273 |  |  | February 20, 2012 | 2012 | Pelton &amp; Crane Company | Modification of the Pelton &amp; Crane Spirit 3000 unit disabled the TLC SAFEdrive feature rendering it ineffective in detecting whether the handpiece is potentially damaged or heating up. | Employee error | Pelton &amp; Crane sent an Urgent Medical Device Recall letter dated February 20, 2012 to all affected customers. The letter identifies the affected product, problem and actions to be taken. The letter states that a Pelton &amp; Crane representative will be contacting the customers to schedule a service appointment with the end user of the device so that the affected TLC SAFEdrive circuit board can be replaced and returned to Pelton &amp; Crane. Customers were instructed to complete the attached Recall Acknowledgement/Return Form and fax it to Pelton &amp; Crane, attention Frank Ray, at (704) 587-7250. The letter also requested customers to mail the original Recall Acknowledgement/Return Form in the enclosed self addressed stamped envelope. |  | 2 | 2 units | (USA) Nationwide Distribution | Terminated | March 28, 2012 | 37 | 2
Z-1161-2012 | 61198 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, versions 7.3, 8.0, 8.1, 8.2, 8.6, 8.9 and 10.0; |  |  | January 30, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian has been identified an anomaly with the Eclipse Treatment Planning System where the calculated dose distribution is not updated or invalidated when a recalculation or reload is performed while the compensator or fluence editor is open. | Device Design | Varian sent an URGENT MEDICAL DEVICE CORRECTION - URGENT FIELD SAFETY NOTICE letter dated January 30, 2012 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter identified the affected product, description of problem, provide recommended user actions, and special instructions for customers outside the USA and Canada. The letter stated that this document will be posted on the customer support website: http://www.MyVarian.com. Customers were instructed to complete the attached Proof of Notification or Receipt Verification Card and return to Varian Medical Systems. For further clarification and questions contact your Varian Customer Support District or Regional Manager. |  | 25604 | 12,077 units ***AMENDED TO 13,527 units*** | Worldwide Distribution - (USA) Nationwide. | Terminated | August 15, 2014 | 928 | 25604
Z-1186-2012 | 60321 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Intact PTH (iPTH) Assay -assay (lot 138) |  |  | October 03, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc | Poor sensitivity and imprecision on ADVIA CentaurÂż systems; Serum samples low bias; EDTA plasma samples low bias | Process control | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated October 2008 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructed customer to immediately discontinue the use of ADVIA Centaur Intact PTH (iPTH) assay kits that contain intact PTH (iPTH) reagent lot 138. It also recommended that for serum samples analyzed on the ADVIA Centaur CP system, the content of the Urgent Field Safety Notice be discussed with their laboratory director to determine the need to review previous test results, conduct patient follow up, and repeat testing for iPTH results that were less than the upper limit of the reference range, and had a companion calcium result that was elevated or at the upper limit of normal. Additionally, a Completion Notification form was attached for customers to complete and return. Customers should contact their local technical support provider for questions regarding this recall. |  | 500 | 100 Test =5768; 500 Test =1778 | Worldwide Distribution-USA (nationwide) and the countries of Canada, Argentina, Chile, Columbia, Egypt, Finland, France, Hong Kong, India, Israel, Japan, Malaysia, Mexico, Netherlands, New Zealand, Poland, South Korea, Saudi Arabia, Singapore, South Africa, Spain, Taiwan, Thailand, Turkey, Uruguay, and Venezuela. | Terminated | March 08, 2012 | 1252 | 500
Z-1179-2012 | 61114 | No | LXB | Hearing Aid. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | SoundBite Hearing System; Microphone Tube Size 0. For use in patients with moderately severe, severe or profound sensorineural hearing loss in one ear and normal hearing in the other or patients with conductive hearing loss where the pure tone average bond-conduction hearing threshold is better than or equal to 25 dB HL. Sonitus Medical, San Mateo, CA. |  |  | February 02, 2012 | 2012 | Sonitus Medical Inc. | Sonitas received reports of stress cracking on the SoundBite Microphone Tube. | Nonconforming Material/Component | Letters were sent to all direct SoundBite customers on February 9, 2012 via certified mail. Customers were asked to remove the old microphone tubes and replace with the new enclosed microphone tubes. Return mail instructions were provided for the old microphone tubes. Customer questions were directed to 650-251-4185. |  | 676 | 676 microphones, all sizes. | Class III Recall - Nationwide Distribution, including the states of CA, AZ, PA, MI, UT, NC, NY, WI, TX, FL and DC. | Terminated | March 23, 2012 | 50 | 676
Z-1184-2012 | 61237 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | Siemens mMLC (ModuLeaf) safety bar and trolley label. |  |  | February 07, 2012 | 2012 | Siemens Medical Solutions USA, Inc | This correction is being performed to update the safety knob of the Siemens brand mMLC (ModuLeaf) to conform to tie IEC 60601-1 safety standard. | Nonconforming Material/Component | The firm, SIEMENS, issued a "Customer Information" letter to its customers. The letter described the product, problem and actions to be taken. Siemens developed a new steel knob that complies with the latest standards; Siemens checked all ModuLeaf mMLC trolleys and if necessary refitted with the safety label and also arranged for an inspection and modification of the customers system- a Release of the safety update UI TH 026/11/S to distribute and hardware update commenced on February 7, 2007 by certified mail.If you have any questions, contact the Regulatory Specialist II at 925-602-8175. |  | 46 | 46 active devices | Worldwide distribution: USA (nationwide) including states of: CO, IA, KY, NJ, OH and TX; and countries of: Algeria, Austria, Belgium, Brazil, China, Egypt, France, Germany, Hungary, India, Italy, Japan, Kuwait, Malaysia, Mexico, P.R., Russian Fed., Saudi Arabia, South Africa, Spain, Turkey, and United Kingdom. | Terminated | August 14, 2012 | 189 | 46
Z-1185-2012 | 61238 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON G5 ventilators with software version V2.0x. |  |  | February 06, 2012 | 2012 | Hamilton Medical Inc | Hamilton-G5 ventilators with software version V2.0x may experience an unintentional change of the displayed ventilation mode under certain conditions. | Software design | Hamilton Medical sent a Medical Device Safety Alert and Correction Action letter dated February 17, 2011 to all affected customers via a traceable method. The letter identified the affected product, problem and immediate actions required by operators, distributor and manufacturer. The letter states that a Hamilton Medical Account manager will contact customers to schedule a convenient time for the Technical Support Department to install the software update. Questions or concerns may be directed to Technical Support Manager at 800-426-6331, Ext 125. |  | 456 | 456 devices in US. | (USA) Nationwide Distribution including the states of: CA, CO, FL, GA, MA, MI, MO, NC, NV, NY, PA, TN, TX, VA and WV. | Terminated | October 23, 2013 | 625 | 456
Z-1194-2012 | 61182 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | SYNCHRON Rheumatoid Factor (RF) Reagent, Part Number: 475136 |  |  | June 20, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman has confirmed that the all in date lots of Synchron Rheumatoid Factor (RF) Reagent might generate false positive results or calibration failures. | Nonconforming Material/Component | Beckman Coulter sent an Urgent Field Safety Notice/Product Correction letter dated June 20, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that Beckman Coulter is developing a new RF assay. In the meantime due to current issues related to shelf life and low end sensitivity, Beckman has decided to take the product off the market effective immediately.Beckman will keep the customers updated on their progress and advise when the new assay becomes available.Customers were instructed to complete and return the enclosed Response Form with in 10 days. Customers with questions regarding this notice can contact their local Beckman Sales representative or contact Customer Support Center at www.BeckmanCoulter.com/CustomerSupport/Support or call (800) 854-3633 in the US and Canada. |  | 22081 | 22,081 units total (12,462 units in US) | Worldwide Distribution - USA (nationwide) and worldwide to foreign countries:Australia, Belgium, China, Costa Rica, Czech Republic, Finland, France,French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy,Japan, Kuwait, Lebanon, Lithuania, Macao, Mexico, Netherlands, NewZealand, Poland, Russian Federation, Serbia, Slovakia, South Africa, Spain,Switzerland, Taiwan, Turkey, United Arab Emirates, and the United Kingdom. | Terminated | September 02, 2014 | 1170 | 22081
Z-1193-2012 | 61092 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | PointCare NOW CD4NOW Gold Reagent Pack, or Gold Pack, |  |  | November 02, 2011 | 2011 | Pointcare Technologies, Inc. | Stability not assured at labeled temperature 2-42ÂżC | Storage | The firm, PointCare, issued a "Customer Service Advisory (CSA-00016)" dated November 2, 2011 to its customers to inform them of the new storage requirements of 2-30ÂżC as opposed to the previously claimed 2-42ÂżC. The letter describes the product, problem and action to be taken. The customers were instructed to correct the labeling on any Gold Reagent Packs and Gold Reagent and ensure that the Gold Reagent Packs and Gold Reagent they currently have in stock is stored within the 2-30ÂżC storage range. Future lots will be labeled with the updated storage temperatures.A member of our customer support team will provide additional information relative to this advisory.Should you have any questions or require further assistance, you may email us at customersupport@pointcare.net or call our offices directly at 1-508-281-6925. |  | 463 | 463 packs | International only: including countries of : Abuja, Haiti, Indonesia, Natal, Nigeria, Senegal, South Africa, Uganda, and United Kingdom. | Terminated | April 26, 2012 | 176 | 463
Z-1191-2012 | 61021 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga UIS 2009 is Microsoft Amalga Unified Intelligence System (UIS), Calculator/data processing module, for clinical use. |  |  | January 11, 2012 | 2012 | Microsoft Corporation Health Solutions Group | The problem occurs when a user filters the data in Amalga datagrid to exclude specific values. If the field specified in the filter criteria contains NULL values, the filtering will exclude those records as well. | Software design | Microsoft Corporation Health Solutions Group sent a "Urgent Field Safety Notice" email dated January 19, 2012, to all affected customers. Notification of customers in China is still pending due to the Chinese New Year Holiday. Customers were instructed to :1. If the user notices varying row counts or inconsistent values in a given column and they are using the exclude filter, they may have to look for alternative ways to filter the data that do not use the above mentioned filter conditions.2. To avoid the indicated issue in your implementation there are two steps to be taken:a. Ensure that all parsers are inserting empty strings () rather than NULLwhen ingesting data into the system andb. Identify all Baseview definitions (i.e. the sql view that a baseview isaccessing) that include an outer join and update those definitions toinclude the ISNULL replacement function [e.g. ISNULL( column_name, )]on any columns which may return NULL due to a non-matching row in theouter join tabAs of January 23, 2012 The firm's members of Services Team have spoken by phone with representative of every customer.Customers should call 425-882-8080 for any questions about this recall. |  | 28 | 28 units | Worldwide Distribution -- USA (nationwide) including the states of AR, CA, KY, HI, FL, MA, MD, MN, NC, NJ, NY, WA, WI and countries China, Switzerland, and United Kingdom. | Terminated | December 12, 2012 | 336 | 28
Z-1192-2012 | 61279 | No | GWG | Neurological endoscope. | Neurology | Neurology | 510(k) | Visionsense Stereoscopic Vision System, Brand name VSII, Model/Part Numbers 172-0030 and 172-0040 --- Classification Name: Neurological Endoscope --- NOTE: Model Numbers 172-0032, 172-0044, 172-0070, 172-0072, 172-0080 were also included in the firm's recall action but these never entered US commerce. These were distributed directly from Israel to the foreign consignees. --- The VSII is intended for viewing internal surgical sites during general surgical procedures and for use in visualization of ventricles and structures within the brain during neurological surgical procedures. These devices do not carry an expiration date. However, they are automatically taken out of service and replaced after 40 uses. |  |  | March 21, 2011 | 2011 | Visionsense, Ltd. | A complaint was received and a unit returned that exhibited the metal and glass shaft window becoming dislodged during scope cleaning. Additionally, during in-process visual inspection of the metal and glass shaft window sub-assembly from a specific batch that was received from the supplier it was noted that the window at the tip of the scope was not seated properly. It was determined that the wi | Nonconforming Material/Component | Visionsense Ltd. sent URGENT DEVICE RECALL and URGENT DEVICE RECALL/FIELD SAFETY NOTICE letter dated March 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer and included a return "fax" form in the recall communication package. The letters were sent via UPS express service. All direct consignees were contacted, and follow up phone calls conducted if needed. The recall required the return of all affected units for inspection and/or repair. Visionsense provided a return authorization number and packing instructions to the consignees. |  | 21 | Domestic - 17 units; Foreign - 4 units | Worldwide Distribution - USA including Arizona, Georgia, Pennsylvania, California, and the country of Turkey | Terminated | March 08, 2012 | 353 | 21
Z-1190-2012 | 61173 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Right Long Tactile Probe, Part #962011S, Medtronic Navigation, Louisville, CO 80027. |  |  | February 06, 2012 | 2012 | Medtronic Navigation, Inc. | When subjected to bending, twisting, and malleting, the probe tips break off when used in the sacrum for which they are not indicated for use. These probes are approved for use in the lumbar vertebra. | Device Design | An Urgent - Product Safety Notice customer letter was sent via FED EX on 2/6/12 and 2/6/12 to direct consignees in the US identifying the affected product and the issue with it. The letter states that Medtronic Navigation is removing the product from clinical use and will provide a sales credit to customers. In the letters to Medtronic field personnel, it states that local sales and service personnel will be identified to remove the affected product from customer sites and return them for scrapping. Customers are to contact a sales representative if they have any questions regarding the letter. |  | 5 | 5 units | Worldwide Distribution -- US and Republic of South Korea. | Terminated | March 26, 2012 | 49 | 11
Z-1187-2012 | 61131 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integra InterFix CT Adapter (CTA) |  |  | January 23, 2012 | 2012 | Integra LifeSciences Corp | Material used in the superior region of the InterFix CT Adapter (CTA) has a different density than that of the Treatment Table Couch Adapter (TTA). Resulting dosimetry can be affected if posterior fields are used. Integra estimates that a target dose decrease of 1.2% would result from a typical 360 degree TomoTherapy arc. | Device Design | Integra sent a Product Technical Bulletin dated January 23, 2012, to all affected consignees via FedEx, email or telephone conference calls. The revised Operator Manual was provided with a Technical Bulletin to all consignees that have been shipped this system. Consignees were asked to distribute the information in the bulletin to members of the Radiotherapy Planning Team within their institution. Customers were also asked to sign and return the enclosed Product Technical Bulletin Acknowledgement Form. For questions regarding this recall call 888-772-7378. |  | 33 | 33 units | Worldwide Distribution - USA including DC, IN, FL, HI, LA, NY, MD, MI, OH,OR, WI, WY,Foreign: Great Britain, Italy, and France | Terminated | October 01, 2012 | 252 | 33
Z-1195-2012 | 61145 | No | CIN | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCESS Ostase |  |  | June 22, 2011 | 2011 | Beckman Coulter Inc. | The Access Ostase reagent lot 021281 shows instability resulting in falsely high patient results. | Process design | A recall communication was initiated on 6/15/2011 with Beckman forwarding an "Urgent Product Correction" (PCA) letter to all their customers who purchased the Access Immunoassay Systems Ostase Reagent Kits. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to discontinue use of the Access Ostase reagent lot identified on the PCA letter and discard all remaining inventory. As stated in the Instructions for Use for Access Ostase, results should be interpreted in light of the total clinical presentation of the patient.All results generated with reagent lot number 021281 where patient results are higher than expected should be evaluated. The affected lot has been removed from Beckman's inventory, and alternate lots will be provided for all current and future replenishment orders or requests.Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions can contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support or call(800) 854-3633 in the US and Canada. |  | 318 | 318 units total (44 units in US) | Worldwide Distribution -- US, including states of FL, IN, VA, CA, and KY and countries of China, Japan, and Panama. | Terminated | March 09, 2012 | 261 | 318
Z-1197-2012 | 61267 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva FD, Software R5.1.x |  |  | February 24, 2012 | 2012 | Philips Medical Systems Nederlands | Displayed skin dose rate is calculated using an incorrect point of reference, resulting in too low displayed values. | Software design | Philips Healthcare issued Urgent Field Safety Notices dated February 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. No action is necessary by the customer/user. A mandatory Field Safety Correction (FCO) Action will be issued to solve the problem. The FCO will consists of a Software update. For further support customers were instructed to contact their Phillips representative.For questions regarding this recall call 978-687-1501. |  | 20 | 45 US; 20 Canada;142 International=207 units | Worldwide Distribution - USA (nationwide) and the following countries:Canada, AustraliaAustriaBelgiumBahrainCzech RepublicDenmarkEstoniaFranceFinlandGermanyHong KongIsraelItalyLatviaLuxemburgMalaysiaMexicoNetherlandsNew ZealandNorwayRussian FedSAUDI ARABIASerbiaSlovakiaSolveniaSouth AfricaSpainSwitzerlandSwedenSyriaThailandTurkeyUnited Kingdom | Terminated | March 26, 2013 | 396 | 59
Z-1198-2012 | 60892 | No | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Flat Proximal Pressurizer. |  |  | January 13, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | The firm identified two complaints have been reported indicating that the product inside the package did not match with the label information. The product was labeled as Flat Proximal Pressurizer P/N 0206-556-000 while the product inside the package was Femoral Pressurizer P/N 0206-566-000. | Mixed-up of materials/components | Stryker Instruments sent an "Urgent Medical Device Recall notification" dated January 13, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to : Immediately check all stock areas or operating room storage and quarantine any affected product found. Mark product as RECALLED PRODUCT . Please indicate on the BRF the quantity of affected Pressurizers you are returning and fill in the form completely. Sign the BRF (even if you do not have any affected product).Note: Your signature on the BRF indicates that you received and understand this notification and are returning all affected product from your facility.Fax the Business Reply Form to Stryker Instruments Regulatory Department,866-521-2762.Upon receipt of the BRF, Stryker will contact you to coordinate the return of allaffected product that you have on hand. Send back all affected product using the pre-paid shipper provided to you by Stryker.Upon receipt of the product, a replacement will be issued to your account |  | 10 | 10 boxes (60 units) | Worldwide Distribution -- Nationwide, including the states of AL and OK and the country of Canada. | Terminated | July 24, 2013 | 558 | 10
Z-1199-2012 | 61179 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Nucletron Adapter and Transfer Guide Tube, part number AL 13005000, a component used with the VariSource" IX HDR Afterloader. |  |  | January 05, 2012 | 2012 | Varian Medical Systems, Inc. | Components of medical device used in brachytherapy may be defective and cause a shifted dose distribution to patients. | Device Design | Varian Medical Systems sent an Urgent Medical Device Correction - Urgent Field Safety Notice letter dated January 5, 2012 to all affected consignees. The letter identified the affected products, description of problem and recommended user actions to be taken. Consignees were requested to cease use of the affected products and contact Varian BrachyTherapy Helpdesk immediately to arrange return of their affected parts. Consignees were requested to return the enclosed proof of notification form. Consignees were instructed to advise the appropriate personnel working in their radiotherapy department of the content of this letter. For future reference, this document will be posted to the Varian Customer Support website: http://www.MyVarian.com. For further clarification contact your local Varian District or Regional Manager. |  | 122 | 122 total units | Worldwide Distribution - USA (nationwide) Including the states of CA, FL, IN, KY, MI, MN, MO, OH, TX, VA and the countries of Canada, Japan and Sweden. | Terminated | November 21, 2012 | 321 | 122
Z-1201-2012 | 61160 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TDX SP Power Wheelchair, Model #'s: TDXSP-CG, TDXSP-CG-GT, TDXSP-MCG, TDXSP-MCG-GT, Distributed By: |  |  | February 10, 2012 | 2012 | Invacare Corporation | Invacare Corporation decided to recall the product because use of this actuator could result in reduced cross sectional current carrying area due to wear/chafing of the wires. | Nonconforming Material/Component | The firm, Invacare, sent an "URGENT: RECALL INFORMATION" letter dated February 10, 2012 to their consignees/customers. The letter described the product, problem and actions to be taken. Invacare has developed retrofit kits intended to address this issue. The consignees were instructed to contact the customers that have potentially affected product and perform the necessary updates. The consignees were also instructed to complete and return the Provider Response Form to order the retrofit kit.If you have any questions, please contact Customer Service at 1-800-333-6900. |  | 102 | 102 wheelechairs | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CT, FL, GA, IA, IL, IN, KY, ME, MI, MN, MO, MT, NC, NE, NJ, NY, OH, RI, SD, TX, VA, WA, and WI; and countries of: Australia, Canada and New Zeland. | Terminated | August 25, 2014 | 927 | 102
Z-1202-2012 | 61156 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S. Surgical System, da Vinci Si and da Vinci Si-e Surgical Systems. |  |  | February 06, 2012 | 2012 | Intuitive Surgical, Inc. | The holding brake, a component of the da Vinci S Surgical System, may allow passive uncontrolled motion due to gravity during specific power-off conditions. | Device Design | Intuitive Surgical sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated February 3, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provided new instructions for use. Additionally, a Device Correction Response Form was enclosed for customers to complete and return via fax to 408-523-1390. Contact Intuitive Surgical Customer Service at 1-800-876-1310 for questions regarding this notice. |  | 1966 | 1966 systems are in commerce at hospital locations. | Worldwide Distribution. | Terminated | May 09, 2012 | 93 | 1966
Z-1204-2012 | 61296 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The Reprocessed Mitek Arthroscopic Wand is considered a soft tissue ablators. Soft tissue ablators are radiofrequency (RF) surgical tools designed for removal and dissection of tissue in arthroscopic surgeries. A generator or controller serves as the power unit. Soft tissue ablators are available in a wide range of sizes and angles; some models include thermal or suction functions. |  |  | June 01, 2009 | 2009 | Stryker Sustainability Solutions, Inc. | On 06/09/2009, Stryker Sustain ability Solutions, formerly Ascent Healthcare Solutions, recalled lots 674201 &amp; 674290 of their Reprocessed Mitek Arthroscopic Wand due to the lots not being tested in accordance with the functional test requirement during the time frame of 04/14/09 to 05/25/09. | Process control | Ascent Healthcare Solutions notified their customers by presenting a letter dated June 1, 2009 from an Ascent Representative. The representative was to recover any products that may have been affected. Questions were directed to 1-888-888-3433. |  | 48 | 48 | Nationwide Distribution (USA), including the states of: TX, NY, MN, CA, and CT. | Terminated | May 07, 2012 | 1071 | 48
Z-1203-2012 | 61275 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Coupler System II |  |  | February 27, 2012 | 2012 | Mizuho Orthopedic Systems Inc | Mizuho determined that current labeling of The Coupler II system required revisions after investigation of a complaint. | Labeling design | Mizuhosi OSI sent an Urgent Medical Device Recall Communication letter dated February 27, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Mizuho OSI provided their customers of the affected product an updated Owner's manual. Customers were instructed to ensure they remove their existing Owner's Manual from use and repalce with the current Owner's Manual. If customers were unsure of the device's use, they should refer to the included Owner's Manual and ensure all users are properly trained, able to assess the device's status prior to and after each use. If customers were unsure of the status of their device, or of its use, they should call 1-800-777-4674 or outside of the USA 00+1=510-476-8199.Customers were asked to respond that the recall has been received, read, understood and communicated to those affected by this activity. A response should be sent via e-mail to couplerII@mizuhois.com.For any questions regarding this recall call 1-800-777-4674 or outside of the USA 00+1=510-476-8199.was sent out on February 27, 2012 to all affected sites. This was after the original was sent to CDRH for review on January 31, 2012, but approval of revisions was not received until 2/27/12. |  | 1077 | 1077 devices | Worldwide Distribution - USA (nationwide) and the countries of Australia, Belgium, Canada, France, Germany, Hong Kong, Israel, Japan, South Korea, Lebanon, Malaysia, Netherlands, New Zealand, South Arabia, Singapore, Germany, South Africa, Spain, Sweden, UA. E., United Kingdom, and Australia. | Terminated | September 25, 2012 | 211 | 1077
Z-1207-2012 | 61103 | No | OJU | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Wet PVP-I Small Procedure Preoperative Skin Prep Tray containing 1 - 59 mL bottle Povidone Iodine Cleansing Solution (Scrub) USP 7.5% and 1 - 59 mL bottle Povidone Iodine Topical Solution (Paint) USP 10% with Powder-Free Esteem Gloves; 20 trays per case; a sterile, latex-free, single use only tray, solution not exposed to sterilant; U.S. components packaged in Mexico except as indicated: Gloves and Cotton-Tipped Applicators made in China, Sponge Stick made in Mexico; CareFusion, McGaw Park, IL 60085 USA; tray catalog number 4420 |  |  | February 07, 2012 | 2012 | Carefusion 2200 Inc | The povidone-iodine (PVP-I) solution bottles in the trays may have loose caps, allowing leakage and also elevated levels of ethylene oxide (EO) sterilant and/or by-products within the PVP-I solution, which may represent a patient safety hazard. | Packaging process control | CareFusion sent e-mails and follow-up Urgent Recall Notices dated February 7, 2012 to all of the affected customers on the same date, requesting that the customers follow the steps outlined below as mitigation until product can be replaced.1. Please immediately destroy/discard or return any affected product you may have in-stock at your facility. Do NOT further distribute these products.2. Forward this RECALL NOTICE to any customers to whom these products were sold notifying them of the potential risk (reference attachment for affected production lot numbers). End users must be advised to discard or return any CareFusion Skin Prep Tray with the noted product/lot number; distributors should manage this return process per their internal procedures. End users MUST NOT use these products. 3. Complete and return the attached Acknowledgement Form indicating that they have received this RECALL NOTICE and that they will communicate the Recall Notice to all customers who have purchased CareFusion Preoperative Skin Prep Tray catalog numbers 4420, 4468, and 4469A. 4. Contact CareFusion Customer Service at (800) 523-0502 ext. 53576 for details on return and/or replacement of stock. |  | 38440 | 38,440 trays | Nationwide Distribution including Arizona, California, Colorado, Connecticut, Florida, Illinois, Missouri, Nebraska, Nevada, New York, Ohio, Puerto Rico, Rhode Island, South Carolina, South Dakota, Texas, Utah and Virginia. | Terminated | June 05, 2015 | 1214 | 1644150
Z-1205-2012 | 61270 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Brand Name: TempSpan CMT Temporary Cement, a temporary cement. |  |  | March 03, 2011 | 2011 | Kerr Corporation | The recall was initiated because Pentron Clinical has confirmed that the TempSpan Temporary Cement may have black plastic material on the surface of the paste, and it may set faster than specified. | Nonconforming Material/Component | Pentron Clinical sent an Urgent Medical Device Recall letter dated April 20, 2011, to all affected customers. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to complete and fax back the enclosed recall form to 1-877-677-8844 in order to confirm receipt of the customer notification. Customers were instructed to contact Pentron Clinical Customer Service at (800) 551-0283, option 1, directly to handle the arrangements of a quick return and replacement.For questions regarding this recall call 714-516-7400. |  | 627 | 627 units total (601 units in the US) | Worldwide Distribution - USA (nationwide) and the countries of Canada, Philippines, Italy, Netherlands, and the UK. | Terminated | March 14, 2012 | 377 | 627
Z-1206-2012 | 61273 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, MedfusionÂż 3500 Syringe Pumps, Software version 5.0.0, Rx Only. |  |  | January 02, 2012 | 2012 | Smiths Medical ASD, Inc. | Pumps rendered inoperable at power up (displaying a blank screen and alarm). | Vendor change control | Smiths Medical ASD, Inc. contacted two consignees by telephone in November 2010 and January 2011. Smiths Medical worked with one of the two customers to correct 540 of the 595 devices (91%) while on-site in January 2011. The remaining devices were corrected by hospital personnel.The other customer was also made aware of the issue, root cause and corrective action. This customer reported to Smiths Medical that they understood the issue, but did not want Smiths Medical to perform the correction on their pumps at the time (January 2011). They preferred to send their pumps to Smiths Medical for repair, as and when needed. As of December 2011, The hospital has returned 67 devices to Smiths Medical.For questions regarding this recall call 651-628-7477. |  | 1459 | 595 (WI )+ 864 (DC) = 1459 | Nationwide Distribution including WI and DC | Terminated | July 25, 2013 | 570 | 1459
Z-1213-2012 | 61090 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SUPERA VERITAS Interwoven Self-Expanding Nitinol Stent Transhepatic Biliary System. |  |  | January 31, 2012 | 2012 | Idev Technologies, Inc. | The firm is taking this action because the device was one of eight that may have been mislabeled with the incorrect size (length). | Labeling mix-ups | IDEV Technologies sent an "URGENT PRODUCT RECALL" letter dated February 15, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to sign the letter and return a copy to the firm via fax to 281-525-2001 to indicate that they have received the notification. Contact the firm at 281-525-2052 for questions regarding the notice. |  | 7 | Seven devices | Nationwide Distribution-USA (nationwide) including the states of AL, IL, NJ, NV, and PA. | Terminated | September 09, 2013 | 587 | 7
Z-1215-2012 | 61093 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra CUSA EXcel Sterile Tips Rx Only |  |  | October 14, 2011 | 2011 | Integra LifeSciences Corp | A small number of non-conforming CUSA Excel 23 KHz Standard Tip and flue (item Number C4601S) were released for distribution that did not meet Integra's in-process requirements for release. It was discovered during in-process inspection that the max power reading exceeded the specification of 80 watts, causing the stroke range of the tip to exceed it required range by up a value up to 0.002 inches | Release of Material/Component prior to receiving test results | Integra sent an Urgent Product Recall Notification letter dated October 14, 2011 to all affected customers via US Mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory, discontinue use, quarantine and return any available affected lot numbers. Customers were instructed to distribute this communication to any customers to whom they have shipped the affected product. Customers were advise to complete the enclosed Recall Acknowledgement and Return Form and return it as soon as possible to Integra Customer Service. For questions contact Integra Customer Service at 609-936-2659. |  | 24 | 24 | Worldwide Distribution | Terminated | October 19, 2012 | 371 | 24
Z-1211-2012 | 61348 | No | LGX | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | MPI Rapid Response HUT Table, Model #RRHUT 1017. The firm name on the label is Medical Positioning Inc., Kansas City, MO. |  |  | March 05, 2012 | 2012 | Medical Positioning Inc | The warning label regarding electrical shock was not placed on the device warning to have the power switch in the "OFF" position prior to connecting the unit to electric power. | Error in labeling | Medical Positioning Inc. sent a "RAPID RESPONSE HUT TABLE LABELING CORRECTION NOTIFICATION" letter dated March 5, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included an Owner's Manual insert page which clearly states the warning and how to avoid the electric shock hazard and a warning tag to attach to the product's power cord near the plug end using the provided wire tie. Contact the firm at 816-595-2356 for questions regarding this notice. |  | 5 | 5 tables | Worldwide Distribution-USA (nationwide) including the states of IN, KS, WI, NY, IL, NC, OH, and MS and the country of Thailand. | Terminated | July 11, 2013 | 493 | 9
Z-1216-2012 | 61204 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips EasyDiagnost Eleva DRF, Fluoroscopic x-ray system |  |  | February 14, 2012 | 2012 | Philips Healthcare Inc. | Misinterpretation of Mirror Icon-Mirror icon on the bottom of each image could be misinterpreted as aRIGHT lead marker for side indication if the image is mirrored within a PACS system | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 8, 2012 to all affected customers. The letter identified the affected product, problem description, hazard involved, actions to be taken by customer/user and actions planned by Philips. The letter contains important information for continue safe and proper use of the equipment. Philips will contact customers to schedule the installation of the updated software. For further information or support contact your local Philips representative at the Technical Support Line 866-767-2822. |  | 60 | 60 units | Worldwide Distribution - USA Nationwide | Open, Classified |  |  | 60
Z-1218-2012 | 60934 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Elecsys 2010 Analyzer Software version 06-07 |  |  | January 31, 2012 | 2012 | Roche Diagnostics Operations, Inc. | Elecsys 2010 customers have reported getting no value and flag 48 results below expected value range after performing automatic dilution for a sample that previously recovered above the measuring range. Incorrect but believable results may occur for the CEA, Rubella IgG or SHBG assays. The other affected assays will generate either no value or a negative result, both of which are not bel | Software design | Roche sent an Urgent Medical Device Correction letter dated January 31, 2012, via UPS Ground (receipt required) to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following: use one of the workarounds outlined in this Urgent Medical Device Correction (UMDC) if you are running an automatic dilution on an Elecsys 2010 analyzer with software version 06-07; complete the attached faxback form (4600-00-0112) and fax it to 1-888-912-8460; and file this UMDC for future reference.For questions regarding this recall call 1-800-428-2336. |  | 607 | Software version 06-07 was installed on Elecsys 2010 analyzers at 13 accounts. | Worldwide Distribution - USA including NY, GA, DE, SD, WI, ME, OR, IA, AND CA and the countries of Austria, Cyprus, Czech Republic, Estonia, France, Germany,Greece, Hungary, Iceland, Italy, Latvia, Norway, Poland, Portugal, Romania, Slovakia, Slovenia, Spain, and Switzerland | Terminated | September 06, 2012 | 219 | 607
Z-1220-2012 | 61007 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | DuoDiagnost, model 707015, Universal RF system |  |  | December 26, 2011 | 2011 | Philips Healthcare Inc. | The device is currently not compliant with FDA's amended design criteria for this system, the product is not commercially available in the US. However, it was inadvertently shipped to Puerto Rico due to a logistical error. | No Marketing Application | Philips sent an Urgent - Field Safety Notice dated December 26, 2011 to the two affected customers. The letter identified the affected product, problem description, hazard involved, actions to be taken by customer/user and actions planned by Philips. The letter states that customers may continue to use the DuoDiagnost system until the new systems are replaced by Philips. A Philips representative will contact customers to arrange for the replacement of current system. For further information or support contact your local Philips representative. |  | 2 | 2 units | USA Nationwide Distributed in Puerto Rico | Terminated | November 01, 2013 | 676 | 2
Z-1217-2012 | 61193 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | FastPlan Treatment Planning System, Model HZ3, Varian Medical Systems, Palo Alto, CA. Varian FastPlan Treatment Planning System is a radiation treatment planning system composed of a high resolution UNIX graphics computer, a printer, a digitizing tablet, and radiation treatment planning application software. |  |  | January 30, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the CT Localization module of the FastPlan treatment planning system. If the module reports a significant difference in tilt or spin on a slice and the use elects to "Discard" that slice, an incorrect calculation of the isocenter location may result. | Device Design | "Urgent Medical Device Correction/Urgent Field Safety Notification" letters were sent on January 30, 2012 by certified mail to consignees in the US and by Federal Express for Canada. International sites will be notified by Varian International Customer Support Managers, who will translate letters if necessary. The notice provided a description of the issue, the actions that can be taken to avoid or mitigate the issue, and the steps Varian is taking to address the issue. Customer service number for US and Canada is: (888) 827-4265. |  | 201 | 201 | Worldwide Distribution - Nationwide (USA) and the countries of Canada, China, Germany, Ireland, Spain, Singapore, Georgia, Hong Kong, India, Macau, Malaysia, Russia, Sri Lanka, Taiwan, Japan, Chile, Columbia, Venezuela and Australia. | Terminated | October 21, 2013 | 630 | 201
Z-1219-2012 | 61250 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Brilliance Workspace Portal is a software component of the Computed Tomography X-Ray System. |  |  | February 23, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Incorrect Standardized Uptake Value ("SUV") values are displayed in two instances:1. After loading PET data (with a Slope value different than 1) into the Quick Review (QR) application from a remote device; and2. During the loading the above mentioned type of PET data into the CT viewer application from a remote device. | Software design | The Philips Healthcare, sent an "URGENT-Medical Device Correction" Field Safety Notice dated February 08, 2012 to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to review the information with all members of their staff and to retain a copy with the equipment Instruction for Use. The letter notifies the customers that a Field Service Engineer will be contacting them to schedule a visit to implement the corrective action by installing the software update version 2.6.1.If you need further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 822 | 822 | Worldwide distributions: USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, VA, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, VT, WI, and WV; and countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech, Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Libya, Malaysia, Malta, Martinique, Mexico, Netherlands, Netherlands Antilles, New Zealand, Norway, Poland, Portugal, Reunion, Russia, Saudi Arabia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uzbekistan, and Venezuela. | Terminated | February 22, 2013 | 365 | 822
Z-1221-2012 | 59992 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Crystal Metal Disposable Blades size Mac 4 (part number 50674, 50674/F, 50674/BMI) and Mac 3 (part numbers 50675, 50675/F, 50675/BMI) |  |  | September 21, 2011 | 2011 | Penlon, Ltd. | Disposable laryngoscope blades incorrectly labeled with wrong size. | Error in labeling | InterMed Penlon sent and Urgent Field Safety Notice letter dated September 21, 2011 to affected customers. The letter identified the affected product, problem and advice on actions to be taken by user. Customers were instructed to complete the Reply Form and send to the address provided. |  | 24940 | 24,940 cartons | Worldwide Distribution - USA (nationwide) in the countries of Belgium, France, Greece, Ireland, Italy, Lithuania, Netherlands, Norway, Portugal, Switzerland, Canada, China, Hong Kong and Malaysia. | Terminated | June 19, 2012 | 272 | 24940
Z-0910-04 | 35768 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Quick-set Plus Infusion Sets. Models MMT-359S6; MMT-359S9; MMT359L6; &amp; MMT-359L9. (differences in models is variation in length, 6 versus 9mm) |  |  | March 02, 2004 | 2004 | Medtronic MiniMed | Possibility of interrupted insulin flow. This posting is for a recall that occurred in March 2004, also posted on the FDA website in 2004. | Component design/selection | Firm will ship all users (customers) one box at a time every three weeks for the next several months. A recall notification letter was issued to users and to healthcare professionals, and distributors. Letters were mailed from March 2-5, 2004. |  |  | Undetermined | Nationwide | Terminated | March 19, 2012 | 2939
Z-1124-2012 | 59944 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova X-ray Imaging Systems. |  |  | May 04, 2011 | 2011 | GE Healthcare, LLC | It was discovered that when the Diamentor M4-KDK is installed on an Innova system the dose display is disabled on the certified Innova monitor and enabled on the Diamentor M4-KDK which is not certified as required by FDA regulation (Subpart J). | Radiation Control for Health and Safety Act | GE Healthcare will alert affected Innova users with a Diamentor M4-KDK unit of this issue and the workaround to mitigate this issue. GE Healthcare will bring affected Innova units in the field into compliance by enabling the Innova certified display and providing affected customers with an Operators Manual Addendum and Service Manual Addendum which will address this topic.GE Healthcare will correct all affected systems by updating the system's Operator and Service Correction manuals and enabling the innova dose display on the Live and Reference monitors and the DLconsole. The Diamentor M4-KDK can still be used to display additional informative dose values. GE Healthcare will, without charge, remedy the defect or bring the product into compliance; the details of which will be included in a subsequent communication to you or through a GE field engineer site visit. GE Healthcare updated the Innova User Instructions, as part of the recall, to indicate to meet regulatory requirements based on the international safety standards IEC 60601-2-43:2000 and/or 21 CFR 1020.30 (k), use as reference the exposure dose displayed on the DL console, live or reference monitor. If you have any further questions or concerns please call 800-437-1171. |  | 19 | 19 | USA Nationwide Distribution | Terminated | March 29, 2012 | 330 | 19
Z-1226-2012 | 61183 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Integra Endoscopic Gastrocnemius Release (EGR) System. The device consists of one (1) unit which is kitted within an instrumentation tray. The Integra EGR System includes a range of instruments for endoscopic approach to the Gastrocnemius Aponeurosis Recession. The cutting instrument features a retracting blade for cutting the gastrocnemius aponeurosis. Associated instruments include an elevator, a cannula, and a cannula obturator, which are all consumable instruments provided sterile. The EGR System is indicated for the treatment of posterior heel cord contracture (Equinus) in those patients who fail to respond to conservative management. |  |  | January 06, 2012 | 2012 | Integra LifeSciences Corp. | The firm received several complaints from customers reporting that the blade of the EGR system may not retract back into the device following surgery. The firm attributes this problem to user error or improper surgical technique on the part of the physician. | Labeling mix-ups | The firm sent a 'A Medical Device Correction Notice' with an attached Acknowledgement and Return Form to all consignees on 1/6/2012 via FedEx and E-mail. The Correction Notice informs customers of the firm's receipt of several complaints regarding the Endoscopic Gastrocnemius Release (EGR) System and the problem with non-retraction of the blade back into the device during surgery. Although no adverse health consequences or patient injuries have been reported to date, the notice is being issued by the firm in order to reinforce the importance of following the additional clarification to the Surgical Technique Manual. According to the firm, this action should serve to reduce the possibility in which the blade may fail to retract into the device. Questions were directed to (609) 936-6832. |  | 850 | 850 units | Nationwide Distribution, including the states of: MI, CA, TX, VT, OK, OH, NC, MN, CT, CO, TN, NY, NJ, MO, FL, WA, AR, MT, VA, IL, PA, NE, GA, AZ, WI, OR, MD, LA, IA, AL, MA, ID, KY, UT, NV, IN and SC. | Terminated | January 23, 2014 | 748 | 850
Z-1225-2012 | 61138 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew BIOSURE HA Interference Screw, 7 mm x 25 mm |  |  | February 06, 2012 | 2012 | Smith &amp; Nephew, Inc. Endoscopy Division | Mislabeled: The device was labeled as a standard thread screw (P/N 72201772) on pouch and box,however the device contained in the packaging is a reverse thread screw (P/N 72201773). | Packaging process control | Smith &amp; Nephew Endoscopy sent an "URGENT-PRODUCT RECALL 1st NOTIFICATION" dated February 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The letter instructs customers to inspect their inventory and quarantine any affected product. For instructions on returning product, contact the Smith &amp; Nephew Endoscopy Division Returns Group at 800-343-5717 ( 0ption 3) or email to endo.andreturns@smith-nephew.com. Additionally, customers were asked to complete the bottom portion of the letter and to fax it to 1-508-261-3620. Contact the Smith &amp; Nephew Endoscopy Division for questions regarding this notice at 1-508-261-3731. |  | 194 | 194 units | Worldwide Distribution-USA (nationwide) and the states of AK, AZ, CA, CO, CT, IA , MD, MI, MN, MO, MS, NE, OH, OK, PA, SC, and VA and the countries of Australia, Canada, Germany, South Korea, Spain, Switzerland, Thailand, and United Kingdom. | Terminated | February 29, 2016 | 1484 | 194
Z-1227-2012 | 61243 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Duet TRS Articulating 60- 3.5MM, Single Use Loading Unit |  |  | February 16, 2012 | 2012 | Covidien LP | Missing component results in the stapler not firing. | Process control | The firm, Covidien, sent a "Medical Device Recall" letter dated February 16, 2012 to customers via federal express. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of Lot# N1H003LX of Duet TRS Articulating Loading Unit and segregate affected product; return inventory to Field Returns Department, 195 McDermott Road, North Haven, CT 06473 and contact customer service at SDFeedback@Covidien.com or (800) 722-8772, option #1 for RGA # prior to returning; and complete and return the Recall Product Return Form via fax to (203) 822-6009 (all affected units must be returned with completed form through the Distributor). If you have any questions or concerns, please do not hesitate to contact your Covidien representative or Quality Assurance at 203-492-5232, or Customer Service at 800-722-877 option 1 customer service. |  | 1158 | 1158 units | Worldwide distribution: USA (nationwide) and country of: Venezuela. | Open, Classified |  |  | 1158
Z-1223-2012 | 61274 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Vanguard Knee Instrumentation Tibial Resection Tower Universal with alignment tower geometry, REF 32487557, LOT 2561334, Biomet UK Ltd, South Wales, UK |  |  | February 14, 2012 | 2012 | Biomet U.K., Ltd. | This recall was initiated following a report that a bag labeled as REF. 32-487557/Lot no 2561334 (Vanguard Tibial Resection head universal w/ alignment tower) actually contained a Ref. 32-487002/Lot no. 2561329 (Vanguard Distal cutting Block w/out handle attachment). Investigation found packaging of the affected product may contain the incorrect item. | Packaging change control | Biomet U.K. sent a Field Safety Corrective Action notice dated February 14, 2012 to all affected customers. The notice identified the affected products, problem and actions to be taken. Customers were instructed to locate, discontinue use and quarantine the recalled products pending return to Biomet. The letter recommended customers provide the notice to any other organization who may have received the affected instruments. Customers were instructed to complete and return without delay the attached "Fax-Back Response Form" acknowledging receipt of notice and indicating the quantity of products to be returned. For questions call +44 (0) 1656 761658, fax: +44 (0) 1656 645454 and e-mail uk.complaints@biomet.com |  | 4 | 4 units | Worldwide Distribution - in the countries of France, Japan, Italy, Netherlands, UK , no devices distributed in the USA | Terminated | November 06, 2012 | 266 | 4
Z-1229-2012 | 61122 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron CX Immuno-Protein Cal, part number: 442840, |  |  | August 26, 2009 | 2009 | Beckman Coulter Inc. | A recall was initiated because Beckman has confirmed that the standardization of the Immuno Protein Calibrator Set has been re-established to IFCC Reference Preparation for Plasma Proteins, BCR-470 beginning with lot number M901330. | Nonconforming Material/Component | Beckman Coulter sent an Urgent Field Safety Notice letter dated August 24, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to share the information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Customers with questions regarding the product correction action were instructed to contact their local Beckman Coulter Representative.For questions regarding this recall call 714-993-8916. |  | 1923 | 1,923 units total (1,882 units in US) | Worldwide Distribution - USA (nationwide) and Canada | Terminated | April 17, 2012 | 965 | 1923
Z-1230-2012 | 61197 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 Systems, Product Code M-4800-01 |  |  | February 09, 2012 | 2012 | Biosense Webster, Inc. | Some CARTO 3 Systems may get radio frequency interference through certain interconnected cable configurations resulting in jumping or disappearing catheter icon and/or error messages during ablation. | Nonconforming Material/Component | Three Urgent Customer Notification recall communications were initiated 2/9/12 with Biosense forwarding the notification to customers with affected backplane cards (electrophysiology lab managers, a notification for customers with affected back plane cards (risk managers), and a notification for customers with affected backplane cards and smarttouch module along with customer acknowledgement forms. The letters provides the customers with an explanation of the problem identified and steps to take in the event that the issue is observed. If the steps provided do not resolve the issue, customers are to contact their Biosense Webster representative. A Biosense Webster representative will contact customers to schedule an appointment for the replacement of the Backplane card.Customers were instructed to complete and return the attached Acknowledgement Form in accordance with the instructions listed on the form. If customers have questions, they can contact their BWI representative. |  | 134 | 134 units total (32 units in US) | Worldwide Distribution -- US, Japan, China, Ireland, Germany, Italy, Slovenia, Indonesia, Poland, Spain, United Kingdom, Switzerland, Belgium, and Canada. | Terminated | September 06, 2012 | 210 | 134
Z-1271-2012 | 61029 | No | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IMMULITEÂż/IMMULITEÂż 1000 CK-MB LKMB --- Catalog number: LKMB1 (100 tests), LKMB5 (500 tests) --- Test code: CMB Color: Dark Blue. |  |  | January 25, 2012 | 2012 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics has observed that the median for the reference range with current kit lots of the CK-MB assay differs from the guidelines published in the Instructions for Use (IFU). This difference applies to all IMMULITE platforms. | Nonconforming Material/Component | Siemens sent an "URGENT DEVICE RECALL" letter dated January 2012 to all affected customers. The letter identifies the product , problem, and actions to be taken by the customers. Siemens recommends that customers reestablish their own medians and reference ranges, in line with the recommendations made in the IFU for these products. A Field Correction Effectiveness Check form was included for customers to complete and return to the firm. Customers with questions or who need additional information should contact their Siemens Technical Solutions Center. |  |  |  | Worldwide Distribution. | Open, Classified
Z-1231-2012 | 61104 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | REF 905463, 2.9 mm Micromax Anchor Long Shaft, Size 2 Blue/White Maxbraid (Lactosorb L-15 resorbable copolymer ) Sterile, EO; Biomet Sports Medicine, Warsaw, IN |  |  | January 25, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action following an internal audit which discovered the component level package insert was included in the product packaging in place of the end level product package insert (01-501072).This meant that the Indications, Contraindications, Possible Adverse Effects, and some Warnings are not present with the packaged device. | Packaging | Biomed Inc. sent an Urgent Medical Device Recall Notice dated January 25, 2012 to distributors, hospitals, and OR managers via US mail. The notice identified the product, description of problem, possible adverse health consequences, and actions to be taken, including directions to immediately locate devices and stop device use ,with directions for return. A copy of the correct package insert was included with directions to report any procedures performed outside of indications for use. Receipt confirmation was requested by calling 800-348-9500, extension 3983 or 3009. The notice included a list of units invoiced to the account with a FaxBack Response form to be filled out and sent to the Attention of Angie Dickson, Biomet, 56 East Bell Drive Warsaw, In 46580. Questions concerning the recall should be directed to 574-371-3983 or 574-371-3009, M-F 8-5PM(EST) |  | 787 | 787 | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Brazil, Chile, Colombia, Israel, Mexico and Netherlands. | Terminated | April 12, 2013 | 443 | 4212
Z-1241-2012 | 61255 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER LH 780 Analyzer, |  |  | December 01, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Auto Stop and Auto Transmit features for Quality Control become disabled after scanning assay values from Coulter 5C and Retic-C Cell Controls assay sheets using the 2D Barcode scanner. | Software design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 30, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers who use the Auto-Stop or Auto-Transmit to ensure the settings are enabled in Control Setup after scanning the assay sheets. Additionally, a Response Form was included for customers to complete and return. Contact Beckman Coulter Customer Service at (800) 526-7694 for questions regarding this notice. |  | 8922 | 8,922 units total (4,073 units in US) | Worldwide Distribution-USA (nationwide) including the Virgin Islands and the countries of Algeria Andorra Angola Antigua and Barbuda Argentina Australia Bahrain Bangladesh Belgium Bermuda Brazil Bulgaria Burundi Cameroon Canada Chile China Colombia Croatia Czech Republic Denmark Ecuador Egypt Ethiopia France Germany Greece Guadeloupe Guam Guatemala Honduras Hong Kong Hungary Iceland India Indonesia Iraq Ireland Israel Italy Japan Jordan Korea, Republic of Kuwait Latvia Lebanon Libyan Arab Jamahiriya Lithuania Macao Malaysia Mayotte Mexico Monaco Morocco Myanmar Namibia Netherlands New Zealand Norway Oman Pakistan Panama Philippines Poland Puerto Rico Qatar Reunion Romania Russian Federation Saudi Arabia Serbia Singapore Slovakia Slovenia South Africa Spain Sri Lanka Sweden Switzerland Taiwan Thailand Trinidad and Tobago Tunisia Turkey United Arab Emirates United Kingdom Uruguay Venezuela Viet Nam Yemen Zimbabwe | Terminated | September 02, 2014 | 1006 | 8922
Z-1239-2012 | 61214 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Ostase QC, Part Number: 37309. |  |  | November 10, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman confirmed that the Access Ostase Calibrator and QC lots do not meet expiration date claims. The Access Ostase Reagent Kit lots identified were quality control tested and released using Access Ostase Calibrator and QC lots that contain the implicated raw material. | Process design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 9, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard all remaining inventory. In addition, a Response Form was enclosed for customers to complete and return. Contact Beckman's Customer Support Center at (800) 854-3633 for questions regarding this notice. |  | 4961 | 4961 units total (1321 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Argentina, Australia, Austria, Belgium, Canada, China, Colombia, Czech Republic, Denmark, France, Georgia, Germany, Greece, Hong Kong, Hungary, India,Ireland, Italy, Japan, Jordan, Lebanon, Liechtenstein, Mexico, Morocco, Netherlands, New Zealand, Panama, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela. | Terminated | June 26, 2012 | 229 | 4961
Z-1214-2012 | 61298 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | D-SPECT Cardiac Scanner System -- COMMON/USUAL NAME: Emission Computed Tomography System --- SPECTRUM DYNAMICS --- Spectrum Dynamics (Israel) Ltd. 22 Bareket St. North Industrial Park, P.O. Box 3033, Caesarea, 30889 Israel |  |  | July 15, 2011 | 2011 | Spectrum Dynamics | The firm received a report related to the detector housing on the D-SPECT Cardiac Scanner. Specifically, it was reported that if the detector housing securing mechanism is compromised, the detector housing may pull away from its secured position on the gantry and potentially come intocontact with a patient. The reported event was discovered during set up and thus did not involve a patient or p | Nonconforming Material/Component | Spectrum Dynamics sent a "Voluntary Field Correction" letter dated July 15, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken.To implement this fix, a BC Technical Field Service Engineer contacted users to schedule a service visit. In the meantime, Spectrum Dynamics did not believe that this issue warranted the discontinuation of use of the D-SPECT Cardiac Scanner. However, in an abundance of caution, if the detector housing appeared to be loose, or appeared to be disconnecting from the gantry, do not use the system. Please contact your local BC Technical Field Service Engineer immediately.Enclosed is a verification form. Please fax the completed form at your earliest convenience to 415.663.4225. For further information or support concerning this issue, please call 925.922.2615. |  | 42 | Domestic- 35 units, Foreign - 7 units | Worldwide Distribution -- USA (nationwide) and the countries of France and United Kingdom | Terminated | April 01, 2015 | 1356 | 42
Z-1270-2012 | 61308 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ System; |  |  | February 14, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Images used in treatment planning may belong to another patient. Risk of treatment to wrong area. Software update to upgrade Siemens installed base to the latest release of MOSAIQ version 2.00X3 SP5, for RT Therapist version 2.1A | Device Design | Siemens sent a Customer Information letters to all affected customers. The Customer information letters inform customers of the release of MOSAIQ 2.00X3 SP5 which specifically addresses a safety issue that is only found whenMOSAIQ is used with COHERENCE Therapist 2.1 MQ (also called RTT2.1A ). This letter is to be kept with Customer information documentation. Update Instruction TH004/12S are being issued with the update. The Software update is delivered and installed by qualified Field Service Engineers. For questions, contact your responsible SIEMENS Service Engineer or Application Specialist. |  | 44 | 44 affected Siemens Customers | Worldwide Distribution - USA (nationwide) including the states of KY, PA, TN, WI, and P.R. and the countries of Australia, Belgium, Bosnia Herzeg. China, Colombia, Croatia, Czech Republic, Germany, India, Ireland, Japan, Lebanon, Malaysia, Mexico, New Zealand, Philippines, Poland, Russian Federation, South Africa, Saudi Arabia, Thailand, Turkey, and United Kingdom. | Terminated | September 21, 2012 | 220 | 44
Z-1275-2012 | 57924 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CRYO-206V and CRYO-206F Procedure Kit, which includes: Cryo Probes, Urethral Catheter, and temperature probes. |  |  | February 27, 2012 | 2012 | Healthtronics, Inc. | HealthTronics has initiated this recall of these products because there have been instances where the temperature probes (aka TempProbes) in the kit may be mislabeled. In some instances, the label near the prove does not match the label at the connector. If the labeling error went undetected, the TempProbes would not be monitoring temperature in the area believed to be under surveillance. | Labeling mix-ups | Endocare sent an Urgent: Device Recall letter dated February 27, 2012 to all affected consignees. The letter identified the affected product, problem and actions to be taken. The letter instructed consignees to immediately cease all distribution, examine their stock and the stock of accounts that have been supplied and remove any affected devices from all inventories. If the product have been further distributed contact Endocare for instruction. If consignees have any of the affected products: 1) Contact Customer Service at 877-367-4873 to obtain a RMA number, 2) Complete the Verification Section on page two, and 3) Return this letter with the affected product to the address provided. For questions about this recall contact Endocare, Inc at 877-367-4873. |  | 9274 | 9274 units | Worldwide Distribution--USA (nationwide) and the countries of Barbados, Brazil, Canada, China, Italy, Jakarta, Malaysia, Madrid Spain, Taiwan, and United Kingdom. | Terminated | March 19, 2013 | 386 | 9274
Z-1269-2012 | 61294 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Touch Panel/Control Unit Carrier of the Swingarm M820/M844 C19, article no. 10448194MO; a component of the Leica M820 C19 Surgical Microscope, article 10448196, and the Leica M844 C19 Surgical Microscope, article 10448197; ceiling mounted surgical microscopes; Leica Microsystems (Schweiz) AG, Max Schmidheiny-Strasse 201, CH-9425, Heerbrugg, Switzerland;; The intended use of the microscope is for improving the visibility of objects through magnification and illumination. It can be applied for observation and documentation and for human and veterinary medical treatment. |  |  | February 27, 2012 | 2012 | Leica Microsystems, Inc. | The screws, which lock the bottom plate to the spindle, pull through allowing the carrier arm with touch panel/control unit of the surgical microscope to disconnect, which potentially could result in them falling. | Device Design | The firm, Leica Microsystems, sent a "Medical Device Correction" letter dated February 27, 2012 to the affected customers via e-mail on the same date. The letter described the product, problem and action to be taken. The customers were informed that the Leica M820 and M844 C19 surgical microscopes require upgrading. An upgrade kit will be provide for replacement of the flange and axis of the touch panel/control unit carrier. The Customers were instructed to NOT USE the affected Leica surgical microscope until the unit has been upgraded by their Leica Microsystems representative and to complete the attached Acknowledgement Form, acknowledging the receipt and understanding of the letter, listing the serial numbers of the unit at their facility and fax it back to Leica at 1-847-236-3747. If they have any questions, contact Director, Global Product Quality Management at 1-847-317-7209. |  | 33 | 33 units | Worldwide distribution: USA (nationwide) including states of: California, Connecticut, Kansas, Missouri, New York, North Dakota, Ohio, Pennsylvania, Texas and Utah; and countries of: Japan and South Africa, | Terminated | May 18, 2012 | 81 | 33
Z-1246-2012 | 61314 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | ENSEALÂż 5 mm Diameter Tissue Sealer G2 14 cm Length Straight Jaw, Model # NSLG2S14 and ENSEALÂż 5 mm Diameter Tissue Sealer G2 25 cm Length Straight Jaw, Model # NSLG2S25. Ethicon Endo-Surgery, LLC Guaynabo, Puerto Rico 00969 |  |  | January 26, 2012 | 2012 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated a voluntary global recall for specific production lots of ENSEALÂż G2 Curved and Straight Tissue Sealers due to two potential issues that may occur related to the activation button: (i) Continuous Activation: The ENSEALÂż device/system may continue to activate when the activation button is released prior to reaching the end of cycle (tone 3). (ii) No Activation | Nonconforming Material/Component | Ethicon Endo-Surgery sent an Urgent Medical Device Recall letter dated February 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop use of the affected product, fill out the attached Business Reply Form (BRF) and return to and return the affected product to:Ethicon Endo-SurgeryATT: ENSEAL G2 Recall 4545 Greek RoadCincinnati, Ohio 45242Customers were instructed to choose one of the following response options:Return the Business Reply Form to their sales representativeCall 1-800-873-3636, Option 6Fax the BRF to 1-513-337-4138For any questions regarding this recall call 513-337-3419. |  | 84 | 84 pieces | Worldwide Distribution - USA including AZ, CA, CT, FL, IL, IN, KS, KY, MA, MI, MO, NC, NY, OH, PA, TN &amp; TX and the countries of Arab Emirates, Austria, Belgium, France, Germany, Great Britain, Italy, Jordan, Kingdom of Saudi Arabia, Lebanon, Portugal, Slovenia, Sweden and Turkey. | Terminated | February 15, 2013 | 386 | 1035
Z-1263-2012 | 60753 | No | MCY | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Wound Management Tray, catalog# 900-540 |  |  | December 08, 2011 | 2011 | Customed, Inc | The firm has determined that there is a possibility that packaging seals may be compromised. The issue could result in an injury to the patient due to product contamination or loss of sterility condition. | Packaging | Customed Inc sent an Urgent Medical Device Recall letter dated December 8, 2011 to distributors via e-mail; then by certified (return and receipt) mail. The letter identified the product, description of problem, and actions to be taken. The letter included the following instructions:1. Product should be removed from sale (Quarantine them) and cease distribution. Products affected shall NOT be used.2. Return the enclosed acknowledgement form to Customed, Inc. Recall Coordinator whether or not the consignee has affected products. 3. If products has been futher distributed, sub-recall products - It is extremely important to notify your customers of this action by including a copy of this leller along with your formal notification. Please, instruct your customer to follow your directions for credit/replacemenl if they purchased the product from you. For qucstions. contact Cuslomed, Inc. Recall Coordinator at 787-622-5151 Ext. 7510. |  | 400 | 400 units | USA Nationwide Distribution - including Florida and Puerto Rico. | Terminated | April 12, 2017 | 1952 | 184527
Z-1268-2012 | 61236 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Vitamin D Diluent 2-pack, |  |  | February 16, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Automatic and system-performed dilution calculation factor for Vitamin D dilutions is incorrect and causesdiluted patient samples to under recover by approximately 50%. | Nonconforming Material/Component | The firm, Siemens Healthcare Diagnostics, sent an "Urgent Device Recall Notice" on February 16, 2012 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to discontinue using the automatic and system-performed dilution features for Vitamin D; to discontinue the use of ADVIA Centaur VitD Diluent, 2-pack (REF 10494100); to manually dilute samples with values greater than the assay range of 150 ng/mL (375 nmoI/L), using the ADVIA Centaur VitD Diluent, 1 bottle (REF 10632114) as stated in the Vitamin D assay Instructions For Use (IFU); to review the contents of the recall notice with their Laboratory Director and retest the previous sample results generated using system- performed dilutions, and complete and return the confirmation fax-back form via fax to TECHNICAL SOLUTIONS CENTER at (302) 631-7597.If you have any questions or need additional information, please contact your local Technical Support Provider or Distributor or call (508) 668-5000. |  | 442 | 442 units | Worldwide distribution: USA (nationwide) and countries of: Columbia, China, Denmark, Egypt, and Hong Kong. | Terminated | October 17, 2013 | 609 | 442
Z-1278-2012 | 61351 | No | KNN | Wheelchair component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Permolock C3, Docking system for Power wheelchair in vehicle. |  |  | February 21, 2012 | 2012 | Permobil Inc | The fasteners used between June 3, 2011 and January 31, 2012 to attach the Permolock C3 pins to wheelchairs manufactured in the United States do not contain the required nylon ring, which allows for a more secure attachment. | Mixed-up of materials/components | The firm, Permobil, sent an "URGENT DEVICE RECALL" letter dated February 27, 2012 via fax to all customers. For those customers where Permobil did not have a valid fax number, the notices were mailed via first class United States mail. The letter describes the product, problem and actions to be taken. Permobil will designate Permobil personnel to complete an update to the Permolock C3 device.If a Permobil representative has not contacted you already, please call 800-736-0925 to schedule the service. |  | 147 | 147 units | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, MD ,MI, MN, MO, NC, NE, NJ, NV, NY, OH, OK, OR, PA, TN, TX, UT, VA, VT, WV, and WY; and countries of: Australia, Canada, France, Netherlands, Sweden, and United Kingdom. | Terminated | August 13, 2013 | 539 | 147
Z-1277-2012 | 61175 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products --- ALT Slides --- [REF] 165 5281 --- [Mfg Symbol] Ortho-Clinical Diagnostics, Inc., Rochester, NY USA. The recalled device is the calibration disk contained within the VITROS Calibrator Kit 3. The is can be identified with the label "VITROS Chemist y Systems --- Universal Calibration Diskette --- [REF] 871 6607 --- Ortho-Clinical Diagnostics" |  |  | February 02, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics received complaints of lower than expected results (&lt; 6 U/L) for samples that were known to be outside of the measuring (reportable) range (&gt; 1000 U/L) when using Calibration Diskette/Assay Data Diskette Data Release Version (DRV) 5653 through 5666 and VITROSÂż Chemistry Products ALT Slides, Generation (GEN) 26 (Product Code 1655281). | Software design | Ortho Clinical Diagnostic sent a Important Product Correction letter dated February 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed customers to: Install Calibration Diskette/Assay Data Diskette DRV 5667 on your VITROS System. Following a successful recalibration, you may resume using VITROSÂż ALT Slides, GEN 26 remaining in your inventory. To please complete and return the Confirmation of Receipt form upon receipt of this notification. Please return this form no later than February 17, 2012. Post this notification by each VITROSÂż System in your facility that utilizes the VITROSÂż ALT Slides, or with the VITROSÂż user documentation. Forward this notification if you have distributed this product outside of your facility.We apologize for the inconvenience this may have caused your facility. If you have any additional questions, please contact Customer Technical Services at 1-800-421-3311. |  | 4306 | 4306 | Worldwide Distribution - USA (nationwide) | Open, Classified |  |  | 4306
Z-1276-2012 | 61312 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Osteonics Shoulder Glenoid Spanner |  |  | November 17, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Two product complaints reported that the Glenoid Reamer Spanner (Lot F3W7017) would not properly connect to the Glenoid Spherical Reamer. | Nonconforming Material/Component | Stryker sent an "URGENT PRODUCT RECALL" letter dated December 21, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Product Recall Acknowledgment Form was attached for customers to complete and return via fax to 201-831-6069. Contact the firm at 201-972-2100 for questions regarding this notice. |  | 5 | five units | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | December 19, 2012 | 398 | 5
Z-1248-2012 | 61356 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic EnTrust, Dual Chamber Implantable, Cardioverter Defibrillator with Atrial and Ventricular Therapies, models D153ATG and D154ATG. The device is indicated for use in ICD patients with atrial tachyarrhythmias, or who are at significant risk of developing atrial tachyarrhythmias. The device is intended to provide ventricular antitachycardia pacing and ventricular defibrillation for automated treatment of life-threatening ventricular arrhythmias. In addition, the device is intended to provide pacing, cardioversion, and defibrillation for treatment of patients with symptomatic, drug-refractory atrial fibrillation and/or life-threatening ventricular tachyarrhythmias. Sterile. Medtronic, Inc., Minneapolis, MN 55432 USA. |  |  | March 06, 2012 | 2012 | Medtronic Inc. Cardiac Rhythm Disease Management | A small percentage of EnTrust ICDs may not meet expected longevity or provide at least three months of device operation between the Elective Replacement Indicator (ERI) and End of Life (EOL) due to a more-rapid-than-expected drop in battery voltage. | Device Design | An "Urgent Medical Device Correction" letter was sent to physicians beginning 3/6/2012. The letter described the issue and provided patient management recommendations. Additionally the letter stated that ongoing updates to the product performance report will be posted on a Medtronic website. |  | 68804 | 68,804 devices worldwide | Worldwide Distribution -- USA, and the countries of: Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Bolivia, Brazil, Canada, Cayman Islands, Chile, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jamaica, Jordan, Kuwait, Lebanon, Macedonia, Martinique, Morocco, Netherlands, Netherlands Antilles, New Zealand, Norway, Oman, Peru, Poland, Portugal, Qatar, Reunion, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | July 27, 2012 | 143 | 68804
Z-1281-2012 | 61420 | No | KZA | Liquid crystal vein locator. | General Hospital | General Hospital | 510(K) Exempt | VeinViewer (R) by Luminetx, Model Number VV1.1 GS, Part No. P00800-L, Luminetx Corporation, Memphis, TN 38104 |  |  | March 14, 2012 | 2012 | Christie Medical Holding Inc | Instrument may tip over if not moved correctly. | Employee error | Christie Medical Holdings, Inc. sent an Urgent Medical Device Correction by letter on March 14, 2012 to all affected customers. The letter identified the product, problem and actions to be taken. The letter includes "Correct Transport" and "Do Not Pull labels to be applied to the unit. Customers were instructed to immediately examine their inventory, quarantine all product subject to correction and apply the labels prior to use. A Service Bulletin is enclosed that should be kept and placed with the User Guide for future reference and a Response Form to be completed and return as soon as possible. For questions contact Technical Service at 901-721-0300. |  | 600 | 600 units | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Australia, China, Brazil, Czech Republic, Canada, Germany, Denmark, Spain, Great Britain, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, South Korea, Kuwait, Mexico, Malaysia, Netherlands, Norway, Philippines, Russian Federation, Saudi Arabia, Singapore, Thailand, Turkey, Taiwan, and United Arab Emirates | Terminated | October 23, 2012 | 223 | 1050
Z-1279-2012 | 61205 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Codman ICP Skull Bolt Kit, Product Code: 82-6638. Single Use, disposable drill kit for gaining access to the cranium to allow for insertion and retention of intracranial cranial pressure sensor, included in the kit. |  |  | February 09, 2012 | 2012 | Codman &amp; Shurtleff, Inc. | Codman ICP Skull Bolt wrench is out of specification and cannot be used to turn the screw on the drill | Nonconforming Material/Component | Codman issued "Urgent - Voluntary Product Recall Notification" letters on 2/9/12 via Fedx in the US and requested accounts examine inventory and return product. An Acknowledgement Form was requested to be returned. Outside the US affiliates were notified to recall to their accounts. Questions were directed to Customer Service (800) 225-0460. |  | 4056 | 4056 kits | Worldwide Distribution - USA, and the countries of Canada, Austria, Brazil, Belgium, Columbia, China, Denmark, Estonia, Spain, Finland, France, Great Britain, Germany, India, Italy, Japan, Lithuania, Netherlands, New Zealand, Oman, Mexico, Portugal, Russia, Singapore, SIovenia, and Switzerland. | Terminated | September 11, 2015 | 1310 | 4056
Z-1287-2012 | 60718 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns Soft-Flow Aortic Cannula; Aortic cannula, 7.0 mm (21 Fr) OD with 3/8" connector, 12" (31 em) long. |  |  | February 23, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During in-house inspection, Terumo Cardiovascular Systems discovered one Soft-Flow Aortic Cannula with a hole in the tip, just above the suture ball. | Nonconforming Material/Component | Terumo sent an URGENT MEDICAL DEVICE RECALL letter dated February 23, 2012 via Federal Express to all US customers. Terumo Foreign Affiliates received the recall notification via email on February 24th, 2012. The letters identified the affected products, reson for removal, potential hazard, affected population and actions to be taken. Customers were instructed to review the Medical Device Removal notice, assure that all users are aware of the notice, confirm receipt of this communication by faxing the attached Customer Response Form to the fax number provided or email to tcvs.recall@terumomedical.com and call Terumo CVS Customer Service for a Returned Goods Authorization (RGA) number. For questions call Terumo CVS Customer Service at 1-800-521-2818. |  | 188530 | 188,282 Cannulae and 248 Custom Tubing Packs | USA Nationwide Distribution including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and the countries of: CANADA, CHILE, India, Japan, Singapore, AUSTRALIA, Malaysia, BELGIUM, BELGIUM, BRAZIL, CHILE, BRAZIL, BRAZIL, and Philippines. | Terminated | April 03, 2013 | 405 | 4718744
Z-1284-2012 | 61278 | No | MIH |  |  | Cardiovascular | PMA | Talent Captivia Stent Graft System - FreeFlow configurations only. |  |  | February 21, 2012 | 2012 | Medtronic Vascular, Inc. | Complaints received related to the inability/difficulty to release the proximal (FreeFlow) bare stent using the tip capture mechanism. | Process design | Medtronic, Inc. sent an Urgent Medical Device Safety Information letter dated February 2012, to all affected customers. The letter identified the product, the problem, and the corrective action taken by Medtronic. Medtronic updated the Instructions for Use (IFU) manual.For questions regarding this recall call 707-566-1110. |  | 1523 | 1523 units | Nationwide Distribution | Terminated | July 12, 2012 | 142 | 1523
Z-1228-2012 | 61268 | No | DSQ |  |  | Cardiovascular | PMA | HeartMate II Left Ventricular Assist System (HM II LVAS) Sealed Outflow Graft Bend Relief. |  |  | February 23, 2012 | 2012 | Thoratec Corporation | A trend of disconnected bend reliefs on the sealed outflow graft used with the HeartMate II LVAS has been identified | Device Design | Thoractec Corporation sent an Urgent Medical Device Correction letter on February 23, 2012 via overnight courier and e-mail to all affected consignees. The Urgent Medical Device Correction notification letters identify the affected products, description of the problem with illustrations, symptoms of problem and immediate actions to taken. All users are asked to review the attached labeling, including new caution statements and share with all surgical personnel. Consignees are instructed to completed and returned the Acknowledgement Form indicating that the information was received and understood. Additionally, an update concerning the Worldwide Medical Device Correction Notification is enclosed with each letter.Questions should be directed to Thoratec Regulatory Affairs at 925-847-8571. |  | 3852 | 3852 units in commerce | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Belgium, Denmark, Germany, Greece, France, Hong Kong, Kuwait, Ireland, Israel, Italy, Lebanon, Liechtenstein, Lithuania, Luxembourg, Malaysia, Netherlands, Turkey, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Thailand and United Kingdom. | Terminated | November 08, 2012 | 259 | 3852
Z-1283-2012 | 61365 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Easy Diagnost Eleva DRF, Software release 4.0 |  |  | March 09, 2012 | 2012 | Philips Medical Systems Gmbh, DMC | Underexposed image due to lag in switching betweenauxiliaries | Software design | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated February 27, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The firm provided a work around for customers to use until a software update can be installed. Customers should contact their local Philips representative for questions regarding this notice. |  | 88 | 88 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Germany, Indonesia, Malaysia, Saudi-Arabia, Singapore, and Thailand. | Terminated | May 23, 2013 | 440 | 88
Z-1282-2012 | 61305 | No | JSC | Microbiological assay culture medium. | Microbiology | Microbiology | 510(K) Exempt | Fildes Enrichment, BBL brand, catalog #211866, in 5 ml tubes,10 tubes/package, manufactured by BD Diagnostic Systems, 250 Schilling Circle, Cockeysville, MD 21030. |  |  | February 20, 2012 | 2012 | Becton Dickinson &amp; Co. | In vitro diagnostic reagent is defective and may fail to enhance the growth of microbiological cultures as intended. | Nonconforming Material/Component | Becton Dickinson (BD) sent an "URGENT PRODUCT RECALL" letter dated February 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard any remaining affected product. BD will issue replacements for the discarded product. Additionally, an Acknowledgment Form was attached for customers to complete and return to the firm via fax to 410-316-4258. Contact BD Technical Services Department for questions regarding this notice. |  | 6930 | 6930 tubes | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, China, Japan, Mexico, and Taiwan. | Terminated | June 18, 2012 | 119 | 6930
Z-1315-2012 | 61282 | [6 empty] | QRS Diagnostic Universal ECG 12 Channel ECG (IEC), Rx only, REF Z-7000-0400, Manufactured for QRS Diagnostic, 14755 27th Ave North, Plymouth, MN 55447 USA. |  |  | February 16, 2012 | 2012 | Qrs Diagnostic | QRS Diagnostic has initiated a recall involving Universal ECG 12 Lead IEC ECG Devices. This recall has been initiated because the chest lead wires on these devices may have been improperly labeled. Incorrect labeling of the lead wires could result in improper lead wire placement on the patient which may result in inaccurate waveforms being displayed to the user and an inaccurate interpretatio | Process control | QRS Diagnostic sent an Urgent Medical Device Recall letter dated February 9, 2012 via e-mail and mail to all affected consignees. The letter identified the affected product, description of the problem, and actions to be taken. If the affected product have been further distributed, consignees are instructed to notify them at once of the product recall. Consignees were instructed to contact Uni-com immediately for a replacement device. Uni-com will provide this replacement at no charge and will arrange to have the device returned to QRS Diagnostic for correction. Additionally, consignees were instructed to complete the attached Medical Device Correction Receipt Confirmation Form and return to QRS Diagnostic via fax at country code =1-763-559-2961 or email support@QRSDiagnostic.com. For questions contact Technical Support at support@QRSDiagnostic.com or by phone 800-465-8408. |  | 13 | 13 devices | Worldwide Distribution and the countries of: Belgium, France, Italy, Mali and Senegal. | Terminated | August 02, 2012 | 168 | 13
Z-1285-2012 | 61363 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Whole body Computed Tomography X-Ray System (Brilliance 64 and Ingenuity CT), Brilliance 64 M/N 728231, Ingenuity CT M/N 728326, Dist. By Philips Medical Systems, Cleveland, OH |  |  | February 17, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances were identified during internal and external validation testing of the Brilliance 64 and Ingenuity CT systems with software version Beta 3 (version 4.0.0.xx368). | Software change control | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 17, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all memebers of their staff who need to be aware of the contents in the communication. Customers were asked to retain a copy of the communication with the equipment instruction for Use. For further information or support concerning this issue, customers should contact their local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center at 1-800-722-9377, option 5.For questions regarding this recall call 440-483-7000. |  | 7 | 7 units | Worldwide Distribution - USA including OH and the following countries: Australia, China, Belgium, Israel &amp; Sweden. | Terminated | February 22, 2013 | 371 | 7
Z-1286-2012 | 61207 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Product is labeled in part: "***REF: 01-7036, LOT 999999, "1.5MM" SYSTEM PLATE L-SHAPE, 3 X 3 HOLE, 100 DEGREE, QTY 00001, CE 0086, 2010-04 *** NON-STERILE PRODUCT ***, LEFT, X-LONG, 0.6 MM TITANIUM, BIOMET MICROFIXATION, 1540 TRADEPORT DRIVE, JACKSONVILLE, FL 32218 USA, 01-736, LOT 999999, QTY 00001, L-SHAPE, 3 X 3 HOLE, 100 DEGREE, LEFT, X-LONG, 0.6 MM, CE 0086.***" |  |  | April 26, 2010 | 2010 | Biomet Microfixation, LLC | On 4/26/10 Biomet Microfixation, Jacksonville, FL initiated a recall of their 1.5 3/3 hole 100 degree left X-long L plate, Model #01-7036, Lot #193690. Product is labeled incorrectly. | Error in labeling | Biomet Microfixation, LLC sent an Urgent Medical Device Safety Alert letter dated April 21, 2010 to all consignees affected. Each consignee was notified by E-mail, letter or telephone call the week of April 26, 2010. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to fill out the attached Inventory Reconciliation sheet and FAX it to the attention of the Global Complaint Coordinator at 904-741-9425. If the affected product have been further distribute, consignees are instructed to notify the recipient of this action by providing the recipient a copy of this notice. This letter must be given to the personnel responsible for receiving field action notices. For questions contact Christy Cain at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 23 | 23 | USA Nationwide Distribution including the states of: California and Texas. | Terminated | April 18, 2012 | 723 | 23
Z-1318-2012 | 60647 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Anesthesia Machines (digital display assembly); |  |  | November 29, 2011 | 2011 | Penlon, Ltd. | Penlon, Ltd has identified the potential for the premature failure of a component in some Electronic Flowmeter sensor display unit PCB assemblies. | Component design/selection | Penlon Limited sent a Technical Bulletin dated November 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. As a precaution a field upgrade part was released that required all existing PCB assemblies to be removed and the updated item installed. The old PCB must be returned to Penlon with the serial number of the affected machine. The required replacement parts were included with the technical bulletin.For further information customers should contact the Technical Support Team at techsupport@penlon.com. |  | 160 | 160 units | Worldwide Distribution - USA (nationwide) and the countries of Peru, Uganda, Kenya, Malaysia, Poland, Saudi Arabia, United Arab Emirates, Philippines, Equador, and Thailand | Terminated | July 03, 2012 | 217 | 160
Z-1317-2012 | 61269 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Version 2.0.0.15, Radiation Therapy Treatment Planning System. |  |  | December 06, 2011 | 2011 | RAYSEARCH LABORATORIES AB | An issue was found with the dose calculation in the RayStation 2.0, that could result in a potential overestimation of the dose. This could lead to the potential under-dosage of a patient. | Device Design | The firm decided to recall and sent Urgent Field Safety Notice Medical Device Correction letters to consignees on 12/06/2011. The letters described the issue, identified the affected product, actions to be taken by the user, a solution, and contact information. Users are to follow the two part workaround, and educate their staff and all users on the workarounds. The letter states that the issue will be resolved in the next version of RayStation. If further information is needed, users are to contact either Freddie Cardel, Director of Customer Support, RaySearch Americas, at 1-877-778-3849 or freddie.cardel@raysearchlabs.com; or Niclas Borglund, Director of Service, RaySearch Laboratories AB, at 46-8-5450-6130 or niclas.borglund@raysearchlabs.com |  | 21 | 21 units (16 domestically and 5 internationally) | Worldwide Distribution -- US, including the states of FL, MA, NJ, and NY and the country of the Netherlands. | Terminated | May 11, 2012 | 157 | 21
Z-1316-2012 | 61222 | No | LBS | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products --- dHDL Slides --- Ortho Clinical Diagnostics, a Johnson&Johnson; company --- [Mfg Symbol] Ortho-Clinical Diagnostics, Inc., Rochester, NY USA --- VITROS dHDL Slides are used to quantitatively measure HDL cholesterol (HDLC) concentration in serum and plasma using VITROS Chemistry and Integrated Systems. --- Product codes: 6801895 (60 slide count); 6802469 (18 slide count) --- indicated to quantitatively measure HDL cholesterol (HDLC) concentration in serum and plasma using VITROS Chemistry and Integrated Systems. |  |  | February 02, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) has revised the VITROS dHDL Slides Instructions for Use (IFU) and removed EDTA plasma as a recommended sample type. | Device Design | Ortho Clinical Diagnostics sent an Important Product Information letter dated February 2, 2012, to all affected customers via Federal Express overnight mail. The letter informed customers of the issue, notifying them of the removal of EDTA plasma as a recommended sample type and providing the revised VITROSÂż dHDL Slides Instructions for Use. J&J; Foreign affiliate consignees were notified by email informing them of the issue on 02 February 2012. Customers were instructed to replace the current pages in their VITROS Chemistry Products Instructions for Use Manual with the updated instructions for Use and Table of Contents.Customers were asked to return the enclosed Confirmation of Receipt form regardless of whether their laboratory uses the affected product. For any questions regarding this recall call 1-800-421-3311. |  | 1366236.9 | 120,225 of 60-slide count, 11.364 of 18-slide count | Worldwide Distribution | Open, Classified |  |  | 1366236.9
Z-1322-2012 | 61170 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BBL (tm) Lowenstein-Jensen Medium Deeps, Catalog #221256, packaged in boxes of ten/600 mL tubes. Product Usage: Lowenstein-Jensen Medium is used for the isolation and cultivation of mycobacteria. The medium tubed as deeps is used for the semi-quantitative catalase test as an aid to the classification of mycobacteria. |  |  | January 30, 2012 | 2012 | Becton Dickinson &amp; Co. | Media fails to perform as intended with quality control organism Mycobacterium kansasii. | Nonconforming Material/Component | The firm, BD Diagnostic Systems, sent an "URGENT PRODUCT RECALL" letter dated January 2012 to Ex-US BD sites via email and its customers via UPS ground shipment on 1/30/12. The letter identified the product, the problem and the action to be taken. The customers were instructed to discontinue use of the recalled lot numbers and discard any remaining packages. Customers were also instructed to complete and return the attached form whether or not they had any inventory remaining via fax to: BD Regulatory Compliance at 410-316-4285. The letter also informed the customers that BD has decided to discontinue this product.If further assistance is needed regarding credit, please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. |  | 3900 | 3900 tubes | Worldwide Distribution - USA (nationwide) including states of: AL, CA, CT, IA, IL, KY, LA, MD, MI, MN, NC, NJ, NY, OH, PA, TN, VA, WA, and WV; and countries of : Belgium and Singapore. | Terminated | April 02, 2012 | 63 | 3900
Z-1323-2012 | 61271 | No | IRQ | Powered heating unit. | Physical Medicine | Physical Medicine | 510(K) Exempt | The LactoSorb Heat/Contouring Pen is a disposable alkaline-battery powered device. |  |  | November 15, 2011 | 2011 | Biomet Microfixation, LLC | On 11/15/2011 Biomet Microfixation, Jacksonville, FL initiated a recall of their Acticort Heat/Contouring Pen, Model #915-3100, Lot #861200. Product is incorrectly labeled | Error in labeling | Biomet Microfixation, LLC sent an Urgent Medical Device Safety Alert letter dated November 15, 2011 via E-mail, letter or telephone call to all consignees affected. The letter identified the affected product, problem and actions to be taken. Each consignee was Instructioned to take the following actions: 1) If you have any unopened boxes of the product in question, please place this notice with it; 2) Please fill out the attached Inventory Reconciliation sheet and FAX to the attention of the Global Complaint Coordinator; and 3) If you have further distributed this product , you must notify the recipient of this action by providing the recipient a copy of this notice. This letter must be given to the personnel responsible for receiving field action notices. For questions please contact Christy Cain at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 42 | 42 | Worldwide Distribution - USA (nationwide) including the state of Texas and the country of Japan. | Terminated | August 04, 2014 | 993 | 42
Z-1321-2012 | 59573 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Base, 220/240V, Catalog 801764, 6 " Roller Pump Occlusion: Reference AA-2011-015 C. indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures. |  |  | March 16, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo CVS has received reports of 27 instances in which the 6 inch Roller Pump jammed during cardioplegia delivery. There has been one report of the pump jamming during a case that concluded with an adverse patient outcome, although there was no indication that the pump jam contributed to the adverse patient outcome. | Device Design | "URGENT MEDICAL DEVICE RECALL" letters were sent on 3/16/12 to all consignees detailing the correction as well as also adding an Addendum to the Operator's manual for the Terumo System 1 to incorporate additional information and warnings for use when operating a 6" roller pump with dual-sized tubing. For questions call 1-800-521-2818. |  | 648 | 648 units | Worldwide Distribution--USA, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, Puerto Rico and the countries of INDIA, CANADA, CHILE, HONG KONG, EGYPT, PAKISTAN, KUWAIT, SAUDI ARABIA, JORDAN, UNITED ARAB EMIRATES, SOUTH AFRICA, SOUTH KOREA, PHILIPPINES, JAPAN, PHILIPPINES, SINGAPORE, AUSTRALIA, MALAYSIA, TAIWAN, RUSSIA, TURKEY, BELGIUM, GERMANY, COLUMBIA, COSTA RICA, ARGENTINA, HONDURAS, COLUMBIA, DOMINICAN REPUBLIC, GUATEMALA, MEXICO, VIETNAM, and THAILAND. | Terminated | May 02, 2013 | 412 | 1642
Z-1178-2012 | 61311 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare EasyDiagnost Eleva DRF 4.0 |  |  | April 01, 2011 | 2011 | Philips Healthcare Inc. | It was discovered in some instances the detector did not reset as quickly as the user prompted a change form one image receptor (auxiliary) to the next, and then back again. | Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter, dated February 28, 2012, was sent to customers. The letter identified the affected product, a description of the issue, the hazard involved, along with the actions to be taken and planned. Customers were to use one of the three work-a-rounds provided. The issue will be corrected with Field Change Order 70600064, which consists of a Software Update 4.0.2SP Service Pack installed on the system. If further information or support is needed, customers can contact Dominic Siewko at 978-659-7936. |  | 96 | 96 | Worldwide Distribution - Including USA. | Completed |  |  | 96
Z-1324-2012 | 61177 | No | MQL | Reproductive media and supplements. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sydney IVF Culture Oil, REF, K- SICO- 50, 50 mL. and K-SICO-200, 200 mL,Store at 2-8 C, Sterile, Cook Medical, William A Cook Australia LTD, |  |  | February 15, 2012 | 2012 | Cook, Inc. | Cook Medical is recalling Sydney IVF Culture Oil due to a number of recent customer complaints that have identified it as a possible cause of embryo degeneration at day 5/6. | Process control | Cook Medical sent an Urgent Recall Notices letter dated February 15, 2012 to all affected customers. The letter identified the recall product, reason for the recall and actions to be taken. The notice included directions to pull the recalled product from stock, isolate it and forward the notice to any sub-accounts. Customers were requested to complete and return the recall acknowlegement and product tracking form by to Fax: 812-829-1801 or Email: julie.cooper@cookmedical.com. For questions call 800-457-4448 x7241. |  | 54 | 54 total from both lots | Worldwide Distribution - USA (nationwide) including the states of IL, KS, KY, MA, NJ. TN, TX, UT, VA, WI, WV, and the country of Barbados. | Terminated | July 03, 2012 | 139 | 54
Z-1319-2012 | 61377 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific Soloist Single Needle Electrode: 18 cm long, 16 gauge, Sterile, packaged with the Soloist Electrode Power Cord. Boston Scientific, Nalick, MA. |  |  | March 02, 2012 | 2012 | Boston Scientific Corporation | Device found labeled with the incorrect expiration dates. The Soloist Single Needle Electrode has been qualified for a 3 year expiration date. Product manufactured between the dates of May 24, 2011 and November 8, 2011 has an incorrect expiration date of 20 years labeled on the pouched needle electrode and 5 years labeled on the carton. | Incorrect or no expiration date | Boston Scientific sent a Urgent Medical Device Recall letter dated March 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed to cease use, quarantine and return unused product with directions to forward the notification to the hospital level. Customers were directed to fill out the Reply Verification Tracking Form and fax to 508Âż683Âż5578, for reimbursement. Returned goods are to be sent to:Boston Scientific CorporationDistribution CenterAttn: QA Returns500 Commander Shea BlvdQuincy, MA 02171We regret any inconvenience that this action may cause and we appreciate your understanding as we take action to ensure patient safety and customer satisfaction. For further questions please call (508) 683-4678. |  | 682 | 682 (1555 used) | USA (nationwide) and the countries of Canada, Columbia, Viet Nam, Germany, France, Great Britain, Italy, Jordan, Kuwait, Netherlands Pakistan, Portugal, S. Arabia, Slovakia, Spain, Sweden and Tunisia. | Terminated | November 14, 2012 | 257 | 682
Z-1224-2012 | 60991 | No | RBS |  |  |  |  | Hamamatsu Corporation, Bridgewater NJ 08807 |  |  | March 30, 2011 | 2011 | Hamamatsu Corporation | It was discovered that the Micro Focus X-ray Source could be experiencing software related problems that would allow the unit to exceed the maximum power setting that was set by the operator. | Software design | Hamamtsu Corporation sent a Warning letter dated December 22, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. CDRH approves your CAP under the following conditions:1.Hamamatsu is able to work with your manufacturer customers to assure the revised manual and new label are provided to your customers' end users.2.Hamamatsu informs CDRH if you receive information which indicates this control defect has occurred. Relevant information includes direct communication from customers or through observation of excessive wear on returned x-ray tubes.For questions regarding this recall call 908-231-0960, ext 2300. |  | 100 | 100 | Worldwide Distribution-- USA (nationwide) and the country of Canada. | Open, Classified |  |  | 100
Z-1212-2012 | 61028 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare MD Eleva system. |  |  | November 19, 2011 | 2011 | Philips Healthcare Inc. | It was discovered there was a software error where the air karma rate is not displayed correctly on the user interface. | Software design | A field correction action will be undertaken to correct the software error in affected systems worldwide. Your proposed action to bring the product into compliance with regulatory requirements is to institute a worldwide field correction to replace the software error in affected systems by updating the current software to version R6.1.2.The timeframe for completion of this field action will be 6 months post field action publication. Preventive actions: The software has been corrected moving forward in subsequent releases for these products.For further questions please call (978) 787-1501. |  | 92 | 92 ( 24 Eleva 11s, and 68 Eleva FDs ) | Worldwide Distribution - USA (nationwide) | Open, Classified |  |  | 92
Z-1334-2012 | 61329 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oridion CO2 sampling lines and water traps under the label: |  |  | February 16, 2012 | 2012 | Oridion Medical 1987 Ltd. | CO2 sampling lines may have paper and plastic particulate present in the packaged units at levels that may exceed visual specification. | Packaging process control | Oridion notified Hospitals, Clinics, EMS with an "Important Communication" letter to clinicians dated February 12, 2012. The letter was intended to make US clinicians aware of an important issue relating to certain Oridion CO2 sampling lines, which may not fully meet quality requirements. In a few cases, paper and plastic particulate may be present in the packaged units at levels that exceed the visual specification. Recommended actions were included in the letter. Questions were directed to 1(781)-972-1252. |  | 2000000 | 2 million | Nationwide Distribution | Terminated | September 13, 2013 | 575 | 2000000
Z-1331-2012 | 61276 | preservative, cytological | LEA | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Surgipath Cyto Jar Fixative; a 120 mL clear, leakproof container prefilled with 45 mL of preservative containing polyethylene glycol for the convenient collection, fixation and preservation of all cytologic specimens, 50 jars per case; for in vitro diagnostic use only; Leica Biosystems, 5205 Rt. 12, Richmond, IL 60071; Made in USA |  |  | February 20, 2012 | 2012 | Leica Biosystems Richmond Inc. | Product was distributed with the incorrect lot number and expiration date. The lot number and expiration date should have been listed at 073013 and instead was noted as 073014 which is beyond the maximum expiration dating of 18 months for this product. | Process design | The firm, Leica Microsystems, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 17, 2012, via FedEx on February 20, 2012 to all affected customers. The letter describes the product, problem, and actions to be taken. A second letter was sent to the customers on February 29, 2012, clarifying that the incorrect lot number and expiration date was 073014, and should have been 073013. The customers were instructed to discontinue use of the product; return the product and complete and return the enclosed Acknowledgement Form via fax to 815-678-2216 within one week of receiving the notification.If you have any questions, contact the Regulatory Affairs Manager at toll free (800) 225-8867 or email: richmond@leica-microsystems.com. |  | 11 | 11 cases | Nationwide distribution: USA including states of: Michigan, Missouri and New York. | Terminated | September 20, 2012 | 213 | 11
Z-1330-2012 | 61375 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Sorin C5 (perfusion system) |  |  | February 23, 2012 | 2012 | Sorin Group Deutschland GmbH | The pump speed control knob is attached to the shaft of an encoder that is used to adjust the speed. Insufficient lubrication within the encoder increases friction over time making it more difficult to turn the pump speed control knob. Excessive friction causes premature wear of the encoder shaft which can result in a stuck shaft. | Nonconforming Material/Component | Sorin Group sent an Urgent Field Safety Notice dated February 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:Before each use of the device during the set-up procedure, rotate all pump speed control knobs back and forth to check for smooth operation. If abnormal friction or any binding is detected, please do not use the system until it is repaired.If an affected product is currently in use with a patient and an increase of friction is detected please prepare for a replacement pump.Please complete and return the attached Confirmation Form by fax to 303-467-6502.Customers were instructed to ensure that the Urgent Field Safety Notice is distributed to all personnel in their organization who need to be aware of the notice. If the product was transferred to a third party, customers were asked to share the information with them as well as Sorin Group USA Customer Service at 1-800-850-2623.For questions regarding this Urgent Field Safety Notice, contact Sorin Group USA Customers Service at 1-800-650-2623.All affected customers and distributors were notified of the Field Correction by certified mail beginning on February 23, 2012. They will be contacted by Sorin Service Group to arrange an appointment to replace affected shaft encoders on site. |  | 39 | 39 | Worldwide Distribution | Terminated | November 27, 2012 | 278 | 1147
Z-1188-2012 | 61194 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Aerolase Corporation. The products involved are the FriendlyLight YAG lasers, a Model Family first identified and reported. |  |  | June 07, 2011 | 2011 | Aerolase | It was discovered that customers were not supplied with a procedure and schedule for calibrating the FriendlyLight YAG lasers. | Radiation Control for Health and Safety Act | Aerolase issued a letter to customers, dated 11/30/2011, identifying the affected product and the reason for the recall. Aerolase is providing, at no charge, the Appendix to the Operating Manual for the LightPod lasers. This Appendix describes the procedure and frequency for checking whether the laser is within calibration, but it is not a procedure for customers to actually calibrate the laser themselves; any attempt to do so by any parties other than a company technician employed by Aerolase would void the warranty. Please contact us any time at 914- 345-8300 with any questions. |  | 176 | 176 | Worldwide Distribution | Open, Classified |  |  | 176
Z-1355-2012 | 61054 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Critical Care Bed Model 2031, 230 V , 50/60 Hz, Max Load = 225 Kg,, Stryker, Portage, MI |  |  | March 12, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Stryker Medical has determined that units manufactured between September 1, 2004 through November30, 2007, did not receive updated preventive maintenance information. (Failure to maintain the bed brakes including the brake plate kits, may result in reduced brake holding force. Reduced braking force may result in a patient fall or caregiver muscle strain, either of which may require medical int | Component change control | Stryker Medical sent an Urgent-Medical Device Correction letter dated March 12, 2012 to all consignees. The letter identified the affected products, problem, updated preventive maintenance information and requested the following immediate actions: 1) Locate the beds listed in this notice, 2) Ensure that the brakes are holding properly, 3) If the brakes are not holding, contact Stryker (1-800-327-0770, option 4) to order the maintenance parts, 3) Continue to periodically inspect brake functionality, as noted on enclosed preventive maintenance schedule/checklist, 4) Order replacement brake components, as needed, 5) Return the enclosed post card to confirm receipt of this notification, 6) If you have loaned or sold any of the beds listed in this letter, please, forward a copy of this notice to the new users and advise us of their new location, 6) If you have disposed of any of the beds and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers. For questions or concerns, please, contact Sandy Cliche at 269-389-6025. |  | 507 | 507 | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of: DEDEDO GU, HAMILTON ON, SANTIAGO, SHATIN, NT , CAPITOL FEDERAL, ITAPEVI SP, ATHENS, OSAKA, SEOUL, MEXICO DF MX, MONTREUX, BALERNA, NORTH POINT, WARSZAWA, CANOVANAS, KOWLOON BAY, SINGAPORE, NEWBURY, and JOHANNESBURG. | Terminated | August 29, 2013 | 535 | 121982
Z-1349-2012 | 61143 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oridion Surestream CO2 Sampling Lines- |  |  | December 01, 2011 | 2011 | Oridion Medical 1987 Ltd. | Surestream CO2 Sampling Lines-Connector may cause breakage of the microswitch inside the CO2 inlet | Device Design | Oridion issued a Technical Note communication in March 2011 to field staff and end users globally. The Technical Note identified the affected product and provided instructions for user on the proper insertion of Surestream products. For questions contact your local distributor or Oridion at technicalsupport@oridion.com |  | 481 | 481 boxes (12025 untis) | USA Nationwide Distribution including the states of: AZ, CA, FL, HI, LA, MA, MO, NJ, NY, OH, OR, TX and WI. | Terminated | June 04, 2012 | 186 | 841
Z-1336-2012 | 61362 | No | GIM | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BacT/ALERT Blood Collection Adapter |  |  | March 05, 2012 | 2012 | bioMerieux, Inc. | The BacT/ALERT Blood Collection Adapter Cap is no longer compatible with Becton Dickinson (BK) Vacutainer Safety Lok Blood Collection Set. The incompatibility is the blood collection set needle backing out of the adapter cap that results in an exposed sheathed needle that may lead to an accidental needlestick. | Nonconforming Material/Component | bioMerieux sent an Urgent Product Correction Notice letter dated March 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions: 1. Ensure the letter and guidance on how to maintain control of the luer connector by securing it between the thumb and forefinger are distributed to all appropriate personnel within their organization. 2. Fill out and return the Acknowledgement Form by Fax to confirm receipt of the recall notice. For additional assistance or questions customers were instructed to contact their local bioMerieux Customer Service Representative at (800) 682-2666, option 3, then option 3 again. Industry customers should call 800- 634-7656, option 3. For questions regarding this recall call 314-731-8537. |  | 274800 | 274,800 boxes (32,862,600 caps) | Worldwide Distribution - USA (nationwide), Australia, Brazil, Canada, Chile, Columbia, Costa Rica, France, China, Indonesia, Korea, Mexico, Peru, Singapore, Ecuador, Dominican Republic, Philippines, Guatemala, Curacao, Thailand, United Arab Emirates, Andorra, Austria, Burkina Faso, Congo, Switzerland, Cameroon, Germany, Algeria, Estonia, Spain, France, Great Britain, French Guiana, Croatia, Israel, India, Italy, Kazakhstan, Lithuania, Latvia, Macedonia, Malta, New Caledonia, Netherlands, French Polynesia, Poland, Portugal, Reunion, Russia, Sweden, Slovenia, Chad, Turkey, South Africa, Saudi Arabia, Congo, CoteDIvoire, Denmark, Djibouti, Finland, Greece, Guadeloupe, Libyan Arab Jamahiriya, Jordan, Morocco, Martinique, Norway, Senegal, Togo, Yemen, Bangladesh, Georgia, Ghana, Uganda, and Qatar. | Terminated | September 04, 2014 | 913 | 274800
Z-1337-2012 | 61367 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | IntelliSpace Portal, M/N 881001, IntelliSpace Portal (ISP) is a multimodality (CT, NM, and MR) thin- client applications server that delivers full diagnostic viewing and clinical applications to the enterprise. IntelliSpace Portal (ISP) is a medical software system that allows multiple users to remotely access IntelliSpace Workspace Portal (ISP) from compatible computers on a network. |  |  | February 22, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances were logged for IntelliSpace Portal running software version 4.0 during in-house random verification testing. When loading a study to Review, after failing to load it through Add to Running Application the Analysis application is not being closed and different (1st loaded study) patient appears in analysis (CCA). | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. Customers were asked to retain a copy with the equipment instruction for use. For any further information or support customers should contact their local Philips representative or local Philips Healthcare office. Customers in North America and Canada should contact the Customer Care Solutions Center at 1-800-722-9377, option 5.For questions regarding this recall call 440-483-7000. |  | 90 | 90 units | Worldwide Distribution - USA including AL, AR, FL, GA, IO, IN, MI, NJ, NY, OH, OK, OR, PA, SC, TN, TX &amp; WA and the countries of Australia, Canada, China, Czech Republic, France, Germany, India, Indonesia, Israel, Italy, Poland, Spain, Sweden, Switzerland &amp; United Kingdom | Terminated | February 22, 2013 | 366 | 90
Z-1340-2012 | 61373 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | RTVue CAM with Corneal Power Measurement, Cornea Power Kit (software) |  |  | February 21, 2012 | 2012 | Optovue Inc. | A few doctors in the US were inadvertently sent the Cornea Power software version kit for international users instead of the domestic version. | Employee error | Field Correction letters were sent to all doctors affected by the issue on February 23, 2012. The letter identified the affected product and the problem. The letter also stated that customers may continue to use the RTVue until the correction was made. Once the correction has been made, customers are to sign and return the attached acknowledgement form. If customers had questions, they are to contact Optovue at 1-866-344-8948 or via mail at 45531 Northport Loop W, Fremont, CA 94538. |  | 8 | 8 | Nationwide Distribution -- CA and WA only | Terminated | April 02, 2012 | 41 | 8
Z-1338-2012 | 61391 | No | LKK |  |  | General Hospital | PMA | Medtronic, Model 8870, Application Software Card is part of the SynchroMedÂż II Drug Infusion System. The System includes the SynchroMedÂż II implantable drug infusion pump 8637-20, 8637-40, N'Vision clinical programmer 8840 and the Software application card 8870. |  |  | March 12, 2012 | 2012 | Medtronic Neuromodulation | Medtronic has confirmed that an algorithm used in the Model 8870 application card software has resultedin nine (9) occurrences of an incorrectly displayed Schedule to replace the pump by date. It is estimated that there are more than 140,000 SynchroMed II pump implants worldwide. A patient with a pump reaching EOS prior to replacement may experience the return of underlying symptoms and/orwit | Software design | Medtronic mailed an "Urgent Medical Device Correction " letter addressed to Dear Healthcare Professional on March 12, 2012. The letter described the problem, Background, Clinical Manifestations, Recommendations, and Additional Information.To ensure there is no interruption in therapy, pump replacements should be scheduled before the ERI alarm sounds and the pump reaches its 7-year end of service.Customers were advised to forward the Urgent Medical Device Correction letter to patients whose name appears on their Pump Replacement List.Reimbursement SupportFor assistance with prior authorization, contact Medtronic Reimbursement at the numbers below:Therapy for Severe Spasticity: (800) 292-2903Pain Therapies: (866) 962-9909For questions regarding the letter call 800-328-0810. |  | 6338 | 6338 | Nationwide Distribution | Open, Classified |  |  | 6338
Z-1339-2012 | 61293 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Optiflux F160NRe Hemodialyzer |  |  | March 06, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Risk of internal blood leaks affecting the hemodialyzers. | Process control | Fresenius North America notified consignees on 3/6/12 by recall notification letter titled: Urgent FMCNA Optiflux F160NRe Hemodialyzer Recall, by certified mail with return receipt. The letter identified the affected product and described the reason for the recall. Customers were instructed to examine their stock immediately to determine whether they have any of the affected lot on hand. If customers have the affected lot, they have been instructed to immediately discontinue use and place all unused units in a secure area for return to Fresenius Medical Care North America. A Reply Response from is to be completed and returned. Customers have been instructed to contact their Fresenius Medical Care Customer Service Team for instructions on how to return the recalled product. Contact 1-800-323-5188. |  | 30576 | 30,576 units | Nationwide Distribution | Completed |  |  | 30576
Z-1351-2012 | 61440 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Straumann Standard Plus Implant 04.8mm Wide Neck |  |  | March 12, 2012 | 2012 | Straumann USA, LLC | Mislabeled: Vial lid label on some of the dental implant vials from this lot incorrectly state 04.8 SP 10 RN instead of 04.8 SP 10 WN | Packaging process control | Straumann sent a "FIELD CORRECTIVE ACTION NOTICE" dated March 12, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product and to return any affected product to the firm. Additionally, a Customer Confirmation Form was enclosed for customers to complete and return. Contact the firm at 978-747-2514 for questions regarding this notice. |  | 346 | 346 units | Nationwide Distribution. | Terminated | May 31, 2012 | 80 | 346
Z-1358-2012 | 61295 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex eXpert DC Intraoral Dental X-Ray System |  |  | December 29, 2011 | 2011 | Gendex Dental Systems | Improper hardware was used during installation of the eXpert DC Intraoral X-Ray system units causing drifting of the tubehead assembly. | Nonconforming Material/Component | Gendex Dental Systems sent a Urgent Medical Device Correction letter dated December 29, 2011, to both affected customers. The Customers were provided background information and details on how the nonconforming installations of the devices would be corrected. Gendex Dental Systems apologizes for any inconvenience caused by this issue, and request your corporation as we correct this problem. The customers were instructed to complete and return the Acknowledgement Form upon receipt of the recall notification letter and fax to (215) 997-5665.For further questions please call (215) 954-0365. |  | 52 | 52 | USA including the states of Iowa and Illinois. | Terminated | June 01, 2012 | 155 | 52
Z-1359-2012 | 61272 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OneTouch Verio IQ Blood Glucose Monitoring System Kit; |  |  | February 21, 2012 | 2012 | Lifescan Inc | An error was discovered in the OneTouch IQ meter software so that the meter turns itself off when a user attempts to view results in the "Results Log" when the log has 256 or a multiple of 256 items to display. | Software design | Lifescan sent an Urgent Medical Device Voluntary Recall letter dated February 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to 1. Identify and sequester all lots of OneTouchÂż Verio"IQ Meters they have in inventory.UPC No. 353885007702 LifeScan Part No. 022267 NDC No. 53885026701Note: Only OneTouchÂż Verio"IQ Meters are subject to this notice. Pleasecontinue to sell all other OneTouchÂż Brand products including OneTouchÂż Verio" TestStrips and all OneTouchÂż UltraÂż Brand Meters, Test Strips and Ancillary Products.2. Communicate this replacement program to their distribution centers and retail stores that have received OneTouchÂż Verio"IQ Meters. Request that they return only OneTouchÂż Verio"IQ Meters per their normal return procedures.3. Once you received all of the OneTouchÂż Verio"IQ Meters to be returned, call Inmar Corp. at 336-631-7744 for a Returned Goods Authorization (RGA) number and product return instructions.When product has been returned, a credit memo for an amount equivalent to the invoiced price of the returned product. No deductions will be allowed.If you have any questions about this action, please contact your LifeScan Sales Representative or call the LifeScan Customer Experience Line at: 1-800-600-7226, Option 1. Customers were advised that they can continue use of their affected product until an updated meter is provided free of charge.For questions regarding this recall call 408-942-5937. |  | 101882 | 101,841 units. 2-27-12 - this figure is an estimate and will be updated when more accurate number is received. | Worldwide Distribution - USA (nationwide) and Canada | Terminated | February 05, 2013 | 350 | 101882
Z-1360-2012 | 61234 | No | ITH | Power traction equipment. | Physical Medicine | Physical Medicine | 510(k) | Removal of Literature/Brochures for Antalgic-Trak. "***What is Antalgic-Trak?***Articulating Non-Surgical Spinal Decompression****Antalgic-Trak By Spinetronics***Note: "Decompression" as used herein refers to unloading due to distraction and positioning.***Antalgic-Trak is an improved form of non-surgical spinal decompression*. |  |  | August 11, 2011 | 2011 | Spinetronics, LLC | On 08/11/2011 Spinetronics, LLC initiated a recall of Antalgic-Trak Product Brochures which contained language and/or claims that were not cleared under the original 510k, and represented a major change or modification in the intended use of the device. | Labeling False and Misleading | Customer notification letters were mailed by certified mail via USPS. The medical user/customer was instructed to either return or destroy any remaining recalled product literature. Customers were asked to complete and return a response form noting the number of available brochures. |  | 1600 | 1600 brochures | Worldwide Distribution, including Nationwide (USA) and the country of Canada. | Terminated | April 20, 2012 | 253 | 1600
Z-1364-2012 | 61488 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Verio |  |  | March 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During product monitoring, Siemens has discovered that the gradient cable connections on a few systems did not meet their specifications after installation. The cable connections overheated, leading to the emission of smoke and possible material damage. | Nonconforming Material/Component | The firm, Siemens Medical Solutions USA, Inc., sent a "Customer Safety Advisory Notice" including Update Instruction MR028/10/S dated March 2012 to affected customers via hand delivery by Siemens Service Engineer . The notice described the product, problem and actions to be taken. The customers were instructed to create a COBR image and to make a copy of the attached "Completion Protocol" fill it out and file in their corresponding binder.Siemens Service Engineer performed an all gradient cable connection visual inspection as well as torque check on the connecting screws at the sites and the factory prior to their release to the field. Siemens also issued a current Update Instruction MR024/11/S. If you have any questions regarding this report, please contact the Quality Engineer, Product Complaints &amp; Field Support at 610-219-4834; email: anastasia.mason@siemens.com or Director, Product Complaints &amp; Field Support at 610-448-4634 or email: josephine.mcbride@siemens.com. |  | 245 | 245 | Nationwide distribution to Washington, DC and the following States: AL, AK, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, and WI. Government Accounts include: BROOKE ARMY MEDICAL CENTER, 3851 ROGER BROOKE DR, FORT SAM HOUSTON, TX 78234; NATIONAL INSTITUTE OF HEALTH, 10 CENTER DRIVE, BETHESDA, MD 20892; NAVAL MEDICAL CENTER, 34800 BOB WILSON DR, SAN DIEGO, CA 92134; US AIR FORCE REGIONAL HOSPITAL, 5955 ZEAMER AVE, ANCHORAGE, AK 99506; VA MEDICAL CENTER, 3350 LA JOLLA VILLAGE DR, SAN DIEGO, CA 92161; VA MEDICAL CENTER, 830 CHALKSTONE AVE, PROVIDENCE, RI 02908; VA MEDICAL CENTER, 109 BEE ST, CHARLESTON, SC 29403; VA MEDICAL CENTER, 10000 BAY PINES BLVD, BAY PINES, FL 33744; VA MEDICAL CENTER, 800 BAILIE DR, AUGUSTA, GA 30901; VA MARTINEZ 612, 150 MUIR RD, MARTINEZ, CA 94553; VA MEDICAL CENTER, 1501 SAN PEDRO SE, ALBUQUERQUE, NM 87108; VA MEDICAL CENTER, 1670 CLAIRMONT RD, DECATUR, GA 30033; VA MEDICAL CENTER - ST LOUIS, 915 N GRAND AVE, ST LOUIS, MO 63106; and VA MEDICAL CENTER HARLINGEN, 2106 TREASURE HILLS BLVD, HARLINGEN, TX 78550. "Other" is National Institute of Health. | Terminated | April 03, 2014 | 738 | 245
Z-1357-2012 | 61333 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | The Polarus Plus Humeral Rod (part number HR-0824-S) is intended as an intramedullary fixation device for humeral fractures. It is sealed in an internal sterile blister pack within a box that is labeled in part: "Polarus Humeral Fixation System Size: 8mm X 240mm Polarus Plus Humeral Rod***Titanium***REF HR-0824-S***LOT 266220***RX only***Manufactured or distributed by: ACUMED 5885 NW Cornelius Pass Road Hillsboro, OR 97124 USA***" |  |  | March 07, 2012 | 2012 | Acumed LLC | The packaging contains the wrong product. It does not contain model HR-0824-S as the package indicates. | Labeling mix-ups | Acumed sent an Urgent Notice Device Recall letter dated March 5, 2012, to all affected consignees. A follow-up notification letter was sent to the domestic customers on March 9, 2012. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were instructed to further notify their consignees about the recall and to cease use of and return the affected product. Consignees were requested to respond using the enclosed response form. The form should be faxed to 503-520-9618. For questions regarding this recall call 503-627-9957, ext. 1302. |  | 12 | 12 rods | Worldwide Distribution - USA (nationwide) and the countries of Italy, Ireland, and Great Britain. | Terminated | July 25, 2012 | 140 | 12
Z-1361-2012 | 61445 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity CT Computed Tomography X-Ray Systems. Brilliance 64 Model# 728231, Ingenuity CT Model # 728326, Distributed By: Philips Healthcare, Cleveland, OH. |  |  | March 12, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare is performing an upgrade to the field for issues related to Philips Brilliance 64 and Ingenuity CT scanner having Software Version 3.5.1. When the user plans a scan with tilt applied, the scan increment changes from the user defined values (noncontiguous) to a contiguous scan. | Software design | . |  | 440 | 440 units | Worldwide Distribution-USA (nationwide) including the states of AK, AR, AZ, CA, CO, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MT, ND, NE, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA and WV, and the countries of Australia, Bahamas, Bangladesh, Belgium, Bolivia, Brazil, Chile, China, Colombia, Croatia, Denmark, Finland, France, Georgia, Germany, Great Britain, India, Indonesia, Iraq, Israel, Italy, Japan, Kazakhstan, Korea, Laos, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Poland, Russia, Slovakia, South Africa, Spain, Sudan, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom and Vietnam. | Terminated | June 18, 2013 | 463 | 440
Z-1365-2012 | 61392 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Version 3.10.00. |  |  | March 05, 2012 | 2012 | Computerized Medical Systems Inc | An unnecessary dose refresh can occur in certain situations which causes inconsistencies in the Beam Doses and the Total Dose. | Software design | An "IMPORTANT SAFETY NOTICE" was issued between 3/5-9/12 via e-mail, FAX, or traditional mail. The Notice described the correct functioning of the system, the problem, when and why the issue occurs, the clinical impact, and the workaround. Patch 3.10.01 is now available. Customers can download their patch at the link provided in the Notice. The Notice also provided a list of offices where customers can contact if they have questions or further support is needed. |  | 44 | 44 | Worldwide Distribution -- USA, including the states of OH, IN, WI, CA, FL, IL, LA, NJ, MA, NC, and MO and the countries of Canada, Australia, Belgium, China, France, Germany, Japan, The Netherlands, New Zealand, Portugal, Switzerland, Taiwan, and the United Kingdom. | Terminated | April 01, 2014 | 757 | 44
Z-1362-2012 | 61327 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Servo-i ventilator system |  |  | December 22, 2011 | 2011 | Maquet Inc. | The Servo-i and Servo-s ventilator systems equipped with specific revision os of the AC/DC convertor can result in electric shock during servicing. | Device Design | Maquet sent a Product Recall Urgent-Medical Device Corrective Action letter dated December 22, 2011 to all affected customers. The letter identified the affected products, patient risk, product issue and actions to be taken. Customers were instructed to complete and return the attached Customer Fax Response form as soon as possible. Maquet Field Service Technicians will update all affected products at all the concerned customer sites. Customers are advised untill Maquet has updated all the affected products (all of which are still under warranty), inform all the concerned biomedical staff about the content of this Device Correction Letter. For technical support call 1-888-627-8383. |  | 3136 | 647 US; 2462 Foreign; 27 Spare Parts Foreign | USA Nationwide Distribution | Terminated | October 22, 2013 | 670 | 3186
Z-1371-2012 | 52174 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | The EM 2400 is a Pharmacy Compounding System consisting of three sub-systems for compounding solutions for intravascular administration. |  |  | June 03, 2008 | 2008 | Baxa Corporation | Remote possibility of an unintended ingredient delivery as a result of a damaged valve actuator on the EM 2400. | Device Design | The firm, Baxa, sent a "Exacta-Mix 2400 Compounder Safety Alert" dated June 3, 2008 to all customers. The alert describes the product, problem and actions to be taken. The customers were instructed to not use the main module if any of the 24 valve actuators are damaged; do not infuse any bags if damage is noticed after they have been pumped; do not try to repair or adjust a damaged or bent valve actuator; immediately alert all compounder users of this issue to ensure that when the device is used, the valve actuators are always inspected for damage;" inspect all valve acutators before setup and after tear-down to ensure they are not damaged or bent. Do not use the compounder if the valve actuators are damaged or bent."; and if a damaged or bent valve is noticed after using the EM2400, immediately hold all pumped bags and contact Baxa Technical Support. If you need additional assistance, please contact Baxa Technical Support (800-678-2292). Technical Support is available 24 hours a day, 7 days a week. |  | 1000 | 1000 units (approximately) | Worldwide distribution: USA (nationwide) and countries including: Australia, Canada, China, New Zealand, and United Kingdom. | Terminated | April 10, 2012 | 1407 | 1000
Z-1367-2012 | 61359 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Extended Brilliance Workstation Computed Tomography X-Ray Systems, M/N 728260, Distributed by: Philips Medical System, Cleveland, OH |  |  | January 26, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | TumorLOC will save the incorrect contour when the user saves the contour while in the edit mode, and the software does not force the user to exit the edit mode before saving. In TumorLOC, when saving while a contour is in edit mode, that contour will shift to the slice corresponding to the most recently drawn contour of the last organ in the Organ list. | Software design (manufacturing process) | Philips sent an Urgent - Medical Device Correction letters to all affected customers. The letter identified the product, what the problem is and under what circumstances it can occur, important information for the continued safe and proper use of the affected equipment, and actions to be taken to minimize the effect of the problem. Philips instructed their customers to put a copy of the letter with their "Instructions for Use" manual. For question contact your local Philips representative or local philips healthcare office. For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site Id or follow the prompts). |  | 49 | 49 units | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY. Product was also shipped to the following countries: Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakstan, Kenya, Kuwait, Libya, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Martinique, Mexico, Netherlands, Netherlands Antilles, New Zealand, Norway, Pakistan, Palestine, Peru, Philippines, Poland, Portugal, Russian Fed, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom &amp; Venezuela. | Terminated | February 22, 2013 | 393 | 1110
Z-1335-2012 | 61499 | No | RFK |  |  |  |  | U-SPECT-II/CT and VECTor/CT cabinet x-ray products. |  |  | April 15, 2012 | 2012 | Milabs Bv | These products have doors which are not equipped with at least one safety interlock which physically disconnects power to the high voltage power supply for the x-ray source. Additionally, these products do not have a certification label and the identification labels do not include the month and year of manufacture. | Radiation Control for Health and Safety Act | A notification will be sent to all affected purchasers by April 15, 2012. The manufacturer will bring the products into compliance with the applicable performance standards within 60 days. Specifically at least one door safety interlock will be replaced per door and revised labels will be produced and applied to the system. These corrections will be made in accordance with MILabs instructions by a MILabs installation and service engineer. |  | 2 | 2 | Nationwide Distribution -- Indiana and North Carolina. | Terminated | August 29, 2012 | 136 | 2
Z-1370-2012 | 61316 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Hologics SecurView DX workstation, software version 7.3.0 |  |  | February 28, 2012 | 2012 | Hologic, Inc. | SecurView DX with software version 7.3.0 defects when used to view mammography tomosynthesis images | Software design | Hologic Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 28, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides instructions on how to re-review the stored images and correct any annotation defect. A SecurView DX 7.3.0 Confirmation of Notification form was enclosed for customers to complete and return to the firm via fax at 866-652-8674. Contact Hologic's Help Desk at 1-877-371-4372 for questions regarding this notice. |  | 7 | 7 units | Worldwide Distribution-Distributed in the state of Texas and the countries Belgium, Brazil, and Canada. | Terminated | May 07, 2012 | 69 | 7
Z-1382-2012 | 60152 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500, 500V. |  |  | February 03, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Replacement of the SBC batteries (system batteries) should be managed as "routine maintenance" instead of as an in-house Service Procedure because some systems may be relying on batteries beyond their useful life. | Labeling design | On 2/3/12 all consignees received an URGENT MEDICAL DEVICE RECALL letter via Federal Express. The letter described the reason for the recall, identified the affected product population, what to do in the event of a failure, the potential hazard, and correction. Terumo CVS will update the Operator's Manual for the CDI 500 with instructions to replace the battery within its useful life. Users are to review the notice and ensure that all users are made aware of the issue. In addition, the Customer Response Form should be completed and returned. Users can contact their local Terumo Cardiovascular Systems representative or call 1-800-521-2818 with questions or concerns, to check the last time system batteries were replaced, or to schedule maintenance for systems approaching 10 years. |  | 3829 | 3,829 total units distributed | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, WV, WA, and WI and the countries of Japan, Germany, Barbados, Saudi Arabia, Columbia, Uruguay, Ecuador, Trinidad, Costa Rica, Paraguay, Panama, Belgium, Jordan, Libya, Mexico, Brazil, Honduras, Canada, Korea, Dominica Republic, El Salvador, Guatemala, Vietnam, Chile, Hong Kong, United Arab Emirates, Singapore, Australia, Malaysia, Taiwan, Philippines, Mexico, Thailand, Vietnam, and South Africa. | Terminated | August 29, 2013 | 573 | 3829
Z-1377-2012 | 61084 | No | DFC | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | Wako Autokit Lp(a) Calibrator, model code# 998-41491, packaged in 1 ml bottles, 2 bottles/kit, and labeled in part ***Manufactured by: Wako Pure Chemical Industries, Ltd. 1-2, Doshomachi 3-Chome, Chuo-Ku Osaka 540-8605, Japan***Distributed by: Wako Chemicals USA, Inc. /Wako Diagnostics |  |  | January 30, 2012 | 2012 | Wako Chemicals USA, Inc. | In-vitro diagnostic kit calibrator is defective and could impact patient profile for cardiovascular disease. | Device Design | Wako Diagnostics notified its customers by telephone on January 26, 2012 and sent an Urgent Medical Device Recall letter dated January 30, 2012 to inform them of the product recall. The notifications identified the affected product, problem and actions to be taken. Consignees were requested to: 1) Complete and sign the attached Return Materials Authorization (RMA) Form to receive replacement material and 2) Package the recalled Product to be returned and send to the address listed on the RMA form. For questions call Technical Support Team at 1-877-714-1924. |  | 10 | 10 kits | The product had been sold to 2 customers in the US. | Terminated | May 02, 2012 | 93 | 10
Z-1375-2012 | 61118 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Assay Protocol File (APF) Software versions for UniCel DxC 600i Systems. |  |  | September 11, 2008 | 2008 | Beckman Coulter Inc. | The recall was initiated because Beckman confirmed a customer report that a timing conflict can occur while aspirating samples for the Access Total BhCG assay. | Software design | Beckman Coulter sent a "PRODUCT CORRECTIVE ACTION" letter dated September 24, 2008 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Customers were instructed to complete and return the enclosed response form. Customers with questions regarding this notification were instructed to contact Technical Support at (800) 854-3633 or their local Beckman Coulter representative. |  | 2948 | 2948 units total (2762 units in the US) | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | June 26, 2012 | 1384 | 2948
Z-1376-2012 | 61436 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OneTouch(R) Ping(R) Insulin Pump glucose management system. The device is intended to deliver insulin therapy. |  |  | March 01, 2012 | 2012 | Animas Corporation | There is an error in the pump software where the OneTouch(R) Ping(R) insulin pump did not allow the user to save changes to the time and date on the pump if those changes were made on February 29, 2012 (Leap Day). Users who did not attempt to change the time or date on their pump on February 29, 2012 are not affected by this issue. | Software design | The firm initiated their recall on March 1, 2012 and sent Urgent Notification letters dated March 6, 2012 to all consignees by UPS with delivery confirmation. Customers were notified of the issue and the appropriate steps to take in order to mitigate the issue. A Web Letter for users was also posted to the Animas web site. Questions were directed to (855) 230-7577. |  | 91239 | 91,239 | Worldwide Distribution, (USA) including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY and the country of Canada and the territory of Puerto Rico. | Terminated | December 24, 2013 | 663 | 91239
Z-1389-2012 | 52024 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Pharmacy compounding system . It is a twenty-four source compounder for multi-source fluid mixing of both macro and micro ingredient down to 0.2 ml quantities. |  |  | March 23, 2009 | 2009 | Baxa Corporation | Software issue with compounding pump resulting in over delivery of component ingredients during routine operations. | Software design | Baxa sent a Safety Alert letter dated March 23, 2009 to all affected consignees. The letter identified the affected product, problem and recommended actions with specific instructions related to the issue. Additionally, the firm will be providing a warning label for the affected product as soon as possible and is currently working on a product upgrade to correct the issue. For assistance contact Baxa Technical Support at 800-678-2292. The firm is also providing a 24 hour a day, 7 days a week Technical Support for the affected product. |  | 1322 | 1322 units | Worldwide Distribution - USA Nationwide including Canada. | Terminated | April 10, 2012 | 1114 | 1322
Z-1386-2012 | 61405 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Aviva Nano Blood Glucose Meter |  |  | March 12, 2012 | 2012 | Roche Diagnostics Operations, Inc. | Accu-Chek Aviva Nano Blood Glucose meters that were manufactured for the purpose of validating a manufacturing process for product to be distributed outside the US were distributed in the US through eBay by an unauthorized third party. These meters were not intended for distribution, were not verified to meet final product specifications and were not distributed to customers with the necessary pr | Employee error | Roche sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated March 14, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to immediately return their ACCU-CHEK Aviva Nano Meter to the firm in an enclosed pre-paid envelope via USPS. Contact Roche Product Removal Services at 1-888-345-5352 between 8:00 am and 5:00 pm for questions regarding this notice. |  | 82 | 82 meter distributed on Ebay | Nationwide Distribution-USA (nationwide) including the states of AZ, NY, RI, TN, NJ, IN, TX, AL, NV, CA, OR, KY, IL, AR, NM, NE, MD, MO, GA, and SD. | Terminated | November 06, 2012 | 239 | 82
Z-1390-2012 | 61393 | No | JFT | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Cortisol Assay |  |  | March 06, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Cortisol Assay reagent kit lots do not meet onboard stability and calibration interval claims | Component design/selection | The firm, Siemens Healthcare, sent an "Urgent Field Safety Notice" via e-mail to all affected customers on March 6, 2012 in the United States and International. The Urgent Field Safety Notice describe the product, problem, and actions to be taken. The notice instructed the customers to perform daily calibration on all Cortisol reagent kits to maintain consistently accurate quality control and patient values; to complete and return the confirmation fax-back form 'Field Correction Effectiveness Check' via fax to the Technical Solutions Center at (302) 631-7597 and forward this notification to whomever they many have distributed this product.If you have any questions, contact Quality Engineer 5- Quality Systems and Compliance at 508-660-8540 or www.siemens.com/diagnostics. |  | 2410 | Total: 40023. 50 Test Kits = 21734;250 Test Kits= 15879;250 REF Test Kits= 2410 | Worldwide distributed: USA (nationwide) and countries of: ARGENTINA, AUSTRALIA, AUSTRALIA, BANGLADESH, BD, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, EGYPT, GERMANY, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, JAPAN, KOREA, MALAYSIA, MEXICO, NEW ZEALAND, PAKISTAN, PHILIPPINES, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, TAIWAN, THAILAND, URUGUAY, VENEZUELA, and VIETNAM. | Terminated | June 26, 2015 | 1207 | 2410
Z-1384-2012 | 61178 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SenDx Medical Inc |  |  | January 30, 2012 | 2012 | Sendx Medical Inc | The product can cause the analyzer to report values outside the published performance specifications for pH, pCO2, cNa+, cK+, cCa2+ and cCl-. | Nonconforming Material/Component | SenDx Medical sent Customers notification letters dated January 30th, 2012, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer. We have recently become aware of a possible condition in a limited number of ABLBO FLEX analyzers configured with a specific Control PCe lerinted Circuit Board), which can cause the analyzer to report values outside the published performance specifications for pH, pCOz, cNa*, cK*, cCa2* and cCl-.Customers were instructed to immediately implement Alternative 1 or 2 until the signal data PCB is replaced. Alterative I:Stop the use of all affected analyzers until corrective actions can be performed on these analyzers.Alternative 2:Verify proper function of the affected analyzers immediately following everypatient sample measurement by measuring a manual QC sample.Customers were asked to complete the fax form acknowledging receipt of the notice and provided a contact for Technical Support at 1-800-736-0600 opt.4. |  | 863 | 819 analyzers and 44 signal data PCBs | One US consignee in state of OH. Other consignees are international such as Australia, France, Netherlands, Denmark, China, United Kingdom, Turkey, Switzerland, Japan, India, Poland, Spain, and Germany. Distribution list submitted with supporting doc.s | Terminated | July 13, 2012 | 165 | 863
Z-1383-2012 | 61483 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Luminos dRF solid state x-ray imager |  |  | February 16, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During regular product monitoring, Siemens became aware of a potential issue with the AXIOM Luminos dRF. Firm became aware of a potential issue with torque wrenches - during a regular scheduled calibration of the wrench, it was discovered that the wrench was out of the specification. The wrench was over-torquing approximately 21% for low-end readings (50 ftlbs) and 9% for high-end readings (250 | Nonconforming Material/Component | Siemens initiated a recall/correction on February 16, 2012 by issuing Update Instruction XP056/11/S to the affected customers. All screws on the unit support on AXIOM Luminos dRF units will be replaced by a Siemens service engineer.For information regarding this recall call 610-219-4834. |  | 2 | 2 | Nationwide Distribution including California | Terminated | November 06, 2014 | 994 | 2
Z-1385-2012 | 61364 | No | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | VariSeed 7.1, software application, part number H62, manufactured by Varian Medical Systems. VariSeed 7.1 is a computer based software application for planning and evaluating prostate brachytherapy procedures. |  |  | January 25, 2012 | 2012 | Varian Medical Systems, Inc. | Brachytherapy device requires the user to enter radioactive source activity, but software may specify unintended units for activity and lead to overdose in patient treatment. | Labeling design | Varian Medical Systems notified direct account healthcare practitioners/ medical physicists with "Urgent Medical Device Correction/Urgent Field Safety Notice" letters by mail on 01/19/2012. Varian's advisory described how to mitigate user error and the firm's intentions to revise Instructions for Use. Technical Support phone numbers were provided for US and international customers. |  | 1377 | 1377 units | Worldwide Distribution | Terminated | November 07, 2012 | 287 | 1377
Z-1392-2012 | 61395 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | BioCare Medical Vina Green Chromogen Kit; 25 mL, 100ml IVD. Kits include: Vina Green Chromogen 1. 0 mL; Vina Green Buffer, 25 mL and Vina Green Dropper Bottle. BioCare Medical, Concord, CA 94520. |  |  | March 13, 2012 | 2012 | Biocare Medical Llc | Reports of sporadic failure of specific lots of Vina Green Chromogen Kit, either through crystal formation that obscures positive staining, color not right, or a complete fading of staining by the chromogen. | Nonconforming Material/Component | "Medical Device Recall Notification" letters were sent on March 13, 2012 to all direct account. Letters informed customers of the reason for recall and the specific lots and catalog numbers involved, along with the reason for recall, how to identify the recalled product, and the actions to be taken. Customers were to immediately discontinue use of the affected product lot(s). They should dispose of the product according to national, state, and local laws and regulations. The enclosed Return Response Form should be completed and returned. Customers will be credited upon receipt of the Response Form. |  | 74 | 74 kits | Worldwide Distribution -- USA, including the states of WI, KY, CA, GA, IN, MA, IL, MD, NC, LA. AZ, TX, NC, NJ, MO, NY, PA, OH, DC, MI, WA, and CO, and the countries of Canada, Sweden, Germany, Italy, and Switzerland. | Terminated | September 20, 2012 | 191 | 74
Z-1391-2012 | 61347 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink(TM) Informatics Systems Software Version 5.0 and Software Version 5.0 Service Packs 1 through 4 used with the Dimension Vista(R) System. |  |  | December 02, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | QC results may be unexpectedly associated with a Pending Control. When this situation occurs, Pending QC Controls maybe be created or reused. | Device Design | Siemens Healthcare Diagnostic Inc, sent a Urgent Field Safety Notice dated December 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were asked to follow these instructions. The first option is Extending the "Minimum Inactive Days" to 999 days. The second option is to proactively monitor the system for unexpected Pending QC Controls.Long Term Action Plan: A corrective and preventive action (CAPA) investigation has been issued. Currently firm is actively investigating.Further questions please call (302) 631-0516. |  | 787 | 787 | Worldwide Distribution -- USA (nationwide) and the countries of CANADA, FRANCE, GERMANY, SPAIN, PORTUGAL, ITALY, JAPAN, NORWAY, AUSTRALIA, AUSTRIA, NEW ZEALAND, BELGIUM, SWITZERLAND, NETHERLANDS, MALAYASIA, SLOVENIA, SOUTH KOREA, DENMARK, SAUDI ARABIA, SLOVAKIA, PUERTO RICO, INDONESIA and TAIWAN. | Terminated | March 21, 2014 | 840 | 787
Z-1387-2012 | 61249 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Elecsys Troponin I Immunoassay |  |  | March 12, 2012 | 2012 | Roche Diagnostics Operations, Inc. | An investigation has confirmed a considerably reduced recovery of Li-heparin plasma samples for Troponin I lot 163176 and Troponin I STAT lot 163177. In the worst case, Li-Heparin sample recovery may be as low as 50% of the serum recovery. Correct serum recovery has been confirmed for the affected lots.The issue is caused by the raw material lot (Poly-L-Lysin) instability. Poly-L-Lysin is the | Nonconforming Material/Component | The firm, Roche Diagnostics Corporation, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated March 12, 2012 to all customers that received the affected lots of Elecsys Troponin I or Elecsys Troponin I STAT. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of the affected product; discard the affected product from their inventory according to their site's local regulations; if their facility has distributed the affected product to other sites, ensure that this letter is provided to those sites; complete and return the attached fax form via fax to 1-888-912-8457 and file this letter for future references.If you have any questions about the information contained in this letter, please contact Roche Diagnostics Technical Support, 24 hours a day, seven days a week at 1-800-428-2336. |  | 739 | 739 kits | Worldwide distribution.Argentina, Kenya, Austria, South Korea, Belgium, Kuwait, Bahrain, Morocco, Brazil, Mexico, Switzerland, Nigeria, Chile, New Zealand, Columbia, Philippines, Germany, Pakistan, Algeria, Poland, Egypt, Qatar, Spain, Tunisia, France, Turkey, United Kingdom, Taiwan, India, Ukraine, Iran, Venezuela, Italy, South Africa, Jordan, and Armenia. | Terminated | February 06, 2013 | 331 | 2027
Z-1394-2012 | 61505 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, IMAGER II Angiographic Catheter, Sterilized with ethylene oxide gas. Made in Ireland, Manufactured for Boston Scientific Corporation, One Boston Scientific Place, Natick, MA 01760-1537. |  |  | March 15, 2012 | 2012 | Boston Scientific Corporation | Boston Scientific Corporation is initiating a Medical Device Recall of certain lots/batches of the Imager" II Catheter (Urology) and the Imager" II Angiographic Catheter . An investigationhas found that, during shipping, some units in the 5-pack box may experience a physical interaction between the carton and the area onthe pouch where the hub is located. This interaction was found to cause pi | Package design/selection | Boston Scientific sent an " Urgent Medical Device Recall" letter dated March 14. 2012. the letter was addressed to "Dear Risk/Materials Manager. The letter described the problem and the product involved in the recall. Advised consignees to cease the distribution of any remaining unused product affected. Advised distributors that the recall notification should be forwarded to their customers. Requested consignees to complete and return the Reply Verification tracking Form enclosed by fax to 508-683-5578.For questions regarding this recall 763-494-1700.Recall Expanded 4/24/12:Imager II Angiographic customers:Eight total number of customers not notified initially (3 USA and 5 APA customers were sent the same initial letter with a a revised send date of April 18, 2012, 329 total number of US Customers notified initially with an incomplete product list were sent a revised notification letter. Urology: Imager II Torque customers:1 US Customer was not notified initially. This customer was sent the same initial letter with a revised send date of April 18, 2012.Fifteen US Customers notified with incomplete product list. These customers were sent a revised notification letter. |  | 406450 | 406,450 (416,555 USA, 389,895 OUS) | Worldwide Distribution - USA (nationwide) and the countries of Arab Emir, Azerbaijan, Austria, Australia, Azerbaijan, Belgium, Bulgaria, Bahrain, Bermuda, Brazil, White Russia, Canada, Switzerland,Chile, China, Colombia, Cyprus, Czech Republic, Germany, Denmark, Ecuador, Estonia, Egypt, Spain, Finland,France, Great Britain, Greece, Guatemala, Hong Kong, Honduras, Croatia, Hungary, Ireland, Israel, Italy,Jordan, South Korea, Kuwait, Lebanon, Lithuania, Morocco, Malta, Mexico, Malaysia, Netherlands, Norway,New Zealand, Oman, Panama, Peru, Pakistan, Poland, Portugal, Qatar, Romania, Serbia, Russian Fed., SaudiArabia, Sweden, Singapore, Slovakia, El Salvador, Syria, Thailand, Tunisia, Turkey, Trinidad, Tobago, Taiwan,Ukraine, Vietnam, South Africa, Dutch Antilles, Sri Lanka, Libya, and Yemen | Terminated | April 14, 2014 | 760 | 417790
Z-1395-2012 | 61408 | No | LPC |  |  | Cardiovascular | PMA | 12Fr SLS II Laser Sheath. The laser sheath is intended for use as an adjunct to conventional lead extraction tools in patients suitable for transvenous removal of chronically implanted pacing or defibrillator leads constructed with silicone or polyurethane outer insulation. |  |  | March 13, 2012 | 2012 | Spectranetics Corp. | Isolated lots of the 12Fr SLS II Laser Sheath can have a split in the outer jacket. | Nonconforming Material/Component | Spectranetics sent Recall Notices to all customers on March 14, 2012. The product issue was described and requested actions were included. Questions were directed to Customer Service at 1-800-231-0978 (US). |  | 615 | 615 | Worldwide Distribution -- USA , and the countries of Australia, Belgium, Canada, Czech Republic, France, Germany, Italy, Ireland, Japan, Netherlands, Russia, Sweden, Switzerland, United Kingdom and Venezuela. | Terminated | May 02, 2012 | 50 | 615
Z-1396-2012 | 61434 | No | KSR | Empty container for the collection and processing of blood and blood components. | Hematology | Hematology | 510(k) | DeRoyal (R) Umbilicup, REF 72-8000, Rx Only, STERILE EO, Manufactured by DeRoyal, Powell, TN 37849 |  |  | January 20, 2012 | 2012 | DeRoyal Industries Inc | Device's needle may become dislodged from the cup during shipment or during use. | Nonconforming Material/Component | DeRoyal Industries sent Recall communication by letter to all affected consignees of the Non-Sterile Bulk and Single Sterile Units on January 20, 2012 and of the Kits and Trays on Febraury 6, 2012. The letters identified the affected products, problem and actions to be followed. Customers were instrructed to complete the enclosed Notice of Destruction form and return this form by fax (865-362-3716) or email (recalls@deroyal.com) within two weeks of the date of this letter. DeRoyal will use your response on the Notice of Destruction form to send replacements as soon as they become available. Distributors were instructed to forward this recall to any of their end users that may have affected product.If you have questions or need assistance with the recall, please call 865.362.1020 or Robby Lockett at 865.362.1034. |  | 2472 | 2472 units | Worldwide Distribution - USA Nationwide and the countries of: Canada, Dominican Republic, South Africa, and the United Arab Emirates. | Terminated | July 11, 2013 | 538 | 9434
Z-1420-2012 | 52120 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxa EVA TPN bags for administering parenteral nutrition solutions. Sold under three brand names: Exacta Mix EVA Bags, Exacta Mix Dual-Chamber EVA Halobag, and Baxa Calibration Bags. |  |  | May 08, 2009 | 2009 | Baxa Corporation | Spike port flor EVA TPN Bags contains low levels of DEHP which could leach out. | Labeling False and Misleading | Baxa sent a Customer Notice letter dated May 8, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Current Baxa product is mislabeled as non-DEHP but this does not present a risk to patients. Baxa is committed to providing products that are non-DEHP. We are working with our supplier to replace the spike port with a non-DEHP as soon as possible. In the meantime, product labeling will be updated to reflect this finding.For further questions please call (303) 690-4204. |  | 17374700 | 347,494 cases/ 50 bags each | Worldwide Distribution -- USA (nationwide) | Terminated | April 13, 2012 | 1071 | 17374700
Z-1401-2012 | 61446 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Non-DEHP Microbore Administration Set with Non-Vented Spike Packaged Assembly |  |  | March 20, 2012 | 2012 | MOOG Medical Devices Group | Ambulatory administration sets used with the Curlin Ambulatory Pumps including the 4000 CMS Ambulatory Pump, 6000 CMS Ambulatory Pump and the Pain Smart IOD Ambulatory Pump may leak during use. | Nonconforming Material/Component | Moog Medical Device Group sent a Medical Device Recall Notification letter dated March 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the affected product from their inventory, contact the distributor or Moog Customer Service at 1-800-970-2337, prompt #7, then return product to distributor or Moog.For additional questions regarding this recall call Moog's Customer Service Department at (800) 970-2337. |  | 11628 | 11,628 | Worldwide Distribution - USA including AZ,,, CA,,, FL,,, GA,,, KS,,, IL,,, MA,,, MD,,, MI,,, OH,,, PA,,, TX,,, VA,,, WI,,, and the countries of Germany,,, Italy,,, Morocco,,, New Zealand,,, and the United Arab Emirates,,, | Terminated | July 26, 2012 | 128 | 43176
Z-1400-2012 | 61370 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | TeleSentry, Multi-parameter Ambulatory Telemetry, Model #TS01. Dist. By: ScottCare Corporation, Cleveland, OH. |  |  | January 19, 2012 | 2012 | Scottcare Corporation | Certain TeleSentry Devices contain a Bad SD Card. When the SD Card goes bad, the TeleSentry Device cannot read or write data to the SD Card. The TeleSentry Device then stops functioning. | Component design/selection | ScottCare sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 19, 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The letter instructs customers to immediately discontinue use and quarantine any affected products. A Recall Acknowledgement Form was enclosed for customers to complete and return via fax to (216) 257-6129. Contact the firm at (216) 362-0550 ext. 113 for questions regarding this recall. |  | 513 | 513 units | Worldwide Distribution-USA (nationwide) including the states of CA, DE, FL, MI, NJ, NV, NY, OH, PA and TX, and the country of Canada. | Terminated | September 27, 2013 | 617 | 513
Z-1432-2012 | 53049 | No | JIF | Ammonia test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SYNCHRON Ammonia (AMM) Reagent |  |  | March 26, 2009 | 2009 | Beckman Coulter Inc. | The Ammonia Reagent lot M808278 may fail calibration or cause a low shift in Quality Control recovery on some SYNCHRON Systems. | Unknown/Undetermined by firm | Beckman Coulter mailed a Product Corrective Action (PCA) Letter on April 13, 2009 dated April 10, 2009 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers to discontinue use and discard appropriately the lot number(s) of the affected product and contact their Customer Service Representative at (800) 526-3821 (US), or (800)463-7828 for reagent replacement. Customers were instructed to complete and return the enclosed response form within 10 days so that Beckman Coulter may maintain records.Customers were told to call the Customer Support Center (Hotline) at (800) 854-3633 with any questions regarding the PCA letter. |  | 3388 | 3,388 kits | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | April 18, 2012 | 1119 | 3388
Z-1430-2012 | 60661 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Double foot switch, type KF 2 1S/1 S-MED-AP, manufactured by the company Seute Schaltgerate. GmbH &amp; Co. KG for Ziehm Imaging GmbH Germany, and used with Ziehm Imaging GmbH Mobile C-arm as a component thereof. Two versions were made: 2 pedal version KF 2 1S/1 S-MED-AP foot switch and 4 pedal foot switch version KF 2 1S/1 S-MED-AP This type of foot switch includes a pair of the dual foot switches assembled on a single metal bracket. The following labels may appear on the foot switch: |  |  | April 29, 2011 | 2011 | Ziehm Imaging Inc | On 04/29/2011, Ziehm Imaging, Inc., Orlando, FL recalled the double foot switch, type KF 2 1S/1S-MED-AP, a component used with Ziehm Imaging GmbH Mobile C-arm due to the potential of unintentional continued activation of fluoroscopic X-ray as the foot switch may not become inactive when the pedal is released. | Component design/selection | Ziehm Imaging, Inc., Orlando, FL sent an Urgent Recall Notice letter dated April 11, 2011, that included a return receipt request, envelope, and a reply questionnaire. to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. They also called consignees. After contact, Ziehm Imaging, Inc. shipped a replacement foot switch with instructions for the customer/end user on how to install the replacement foot switch. If no signed confirmation was received, Ziehm Imaging, Inc. visited the consignee and replaced the foot switch. Customers were asked to fax the completed response form to 866-839-7893. |  | 49 | 49 units | Nationwide Distribution including Florida, South Carolina, Georgia, Alabama, Virginia, Ohio, Massachusetts, New York, Pennsylvania, Kentucky, Connecticut, Rhode Island, Indiana, Illinois, Nebraska, Minnesota, Arizona, California, Oregon, and Washington. | Terminated | April 18, 2012 | 355 | 49
Z-1429-2012 | 61455 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Norco Folding Shower Bench; |  |  | March 20, 2012 | 2012 | North Coast Medical Inc | Due to a consumer complaint, North Coast Medical determined that due to a design flaw, there was not enough plastic support material to retain bracket bolts when subjected to a force greater than 150 lbs. | Device Design | North Coast Medical, Inc. sent a Recall Notification Letter dated March 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify and discard the affected product. Complete the enclosed Form and return the form even if they no longer have the bath bench. Return the form by fax to 408-775-1430. Customers should contact the Quality Department if they have any questions, concerns or need any assistance following the instructions.For questions regarding this recall call 800-821-9319, ext 121. |  | 370 | 370 benches | Nationwide Distribution | Terminated | July 20, 2012 | 122 | 370
Z-1428-2012 | 61352 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | PantaÂż Arthrodesis Nail Support device, 510 (k) #K050882. One (1) unit of product is kitted within an instrumentation tray. The product is available in six (6) sizes. All nails are color-coded for easy size identification. Bony fixation is achieved using two (2) tibial screws, two (2) calcaneal screws and one (optional) talar screw. The nail, cross-locking screws and end cap are manufactured from a titanium alloy: TI-6AI-4V ELI ISO 5832-3, ASTM F136. The PANTA Arthrodesis Nail is intended for use in tibiotalocalcaneal arthrodesis and treatment of trauma to the hindfoot and distal tibia. |  |  | May 19, 2011 | 2011 | Integra LifeSciences Corp. | The firm received multiple complaints from their user customers (surgeons) who reported encountering difficulty inserting a calcaneal or tibial screw through the holes of the PANTA nail. | Use error | Integra issued a Field Safety Notice with an attached Acknowledgement and return Form to all Integra Sales Representatives and Customers via Fed Ex and E-mail on May 19, 2011. The Field Safety Notice informs the customers of the firm's receipt of information concerning several reported incidents that lead to the misalignment of the calcaneal of tibial when used with the Panta Arthrodesis Nail Support Device. The notice goes stated that Integra determined that the misalignment of these screws may be attributed to improper surgical technique and to that end, the Field Safety Notice is being issued in order to reinforce the importance of following the recommended guidelines our lined in the product's Surgical Technique Manual (a copy of the referenced guidelines are provided with the Field Safety Notice (FSN). The FSN goes on to direct the reader/customer to applicable section sections of the manual to which particular attention should be directed in order to minimize the potential for misalignment of the screws. The customers are asked to distribute the information provided to the surgical teams at their various institutions. Moreover, they are also asked to sign and return the enclosed Field Safety Notice Acknowledgement and Return Form so as to document their receipt of the notification and the material provided. Should the customer(s) have questions, they are directed to contact Sean Luland, Director of Marketing-Lower Extremity Reconstruction at Integra LifeScience at: 513-533-7929, for additional information.For questions regarding this recall call 609-275-0500. |  | 64 | 64 units were distributed | Nationwide Distribution including MD, IL, NU, CA, GA, CO, OH, SD, AR, FL, TX, MA, MO, NC, OK, KY, NJ, UT, VA, MN, DC, PA, WA, CT, AZ, IL, MI, AL, NE, TN, NM, NV, LA, OR, DE, WI, ND, NH, MS, SC, IN and Puerto Rico. | Terminated | February 19, 2013 | 642 | 64
Z-1424-2012 | 61304 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Attune Impaction Handle Warsaw, IN 46582 |  |  | March 16, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a Recall Notice for nine lots of the Attune" Impaction Handle due to the potential for the handle trigger to break during surgery. The Attune Impactor is provided as a part of a set of tools and is designed specifically for the installation of the Attune Knee | Device Design | DePuy Orthopaedics, Inc. sent an URGENT INFORMATION - MEDICAL DEVICE RECALL NOTICE letter dated March 16, 2012 to all consignees affected which extended to the Depuy Distributor level, surgeon and hospital level. The affected DePuy Distributors were notified via email on March 16, 2012. The sales representatives notify hospitals and surgeons by mail or in person with written communication. The letter identified the affected product, reason for recall notice, clinical implications and actions to be taken. Consignees were instructed to complete and return the Response cards to DePuy by fax to 574-372-7567 or email to:kseppa@its.jnj.com. When the new impaction handles are available, the DePuy sales representative will be responsible for providing the new handles, removing and returning the affected handles. For product-related questions, please contact your local DePuy Orthopaedics sales representative. For clinical-related questions from surgeons, please contact DePuy's Scientific Information Office at 1-888-554-2482 (M-F; 8 am - 5 pm EST.) For questions about recall information provided, please contact Katie Seppa, Manager of Customer Quality, 574-372-7333 (M-F; 8 am - 5 pm EST.) |  | 136 | 136 impaction handles | USA Nationwide Distribution including the states of: MA, MD, IL, IA, MN, WA, CA, FL, NC, VA, and IN. | Terminated | January 09, 2014 | 664 | 136
Z-1427-2012 | 61538 | [6 empty] | Hudson RCI, BiteGard Oral Bite Block, Single Patient Use, Rx Only, Teleflex Medical, Research Triangle Park, NC. The Teleflex Medical BiteGard Molar Bite Black is a disposable single patient use device that fits between the patients molars to prevent the patient from closing their mouth entirely and from obstructing the Laryngeal Mask Airway (LMA) tube, in the case where one is used. |  |  | April 03, 2012 | 2012 | Teleflex Medical | Teleflex Medical received complaints that the bite block can become separated from the handle. If separation occurs, medical intervention may be required to remove the bite block from the patient's trachea or esophagus/GI tract. | Process control | Consignees received "Urgent Medical Device Recall Notification" letters on or about 04/03/2011. The letters described the issue and provided recommended actions regarding remaining product and returns. Questions were directed to Customer Service at (866) 804-9881 (US). |  | 333724 | 333,724 units | Worldwide Distribution, including Nationwide (USA) and the countries of Canada, Singapore, China, Denmark, Netherlands, India, Australia, Japan and the United Kingdom. | Terminated | May 20, 2014 | 777 | 333724
Z-1425-2012 | 52748 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Total Bilirubin, 6L45-20 and 6L45-40. |  |  | July 09, 2009 | 2009 | Abbott Laboratories Inc. | Abbott Laboratories has issued a recall on the Clinical Chemistry Total Bilirubin due to an under-recovery on one clinical neonatal specimen resulting in a value less than the assay linear limit. | Material/Component Contamination | Abbott Laboratories plans to release a Product Correction letter to customers instructing customers to follow the steps provided in order to revise their Total Bilirubin assay parameters. The letter explained the reason for the recall, along with the patient impact. Customers were to keep the letter as part of their laboratory records and also forward a copy to any other laboratories that they may have forwarded Total Bilirubin kits. U.S. Customers should contact Customer Service at 1-877-4ABBOTT, if they have questions. |  | 11282 | 11,282 | Worldwide Distribution -- USA, Guatemala, Chile, Brazil, Canada, Germany, Hong Kong, Thailand, Puerto Rico, Australia, New Zealand, Trinidad &amp; Tobago, Bermuda, Curacao, Cayman Island, and Jamaica. | Terminated | March 07, 2014 | 1702 | 11282
Z-1431-2012 | 61547 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Zuma Mobility System Models: Zuma Mobility Assist 412045-01 and Zuma Mobility Trainer 412046-01. |  |  | April 06, 2012 | 2012 | KCI USA, Inc. | The Zuma Mobility Assist and Trainer have the potential to bind during lowering operation. | Device Design | KCI sent an Urgent - Voluntary Medical Device Correction notification letters dated April 6, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. KCI will contact affected customers to schedule a time to upgrade the device. For questions contact KCL's Customer Technical Service Center at 1-800-275-4524. Select option 3 followed by option 2. |  | 4 | 4 units | USA Nationwide Distribution including the states of ID and FL | Terminated | June 19, 2012 | 74 | 4
Z-1426-2012 | 61342 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PFC Steinman Pin/Drill Pack Sterile (WPC 667-2012) Drill Bit and Steinman Pin Packet (1 unit) |  |  | March 12, 2012 | 2012 | DePuy Orthopaedics, Inc. | Sterility of the pins is at risk because product packaging, which is designed to include 2 pouches (inner pouch and outer pouch) around the pin mounting card only, contains a single pouch, which could compromise product sterility. | Packaging process control | The firm, DePuy Orthopedics Inc., sent an "URGENT INFORMATION - MEDICAL DEVICE RECALL NOTICE" dated March 12, 2012 to its customers via mail or in person with written communication. The notice described the product, problem and actions to be taken.The customers were instructed to cease use or distribution of recalled products and to return them to DePuy in Warsaw, IN. The sales representatives are expected to aid customer in returning of recalled products. Distributors and sales representatives are to return the recalled products through the normal DePuy Returns process, to attention of Returns and marking H12-01 on the outside of the box.The customers were also instructed to complete and return the Reconciliation Form to their Depuy Sales Representative or fax to 574-372-7567. Follow-ups for the Reconciliation Form for non-responders will be conducted with the Distributors until the form is returned. For Hospitals that are now out of business or refuse to sign the response card, the Distributor will indicate this on the Reconciliation form and return.For questions about the recall information provided, please contact Manager of Customer Quality, at 574-372-7333 (M-F; 8am - 5pm EST). For product-related questions, please contact your local DePuy Orthopaedics sales representative. For clinical-related questions from surgeons, please contact Depuy's Scientific Information Office at 1-888-554-2482 (M-F; 8am - 5pm EST). |  | 4800 | 40 of each lot, 120 total | Nationwide distribution: USA including states of: FL, LA, TX, OR, CA, AZ, KY, GA, MS, and PA. | Terminated | March 06, 2013 | 359 | 4800
Z-1421-2012 | 61403 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Sterile wound closure strips, packaged as follows: |  |  | February 07, 2012 | 2012 | Deroyal Industries, Inc. Lafollette | Packages of Episeal wound closure strips and surgical kits containing them were recalled because the adhesive backing on Episeal wound closure strips may not release from the backing rendering the product unusable. | Nonconforming Material/Component | DeRoyal Industries sent a Recall of Episeal Wound Closure Strip letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to take the following action:1. Please identify and quarantine any affected inventory using the product and lot numbers listed above.2. Render affected product unusable, destroy and discard according to your facilitys guidelines.3. If you are a distributor, forward this recall to any of your end users that may have affected product or provide us with the list and we will contact them.4. Please complete the enclosed Notice of Destruction form, indicating the affected products you found in your inventory and destroyed. PLEASE COMPLETE AND RETURN THIS FORM EVEN IF YOU NO LONGER HAVE THE PRODUCT IN INVENTORY.Return this form by fax (865-362-3716) or email (recalls@deroyal.com) within two weeks of the date of this letter. DeRoyal will use your response on the Notice of Destruction form to send replacements as soon as they become available.If you have questions or need assistance with the recall, please contact DeRoyal Customer Support at 1-800-251-9864.We apologize for inconvenience this may cause you |  | 126720 | 126,720 units | Worldwide Distribution -- USA (nationwide) | Terminated | July 10, 2013 | 519 | 205661
Z-1433-2012 | 61465 | No | JGJ | Magnesium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista(R) MG Flex(R) Reagent Cartridge |  |  | February 17, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed a negative shift of approximately 0.2mg/dL [0.08 mmol/L] in patient and QC samples when using Vista Magnesium | Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notification letter dated February 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter explained the correction needed and instructed customers to forward the letter to anyone they may have distributed the affected product.For questions or concerns call 800-441--9250.For questions regarding this recall call 301-631-6299. |  | 9178 | 9,178 | Nationwide Distribution including Puerto Rico and within the US to Washington, DC and the following States: AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Puerto Rico. | Terminated | March 10, 2015 | 1117 | 9178
Z-1437-2012 | 61622 | No | KDL | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ViaSpan Cold Storage Solution 1000mL Bag distributed under the Barr/Duramed Pharmaceuticals label (10 x 1000mL bags). |  |  | March 30, 2012 | 2012 | Teva Pharmaceuticals USA, Inc. | Lack of assurance of sterility | Material/Component Contamination | Teva Pharmaceuticals USA sent an "URGENT MEDICAL DEVICE RECALL-RETAIL LEVEL" letter dated March 30, 2012. The letter identifies the problem, product, and actions to be taken by the customers. The letter instructs customers to check their inventory and to discontinue distribution of the affected product. A Stock Response Form was enclosed for customers to complete and return via fax to 817-868-5362. Contact Teva Customer Services at 800-545-8800 for questions regarding this notice. |  | 4070 | 4,070 | Nationwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CT, DE, FL, GA, IA, IL, KS, KY, LA, MA, MD, MN, MO, NC, ND, NE, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV and Washington, D.C. | Terminated | October 25, 2012 | 209 | 4070
Z-1434-2012 | 55440 | No | CIX | Albumin test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Albumin BCG (AlbG) List Number 7D53-21 |  |  | March 10, 2010 | 2010 | Abbott Laboratories Inc. | Observed particulate matter in some reagent cartridges. | Material/Component Contamination | Abbott Laboratories sent a Product Recall notification letter dated March 10, 2010, to all affected customers. The notification letter contained an explanation, patient impact, neccassary actions, and contact information.Customers were instructed to determine if they currently have any of the affected product in their inventory. Discontinue use and destroy any remaining inventory of the affected product in accordance with their institution's policies and procedures. Completed the enclosed Customer Reply Form. If customers forwarded any of the affected product they should provide a copy of the notification letter to their customers.For any questions customers should call Customer Service at 1-877-422-2688.For questions regarding this recall call 626-440-0705. |  | 2718 | 2718 (U.S. 652) | US nationwide and internationally to Germany, Hong Kong, Australia, New Zealand, Trinidad/Tobago, and Jamaica. No dsitribution of Canada or Mexica. | Terminated | April 18, 2012 | 770 | 2718
Z-1435-2012 | 61620 | No | HPZ | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | MEDPOR PLUS SST EZ 22mm Sphere Implant, Howmedica Osteonics Corp, 15 Dart Road, Newnan, GA 30265. MEDPOR PLUS Orbital Implants are intended to restore the volume of an enucleated or eviscerated eye globe. |  |  | March 20, 2012 | 2012 | Howmedica Osteonics Corp dba Stryker Craniomaxillofacial | The MEDPOR PLUS SST EZ 22mm sphere implants are not within specification of 22mm +/- 8, therefore implant is not easily released from the syringe. | Process control | The sole distributor was notified by "Urgent Product Recall" letter on March 20, 2012. "Urgent Medical Device Recall" letters were sent to Stryker Customers on March 26, 2012. The letters described the product issue, hazards, and actions recommended to mitigate the risk. Customer number provided: (800) 962-6558. |  | 31 | 31 units | US Distribution, the state of Michigan. | Terminated | August 10, 2012 | 143 | 31
Z-1441-2012 | 49059 | No | LGW |  |  | Neurology | PMA | OMG (Observational Mechanical Gateway) Connector - A Model Number SC-9305 |  |  | April 22, 2008 | 2008 | Advanced Bionics Corporation | The recall was initiated because there is a possibility that the OMG Connector-A is incorrectly packaged with an OMG Connector-M Instruction manual. | Labeling mix-ups | Algotec Limited initiated a Field Safety Notice Urgent Medical Device Correction letter dated April 22, 2008, with Boston Scientific forwarding a recall letter to one customer who received the recalled product. The letter instructed the consignee to replace the previously packaged manual with the OMG Connector-A manual that was supplied with the notification. The replacement of the manuals is to be confirmed via the Verification form.Customers were instructed to complete and fax the Verification Form to their local Boston Scientific Customer Service to Fax 00 44 1442 411732 by 06 May 2008.For questions regarding this recall call 661-949-4481. |  | 6 | 6 Units | Worldwide Distribution including Great Britian | Terminated | April 19, 2012 | 1458 | 6
Z-1440-2012 | 49410 | No | GZJ | Transcutaneous electrical nerve stimulator for pain relief. | Neurology | Neurology | 510(k) | Synaptic 3600 and 3800 with Cranial Electrical Stimulation manufactured by Synaptic USA, Aurora, CO. |  |  | July 18, 2008 | 2008 | Synaptic Corp | Device was distributed without marketing approval. | PMA | Synaptic Corp, sent a Urgent Medical Device Recall letter dated July 14th, 2008, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.The customer was instructed to return your 3600 and or 3800 devices so that the cranial stimulator can be removed and destroyed. Please call 1-800-685-7246 to schedule pick up of your device..Complete the enclosed Confirmation Form and return it ( postage paid) to Synaptic Corp. It is important that you return the attached response confirmation form even if you do not have the (CES) module in your model.Should you have any questions about this recall please call 1-800-685-7246 ext. 101. |  | 59 | 59 units | Worldwide Distribution -- USA (nationwide) and the country UK. | Terminated | April 19, 2012 | 1371 | 59
Z-1438-2012 | 61089 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF. CP113456, MAK OSS No Head Lock Pin, NO HEAD / WITH PLUG CO-CR-MO ALLOY/ARCOM UHMWPE |  |  | March 20, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action after receiving complaints of the MAK OSS No Head Lock Pin dislodgingfrom its intended position (post-initial surgery). An investigation was initiated and the available lots ofthe Lock Pin were evaluated and measured against print specifications. | Component design/selection | Biomet sent an "URGENT MEDICAL DEVICE SAFETY ALERT" letter dated March 20, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to continue monitoring implanted patients for disassociation of the pin component and report to Biomet. The Alert included a complete listing of device invoice history and patient information (if supplied). Contact the firm for confirmation of receipt of the notice and for questions at 574-372-3983. |  | 64 | 64 units distributed | Distributed in California and Pennsylvania. | Terminated | February 07, 2013 | 324 | 64
Z-1443-2012 | 61453 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PREMIER(R) Anterior Cervical Plate System Adjustable Drill Stop, REF 6905712, QTY: 1 EA, Rx only, Medtronic Sofamore Danek USA, Inc., 1800 Pyramid Place, Memphis, Tennessee 38132 |  |  | March 13, 2012 | 2012 | Medtronic Sofamor Danek USA Inc | Drill stops may be bypassed when the device is used with powered drilling instruments. | Device Design | Medtronic notified their US Sales force on March 5, 2012, of the upcoming field action. The written "Voluntary Recall" communications were disseminated March 13 and March 14 to the US consignees. The letter identified the affected products, problem and actions to be taken. Customers are instructed to 1) disseminate this information to personnel within their facility as appropriate; 2) review work area, district sales office, and sample inventory for affected product; 3) complete the attached Risk Manager Questionnaire and return any recalled product; and 4) Sign and acknowledge the attached Sales Rep Questionnaire indicating receipt of this letter. For questions or concerns, contact your Medtronic Sales Representative or Global Quality Department at 800-876-3133, extension 3197. |  | 1411 | 1411 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Barbados, Brazil, Netherlands, Hong Kong, Japan, Korea, Columbia, Chile, Canada, India, Jamaica, Switzerland, Costa Rica, Ecuador, and Mexico. | Terminated | April 30, 2013 | 413 | 4198
Z-1439-2012 | 61188 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Optilock, Locking Screw 3.5 mm x 28 mm, non sterile, REF 37728, Biomet Trauma Parsippany, NJ 97054. |  |  | March 15, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action after a complaint was received that item 36728, Non-Locking Screw 3.5mm x 28mm was in a package labeled as 37728, Locking Screw 3.5mm x 28mm. | Labeling mix-ups | A March 15 2012 Urgent Medical Device Recall Notice was sent to Distributors, Operating Room Managers, and Physicians with directions to locate and remove recalled items with instructions to provide the recall notice to any sub-accounts. The letter also identified the affected product and explained the reason for the recall. OR Managers and Physicians are to confirm recall notice receipt by calling 800-348-9500, extention 3983 or 3009 and make arrangements for product return. Distributors are to complete and FAX back the response tracking form to the attention of Angi Dickson, Biomet @ 574-372-1638. Questions should be directed to 574-372-3983 or 574-371-3009 M- F, 8 AM- 5 PM, (ET). |  | 45 | 45 | Worldwide Distribution -- USA, including NJ, TX, VA, MI IL, KT, &amp; Puerto Rico, and country of the Netherlands. | Terminated | August 29, 2013 | 532 | 45
Z-1445-2012 | 61390 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific LATITUDEÂż Patient Management System Communicator, Model 6476, Distributed by Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 55112-5798, Made in Malaysia. The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Guidant or Boston Scientific pulse generator and transfer data to a central database. |  |  | February 24, 2012 | 2012 | Boston Scientific CRM Corp | Boston Scientific LATITUDEÂż Patient Management System Communicators were shipped to the incorrect patients. | Process control | The patients were notified via phone starting on 2/24/12 that their communicator will not be able to activate successfully and that their communicator and sensors will be replaced. If the patient could not be reached via phone or voice mail, a Boston Scientific letter dated 2/29/12 was sent to the patient. The letter described the problem, explained What is a LATITUDE Communicator and What should you do. Questions were directed to Patient Services at (866) 484-3268. |  | 8 | 8 | Nationwide Distribution, including the states of: AL, CA, FL, IL, IN, MO, NJ, NY, OH, OR, TX VA, and WA. | Terminated | August 02, 2012 | 160 | 8
Z-1446-2012 | 61637 | No | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Triage TOX Drug Screen (MTD) Panel, Model # 94400. |  |  | March 12, 2010 | 2010 | Alere San Diego dba Biosite Innovacon Hemosense Inc | The recall was initiated because Biosite Inc. received an increase in customer calls regarding unexpected positive THC results when using the Triage TOX Drug Screen. | Nonconforming Material/Component | Biosite sent a "PRODUCT ADVISORY NOTICE" letter dated March 12, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter states that customers may continue to use the product. Customers were instructed to complete and fax an enclosed verification form to Technical Services at 1-858-695-7100. Customers with any questions about the information contained in the notification were instructed to contact Biosite Inc. at (888) 246-7483. |  | 1258002 | 1,258,002 kits | Worldwide Distribution-USA (nationwide) and the countries of Canada and European Union. | Terminated | April 23, 2012 | 773 | 1258002
Z-1447-2012 | 50072 | No | CIT | Aspartate amino transferase (AST/SGOT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT/AEROSET Activated Aspartate Aminotransferase |  |  | June 30, 2008 | 2008 | Abbott Laboratories Inc. | This letter is to inform you that some samples may generate A-AST results of less than (&lt;) 5 U/L or results within the reference range when actual concentration exceeds the non-Flex linearity of 1985 U/L. In other cases the following errors may be generated.-ARCHITECT cSystems: Error Code 1053 (Unable to calculate result, rate reaction linearity failure)-AEROSET: RL% Error CodeRevised P | Nonconforming Material/Component | Abbott Laboratories sent a Product Correction letter dated August 5, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:Update your assay parameters for A-AST1. Follow the steps in the appropriate attachment: - Use Attachment A for ARCHITECT cSystems - Use Attachment B for AEROSET.2. All A-AST results of less than (&lt;) 5 U/L should be evaluated by comparing to other clinically relevant assaysRetain a copy of this letter for your laboratory records.For questions regarding this recall call 626-440-0705. |  | 1646 | 1,646 | Worldwide Distribution including Germany | Terminated | April 23, 2012 | 1393 | 1646
Z-1452-2012 | 60774 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Sarns Modular Perfusion System 8000 Catalog number 148689 or 78-8067-3878-3 Temperature-Pressure Board, Cardioplegia Monitor (service part only) |  |  | March 21, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo received multiple reports of malfunctions of the Cardioplegia and arterial monitors for the Sarns Modular Perfusion System 8000. Upon review of complaint investigation, the malfunctions are caused by or could have been caused by soldering anomalies on the circuit board of the monitor. The compromised solder joints could cause intermittent or loss of display, or could cause monitor reset, po | Device Design | TERUMO sent an URGENT MEDICAL DEVICE RECALL- SAFETY ADVISORY to all consignees on March 21, 2012 via Federal Express. Terumo issued the safety advisory to remind users how to intervene if they experienced intermittent or complete loss of pressure display on the arterial monitor or cartiplegia monitor, or a related pump stop. Customers were asked to review the Medical Device Correction notice, assure that all users are aware of the notice, confirm receipt of the communication by faxing the attached Customer Response Form to 1-800-292-6551.For questions regarding this recall call 1-800-521-2818. |  | 3022 | 3,022 | Worldwide Distribution - USA including AK, AL , AR, AZ , CA, CO , CT, DC, DE , FL, GA, HI, IA , ID , IL, IN , KS, KY , LA, MA , MD, MI, MN , MO, MS , MT, NC , ND, NE, NH, NJ, NM , NY, OH , OR, PA , RI, SC , TN, TX , VA, WA , WI, WV, WY, and Puerto Rico and the countries of PERU, Philippines, REPUBLICA, SANTIAGO, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Tujunga, United Arab Emirates, Uruguay, Venezuela, Vietnam, Wooddale, and Yemen. | Terminated | July 11, 2013 | 477 | 3022
Z-1449-2012 | 60595 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Ref, S331140 Selex M2A Magnum Modular Head, 40 mm head diameter plus 3 mm neck,1 Taper, Sterile, Biomet Orthopedics, Warsaw, IN |  |  | September 13, 2011 | 2011 | Biomet, Inc. | Devices with a specification for a +6mm neck length were incorrectly labeled for a +3mm neck length specification. | Nonconforming Material/Component | Biomet contacted consignees with device inventory by telephone on September 13, 2011 requesting all associated product be quarantined immediately. An Urgent Medical Device Recall Notice was sent via U.S. mail March 23, 2012 to consignees with directions to immediately locate, discontinue use of the affected product and return to Biomet. The letter instructed consignees to carefully follow the instructions on the enclosed "FAX Back Response Form" and fax a copy of the Response Form to 574-372-1683 prior to return of product. The notification letter to consignees contained directions to confirm receipt of notice by calling 800-348-9500, extension 3983 or 3009. Questions concerning recall are directed to 574-372-3983 or 574371-3009 M-F 8 A.M. to 5 P.M. (ET). |  | 5 | 5 devices | USA Nationwide Distribution including the states of: MI, NC., OK and WI. | Terminated | February 21, 2013 | 527 | 5
Z-1448-2012 | 60865 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler Software, User Manuals, and Handi-Guides for acute and chronic peritoneal dialysis. Software versions earlier than 2.8. |  |  | March 30, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Increased risk of intraperitoneal volume (IIPV), referred to as Overfill | Software design | Fresenius Medical Care sent a Urgent Product Information letter dated March 30, 2012, and April 3, 2012, to all affected customers. The letter was sent via Certified Mail, Return Receipt Requested. The letter identified the product the problem and the action needed to be taken by the customer.Recall to occur in 2 phases:Phase I: Fresenius Medical Care notified all clinics with patients using Liberty Cyclers with software versions earlier than 2.8 that these patients are receiving updated User Manuals and Handi-Guides. These new materials will include updated descriptions of the safety warnings included with the materials provided with the release of software version 2.8. Clinics will also be informed that their patients will be receiving replacement cyclers with updated software within 12 months. These notifications via telephone and certified letter (return receipt requested). Following clinic notification, Fresenius Medical Care will contact patients operating Liberty Cyclers with software versions earlier than 2.8 by phone, informing them that they will be receiving updated User Manuals and Handi-Guides. In addition, these patients will be informed that they will be receiving replacement cyclers with updated software at some point in the future and that Fresenius Medical Care anticipates that the full timeline of cycler replacements will occur over several months. A certified letter return receipt requested will be sent to all patients after the initial phone contact containing the updated User Manuals and Handi-Guides. Phase II: Liberty Cyclers with updated software will be issued to customers using Liberty Cyclers with software versions earlier than 2.8 as inventory is available.Please complete and return the enclosed Reply Form, indicating receipt and understanding of this communicationFurther questions please call (1-800) 227-2572. |  | 10146 | 10,146 units | Worldwide Distribution -- USA (nationwide) | Completed |  |  | 10146
Z-1399-2012 | 61553 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Automatic Mobile X- Ray (AMX) series: Brivo XR285 15KW, Brivo XR285 30KW, Optima XR200 15KW, Optima XR200 30KW, Optima XR220 15KW, and Optima XR220 30KW. |  |  | December 23, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a discrepancy on the x-ray source assembly of the mobile radiographic product, related to the light field to x-ray field edge alignment. | Radiation Control for Health and Safety Act | GE Healthcare will issue an Important Electronic Product Radiation Warning letter to customers. The letter will describe the defect and related hazards, along with the safety instructions, affected product, and product correction GE Healthcare will, without charge, remedy the defect or bring the product into compliance with each applicable Federal and IEC standard in accordance with an approved plan, the details of which will be included in a subsequent communication to customers or through a GE Healthcare field engineer site visit. A GE Healthcare Service Representative will update the external tube port on the source assembly to address the issue.Questions or concerns can be directed to the Call Center at 800-437-1171 or your local GE Healthcare Service Representative. |  | 130 | 130 installed in US | USA (nationwide) | Terminated | January 16, 2014 | 755 | 130
Z-1458-2012 | 61457 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | CARE CLINER 6550 (STD)/6570 (X-LG) with Drop Arm products are labeled in part: "***CARE CLINER 6550/6570 (w/ NYLON CASTERS)***Winco ISO 9001-2008 Certified***All of our best features, plus the added versatility of drop arms***Leave it to Winco to combine the best clinical chair features in the industry with the convenience of dual drop-arms! Our drop-arm care cliners make assisting your patients easier. Our side-panels lift, making clean-ups quick and easy.***KEY FEATURES***Standard head-rest cover***Padded drop-down armrests***Dual fold down side-tables***Easy-to-clean side panels***Trendelenburg Positioning***1.800.237.3377***5516 SW 1st Lane***Ocala, Florida 334474***www.wincomfg.com***". |  |  | February 06, 2012 | 2012 | Winco Mfg., LLC | Winco Manufacturing LLC is recalling the Care Cliner, XL Care Cliner, Drop Arm Care Cliner, Elite Care Cliner, and XL Elite Care Cliner chairs with nylon casters distributed between 04/01/2009 and 08/31/2009. Certain 5" nylon casters shipped with the chairs may be defective. There is a potential of the caster's axle working itself out of the assembly under unconventional conditions. Should this | Device Design | Winco Mfg. sent an "URGENT DEVICE RECALL" letter dated February 1, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to complete the Caster Recall Replacement Form and return it to the firm's Customer Service Department via fax to 352-854-9544. Contact the Customer Service Department at 800 237-3377 for questions regarding this notice. |  | 1311 | 1,311 Casters (437 Chairs) | Worldwide Distribution-USA (nationwide) and the countries of Canada and Singapore. | Terminated | July 27, 2012 | 172 | 5132
Z-1454-2012 | 61482 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4 |  |  | January 16, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed customer complaints on Vista(R) software version 3.4 for Below Manufacturer Assay Range error flags associated with QC and patient test results that are within the assay range as defined in the Instructions For Use for the respective method. This issue is isolated to Vista(R) software version 3.4. Customers who have not made any changes to the default method configurations | Software change control | Siemens sent an Urgent Field Safety Notice dated December 19, 2011 to all affected customers alerting them to Invalid Below Manufacturer Assay Range Flags on Dimension Vista(R) Systems with Software Version 3.4. The UFSN included instructions on how to correct the situation and also included an effectiveness check form to be faxed to the firm at 302-631-8467. On January 16, 2012, the firm issued an additional UFSN alerting consignees to Invalid Above and Below Manufacturer Assay Range Flags. This notice also included directions and a Field Correction Effectiveness Check form to be completed and faxed to 302-631-8467.For technical questions customers were instruced to call 800-441-9250.For questions regarding this recall call 302-631-6299. |  | 1683 | 1,683 | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland &amp; Baltics, France, Germany, Italy, Japan, South Korea, Malaysia, Netherlands, New Zealand, Norway, Portugal, Puerto Rico, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland and United Kingdom. | Terminated | March 05, 2014 | 779 | 1683
Z-1455-2012 | 61346 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Quintex Cervical Plating System - Screws. The Quintex Cervical Plating System is intended for the treatment of cervical spinal instability. The Quintex Cervical Screws are available in three different versions: semi-constrained, constrained, and dynamic. The screws all have a head diameter of 4.0mm and vary in length from 10mm to 22mm. Rescue screws are also available with a 4.5mm head diameter and lengths of 11-19mm. |  |  | January 24, 2012 | 2012 | Aesculap Implant Systems | The firm has received complaints that the locking ring that sits in the head of the screw can come out of the screw head when engaged with the screwdriver if the screwdriver is improperly angled. | Nonconforming Material/Component | The firm initiated this recall by sending an "Important Recall Notification" letter to consignees on January 20, 2012. This letter instructed consignees not to use the product and requested that all stock be returned to the recalling firm promptly. Customers may contact 1-(866) 229-3002 for questions. |  | 891 | 891 | Nationwide Distribution, including the states of CA, CO, FL, MO, and TX. | Terminated | December 13, 2012 | 324 | 891
Z-1462-2012 | 61514 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura X-ray System. |  |  | April 02, 2012 | 2012 | Philips Healthcare Inc. | The monitor Ceiling suspension may fail, which may cause the monitor to drop. | Process control | Philips sent an "URGENT-FIELD SAFETY NOTICE" dated March 27, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides information for continued safe and proper use of the equipment. Contact your local Philips representative at 866-767-2822 for further support regarding this issue. Recall expanded 5-2013 to additional 460 units distributed worldwide. |  | 145 | 145 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Algeria, Argentina, Belgium, Brazil, Canada, Czech Rep, Chili, China, Columbia, Denmark, Egypt, France, Germany, India, Indonesia, Italy, Japan, Malaysia, Mexico, Morocco, Netherlands, Oman, Russia, Saudi Arabia, Singapore, South Korea, Spain, Thailand, United Kingdom, and Vietnam. | Terminated | May 09, 2014 | 767 | 145
Z-1456-2012 | 61473 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Certain Low-Profile, Angled Abutment, Model Number ILPAC3217. Product is labeled in part: "***Biomet 3i***4555 Riverside Drive , Palm Beach Gardens, FL 33410 USA***REF ILPAC3217***LOT 2011091179***Certain Low Profile 17 Abutment***3.4mm(D) X 2mm(H)***Abutment:; Pillar; Abutment; Abutment; Pilastro; Pillar***CE dose***Use by 2016-09***Sterile using Radiation***RX Only***Do not reuse***Consult accompanying documents***Do not re-sterilize***Do not use if package is damaged***ILPAC3217***2011091179***"P-LBLNY3P***Rev. D***". |  |  | January 10, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i initiated a recall for a small percentage of the 33mm (ILPAC3217) abutment packages that may contain the 5mm (ILPAC5217) abutment. | Material/Component Contamination | Biomet 3i , LLC sent Urgent: Medical Device Recall notices on 1/10/2012 to each customer in via e-mail, facsimile, telephone or postal mail. The letters identified the affected product and the reason for the recall. Each customer was instructed to check their respective inventory for the affected product and return any unused product to Biomet 3i at the address provided. Biomet 3i will issue replacement product at no charge. Questions or concerns should be directed to 1-800-342-5454. |  | 400 | 400 | Worldwide Distribution -- Canada, Chile, Columbia, France, Germany, Hong Kong, India, Israel, Italy, Japan, Mexico, Netherlands, Norway, Poland, Singapore, South Africa, Spain, Sweden, Switzerland, UK, US, and Vietnam. | Terminated | April 30, 2012 | 111 | 400
Z-1470-2012 | 61409 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS&amp;amp;amp;amp;amp;reg; 5600 Chemistry System, Product Code: 6802413 |  |  | March 15, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho-Clinical Diagnostics is recalling VITROS Chemistry Systems due to possible internal hard drive failure. | Nonconforming Material/Component | Ortho Clinical Diagnostics (OCD) sent an Important Product Correction Notification letter dated March 15, 2012 via FedEx overnight courier to all affected customers. Foreign affiliates were informed by email on 15 March 2012 of the issue and instructed to notify their consignees of the issue and actions. The letter identified the affected products, problem and actions to be taken. The letter instructed customers to: 1) Consider performing the backup procedure more frequently to help prevent the loss of data on your system; 2) Contact OCD Customer Technical Service representatives if a hard drive failure occurs on your system; A Field Engineer will be dispatched to replace the hard drive on your system; 3) Complete and return the Confirmation of Receipt form upon receipt of this notification; and 4) Post this notification by each VITROSÂż System in your facility or with the user documentation. For questions, refer to the Questions and Answers section. For additional questions, contact Customer Technical Services at 1-800-421-3311. |  | 417 | 417 | Worldwide Distribution | Open, Classified |  |  | 785
Z-1463-2012 | 61472 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Bolus Cord; a remote bolus cord and switch for use with GemStar 7 Therapy and Pain Management Pumps, list 13000, 13150, 13087 and 13088; Hospira, Inc., Lake Forest, IL 60045; List Number 13027; |  |  | March 28, 2012 | 2012 | Hospira Inc. | Hospira has received customer reports of bolus delivery failures when using the GemStar bolus cord. Possible failures may be a result of cracked or damaged bolus cords which may result in an open or short circuit of the Bolus Cord. | Nonconforming Material/Component | Hospira, sent an "URGENT DEVICE FIELD CORRECTION RECALL" letter dated March 28, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to take care when removing the bolus cord from the base of the GemStar unit, inspect the cord for damage, discontinue using any damaged cords, and to call Hospira at 1-800-241-4002 for replacement. A Reply Form was enclosed for customers to complete and return via fax to 1-866-912-7085 or e-mail to: Hospira2627@stericycle.com. Contact Hospira Customer Care at 1-877-946-7747 for technical assistance. |  | 87095 | 47,127 cords - U.S. and 39,968 cords - ROW | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Guatemala, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Luxembourg, Malaysia, Malta, Mexico, Netherlands, New Zealand, Nola, Norway, Oman, Philippines, Portugal, Qatar, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay and Venezuela. | Open, Classified |  |  | 87095
Z-1465-2012 | 61510 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MEDITECH Anatomical Pathology Software Client Releases- |  |  | March 30, 2012 | 2012 | Medical Information Technology, Inc. | Meditech Anatomical Pathology software is being recalled because data on a pathology specimen may be inadvertently deleted. | Software design | MEDITECH sent a notification beginning on January 19, 2012 to customers who are LIVE with MEDITECH Client Server Anatomical Pathology 5.6 and higher. This notification was made via e-mailed task updates that informed customers of the software issues and the code corrections that are available. |  | 280 | 280 | Worldwide Distribution - USA Nationwide and the countries of Bahamas, Canada and United Kingdom. | Terminated | August 13, 2014 | 866 | 280
Z-1466-2012 | 61516 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage CardioProfiler Panel, PN 97100CP. The test is used as an aid in the diagnosis of myocardial infarction (injury), an aid in the diagnosis and assessment of severity of congestive heart failure, an aid in the risk stratification of patients with heart failure, and as an aid in the risk stratification of patients with acute coronary syndromes. |  |  | March 05, 2012 | 2012 | Alere San Diego | Alere San Diego, Inc. is initiating a voluntary recall of the Alere Triage CardioProfiler Panel PN 9700CP, Lot W49569 and the Alere Triage Profiler SOB Panel PN 97300, Lot W48990. Alere San Diego, Inc. has determined that these lots have an increased frequency of Troponin I results &gt;0.05ng/mL for samples which are found to be below 0.05ng/mL upon additional testing. | Process control | A recall communication was initiated on 3/15/12 with Alere San Diego, Inc. forwarding a "Urgent Medical Device Recall" letter to all their customers who purchased the Alere Triage CardioProfiler Panel PN 97100CP, Lot W49569 and Alere Triage Profiler SOB Panel PN 97300, Lot W48990. The letter informed the customers about the problems identified and the actions to be taken. Customers with any questions about the information contained in this notification should contact Alere San Diego, Inc. at (877) 308-8287. |  | 1080 | 1080 kits | Worldwide Distribution, including Nationwide (US) and the countries of Brazil, China, India and Kuwait. | Terminated | December 04, 2012 | 274 | 2396
Z-1464-2012 | 61528 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Systems with Eleva software version 2.x and stitching option |  |  | April 03, 2012 | 2012 | Philips Healthcare Inc. | Potential misdiagnosis due to improperly automatically stitched images, improper image ruler handling. | Labeling design | The firm, Philips Healthcare Imaging, issued an "URGENT-FIELD SAFETY NOTICE" dated March 30, 2012 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to follow the instructions carefully and to not use the Philips Image Ruler for measurements. The firm plans to supply an Addendum to the Instructions For Use and a Philips Service Engineer will contact the customer as soon as the Field Action Kit is ready to be implemented.If you need any further information or support concerning this issue, please contact your local Philips representative or call 1-866-767-2822. |  | 606 | 606 units | Worldwide distribution: USA (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech. Rep, Denmark, Finland, France, Germany, Guadeloupe, India, Indonesia, Iran, Iraq, Italy, Japan, Malta, Mexico, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Salkowski, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Ver.Arab.Emir. | Terminated | March 13, 2014 | 709 | 606
Z-1475-2012 | 60610 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | UMP Personal Sentry Deluxe Pull String Monitor, Model #91650. |  |  | December 06, 2011 | 2011 | Stanley Security Solutions, Inc. | Monitor gives false alarms | Nonconforming Material/Component | Stanley Healthcare Solutions sent an "IMPORTANT RECALL NOTICE" dated December 8, 2011. The letter identifies the product, problem, and the actions to be taken by the customers. The letter instructs customers to return affected product through Customer Service and to have the appropriate authority sign the Acknowledgement Form. Customers with questions can call UMP Customer Service at 1-800-206-1045. |  | 592 | 592 units | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | September 24, 2013 | 658 | 592
Z-1473-2012 | 61535 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | ASP 6025 Vacuum Tissue Processor; an automated modular tissue processor; Leica Biosystems Nussloch GmbH, Heidelberger Str. 17-19, D-69226 Nussloch, Germany; The intended use of this device is as an automated system used to process tissue specimens for examination through fixation, dehydration, and infiltration. |  |  | March 30, 2012 | 2012 | Leica Microsystems, Inc. | There is a potential malfunction with the device in autorotation mode, which can lead to incorrect reagent handling of the device. The end result may lead to damage or to loss of tissue specimens. | Software design | Leica Microsystems sent a Urgent Medical device Recall letter dated March 30, 2012, to all affected customers in the U.S. on the same date. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed of the potential malfunction of the device in the autorotation mode, which could lead to damage or loss of tissue specimens. The letter provided instructions on what the customer should do with the recalled product in the interim before the software fix can be installed. They were instructed to follow the steps below: 1. Remove and dispose all reagents including Paraffin and clean all bottles and Paraffin bathes.2. Conduct a "Smart Clean" (IFU, chapter 6.1.4).3. In Submenu "System Setup", here at "Program options", respectively "Station sequence", customer shall switch from "Concentration" to "Auto by RMS".4. Prestored Xylene protocols shall be used instead of autorotation protocols (IFU, chapter 5.6.2 "Time - optimized Xylene infiltration programs"). Here you find the five validated programs which need to be copied to Favorite programs (IFU, chapter 5.5.3).5. Use the Reagents, bottle set up from "Time optimized Xylene programs" (IFU, chapter 5.6.2., page 134) and implement them into submenu "Reagents" (IFU, chapter 5.2. following pages).6. In submenu "The RMS system adjusting warning thresholds rules" (IFU, chapter 5.2.2) the validated reagent thresholds need to be implemented into Submenu "Reagents" (IFU, chapter 5.2).7. Current Reagent management values are visible within Submenu "Reagent Status" (IFU, chapter 5.2.3)The customers were also informed that their Leica Microsystems representative will them to make the necessary arrangements to complete a software upgrade at no charge. The customers were also requested to sign and date the enclosed acknowledgement form and fax it back to Leica at 847-236-3747 or e-mail it to LMGRA@leica-microsystems.com within 10 working da |  | 67 | 27 units - U.S., 40 units - ROW | Worldwide Distribution -- USA (nationwide)including states of: California, Florida, Georgia, Kentucky, Massachusetts, Minnesota, Mississippi, Nebraska, New York, North Carolina, the District of Columbia; and countries of: Australia, Belgium, France, Germany, China/Hong Kong, Italy, Japan, the Netherlands, Singapore, Switzerland, Turkey and United Kingdom. | Terminated | October 10, 2012 | 194 | 67
Z-1474-2012 | 61328 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips eCareManager System, Model 865325, software versions 3.7, 3.7.1 and 3.8. |  |  | February 22, 2012 | 2012 | Visicu, Inc. | Three versions of software for the eCareManager failed to display anticoagulants on the medications screen. The problem arose due to an unexpected change in the National Drug Data File. | Software change control | Philips notified all affected customers via a preliminary email message advising them of the issue. An URGENT - Field Safety Notice follow-up letter was mailed to all customers on February 22, 2012. This letter identified the affected product, described the problem and circumstances in which it can occur; recommend actions the customer should take to prevent risks for patients, and list the actions planned by Philips to correct the problem. For questions contact Philips eICU Customer Support 877-374-2872. |  | 12 | 12 | Nationwide Distribution (USA) including states of MA, MO, WI, AR, AZ, PA, SD, VA, and IN. | Terminated | July 06, 2012 | 135 | 12
Z-1476-2012 | 60698 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Micro-Tech Informer Plus Silver Monitor, Model 81830. The firm name on the label is Stanley Security Solutions, Inc., Lincoln, NE. |  |  | December 13, 2011 | 2011 | Stanley Security Solutions, Inc. | Connection between the call cord and the receptacle may not make a correct connection resulting in a continuous or intermittent (false) signal to the nurse call system | Component design/selection | Stanley Healthcare Solutions, Inc. contacted affected customers via phone and email on December 13, 2011. A follow-up certified letter was sent on December 14, 2011. Customers were additionally sent and "Important Recall Notice letter on December 16, 2011 after errors in the original letters were detected. The letter identified the affected product, problem, actions to be taken, and included a reply form for the customer to complete and return. The recalling firm requested their customers to stop using the monitor and contact them for return of the unit. Contact information (phone, fax, &amp; email) was included. For question contact Micro-Tech Customer Service at 1-800-206-1045. |  | 824 | 824 units | Worldwide Distribution - USA Nationwide including the states of CT, FL, KY, MN, NC, NE, NH, NJ, NY, PA, TX and the country of Canada (5 units). | Terminated | September 24, 2013 | 651 | 824
Z-1478-2012 | 61658 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | REF 28234 Troch Nail, Long 11MM, Right 34 CM |  |  | March 30, 2012 | 2012 | Biomet, Inc. | The Trochanteric Nail, lot 486130, wasmislabeled as a right nail orientation but the contents are actually a left nail orientation | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall Notice dated April 18, 2012 to all affected consignees. The letter identified the affected product, reason for the recall including risks involved and actions to be taken. The notice provided directions to immediately locate and return the affected product. Receipt Confirmation was requested by calling 800-348-9500 Ext 3009 or 3983. FAX Back Response Form should be completed and returned to 574-372-1683. Questions should be direct to 574-371-3009 or 574-372-3983, M-F, 8 AM to 5 PM. |  | 30 | 30 units in total | Worldwide Distribution - Nationwide (USA) | Terminated | March 08, 2013 | 343 | 30
Z-1477-2012 | 61635 | No | KPJ | Topical oxygen chamber for extremities. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | EPIFLO; Transdermal continuous oxygen therapy device; single use; disposable medical device. Model Number EPIFLO-15. Mfr.By: Neogenix, LLC, dba Ogenix, 3401 Enterprise Pkwy, Suite 340, Beachwood, OH 44122. |  |  | April 16, 2012 | 2012 | Ogenix Corporation | Ogenix Initiated a voluntary recall for the EPIFLO-15 devices due to a mislabeling of the expiry date for this lot. Instead of March 2013, this lot carried May 2013 as the expiry date. | Labeling mix-ups | Ogenix sent a letter dated March 16, 2012 to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. Customers were instructed to confirm the status of the devices and ship product with the recalled lot back to Ogenix and complete the form provided with the letter. |  | 38 | 38 units | Worldwide Distribution-USA (nationwide) including the states of FL, MN, OH and the county of Turkey. | Terminated | February 19, 2013 | 309 | 38
Z-1479-2012 | 61604 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew 560P High Definition Camera Control Unit |  |  | March 29, 2012 | 2012 | Smith &amp; Nephew, Inc. Endoscopy Division | May experience intermittent loss of videodue to a defective board component | Nonconforming Material/Component | Smith &amp; Nephew Endoscopy Division sent an "URGENT PRODUCT RECALL 1st NOTIFICATION" letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory and to complete and return the bottom portion of the letter via fax at 405-917-8650. Contact the firm at 405-917-8556 for questions regarding this notice. |  | 239 | 239 units | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, Canada, Denmark, Finland, France, Germany, Italy, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, and UK Asia-pacific: Australia, China, Hong Kong, India, Japan, Singapore, South Korea, and Thailand, Mexico, Puerto Rico, South Africa, Turkey, and Venezuela. | Terminated | January 13, 2016 | 1385 | 239
Z-1480-2012 | 61653 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Churchill Medical Systems Winged Infusion Set, 22Ga, 1", |  |  | April 16, 2012 | 2012 | Churchill Medical Systems, Inc. | Mislabeled: MBS-2210 code on case label -22Ga Winged Infusion Set, ; however, the product label states 20Ga MBS-2010 code listed on devicelabel. | Labeling False and Misleading | The firm, VYGON, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 16, 2012 to its customers via overnight, courier service. The letter described the product, problem and actions to be taken. The customers were instructed to check their stock for the products included within the scope of the recall; count and document their affected inventory; contact Vygon's Customer Service Department at 1-800-473-5414 for return merchandise authorization number; return the affected product; complete and immediately return the Acknowledgement and Inventory Return Form via fax to 215-672-6740 and send the original form with their shipment . Distributors were instructed to communicate this recall to their customers and provide them with the Acknowledgement and Inventory Return Form for completion and return.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by e-mail at customerservice@yygonus.com. |  | 552 | 552 units | Worldwide distribution: USA (nationwide) including states of: IN, KS, KY, NY,MA, MI , and TX; and country of: Canada. | Terminated | March 28, 2013 | 346 | 552
Z-1485-2012 | 61602 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | VALOR(TM) TTC FUSION, INSTRUMENT KIT, REF 4150-KIT1, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002. |  |  | April 05, 2012 | 2012 | Wright Medical Technology Inc | The possibility exists that the handle of the targeting guide may rotate and/or break under normal usage. | Device Design | Wright Medical Technology, Inc. sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated April 2, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to check their stock and to return the affected product to the firm. Additionally, a Product Recall Verification Form was enclosed for customers to complete and return via fax to 901-867-7401. Contact the firm at 901-867-4324 for questions regarding this notice. |  | 133 | 133 units | Worldwide Distribution-USA (nationwide) and the country of Australia. | Terminated | January 28, 2013 | 298 | 133
Z-1481-2012 | 61411 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coloplast, Restorelle DirectFix, Posterior Polypropylene Mesh, Catalogue Number 501460, Rx Only, Manufacturer, Coloplast A/S 3050 Denmark |  |  | March 02, 2012 | 2012 | Coloplast Manufacturing US, LLC | Coloplast is conducting a field correction based on the conclusion that two lots of Restorelle DirectFix Posterior mesh were mislabeled. The incorrect size was printed on the patient label. Specifically the patient label depicts the size as 16x11 cm, however, the correct size is actually 14x15 cm. It is important to note that both the product and box labels were correct. | Error in labeling | Coloplast sent a notification letter dated March 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were informed their coloplast representative is providing replacement device (s) and new patient labels. For further questions please call 1 (800) 258-3476. |  | 134 | 134 | USA ( nationwide) including the states of AL, AZ, AR, CA, CO, FL, GA, ID, IL, KY, MD, MS, MO, NY, PA, TN, VA,WA and WI. | Terminated | October 11, 2012 | 223 | 134
Z-1488-2012 | 61464 | No | EQN | Nasopharyngoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | BOMImed Reusable Integrated Fiber Optic Laryngoscope Blades, MacIntosh and Miller styles - all sizes. Instrument is used to aid with intubation. |  |  | December 08, 2010 | 2010 | Bomimed Inc | On 9/1/2010 BOMImed, Inc. Winnipeg, MB, Canada initiated a recall of their BOMImed Integrated Reusable Laryngoscope Blades (Macintosh &amp; Miller styles - all sizes), Model #OL-32D0 to OL-32D4 &amp; OL-32E00, OL-32E0, OL-32E1, OL-32E2, OL-32E3, OL-32E4. | Process design | BOMImed sent a "Medical Device Safety Alert &amp; Correction Action" letter by Fax and Mail on 11/25/2010 to the consignee. The product issue was described and recommended actions to mitigate risk were provided. Contact number is 1-800-667-6276 ext. 234. |  | 30 | 30 units | US Distribution to the state of California. | Terminated | May 08, 2012 | 517 | 30
Z-1486-2012 | 61545 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens-branded linear accelerator with a 550 TxT Treatment Table with a serial number &lt;_ 1894. Siemens brand ARTISTE, ONCOR and PRIMUS Linear Accelerator systems include the 550 TxT" Treatment Table. |  |  | April 03, 2012 | 2012 | Siemens Medical Solutions USA, Inc | SIEMENS has received reports where a dropping down of the 550 TxT" Treatment Tablewas observed while the table was being moved downwards during operation by means ofthe hand control. | Software design | Siemens sent Customer Information letters to all affected customers beginning April 3, 2012 via Siemens Field Service Engineers. The letter identified the affected product, what is the issue and what has Siemens done to address this issue. The letter included an Update Instructions TH006/12/S - TxT Table Software Update VD00A. SIEMENS has arranged for a modification of the affected system. A SIEMENS Service Representative will ensure that the new software version VD00A is installed on affected system. The letter instructs customer to please include this Customer Information in their Digital Linear Accelerator System Owner Manual chapter Safety Advisory Letters where it should remain. |  | 624 | 624 active devices | Worldwide Distribution - USA (nationwide) including P.R and the countries of: Africa, Angola, Argentina, Australia, Austria, Belgium, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Italy, Japan, Kuwait, Lebanon, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Republic of Korea, Russian Federation, Saudi Arabia, Serbia, South Thailand, Sri Lanka, Spain, Switzerland, Trinidad, Tobago, Turkey, Ukraine and United Kingdom, . | Terminated | March 14, 2013 | 345 | 624
Z-1489-2012 | 61613 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | EG-3870UTK, Pentax Ultrasound Video Gastroscope. |  |  | March 14, 2012 | 2012 | Pentax Medical Company | Endoscope ultrasound balloon evacuation tube may become clogged with patient body fluid if users fail to properly follow instructions for pre-procedure check, operating the suction valve during procedure and pre-cleaning procedure. | Equipment maintenance | Pentax sent an "URGENT MEDICAL DEVICE RECALL" letter on March 14, 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. A response form was included for customers to complete and return to the firm. Contact the firm at 800-431-5880 ext. 2064 for questions regarding this notice. |  | 194 | 194 units | Nationwide Distribution. | Terminated | July 21, 2015 | 1224 | 194
Z-1483-2012 | 61656 | Yes | NIK |  |  | Cardiovascular | PMA | Cardiac Resynchronization Therapy (CRT) Leads |  |  | April 03, 2012 | 2012 | St Jude Medical CRMD | Visual observations of externalized conductors due to abrasion of the silicone insulation in the distal portion of these QuickSite and QuickFlex leads. | Device Design | A recall communication was initiated on 4/3/12 with St. Jude Medical forwarding a Medical Device Advisory to all doctors who have received the St. Jude QuickSite and QuickFlex Left Ventricular CRT leads models 1056T, 1058T, 1156T, and 1158T. The letter provides the doctors with the problems identified and actions to be taken. Doctors with any questions or concerns are instructed to contact their local St. Jude Medical representative or their Technical Services Department at (800) 722-3774.A Physician Device Advisory Notice for the QuickSite and QuickFlex Left Ventricular CRT Leads was issued on 4/3/12 to physicians who have received the recalled products. The notice provides the physicians with the problem identified and actions to be taken. Physicians are instructed to return any packaged product in hospital inventory to SJM Sales Representatives. For implanted product, due to the nature of the issue, externalized conductors may be observed with no electrical or performance anomalies. Therefore, prophylactic explant is not recommended. It is recommended that physicians continue to monitor their patients' implanted system at regularly scheduled intervals. A Medical Advisory and Inventory Management Information letter was sent on 4/3/12 to hospital administrators to notify them of the problems identified and the actions to be taken. |  | 101000 | 101,000 units (59,000 units in US) | Worldwide Distribution -- Algeria, Argentina, Australia, Austria, Bahrain, Balkan, Baltic, Belarus, Belgium, Bolivia, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Malta, Mauritius, Montenegro, Morocco, Netherlands, New Caledonia, New Zealand, Norway, Oman, Paraguay, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Senegal, Serbia, Serbia Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Tunisia, Turkey, UK, United Arab Emirates, US, Venezuela, and Yemen. | Terminated | July 17, 2012 | 105 | 101000
Z-1492-2012 | 61413 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Optima MR450W with GEM. |  |  | April 12, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare is conducting a recall on Optima MR450W with GEM. During scanning of a patient, if an improperly torqued gradient connection exists, it may cause the gradient lead to produce excessive heat. If a second and compounding issue of improperly routed Rear Pedestal Air Supply Hose is also present, routed in such a way to cause the hose to come in contact with the overheated gradient lead | Nonconforming Material/Component | GE Healthcare is not sending a Communication Letter to customers. There is no plan to contact customers with an affected Optima MR450W with GEM system. A GE Healthcare Field Service Engineer will be dispatched to sites to inspect and correct all potentially affected systems. As an effectiveness check, GE Healthcare will track all service requests for the field action to completion. For any questions or additional information, call 262-548-2608. |  | 41 | 41 | Worldwide Distribution - USA (Nationwide) including the states of :CO, FL, IA, MA ,MN, NJ, NY, NC, OH, TX. and the countries of BRAZIL., BRITAIN, FINLAND, FRANCE, GREAT GERMANY, ITALY, SAUDI ARABIA, SWEDEN, SPAIN, KUWAIT, and SWIITZERLAND. | Terminated | August 09, 2012 | 119 | 41
Z-1491-2012 | 61680 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeShield Primary Plumset, CLAVE port, CLAVE Y-Site, 104 inch, 15 drops/mL, 48 sets per case; a sterile Rx device; Hospira Inc, Lake Forest, IL USA. Made in Costa Rica; product 120300412, list number 12030-12. |  |  | April 25, 2012 | 2012 | Hospira Inc. | The Lifeshield Primary Plumsets may cause a cassette test failure alarm when the set is loaded onto the infusion pump. A set which results in a cassette test failure alarm may result in delay in therapy, requiring medical intervention. | Process control | Hospira, sent an "URGENT DEVICE RECALL" letter dated April 25, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to check their inventory and immediately quarantine any affected product to return for credit. A Reply Form was enclosed for customers to complete and return via fax to 1-888-386-2073 or e-mail to: Hospira8503@stericycle.com. Contact Hospira Customer Care at 1-877-946-7747 for information regarding product availability. |  | 145875 | 145,875 sets | Worldwide Distribution - USA (nationwide) including the states of: Alabama, Colorado, Georgia, Illinois, Indiana, Iowa, Kentucky, Maryland, Massachusetts, Michigan, Minnesota, Missouri, North Carolina, North Dakota, Oklahoma, Pennsylvania, Wisconsin and the country of Australia. | Open, Classified |  |  | 145875
Z-1487-2012 | 53340 | No | FFX | Gastrointestinal motility monitoring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Insight rev H software. |  |  | February 23, 2009 | 2009 | Sandhill Scientific, Inc | Software error may produce incorrect esophageal contraction readings. | Software change control | Sandhill Scientific sent a Urgent Product Recall letter dated February 23, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were informed All Suite revision H, BioVIEW Analysis Software rev 5.5.1 are affected. Replacement of the Suite revision H, BioVIEW Analysis Software rev 5.5.1, will be provided at no cost to you. Sandhill Scientific will pay for shipping/handling charges associated with replacement of this software.We anticipate that by the first week of March, we will be able to provide you with the Software that resolves this problem.For further questions please call toll free 800-468-4556. |  | 60 | 60 programs | Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, FL, IA, IL, MA, MD, MI, MN, NC, ND, NE, NM, NY, OH, OR, PA, SC, SD, TX, UT, VA and WA and the countries of Belgium, Canada, Malaysia, Mexico, New Zealand, South-Korea, Switzerland and United Kingdom. | Terminated | May 07, 2012 | 1169 | 60
Z-1493-2012 | 61580 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE or ONCOR Linear Accelerators in combination with 160MLC (Multi Leaf Collimator). Linear Accelerators that may be affected are: ARTISTE MV System; ONCOR Avant Garde ONCOR Impression ONCOR Impression Plus ONCOR Expression. Manufactured by Siemens AG, Medical Solutions, Kemnath, Germany; Registered by Siemens Medical Solutions, Concord, CA. ARTISTE and ONCOR family of Linear Accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. |  |  | April 03, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety risk related to the behavior of the 160 MLC when using the "Motion Stop" button followed by "Reset" several times during the treatment of one beam or one segment (issue 1). Additionally, there may be a potential for collision if the operator does not ensure clearance of the extended flat panel positioner for all gantry rotation angles (issue 2). | Labeling design | Siemens sent a Safety Advisory Notice with Update to all affected customers. The notice informed the customers of the risk of pressing the Motion Stop button repeatedly during treatment. Work Steps were provided for customers to follow. |  |  |  | Worldwide Distribution - USA including the states of CA, DE, FL, IL, KY, LA, MA, MI, NE, NJ, NY, OH, PA, SD, TN, UT, WI, and WV | Terminated | July 25, 2012 | 113
Z-1496-2012 | 61709 | No | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Doctors Research Group Kryptonite Bone Cement packaged in 5cc and 10 cc Kits: |  |  | April 06, 2012 | 2012 | Doctor's Research Group, Inc. | Bone cement strength and stiffness fall below specification at body temperature | Component design/selection | Doctors Research Group (DRG) issued a Voluntary Product Recall Notification letter dated April 2012 to all affected customers via email or Federal Express. The letter identified the affected product, problem and actions to be taken. Customers were instructed to identify, segregate and return any affected products in their possession. Customers were asked to fill out and return the enclosed verification form Distributors were requested to notify their customers. For questions contact DRG at 203-262-9335. |  | 3100 | 3100 kits | USA Nationwide Distribution | Terminated | July 14, 2015 | 1194 | 3100
Z-1499-2012 | 61610 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Dried Overnight Gram Negative Panels; |  |  | June 16, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Three lots were confirmed to not meet performance claims due to random non-biological contamination which is only apparent (precipated) after inoculation of the panels | Process control | Customer Notification letters were sent to all customers of the three MicroScan panels by traceable mail. The letter identified the affected product and asked customers to discard any remaining inventory in addition to contacting customer service for replacement product. The letter also recommended that customers consider the need to review previous test results, conduct patient follow-up, and or repeat tests as needed. Customers were to complete and return the customer reply card to confirm that they received and understood the recall communication. |  | 683 | 683 boxes (20 panels per box) | Worldwide Distribution -- US, Turkey, Poland, and Great Britain. | Terminated | May 08, 2012 | 327 | 2050
Z-1500-2012 | 61646 | No | LJY | Mumps virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | AtheNA Multi-Lyte MMV Test System- A93111G |  |  | March 08, 2012 | 2012 | Zeus Scientific, Inc. | Select lots of AtheNA Multi-Lyte MMV Test System have lost reactivity and may produce false positive results in seronegative patients who have not been previously affected with or vaccinated against Mumps. | Nonconforming Material/Component | The firm, Zeus Scientific, sent an "URGENT NOTICE; Voluntary Product Recall" letter dated March 8, 2012 their customers via Fed Ex. The letter described the product, problem and actions to be taken. The customers were instructed to remove and discard any remaining inventory of the listed product; if you are not the end-user and/or you have distributed these products to customers, notify them accordingly; and complete, sign and return the verification notice via fax to: 908-526-2058 or email a PDF to support@zeusscientific.com.If you have any questions, call 908-526-3744 or 800-286-2111. |  | 180 | 180 kits | Nationwide distribution. | Terminated | July 03, 2013 | 482 | 180
Z-1495-2012 | 61693 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | True Beam; |  |  | April 09, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified in TrueBeam when using respiratory gating with patients in the prone position. The gating curve is inverted when a prone patient plan is loaded into TrueBeam. | Device Design | Varian sent an Urgent Medical Device Correction - Urgent Field Safety Notice dated April 13, 2012 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter identified the affected product, description of problem, details and recommended user actions. Varian will provide a software version that will mitigate the issue described in the letter and Varian Customer Support will notify each affected facility when the software is available. The letter instructs customers to retained a copy of this notice along with the most current product labeling. For future reference, this document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com. For questions contact Varian Oncology Help Desk at 1-888-827-4265 US and Canada. |  | 423 | 207 devices ***12-12-12 AMENDED TO: 216 UNITS*** | Worldwide Distribution - USA Nationwide and the countries ofAustralia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Italy, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Thailand, Turkey and United Kingdom | Terminated | April 16, 2014 | 737 | 423
Z-1494-2012 | 61578 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens-branded ONCOR" or PRIMUS" linear accelerator with an OPTIVUE" or MVision" imaging system; |  |  | April 04, 2012 | 2012 | Siemens Medical Solutions USA, Inc | The firm became aware of a potential issue with Siemens-branded ONCOR or PRIMUS linear accelerator with an OPTIVUE or MVision imaging system that has the potential collision problem of the imaging panel if the panel does not retract before the gantry rotates; this may result in patient injury. | Process control | Siemens sent an Update Instruction containing the Customer Information letters beginning April 4, 2012 delivered by Siemens customer Service Group or by certified mail to all affected customers. The letter identified the affected product, what is the issue and what has Siemens done to address this issue. The letter instructs customer to include this Customer Information in their Digital Linear Accelerator System Owner Manual chapter Safety Advisory Letters where it should remain. |  | 73 | 73 units | Worldwide Distribution - USA Nationwide | Terminated | September 17, 2012 | 166 | 73
Z-1504-2012 | 53828 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Caridian BCT COBE Spectra System Blood Warmer, Catalog number 950000000. |  |  | October 26, 2011 | 2011 | CaridianBCT, Inc. | Potential for air entry into the fluid pathway of the COBE Spectra due to a leak at the return luer connection to an elevated bloodwarmer when directions for use are not followed. | Device Design | Caridian BCT sent a letter to customers dated June 2011. It was a Field labeling addendum to the COBE Spectra Apheresis System Operator's Manual stating: "When connecting a blood warmer tubing set to the return line, ensure that the tubing connection is tight. Put the luer connection no higher than 20 in (50 cm) above the return access to prevent the possibility of air entering the tubing." This correction was distributed to all consignees and is included with the current version of the owner's manual. |  | 2459 | 2459 machines | Worldwide distribution including USA, Canada, Algeria, Argentina, Australia, Austria, Bahamas, Belgium, Brazil, Chile, China, Colombia, Czech Republic, Denmark, Netherlands, UK, Hong Kong, Hungary, India, Italy, Iran, Ireland, Israel, Jordan, Korea, Kuwait, Lebanon, Libya, Luxembourg, Egypt, Finland, France, Germany, Japan, Malaysia, Mexico, Norway, Oman, Pakistan, Philippines, Poland, Portugal, Romania, Russia, Spain, Saudi Arabia, South Africa, Sweden, Switzerland, Taiwan, Thailand, Ukraine, Uruguay, Vatican, Vietnam, and Greece. | Terminated | May 11, 2012 | 198 | 2459
Z-1503-2012 | 61384 | No | GHM | Sickle cell test. | Hematology | Hematology | 510(k) | Sickle- Scan Hemoglobin S Cell Reagent Set, Catalog No. 13-2101 100 Tests For the Determination of Hemoglobin S Two bottles Sickle Cell Solubility Buffer, 100 ml , 2 vials Sickle Cell Solubility Powder, Size 100 tests, Store 20-30 C, Distributed by True-Medix Diagnostic Brooklyn, NY 11228, 800-952-3349. |  |  | February 23, 2012 | 2012 | MichClone Associates, Inc. | Two complaints concerning the powder vial included in the 100 test kit has been observed producing pressure while sealed. When the powders are opened, the pressure from within the vial releases a small amount of powder into the air.This powder. if released near the face and inhaled can cause minor irritation resulting in sneezing and nasalcongestion. Investigation found no affect on the perf | Nonconforming Material/Component | Michclone Associates, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 15, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to identify affected lots, notify sub accounts, and complete and return the response form to arrange return of unused product and replacement . Forms can be returned by FAX 248(307)-9498, mail: Michclone Associates Inc. 680 Ajax Drive, Madison Heights, MI 48118 or Email: toben@greatlakesdiag.com. Questions and concerns should be addressed to Technical Support Department at 248-307-9494. |  | 400 | 400 kits | Nationwide Distribution. | Terminated | December 13, 2012 | 294 | 770
Z-1505-2012 | 61675 | No | EBZ | Ultraviolet activator for polymerization. | Dental | Dental | 510(k) | DISPOSA-SHIELD. The Disposable Curing Shield For Your Curing Light. The disposable Disposa-Shield are sold in different sizes from 6 mm, 8 mm, 9 mm, 11 mm and 13 mm. |  |  | April 23, 2012 | 2012 | O'Ryan Industries Inc | Disposa-Shield Disposable Curing Shield is recalled due to lacking 510K, PMA or IDE, and inadequate quality system. | No Marketing Application | The firm, O'Ryan Industries Inc., began to contact consignees by telephone and sent email on April 25, 2012. O'Ryan sent a recall notification letter dated May 7, 2012 to their customers through the Purchasing Agents and Order Department Personnel. The letter described the product, problem, and actions to be taken by the customers. Customers were advised to return the Disposa-Shield back to O'Ryan Industries.Customers can call O'Ryan Industries at 800-426-4311 or fax to 360-892-6742. |  | 11376 | 11,376 units | Nationwide distribution: USA including states of NC, MN, OR and WI. | Terminated | August 10, 2012 | 109 | 11376
Z-1520-2012 | 61577 | No | LGC | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | Product is labeled in part: "***is immunosimplicity***HSV 1&amp;2 IgM***Enzyme Immunoassay Test Kit***For the Detection of HSV 1&amp;2 IgM antibodies***Catalog No: 720-350***Exp. Date:***Contains Sufficient for 96 Tests***C - Corrosive Harmful***Lot No: ***For In Vitro Diagnostic Use***Caution, consult accompanying documents.***Manufacturer***Diamedix***A Subsidiary of IVAX Diagnostics, Inc. MIAMI, FL 33127 USA-1-800-327-4565***Biohazard***" |  |  | January 16, 2012 | 2012 | Diamedix Corporation | On 01/16/20112 Diamedix initiated a recall of IS-HSV 1&amp;2 IgM Test Kit, Catalog # 720-350, Lot 60701, Exp. June 30, 2012. The Positive Control for this lot is Out of Specification. | Nonconforming Material/Component | Diamedix sent Recall notification letters dated January 16, 2012, to all affected customers by Certified Mail. Consignees were asked to review their inventory and then contact Diamedix to request replacement product. Customers were instructed to call 305-324-2314 with questions. For questions regarding this recall call 305-324-2560. |  | 442 | 442 kits | Nationwide Distribution and Puerto Rico. | Terminated | May 18, 2012 | 123 | 442
Z-1508-2012 | 61713 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Product is labeled in part: Inner Label "***SMART Control Transhepatic Biliary Stent***REF Cat. No.***Lot No.***Use By***Cordis S.M.A.R.T. CONTROL Nitinol Stent Transhepatic Biliary System***Assembled in Mexico***Cordis Corporation 14201 North West 60th Avenue Miami Lakes, Florida 33014, USA***Cordis S.M.A.R.T. CONTROL BILIARY***Cordis a Johnson &amp; Johnson company***Endovascular***" |  |  | April 09, 2012 | 2012 | Cordis Corporation | Cordis Corporation is recalling 32 lots of Cordis S.M.A.R.T. Nitinol Stent Transhepatic Biliary System and S.M.A.R.T. Control Nitinol Stent Transhepatic Biliary System. There is a potential for a sterility breach due to holes or tears in the Tyvek packaging. | Packaging process control | Cordis Corp sent a " Urgent Medical Device Recall (Removal) letter dated April 9, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A Response Form was provided with each notification. Customers were asked to complete and return Response Form. The customers were instructed to :" Immediately identify and set aside all product listed above in a manner thatensures the affected product will not be used." Review, complete, sign and return the enclosed Acknowledgement Formdirectly to Cordis at the fax number on the form." Either return any affected product with reference to the RA# on the form, orcontact your local sales representative, or contact Customer Service (see Available Assistance section) to facilitate return of the affected product.Replacement product will be provided in a short time frame. Credit will beprovided for customer-owned inventory upon request." Pass on this notice to anyone in your facility that needs to be informed." If any product listed below has been forwarded to another facility, contact thatfacility to arrange return." Maintain awareness of this notice until all product listed above has beenreturned to Cordis.For questions related to the issue or product replacement, please contactyour local sales representative, or contact the special Customer Servicenumber (800) 551-7683, Monday through Friday from 8:30 AM to 8:00 PMEDT. |  | 278 | 278 units | Worldwide Distribution-- USA (nationwide) and the country Canada. | Terminated | December 24, 2013 | 624 | 1262
Z-1521-2012 | 61685 | No | JXE | Nerve conduction velocity measurement device. | Neurology | Neurology | 510(k) | Ultra-DriveTool, 6 MM Plug Puller, REF 423868. Recommended Power Setting 70 For Single Use Only, Sterile, Biomet Orthopedics, Warsaw, IN 46581 |  |  | April 18, 2012 | 2012 | Biomet, Inc. | Incorrect product picture label was on the product box. | Process design | Biomet Orthopedics sent an Urgent Medical Device Recall Notice dated April 18,2012, to all affected customers. The letter identified the reason for the recall including directions to immediately locate, discontinue use of the device, and return product to Biomet, 56 East Bell Drive, Warsaw, IN 46580. Consignees were asked to confirm receipt of the notice by calling 1-800-348-9500 extension 3983 or 3009. A FAX Back Response Form was to be filled out and sent within 3 days of receipt to 574-372-1683. Questions related to the notice should be directed to 474-372-3983 or 574-371-3009, Monday- Friday 8 a.m. to 5 p.m.For questions regarding this recall call 574-371-1021. |  | 30 | 30 units | Worldwide Distribution - USA (nationwide) and Europe, Austrailia | Terminated | April 18, 2013 | 365 | 30
Z-1509-2012 | 61630 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ARCHITECT STAT Troponin-I Reagent Kit; list numbers 2K41-28 (500 test kit); an in-vitro diagnostic test kit; each kit contains bottles of microparticles, conjugate and assay diluent; Abbott Laboratories, Abbott Park, IL 60064 |  |  | April 11, 2012 | 2012 | Abbott Laboratories | The lot of Troponin-I is demonstrating a shift in expected results in some cases, and varies from kit to kit. A concentration change of up to +/-55% could occur when switching kits without recalibrating. | Material/Component Contamination | The firm, Abbott Laboratories, sent a "Product Recall Immediate Action Required"letter dated April 11, 2012 to all Troponin-I customers who received the affected reagent lot internationally. The customers were directed to determine if they are currently using and/or have inventory of lot 74264UN11; to discontinue use and destroy any remaining inventory of the lot; to follow their facility's policies and procedures for communicating with the heath care providers they serve; to forward a copy of the letter to any laboratory to whom they may have forwarded the reagent; to retain a copy of the letter for their records and to complete and return the Customer Reply Form (sent to Abbott international affiliates via e-mail on 4/11/2012) via fax to: 1-800-777-0051 or e-mail: QAGCO@abbott.com. Any questions were directed to their local Customer Service representative or call the field action coordinator at 847-938-1923 or email: albert.chianello@abbott.com. |  | 8723 | 8,723 kits | International distribution only: including countries of: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bosnia &amp; Herze, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macedonia, Mali, Malaysia, Mauritania, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Paraguay, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela and Vietnam. | Completed |  |  | 8723
Z-1516-2012 | 61661 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus(R) CS Acute care(TM) NT - proBNP TestPak reagent (CPBNPM) |  |  | April 05, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed a low frequency of calibration failures due to slope errors that are not resolved with routine maintenance with the specific SCS Pro-BNP lots listed. This issue only causes the slope error failures but does not affect patient results when the calibration is within specifications. | Component design/selection | The firm, SIEMENS, sent a "Urgent Field Safety Notice" dated March, 2012 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to discontinue use of the lots once revised expiration date occurs; contact their local Siemens Customers Service Center at 888-588-3916 for replacement; complete and return the FIELD CORRECTION EFFECTIVENESS CHECK form via fax to the TECHNICAL SOLUTIONS CENTER at (302) 631-8467 as soon as possible, and forward this notification to anyone to whom they may have distributed this product. The remaining inventory at the firm was destroyed.If you have any technical questions regarding this information, please contact the Siemens Technical Solutions Center at 800-405-6473. |  | 1899 | 1,899 cartons | Nationwide distribution: USA including states of: AK, AR, CA, CO, CT, DC, FL, HI, ID, IN, IL, IA, KS, KY, LA, ME, MI, MN, MS, MO, MT, NE, NH, NM, ND, NY, OH, OK, OR, PA, SD, TX, UT, VT, WA, WV, WI, and WY. | Terminated | March 19, 2013 | 348 | 1899
Z-1514-2012 | 61481 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Zimmer Dental Trabecular Metal Implant, TMM4B11, 4.1mm x11.5 MM; 3.5 mm implant TM, MTX Rx, Sterile R, Zimmer Dental, Carlsbad, CA 92008, USA |  |  | March 22, 2012 | 2012 | Zimmer Inc. | This Zimmer Dental voluntary device recall resulted from an investigation into the February 2012 complaint involving an apical tip of a 4.1mm D Trabecular Metal Implant which separated from the implant assembly during surgery on a patient with a dense (Type D1), thick, inferior border. The recall correction includes adding precautions to the Instructions for Use regarding the placement of a 4.1mm | Labeling Change Control | Zimmer sent an Urgent Medical Device Correction notification letter dated April 24, 2012, with an attached Technical Bulletin listing precautions present in the new Informaiton for Use sent to distributors and customers via E-mail and FED EX. Notifications identified the issue and risks found with the device listing responsibilties, precautions, and procedural precautions. Notifications contained directions to forward information to any sub accounts and clinicians. Zimmer had telephone follow-up with each distributor and customer to assure information was received and is understood. Questions should be directed to Technical Service (800)8511-7019 orCustomer Service(760) 929-1300 or Regulatory Affairs at 1-800-854-7019 .For questions regarding this recall call 574-372-4807. |  | 233 | 233 | Worldwide Distribution - USA (nationwide) and the countries of Germany, Italy, France Egypt, Spain, Russia, Bulgaria, Turkey, Israel, UAE and CANADA | Terminated | December 03, 2012 | 256 | 1209
Z-1506-2012 | 61517 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Siemens Axiom Sensis XP diagnostic programmable computer |  |  | February 22, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction in the connection of a Sensis EP system with Carto(R) 3 dedicated cable set (part number 14411343) with the Carto 3 system from Bionsense Webster. | Nonconforming Material/Component | Siemens sent a recall/correction letter dated February 22, 2012, via Update Instruction AX049/11/S of this product to all affected customers. The letter informed customers of the potential issues and provides additional information on the implemented resolution. An Update Instuction AX048/11/S will be available immediately. The update instruction provides a new version of the Carto 3 addendum to the Operator Manual and the exchange of the previous stim cables with the latest cable.For questions call 610-219-4834. |  | 3 | 3 | Nationwide Distribution including FL, NC, and SC. | Terminated | January 24, 2014 | 702 | 3
Z-1519-2012 | 61522 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Instrumentation Laboratory (IL)HemosIL von Willebrand Factor Activity Assay Kit |  |  | April 11, 2012 | 2012 | Instrumentation Laboratory Co. | Falsely elevated levels may be reported with the HemosIL von Willebrand Factor Activity assay. | Nonconforming Material/Component | Instrumentation Laboratories mailed Recall Letter via USPS to US and Canadian customers on April 11, 2012, through Beckman Coulter. The customer mailing and tracking of the Urgent Recall Notification and Recovery Procedure in the United States and Canada will be handled through the distributor, Beckman Coulter. Accounts are requested to check inventory for the product lots of Hemosll von Willebrand Factor Activity, PN 0020004700, listed in the table and document their destruction on the providedtracking form and return the completed form. Questions or concerns regarding this notification, please call Beckman Coulter Technical Service at (800) 526-7694.On March 29, an e-mail communication was circulated to the IL Affiliates in Europe and Japan requesting them to translate the Urgent Field Safety Notification [title of document outside US and Canada] and product destruction tracking form, mail the notification to their customer base and notify theirCompetent Authorities..Fpr questions regarding this recall call 781-861-4467. |  | 9320 | 9320 Kits | Worldwide Distribution - USA (nationwide) and the countries of Canada, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELARUS BELGIUM, BRAZIL, CHILE, CHINA, CZECH REPUBLIC, DENMARK FRANCE, GERMANY, GREECE, HUNGARY, INDIA, IRAN (ISLAMIC REPUBLIC OF), ISRAEL, ITALY, JAPAN, JORDAN, KOREA, REPUBLIC OF KUWAIT LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS NORWAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TURKEY, and the UNITED ARAB EMIRATES | Terminated | August 24, 2012 | 135 | 9320
Z-1517-2012 | 61681 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 8" x 4" 16-ply X-Ray Gauze, hospital fold, banded in 10's; 200 pieces per bag, 10 bags per case; Distributed By: International Medsurg Connection, Schaumburg, IL USA 1-847-619-9926; Made in China; item IMSGZX8416-10 |  |  | April 09, 2012 | 2012 | International Medsurg Connection, Inc. | The lots of X-Ray Gauze are contaminated with foreign material, including remnants of cotton seeds and lint. | Process control | International Medsurg Connection notified their direct accounts via telephone, e-mail and/or fax on April 9, 2012, informing them of the foreign material contamination of the gauze, and requesting them to quarantine all affected lots in inventory, cease distribution of the affected product, and notify any customers at the retail level to remove and discard these products. Any questions were directed to (847) 619-9926. |  | 276 | 276 cases | Nationwide Distribution including Illinois, South Carolina, Texas and Virginia | Open, Classified |  |  | 291
Z-1541-2012 | 53196 | No | KZE | Nonelectrically powered fluid injector. | General Hospital | General Hospital | 510(k) | Trima Accel Automated Blood Collection System, catalog number 917000000, Caridian BCT, Lakewood, CO 80215 USA. |  |  | June 01, 2010 | 2010 | CaridianBCT, Inc. | Labeling clarification and software updates to include fixes for end of run summaries, increasing boot reliability, corrections to the double red blood cell count collection TBV limit, and to be in compliance with the most current FDA guidance for post donation platelet count targets. | Software design | CaridianBCT sent a Letter notification dated March 2011 to consignees The letter identified the upcoming software updates to version 6.0.1 as well as operations manual changes. For questions contact your CaridianBCT Representative, the CaridianBCT Support Center at +1-877-339-4228 or 1-303-231-4357 or your local CaridianBCT Customer Service office. |  | 2278 | 2278 units | USA Nationwide Distribution | Terminated | July 25, 2012 | 785 | 2278
Z-1539-2012 | 52454 | No | GWG | Neurological endoscope. | Neurology | Neurology | 510(k) | Mitaka Point Setter, holding system for surgical instruments. |  |  | February 27, 2009 | 2009 | Mitaka USA, Inc. | Use of a potentially weaker spring in the device than required to meet specifications to hold the Point Setter ARM in place during surgery. | Component design/selection | Mitaka USA sent a letter dated to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter included instructions for returning the product. Contact the Customer Support Department at 800-421-0837, ext. 7577 for questions regarding this notice. |  | 18 | 18 units (3 internationally and 15 within the U.S.) | Worldwide Distribution-USA (nationwide) including the states of AZ, FL, CA, SC, OH, WA, TX, VA, MI, MO, GA, and NC and the country of Canada. | Terminated | May 22, 2012 | 1180 | 18
Z-1530-2012 | 61644 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Terumo Custom Cardiovascular Procedure Kit-PEDIATRIC PACK |  |  | April 12, 2012 | 2012 | Terumo Cardiovascular Systems Corp | Loose connections (i.e., connections between PVC circuit tubing and Oxygenator ports) may result in leaks and patient blood loss | Nonconforming Material/Component | Terumo Cardiovascular Systems issued an Urgent Medical Device Recall notification on 4/12/12 to all user facilities/consignees who have received affected devices. The letter identified the affected product, stated the reason for the recall, the potential hazard involved, and the correction. Users are to review the letter, ensure that all users are made aware of the issue, discontinue use of the affected product, and return it to Terumo CVS. In addition, users are to return the response form to the manufacturer. Affected units will be replaced with new units. Contact Terumo CVS Customer Service @ 1-800-521-2818 or tcvs.recall@teruomedical.com |  | 37 | 37 units | Worldwide Distribution -- US, including the states of AL, AZ, CA, MO, and NY and the country of Canada. | Terminated | January 23, 2013 | 286 | 843
Z-1522-2012 | 61786 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | The Evolution steam sterilizers, manufactured by Steris Corporation, utilize a loading system to assist the operator in transporting, loading and unloading items to be sterilized. It consists of two major components: |  |  | March 23, 2012 | 2012 | Steris Corporation | Through Customer feedback and field service experience, STERIS learned that operators may experience difficulty when inserting or removing equipment from the Evolution sterilizer using the transfer carriage due to the following issues with the transfer carriage. 1) Deformation of the latch pin or cracked ball bearings preventing the roller from turning.2) The rollers on the ends of the pivot a | Device Design | Steris Corporation sent an Urgent Voluntary Field Correction Notice dated March 23, 2012 via FedEx with tracking numbers for delivery confirmation to each affected customer. The letter identified the affected product, problem and actions to be taken. The letter stated a STERIS Field Service Representatives will contact each affected customer to schedule a service visit to upgrade the transfer carriages. For technical questions or questions regarding STERIS visit to your facility, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 523 | 523 | Worldwide Distribution - USA Nationwide including AL, AZ, CA, CO, CN, DE, FL, GA, IL, IN, KA, KY, LA, MD, MA, MI, ME, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, &amp; WY and the countries of Brazil and Canada. | Terminated | June 11, 2013 | 445 | 523
Z-1536-2012 | 61527 | No | LLI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | LUMEX A Graham-Field Brand, Bed End Sert. Low/Full-electric, Plastic Panels, Item 690-7604-000 Manufactured for : GF Heath Products, Inc., Atlanta, Georgia 30360 Made in China. |  |  | April 17, 2012 | 2012 | GF Health Products, Inc./d.b.a.Basic American Metal Products | GF Health Products, Inc., is conducting a recall on the Lumex Patriot Full Electric Homecare "Bed Ends" manufactured 3/27/09, 5/9/09, 5/30/09 and 6/20/09. The supplier deviated from the print specifications of a machined pin to a sheared pin. This deviation has the potential to provide room for the pin to move during use. If failure were to occur, the bed could descend rapidly causing non-typic | Component change control | GF Health Products, Inc. sent an "URGENT CORRECTION AND REMOVAL" letter dated April 17, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately cease use of the affected product, secure the product, and to prevent the product from further distribution to customers and end users. Contact the firm at 1-800-556-6508 for questions regarding this notice. |  | 3095 | 3095 pairs of bed ends | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, and WI and the countries of EL SALVADOR, CHILE, GUATEMALA, TRINIDAD, TOBAGO, and BAHAMAS. | Terminated | July 15, 2013 | 454 | 3095
Z-1537-2012 | 61636 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | HemaSyst, Dual Cannula Tip with Adapter 16 ga. x 4" (10.2 cm). Product is labeled in part:"***Pouch label***HAEMACURE***HemaSyst***tm***Contents: One ***DUAL CANNULA TIP***WITH ADAPTER***16 ga. x 4" (10.2 cm)***REF H-1604*** CE 0297***LOT 111144***2014/10***40680b/20004-02***CAUTION: U.S. Federal Law restricts this device to sale by or on the order of a licensed physician (or properly licensed practitioner).***Contents of unopened undamaged package are sterile***CE 0297***Manufactured by: ***Micromedics*** St. Paul, MN 55121-1385 USA***Distributed by:***Haemacure Corporation***Sarasota, FL 34236 USA***Toll Free 1-877-872-4583***www.haemacure.com***Box label***Contents: 5 packages of 1 each***DUAL CANNULA TIP***WITH ADAPTER***16 ga. x 4" (10.2 cm)***REF H-1604***CE 0297***STERILE EO***LOT 11144***2014/10***40681B/20544***". |  |  | November 16, 2009 | 2009 | Haemacure Corporation | This product has been recalled due to the lot number being different on the box and pouch. The lot number on the box label was incorrectly printed as 11144 instead of 111144. The pouch label referenced the correct lot number, 111144. | Labeling mix-ups | Haemacure Corporation, Sarasota, FL initiated letters via e-mail on 11/16/2009 with delivery and read receipt confirmation to their two consignees. The letters identified the affected product and stated the reason for the recall. The letter also states that the firm will replace each order upon receipt of the mislabeled boxes at no charge to the customer. Customers were asked to return their unused devices. Questions should be directed to customer service at 1-877-872-4583. |  | 7 | 7 boxes | Nationwide Distribution -- Illinois and California. | Terminated | May 14, 2012 | 910 | 7
Z-1540-2012 | 61629 | No | MTF |  |  | Immunology | PMA | Dimension Vista(R) Total Prostate Specific Antigen Flex(R) reagent cartridge |  |  | April 02, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed a report of leaking Flex(R) reagent cartridge lot 11300BB from Dimension Vista(R) TPSA; when a leaking Flex(R) is used, patient test results may be falsely depressed. | Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notification letter dated April 14, 2012 to all Dimension Vista(R) System customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and discard any remaining inventory of the affected product. Customers were asked to complete the attached form and fax it to 302-631-8467 to indicate that you have received the notification. For technical questions or concerns, contact the Siemens Technical Solutions Center at 800-441-9250. |  | 1103 | 1,103 cartons | USA Nationwide Distribution including the states of: AL, AZ, CA, CO, CT, FL IL, LA, MD, MI, MN, MS, MT, NC, NH, NJ, NV, NY, OH, PA, SD, TN, TX, UT, WI ,WV including Puerto Rico. | Terminated | April 02, 2013 | 365 | 1103
Z-1577-2012 | 61468 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00-5976-020-17, NexGen CR Complete Knee Solution, Cruciate Retaining (CR) Articular Surface Anterior Constrained AC, Size Purple, AE Micro 17mm Height , Sterile R, Rx ,Zimmer, Inc. Warsaw, IN |  |  | February 24, 2012 | 2012 | Zimmer Inc. | Zimmer has received 5 complaints where a NexGen CR micro articular surface was used with a standard CR femur, even though the compatibility chart and product labeling indicates that these components are NOT compatible. Investigation identified additional instances where a CR micro articular surface and standard CR femur were sold together on the same purchase order. CR micro articular surfaces are | Labeling design | Zimmer sent an Urgent Correction Notice dated March 5, 2012, to all affected customers via telephone, E-mail and letter notifying users of the appropriate compatibility of micro sizes of NexGenÂż CR and CR-Flex Articular Surfaces and informing them of known cases of incompatible devices being implanted, including associated risks. Customers and surgeons were asked to acknowledge receipt by calling 1-888-912-7349. Questions and concerns or help in notiying accounts about this correction were to be addressed to 1-888-548-8514.For questions regarding this recall call 574-372-4807. |  | 83322 | 83,322 all sizes | Worldwide Distribution - USA (nationwide) and the countries of Poland, Taiwan, Denmark, Korea, Singapore, China, Canada, Australia, Japan, France, Italy, UK, Germany, Spain, Belgium, Portugal, Switzerland, Sweden, India, Thailand, Peru, The Netherlands, Malaysia, and Austria. | Terminated | April 18, 2014 | 784 | 83322
Z-1542-2012 | 61475 | exerciser, non-measuring | ION | Nonmeasuring exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Norco Shoulder Pulley |  |  | July 11, 2011 | 2011 | North Coast Medical Inc | Potential for plastic D-ring to break during use and cause injury. | Device Design | North Coast Medical began contacting customers of the recall on July 13, 2011 by telephone. A phone script was used to inform customers of the affected product, problem, and actions to be taken by the customers. Information was provided for customers to return affected product. Affected devices were re-worked and inspected. Contact the firm by e-mail at www.ncmedical.com for questions regarding this recall. |  | 2364 | 2364 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Canada, Switzerland, Germany, Great Britain and the Netherlands. | Terminated | September 06, 2012 | 423 | 2364
Z-1543-2012 | 61699 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PATIENT SPECIFIC INSTRUMENTS |  |  | March 23, 2012 | 2012 | Materialise USA LLC | March 23, 2012 a field representative reported patient specific -0242-L tibia guide case packaging contained a different patient specific -0864-R tibia guide. The two patient specific products had been switched in packaging. Neither products had been delivered to the end user. | Labeling mix-ups | Materialise telephoned both field representatives on March 23, 2012 and confirmed recalled product mix-up. Materialise requested both cases be returned and replacements were provided. No affected product was received by the end user. |  | 1 | 1 unit | Nationwide Distribution including the states of Ohio and Minnesota. | Terminated | January 23, 2013 | 306 | 1
Z-1546-2012 | 61500 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | RF IgA Calibrator, a component of the EL-RF/3 Rheumatoid Factor Test Kit; an in vitro diagnostic; TheraTest Laboratories, Lombard, IL 60148; component catalog #12833, kit catalog #303-305. The TheraTest EL-RF/3 is intended as an aid to the diagnosis of Rheumatoid Arthritis (RA). |  |  | March 29, 2012 | 2012 | Thera Test Laboratories, Inc. | The RF IgA Calibrator component of the EL-RF/3 Rheumatoid Factor test system may be unstable, resulting in a drop in the RF IgA calibrator OD value which could invalidate an RF IgA test run. | Nonconforming Material/Component | TheraTest sent "Urgent Device Recall" letters dated March 28, 2012 to customers that received the affected kit, requesting them to return or destroy the remainder of the product. TheraTest issued a replacement kit to all customers that had unused product as well as any material they consumed which resulted in a failed run. Customers were also instructed to complete the enclosed confirmation form and fax the completed form back to TheraTest. Any questions were directed to TheraTest at 800-441-0771. |  | 207 | 207 kits | Nationwide Distribution (US) | Open, Classified |  |  | 207
Z-1545-2012 | 61804 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Instructions For Use: "***SmartLOX***Instructions for Use Cervical Plate System***STERILIZATION***The SmartLOX Cervical Plate System components are supplied clean and not sterile. All implants and instruments should be cleaned and sterilized prior to surgery. AORN recommended practices for in hospital sterilization should be followed. Sterilization testing of components has shown the following recommendations for sterilization are effective Method: Steam***Cycle: Pre vacuum***Temperature 270 F (132 C)***Exposure Time: 3 minutes***Contact Information: Captiva Spine Inc. 967 Alternate AA #1 Jupiter, FL 33477. Tel: 877-772-5571***CAPTIVA SPINE***". |  |  | March 22, 2012 | 2012 | Captiva Spine, Inc | Captiva Spine has initiated a correction of the SmartLOX Instructions for Use (IFU) for SmartLOX Cervical Plate System. The pre vacuum sterilization cycle [270 (132 C) for 3 minutes] has not been validated. | Labeling False and Misleading | Captiva Spine sent a "VOLUNTARY CORRECTION NOTIFICATION" letter dated March 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides steps for the correction and a verification form for customers to complete. Customers should contact the firm at 877-772-5571 for questions concerning this notice. |  | 14 | 14 units | Nationwide Distribution including the states of FL, IN, UT, GA, PA, TX, AL, NC, MI, and MN. | Terminated | November 18, 2013 | 606 | 14
Z-1612-2012 | 61418 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Medical Naturalyte Acid Concentrate for bicarbonate Dialysis 45x , 55 gallons (208.2 Liters) |  |  | March 16, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Sodium content does not meet product requirements | Process control | Fresenius Medical Care North America sent an "URGENT PRODUCT RECALL" letter dated March 28, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their stock immediately to determine whether they have any affected product. A fax back form was included for customers to complete and return. Contact FMCNA Customer Service Team at 1-800-323-5188 for questions regarding this notice. |  | 154 | 154 units | Nationwide Distribution. | Terminated | January 28, 2014 | 683 | 154
Z-1614-2012 | 59861 | No | NBE |  |  | General &amp; Plastic Surgery | PMA | Coseal Surgical Sealant. |  |  | July 28, 2011 | 2011 | Baxter Healthcare Services L.P. | The recall is being conducted as a precautionary measure due to out of specification results observed at the 24 months time point (out of 24-month licensed shelf life to expiration) during a Coseal stability study. The parameter that is out specification is an indicator of possible failure for the product to gel appropriately . Coseal's failure to gel does not represent risk for the patient's life | Device Design | Baxter Healthcare sent an"URGENT PRODUCT RECALL" letter dated July 27, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately stop the usage/distribution of the affected product and to quarantine any units that are in inventory. Instructions for obtaining replacement product were attached to the letter. Contact the firm at 1-800-423-2090 for questions concerning this recall. |  | 11160 | 11160 | Nationwide Distribution. | Terminated | May 31, 2012 | 308 | 11160
Z-1615-2012 | 61719 | No | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | QuietCare, Model Numbers QC200403, v.02; QC200403, v.03; QC200403, v.04; QC200405 (USA) and QCE200403, v.02; QCE200403, v.03 (UK). Manufacturer Living Independently Group, Inc., New York, NY. QuietCare is intended for use in monitoring the environmental conditions and activity (motion) of an individual living in a senior housing community. |  |  | July 30, 2011 | 2011 | Intel-GE Care Innovations LLC | The QuietCare production system experienced a partial outage on 7/29/11. As a result of the partial outage, some monitoring data about residents' movements was not processed in the production system, and alerts were therefore not sent to caregivers. The problem was identified and corrected and the unprocessed data was recovered and processed. However, because of the time lag in processing, a fe | Employee error | Initial notification to the affected facilities was done by telephone or email on 7/30/11. This was followed up with Urgent Field Notification letters to the customers on 8/12/11 via e-mail. For questions, customers were instructed to contact Customer Support at 1 (866) 772-8243. |  | 164 | 82 units affected by the outage: Domestic - 80 units, Foreign - 2 units | Worldwide distribution - USA and the United Kingdom. | Terminated | May 21, 2012 | 296 | 164
Z-1616-2012 | 61803 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S1000 ultrasound systems running software version 1.0, Siemens Medical Solutions, USA, Inc. Mountain View, CA. The product is indicated for use as an ultrasound imaging system. |  |  | April 25, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale. | Software design | A field corrective action was initiated on 04/25/2012 under Siemens update program US019/12/S in which a "Customer Safety Advisory Notice" was sent to all customers notifying them of the problem and the recommended actions to avoid potential risk. A follow-up field corrective action will be initiated in which the affected software will be updated in the field. |  | 45 | 45 | Worldwide Distribution -- USA, including the states of TX, NJ, GA and the countries of Germany, United Kingdom, Republic of Korea, Australia, Japan, Poland, Israel, Singapore, France and Saudi Arabia. | Terminated | October 19, 2012 | 177 | 45
Z-1617-2012 | 61694 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam and TrueBeam STx V1.0, 1.5, 1.6.95 and below. TrueBeam Radiotherapy Delivery System is intended to provide stereotactic radiosurgery and precision radiotherapy anywhere in the body where radiation treatment is indicated. Varian Medical Systems, Palo Alto, CA. |  |  | April 09, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian received a report involving a Gantry collision and is sending a notification to remind users of the collision protection tools available with TrueBeam and TrueBeam STx, including actions the operator should be aware of. | Labeling design | An "Urgent Medical Device Correction / Field Safety Letter" was distributed to all affected users on April 9, 2012, with a description of the problem and user corrective action steps. The Letter was also distributed to the Varian Sales, Marketing, and Service organizations. Customers in the US and Canada may contact (888) 827-4265 for questions.*** An additional letter was sent on 11/5/12 to additional consignees identified by Varian as being affected the the recall but not previously notified. |  | 2484 | 207 devices 12/7/12 - amended to 216 devices. | Worldwide Distribution -- USA, and the countries of Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Netherlands, New Zealand, Norway, Portugal, Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, and the United Kingdom. | Terminated | April 16, 2014 | 737 | 2484
Z-1613-2012 | 61657 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Galil Medical, 17G Cryoablation Needle Holder FPRPR 4014, Rx only, Sterile EO. |  |  | April 12, 2012 | 2012 | Galil Medical, Inc. | Galil Medical is recalling 17G Cryoablation Needle Holder, FPRPR 4014. This voluntary recall has been initiated due to the possibility of excessive bioburden on the Needle Holders. | Material/Component Contamination | Galil Medical sent an "Urgent Medical Device Recall" letter dated April 12, 2012 to all affected customers. The letter identified the affected product, problem, actions to be taken. The letter instructs customers to return all unused affected products, use a revised needle preparation procedure and forward and communicate this Recall Notice and instructions. Customers were instructed to complete and return the attached Response Form with returned product or fax to 1+877-510-7757. For any questions contact Galil Medical Customer Service. |  | 26 | 26 kits | Worldwide Distribution - USA (nationwide) including the states of CA, NY, NC, WA and the countries of CANADA, DENMARK, FRANCE, NETHERLANDS, TAIWAN and UNITED KINGDOM. | Terminated | August 16, 2012 | 126 | 26
Z-1618-2012 | 61715 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cutting Edge Acetabular Spherical Reamer 44 M, Stryker Howmedica Osteonics Howmedica Osteonics Corp. 325 Corporate Drive Mahwah, NJ 07430 USA. |  |  | December 23, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | A 46mm acetabular reamer was lasermarked as a 44m reamer and distributed to the field. | Device Design | Stryker sent an "URGENT PRODUCT RECALL" letter dated December 29, 2011 to all affected customers. Notificationletters and Product Accountability Forms were sent to Stryker accounts via FedEx on January 3, 2012 with return receipt.Hospital Risk Management and Chief of Orthopaedics Notification letters and Product Recall Acknowledgement Forms were sent on January 10, 2012.Surgeon Notification letters and Product Recall Acknowledgement Forms were sent on January 10, 2012.Notification sent via e-mail toCanada.EMEA, LatinAmerica.Asia Pacific on December 23, 2011.The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to return the affected product. A Product Recall Acknowledgment Form was attached to the letter for customers to complete and return. Contact the firm at 201-972-2100 for questions regarding this recall. |  | 134 | 134 units | Worldwide Distribution. | Terminated | July 17, 2013 | 572 | 134
Z-1623-2012 | 61819 | No | NPL | Bone grafting material. | Dental | Dental | 510(k) | a translucent, rectangular, resorbable, collagen membrane sheet derived from bovine tissue. |  |  | February 28, 2012 | 2012 | Kensey Nash Corp | Recall has been initiated due to concerns regarding the sterility of the product. The firm initiated the recall because of use of inadequately sterilized product might result in patient infection, | Employee error | Reimser, Inc. sent an "URGENT PRODUCT RECALL-COLLAGUIDE COLLAGEN DENTAL MEMBRANE" letter dated March 12, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to cease use, secure and quarantine affect product. Additionally, a Recall Acknowledgement Form was attached for customers to complete and return to the firm via fax to 1-800-505-4982. Contact the firm at 484-713-2197 for questions regarding this recall. |  | 13189174161 | total all sizes: 6387 units; p/n 20650-03 - 765 units | Worldwide Distribution. | Terminated | April 11, 2013 | 408 | 13189174161
Z-1624-2012 | 61801 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Pelleve Non-Ablative Wrinkle Treatment Handpiece, Product Codes D110FSHP, D115FSHP (reusable) --- Pelleve --- Elegance through science --- Contents: One (1) Pelleve Handpiece (Reusable) --- Non-ablative treatment of mild to moderate facial wrinkles and rhytids for skin phototypes I-IV. |  |  | May 04, 2012 | 2012 | Ellman International, Inc. | Ellman International is recalling its Pelleve reusable handpiece because of possible cable failure and sparking. | Equipment maintenance | Ellman will notify all consignees initially via UPS shipment with an Urgent: Medical Device Recall letter containing notification of the recall and instructions. The letter identifies the affected product, problem, and instructs customers to identify and return the handpieces in the enclosed prepaid mailer. Customers are to complete and return the Product Recall Verification Form and fax it to the number provided. In addition, customers are provided with a toll-free number, 1-877-522-7141, to call in the event that they have questions or require additional instruction. Customers will initially be provided with a fixed-life disposable GlideSafe handpiece and may choose to receive a replacement reusable handpiece with modified (more robust) cable, or additional GlideSafe handpieces. |  | 1088 | 1,088 units | Worldwide Distribution | Terminated | February 12, 2016 | 1379 | 1088
Z-1619-2012 | 61343 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Tibial Alignment Handle |  |  | October 10, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker received a Product Experience Report (PER) that a swivel block was improperly oriented on a Triathlon Tibial Alignment Handle, catalog no. 6541-2-807. | Device Design | Stryker sent an Urgent Product Field Action letter dated October 10, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Stryker Sales Representative to arrange for return of the affected product . Customers were asked to fax back the attached Poroduct Field Action Acknowledgement Form as soon as possible to 201-831-6069. For any questions customers were instructed to call 201-972-2100.For questions regarding this recall call 201-831-3523. |  | 6126 | 6,126 Units | Worldwide Distribution - USA (nationwide) and the countries of Asia Pacific, Canada, Europe, Middle East and Africa (EMEA), and Japan | Terminated | September 10, 2014 | 1066 | 6126
Z-1625-2012 | 61859 | No | KDH | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 50 Portal Vein Cannula REF 61.110.10.030 Size: Fr 10 Length: 12 Inches LL Connector LOT 120493 NON STERILE MEDDEV USA, INC. 100 Highland Park Village/Suite 200 DALLAS, TEXAS 75205 |  |  | May 10, 2012 | 2012 | MEDDEV USA, INC. | The device was marketed without a 510(k) and is not sterile. | No Marketing Application | MEDDEV USA, Inc decided to recall and notified their customer by phone on 05/10/12. Customers were instructed to return the entire shipment to MEDDEV .For questions regarding this recall call 214-692-3559. |  | 20 | 20 units | Nationwide Distribution including Illinois | Terminated | August 15, 2013 | 462 | 20
Z-1626-2012 | 61860 | [6 empty] | 8F Plastic Dignity Mid-Sized CT Port w/Silicone Filled Suture Holes and w/Attachable ChronoFlex Polyurethane Catheter. Catalog # MRCTI8000S. The Power Injectable, Implantable Infusion Port is an implantable single fluid reservoir port with a polyurethane (8F) catheter attachable for application by the inserting physician. |  |  | April 12, 2012 | 2012 | Medical Components, Inc dba MedComp | 50 trays of Dignity Low Profile CT Port were labeled as Dignity Mid-Sized CT Port. | Employee error | On April 12, 2012, MedComp sales representatives and the foreign distributor were notified via e-mail with a "Product Alert" notification. The issue was described and requested actions were provided to recipients. |  | 50 | 50 trays | Worldwide Distribution - - USA, including the states of CA, NJ, AZ, PA, TN and the country of Spain. | Terminated | July 18, 2013 | 462 | 50
Z-1627-2012 | 61750 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The ECHELON 60 Endoscopic Linear Cutter - Straight, compact ECHELON 60 Endoscopic Linear Cutter - Straight, and the Long ECHELON 60 Endoscopic Linear Cutter - Straight are sterile, single patient use instruments that deliver staples while simultaneously dividing tissue between rows. ECHELON 60 instruments with white, blue, gold, and green reloads deliver two triple-staggered rows of staples. The instruments have a staple line that is approximately 60 mm long and a cut line that is approximately 57 mm long. A staple retaining cap on the reload protects the staple leg points during |  |  | March 02, 2012 | 2012 | Ethicon Endo-Surgery Inc | There is a possibility of a component breakage, which may cause the device to stay locked on tissue after firing. If the component exhibits this issue, the device does not open either with the anvil release button or with the manual firing release lever. | Vendor change control | Ethicon Endo-Surgery sent an Urgent Device Recall letter dated March 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately to determine if they have the affected product and to remove the affected product. Fill out the Business Reply Form and fax it to 1-513-337-4138 within 3 business days, even if they do not have the affected product in their inventory. Ethicon will automatically replace any affected product returned by June 29, 2012. Customers were asked to share this information with appropriate staff.If customers should need clinical or product support call 1-800-873-3636, option 6, or contact their local sales representative.For questions regarding this recall call 513-337-7000. |  | 3120 | 3120 | Worldwide Distribution - USA including AL, AR, CA, CO, FL, IA, IL, KS, LA, MD, MO, MS, NC, NM, NY, OH, SC, TX &amp; VA; and the countries of Argentina, Ecuador, Mexico, Australia, Egypt, Singapore, Belgium, Greece, Slovenia, Brazil, India, South Africa, China, Israel, United Arab Emirates, Colombia, Japan &amp; Venezuela. | Terminated | March 22, 2013 | 385 | 3120
Z-1630-2012 | 61444 | No | MKQ |  |  | Pathology | PMA | BD Focal Point"100V (Refurbished), Catalog number: 490651 |  |  | October 06, 2011 | 2011 | Tripath Imaging, Inc. | The cable tie mounts that adhere to the inside of the FocalPoint do not adhere as intended and can therefore become detached. | Component change control | BD Diagnostics first notified consignees by letter on October 6, 2011. BD Diagnostics also sent a letter dated February 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Field Representatives were advised to inspect all cables, adhesive cable ties and connections to ensure they are secure. Inspections and repairs, if necessary, would be carried out at the next service call for each instrument, or no later than the next regularly scheduled preventative maintenance visit. For questions consignees were instructed to contact BD Diagnostics Technical Support at 1-877-822-7771 option 2 in the US, their local distributor, or their local BD office for any questions.For questions regarding this recall call 918-563-1624. |  | 1 | see line 1. | Worldwide Distribution - USA including AZ, CT, IL, IN, FL, MN, NC, OK, OR, and TX and the countries of Australia, Belgium, Brazil, Canada, Japan and Thailand | Terminated | December 03, 2012 | 424 | 73
Z-1628-2012 | 61663 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Silicone Cuffed Wire Reinforced Extended Connect Pediatric Perfect Fit Neck Flange Tracheostomy Tube, Size 4.0 mm I.D. Catalog Number 353-040. |  |  | April 18, 2012 | 2012 | Arcadia Medical Corporation | A defect in 3 tracheostomy tubes. The defect observed was a separation in the bond between the proximal junction of the neck flange and the cannula. | Process control | Arcadia notified the 3 consignees with a MEDICAL DEVICE VOLUNTARY RECALL notification dated 4/18/12. The notification identified the affected product, along with the reason for the recall. Consignees were instructed to cease distribution/use of any unopened product that they have remaining and return to firm following the steps provided. Arcadia will replace the product at no charge to the customer. Consignees were to confirm receipt of the recall notification by emailing Trena Depel, trena@arcadiamedical.com.Phone: 925-918-0595Fax: (415) 762-5489 |  | 1103307 | 01103307 (50 units) and 01107055 (73 units) | Worldwide Distribution - US, including Mariemount, OH and countries of UK and France. | Terminated | August 13, 2012 | 117 | 1103307
Z-1620-2012 | 61568 | No | LLL | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Product is labeled in part: "***is immunosimplicity***ENA-6 Screen***Enzyme Immunoassay Test Kit***Catalog No: 720-230***Exp. Date: 2012-05-31***Contains Sufficient for 96 Tests***C - Corrosive Harmful***Lot No: S0000***For In Vitro Diagnostic Use***Caution, consult accompanying documents***Manufacturer DIAMEDIX***a Subsidiary of IVAX Diagnostics, Inc.***Miami, FL 33127 USA-1-800-327-4565***Biohazard***" |  |  | June 08, 2011 | 2011 | Diamedix Corporation | On 06/08/2011 Diamedix Corporation initiated a recall of Is-ENA-6 Screen Test Kit, Catalog # 720-320 Master Lot S0000, Exp. Date May 31, 2012. The labeling on the kit box has an expiration date of 2012-05-31 while the components have an expiration date of 2012-04-30 for an incorrect expiration date. | Error in labeling | Diamedix Corporation sent a Customer notification letter dated June 8, 2011, to all affected customers via Certified Mail. Customers were instructed to search their inventory and to return affected product to Diamedix.Diamedix apologizes for any inconvenience this has caused our valued customers. Please contact the Diamedix Technical Service Department at 1-800-327-4565 or 305-324-2300 to arrange for the return and replacement of any kits you might still have of the lot in question. Please rest assured that Diamedix values you as a customer and is in the process of taking the proper corrective and preventive actions. |  | 312 | 312 Test Kits | Worldwide Distribution -- (USA) nationwide including the states of FL, NY, NC, MS, KY, MA and Puerto Rico and the countries of Italy, Chili and Venezuela. | Terminated | November 15, 2013 | 891 | 312
Z-1647-2012 | 48179 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | EZ STEER Catheter D-1267-04-S / 8mm Non-Navigation Bi-Directional Catheter |  |  | December 15, 2007 | 2007 | Biosense Webster, Inc. | The recall was initiated because the catheter unit was shipped with a label expiration date of 9/30/2010, which is 11 months beyond the correct expiration date of 10/31/2009. | Incorrect or no expiration date | The firm, Biosense Webster, sent a "Medical Device-Voluntary Field Removal" letter dated December 17, 2007 to its customer. The letter identified the product, problem and an action to be taken. The customer was instructed to complete and return the Product Voluntary Field Removal Certification Form via fax to Quality Compliance at (909) 839-7207or mail using the enclosed self-addressed stamped envelope. If you have additional questions, please follow up with your local Biosense Webster representative.or contact Customer Service at (909) 839-8500. |  | 1 | 1 unit | Nationwide distribution: CA only. | Terminated | May 23, 2012 | 1621 | 1
Z-1645-2012 | 61770 | No | MTG |  |  | Immunology | PMA | Access Hybritech Free PSA, Part Number: 37210 |  |  | October 11, 2011 | 2011 | Beckman Coulter Inc. | A recall was initiated because Beckman Coulter has confirmed a change to the recommended diluents for the Access Hybritech free PSA reagent. | Process design | Beckman Coulter sent a Product Correction letter dated October 11, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed regarding a change to the recommended diluents for the Access Hybritech free PSA reagent. Patient safety is not affected by this change." Do not dilute elevated free PSA patient samples with Wash Buffer II." Do not dilute elevated free PSA patient samples using the UniCel DxI onboard dilution feature." Dilute elevated free PSA samples with Access Hybritech free PSA Calibrator S0 only." A review of historical patient results is not required because there is no effect on clinical interpretation.* The UniCel DxI onboard dilution feature is only available in limited geographies." An updated version of the UniCel DxI Assay Protocol File (APF) will remove the Access Hybritech free PSA (dfPSA) onboard dilution assay from the list of available assays." The Access Hybritech free PSA IFU will be updated to remove Wash Buffer II as an acceptable diluent for elevated patient samples. Correct any procedure manuals or other documents created in your laboratory that list Wash Buffer II as an acceptable diluent for the Access Hybritech free PSA assay.Please share this information with your laboratory staff. If you have forwarded any of the affected product listed above to another laboratory, please provide a copy of the letter to them.Please complete and return the enclosed response form within 10 days so we are certain you received this important communication.If you have any questions regarding this notification, please contact our Customer Support Center" Via our website, http://www.beckmancoulter.com/customersupport/support" Via phone, call 1-800-854-3633 in the United States and CanadaOutside the United States and Canada, contact your local Beckman Coulter Representative. We apologize for the inconvenie |  | 90965 | 90,965 units total (24,951 units in US) | Worldwide Distribution -- USA (nationwide ) and the countries of Algeria, Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia, Herzegovina, Brazil, Bulgaria, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Djibouti, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, French, Polynesia, Georgia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Lebanon, Libyan, Arab, Jamahiriya, Macao, Mexico, Moldova, Republic of Morocco, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Panama, Philippines, Poland, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab, Emirates, United Kingdom, Uruguay ,Venezuela, Viet Nam, Yemen and Zimbabwe. | Terminated | March 18, 2013 | 524 | 90965
Z-1646-2012 | 61809 | No | OUO | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/MRI System v.3.7.1, 510(k) #K103483, Model #882380. |  |  | April 13, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were identified and corrected by Philips Medical for their Ingenuity TF PET/MRI Systems running software version 3.7.1. The firm states that these non-conformances stemmed from software and/or documentation errors, and that software corrections are being performed in order to 'improve the overall quality and functionality of the Ingenuity TF CT/MRI S | Software design | The firm, Philips Medical, sent an 'URGENT - Medical Device Correction' notification letter dated April 13, 2012 to all affected customers via Certified mail. The letter informs the customers of the firm's upcoming software update. It tells the customers that the software for the Ingenuity TF PET/MR system unit will be upgraded from version 3.7.1 to version 3.7.2 and that the update/upgrade will provide a number of improvements to the overall quality and functionality of the Ingenuity TF PET/MR System. The notification document states that Philips will be installing the system update to version 3.7.2 software through a Field Change Order. This will be implemented free-of-charge and will include complementary PET/MR application training. For further information, call toll-free telephone number which they can call 1-800-722-9377 and follow the menu options to reach a Customer Solutions Engineer. |  | 6 | 8units (2 units in U.S. &amp; 6 units in other countries) | Worldwide distribution: USA (nationwide) including states of: OH and NY; and countries of: China, Finland, Germany, South Korea, Spain, and Switzerland. | Terminated | August 05, 2013 | 479 | 6
Z-1644-2012 | 61746 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, Versions 4.30.00-4.64.00. |  |  | April 30, 2012 | 2012 | Computerized Medical Systems Inc | XiO will underestimate the dose delivered when the closed leaf gap is set to 0mm on XiO and a value other than 0mm is set on the accelerator. | Software design | Computerized Medical Systems, Inc. sent an Important Safety Notice the week of April 30, 2012, by regional support managers via e-mail, fax, or regular mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. North American customers were instructed to go to www.elekta.com, select the SupportPlus Login at the top of the page, enter their portal, select Downloads/Updates to download the latest patches.For Customers outside North America, go to www.elekta.com,select the SOFTWARE button on the left, Treatment PlanningSoftware, then Contact TPS Support tab, Download TPSSoftware Patches..............................................................Customers were also instructed to distribute this notice to any and all users of CMS software at their organization who are potentially affected by this issue.For questions regarding this recall call 408-380-8023. |  | 1652 | 1,652 sites received the versions | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | April 01, 2014 | 701 | 1652
Z-1651-2012 | 61812 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MLCi/MCLi2. |  |  | April 19, 2012 | 2012 | Elekta, Inc. | The Y2 diaphragm block was found loose. | Nonconforming Material/Component | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated April 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to carry out the recommended weekly physics check of the X-ray and light field size calibration. If an error is found, customers are instructed to contact their local Elekta service engineer. |  | 386 | 386 | Nationwide Distribution including Puerto Rico. | Terminated | October 06, 2014 | 900 | 386
Z-1648-2012 | 61811 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Accessories for the Captiva Spine SmartLOX Cervical Plate System. Products are labeled in part: |  |  | November 22, 2011 | 2011 | Captiva Spine, Inc | Captiva Spine is recalling the SmartLOX Cervical Plate components 12 &amp; 14 mm Drill Bits, Single Barrel Variable Drill Guides, and Double Barrel Fixed Drill Guides. The components are part of the instrumentation kit used to complete an anterior screw fixation of the cervical spine. There is a potential risk for the drills to not properly fit through the drill guides due to an overlap in the dime | Nonconforming Material/Component | Captiva Spine notified customers of the recall by letter on November 23, 2011, and advised that the affected products would be replaced. Customers were asked to return the drill guides and unused drills with the provided return labels.For questions customers should call 877-772-5571.For questions regarding this recall call 561-277-9480. |  | 212 | 212 units | Nationwide Distribution including TX, IN, UT, AZ, GA, CA, AL, MI, and MN. | Terminated | November 18, 2013 | 727 | 317
Z-1650-2012 | 61701 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Celsus prosthetic foot, Product Number 103658-00. The Celsus is intended for low impact level lower limb prosthesis users weighing up to 300 lbs. |  |  | April 05, 2012 | 2012 | College Park Industries, Inc. | An incorrect shorter part ( pyramid) was installed in the foot. A foot of this size and configuration requires a longer pyramid. Customers requiring the extra firm configuration may have a minimal risk of component failure, which could be implicated in a potential fall. | Mixed-up of materials/components | All customers who received affected product were initially contacted via phone on April 5, 2012 and made aware of the recall, including instructions to notify customers immediately and arrange replacements. A "Recall Notice" letter dated April 24, 2012 was sent to customers describing the problem and the recommended actions. Contact reference person is Kimberly Light, 1 800-728-7950, info@college-park.com. |  | 11 | 11 | Worldwide Distribution - USA, including the states of CA, MA, ME, WA, TX, VA, AZ and the country of England. | Terminated | July 31, 2012 | 117 | 11
Z-1653-2012 | 61732 | No | MUI | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Other-Sonic Generic Ultrasound Transmission Gel 5 Liters Hypoallergenic |  |  | March 02, 2012 | 2012 | Pharmaceutical Innovations, Inc. | Other-Sonic GENERIC ULTRASOUND TRANSMISSION GEL is being recalled due to possible microbial contamination. An FDA Safety Communication was issued for this product on April 18, 2012. | Employee error | Phamaceutical Innovations Inc. sent an Urgent Message Notice letter dated March 02, 2012 with a followup letter dated March 9, 2012 and an Inventory Withdrawal letter dated March 23, 2012 via email and fax to all affected customers. Customers were instructed to examinne inventory, refrain from distribution and return remaining inventory for a replacement. For questions call 973-242-2900. |  | 40164 | 40,164 units in distribution worldwide - unable to obtain units in distribution Nationwide | Worldwide Distribution - United States (nationwide) including the states of: FL, MA, MS, NJ, OH, TN and VA, and the countries of: Canada and Australia. | Open, Classified |  |  | 40164
Z-1656-2012 | 61754 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | ECMO Heaters Model 333W, 115V - Part #86135, and 230V &amp; 240V - Part #86136, Cardiopulmonary by-pass temperature controllers. The ECMO Heaters are used in conjunction with a heat exchanger to warm the fluid temperature. The device is composed of a heater, a circulating pump, fan, safety high limit, and a microprocessor board. Water is heated and pumped from the unit to the heat exchanger. The device can be set so that the operator can monitor the actual water temperature or, with the use of a 400 series probe, the fluid (external) temperature. |  |  | February 03, 2012 | 2012 | Cincinnati Sub-Zero Products Inc | Design error. The probe jack contacts (terminals on certain of the firm's ECMO Heater devices are located too close to the solid state relay contacts (terminals) which can lead to the terminals coming into contact with one another. If this situation occurs, a Service Technician, Operator/User, and/or Patient could be exposed to electric shock during device servicing or use. This situation was ori | Device Design | Cincinnati Sub-Zero (CSZ) issued an 'URGENT Medical Device Field Action Letter to their customers dated January 31, 2012. The letter was mailed out starting on February 3, 2012. The letter includes: a description of the product and problem, describes the affected Serial Number Range subject to recall/correction and instructs the customers to discontinue using the affected serial-numbered units.The customers are instructed to return the product to their Service Department or Biomedical Engineering Dept. for correction and to complete and mail to CSZ the attached 'Field Action Form.' For additional information, the firm can be reached on a toll-free # 1-800-989-7373. |  | 298 | 298 units (115V=295 and 230V/240V=3) | AL, AR. AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL,K IN,, KS, LA, MA, MD, MI, MN, MO, NC, NE, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, and WI | Terminated | February 26, 2013 | 389 | 298
Z-1654-2012 | 61621 | Yes | FIR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PlasmaPrO" |  |  | November 29, 2010 | 2010 | Apheresis Technologies, Inc. | Apheresis Technologies, Inc., Palm Harbor, FL, initiated a voluntarily withdrawal on November 29, 2010 of their five (5) units of the Plasma Pro Plasma Pump and later classified as a recall (removal) rather than a withdrawal, Model # PP-05, Serial Numbers: E08003041, E08003042, E08003043, E08003044 and E08003045. | PMA | Apheresis Technologies sent a Letter by FedEx to their two consignees. The letter notified customers of the voluntary recall initiated, affected product and description of the problem. For questions call 727-784-2353 or 1-800-382-5556. |  | 5 | 5 units | US Nationwide Distribution including the states of Colorado and Utah. | Terminated | May 25, 2012 | 543 | 5
Z-1655-2012 | 61825 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions, GuideLiner V2 Catheter, Rx only, Sterile EO, Catalog Number 5571, 6F, Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369. |  |  | January 17, 2012 | 2012 | Vascular Solutions, Inc. | Vascular Solutions has became aware of a potential defect with their 6F GuideLiner V2 catheters (Model 5571) with lot number 555038. Use of this product could result in the catheter shaft separating from the pushwire on the device. This may cause a procedural delay and/or injury possibly requiring medical intervention. | Material/Component Contamination | Vascular Solutions sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 17, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product , immediately remove the product and place the product in a secure area. Additionally, customers were asked to complete and return a Field Action Customer Inventory Form to the Customer Service Department via fax at 763-656-4251 or e-mail at customerservice@vasc.com. Customers may contact their local sales representative for questions concerning this recall. |  | 330 | 330 | Nationwide Distribution including the states of AL, AZ, AR, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LAM ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, WA, WV, and WI. | Terminated | July 19, 2012 | 184 | 330
Z-1657-2012 | 61692 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | Becton Dickinson Phoenix System Software, catalog # 448035/448046, component of BD automated microbiology system, and used with BD Gram Positive MIC or MIC/ID panels. |  |  | April 10, 2012 | 2012 | Becton Dickinson &amp; Co. | BD identified an increase in vancomycin-resistant strains that may exhibit erroneous results on all PMIC or PMIC/ID panels and could lead to ineffective treatment regimen in affected patients. | Nonconforming Material/Component | Becton Dickinson (BD) sent an Urgent Field Correction Action letter dated April 2012 and an Update letter dated April 18, 2012 to all affected customers via e-mail and UPS overnight delivery. The letter identified the affected products, problem and actions to be taken. The letter provided specific workaround instructions for device users pending software upgrades. A BD Service Engineer will contact affected customers to arrange for the installation of the updated software when it becomes available. The letter included an Acknowledgement Form to be completed. For questions contact Technical Services at 800-638-8663. |  | 3454993 | 3,454,993 total products (U.S.) | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, AZ, CA, CO, CT ,DC ,FL , GA, IA, IL, IN, KY, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, AND WY. and the countries of Canada, Australia, Philippines, Thailand, Malaysia, India, New Zealand, Belgium, Singapore, Japan, Taiwan, China, Chile, Colombia, Korea, Brazil and Venezuela. | Terminated | August 08, 2013 | 485 | 3454993
Z-1673-2012 | 61540 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | HPBLUE: HARMONICÂż Blue Hand Piece; HP054: HARMONICÂż Hand Piece; Non-Sterile products are individually packed in a poly bag and then placed into an individual carton. Mfr. By: Ethicon Endo-Surgery, LLc, 475 Calle C, Guaynbo, PR 00969. |  |  | March 16, 2012 | 2012 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated a voluntary recall due to some lot numbers of the HARMONICÂż Hand Piece (HP054) and the HARMONICÂż Blue Hand Piece (HPBLUE), which impacts their compatibility with the HARMONICÂż Generator 300 (GEN04). The problem occurs after successful set up and initial use of HARMONICÂż Hand Pieces (HP054 and HPBLUE) with the Generator 300 (GEN04). The procedure is interrupted when | Nonconforming Material/Component | Ethicon Endo-Surgery Inc sent an URGENT: Product Correction Notification letter dated March 2012 to all affected customers. The letter identified the product, the problem, the action to be taken by the customer. Labels were enclosed that are used to identify the recalled product. Customers were instructed to fill out the Business Reply Form and fax it to Stericycle at 1-888-912-7091 within three business days for a credit.For questions regarding this recall call 513-337-3419. |  | 1896 | 1896 instruments | Worldwide Distribution-USA including: AK,AL,AR,AZ,CA,CO,CT,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NCND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,RI,SC,SD,TN,TX,UT,VA,VT,WA,WI,WV, and WY.and the countries of: Argentina, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Cyprus, Ecuador, Egypt, Greece, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Korea (South),Kuwait, Lebanon, Malaysia, Mexico, Netherlands, Antilles, Pakistan, Panama, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Taiwan, Thailand, Turkey, United Arab Emirates Uruguay, Venezuela &amp; Yemen. | Terminated | March 05, 2013 | 354 | 1896
Z-1675-2012 | 61760 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Essenza, Nuclear Magnetic Resonance Imaging System |  |  | March 30, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is an error in the compatibility data sheet which is part of the System Owner Manual | Error in labeling | Siemens issued a Customer Safety Advisory Notice letter to affected customers on March 30, 2012. The letter identified the affected product, problem, potential risk and what steps can the user take to avoid the potntial risk of this issue. Siemens is releasing corrected pages of the Operator Manual in order to deliver clear guidance and recommendation for the users. Customers are instructed to ensure that the safety advisory and updated pages are replaced in the System Owner Manual. |  | 97 | 97 | US Nationwide Distribution - including the states of AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY, LA, MS, MO, NE, NJ, NY, NC, OH, OR, PA, TN, TX, WA, WI, WY and Puerto Rico. | Terminated | April 03, 2014 | 734 | 97
Z-1436-2012 | 52013 | No | LNX |  |  | General Hospital | 510(k) | Abacus TPN Calculating Software |  |  | January 19, 2009 | 2009 | Baxa Corporation | Ordering salt-based ingredients on an ion-based template resulted in medication errors when calculating amounts of ingredients required. | Labeling design | Baxa Corporation sent a "Safety Alert" letter dated January 19, 2009 via FedEx to all affected customers. The letter described the affected product, problem, and actions to be taken. Customers were instructed to ensure that users create either a salt-based template OR an ion-based template for ordering, and that both methods of ordering are not used in one template. Customers were provided a list of safe practices to follow on the enclosed ISMP Safety Alert. Customers were advised upon receipt of the Safety Alert that Baxa Technical Support will review and address any issues with the customer on their configuration setup for calcium and magnesium. For questions contact Baxa Technical Support at 1-800-678-2292. |  | 667 | 667 units | Worldwide Distribution - US (Nationwide) and the countries of: Canada, China, Australia, Panama and Columbia. | Terminated | September 12, 2012 | 1332 | 667
Z-1674-2012 | 61532 | No | KMA | Powered nasal irrigator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Water Pik Sinusense battery operated device including a blue semi-transparent water reservoir, a white hand held button initiated injector device, and a soft seal nozzle affixed to the white injector. The product is sold individually in white, blue red and orange cartons with a photo of the device as well as a stream of water flowing through it. The contents of the container include the device as well as a package insert with utilization instructions. |  |  | April 03, 2012 | 2012 | Water Pik, Inc. | Customer complaints related to cross contamination of sinus rinse solution between the battery compartment and solution reservoir used for rinsing nasal passages. | Device Design | Water Pik has generated a Press Release letter dated May 2012. The letter included safety information and instructions for the consumer to contact Waterpik for a replacement model if they experience water intrusion into the battery compartment of their device. For additional information go to www.waterpiksinusense.com/exchange. |  | 105257 | 105257 individual units | US Nationwide Distribution. | Terminated | August 07, 2012 | 126 | 105257
Z-1677-2012 | 61338 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoration HA Reduced Neck 255 mm Long Stem |  |  | October 27, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | An incorrect laser mark was noted on specific lots of product, incorrectly implying a stem size of 9, when the product is actually size 8. | Error in labeling | Stryker sent an "URGENT PRODUCT RECALL" letter dated October 27, 2011 and October 28, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs the customers to return the product to the firm. An Acknowledgement Form was included for customers to complete and return. Contact the firm at 201-972-2100 for questions regarding this recall. |  | 31 | 31 Units | Nationwide Distribution and the country of Canada. | Terminated | August 22, 2013 | 665 | 31
Z-1678-2012 | 61814 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Altrus Energy Source REF 60-9500-120 Tissue Fusion Energy Source ConMed Corporation 525 French Rd, Utica, NY 13502, MDSS GmbH, Germany. |  |  | May 07, 2012 | 2012 | ConMed Electrosurgery | An investigation of the ConMed Electrosurgery Altrus Energy Source revealed that the device could have been calibrated incorrectly resulting in error messages or delays in surgery. | Equipment maintenance | ConMed Electrosurgery sent an "URGENT: MEDICAL DEVICE" letter dated May 2, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to review their inventory for the affected product and to complete and return a Business Reply Form to the firm via fax to 303-699-9854. |  | 122 | 122 units | Nationwide Distribution and the country of Canada. | Terminated | June 19, 2012 | 43 | 122
Z-1676-2012 | 61756 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | SOLUS Anterior Lumbar Interbody Fusion Spinal |  |  | March 19, 2012 | 2012 | Alphatec Spine, Inc. | Alphatec Spine has received several complaints (n+4) on the Alphatec Solus Distal Deployment Wrench for twisting or bending, which has been discovered either during or after instrumentation use. A fifth complaint received in which it was reported that the spline tip of the Alphatec Solus Proximal Deployment Wrench had sheared from the instrument upon use. | Device Design | Alphatec Spine sent an Urgent Medical Device Recall Notification dated March 30, 2012, to all affected customers. The notification letter informed customers of the removal/recall and included instructions on what to do with the recalled products. Initial letter was delivered via FedEx, and followed by telephone call with a total of 4 attempts. Customers were instructed to return all of the affected product. Upon receipt of of the Medical Device Recall Notification, customers were instructed to contact Alphatec Spine's Customer Service Department immediately for instructions on how to return the instruments and fill out the lower portion of this letter to confirm that they have read the notification and have taken all necessary removal actions: Customers should return a signed copy of the form to:Alphatec Spine, Inc.5818 El Camino Real,Carlsbad, CA 92008ATTN: Rand JadanForms may also be returned via email to rjadan@alphatecspine.com or fax to 760-431-0289.For questions regarding this recall call 760-431-9286. |  | 18 | 18 devices | Worldwide Distribution - USA (nationwide) and Germany | Terminated | June 20, 2012 | 93 | 18
Z-1679-2012 | 61821 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Altrus Thermal Tissue Fusion, Tissue Fusion Handpiece 5 mm O.D., 36 cm Shaft, Catalog Number 60-9522-001, Sterile/EO. |  |  | May 07, 2012 | 2012 | ConMed Electrosurgery | The devices were distributed prior to the firm completing performance qualifications for their package sealing and adhesive bonding. | Process control | ConMed Electrosurgery sent an "URGENT: DEVICE RECALL" letter dated May 2, 2012 to all affected customers. The letter identifies the product, problem,and actions to be taken by the customers. The letter instructs customers to review their inventory and to complete and return a Business Reply Form via fax to 303-699-9854. |  | 153 | 153 units | Class II Recall-Nationwide Distribution-including the states of OH, GA, IL, PA, AZ, KY, FL, NY, MO, TN, TX, AL, IN, SC, MD, CT, NJ, and CA. | Terminated | June 05, 2012 | 29 | 237
Z-1684-2012 | 61808 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial |  |  | April 23, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale. | Software change control | Siemens initiated a Customer Safety Advisory Notification letter Field correction on April 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers' local Customer Service Engineer will install an update to their system as soon as it becomes available. For questions customers were instructed to contact their local service support person. Customers were also asked to pass the notification leter to all those within their organization who need to be aware of the issue until the corrective action is completed.For questions regarding this recall call 650-694-5747. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Austria, Australia, Belgium, Canada, Switzerland, Chile, China, Germany, Denmark, Ecuador, Spain, France, Great Britain, Greece, Hong Kong, Hungary, Indonesia, Israel, India, Italy, Jordan, Japan, South Korea, Kuwait, Lebanon, Myanmar, Mexico, Malaysia, Netherlands, Norway, New Zealand, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Sweden, Slovenia, El Salvador, and Thailand. | Terminated | July 16, 2013 | 449
Z-1661-2012 | 61895 | No | FPK | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Epidural Admin Set with Non-Vented Bag Spike and 0.2 Micron Filter Package Assembly, product code: 340-4126, Method of Sterilization using Ethylene Oxide. |  |  | May 18, 2012 | 2012 | MOOG Medical Devices Group | Moog Medical Device Group became aware on March 23, 2012 via customer complaint, of a reversed pump segment on an Administration Set. This malfunction could reverse the flow of fluid or medication backwards from what was intended. | Process design | The firm, Moog Medical Device Group, sent an "Urgent Field Safety Notice" dated May 21, 2012 via certified letter and a press release direct to consignees/customers. The notice describes the product, problem and actions to be taken. The customers were instructed to do the following: If purchased by distributor -remove suspect product from inventory; contact your distributor to arrange replacement product and request a shipping return label as per the distributor's process and ship product to their location.If purchased by Moog-remove suspect product from inventory, contact Moog Customer Service at 1-800-970-2337, prompt #7 and Moog customer service will provide a call tag for convenient product return.For additional information or clarification, please contact the Post Market Surveillance Manager at (801) 264-1001, ext. 224 or via email at: lbrewer@moog.com. |  | 65160 | 65,160 | Nationwide distribution: USA including states of: AZ, CA, FL, IL, KS, MA, MI, MD, NC, ND, NY, OK, OH, and TX. | Terminated | February 22, 2013 | 280 | 544900
Z-1683-2012 | 57371 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Terumo BCT Lakewood, CO 80215, Catalog No 61000 Spectra Optia Apheresis System, Country of Origin: US. |  |  | May 04, 2012 | 2012 | CaridianBCT, Inc. | Defective product may result in decreased efficacy. | Component design/selection | Terumo BCT sent a Voluntary Medcal Device Safety Alert letter dated May 18, 2012 to each of their consignees via certified letter. The letter identified the affected product, problem, actions to be taken and instructions for healthcare providers until a corrective action is identified and implemented. Customers were instructed to share this safety alert with all those who need to be aware and complete the attached Acknowledgement Form and fax to +1.303.876.9277 or email to Terumo BCT. For questions contact your Terumo BCT Representative, the Terumo BCT Support Center at (U.S. Toll Free +1.877.339.4228) or +1.303.231.4357, or your local Terum BCT Customer Service Office. |  | 372 | 372 units | US Nationwide Distribution and the country of Canada. | Terminated | November 27, 2012 | 207 | 372
Z-1690-2012 | 61762 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 16 - Model #728246, CT Scanning System is subject to recall/correction |  |  | March 26, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Component defect. Following their receipt of a customer complaint, Philips Medical determined that the metal blocks which support the scissor slide rails underneath the scanning table (couch) of several of the firm's CT scanning systems were not manufactured to specification and may crack. | Nonconforming Material/Component | Philips issued an 'URGENT Medical Device Correction' letter dated March 27, 2012 to all affected customers. The letter identified the affected products, problem, hazard involved, how to identify affected products, and actions to be taken by customer/user. The letter informed customers that Philips intends to dispatching service personnel to each customer site in order to correct this issue. In the meantime, the letter recommend customer to contact their Philips representative if they hear unusual noises when moving the couch vertically, or if there is unintended vertical motion of the couch. For questions contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center 1-800-722-9377, option 5: Enter Site ID or follow the prompts. |  | 10 | 10 system units were distributed in the U.S. | Worldwide Distribution - U.S. (Nationwide) including the states of: CA, FL,, HI, IA, IN, KY, LA, MA,, MI, MO, MS. NM, NV, OH, PA, PA, SC, TN, TX, WA, and Puerto Rico | Terminated | February 22, 2013 | 333 | 35
Z-1689-2012 | 61779 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Philips Practix Convenio Mobile X-ray System |  |  | May 02, 2012 | 2012 | Philips Healthcare Inc. | Unexpected Low-level X-ray pulse | Device Design | Philips sent a Urgent Field Safety Notice dated April 26, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to please review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of the communication. Please retain a copy with the equipment instruction for Use. A Philips Service Engineer will contact account as soon as the Field Action Kit is ready to be implemented. Communicate with Philips with regard to this program, please reference Field Change Order 70400049.Further questions please contact your local Philips representative at 1-800-722-9377.On 2/6/15, Philips reported that the recall was extended to an additional 88 units that were unit stock that did receive the correction. Philips issued recall notifcation letter (FCO70400049) on 1/20/2015 to users. Control boards will replaced by a Philips Service Engineer. |  | 76 | 10 US; 66 internattional | Worldwide Distribution-- USA ( nationwide ) including the states of IA, LA, LA, ND, NY, OH and TX and the countries of Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Germany, Hong Kong, Indonesia, Italy, Malaysia, Mexico, Netherlands, Philippines, Qatar, Russia, Saudi Arabia, Singapore, South Africa,Thailand and Ukraine. | Open, Classified |  |  | 76
Z-1688-2012 | 61858 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance iCT and iCT SP are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | May 07, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | On 3/5/2012, Philips Healthcare was notified of one incident involving a Brilliance iCT scanner that resulted in damage to the gantry and slight injury to the user when the heat exchanger detached from the cooling unit due to some missing welds. | Nonconforming Material/Component | Philips Healthcare sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated May 8, 2012 to all affected customers. The letter identifies the product and problem. No action is required by the customers. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. Philips Healthcare will inspect the cooling units on all iCT and iCT SP systems for this issue and if the welds on a cooling unit do not pass the inspection, Philips will replace the cooling unit free of charge. |  | 372 | 372 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, DC, DE, FL, GA, IA, IL, IN, KY, MA, MD, MN, MO, MS, NC, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, VT, WA and WV and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Great Britain, Greece, Hungary, Iceland, India, Indonesia, Iraq, ISRAEL, Italy, Japan, Jordan, Korea, Lebanon, Mexico, Netherlands, New Zealand , Norway, Philippines, Poland, Romania, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Arab Emirates and Uzbekistan. | Terminated | February 26, 2013 | 295 | 372
Z-1686-2012 | 61672 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia E Series SPECT System; Single-Photon Emission Computed Tomography. The Siemens Symbia E system is a high performance dual detector variable angle system with flexible patient positioning for planar imaging, whole-body scans, and SPECT studies. |  |  | May 15, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | Siemens has identified a potential patient safety issue for e.cam and Symbia E camera systems that have had the radial drive motor assembly replaced during a Service action. There is a remote chance that during the motor replacement the radial drive gear may have been replaced incorrectly. If the gear is installed incorrectly the mechanical brake may not engage creating a patient hazard if the sys | Employee error | On May 15, 2012, Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated April 11, 2012 to all affected customers. The letters informed customers of the potential patient safety issue for e.cam and Symbia E camera systems that have had the radial drive motor assembly replaced during a Service action. Customers were informed that a Siemens Service Engineer will be inspecting the radial drive gear on their camera system to ensure it is functioning correctly. For questions or assistance they were instructed to contact Siemens Medical Solutions USA, Inc. at 800-888-7436. |  | 9 | 9 units | Worldwide Distribution, US (Nationwide) and the countries of Australia, Belgium, Canada, China, France, Germany, Italy, Japan, Republic of Korea, Mexico, Netherlands, Portugal, Spain, Taiwan and the United Kingdom. | Open, Classified |  |  | 146
Z-1695-2012 | 61908 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Monoject Insulin Safety Syringe 3/10mL 30 X 5/16" |  |  | May 21, 2012 | 2012 | Covidien LLC | Safety shield may separate from the syringe, exposing the needle, which could potentially result in a needle stick | Equipment maintenance | Covidien sent an Urgent Product Recall letter dated May 21, 2012, via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop using any affected product. Distributors were requested to notify their customers. Customers were instructed to complete the Monoject Insulin Safety Syringe recalled product return form even if they have no product to return. The form should be faxed to 1-800-895-6140 or e-mailed to sdfeedback@covidien.com. Customer Service questions regarding the recall please use one of the following methods.Telephone:1-800-722-8772 option I (Weekdays, 8am to 6:30pm Eastern)e-mail:sdfeedback@covidien.com |  | 77500 | 77,500 syringes | Worldwide Distribution - USA (nationwide) and the countries of Canada, Australia, Belgium, and Panama. | Terminated | June 10, 2013 | 385 | 28069500
Z-1699-2012 | 61790 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System Software Version 1.6 &amp; Below |  |  | April 26, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling VITROS 5600 Integrated System Software Version 1.6 and below because of an anomaly that may inadvertently add a surfactant to a CurveTip position. | Software design | Ortho Clinical Diagnostics (OCD) sent an Urgent Product Correction Notification letter dated April 26, 2012, via FedEx overnight courier to US consignees with VITROSÂż 5600 Systems. The letter identified the affected product, problem and actions to be taken. Foreign affiliates were informed by email on April 26, 2012 of the issue and instructed to notify their consignees of the issue and actions. OCD recommend customers to group all DAT assays into a single run as described in the letter, until the next Software Version 2.0 is installed on their VITROSÂż 5600 System. Customers were instructed to complete and return the Confirmation of Receipt form. For questions contact Customer Technical Services at 1-800-421-3311 |  | 1357 | 1357 units | Worldwide Distribution - US (nationwide) | Open, Classified |  |  | 1357
Z-1698-2012 | 61623 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Transforminal Lumbar Interbody Fusion (TLIF) PEEK Inserter used in the Genesys Spine Apache" Interbody Fusion System. |  |  | August 24, 2011 | 2011 | Genesys Orthopedic Systems, LLC | The inserters have the potential to break. | Reprocessing Controls | Genesys Orthopedic Systems, LLC decided to recall the affected product and called their customers on August 24, 2011. The firm removed the inserters from distribution from 8/24-31/2011. |  | 4 | 4 inserters | Nationwide Distribution including Florida and Texas. | Terminated | December 04, 2012 | 468 | 4
Z-1700-2012 | 61755 | No | CFP | Immunoreactive insulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE --- Insulin --- Catalog Number: LKIN1 (100 test), LKIN5 (500 tests) --- Test Code: INS, Color: Orange |  |  | April 30, 2012 | 2012 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics has observed a negative bias in the Immulite Insulin assay. | Component change control | Siemens sent an Urgent Device Recall letter dated April 2012 via fax to affected U.S. customers and an Urgent Field Safety Notice to customers outside the U.S. The letter identified the affected kit lots, reason for voluntary recall, risk to health and actions to be taken by the customer/user. Customers were instructed to discard any inventory of the affected kit lots and complete and return the attached Field Correction Effectiveness Check questionnaire within thirty (30) days to request a no-charge replacement kits. Customers were requested to post this field correction by each IMMULITE system to alert users to the situation. For questions contact your Siemens Technical Solutions Center. |  |  |  | Worldwide Distribution - US (nationwide) | Terminated | March 16, 2017 | 1781
Z-1702-2012 | 61673 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga UIS 2009 (previously known as Microsoft Amalga Unified Intelligence System), Calculator/data processing module, for clinical use. |  |  | April 12, 2012 | 2012 | Microsoft Corporation Health Solutions Group | Microsoft Amalga UIS 2009 has an issue with the display of lab results in the Lab Details window of the lab result viewer. The Lab Details view does not freeze/pin the row header when the user scroll horizontally. | Software design | Microsoft Corporation Health Solutions Group sent an "URGENT FIELD SAFETY NOTICE" dated April 12, 2012 to all affected customers via e-mail. The letter identifies the product, problem, and actions to be taken by the customers. The firm plans to revise its labeling for Amalga to clarify that the row headers will not be visible as the user scrolls horizontally. Customers were asked to complete and return to the firm a response form via fax at 425-706-7329 or e-mail to vlpigin@microsoft.com. |  | 13 | 13 units | Worldwide Distribution-USA (nationwide) including the states of AR, MA, MD, NC, NY, and WA and the country of China. | Terminated | December 12, 2012 | 244 | 13
Z-1707-2012 | 61823 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Heparin Lock Flush Solution 100 Unit/mL, USP |  |  | May 07, 2012 | 2012 | Excelsior Medical Corp | Some lots of Heparin Lock Flush, USP Syringes (3 mL fill in 10 ml syringe, 4 mL fill in 10 ml syringe, 5 mL fill in 10 ml and 6 ml in 10 ml syringe), failed Anti-Factor IIa assay for potency. | Nonconforming Material/Component | Excelsior Medical sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue using the affected product and to complete and return an attached Customer Reply Form via fax to (866) 807-3714. Contact Excelsior Medical Customer Service at (800) 487-4276 for questions regarding this recall. A second letter which included additional lots was sent out on 5/25/2012 with the date of 5/24/2012. A return response form was also included. |  | 34230500 | 34,230,500 units (596 lots) | Worldwide Distribution-USA (nationwide) and the countries of Brazil, Canada, and Colombia. | Terminated | September 12, 2013 | 493 | 34230500
Z-1706-2012 | 61766 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Xia 3 Anti-Torque Key Manual Surgical Instrument; Catalog Number 48237026. |  |  | January 12, 2012 | 2012 | Stryker Spine | Stryker received four reports regarding a handle which separates at the point where it is welded to the shaft on the Xia 3 Anti-Torque Key during surgery. | Mixed-up of materials/components | Stryker Spine sent an Urgent Product Recall letter dated January 12, 2012 with an attached Customer Response forms via FEDEX to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory to identify the affected product and return using the prepaid mailing label provided. The letter asked customers to complete the Customer Response Form and fax to 201.760.8370. For questions call 201.760.8298. |  | 95 | 95 units | US Nationwide Distribution | Terminated | March 21, 2013 | 434 | 95
Z-1703-2012 | 61834 | No | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Brand Name: QuietCare, Model Numbers QC200403, v.02; QC200403, v.03; QC200403, v.04 -- Classification Name: System, Environmental Control, Powered. -- Firm on label: Manufacturer Living Independently Group, Inc., A wholly owned subsidiary of General Electric Company, PO Box 1466, New York, NY 10116 USA, 1-866-772-8243 --- The system is a passive behavior-monitoring device. QuietCare uses strategically placed, unobtrusive, passive wireless infrared pyroelectric sensors in the senior's residence to detect their movement. Data from these sensors are sent to a base station that periodically transmits the information to a centralized secure file server through standard phone lines over a toll-free number. Sophisticated algorithms analyze the data, learning each individual's normal patterns of behavior. When the system detects significant deviations from these baselines it automatically alerts designated caregivers. QuietCare also generates alerts of late wake-ups, potential bathroom falls and if household temperatures become dangerously low or high. The behavioral changes that the system tracks can often serve as subtle indicators of the potential emerging health problems. These changes include: nighttime bathroom usage, entering or exiting the bedroom at a different time than usual, change in the number of times the refrigerator door is opened/closed change in the number of times the medicine cabinet door is opened/closed. The intended users of the QuietCare device are the caregivers at the Assisted Living Facilities and Independent Living Facilities. The monitored individuals (seniors) do not actively operate the device and the device function is not dependent upon the senior resident's ability to use the device. --- Class 2 device, 510(k) exempt, Device Listing Number D117156. |  |  | August 01, 2011 | 2011 | Intel-GE Care Innovations LLC | On 7/29/11, the QuietCare production system experienced a partial outage from about 15:00 hours until 22:30. During the engineering investigation into the cause of the service disruption, it was discovered that some residents did not have the wander and night motion features properly activated, and alerts were not being transmitted. Failure to generate alerts in a timely manner could result in a | Employee error | Intel-GE Care Innovations sent an "URGENT FIELD NOTIFICATION" e-mail dated August 12, 2011 to all affected customers. The notification identifies the product and problem. No action is required by the customers. Contact Customer Service at (866) 772-8243 for questions regarding this notice. |  | 61 | 61 units were identified as affected | Nationwide Distribution-USA (nationwide) including the states of Arizona, California, Connecticut, Idaho, Illinois, Minnesota, New Jersey, New York, Ohio, Pennsylvania, Texas, and Utah. | Terminated | June 05, 2012 | 309 | 61
Z-1705-2012 | 61919 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Product is the Spacelabs Medical Xprezzon Bedside Monitor, Model 91393. |  |  | May 22, 2012 | 2012 | Spacelabs Healthcare, Llc | The screw used to connect the external power supply to the Spacelabs Medical Xprezzon Bedside Monitor, Model 91393, may damage an internal capacitor. The damage may result in an inability of the monitor to power on, power off, or involuntarily shut down. | Device Design | The firm, Spacelabs Healthcare, sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 22, 2012 via return receipt requested, to US customers. On May 29, 2012 a translated customer letter was emailed to all international subsidiaries and distributors of record. The letter describes the product, problem and actions to be taken. The customers were instructed to notify staff of the potential failure and to only tighten the power supply screw with their fingers. In addition, the firm will schedule a time to upgrade the monitors at the customers site.If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 1898 | 1898 monitors (Domestic: 1188) (International: 710) | Worldwide distribution: USA (nationwide) and countries including: Australia, Canada, Chile, China, Czech Republic, Ecuador, France, Germany, Greece, Hong Kong, India, Italy, Kuwait, Malaysia, Netherlands, Panama, Poland, Portugal, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Switzerland, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | May 07, 2013 | 350 | 1898
Z-1704-2012 | 61550 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Knee walker, one unit per box. Manufacturer Kenstone Metal (Kunshan) Co., LTD. The intended use is to provide additional mobility to individuals who have difficulty walking. |  |  | April 06, 2012 | 2012 | Invacare Corporation | Adverse event data prompted testing of device, which revealed that the rod supporting the knee pad was not made to specifications. Upon placing a load on the knee pad, the rod may break or bend causing the walker to collapse. As a result, the user may fall. | Nonconforming Material/Component | Invacare notified all consignees and distributors with a "Urgent: Recall Information" letter dated May 16, 2012. The product issue was described and recommended actions were provided to customers. Questions were directed to 1-888-257-7912. |  | 1640 | 1640 | Worldwide Distribution--USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and the country of China. | Terminated | March 01, 2013 | 329 | 1640
Z-1772-2012 | 61767 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 6.0 MM DIAMETER 50 MM THREAD LENGTH EXTERNAL FIXATION HALF PIN HA TI-6AL-4V, REF 71932962, QTY: (1), STERILE R, Smith &amp; Nephew, Inc., Memphis, TN 38116 |  |  | April 13, 2012 | 2012 | Smith &amp; Nephew Inc | Plastic trays which encase these products have burrs on the hinged packaging which can puncture the single barrier sterile pouch, voiding sterility. | Packaging | Smith &amp; Nephew Sales Representatives and Smith &amp; Nephew International Distributors were notified of the problem with the affected product and the recall via email on April 13, 2012. All affected direct consignees were notified via Fed Ex on April 13, 2012. The Urgent Medical Device Recall letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease use and distribution of the product and quarantine them immediately for return. Customers were instructed to complete the Inventory Return Certification Form included with the letter. Customers should fax or email the form to 901-566-7975 or fieldactions@smith-nephew.com. The affected product should be returned to:Smith &amp; NephewAttn: Global Field ActionsSafety Affairs1450 Brooks RoadMemphis, Tennessee 38116If the affected product was distributed further customers were instructed to provide a copy of the recall letter to these customers. |  | 4 | 4 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Malaysia, Germany, Japan, Portugal, United Arab Emirates, Spain, France, South Africa, Norway, Sweden, Denmark, Switzerland, Belgium, China, The Netherlands, Finland, Mexico, Italy, UK, Venezuela, Colombia, Thailand, Chile, South Korea, India, and Turkey | Terminated | September 22, 2015 | 1257 | 112205
Z-1773-2012 | 61954 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Sterrad Cyclesure 24 Biological Indicator, P/N: 14324 |  |  | February 17, 2012 | 2012 | Advanced Sterilization Products | The recall was initiated by Advanced Sterilization Products because they recently determined that certain lots of raw materials that were rejected during incoming inspection acceptance test were used for the manufacture of a specific lot range of Sterrad Cyclesure 24 Biological Indicator. | Process control | Advanced Sterilization Products a Johnson &amp; Johnson company sent an Urgent Product Recall letter dated March 16, 2012, to all affected customers. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to read the "Details on the Affected Devices/Description of the problem" section in the letter. Customers were instructed to immediately identify and set aside all product listed above in a manner that ensures product will not be used. The lot number and the part number are printed on the label placed around the ampoules. Customers were instructed to return any affected product in accordance with the "Product Return Instructions" section of the letter. Customers were instructed to complete the requested information on the enclosed postage paid business reply card and return it to Stericycle even if they do not have the affected products listed on the letter in stock. Customers with questions were instructed to contact Stericycle at (877) 492-4795.For questions regarding this recall call 949-453-6400. |  | 2284 | 2,284 cases (137,040 Biological Indicators) | Worldwide Distribution - USA (nationwide) and the countries of AR, AU, BE, BR, CA, EC, HK, JP, MX, and PR | Terminated | October 17, 2013 | 608 | 2284
Z-1783-2012 | 61925 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Accessories for the Captiva Spine FuseLOX Lumbar Cage. Products are labeled in part: "FuseLOX Lumbar Cage Lumbar Interbody Fusion Device***Description:***Part Number:***Lot Number:***Quantity: 1***Material: PEEK-OPTIMA LT1, Tantalum ASTM F560***Non-Sterile-Single Use Only (See package insert for sterilization parameters and labeling.)***CAUTION: Federal (United States) law restricts these devices to sale by or on the order of a physician.***Manufactured and distributed by: CAPTIVA SPINE***captivaspine.com***967 Alternate A1A #1 Jupiter, FL 33477 Tel: 877-772-5571 Fax: 866-318-3224 Email: info@captivaspine.com***" |  |  | April 27, 2012 | 2012 | Captiva Spine, Inc | Captiva Spine, Inc. initiated a recall of FuseLOX Lumbar Implants, Trials, and Inserters after FDA advised the firm that there is inadequate support for implemented design changes to the cleared device Pivotec Lumbar Interbody Fusion Device. | Finished device change control | Captiva Spine sent a Voluntary Correction Notification letter dated May 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and return affected devices. Customers were also instructed to complete and return a Tracking/Verification Form.For questions customers were instructed to call 877-772-5571.For questions regarding this recall call 561-277-9480. |  | 136 | 136 units | USA (nationwide) including FL,TX, UT, and NJ. | Terminated | December 24, 2013 | 606 | 254
Z-1788-2012 | 37654 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Revolution XR/d; GE Healthcare, 3000 North Grandview , Waukesha, WI, 53188 |  |  | May 31, 2006 | 2006 | General Electric Med Systems LLC | An incident was reported that an operator's finger was pinched between the fixed roller and the gliding tabletop of a Proteus XR/a Table while moving the tabletop with the thumb placed over its edge and fingers underneath | Device Design | GE Healthcare sent an "URGENT SAFETY INFORMATION" letter dated May 31, 2006 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included new operator instructions and warning labels. |  | 3796 | 3,796 units | Worldwide Distribution. | Terminated | July 31, 2012 | 2253 | 3796
Z-1791-2012 | 61866 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance iCT and iCT SP are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | May 07, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips is issuing a software update (version 3.2.4) to the Brilliance iCT and iCT SP systems currently running software version 3.2.0, 3.2.1 or 3.2.3, addressing a number of non-conformances. A review was conducted of all nonconformances that are fixed by the software upgrade. None of these nonconformances result in unacceptable hazards.The nonconformance documented below had the highest sc | Software design | Philips sent an Urgent - Medical Device Correction letter dated May 8, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. Customers were instructed to review the Release Document provided them with information about the software update, new feature and enhancements. For North America and Canada customers should contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Side ID or follow the prompts.For questions regarding this recall call 440-483-4918. |  | 241 | 241 | Worldwide Distribution - USA including AL, CA, DC, DE, FL, GA, IA, IN, MA, MD, MN, MS, NC, NY, OH, OR, PA, TN, TX, VT, WA &amp; WV and the countries of Argentina, Australia, Belgium, Brazil, Canada, China, Czech Republic, Denmark, EGYPT, France, Germany, Greece, Hungary, ICELAND, India, Indonesia, Israel, ISRAEL, Italy, Japan, Jordan, Lebanon, Netherlands, New Zealand, PHILIPPINES, POLAND, Portugal, Qatar, Russia, SINGAPORE, South Korea, SPAIN, SWEDEN, Switzerland, Taiwan, THAILAND, Turkey, UNITED KINGDOM &amp; Uzbekistan. | Terminated | July 25, 2013 | 444 | 241
Z-1792-2012 | 61853 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Lipofilter 3000, model ASP-CAN-2, labeled in part ***MicroAire, 1641 Edlich Drive, Charlottesville, VA 22911, USA*** |  |  | April 26, 2012 | 2012 | MicroAire Surgical Instruments, LLC | Medical device was not manufactured under Good Manufacturing Procedures and a user may be unable to sterilize the device according to procedures. | Component change control | MicroAire Surgical Instrument sent a Device Removal Notice letter dated April 26, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were advised to recover the affected product from inventory and quarantine. Customers were instructed to contact MicroAire Customer Service Department at 800-975-8000 or 434-975-8000 to arrange for return and replacement of the affected product. For questions call 434-975-8336 or email toddm@microaire.com. |  | 15 | 15 devices | US Nationwide Distribution including the states of CA, OH and NY. | Terminated | November 21, 2012 | 209 | 15
Z-1794-2012 | 61951 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PlasmaJet Open Surgery Handpiece (Part of the PlasmaJet Neutral Plasma Surgery System) |  |  | April 27, 2012 | 2012 | Plasma Surgical Inc. | Potential failure in the glued seal between the stem and the tip of the handpiece leading to the possible leakage of a very small drop of distilled water coolant from the handpiece in use. | Nonconforming Material/Component | The firm, Plasma Surgical, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 27, 2012 to its Consignees/Customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their physical inventory; locate and quarantine or remove from service all identified product; and complete and return the attached CUSTOMER RESPONSE FORM via fax to designated persons at: (404) 678-4395 or mail to: Plasma Surgical Inc., 1125 Northmeadow Parkway, Suite 100, Roswell, Georgia 30076 (even if they no longer have any of the product in their inventory). A Plasma Surgical Representative will contact the customers to arrange for the return of the products and product replacement, should they have any of the physical inventory (product).Should you have any queries concerning this matter, please do not hesitate to contact, Vice President of Sales and Marketing, imawhinney@plasmasurgical.com or call (678) 578-4390. |  | 36 | 36 units | Worldwide distribution: USA (nationwide) including states of: CA and NY; and country of: United Kingdom. | Terminated | November 01, 2012 | 188 | 36
Z-1796-2012 | 61717 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Integra Mozaik Putty |  |  | October 11, 2011 | 2011 | Integra LifeSciences Corp. | The Instructions for Use {HIFU} {Putty Part Number 7700713001 and Strip Part Number 770151300 l for CE marked Mozaik Strip and Putty products contains an incorrect translation of the term "bone marrow aspirate". The term is incorrectly translated in Spanish and French, | Employee error | Integra LifeSciences sent a Field Safety Notice letter dated October 11, 2011 to their sole foreign consignee. The letter identified the affected products, problem and actions to be taken. The consignee was asked to remove the existing Instructions for Use (IFU) and replace it with the corrected IFU. The consignee was instructed to complete and return the attached Distributor Rework and Acknowledgment Forms provided with the customer notification letter as documentation that instructions were followed and all required actions were completed. For any questions, call: +33 (0)4 37 47 5900 , fax: +33 (0)4 37 47 5999; or email: carole.robin@integralife.com. |  | 42 | 42 units | Worldwide Distribution | Terminated | February 14, 2013 | 492 | 64
Z-1802-2012 | 61945 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Churchill Medical Systems, Inc. A VYGON Company PICC LINE TRAY Kit |  |  | May 11, 2012 | 2012 | Vygon Corporation | Manufacturer issued a voluntary recall on multiple lots of VY(TM)Flush Heparin Lock Flush Syringes therefore Vygon has subsequently initiated a recall on Vygon branded Heparin Lock Flush Syringes and kits containing the syringes. | Nonconforming Material/Component | VYCON sent an URGENT MEDICAL DEVICE RECALL letter dated May 10, 2012, to all affected customers. The firm expanded their recall on May 25, 2012 by sending a second letter with additional lots and products. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory and quarantine any affected product. Customers should contact Vycon's Customer Service Department at 1-800-473-5414 for a Return Merchandise Authorization number. Customers were instructed to enter the number on the space provided on the Recall Acknowledgment and Inventory Return Form, sign and date form and fax to Vycon at 215-672-6740. Customers should return the affected product as well as the original signed Recall Acknowledgment and Inventory Form using the instructions from Customer Service involving product pick-up. All inventory will be credited. If product was further distributed the Urgent Medical Device Recall letter should be provided to their customers.For questions regarding this recall call 215-672-6740. |  |  |  | Nationwide Distribution including AZ, CA, DE, FL, IL, IN, IA, KS, KY, MD, MA, MI, MN, MO, NJ, NY, OH, PA, UT, VA, WA, WV, and WI. | Terminated | March 11, 2013 | 304
Z-1490-2012 | 61355 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Sigma Spectrum VOLUMETRIC INFUSION PUMP, with or without Master Drug Library, Part Number 35700 -- Common/Classification Name: Infusion Pump -- Class 2 medical device -- SIGMA LLC, Medina, NY 14103 |  |  | July 06, 2011 | 2011 | Sigma | Recall expansion; the infusion pumps have the potential to fail causing inaccurate flow conditions during use. These conditions range from back flow to free flow, which could result in over-infusion. Sigma's evaluation of subsequent complaints involving failed bearings indicated that the loss or degradation of bearing lubricant to be a more significant contributor to bearing failure than initiall | Nonconforming Material/Component | Sigma sent an URGENT Medical Device Recall Notifications letter dated July 6, 2011 via UPS mail the consignees informing them of the expansion of the Spectrum Infusion Pumps recall. Some consignees were notified by telephone. The letter identified the affected product, problem, actions to be taken, clinical precautions and recommendations. Additional information regarding this voluntary recall is available on Sigma's web site www.sigmapumps.com/Recall. Customers were provided the following instructions: (1) Confirm that you have the affected serial number(s) ranging from 700000 to 794213 in inventory using the Customer Response Form. Refer to attached list of known affected devices in your possession. Revise this list as required. (2) Immediately remove affected pumps from neonatal use until B. Braun BC 1000 back check valves are employed. (3) Respond to this notification within five (5) working days of receipt by faxing the enclosed Customer Response Form to the number noted on the form, calling (866) 482-2893 between the hours of 8am and 8pm (EDT) or e-mailing a scanned version of form to sigma2407@stericycle.com. (4) Within three weeks of receipt of the Customer Response Form, (which verifies the device serial numbers in your facility), SIGMA will issue a customized recall report for each facility identifying the specific affected serial numbers, SIGMA internal contacts, and the plan for coordination of pump return and repair. For any adverse reactions experienced with the use of this product should be reported to SIGMA at (800)-356-3454 ext. 300 and to FDA MedWatch either by: Phone at 1-800-FDA-1088, by fax at 1-800-FDA-0178; Mail at MedWatch, FDA 5600 Fishers Lane, Rockville, MD 20857, or MedWatch website by linking to: https://www.accessdata.fda.gov/scripts/medwatch/medwatch-online.htm. |  | 63225 | 63,225 pumps total (corrected number): Domestic - 63,032 pumps; Canada - 29 pumps; Okinawa - 164 pumps | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of Canada and Okinawa | Open, Classified |  |  | 63225
Z-1538-2012 | 61830 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Practix Convenio Mobile X-ray system |  |  | October 19, 2011 | 2011 | Philips Healthcare Inc. | When the system is switched off or when the tube arm is parked in the down position (X-ray generator switch off), a radiation pulse (50 kV, 2.5 pGy) is generated inadvertently. .When the system is switched off or when the tube arm is parked in the down position (X-ray generatorswitch off), a radiation pulse (50 kV, 2.5 pGy) is generated inadvertently. The failure happened whencomponent supplie | Radiation Control for Health and Safety Act | Philips Healthcare sent a proprosed corrective action letter dated May 3, 2012, to FDA. The proposed corrective action plan is to bring the product into compliance with regulatory requirements by issuing a field corrective action which should be completed by Philips by August 1, 2012.For questions regarding this recall call 978-687-1501. |  | 11 | 11 - US | Nationwide Distribution | Open, Classified |  |  | 11
Z-1810-2012 | 61848 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | T4-Hytrel Zipper Toga |  |  | June 01, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | The clear tape, that is intended to aid in the prevention of patients fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga, was not applied during the manufacturing process. Stryker has received 5 complaints associated with the defective seams which allowed bleed through of bodily fluids. | Process control | Stryker sent the May 29,2012 URGENT MEDICAL DEVICE RECALL NOTIFICATION and the June 14, 2012 REVISED URGENT MEDICAL DEVICE RECALL Notice to all consignees. The letters identified the affected products, problem and actions to be taken. The Revised letter instructed customers to find, quarantine, make product unusable and destroy all implicated product . Replacement products may not immediately available. Customers were asked to complete and sign the Business Reply Form (BFR) and fax to Stryker Instruments Regulatory Department at 866-521-2762. If the affected products have been further distributed, forward this notification letter and the attached BRF to all affected locations. Questions should be addressed to Angela Ragainis, 269-389-4354; angela.ragainis@stryker.com |  | 16 | 16, 846 (last 3 years) 35,492 all years | Worldwide Distribution - US (nationwide) including the states of: AK, AL, AZ, CA, CO, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV and the country of SWITZERLAND | Terminated | April 14, 2014 | 682 | 16
Z-1812-2012 | 61737 | No | MKQ |  |  | Pathology | PMA | PrepStain GYN 480 Disposable Test Kit with PreCoat Slides (Catalog #490523 |  |  | March 22, 2012 | 2012 | Tripath Imaging, Inc. | The PrepStain Syringing Pipettes may not dispense the minimum volume of 7.5 mL of the sample due to leaking. This event resulted in an increase in complaints of leaking syringing pipettes | Component design/selection | BD Diagnostics sent a "BD PrepMate Syringing Pipette" letter dated March 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The firm encouraged customers to refer to the Operator's Manual and inspect the centrifuge tubes to verify that the correct amount of sample has been transferred. Customers experiencing this issue were instructed to discard the pipettes and contact BD Diagnostic Technical Support for a replacement. Contact the firm at 1-877-822-7771 for questions regarding this notice. |  | 46112784 | 46,112,784 for all products in Recall Event | Worldwide Distribution. | Terminated | December 03, 2012 | 256 | 46112784
Z-1818-2012 | 61926 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Software, software release 2.0.2 and 2.0.2SP1 including mirror icon |  |  | May 22, 2012 | 2012 | Philips Healthcare Inc. | With Digital Diagnost (DIDI) R 2.0.2 a mirrored "R in a circle" is burnt into the image (if mirrored) in the lower right corner of the image. When such an image is mirrored back inside the PACS system, an unmirrored "R in a circle" appears in the lower left corner. This can be mistaken for a "right patient side" marker, although this can appear on the left patient side | Software change control | Philips Healthcare issued a Field Safety Notice on 5/22/12 advising users of the error. A software update (Rel 2.1.2) will issue to correct the problem. Contact your local Philips representative. 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 752 | 752 units | Worldwide Distribution - USA (nationwide) and the countries of:NationwideForeign: CanadaAFGHANISTANALBANIAALGERIAANGOLAARGENTINAARMENIAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHAMASBAHRAINBANGLADESHBARBADOSBELARUSBELGIUMBENINBHUTANBOLIVIABOSNIA&HERZEGOVINA;BRAZILBRUNEI DARUSSALAMBULGARIABURUNDICAMBODIACAMEROONCANARY ISLANDSCHADCHILECHINACOLOMBIACONGOCONGO,DEMOCRATICCOSTA RICACOTE D'IVOIRECROATIACUBACYPRUSCZECH REPUBLICDENMARKDJIBOUTIDOMINICAN REPUBLICECUADOREGYPTEL SALVADOREQUATORIAL GUINEAERITREAESTONIAETHIOPIAfake test ssdFINLANDFRANCEFRENCH GUIANAGABONGAMBIAGEORGIEGERMANYGHANAGREECEGUADELOUPEGUAMGUATEMALAGUINEAGUYANAHONDURASHONG KONGHUNGARYICELANDINDIAINDONESIAIRANIRAQIRELANDISRAELITALYJAMAICAJAPANJORDANKAZAKSTANKENYAKUWAITKYRGYZSTANLAO PEOPLE'S DEM.REPLATVIALEBANONLIBERIALIBYAN ARAB JAMAHIRILIECHTENSTEINLITHUANIALUXEMBOURGMACAUMACEDONIAMALAWIMALAYSIAMALIMALTAMARTINIQUEMAURITANIAMAURITIUSMETRACOMMEXICOMOLDOVAMONGOLIAMOROCCOMOZAMBIQUENAMIBIANEPALNETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANIGERIANORTH KOREANORWAYOMANPAKISTANPALESTINEPANAMAPARAGUAYPERUPHC BESTPHILIPPINESPMCC BU CanadaPMCC BU USAPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIARWANDASAUDI ARABIASENEGALSERBIA&MONTE; NEGROSIERRA LEONESINGAPORESLOVAKIASLOVENIASOMALIASOUTH AFRICASOUTH KOREASPAINSRI LANKASUDANSURINAMESWEDENSWITZERLANDSYRIAN ARAB REPUBLICTAIWANTAJIKISTANTANZANIATHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYTURKMENISTANUGANDAUKRAINEUnidentified CountryUNITED ARAB EMIRATESUNITED KINGDOM URUGUAYUZBEKISTANVENEZUELAVIET NAMYEMENZAMBIAZIMBABWE | Open, Classified |  |  | 752
Z-1816-2012 | 62130 | No | GAC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Howmedica Osteonics Accolade Offset Rasp Handle |  |  | December 29, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware of the potential for the above noted Accolade Rasp Handle Assembly to fracture upon use. | Device Design | Stryker Orthopaedics sent an email notification to all consignees on December 29, 2011, which was followed by an Urgent Product Recall letter dated January 4, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the attached Product Recall Acknowledgement Form and fax back to 201-831-8089 within 5 days. Customers were also asked to contact the hospitals in their territory that have the affected product to arrange return of the product. Customers were instructed to return the affected product to the attention of Regulatory Compliance, Stryker Orthopedics, 325 Corporate Drive, Mahwah, New Jersey 07430 with the fluorescent orange PRODUCT REMEDIATION sticker attached indicating the Product Remediation # RA 2011-161. Mark the outer box with the words "Product Recall."For questions customers were instructed to call 201-972-2100.For questions regarding this recall call 201-831-5970. |  | 41 | 41 units | Nationwide Distribution | Terminated | June 13, 2013 | 532 | 41
Z-1787-2012 | 61639 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aestiva/5 7900 SmartVent, anesthesia machine. Used in Datex-Ohmeda Aestiva/5 Anesthesia Systems. |  |  | April 10, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the interlock slide mechanism missing from the 3-vaporizer manifold of your Aestiva 7900 Anesthesia Device. This Potential for two (2) vaporizers delivering agent at the same time. This could manifest as over delivery of a single agent if both vaporizers contain the same agent or as delivery of more than one agent. This m | Process control | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 10, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact GE Healthcare Customer or Technical Support at 800-345-2700 for questions regarding this notification. |  | 8 | 8 | Nationwide Distribution - including the states of AL, FL, GA, MO, and VA. | Terminated | July 09, 2012 | 90 | 8
Z-1822-2012 | 62120 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Peak SE Primer Kits (4 syringes per kit) manufactured by Ultradent Products, Inc, South Jordan, UT, Self Etch Primer. The product is contained within prefilled, labeled syringes in a transparent plastic container identified with an aqua, red, orange, blue and black label with black and blue print writing. Product code: 5135, 872.3200 KLE, liquid and powder mixture is a syringe-in-a-syringe delivery system. |  |  | May 24, 2012 | 2012 | Ultradent Products, Inc. | Ultradent is recalling lot B6Z4K of Peak SE Primer as a result of a field complaint, dated May 10, 2012, stating that a syringe of Peak SE Primer was mislabeled as Peak LC Bond Resin. These products are used by Dental professionals for bonding needs in restorative dental procedures. | Labeling mix-ups | Ultradent Products Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 5, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discard any affected product from their 5135 kit. Customers were asked to complete and return a Recall Return Response form to the firm by fax at 801-553-4609. |  | 172 | 172 kits (4 syringes per kit) | Worldwide Distribution. | Terminated | August 15, 2012 | 83 | 172
Z-1821-2012 | 61877 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Philips Easy Diagnost |  |  | May 11, 2012 | 2012 | Philips Healthcare Inc. | If a lead ruler is not positioned in the same plane as the anatomical image detail to be measured, measured object distances (using the ruler) can be wrong by a corresponding factor. If, additionally, the measurement tool is not used to check the ruler, or is used in the right way (ie: calibration), such calibration error is not detected. As a result of this, a surgical intervention may be execut | Device Design | Philips sent an Urgent - Field Safety Notice letter dated May 11, 2012 to their affected customers. The letter identified the affected product, what the issues are, under what circumstances they may occur and the actions to be taken to avoid or minimize the occurrence of the issue Philips plans to supply an Addendum to the Instruction for Use. A Philips Service Engineer will contact the affected customers as soon as the Field Action Kit is ready to be implemented. For questions contact your local Philips representative: Technical Support Line 1-800-722-9377. |  | 6 | 6 | 6 US customers | Open, Classified |  |  | 6
Z-1820-2012 | 61868 | No | IRO | Therapeutic vibrator. | Physical Medicine | Physical Medicine | 510(K) Exempt | USA-002 Rev A brochure is a two sided tri-fold printed marketing brochure describing the approved labeling indications for use, precautions and general operational information of the Exer-Rest device, Models AT3800 and AT4700. |  |  | May 07, 2012 | 2012 | Non-Invasive Monitoring Systems, Inc. | Non-Invasive Monitoring Systems Inc (NIMS) is voluntarily recalling promotional materials which contain unsubstantiated direct and indirect claims. | Labeling False and Misleading | NiMS sent an Urgent Promotional Material Recall letter dated May 29, 2012, to all affected customers. NIMS contacted the 18 facilities by phone starting on May 29, 2012 and sent the 73 end users letters dated May 29, 2012. NiMS instructed customers to immediately examine their inventory and quarantine any of the US-001 or USA-002 brochures. If customers further distributed brochures customers were instructed to notify their customers at once by including a copy of the recall notification letter. Customers were requested to destroy all Exer-Rest promotional materials. For any questions customers were instructed to call 305-575-4204.For questions regarding this recall call 305-575-4201. |  | 91 | USA-002 Rev A brochure: 91 | Nationwide Distribution | Terminated | July 16, 2012 | 70 | 3756
Z-1778-2012 | 61843 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | NIDEK Medical NUVO LITE MODEL 525 OCSI, 115V -60Hz - 330W. |  |  | May 01, 2012 | 2012 | Nidek Medical Products Inc | Capacitor failure may result in a fire hazard and loss of supplemental oxygen. | Component design/selection | Nidek Medical Products Inc sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 1, 2012 to affected customers. The letter identifies the product, problem, and the actions to be taken by the customers.The letter included a response form for customers to complete and return to the firm. Contact the the firm at 205-856-7200 ext. 242. |  | 730 | 730 units | Worldwide Distribution-USA (nationwide) and the countries of ALBANIA, ALGERIA, AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BULGARIA, CAMBODIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ENGLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONG KONG, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, IVORY COAST, JAPAN, KAZAKHSTAN, KINGDOM OF SAUDI ARABIA, KUWAIT, LEBANON, LIBYA, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, MYANMAR, NAMIBIA, NEPAL, NEW ZEALAND, NIGERIA, PAKISTAN, PERU, PHILIPPPINES, PORTUGAL, REPUBLIC OF CHINA, REPUBLIC OF SOUTH AFRICA, RUSSIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SOUTH AUSTRALIA, SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, THE NETHERLANDS, TOGO, TUNISIA, TURKEY, UGANDA, UNITED ARAB EMIRATES, and VIETNAM. | Terminated | March 21, 2017 | 1785 | 46972
Z-1824-2012 | 57878 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | HEPARIN LOCK FLUSH SOLUTION, USP, 10 USP Units/mL, 1 mL Multiple Dose Vial, plastic vial, vial size 3 mL, NDC 63323-544-01, Product Code 504401, Rx only |  |  | May 20, 2010 | 2010 | App Pharmaceuticals Llc | CGMP Deviations: Incomplete documentation associated with test results. | Release of Material/Component prior to receiving test results | APP sent an URGENT DRUG RECALL LETTER and Response Form dated May 20, 2010, to all affected consignees. The consignees contacted were at the retail/ medical facilities/hospital IeveIs, including intermediate wholesale IeveIs. Direct distributor consignees were instructed to notify their customers immediately of this recall action, and direct them to discontinue distributing or dispensing the affected lots, plus return the affected products and lots to APP Pharmaceuticals, LLC in Bensenville, IL. Direct consignees with affected lots on hand were instructed to immediately discontinue distributing or dispensing the affected products and return them to APP Pharmaceuticals, LLC in Bensenville, IL 60106. Questions regarding the information on how to return product can be made to the firm's Quality Assurance Department at 1-866-716-2459.For questions regarding this recall call 847-969-2700. |  | 133150 | 133,150 vials | Nationwide Distribution and Puerto Rico | Terminated | June 20, 2012 | 762 | 637651
Z-1823-2012 | 62125 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow International, Inc. 7FR 3-Lumen Arrowg+ard Blue Plus(R) Pressure Injectable Central Venous Catheter - Catalog number CDC-42703-XP1A |  |  | June 07, 2012 | 2012 | Arrow International Inc | Arrow is recalling the product due to the possibility of voids in the catheter hub which could lead to an interlumen crossover. | Nonconforming Material/Component | Arrow International sent a Urgent Medical Device Recall Notification letter dated June 11, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to take the following action:1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog and lot numbers listed.2. To return product complete the enclosed Recall Acknowledgement Form and fax to 1-800-343-2935, Attn: Customer Service. This will allow us to document the amount of the product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization ( RGA) Number and will provide instructions for the return of product to Arrow International.3. If you have no affected stock, please complete t he enclosed Recall Acknowledgment Form and fax it to 1-800-343-2935, Attn: Customer Service. This will allow us to document your receipt of this letter.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-800-343-2935. |  | 1085 | 1,085 | USA (nationwide) AL, AZ, AR, CA, CO, FL, GA, IL, IN, KS, LA, MA, MI, MO, NV, NH, NC, OH, PA, TX, VA, WA, and WI. | Terminated | April 03, 2014 | 665 | 1085
Z-1828-2012 | 61376 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Implant, Dental, Endosseous, Acid Etched, ICE, Osseotite, Internal Connection. . |  |  | November 04, 2011 | 2011 | Biomet 3i, LLC | On November 3, 2011 Biomet 3i, Palm Beach Gardens, FL initiated a recall of their Osseotite Certain Implant, Model # IOSM313, Lot #2010111529. Dental implant lot may not have an internal thread. | Device Design | Biomet 3i sent an Urgent: Medical Device Recall letter dated November 04, 2011 via e-mail, facsimile, telephone or postal mail to all consignees affected. The letter identified the affected product, problem and actions to be taken. Each customer was instructed to check their respective inventory for the affected product and return any unused item(s) to Biomet 3i for prompt replacement. For question or concerns contact Biomet 3i at 1-800-342-5454. |  | 219 | 219 units in total | Worldwide Distribution - US(nationwide) and the countries of: Brazil, Canada, Columbia, Italy, Lebanon, Spain, Switzerland, and UK. | Terminated | June 25, 2012 | 234 | 219
Z-1835-2012 | 61742 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | SARNS FLEXIBLE ARTERIAL CANNULA: 6.7 MM (20 FR) 00 WITH 3/8" CONNECTOR,SUTURE RING, 9.5" (24 CM) LONG. |  |  | May 09, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During production of one lot of the Sarns High-Flow Aortic Arch Cannula, Terumo Cardiovascular Systems (Terumo CVS) found a foreign substance on the outer surface of some cannulae connectors.Terumo CVS' preliminary investigation found that the substance can be dislodged from the cannula surface, the substance was likely deposited during the molding process, but the exact composition of the parti | Nonconforming Material/Component | Terumo sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 9, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter advises customers to discontinue use of the affected product and to return the product to the firm. A Response Form was attached for customers to complete and return. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this recall. |  | 133793 | 133,793 (each) | Worldwide Distribution-USA (nationwide) including the states of AL , AR , AZ, CA , CO , CT , FL , GA , HI , IA , IL , IN , KS , KY , LA , MA , MD , ME , MI , MI , MN , MO , MT , NC , ND , NE , NH , NJ , NM , NV , NY , OH , OK , OR , PA , RI , SC , SD , TN , TX , UT , VA , WA , WI, and WY and the countries of Australia, Belgium, Brazil, Canada, Chile, Germany, Honduras, Hong Kong, India, Japan, Mauritius, Nicaragua, Philippines, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand, Trinidad and Tobago, United Arab Emirates (UAE), and Vietnam. | Terminated | July 11, 2013 | 428 | 133793
Z-1831-2012 | 61913 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | The product is Pathfinder SL, version 1.6.0.3057 software used with the Spacelabs Healthcare Pathfinder SL Holter Analyzer, version 1.6.0. |  |  | May 11, 2012 | 2012 | Spacelabs Healthcare, Llc | The firm received reports of the Pathfinder LS software crashing after editing a batch of Holter Analyzer records. The software will stop with an error message requiring a restart. In addition, QTc values will not update after moving the marker. | Software design | Spacelabs Healthcare, Llc sent a "MEDICAL DEVICE CORRECTION" letter dated May 11, 2012 via mailed, return receipt requested, to US customers. The customer letter (translated as necessary) was emailed to all international subsidiaries and distributors of record. The letter identified the affected product, problem and actions to be taken. Consignees are advised to turn the power off and then on to restore normal functionality. Updated software will be provided upon completion of validation. For questions contact Spacelabs Healthcare at 1-800-522-7025 and select 3 for Technical Support. |  | 114 | 114 | Worldwide Distribution - US (nationwide) including the states of Nevada and Vermont, and the countries of AUSTRALIA, AUSTRIA, CANADA, CYPRUS, GERMANY, ICELAND, ISRAEL, POLAND, ROMANIA, SOUTH AFRICA, SWITZERLAND, THAILAND, TURKEY and UNITED KINGDOM. | Terminated | October 22, 2012 | 164 | 114
Z-1829-2012 | 62052 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Multi Diagnost Eleva II with swivel cable |  |  | May 21, 2012 | 2012 | Philips Healthcare Inc. | MultiDiagnost Eleva X--Ray system may have damage to System Cables from repeated movement | Device Design | Philips Healthcare Imaging issued Urgent Field Safety Notice dated May 21, 2012, to accounts informing an inspection will be performed of the unit. The Cable guiding plate will be replaced. If the cables show any damage the complete cable set wll also be replaced. No action is be taken by the customer. For further information customers should call 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 433 | 433 units | Worldwide Distribution - USA (nationwide)and the countries of:NationwideForeign: AustraliaAustriaBelgiumBharainCanadaChinaColumbiaCzech RepublicDenmarkDutch AntillesEgyptEstoniaFinlandFranceGermanyHong KongIrelandIsraelItalyLatviaLebanonLuxembourgMalaysiaMexicoNetherlandsNew ZealandNorwayPhilippinesQatarRussian Fed.Saudi ArabiaSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandSyriaTaiwanThailandTurkeyUnited Kingdom | Open, Classified |  |  | 433
Z-1843-2012 | 61946 | No | MXM |  |  | Neurology | PMA | Olympic Cool-Cap System is intended to provide treatment for |  |  | May 21, 2012 | 2012 | Natus Medical Incorporated | Olympic Cool-Cap Cooling Module's power supply failed during treatment. The power supply failure can happen without prior indication and results in the unavailability of the system to begin or continue cooling treatment. | Component design/selection | Natus Medical Incorporated sent a URGENT: MEDICAL DEVICE FIELD CORRECTIVE ACTIONS Olympic Cool-Cap System letter dated May 21, 2012, to all affected customers. Fed-Ex to all U.S. customers, specifically to Biomedical department, NICU, with affected devices. The letter identified the product the problem and the action needed to be taken by the customer.In addition, Natus also sent the Urgent Field Safety Notice-Olympic Cool-Cap System (DOC-007459) dated May 2012 to all US customers and International Distribution Partners who were affected by both Frozen Screen and Power Supply field corrective action.Customers were instructed to: " During all routine checks of infants undergoing treatment, look at the clockdisplayed on the screen to verify that the seconds counter on the clock isadvancing. The time displayed should continue to advance if the system isoperating correctly." If the system experiences a screen-freeze during a cooling session, the Olympic Cool-Cap system can be rebooted and cooling resumed." Ensure that all the personnel in your facility that use the Olympic Cool-Capsystem are advised of this Field Safety Notice (see attached).The U.S. customers are advised of the issue specifically stated in Action #2 in the letter and they are also informed that Natus will contact them within the next 30 to 120 days to schedule the power supply replacement in their affected system(s).Natus requires that the Distribution Partners to notify their customers with copies of the Natus recall letter and the Field Safety Notice (DOC-007459). Specifically these two documents should be addressed to both the Biomedical Department and the Neonatal Intensive Care Unit (NICU).Customers with questions about this recall should call Natus Technical Service at 888-496-2887 or email to TService@natus.com. |  | 181 | 181 units | Worldwide Distribution -- USA (nationwide) including the states of : AK; AR; AZ; CA; CO; DE; FL; GA; HI; ID; IL; IN; KY; MA; MI; MN; MS; ND; NE; NJ; NY; OH; PA; SC; SD; TN; TX; and WA. and the countries of : Argentina; Australia; Austria; Canada; Egypt; Germany; Hong Kong; Indonesia; Iran; Italy; Kuwait; Poland; Russia; Saudi Arabia; Singapore; Slovenia; South Africa; South Korea; Spain; Taiwan; Thailand; Turkey; United Arab Emirates; and United Kingdom. | Terminated | February 21, 2014 | 641 | 181
Z-1844-2012 | 61942 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien ROTICULATOR 55-3.5 Single Use Stapler |  |  | May 09, 2012 | 2012 | Covidien LP | Potential for the sterility barrier to be compromised. | Package design/selection | The firm, Covidien, sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 16, 2012 via federal express to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify and quarantine all affected inventory (note: units from affected lots may have been incorporated into Covidien BEST PRACTICE KITS; evaluate kits for affected lots). All affected ROTICULATOR units must be returned. The customers were also instructed to complete and return the attached ROTICULATOR Single Use Stapler Recalled Product Return form via fax to: (800) 895-6140 or email to: SDFeedback@Covidien.com, even if you do not have any of the affected units, and include a copy of completed form with products being returned. Customers who received product directly from Covidien should contact Covidien Customer Service at SDFeedback@Covidien.com or (800) 722-8772, option 1, prior to returning affected products to Covidien, Field Returns Department, at 195 McDermott Road, North Haven, CT 06473; if received from distributor, return form and contact distributor directly.If you have any questions or concerns, please do not hesitate to contact Covidien Representative or Covidien Quality Assurance at (203) 492-5232. |  | 50698 | 50698 units | Worldwide distribution: USA (Nationwide) and countries including; Africa, Argentina, Australia, Belgium, Brazil, Canada, China Costa Rica, CROATIA, Czech Republic,Dominion Republic, ESPAÂżA France, Germany, Greece Guadeloupe, Guatemala, Honduras Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Japan, Korea ,Malaysia, Mexico, Peru, Portugal, Puerto Rico, Spain, Sri Lanka, Switzerland, Taiwan , Thailand, and UK. | Terminated | February 12, 2014 | 644 | 123856
Z-1849-2012 | 61741 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Medtronic 0-arm Imaging System with Software Version 3.1.1, Catalog Numbers: BI-700-00027-120R and BI-700-00027-120. The 0-armÂż Imaging System is a mobile x-ray system designed for 2D fluoroscopic and 3D imaging. |  |  | October 12, 2010 | 2010 | Medtronic Navigation, Inc. | Software Version 3.1.1 does not meet navigational accuracy. | Software design | Medtronic Navigation contacted customers by telephone on 10/12/2010 and with "Urgent Field Safety Notice" follow up letters dated 10/14 or 10/15/2010. Sites were visited by Medtronic Service Representative to complete navigational accuracy testing and either replace or correct the unit. |  | 3 | 3 units | Nationwide Distribution including the states of CO, IN, and MN. | Terminated | June 21, 2012 | 618 | 3
Z-1850-2012 | 61935 | No | FDI | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Profile Single-Use Pediatric Snare 11mm |  |  | May 16, 2012 | 2012 | Boston Scientific Corporation | Difficulty in extending snare loop from the catheter | Process control | Boston Scientific issued an Urgent Medical Device Recall Action letter via Federal Express Priority mail on May 16, 2012. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form. Scan and fax to 508-683-5578.For questions regarding this recall call 508-683-8000. |  | 2760 | 2760 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Australia, Austria, Belgium, France, GB, Germany, Italy, Netherlands, Qatar, Spain, Sweden, and Switzerland | Terminated | February 28, 2013 | 288 | 6070
Z-1853-2012 | 61761 | No | JOZ | Automated platelet aggregation system. | Hematology | Hematology | 510(k) | VerifyNow P2Y12 Assay, Part Number: 85064, All reagent lots. |  |  | February 06, 2012 | 2012 | Accumetrics Inc | The recall was initiated by Accumetrics because of a field correction to the VerifyNow System; specifically the results reported when running the VerifyNow P2Y12 Test. Currently the VerifyNow P2Y12 Test reports three values: PRU, BASE and % Inhibition. The change implemented is the elimination of the BASE and % Inhibition results. There will be no change to the reporting of the PRU result whic | No Marketing Application | Accumetrics, Inc. sent an Urgent Field Correction notification letter and an Ackknowledgement - fax back form to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter informed customers that an Accumetrics representative will becontacting them to schedule an appointment to update their VerifyNowinstrument software. The letter stated that Accumetrics realizes that procedures and report format may need to be revised as software changes is implemented, and will provide assistance to help in making this transition. For any questions or concerns, call at 1-800-643-1640 ext. 2; oremail at: support@accumetrics.com. |  | 26075 | 26,075 units | US Nationwide Distribution | Terminated | November 27, 2012 | 295 | 26075
Z-1856-2012 | 62127 | No | MCN |  |  | Obstetrics/Gynecology | PMA | seprafilmÂż Single Site ADHESION BARRIER |  |  | June 13, 2012 | 2012 | Genzyme Corporation | Sterility of product may be compromised due to packaging defect | Packaging process control | Genzyme issued an Urgent: Medicatical Device Product Recall Notification Letter dated June 13, 2012 to all affected customers. Stericycle Expert Solutions, will be acting on behalf of Genzyme to coordinate notifications, communications, and logistics for returns of all product subject to this recall. The letter identified the affected products, problen and actions to be taken. Customers are requested to check their inventory and return all affected products per the instructions provided. Customers are instructed to complete the enclosed Business Reply Form and return via fax transmission to 866-912-2519. For questions call Stericycle at: 877- 473-7069. |  | 1285 | 1285 units | Worldwide Distribution - US (nationwide) and the country of S. Korea | Terminated | July 24, 2013 | 406 | 5165
Z-1861-2012 | 61712 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Rose Bed Model (FL14E) |  |  | May 17, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | An adverse trend of scale and bed exit service reports/complaints were detected for the GoBed+, GoBed II, Rose, and MA204 model beds. Investigation found an abnormally high failure rate of loadcells, an integral component in scale and bed exit function. The scale faults are manifested as a fluctuating scale reading, a scale error causing the scale system to be disabled until serviced, or an inac | Component design/selection | The firm, Stryker Medical, sent an "URGENT- Medical Device Correction" letter dated May 17, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to observe the beds for any malfunctions according to Preventive Maintenance Checklists in maintenance manual; report scale or bed exit malfunctions to Stryker (1-800-327-0770, option 2) to arrange for field service or to order service parts, at no charge. The customers were also instructed to forward this notice and advise Stryker of any subaccounts if the product was further distributed. A postage paid return postcard is to be filled for tracking and to make arrangements for repair. If you have any urgent questions or concerns, call 269-389-6466, Monday-Friday 8a.m-5p.m. (ET). |  | 999 | 999 | Worldwide distribution:USA (nationwide) and countries including: Argentina, Bertec, Brasil, Brazil, Canada, Chile, China, India, Latin America, Mexico,Osteonics SA, Polska, Portugal, Puerto Rico, and Singapore. | Terminated | May 12, 2014 | 725 | 27319
Z-1858-2012 | 61899 | No | NAB | Nonresorbable gauze/sponge for external use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hartmann Sterilux Ref 56910000 Sterile Premium gauze sponge 4"x4" 12-ply 24 boxes/case 1200 total sponges |  |  | May 14, 2012 | 2012 | Hartmann USA, Inc | Micro holes were detected in the laminate of the sterile barrier system of sterilized Hartmann Econolux and Hartmann Sterilux gauze products | Material/Component Contamination | The firm, HARTMANN USA, sent an "URGENT: RECALL NOTICE" dated May 14, 2012 to its Consignees/Customers. The notice described the product, problem and actions to be taken. The customers were instructed to examine their inventory, block further shipments of affected items and retrieve additional affected stock from their customers; to confirm initial receipt of the notice by emailing recall@harmanninfo.com; complete and return the Affected Products and Stock Levels form via fax to: HARTMANN USA customer service at (1-803-325-7606); contact customer service to obtain a Returned Goods Authorization; and take adequate measures to ensure this notice is distributed appropriately.If you have any further questions or issues related to this recall, please contact HARTMANN USA customer service or your sales representative at 1-800-243-2294 or email: recall@hartmanninfo.com. |  | 604711272 | Lot 102147399 - 421 Cases; Lot 102250396 - 360 Cases; 200105390 - 522 Cases; 200206394 - 390 Cases | Nationwide distribution: USA including states of: AZ, CA, FL, GA, KY, MO, MS, NC, PA, SC, TN, TX, and VA. | Terminated | October 06, 2014 | 875 | 1313561759
Z-1865-2012 | 62016 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC Uroview 2800, Model number A349855. |  |  | November 20, 2006 | 2006 | GE OEC Medical Systems, Inc | Recalled in November 2006 due to several intermittent potential safety issues including navigation tracking inaccuracy, system Lok-up, incorrect image display, incorrect dose area calculated air kerma in "film mode" and "digital spot mode", and the system does not reboot. | Software design | On 11/20/06 two written notifications were sent to consignees addressed to the Facility Administrator and the Radiology Administrator. The notifications advised them that GE Healthcare had identified several intermittent potential safety issues that may occur with the product. Interim solutions were provided for each issue. The letter stated that GEHC OEC would contact customers and remedy the issues free of charge when a permanent solution became available. An Urgent Recall Notice, dated April 13, 2007, was also sent to customers. Customers were to contact the service team at 800-874-7378 for further information, if there were any questions or concerns. |  | 717 | 717 units | Worldwide Distribution. | Terminated | June 25, 2012 | 2044 | 717
Z-1864-2012 | 62331 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Advanta 2 Bed. |  |  | June 18, 2012 | 2012 | Hill-Rom, Inc. | During an evaluation of Advanta 2 siderail samples returned from the field it was found that the screws used to attach the siderail were being stripped from the plastic cavity of the siderail. | Process design | Hill-Rom sent an "URGENT FIELD SAFETY NOTICE" dated May 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included a Customer Response Form for customers to complete and return to the firm. Contact Hill-Rom Technical Support at 800-445-3720 for questions regarding this notice. |  | 7843 | 7843 devices | Worldwide Distribution-USA (nationwide) and the countries of Canada, France, and Singapore. | Terminated | June 06, 2014 | 718 | 7843
Z-1866-2012 | 61955 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Formula Shaver Handpiece (with buttons); |  |  | May 21, 2012 | 2012 | Stryker Endoscopy | Stryker Endoscopy determined that there may not have been adequate determination of required routine checks to ensure accuracy and fitness for use of specific serial numbers of Formula and 180 Shaver Hand Control. | Nonconforming Material/Component | The firm, Stryker Endoscopy, sent an "Urgent: Device Removal" letter dated May 21, 2012 to all affected customers via Federal Express. Stryker Sales Reps were notified by e-mail on the same day. Stryker international sites were notified by e-mail on May 24, 2012. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their inventory for part and serial number; quarantine any of the affected product, call Stryker Endoscopy's Customer Service at 1-800-624-4422 for replacements and return affected product.If you have any questions regarding this letter please contact us at shavers@stryker.com or call 1-800-624-4422. |  | 34 | 34 devices | Worldwide distribution: USA (nationwide) including states of: CA, FL, IA , MI, NV, NC, OK, TX, and WI; and countries of: China and Great Britain. | Terminated | September 20, 2012 | 122 | 34
Z-1826-2012 | 62108 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius NaturaLyte Liquid Acid Concentrate |  |  | March 29, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Risk of Alkalosis with acetate containing dialysis acidconcentrates | Labeling False and Misleading | Fresenius Medical Care sent an "IMPORTANT PRESCRIBING INFORMATION" letter dated March 29, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Reply Form was enclosed for customers to complete and return via fax to 781-699-9635. Contact Customer Service at 1-800-323-5188 or Medical Information at 1-855-616-2309 for questions regarding this recall. |  |  |  | Nationwide Distribution and the country of Guam. | Terminated | July 20, 2017 | 1939
Z-1867-2012 | 61768 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | BCS(R) Behring Coagulation System, Device catalog No. OVIO03 |  |  | December 01, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed that if the Cleaner SCS vial becomes empty after processing a test for Antithrombin assay using the INNOVANCE(R) Antithrombin kit, and the test immediately following is an APTT based clotting assay several results may be shortened. | Device Design | An Urgent Field Safety Notice dated December 2010 was issued to all BCS(R) System owners. Firm conducted a voluntary field corrective action to notify BCS(R) System owners and to install a software "patch" to resolve the issue. Until the software patch was installed on the BCS(R) System, customers were instructed in the Safety Notice that "the potential issue of carryover can be avoided if the Cleaner SCS vial is kept filled and not allowrd to run empty." |  | 186 | 186 | Worldwide Distribution - US (nationwide) including the states of: AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IO, KS, KY, LA, MD, MA, MI, MS, MN, MT, NE, NH, NJ, NY, NC, ND, OH, OK, OR, PA, TN, TX, UT, VA, WA, WV, and WI and the countries of: Australia, France, Canary Islands, Spain, Belgium, Germany, Switzerland, Netherlands, Italy, Guadeloupe, Ceuta, Austria, Slovenia, Croatia, Saudi Arabia, Turkey, Greece, and San Marino. | Terminated | March 05, 2014 | 1190 | 186
Z-1830-2012 | 61915 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow Two-Lumen Central Venous Catheterization Set with Blue FlexTip(R) ARROWgard Blue(R) Catheter |  |  | May 03, 2012 | 2012 | Arrow International Inc | Lidstock of product does not contain chlorhexidine contraindication and contains wording "contains no medication" on the label. Product insert does reflect medicated status. | Labeling mix-ups | Arrow International sent an Urgent Field Corrective Action letter dated May 3, 2012 to their US consignee. The letter identified the affected product, problem and actions to be taken. The letter instructed the consignee to quarantine any affected inventory and relay the information to their Arrow Sales Representative. The letter stated that the Arrow Sales Representative would affix new labels to each kit within their inventory. The consignee was instructed to sign the Acknowledgement Form provided by the Sales Representative which would then be submitted to Arrow. An Urgent Field Safety Notice letter dated May 15, 2012 was sent to foreign consignees with field safety corrective action instructions to follow. If you require additional information or clarification regarding this matter contact your sales representative. |  | 634912 | 634,912 | Worldwide Distribution - United States (nationwide) include the state of Texas. and the countries of: ARMENIA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, CANADA, CHILE, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, HUNGARY, INDIA, ITALY, IRELAND, KENYA, KUWAIT, KYRGYZSTAN, LATVIA, LEBANON, LITHUANIA, LUXEMBOURG, MONTENEGRO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, PORTUGAL, QATAR, RÂżUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, TURKEY, UKRAINE, , UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | March 22, 2013 | 323 | 634912
Z-1868-2012 | 62238 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | MRCTT80041 MedComp 8F Plastic Pro-Fuse(R) CT Low Profile Port with Attachable Chronoflex(R) Polyurethane Catheter. |  |  | April 02, 2012 | 2012 | Medical Components, Inc dba MedComp | Product was packaged with the incorrect port and does not match the label. The kit is labeled as an 8F Low Profile Pro-Fuse CT Port. The kit contains an 8F Midsize Dignity CT Port. | Packaging process control | MedComp sent a "PRODUCT ALERT" letter dated April 2, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The firm requested the return of all affected product. Contact the firm at 215-256-4201 ext. 2225 for questions regarding this notice. |  | 30 | 30 | Worldwide Distribution-USA (nationwide) including the states of GA and WA and the countries of Ireland and Turkey. | Terminated | May 20, 2013 | 413 | 30
Z-1896-2012 | 62333 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx monitor/defibrillators |  |  | June 20, 2012 | 2012 | Philips Healthcare Inc. | HeartStart MRx Unexpected Pads/Paddles ECG Failure | Process control | Philips issued an Urgent Voluntary Medical Device Correction/Field Safety Notice letter on June 20, 2012. The letter informs customers of the issue, identifies details (including models and software versions applicable) of units affected, gives instructions on actions to be taken until the correction is available, and identifies what action Philips plans to eliminate or remedy the issue. The letter states that a Philips Healthcare representative will contact customers to arrange for the replacement of the Power PCA in all the affected devices.Philips will complete the installation free of charge. For questions or support concerning this issue contact your local Philips representative at 1-800-722-9377. |  | 2914 | 2914 units | Worldwide Distribution - US (nationwide) and the countries of: AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHINA, COLOMBIA, EL SALVADOR, ETHIOPIA, FRANCE, GERMANY, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA, REPUBLIC OF, LEBANON, LIBYAN ARAB JAMAHIRIYA, MEXICO, MOROCCO, NETHERLANDS, NETHERLANDS ANTILLES, NICARAGUA, NORWAY, PERU, PHILIPPINES, POLAND, QATAR, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, THAILAND, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Terminated | March 16, 2017 | 1730 | 2914
Z-1869-2012 | 61889 | No | CDD | Vitamin B[bdi1][bdi2] test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension Vista(R) System B12 Flex(R) Reagent Cartridge (K6442) |  |  | May 08, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed that the presence of Intrinsic Factor Blocking Antibody in some patients can result in a false and significant elevation in the reported concentration of Vitamin B12 on the Dimension Vista System. This issue affects all lots of B12 Flex reagent cartridges on the Dimension Vista System. | Device Design | Siemens Healthcare Diagnostics Inc. sent a Urgent Device Recall letter to all affected customers dated May 2012. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed customers to immediately discontinue use of the device, quarantine and dispose of all inventory, ensure that your Medical Director reviews this letter as well as the attached physician letter. Please complete the attached Field Correction form and fax to 302-631-8467. Further questions please call 302-631-6311. |  | 59735 | 59,735 | Worldwide Distribution -- USA (nationwide) and the countries of Austria, Belgium, Czech Republic, Denmark, Finland, France, Germany, Great Britian, Italy, Netherlands, Norway, Portugal, Slovakia, South Africa, Spain, Switzerland, United Arab Emirates, Canada, New Zealand, and South Korea. *** Center Recommended Depth is Retail*** | Terminated | April 03, 2014 | 695 | 59735
Z-1898-2012 | 62272 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Adult Divided Cannula end tidal C)2 sampling w/male luer-lock connector. |  |  | May 30, 2012 | 2012 | SALTER LABS | A complaint was received, and later verified, that product labeled as SO-1296 with male luer lock connector actually contained product with a female luer lock connector. | Packaging | The firm, Salter Labs, called their customers to discuss the recall and followed up with a "PRODUCT RECALL" letter dated May 29, 2012. The letter was distributed to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory and return any of the lots in their possession to Salter Labs; contact Salter Labs Customer Service at 1-800-235-4203 to arrange for return of the product, and complete and return the attached form/record via fax to: 661-854-6816 or Toll Free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 242 | 242 cases | Nationwide distribution: USA including states of: FL, IN, KS, PA and TX. | Terminated | July 27, 2012 | 58 | 242
Z-1902-2012 | 62017 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | MULTIGENT Acetaminophen for use with Architect/Aeroset |  |  | May 29, 2012 | 2012 | Sekisui Diagnostics Llc | Discoloration of the acetaminophen enzyme reagent (R1), generates a high calibration factor and erratic control recovery. | Storage | Sekisui Diagnostics Llc sent an "IMPORTANT PRODUCT FIELD CORRECTION" letter dated May 29, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately stop using the affected product and to destroy any remaining inventory. A Confirmation of Notification Form was attached for customers to complete and return via fax to 902-628-6504. Contact the firm at 902-628-0984 for questions regarding this notice. |  | 1600 | 1600 kits | Illinois | Terminated | November 14, 2012 | 169 | 1600
Z-1897-2012 | 62056 | No | LMF | Absorbable hemostatic agent and dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | Ethicon Surgiflo Hemostatic Matrix Kit with Thrombin Plus Flex Tip |  |  | April 03, 2012 | 2012 | Ethicon, Inc. | The lyophilized thrombin component of the Surgiflo Hemostatic Matrix kit with Thrombin has the potential to be non-sterile due to an undetected air leak which occurred during the manufacturing process. | Packaging process control | Ethicon Inc. sent a Urgent Voluntary Product Recall letter dated April 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If you have any product subject to this voluntary recall, please remove it from your inventory immediately and return it per the following instructions.Return product to Stericycle using the enclosed pre-paid shipping label.Shipping fees wil be charged directly to Ethicon through the use of this shipping label. No reimbursements will be made for the shipping charge if this label is not used. Extra shipping labels may be obtained by calling 1-866-338-2611. Credit will be issued for impacted product that is returned.Further questions please call 1-866-338-2611, or contact your Ethicon Biosurgery Sales Representative.Thank you for your coorporation and immediate assistance.. |  | 27396 | 27,396 kits | USA (nationwide) | Terminated | April 16, 2014 | 743 | 27396
Z-1871-2012 | 61670 | No | MMD | Powered toothbrush. | Dental | Dental | 510(K) Exempt | Colgate Adult Motion Whitening Battery Toothbrush |  |  | March 23, 2012 | 2012 | Colgate Palmolive Company | Brush battery cap was forcibly blown off at the end of the brush handle. | Nonconforming Material/Component | Colgate-Palmolive Company sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 22, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter requests the termination of sale and segregation of the affected product for return. Contact your Colgate-Palmolive representative to coordinate return of the affected product or if you have any questions concerning this matter. |  |  |  | Nationwide Distribution. | Terminated | June 19, 2013 | 453
Z-1895-2012 | 61952 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bigliani/Flatow The Complete Shoulder Solution Trabecular Metal Glenoid Instrument Drill with Stop, 6 mm; the products are medical instruments intended for use when performing surgery of the shoulder. |  |  | May 12, 2012 | 2012 | Zimmer, Inc. | Bigliani/Flatow The Complete Shoulder Solution Trabecular Metal Glenoid Instrument Drill with Stop, 6 mm. Catalog number 47-4301-031-00 was actually packaged and labeled as catalog number 47-4307-031-00, lot 62071190. | Mixed-up of materials/components | Zimmer sent URGENT DEVICE REMOVAL NOTIFICATION emails to distributors on May 12, 2012 to remove the product from all locations in their territories. This was an initial notification to prevent further distribution and use of the product. Zimmer Inc. sent follow-up notifications dated June 7, 2012 to the distributors, surgeons and hospital user accounts describing the product issue and providing recommended actions. Customers were directed to (877) 946-2761 for assistance or questions. |  | 87 | 87 units | Nationwide: TX, FL, MS, MI, OH, KS, TN, KY, OK, UT, CO, CA, IL, NC, NJ, ND, and WI. | Terminated | September 12, 2013 | 488 | 87
Z-1875-2012 | 61731 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns Antegrade Cardioplegia Cannula: ROOT INFUSTION VENT CATHETER WITH 12 GAUGE BLUE NEEDLE RECTANGULAR FLANGE, 6" (15 CM) LONG |  |  | May 09, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During production of the Sarns Antegrade Cardioplegia Cannula, Terumo Cardiovascular Systems (Terumo CVS) identified the presence of a foreign substance on the inner surface of some cannulae tips.Terumo CVS' preliminary investigation found that the substance can be dislodged from the cannula surface. The substance was likely deposited during the molding process, but the exact composition of the | Material/Component Contamination | Terumo sent an Urgent Medical Device Recall letter dated May 9, 2012 to All US first consignees will receive a removal notification letter via Federal Express. The letter identified the affected products, reason for correction and actions to be taken. Customers were instructed to discontinue use of affected products and return for replacement. The letter advised customers to assure that all users are aware of this notice and to fill out and return the attached Response Form. For questions call Terumo CVS Customer Service at 1-800-521-2818 or by fax at 1-800-292-6551. |  | 53139 | 53,139 (each) in total | Worldwide Distribution - US (nationwide) including the states of: AL, AR , AZ , CA, CT , FL, IA , IL, IN , KS , LA , MI , NM, MO , NC, NE , NJ , NV, NY , OH, OK , PA, RI , SC , TN, TX, WA , WI, WV and the countries of: BELGIUM, BRAZIL, Byelorussian SSR, CAMBODIA , Dominican Republic, Honduras, Japan, Malaysia, Mexico, Nicaragua, Saudi Arabia, Singapore, South Africa, Thailand, Trinidad and Tobago, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | April 24, 2013 | 350 | 53139
Z-1903-2012 | 61966 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS Panels Creatinine test strips. |  |  | May 15, 2012 | 2012 | Polymer Technology Systems, Inc. | During a routine product investigation, a creatinine sample with a value less than 1.0 mg/dl was tested on a version 2.60 CardioCek PA meter and the meter provided a result of 0.08mg/dl, which is not the reportable measuring range for the CardioChek PA meter for the creatinine assay. | Software design | Polymer Technology Systems, Inc sent a URGENT FIELD CORRECTION letter dated May 15, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed the customers to discontinue use and dispose of the product. PTS Inc. will provide appropriate replacement for any product that is properly identified on the attached Customer Acknowledgement and confirmation letter,If you have any questions please contact Customer Service toll-free at 1-877-870-5610, or +1-317-870-5610 ( outside the US ) |  | 1650 | 1650 vials | Worldwide Distribution -- Nationwide Distribution including the states of CA and VA., and the countries of China, Poland, Ireland and France. | Terminated | July 20, 2012 | 66 | 1650
Z-1901-2012 | 61943 | No | BZE | Breathing system heater. | Anesthesiology | Anesthesiology | 510(k) | AirLife Isothermal Breathing Circuit Accessory - Infant Patient Y Connector; an Rx device for single patient use; 30 connectors per box; CareFusion, 22745 Savi Ranch Parkway, Yorba Linda, CA 92887 USA, Made in Mexico |  |  | May 29, 2012 | 2012 | Carefusion 2200 Inc | The Y adapter within the breathing circuit may spontaneously crack, causing a leak in the closed ventilation system. This leak could lead to a decrease in the tidal volume delivered to the ventilated patient. | Equipment maintenance | The firm, CareFusion, sent an "URGENT RECALL NOTICE" dated May 29, 2012 to all of the affected customers on the same date via registered return receipt mail. The notice described the product, problem and actions to be taken. The customers were instructed to immediately destroy any affected product in-stock at their facility and to NOT further distribute these products. NOTE: The notice stated "Carefusion is removing from the the market all lot numbers less than lot #0000414257 or begin with the letter 'Y' for all Infant Breathing Circuit product codes listed in the provided attachment".The customers were also instructed to not return the product to Carefusion. Return the product directly to: Owens and Minor Health Care Logistics at 1651 California Street, Suite C, Redlands, CA 92374, Attn: Recall Coordinator, for immediate replacement.NOTE: The notice the stated "Failure to follow these important measures may result in a potential patient safety risk".In addition, the customers were instructed to forward the recall notice to any customers to whom these products were sold, notifying them of the potential risk, advising them to discard or return any CareFusion AirLife Infant Breathing Circuits with the noted product/lot numbers for replacement; and to complete and return a enclosed reply form indicating the amount of inventory in stock, amount destroyed and amount returned. The distributors were requested to complete and return the attached Acknowledgement Form indicating that they have received the recall notice and that they communicated the Recall Notice to all customers who have purchased the affected breathing circuits via Fax to (847) 473-7774 or email to David.Wehrheim@Carefusion.com. They were also instructed to contact CareFusion Channel Management at (866) 802-6320 for further details on return and/or replacement of stock.If you have any questions related to this information, please call (847) 473-7495 or email David.Wehrheim@Caref |  |  |  | Worldwide distribution: USA (nationwide) including Puerto Rico and internationally to: Canada, Egypt, France, Germany, Great Britain, Italy, Kuwait, Mexico, Portugal, Saudi Arabia, Spain, Taiwan and United Arab Emirates. | Terminated | January 22, 2016 | 1333
Z-1924-2012 | 61749 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LUSTER HIP SZ 10 |  |  | May 16, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a Field Safety Notice for the LUSTER Cemented Polished Femoral Hip Stems and the ULTIMA LX Cemented Polished Femoral Hip Stems that were implanted with any metal-on-metal (MoM) articulation (including a head of CoCr and an acetabular cup where the cup is of an all metal construction (monoblock or modular)).The LUSTER Cemented Polished Femoral Hip Stems and the U | Device Design | DePuy sent an "URGENT RECALL INFORMATION-FIELD SAFETY NOTICE" letter dated May 16, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 574-372-7333 for questions regarding this notice. |  | 69 | 85) LUSTER Cemented Polished Femoral Hip Stems in 69 lots | Nationwide Distribution-including the states of CA, CO, IA, NC, OH, TX, WA, WI, and WV. | Terminated | March 05, 2013 | 293 | 5064
Z-1905-2012 | 61555 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TDX SP Power Wheelchair, TDX SR Power Wheelchair, TDX SI Power Wheelchair. Model Number(s): TDXSI, TDXSI-2, TDXSI-2-S, TDXSI-CG, TDXSI-HD, TDXSI-HD-S, TDXSIV, TDXSIV-2, TDXSIV-2-S, TDXSIV-HD, TDXSIV-HD-S, TDXSP, TDXSPBASE, TDXSP-CG, TDXSPEURO, TDXSP-MCG, TDXSP-MCG-GT, TDXSR, TDXSR-CG, TDXSR-CG-HD, TDXSR-MCG, FDX, FDX-CG, FDXEURO-14, FDX-MCG. Mrf. By: Invacare Corporation, 1200 Taylor Street, Elyria, OH 44035. |  |  | April 06, 2012 | 2012 | Invacare Corporation | Invacare Corporation decided to recall the product because of a potential risk of fire and serious injury. The post fuse assemblies could potentially have a short between the positive ring terminal and the terminal battery bracket which would create an alternate current path around the fuse. | Employee error | Invacare sent an "URGENT: RECALL INFORMATION" letter dated June 7, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. THe letters instructs customers to examine their inventory for the affected product and quarantine any product found. AProvider Response Form was enclosed for customers to complete and return via fax to 888-345-0289. Contact the firm at 888-641-9738 for assistance with this recall. |  | 680 | 665 affected power wheel chairs and 15 post-fuse assembly service parts have been identified. | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and Puerto Rico and the countries of Australia, Germany, New Zealand. | Terminated | May 20, 2014 | 774 | 680
Z-1904-2012 | 61979 | No | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ARCHITECT CA 19-9XR Reagent Kit; ARCHITECT CA 19-9 Reagent Kit, list 291-20 (4 x 100 tests), 2K91-25 and 2K91-27 (1 x 100 tests); |  |  | May 30, 2012 | 2012 | Abbott Laboratories | The six ARCHITECT CA 19-9XR reagent lots are demonstrating a shift up in patient results. The Abbott Controls however do not detect the upward shift and will test within range. | Nonconforming Material/Component | The firm, Abbott Laboratories, sent a "Product Recall Immediate Action Required" letter dated May 30, 2012 to all Architect CA 19-9XR customers who received the affected reagent lots. The customers were directed to immediately discontinue use and destroy any remaining inventory of the lots listed; to review the patient results obtained with the affected lots with the treating physician or heath care provider to ensure they are aware of the potential analytical issue and can determine whether re-evaluation of the patient is necessary; to forward a copy of the letter to any laboratory to whom they may have forwarded the reagent; and to retain a copy of the letter for their records and to complete and return the Customer Reply Form via fax to: 1-800-777-0051 or e-mail: QAGCO@abbott.com. For questions were call your local Customer Service representative or call the field action coordinator at 847-938-1923. |  | 8672 | 316 kits - USA, 8,356 kits ROW | Worldwide Distribution - US (nationwide) the states of: Alabama, Arkansas, Arizona, California, Florida, Iowa, Indiana, Louisiana, Michigan, Minnesota, New Jersey, Nevada, New York, Ohio, Oklahoma, Pennsylvania, South Dakota, Texas, Virginia, Washington, including Puerto Rico, and the countries of Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Barbados, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Gaza &amp; Jericho, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kyrgyzstan, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Mexico, Montenegro, Morocco, Netherlands, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Togo, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Vietnam. | Open, Classified |  |  | 8672
Z-1817-2012 | 61253 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | TX1 Tissue Removal System Console. Model: TX1 |  |  | February 22, 2012 | 2012 | American Optisurgical Inc | Two reasons for recall.1. Incorrect pressure relief valve installed inside console, leading to a remote chance of a risk to health. 2. Technical Specifications, under Irrigation, the Operating Pressure value is incorrect. | Nonconforming Material/Component | The firm, American Optisurgical Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter via email on February 23, 2012. An updated "URGENT: MEDICAL DEVICE RECALL" recall letter dated March 1, 2012 was sent via email on March 1, 2012 to its customer. This updated recall letter described the product, problem and actions to be taken. The letter notified the two reasons for recall: " An incorrect pressure relief valve has been installed inside the console. Should the regulator that controls pressure delivered to the inflation cuff fail, the pressure relief valve would not open until a pressure of 10psi was achieved, which may present a health risk". "The Operator's Manual included with the system, Rev. 1, has a typographical error. Specifically, in Section E. Technical Specifications, under Irrigation, the Operating Pressure value is incorrect". The updated recall included the above second reason for recall. The firm instructed the customer, for replacement of the pressure relief valve, and to return the revision 1 operator manual to the company's address using a pre-paid FedEx account. For questions call 949.580.1266 M-F, 7:00am to 5:00pm PST. |  | 5 | 5units | Nationwide distribution: MN only. | Terminated | June 15, 2012 | 114 | 5
Z-1793-2012 | 61332 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Columbus Tibia Plateau (component of the Columbus REVISION Knee System) |  |  | February 21, 2012 | 2012 | Aesculap Implant Systems | During an internal quality inspection, the tibia plateau was found to be out of specification. There is a potential for misalignment between the tibia plateau and attachment of the extension stem. | Nonconforming Material/Component | Aesculap Implant Systems (AIS) initiated their recall of this product by first telephoning consignees and then on February 21, 2012 by sending an IMPORTANT RECALL NOTIFICATION letter dated February 21, 2012 to all consignees. The letter identified the affected product, problem and actions to be taken. The notification told consignees not to use the affected product and to return any stock to AIS (USA) promptly. For questions contact Scott Stephan at (610) 984-9239 or Christian Gabriel at (610) 984-9300. |  | 8 | 8 | Nationwide Distribution including PA and TX. | Terminated | January 15, 2014 | 694 | 8
Z-0857-2012 | 60853 | No | JEQ | Powered toothbrush. | Dental | Dental | 510(K) Exempt | Arm and Hammer Spinbrush Dual Action Powered Toothbrush Pro-Clean, various models. |  |  | December 22, 2011 | 2011 | Church &amp; Dwight Inc | Church and Dwight plans to take actions to best convey to the public the potential small parts choking hazard with using a Spinbrush beyond the recommended 3-month brush head life. | Device Design | Church &amp; Dwight Co., Inc. proposed a consumer education media campaign starting on December 22, 2011 with the following statement:"Safety Notice: Please remember to replace your brush head after 3 months of use, or if the brush is damaged or the parts become loose. Extended usage, loose parts or excessive wear could lead to brush head breakage, generation of small parts and possible choking hazard. Inspect the brush for loose parts before use."The campaign will be conducted in the following media and anticipate that the combined consumer reach of all of these vehicles to be approximately 144 million people an average of 4 times.- Television campaign will be run during the entire Q1 of 2012 (January -March). -Print ads that will appear in the December 30th issues of People, Star, US Weekly and Entertainment Weekly. -Church &amp; Dwight Website will contain the Safety Alert notice starting December 22nd. (www.spinbrush.com)Consumer Relations IVR (Interactive Voice Response) to consumers calling the Spinbrush toll-free numbers). (800-353-3384 and 800-5610752).4-30-121) A print advertisement (Attachment A) containing the Sonic Safety Notice that will appear in each magazine (Star, People, US Weekly, and Entertainment Weekly) as noted:Five total print insertions are planned (on sale dates listed):I. 4/30 -Star2.5/14 -Us Weekly3.5/28 -People4.6/11 -Entertainment Weekly5.6/25 -Star2) A disk containing the 15 second commercial containing the safety notice.3) A document (Attachment B) of the National TV slots showing when the commercials will run in the quarter, and another with the shows during which they will air. TV ads will begin running the week of April 16th. The DRTV schedule is determined weekly and we are informed after the fact when the ads where shown. The DRTV stations are included at the bottom of the National TV listing." In total, combined print and TV are estimated to reach 163MM impressions (measurement target: women |  | 22783000 | 22.783 million units | Worldwide Distribution. | Terminated | April 14, 2015 | 1209 | 41417300
Z-1927-2012 | 62283 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | AirLifeÂż MiniSpacerÂż MDI Dispenser Adapter. Dual-Spray MiniSpacerÂż Metered Dose Inhaler (MDI) Dispenser/Adapter with 15 mm connections. |  |  | June 06, 2012 | 2012 | Thayer Medical Corporation | Internal data identified a molding defect in the 15 mm O.D. / I.D. MiniSpacer meter dose inhaler (MDI) adapter that results in an incomplete fill of the aerosol dispensing pathway. | Device Design | Thayer Medical sent an Urgent Medical Device Recall letter dated June 7, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine any affected product. Customers were also instructed to destroy and/or dispose of the product in a proper manner and certify to Thayer Medical that they have done so by complelting the attached Certificate of Destruction. The Certificate of Destruction should be sent to Thayer Medical via fax or email. If customers prefer to return the product to Thayer Medical they should contact Customer Service at info@thayermedical.com or fax 1-520-790-5854 for instructions.For questions regarding this recall call 520-790-5393, ext 2210. |  | 19340 | 19340 | Worldwide Distribution - USA including CA, KY and Europe. | Terminated | January 15, 2013 | 223 | 19340
Z-1936-2012 | 61998 | No | LJG | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Microbiologics 217 Osseo Avenue North, St. Cloud, MN 56303, Gram Stain Control Slides, SLO3-10, ATCC Licensed Derivative, MediMart Europe, 11 rue Emile, Zola B.P. 2332, 38933 Grenoble Cedex 2 France. |  |  | May 29, 2012 | 2012 | Microbiologics Inc | Microbiologics is recalling a number of lots of KWIK-QC Gram Stain Slides. This product gram negative control organism (E.coli) is not properly fixated on the slide, which may cause the user to unintentionally contaminate the patient area of the slide by flooding the control well with stain. This contamination could potentially lead to misdiagnosis and improper treatment of a patient. | Nonconforming Material/Component | Microbiologics sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 30, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers not to use or distribute the affected product and to quarantine and return to the firm any remaining product. A Customer Response Form was attached for customers to complete and return to the firm. Contact Customer Service at 320-229-7057 for questions regarding this recall. |  | 13720 | 13,720 slides (10 slides per box) | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, GA, IL, IA, KY, ME, MA, MI, MN, MS, NJ, NM, NY, PA, PR, SC, TX, VT, and VA. and the countries of SWITZERLAND, ITALY, ISRAEL, IRELAND, COLOMBIA, SWEDEN, INDIA, SINGAPORE, SERBIA, SAUDI ARABIA, CANADA, TAIWAN, SOUTH AFRICA, FRANCE, BRAZIL, MALTA, MACEDONIA, MALAYSIA, FINLAND, ECUADOR, VENEZUELA, REPUBLIC OF KOREA, NETHERLANDS, PERU, UNITED ARAB EMERITES, TURKEY, POLAND, and LEBANON. | Terminated | January 08, 2013 | 224 | 13720
Z-1941-2012 | 57350 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Victory 9 with Power Seat (SC609PS) |  |  | November 26, 2010 | 2010 | Pride Mobility Products Corp | The main harness connector was manufactured without the required back shield sealing protection. As a result, there is the potential for water to seep through the nonconforming harness. | Nonconforming Material/Component | Three (3) pieces of information was distributed in the notification. They are: (1) Letter of Correction to the Provider (2) Letter of Correction for the Consumer (provider to issue to consumer) (3) Product Insert for Victory 9 Controller Interface Harness Replacement. The letter of Correction for Victory 9 with Power Sear (SC609PS), dated November 2010, identified the affected product and the reason for the correction. Customers were to record the serial number(s) and contact their provider's office to make an appointment to have the harness replaced. |  | 186 | 186 | Worldwide Distribution | Terminated | March 22, 2013 | 847 | 186
Z-1933-2012 | 62137 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PathFinder NXT Pivoting Percutaneous Rod Holder |  |  | June 05, 2012 | 2012 | Zimmer, Inc. | Complaints have been reported where, during surgery, the surgeon was unable to loosen orremove the inserted rod from the rod holder instrument. The complaint investigation showedthe hex mating feature of the set screw was broken and material missing. There were noreports of fragments of the set screw being left in the patient or that there was any patient oruser injury, but these events we | Device Design | The firm, Zimmer Spine, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 5, 2012 to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the notification and ensure they are aware of the content; to immediately remove and discontinue use of Pivoting Percutaneous Rod Holder (part number 3564-1) from Pathfinder NXT instrument sets in their possession and follow the PathFinder NXT Surgical Technique; to notify their Zimmer Spine customer service representative to obtain return authorization; and return the complete Surgeon Contact Certification Form within 10 working days from receipt of notification to ronald.musselman@zimmer.com or Fax to 1-512-258-0995.If you have any questions, call Zimmer Spine Customer Service at 866-774-6368. |  | 77 | 77 units | Worldwide distribution: USA (nationwide) and countries of: Australia, France and Germany. | Terminated | February 14, 2013 | 254 | 77
Z-1934-2012 | 62240 | No | HWW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S4 Cervical (S4C) Revision Instruments (FW099R and FW101R). The S4C Revision Instruments assist in the removal of an occipital cervical screw during revision surgery. |  |  | March 07, 2012 | 2012 | Aesculap Implant Systems LLC | Aesculap Implant Systems, LLC has initiated a voluntary removal of the S4 Cervical (S4C) Occiput Revision Instrument that is part of the S4C Occipital Instrument Set. The instruments may not thread into the screw properly during a revision surgery due to tolerance issues between the two instruments. | Component design/selection | The firm initiated their recall of this product by issuing an "Important Product Removal Notification" letter to all consignees dated March 7, 2012. The letter described the problem, advised them not to use the product, and instructed consignees to return the instruments to the recalling firm with the attached Distribution Inventory Sheet. Customer Service may be contacted if questions at 1-866-229-3002. |  | 15 | 15 | Nationwide distribution, including the states of CA, CO, IL, LA, NV, OH, and TX . | Terminated | August 08, 2012 | 154 | 15
Z-1940-2012 | 61874 | No | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Depuy Spine Side Fire Needles 13G X 4IN Side-Fire Needle, part of the Confidence Spinal Cement System |  |  | May 04, 2012 | 2012 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Additional Instructions for Use and Surgical Technique for the Confidence Needle due to reports of breakage with the side fire needles | Labeling design | Depuy Spine sent a "CORRECTION NOTICE" to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Correction Notification Form was attached for customers to complete and return via fax to 508-828-3070. For questions about Confidence Needles, please contact your local DePuy Spine Sales Representative. If you need assistance in locating your local Sales Representative please contact JJHCS CIS @ 1 -800-255 -2500 . |  | 6979 | 6979 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, Austria, China, Croatia, Czech Republic, France, Germany, Greece, Hungary, Italy, Lebanon, Netherlands, New Zealand, Poland, Russian Federation, Singapore, South Africa, Spain, Switzerland, Turkey, UAE, and UK. | Completed |  |  | 8896
Z-1935-2012 | 62447 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Vacuette Quickshield Complete Plus, 21G x 1 (0.8 x 25 mm), Greiner bio-one. |  |  | April 02, 2012 | 2012 | Greiner Bio-One North America, Inc. | : The Quickshield Complete Plus may have a defect (split at the base), which can cause the green safety shield attachment of the Quickshield holder to be loose and unstable. | Process control | greiner bio-one sent a notification letter dated April 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the attached Product Disposition Site Confirmation form and fax to 800-726-0052 or email to patech@gbo.com. This form will confirm the affected product has been discarded/destroyed. The firm will replace the product after the completed form is returned to them. Customers with questions should call 800-515-8112.For questions regarding this recall call 704-261-7823. |  | 263.5 | 19.5 cases (5,850 pieces) and 244 cases (73,200 pieces)= 263.5 (73258 pieces total) | Nationwide Distribution including IL, MS, OH, PA, TN and VA. | Terminated | February 27, 2013 | 331 | 263.5
Z-1943-2012 | 61525 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Single Shot Epidural Trays, Sterile, Integra Pain Management, West Valley City, Utah 84119, each kit contains sterile instruments, drugs and other medical materials that will support the administration of spinal anesthesia for use by a surgeon, nurse or other licensed medical professional. Model number is 3403644. |  |  | June 07, 2012 | 2012 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences is recalling one lot of Single Shot Epidural Trays as the product was inadvertently shipped missing injectable components. | Packaging process control | Integra sent an Urgent - Medical Device Recall letter dated June 7, 2012, to one domestic consignee via Federal Express with a recall acknowledgement form followed up with a telephone call. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the Recall Acknowledgement Form including quantity of affected product in their inventory. Customers were also instructed to contact Integra Pain Management Customer Service to arrange for returns of all affected product. Product reconciliation will be performed to account for all affected products returned. Returned product will be reprocessed to correct for the missing components.Customers with questions were instructed to call Integra Pain Management Customer Service at 1-800-241-2210.For questions regarding this recall call 609-936-2311. |  | 40 | 4 cases x 10 units per case | Nationwide Distribution including MA | Terminated | August 07, 2012 | 61 | 40
Z-1942-2012 | 62270 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Philips Xcelera Connect, Software R2.1 L 1 SP2, an interface engine for data exchange. |  |  | June 14, 2012 | 2012 | Philips Healthcare Inc. | Xcelera Connect R2.1 L 1 SP2 , incomplete information arriving from unformatted reports interface | Software design | Philips Healthcare sent an "URGENT-FIELD SAFETY NOTICE" letter dated June 1, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Philips representative will contact customers when the upgrade becomes available. Contact your local Philips Representative for questions concerning this issue. |  | 226 | 226 units | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Chile, Denmark , Estonia, Finland, France, Germany ,Hong Kong, Ireland, Italy, Malaysia, Netherlands, Poland, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Taiwan, United Kingdom, and Utd. Arab. Emir. | Terminated | March 14, 2017 | 1734 | 226
Z-1957-2012 | 60341 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Central Control Module for System 1: System 1 Base, 100/120V |  |  | June 22, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo CVS has received reports of the Central Control Monitor (CCM) for Terumo Advanced Perfusion System 1 losing partial or full functionality. Symptoms that have been associated with a malfunctioning CCM include: full or partial loss of display, loss of control functions; inability to distinguish various status, alert, alarm conditions, the inability to operate the touch screen, or the CCM | Device Design | The firm, TERUMO, sent an "URGENT MEDICAL DEVICE RECALL" notice dated June 22, 2012 and addendum to the Operator's Manual via Federal Express to its customers. The letter described the product, problem and actions to be taken. The notice states to the customers "Please note that we are not recommending that you stop using your Central Control Monitor". The customers were instructed to review the Medical Device Recall notice; assure that all users are aware of this notice; place the Addendum at the beginning of the Central Control Monitor Section (page 4-1) of the Operator's Manual and complete and return the attached Customer Response Form to the fax number indicated on the form. If you have any questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818; customer service hours are from Monday-Friday, 8 AM - 6 PM ET. |  | 992 | 992 units | Worldwide distribution: USA (nationwide) and countries including: Argentina, Belgium, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates (UAE), and Vietnam. | Terminated | March 06, 2013 | 257 | 1640
Z-1950-2012 | 62470 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Drager PT 4000 Phototherapy System |  |  | July 02, 2012 | 2012 | Draeger Medical Systems, Inc. | A routine internal quality check found that the Drager PT 4000 Phototherapy System is fitted with a 16 Amp rated fuse instead of the specified 1.6 Amp rated fuse. This may mean that the internal fuse would not trip promptly in the event of an internal short-circuit. The possible consequence is the wall outlet circuit breaker would trip first on the device may severely overheat. There have not b | Other | Drager Medical Systems sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 2012 to all affected customers. The letter identifies the product, problem, and actions to bee taken by the customers. In addition, a Customer Reply Card was included for customers to complete and return. Contact the firm at 800-543-5047 for questions regarding this recall. |  | 65 | 65 | Nationwide Distribution-USA (nationwide) and Puerto Rico. | Terminated | July 01, 2013 | 364 | 65
Z-1945-2012 | 61807 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Dual Incu i, (Atom Infant Incubator Model 100) Atom Medical Corporation 3-18-15 Hongo, Bunkyo-Ku, Tokyo, Japan 113-0033 Usage: An incubator for infants. |  |  | May 01, 2012 | 2012 | Atom Medical Corporation | There is a potential for the device's canopy to move unintentionally when using the incubator mode or radiant warmer mode. | Device Design | Atom Medical sent an Urgent Field Safety Notice dated May 1, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately identify all affected product in their inventory and quarantine them until the rework was complete. Customers were also asked to identify any customers to whom they further distributed the product and to contact their customers at once and arrange a time in whcih the filed service technician could perform the rework. Customers were asked to return the Recall Response Forms to Atom Medical (kirautc@atomed.co.jp, tsugino@atomed.co.jp,or matsumoto@atomami.co.jp). All forms should be completed and returned by January 31, 2013.For questions regarding this recall call 512-327-9997. |  | 377 | 377 units | Worldwide Distribution - USA including PA. Internationally to Japan, Kuwait, Saudi Arabia, Lenanon, Italy, Spain, China, East Europe, Slovakia, Poland, Czech Republic, Ukraine, Egypt, Russia, Algeria, Colombia, Germany, Malaysia, Switzerland, Australia, Europe, Portugal, Iran, Israel, Philippines, UAE, Korea, and Netherlands | Terminated | January 31, 2013 | 275 | 377
Z-1946-2012 | 61832 | No | DSK | Blood pressure computer. | Cardiovascular | Cardiovascular | 510(k) | Safedraw Venous Close Loop System, Blood Sampling Set, packaged in a tyvek bag, Model/Part Number 688695, Argon Critical Care Systems, Singapore Pte. Ltd., 198 Yishun Avenue 7, Singapore 768028. |  |  | May 04, 2012 | 2012 | Argon Medical Devices, Inc | Some of the kits may exhibit a potential defect with the pouch seal. The pouch issue may compromise the sterile barrier. | Packaging | Argon Medical Devices sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Recall Response Form was attached to the letter for customers to complete and return to the firm via fax to 903-577-9393. Contact the firm at CRS@argonmedical.com for questions regarding this recall. |  | 10 | 10 pieces | Distributed only in Missouri. | Terminated | January 14, 2013 | 255 | 10
Z-1956-2012 | 61793 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | CIVCO VirtuTRAX Instrument Navigator, Reusable non-sterile general purpose driveBAY, trakSTAR, electromagnetic sensor (8mm) with VirtuTRAX Instrument navigator and (8.9 tapered to 1.3 x 147cm) telescopically-folded CIV-Flex sensor covers (5), Ref. 610-1080. The firm name on the label is CIVCO, Kalona, IA. The kit contains sensor component part number 410-240. |  |  | May 07, 2012 | 2012 | Civco Medical Instruments Inc | The general purpose electromagnetic sensor was incorrectly calibrated causing potential variability in the projected position of the needle path with the ultrasound device. | Component design/selection | CIVCO sent an Urgent Medical Device Recall letter dated May 3, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to return all general purpose sensors as the product may not function as intended. To return the affected product customers should contact the OEM Sales Coordinator at 319-248-6521 to schedule the return of their device. Upon receipt of the snesor, CIVCO will arrange for recalibration. Once recalibration is completed, the sensor will be returned to the customer at no charge. Customers were asked to forward this communication to all their customers who purchased the affected product. Once CIVCO has received the customer list, they will contact the end-user to request they return their general purpose sensor directly to CIVCO for recalibration. For questions customers should call 800-441-6741.For questions regarding this recall call 319-656-4447. |  | 24 | 24 kits | Worldwide Distribution - USA including WI, CA, MA, IN, WA, and NJ. Foreign distribution was made to Germany, Canada, Denmark, Hong Kong, Ukraine, Japan, Italy, Korea, and France. | Terminated | January 07, 2015 | 975 | 384
Z-1951-2012 | 61529 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AEM Disposable Electrodes, Device, Electrosurgical, Cutting and Coagulation and Accessories. AEM Disposable Electrodes used as electrosurgical accessories intended for ablation, removal, resection and coagulation of soft tissue in open, endoscopic and laparoscopic surgical procedures channeling monopolar high-frequency electrical current from compatible electrosurgical generators. The device is model FC03XX series. |  |  | June 15, 2012 | 2012 | Encision, Inc. | Encision is performing a field correction to their AEM Disposable Electrodes due to two consumer complaints reporting unanticipated alarms from the AEM monitor while using the device. This issue could result delays or complications during surgical procedures. | Device Design | Encision sent an Urgent Medical Device Recall letter dated June 13, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to lubricate the contact on the instrument cord with a conductive gel (to be supplied by Encision) prior to making the connection of the cord to the instrument. Instructions for applying the gel were attached. This corrective action was temporary; allowing customers to use their inventory of devices and for new product in the short term, until a design correction can resolve the problem. Customers would be contacted by their Encision Sales Representative to review this information for their specific situation and train their staff on its correction. For questions customers should call 303-339-6917 or Customer Service at 800-998-0986.For questions regarding this recall call 303-339-6917. |  | 48 | 48 boxes | Worldwide Distribution - USA including CA, TX, GA, NY, NC, OH), and New Zealand. | Terminated | August 07, 2012 | 53 | 48
Z-1925-2012 | 62118 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | FLOW-I Anesthesia System |  |  | February 20, 2012 | 2012 | Maquet Cardiovascular Us Sales, Llc | A technical alarm may be generated on the FLOW-i system when using the MAN/AUTO switch to change ventilation modes from manual to automatic or from automatic to manual if the switch is not fully engaged in the "on" or "off" position but remains in the "in-between position. | Device Design | The firm, Maquet Getinge Group, sent an "URGENT DEVICE FIELD CORRECTION" letter dated February 20, 2012 to its customers. The letter describes the product, problem and action to be taken. The letter state that a Maquet Service Representative will upgrade the customers software in the FLOW-i Anesthesia System to the current revision and also provide the customers with the corresponding user's manual. Note: This upgrade includes other minor improvements. The customers were instructed to sign a document to verify satisfactory completion of the work once the software upgrade was completed.If you have questions or require additional information, please contact your local MAQUET representative or Technical Support at 888-627-8383.A press release was issued on 7/6/2012. |  | 284 | 284 units | Worldwide distribution: USA (nationwide) state of GA only and countries of: Argentina, Australia, Austria, Belgium, China, Czech Republic, Ecuador, Finland, France, Germany, India, Ireland, Israel, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Russia, Serbia, Singapore, South Africa, Spain, Sweden incl Maquet Critical Care, Switzerland, Thailand, United Arab Emirates, and United Kingdom. | Terminated | September 09, 2015 | 1297 | 284
Z-1929-2012 | 61927 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage Cardiac Panel PN 97000HS |  |  | May 22, 2012 | 2012 | Alere San Diego, Inc. | Certain lots of the affected products may have significantly decreased precision relative to the package insert, which could result in an increased frequency of false positive or false negative results. These false positive and false negative results are unpredictable within lots and may not be detected by Quality Control testing. | Process design | Phase 1 of the recall included a recall communication being initiated on May 22, 2012 with Alere forwarding a recall letter to all their customers who purchased the Triage CardioProfiler Panel, Triage Cardiac Panel, Triage Profiler SOB Panel, Triage BNP, and Triage D-dimer. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to discontinue use of the affected product, and to complete and fax the enclosed verification form within 10 days. Phase 2 of the recall included a recall communication letter being forwarded to customers on 6/11/12. The second letter included additional lots of Triage products that go beyond the lots included in an earlier recall notification letter. The letter informed the customers that the additional lots have significantly decreased precision relative to the package insert which could result in an increased frequency of false positive or false negative results. A third letter dated June 12, 2012 included 3 more lots for the Triage Cardio Profiler Panel PN 97100CP and Alere Triage Profiler SOB Panel PN 97300. The letter informed the customers that the additional lots have an increased frequency of Troponin I results &gt;0.05 ng/mL for samples which are found to be below 0.05 ng/mL upon additional testing.Customers with questions about the information contained in the notifications, were instructed to contact Alere San Diego at (877) 308-8287 or at 9975 Summers Ridge Road, San Diego, CA 92121. |  | 6044 | 6,044 kits | Nationwide Distribution | Terminated | October 18, 2013 | 514 | 60074
Z-1948-2012 | 62116 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Chorus Mobile Patient Lifter; manufactured by Arjo Ltd., St. Catherine Street, Gloucester, GL1 2SL, England, U.K.; Models KKB5060-21US, KKB5062-21US and KKB5080-21US. |  |  | June 29, 2012 | 2012 | Arjo, Inc. dba ArjoHuntleigh | ArjoHuntleigh has received a limited number of customer reports concerning injuries related to an exposed metal support bracket when the proactive knee pad assembly is removed. | Labeling False and Misleading | ArjoHuntleigh sent an "URGENT FIELD CORRECTION RECALL" notice dated June 25, 2012, to all affected customers via UPS Ground on June 29, 2012. The notice informed the accounts of the potential for severe patient injury if the proactive knee pad support is removed from the lift without the removal of the knee pad support bracket. The accounts were reminded that the entire knee pad and support bracket assembly must be removed from the lift when required to do so for exercise. An additional warning statement with diagrams was included in the notice. The accounts were requested to disseminate the information to all users of the affected lifts, to include a copy of the notice in the Instruction for Use Manuals for their lifts, and to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 3234114 | 1,989 units in the U.S., 1,626 units in the rest of the world | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Belgium, Bermuda, Canada, Cyprus, Czech Republic, Denmark, Dubai, Estonia, Finland, France, Germany, Gibraltar, Great Britain, Greece, Hong Kong, Iceland, Ireland, Israel, Italy, Japan, Kenya, Korea, Luxembourg, the Netherlands, New Zealand, Northern Ireland, Norway, Poland, Portugal, Quatar, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland and Turkey | Open, Classified |  |  | 22492034
Z-1944-2012 | 62302 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information Systems interfaced to Elekta Treatment machines. |  |  | June 08, 2012 | 2012 | Impac Medical Systems Inc | Impac Medical became aware of a problem with the recording of the Wedge Motor Units in MOSAIQ when it resumes after a machine terminates/faults, the remaining Wedge MU was being erroneously set to zero. | Software design | A User Notice began being distributed on June 8, 2012. This will be followed by contact at each site by IMPAC support to determine if the site is affected by this issue, and if so, the upgraded version will be suggested. The Notice identified the affected device, the problem, the clinical impact, along with the interim mitigations and the final resolution. Users are encouraged to monitor daily delivery and recordings of treatments to be sure that the Wedge Motor Units treatments are properly delivered. The letter also recommends that treatment charts be checked weekly and that users not override any parameters in MOSAIQ until the machine setup has been completed. Users are to contact Elekta Support for the upgrade. |  | 11 | 11 affected sites | Worldwide Distribution -- USA, including the states of VA, CT, GA, NJ and NY and the countries of Belgium, Denmark, France, and the United Kingdom. | Terminated | June 19, 2013 | 376 | 11
Z-1979-2012 | 62389 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical HOTLINE 3 Blood and Fluid Warmer, |  |  | June 27, 2012 | 2012 | Smiths Medical ASD, Inc. | HOTLINE 3 Blood and Fluid Warmer( Reorder No. HL-390 ) disposable is not correctly installed , the microswitch may become damaged and unusable | Device Design | Smiths Medical sent an Urgent Medical Device Correction -Field Safety Notice letter dated June 21, 2012 via U.S. Mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter included an Urgent Medical Device Field Safety Notice Acknowledgement Form ("Confirmation Form") that customers were instructed to complete and return to Smiths Medical for carrying out the correction and affectivity tracking purposes. The letter stated that a Smiths Medical Representative will visit the affected sites to correct the units with a shelf upgrade kit, which includes instructions and all the parts necessary to make the correction to their affected devices. Customers were provided with a form to record the serial numbers of the devices corrected so that this can be returned to Smiths Medical. For questions regarding this information call Smiths Medical Quality Department at 1-781-763-9335.. |  | 2247 | 2247 units | US Nationwide Distribution | Terminated | February 20, 2014 | 603 | 2247
Z-1974-2012 | 62319 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-Stat Portable Clinical Analyzer; Abbott Point of Care, Princeton, NJ 08540 |  |  | September 08, 2011 | 2011 | Abbott Point Of Care Inc. | Some i-Stat Analyzers become uncomfortably hot to touch in the are of the battery. This is a recall extension. | Device Design | Abbott Point Care sent an Urgent Product Recall letter dated September 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to confirm the number of affected product at their facility that were using the Old Battery Carrier. Fill out the affected Buesiness Reply Card and submit their facility information as per the instructions on the card. Upon the receipt of the New Battery Carriers, replace the Old Battery Center with the New Battery Center. Customers were instructed to immediately dispose of all Old Battery Carriers per local regulations. If customers forward the affected product to another facility, they should provide a copy of the recall letter to them and advise them to replace the carriers if required. For any questions call 800-366-8020, Option 1 or your Abbott Point of Care representative.For questions regarding this recall call 609-454-9272. |  | 857258 | 06F16-02-1295 UNITS; 210000-1076 UNITS; 210001-447 UNITS; 210002-5359 UNITS; 210003-9051 UNITS | Worldwide Distribution - USA (nationwide) | Terminated | January 28, 2013 | 508 | 2577875
Z-1967-2012 | 52194 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Rigid Suction Wand, sterile. Model numbers S033 and SPC2081. |  |  | June 06, 2012 | 2012 | Edwards Lifesciences, LLC | Edwards is recalling Suction Wand models SPC2081 and S033 because the plastic tip attached to the metal suction wand may become detached during routine use. | Process control | Edwards Lifesciences sent an Urgent Field Safety Notice Product Recall letter dated June 2012 to all affected customers. The letter identified the affected product, description of the issue and actions to be taken. Customers were instructed to examine their inventory for the affected product, quarantine and return any remaining in stock and complete the enclosed acknowledgment form with fax instructions to assure that the notification was received. For question call Edwards Customer Service at (800) 424-3278 from 8:00AM-4:30PM Pacific Time. |  | 1524 | 1524 units | Worldwide Distribution - US (nationwide) including the states of CT, TX, CA, AZ, MO, and MD and the countries of Europe and Japan. | Terminated | September 10, 2012 | 96 | 1524
Z-1964-2012 | 61231 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 |  |  | June 22, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received two complaints related to the Gas System for the Terumo Advanced Perfusion System 1. Both were reported instances of users experiencing difficulty adjusting the System 1 Gas System local flow control knob. | Error in labeling | Terumo CVS sent an "URGENT MEDICAL DEVICE RECALL-SAFETY ADVISORY" letter dated June 22, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this recall. |  | 992 | 992 Units | Worldwide Distribution-USA (nationwide) and the countries of ARGENTINA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | February 19, 2013 | 242 | 2735
Z-1970-2012 | 62373 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base 100/120V |  |  | June 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received reports of various roller pump malfunctions of the Terumo Advanced Perfusion System 1 that could result in the inability to use the pump and would then require replacement with back-up equipment. Description of malfunctions:" Unplanned pump stop or pause" Changes in pump speedIf it is necessary to replace a roller pump during cardiopu | Device Design | Terumo CVS sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to complete and return a Customer Response Form to the firm via fax at 1-800-292-6551. Contact the firm at 1-800-521-2818 for questions regarding this recall. |  | 992 | 992 units | Worldwide Distribution-USA (nationwide) and the countries of ARGENTINA, AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatamala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | March 01, 2013 | 249 | 1640
Z-1972-2012 | 62441 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | enFlow IV Fluid Warmer (510k name: EFLOW IV FLUID WARMER, MODEL |  |  | February 20, 2012 | 2012 | Vital Signs Devices, a GE Healthcare Company | The EnFlow Warmer Strapdoes not meet the requirements of the biocompatibility standard (ISO 10993) for products that mayhave contact with skin for less than 24 hours. | Device Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included customer instructions and a Customer Response Form to be completed and returned to the firm. Contact Customer Service at 1-800-932-0760 for questions regarding this recall. |  | 11803 | 11803 cases | Nationwide Distribution. | Terminated | August 19, 2013 | 546 | 11803
Z-1960-2012 | 61591 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Advantage CT/PET Fusion for Advantage Windows Workstation with Volume View CT Colonography/Navigator 2 software. (Version 4.0, 4.1,or 4.2) |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue due to incorrect alignment of fused CT and PET images associated with the CT/PET fusion capabilities of your Volume Viewer and CT/PET Fusion application. Volume Viewer allows two studies (one PET and one CT) that are registered in the same spatial domain to be viewed simultaneously in fused mode. The two images should be aligned to th | Software design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides safety instructions for customers to use until a software upgrade is available. Contact the firm at 800-437-1171 for questions regarding this recall. |  | 70 | 70 | Worldwide Distribution-USA (nationwide) including the states of CA, FL, IL, MA, NJ, NY, SC, and TX and the countries of TURKEY, TUNISIA, TAIWAN, SPAIN, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, KENYA, JAPAN, ITALY, IRELAND, INDIA, GREAT BRITAIN, GERMANY, FRANCE, EGYPT, CHILE, CHINA, CANADA, BRAZIL, BELGIUM, AUSTRIA, and ARGENTINA. | Terminated | December 06, 2012 | 196 | 70
Z-1969-2012 | 62367 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base 220/240V |  |  | June 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received reports of various roller pump occlusion malfunctions that could result in the inability to use the pump and would then require replacement with back-up equipment. Occlusion malfunctions includestiff/difficult to adjust occlusion, loss of occlusion setting, unequal roller-to-roller occlusion and inability to adjust the occlusion.If it is | Equipment maintenance | The firm, Terumo, sent an "URGENT MEDICAL DEVICE RECALL" notification dated June 25, 2012 and an update to the Operator's Manual for the Terumo Advanced Perfusion System 1 via Federal Express to its customers. The notification describes the product, problem and actions to be taken. The customers were instructed to review this Medical Device Recall notice; assure that all users are aware of this notice; place the addendum at the beginning of the Roller Pumps Section (page 5.1) of the Operator's Manual, and complete and return the attached Customer Response Form to the fax number/email address indicated on the form. Note: Terumo is continuing its investigation into the root cause of the malfunctions, and may initiate further action pending results of the investigation.If you have any questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818; customer service hours: Monday-Friday, 8 AM - 6 PM (ET). |  | 648 | 648 units | Worldwide distribution: USA (nationwide) and countries including: Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | September 18, 2015 | 1180 | 1640
Z-1961-2012 | 62346 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker brand Arthroscope and Hardware Set Tray (sterilization tray); Model/Part Number: 233-032-116; Product is manufactured and distributed by Stryker Endoscopy, San Jose, CA |  |  | June 19, 2012 | 2012 | Stryker Endoscopy | The gravity steam, Ethylene Oxide (EO) and STERRAD 100S sterilization methods are being removed from the "Instructions for Use"; due to the firm's evaluation that identified original validation was not done at worst case conditions. | Process control | Stryker Endoscopy sent a Urgent Device Correction letter dated June 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to:1. Discard any old Instructions for Use for your Arthroscope and Hardware Set Tray.2. Download the new IFU from the URL link: http://stryker.com/arthrotray .Complete and sign the enclosed acknowledgement of receipt form and fax to(408) 754-8378 or scan in and email to trays@stryker.com.* It is important to send the self addressed confirmation form back to Stryker.Stryker is committed to delivering the highest quality products to its customers. We sincerely apologize for any inconvenience this action may cause.If you have any questions regarding this letter please contact us at 1-800-624-4422 or via email at trays@stryker.com.. |  | 6388 | 6388 units | Worldwide Distribution--USA (nationwide) | Terminated | October 23, 2012 | 126 | 6388
Z-1965-2012 | 61480 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base, 100/120V |  |  | June 22, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Advanced Perfusion System 1 battery charge is being depleted faster than expected during periods of system shutdown. The System 1 may falsely indicate a longer battery life than actual battery time remaining if not subjected to regular use, such as would occur with a backup system. This may result in unexpected system shutdown. Prolonged battery depletion may cause permanent battery damage. | Device Design | TERUMO sent an Urgent Medical Device Recall letter dated June 22, 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to review the Medical Device Recall Notice, assure that all users are aware of the notice. Consignees should place the Addendum at the beginning of the System Base Section (page 3-1) of the Operator's manual, confirm receipt of this communication by faxing the attached Customer Response Form to the fax number indicated on the form. All System 1 product is on quality hold. Terumo CVS will release product including batteries as needed only in the event that customers require service. All affected System 1 customers will receive a safety notification and an addendum to the Operator's Manual. There will be a field correction but no product returned for this recall.Customers with questions should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 993 | approx. 993 units | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV.Internationally to: Republic, Egypt, Germany, Guatamala, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam, | Terminated | November 08, 2013 | 504 | 1625
Z-1959-2012 | 62406 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | CSI Stealth 360-degree Orbital Atherectomy Device (OAD), Part Number 70056-11, Model Number PRD-SC30-125S. Sterilized with Ethylene Oxide. The Stealth 360-degree Orbital Atherectomy Device (OAD) is part of the Stealth 360-degree Orbital PAD system. |  |  | June 18, 2012 | 2012 | Cardiovascular Systems, Inc. | CSI has initiated a recall on specific lot numbers of Stealth 360-degree Orbital Atherectomy devices which have the potential for an insufficient crown solder bond that may result in the crown becoming detached from the driveshaft. | Employee error | Phone calls to affected sites began June 15, 2012 to quarantine the product. These same sites were also notified with an Urgent Medical Device Recall letter, dated June 18, 2012, via FedEx and e-mail (where available) on June 18, 2012. The letter notified each customer to remove from use and return the recalled product. The customer was instructed to complete a Customer Acknowledgement Form and return it with the device(s) being returned in the Returned Goods Authorization kit provided. Customers with questions can contact customer service at 1-877-274-0901.For any questions regarding this recall call 651-259-2805. |  | 27 | 27 | Nationwide distribution including CA, FL, IL, MI, MN, PA, and TX. | Terminated | July 19, 2012 | 31 | 27
Z-1983-2012 | 61151 | No | LGW |  |  | Neurology | PMA | Eon Mini Neurostimulation (IPG) System (Model 3788); |  |  | December 19, 2011 | 2011 | Advanced Neuromodulation Systems Inc. | The firm has received 112 complaints of the Eon Mini IPGs that lost the ability to communicate or recharge due to a workmanship issue resulting in loss of pain relief and subsequent explant. The firm has determined the cause of these reports to be related to process variances in the positioning of the internal battery and printed circuit board, causing a short, and therefore, prematurely depletin | Process design | St. Jude Medical sent an Important Medical Device Recall letter dated December 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were issued the following instructions:For unimplated inventory of affected product customers should contact their SJM Representative to have the device returned to St. Jude Medical. A replacement device will be provided at no additional cost to them.For implanted affected product it is recommended that customers do not unnecessarily explant the devices associated with this advisory if the IPGs are functioning as intended. If there is a sudden loss of power or if the duration between recharges becomes significantly shorter, customers should contact their St. Jude Medical Representative. If device replacement is required due to IPG failure related to this voluntary recall notice, St. Jude Medical will provide a replacement IPG at no charge.For questions regarding this recall call 972-309-2154. |  | 25255 | 25, 255 devices | Worldwide Distribution - USA (nationwide) and internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Mexico, Netherlands, Puerto Rico, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | July 24, 2015 | 1313 | 25255
Z-1980-2012 | 62325 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Pathfinder NXT Fixed Percutaneous Rod Holder, Part numbers 3573-1 (angle handle) , Stainless Steel / Silicone Rubber, Rx only, non sterile. |  |  | June 05, 2012 | 2012 | Zimmer, Inc. | Zimmer is recalling Pathfinder NXT Fixed Percutaneous Rod Holder, because complaints have been received that the rod holder tip may crack or break when the rod is manipulated during surgery. | Device Design | Zimmer sent an Urgent Medical Device Correction and Removal notification dated June 5, 2012 to distributors and all affected accounts of Zimmer Spine. The notice identified the affected products, problem and actions to be taken. Distributors were asked to hand deliver Recall Notices to Surgeons and return the completed Surgeon Contact Certification Form within 10 working days to ronald.musselman@zimmer.com or FAX to 1-512-258-0995. Notices to Surgeons and User facilities included revised directions/precautions for the Instructions for Use and attached Surgical Techique. Devices will be removed once newly design parts become available.For any related questions or assistance about this please contact Zimmer Spine Customer Service at 866-774-6368. |  | 249 | 249 ( angle handle and straight) in total | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Australia, Belgium, France, Gelmany, Italy, Spain, Switzerland and the United Kingdom. | Terminated | September 03, 2014 | 820 | 249
Z-1982-2012 | 61923 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics, Definition PM Femoral Hip Stem Catalog Numbers 6265-3-099 through 6265-3-107. Howmedica Osteonics Corp. 325 Corporate Drive, Mahwah, NJ 07430. A subsidiary of Stryker Corp. Stryker France, ZAC Satorlas Green Pusignan, Av De Satolas Green 69881 Meyzeiu, Cedex France. Made in Ireland. |  |  | December 29, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware of the potential for cracking of the PMMA coating on their PM COCR Hip Stem product. | Employee error | Stryker Orthopaedics sent an Urgent Product Recall letter dated January 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Stryker Sales Representative to arrange for return of any affected product. Customers were also asked to fax back the attached Product Recall Acknowledgement Form as soon as possible to 201-831-6069. For any questions customers should call 201-972-2100.For any questions regarding this recall please call 201-831-5970. |  | 16913 | 16,913 units | Nationwide Distribution | Terminated | March 19, 2015 | 1176 | 16913
Z-1984-2012 | 62138 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Pudenz Flushing Valve with Integral Connectors, catalog number NL850-1330. Integra Neurosciences, Plainsboro, NJ. Device is used for the treatment of hydrocephalic patients when shunting cerebrospinal fluid from the ventricles of the brain. |  |  | April 04, 2012 | 2012 | Integra LifeSciences Corp | The pouch label of catalog NL850-1330 indicates an incorrect model and closing pressure of the device. The incorrect label reads: "Model: 12 mm-Medium Pressure" instead of "Model: 16 mm-Low Pressure". | Labeling Change Control | The firm initiated a voluntary recall by notifying the two consignees initially via email and then with an "URGENT: MEDICAL DEVICE RECALL" dated April 5th or 9th, 2012. The letter informed them of the labeling error and provided instructions regarding affected product. The International Customer Service number was provided 609-936-5400. |  | 39 | 39 units | International Only Distribution, including the countries of Thailand and Japan. | Terminated | August 02, 2012 | 120 | 39
Z-1989-2012 | 62397 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | The Axcess Cranial Screw (sometimes referred to as the Axcess Cranial Pin) is a self-tapping, self-drilling screw and is an accessory used in conjunction with the Axcess Cranial Tri-Pin Kit. |  |  | November 08, 2006 | 2006 | GE OEC Medical Systems, Inc | GE Healthcare is aware of a few instances where the Axcess Cranial Screw has broken during insertion. There is a possibility that breaking of the screws could lead to a health hazard. | Device Design | GE Healthcare sent an "URGENT REMOVAL NOTICE" dated November 8, 2006 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact GE Healthcare Customer Service at 1-800-874-7378 for questions regarding this recall. |  | 15 | 15 kits | Distributed in Washington, D.C., the states of AL, CA, FL, NY, OR, and VA | Terminated | July 11, 2012 | 2072 | 15
Z-1988-2012 | 62107 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Intellispace Portal software. |  |  | May 23, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips is issuing a software update to the Intellispace Portal systems to address a number of non-conformances. The most frequent and severe non-conformance identified involves the portal setting of "PAC Integration". This setting has a feature that allows the user to launch images of multiple patient studies into an application, but does not display a warning message informing the user that diff | Software design | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact your local Philips representative or the Customer Care Solutions Center at 1-800-722-9377, option 5, for questions regarding this recall. |  | 96 | 96 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, FL, GA, IA, IN, MI, NC, NJ, NY, OK, PA, SC, TN, TX, and WA and the countries of Australia, Austria, Canada, China, Czech Republic, Denmark, and France. | Terminated | March 28, 2013 | 309 | 96
Z-1990-2012 | 62332 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Button Switchpens, w/Extendable Needle Electrode 2 mm Exposure Single Use Electrosurgical./Extendable Pencils, 510(k) #K936304, Model Number: 75820. The device is packaged in a sterile Tyvek peel pouch within a cardboard box which contains 24 pouches per box. |  |  | June 05, 2012 | 2012 | PSC Industries Inc | The carton label for this device has the manufacturer symbol where the expiration date symbol should be. Firm officials met with FDA on 5/17/2012. At which time, they reported various labeling errors for their Button Switch Pen Electrode Electrosurgical /Extendable Pencils. | Labeling mix-ups | Olsen Medical sent an "URGENT:MEDICAL DEVICE RECALL" letter dated June 5, 2012. The letter identifies the product, problem, and actions to be taken by the customers. The letter asks that the customer immediately examine their inventory for the presence of the suspect product and quarantine the product, if found. Customers who may have further distributed the recalled product are asked to contact their (sub-account) customers and notify them of the recall. The customers are also asked to call a toll-free telephone number in order to received a Return Material Authorization (RMA) number and so identify if they would like credit or replacement inventory upon returning the product. The customer is also asked to complete and return the enclosed 'Recall Response Form' and to fax the completed form to the fax number on the form or to send the completed from back to the recalling firm via parcel post. |  | 50 | 50 units | Nationwide Distribution-including DC and the states of CA, GA, IN, KY, LA, MT, TX, UT, and WI. | Terminated | January 17, 2014 | 591 | 650
Z-1987-2012 | 62005 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IntelliSpace Portal software and Extended Brillance Workspace software |  |  | May 23, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The Intellispace Portal software and Extended Brillance Workspace software may have significantly different results in the colors and size of the affected area when the customer uses the "Time Sensitive" and "Time Insensitive" methods. | Software design | Philips Healthcare sent a Urgent Medical Device Correction letter dated May 31, 2012, to all affected customers. The letter was sent via certified mail to each customer. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. For any further information or support, customers can contact their local Philips representative or local Philips Healthcare office at (1-800-722-9377 Option 5 : Entersite ID or follow the prompts).Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. The user was informed that a field service representative will visit the facility to install the new software.Philips apologizes for any inconveniences caused by this problem. |  | 38 | 38 | Worldwide Distribution -- Nationwide Distribution including the states of AZ, OR, PA and TX and the countries of Australia, Belgium, Canada, Denmark, Finland, France, Germany, India, Indonesia, Israel, Russia, South Africa, South Korea, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | September 30, 2013 | 495 | 38
Z-1986-2012 | 62342 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx M3536A Defibrillator/Monitor with Software R.03.03 |  |  | June 20, 2012 | 2012 | Philips Healthcare Inc. | Defibrillator/monitor cannot analyze an ECG report during acquisition of 12 lead ECG | Nonconforming Material/Component | Philips Healthcare issued an Urgent Voluntary Medical Device Correction letter on June 20, 2012, to all affected customers. The Notification informs customers of the issue, identifies details (including models and software versions applicable) of units affected, instructions on actions to take while awaiting the correction, and identifies what action Philips plans to eliminate or remedy the issueA Philips Healthcare representative will contact customers to arrange for the updated R.03.03 software installation in all the affected devices.For questions customers were advised to call 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 35 | 35 units | Worldwide Distribution - USA including FL, IN, MA, M, NY, OH, TN, TX, and VA. Internationally to AUSTRALIA, HONG KONG, and INDIA | Terminated | August 14, 2013 | 420 | 35
Z-2000-2012 | 62128 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Intellispace Portal software |  |  | May 22, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | In the Multi-Modality Tumor Tracking application, when changing the contours of the lesion using the editing mode "edit contour" tool, the measurements are not recalculated causing discrepancies with the lesion measurements. | Software design | Philips Healthcare sent an Urgent-Medical Device Correction letter dated May 23, 2012, via certified mail to each consignee. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. The letter informed affected customers that a Philips Service Engineer will contact them for implementation of the software update on the affected systems free of charge. For any further information or support, customers can contact their local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 44 | 44 | Worldwide Distribution - US (nationwide) including the states of: AR, LA, MI, NC, NY, OH, OK and PA and the countries of: Australia, Canada, Denmark, France, Germany, Israel, Netherlands, Spain and Switzerland. | Terminated | March 28, 2013 | 310 | 44
Z-1997-2012 | 61958 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bayonet Forceps 8.25 (209.6 mm) 0.5mm Tips, Sterile Single Use Surgical/Bipolar Forceps, Model #20-13420, 510 (k) #K884175. The surgical device is package one (1) per Tyvek peel pouch with ten (10) Tyvek peel pouches per box. |  |  | June 01, 2012 | 2012 | PSC Industries Inc | Firm officials reported to CIN-DO that for this device, in the labeling on the inner Tyvek peel pouch, the letter 'E' in Electrosurgical was turned into an 'X' and the number '5" of the (mm tips) was turned into a '0'. This affected the description and some of the text of the labels. However, it did not t affect the Part Number and Lot Number on the label. | Labeling mix-ups | The firm, Olsen Medical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 5, 2012 to their customers via Certified mail along with representative labeling to illustrate the specific labeling issue associated with the various types and/or sizes of Bayonet Forceps surgical instruments that the customers received. The letter described the product, problem and actions to be taken. The letter instructed the customers to immediately examine their inventory for the presence of the suspect product and quarantine the product, if found; to discontinue use immediately and promptly call (800) 251-3000 for a 'Return Material Authorization (RMA) number; to complete and return the enclosed 'Recall Response Form' and fax the completed form to the fax number on the form or to send the completed form back to the recalling firm via parcel post, and for customers who may have further distributed the recalled product are asked to contact their (sub-account) customers and notify them of the recall.If you have any questions call (615) 964-5515 M-F 8:00 am - 5:00 pm CST. |  | 10 | 10 units | Worldwide distribution: USA (nationwide) including states of: CA, CT, FL, GA, MD, MI, NC, OR, SC, and TX and countries of: Canada and Philippines. | Terminated | September 15, 2015 | 1201 | 267
Z-2007-2012 | 62434 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco |  |  | June 21, 2012 | 2012 | Elekta, Inc. | Incorrect patient shift directions when the Setup reference dialog is printed out when the DICOM coordinates option is enabled for reports. | Software design | Elekta sent an Important Safety Notice to affected customers. The notice identified the affected product, problem, clinical impact and actions to be taken. The notice inform customers of how the issue will be resolved and instructions to follow when it is available. |  | 38 | 38 | US Nationwide Distribution - including the states of: CA, CT, FL, IL, IN, KY, MI, MN, MO, NJ, NY, NC, OH, OK, OR, PA, WA, and WI | Terminated | April 01, 2014 | 649 | 38
Z-1985-2012 | 62112 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | InstaTrak with Multiple Dataset Navigation, 892.2050 System, Image Processing System, Model Number IT3500 Plus. |  |  | September 15, 2006 | 2006 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. On October 11, 2006, GE Healthcare recalled GE OEC InstaTrak 3500 Plus System with Software version 5.2, the surgical Navigation and Visualization Application due to software related issues. | Software design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC sent an Urgent Safety Advisory Notice dated October 11, 2006 to all affected customers. The notice identified the affected product, problem and actions to be taken. Customers were informed that a GEHC OEC would be contacting their facility to verify the current software version installed at their site. If the issue affects their facility a software upgradewill be provided at no cost. Customers were advised to excercise caution when using the affected product until the 5.2.1software upgrade is completed. For question contact GE Healthcare Customer Service at 1-800-874-7378, option 1 or 1-801-536-4688, option 1.For questions regarding this recall call 801-536-4516. |  | 86 | 86 units | Worldwide Distribution - US (nationwide) | Terminated | August 21, 2012 | 2167 | 86
Z-2002-2012 | 62433 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal SIM |  |  | June 21, 2012 | 2012 | Elekta, Inc. | Incorrect patient shift directions when the Setup reference dialog is printed out when the DICOM coordinates option is enabled for reports. | Software design | IMPAC Medical Systems, Inc. sent an Important Safety Notice LINFFM0001/1.0 on June 21, 2012, to consignees on concerning a patch to the software. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to not use the DICOM coordinates option for Reports when printing the Setup Reference Diaglog. Customers were instructed go to www.elekta.com, select the SupportPlus Login at the tope of the page, enter their portal, select Downloads/Updates to download the latest patches.For questions regarding this recall call 770-670-2548. |  | 38 | 38 units | Nationwide Distribution including CA, CT, FL, IL, IN, KY, MI, MN, MO, NJ, NY, NC, OH, OK, OR, PA, WA, and WI. | Terminated | April 01, 2014 | 649 | 38
Z-2001-2012 | 62381 | No | OGK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | Enforcement Discretion | superDimension Systems AAS00016-xx with Software Version 4.0-4.9. |  |  | June 18, 2012 | 2012 | superDimension, Inc | superDimension, Inc., is conducting a recall of certain superDimension Systems operating Software Version 4.0 - 4.9. The affected systems are sold under the product code AAS00016-xx.The products being recalled have been found to malfunction when connected to the superDimension Locatable Guide (LG). Specifically, the Systems using Software Version 4.0-4.9 will not recognize the Locatable Guide | Device Design | superDimension sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 15, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter advised customers to discontinue use of the system and requested consignees to complete and return the acknowledgement receipt form via fax to 763-210-4098. Contact the firm at 763-210-4057 for questions regarding this recall. |  | 99 | 99 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of Alabama, Alaska, Arizona, Arkansas, California, Colorado, Connecticut, Delaware, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New Mexico, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, South Dakota, Tennessee, Texas, Vermont, Virginia, Washington, West Virginia, and Wisconsin and the countries of Canada, Australia, Austria, Belgium, Czech Republic, Denmark, France, Germany, Hungary, Italy, Netherlands, South Africa, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | February 28, 2013 | 255 | 99
Z-2012-2012 | 62281 | No | OYG | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Presage ST2 Assay, catalog numbers BC-1065 and BC-1065E. |  |  | February 28, 2012 | 2012 | Critical Diagnostics | The recall was initiated because Critical Diagnostics has confirmed that it is possible to obtain inaccurate test results for specimens when using the Presage ST2 Assay kits. | Component design/selection | Critical Diagnostics sent a Medical Device Recall letter dated March 5, 2012, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer.Customers were instructed to please use Federal Express (FedEx) to return the kits, and charge to account: 305599069 Replacement kits will be sent to you.Customers with questions about this notice were instructed to contact Critical Diagnostics' customer support at (877) 700-1250 ext 3. |  | 33 | 33 units | USA including the state of California and the countries of Austria, Spain, United Kingdom, and Italy. | Terminated | July 16, 2012 | 139 | 33
Z-2008-2012 | 62457 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom surgical kits and trays: |  |  | June 16, 2012 | 2012 | DeRoyal Industries Inc | Surgical kits and trays contain Stryker Hytrel Togas which were recalled because the clear tape, that is intended to aid in the prevention of patient's fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga, was not applied during the manufacturing process. | Nonconforming Material/Component | DeRoyal initiated the recall on 6/16/2012 by notifying customers by overnight delivery of letters describing the problem, instructing the customer how to handle inventory on hand, and requesting the customer to destroy the product, fill out the provided notice of destruction form, and return the form to DeRoyal.For questions customers were instructed to call 865-362-1020 or 865-362-1034.For questions regarding this recall call 865-362-2334. |  | 303 | 303 units | Nationwide Distribution including SD and NE. | Terminated | November 09, 2012 | 146 | 303
Z-2011-2012 | 62024 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bipolar Forceps model number: S02-XXX, S03-XXX, P02-XXX, and P03-XXX. "Instructions For Use" (QF7.5.1A.R01 thru QF7.5.1A.R06 Care and Handling of Bipolar Forceps. |  |  | May 23, 2012 | 2012 | Stingray Surgical Products, LLC | On 05/23/2012, Stingray Surgical Products, LLC., Boca Raton, FL initiated a recall initiated due to "Instructions For Use" of Bipolar Forceps not being cleared by FDA and incorrect "Instructions For Use" . The sterilization instructions have been altered without the clearance of the FDA. | Packaging | Stingray Surgical Product sent an URGENT Instructions for Use Recall letter dated Mary 23, 2012, to all distributors. The letter identified the product, the problem, and the action to be taken by the client. Distributers were requested to convey the recall to their users along with fowarding copies of the firm's recall notification letter. IFU's will be destroyed or discarded by either user or distributor since returns will not be necessary. Distributors were asked to complete the form at the bottom of the letter and return via e-mail or fax. For questions clients were instructed to call 561-210-7582.For questions regarding this recall call 561-210-7582. |  | 6326 | 6326 | Worldwide Distribution - USA (nationwide) | Terminated | August 06, 2014 | 805 | 6326
Z-2010-2012 | 61763 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 |  |  | June 21, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received reports of several issues that manifest themselves as communications or power failures in the system network for Terumo Advanced Perfusion System 1. Issues that affect a single pump or module are:" The appearance of a red X on a single pump or module icon on the CCM " The appearance or a single '?' symbol on a pump or module icon on the CCM" The fai | Nonconforming Material/Component | Terumo sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 21, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. CUSTOMER INSTRUCTIONS1. Review this Medical Device Recall notice.2. Assure that all users are aware of this notice.3. Confirm receipt of this communication by faxing, or emailing the attached Customer Response Form to the fax number/ email address indicated on the form.Terumo CVS Customer Service-1-800-521-2818Fax-1-800-292-6551.Customer Service Hours:Monday - Friday, 8 AM - 6 PM EST |  | 648 | 648 units | Worldwide Distribution-USA (nationwide) including the states of , AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV and the countries of ARGENTINA, AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | June 10, 2013 | 354 | 1640
Z-2014-2012 | 61463 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | EtacÂż Ono Walker, EtacÂż Avant Walker, and EtacÂż Salsa Walker. |  |  | April 06, 2012 | 2012 | Etac Supply Center Ab | The recall has been initiated because there is the potential that users may fall and become injured, with the potential for serious injury, including bone fracture. One death occurred in a user fall incident, but it is unclear whether the fall was caused by the loss of a wheel. Among the walkers distributed worldwide, EtacÂż has received a total of 49 reports from Europe and Japan of wheels falling | Device Design | Etac sent an "URGENT DEVICE RECALL " letter dated March 20, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately discontinue selling and distributing the affected product. A Reply Postcard was enclosed for customers to complete and return. Contact the firm at 847-277-7705 for questions regarding this recall. |  | 3187643123 | 204,323 units of the Etac Ono Walker, 15,601 units of the Etac Avant Walker, and 6,771 units of the ETAC Salsa Walker | Worldwide Distribution-USA (nationwide) and the countries of United Arab Emirates, Austria, Australia, Belgium, Canada, Switzerland, China, Colombia, Germany, Denmark, Estonia, Spain, Finland, France, Great Brittain, Greenland, Greece, Hungaria, Ireland, Israel, Iceland, Italy, Japan, South Korea, Kuwait, Lithuania, Luxembourg, Latvia, Malaysia, Netherlands, Norway, Poland, Portugal, Rumania, Russia, Saudi Arabia, Sweden, Slovenia, Slovakia, and Thailand. | Terminated | June 18, 2013 | 438 | 3187643123
Z-2013-2012 | 62329 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott Molecular Bar Code Scanner User's Guide; instructions for setting up the Bar Code Scanner LN 6L89-01, a component of the Abbott m2000 Instrument System; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list number 06L88-02. |  |  | January 31, 2012 | 2012 | Abbott Molecular | In the Abbott Molecular Bar Code Scanner User's Guide, List Number 6L88-02/09, the section Modifying the Keyboard Format for Bar Code Scanner LN 6L89-01 contains incorrect bar codes that can prevent the correct configuration of the bar code scanner language format.The Abbott m2000 System is intended for use in performing nucleic acid testing in clinical laboratories. It is comprised of the Ab | Error in labeling | Abbott Molecular Inc. sent an Urgent Field Safety Notice letter dated January 31, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to replace the affected pages of the Bar Code Scanner User's Guide with the replacement pages provided with the letter, and immediately perform the Modifying the Keyboard Format section with these replacement pages. The accounts were requested to complete the enclosed reply form and fax it to (847) 775-6728 to acknowledge receipt, understanding and implementation of the information provided in the notice. |  | 100 | 100 units | California, Florida, Maryland, Oregon and Virginia, and internationally to Belgium, Brazil, China, France, Germany, Hong Kong, Japan, Malaysia, Mexico, Qatar, Switzerland and United Kingdom | Terminated | December 04, 2012 | 308 | 100
Z-2015-2012 | 62123 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Invasive Blood Pressure Care Cables. |  |  | June 12, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware through complaints of a potential safety issue with its Invasive Blood Pressure Cable made for use with a Spectramed / BD invasive blood pressure transducer. Due to use of an incorrect O-ring, the mating of the cable with the transducer is faulty. This situation inhibits blood pressure measurement from being monitored correctly. | Nonconforming Material/Component | GE Healthcare sent a Urgent Medical Device Correction letter dated June 12, 2012, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Inforamtion. Advised customers to discontinue use, further distribution and to contact their customers. Please return any defective Invasive Blood Pressure Cables by isolating all affected product and completing the attached Confirmation form/fax back per its instructions. This step is required to confirm receipt of communications with all customers.If you have any questions or concerns regarding this notification, please contact Customer Service at +1-800-932-0760 (Domestic &amp; International). Hours of Operation: 8:00 am EST to 6:00 pm EST. . |  | 279 | 279 | Worldwide Distribution -- USA (nationwide) including the states of AK, CA, IL, MS, NJ, PR, SD, TN, TX, and VA. and the countries of ARGENTIAN, AUSTRALIA, BELGIUM, CHINA, GERMANY, ALGERIA, ECUADOR, EGYPT, FINLAND, UNITED KINGDOM, JORDAN, JAPAN, MOROCCO, MAURITUS, NETHERLANDS, OMAN, PERU, SEBIA, SUDI ARABIA, SINGAPORE, EL SALVADOR, TURKEY, URUGUAY and SOUTH AFRICA. | Terminated | December 05, 2013 | 541 | 279
Z-2006-2012 | 62349 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion EnVe Ventilator |  |  | June 19, 2012 | 2012 | CareFusion 203, Inc. | CareFusion identified that the EnVe Ventilator may not hold the set Positive End Expiratory Pressure (PEEP) value either intermittently or continuously or it may alarm for a Patient Circuit Fault. The unit will activate either a Low PEEP or Patient Circuit Fault Alarm both audibly and Visually to alert the healthcare professional. Ventilation delivery to the patient may be compromised. | Device Design | CareFusion sent an Urgent Product Recall letter with a return receipt dated June 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Required Action for Users: If the ventilator resumes normal ventilation (intermittent leak) the audible alarm stops. To clear the alarm indicator on the ventilator display, enter Alarm Messages tab, push alarm Reset to clear display. If the ventilator does not resume normal ventilation (continuous leak) provde an alternative method of ventilation to the patient. CareFusion will follow-up with a phone call to coordinate the correction of the devices. For questions customers should call 1-800-554-8933. |  | 266 | 266 | Nationwide Distribution including AZ, CA, CO, FL, IL, KY, ME, MI, MO, MN, MT, ND, NJ, NM, NV, NY, OK, PA, SC, SD, TX , UT, WI, and WV. | Terminated | February 11, 2013 | 237 | 266
Z-2016-2012 | 62051 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker T2 Arthrodesis Nailing System. Manufacturer: Stryker Trauma GmbH, Professor-Kuntscher-Str. 1-5, 24232 Schonkirchen Germany. Distributed in the USA by Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430. |  |  | January 26, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has become aware that there exists a potential for breach of the sterile barrier of packaging associated with certain lots of T2 Knee Arthrodesis Nails. | Packaging | Stryker Orthopaedics sent Recall letters/Customer Response forms dated 1/26/2012 and 1/30/2012 on 1/27/2012 and 1/31/2012 via FEDEX to their branches, hospitals and surgeon accounts respectively. The letter identified the product, the problem, and the aciton to be taken by the customer. Customers were instructed to contact their Stryker representative to arrange for return of any affected product. Customers were asked to fax back the attached Product Recall Acknowledgment Form as soon as possible to 201-831-6069..For any questions customers should call 201-972-2100.For questions regarding this recall 201-831-5970. |  | 532 | 532 units. | Nationwide. | Terminated | July 17, 2013 | 538 | 532
Z-2019-2012 | 62437 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Xprezzon Bedside Monitor, Model 91393, installations with Option C. It is a patient monitor with arrhythmia detection or alarms. |  |  | June 12, 2012 | 2012 | Spacelabs Healthcare, Llc | There were reports that a certain combination of key presses will cause the module to disconnect from the Xprezzon bedside monitor, model 91393, and the patient parameters will drop from the display. The parameters return after power cycling the monitor. | Software design | Spacelabs Healthcare sent an Urgent Medical Device Correction letter dated June 29 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to advise their staff of the potential failure of the XPREZZON Bedside Monitor. Spacelabs will contact customers to schedule a convenient time for their hospital to upgrade their affected monitors, at no cost. On July 6 2012, a customer letter (translated as necessary) URGENT - MEDICAL DEVICE CORRECTION Spacelabs Healthcare XPREZZON Bedside Monitor will be emailed to all international subsidiaries and distributors of record. Customers with questions can contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support.For questions regarding this recall call 800-522-7025, ext 2. |  | 21 | 21 units | Worldwide Distribution - USA including Florida, Iowa, Oklahoma, South Caroline and the countries of India, Germany and Panama. | Terminated | November 28, 2012 | 169 | 21
Z-2017-2012 | 62451 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MicroCool Breathable Impervious Surgical Gown, (Large and X-Large), Sterile, Made in Honduras. |  |  | September 17, 2009 | 2009 | Kimberly-Clark Corporation | Complaints were received of gowns pilling or balling, linting and abrades, blood strikethrough and fibers falling off the gowns. | Process control | Kimberly Clark sent an Urgent Medical Device Recall letter dated September 17, 2009. to all affected consignees. The Recall notification of the expanded recall was sent to consignees on November 17, 2009. Customers were instructed to destroy all affected product remaining in their inventory. For questions regarding this recall call 770-587-7400. |  | 79941 | 79,941 units | Nationwide Distribution including AL, FL, GA, KY, LA, MS, NC, OH, SC, TN, and VA | Terminated | July 19, 2012 | 1036 | 79941
Z-1992-2012 | 62115 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Infusomat(R) Space Pump (US Version) with software version 686G030103 (commonly termed 'G03'), 686E030003, 686F030005, 686F030007, 686G030002, and 686G030102. |  |  | March 23, 2012 | 2012 | B. Braun Medical, Inc. | Firm became aware of the potential for breakage of the anti free flow clip catch, located on the inside of the pump door, when the IV set anti free flow clip is incorrectly inserted into the pump and the pump door forced closed. If the clip catch is broken and the door opened, free flow protection is still ensured. However, after the set is manually removed from the pump without the closure of t | Device Design | B. Braun Medical, Inc. sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated March 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to identify the affected product in their inventory and to complete and return the Product Acknowledgement Form via fax to (610)849-1197. Contact the Customer Support Department at (800) 626-7867 for assistance with returning any product or if you have questions regarding this notice. |  | 38225 | 38,225 pumps | Nationwide Distribution-including Puerto Rico and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, and WI. | Terminated | November 03, 2014 | 955 | 38225
Z-2022-2012 | 62337 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | LMA Esophageal Intubation Detector (EID), Catalog Number EID100 |  |  | May 23, 2012 | 2012 | LMA North America Inc | The recall was initiated because LMANA has received a complaint that the Esophageal Intubation Detector (EID100) has a potential to create a false positive if used on a patient during confirmation of intubation. | Process control | LMA North America Inc. sent a recall letter dated May 23, 2012, for US customers, and May 24, 2012, for International customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to fill out the bottom portion of the recall letter and promptly fax it back to LMA North America at (858) 622-4130. If customers have units of EID100 from the affected lots remaining in stock, customers were instructed to arrange for immediate return by contacting LMANA's customer service department at (800) 788-7999, menu selection #2. For further questions please contact Customer Service Representative at (800) 788-7999. |  | 18799 | 18,799 units | Worldwide Distribution--USA (nationwide) and the countries of Canada, United Kingdom, ZA, and Australia. | Terminated | December 26, 2012 | 217 | 18799
Z-2023-2012 | 62020 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | EkoSonic MACH4 Endovascular Device (Catalog #500-55106, 500-55112; 500-55118; 500-55124; 500-55130; 500-55140; 500-55150; 500-54106; 500-56130; 500-56140; 500-56150) is a catheter, continuous infusion. it consists of the Intelligent Drug Delivery Catheter (IDDC) and the MicroSonic Device (MSD). |  |  | May 24, 2012 | 2012 | EKOS Corporation | EKOS Corporation have recently received several complaints from users where the radio-opaque marker bands on the Intelligent Drug Delivery Catheter (IDDC) were dislodged during a procedure and report of MicroSonic Device (MSD) breaking during insertion or operation. | Device Design | Starting May 25, 2012 to May 31, 2012, EKOS Corporation sent "New warnings in EkoSonic Endovascular Device Instruction for Use" letter, dated May 24, 2012, to all affected customers.Attached with the notification letter is the revised EKOS MACH4 Endovascular Device Instructions for Use in Peripheral Vasculature. Customers are informed of the complaints about the radio-opaque marker bands on the Intelligent Drug Delivery Catheter (IDDC) were dislodged during a procedure and reports of the MicroSonic Device (MSD) breaking during insertion or operation.Customers with questions about the revised Instruction for Use, please contact EKOS Corporation at 425-415-3132. |  | 21892 | 21,892 units | Worldwide Distribution - USA including OH, KY, and TN and the countries of Switzerland, Germany, Austria, Czech Republic, Slovakia Republic, Poland, Netherlands, Belgium, Luxembourg, United Kingdom, Republic of Ireland, Slovenia, Costa Rica, Argentina, Brazil, Dominican Republic, Taiwan, Korea, Greece, Mexico, Italy, Uruguay, Paraquay, Yugoslavia, New Zealand, India, Venezuela, South America, Norway, Sweden, Denmark, Turkey, Great Britain, Israel, Spain, Portugal, Greece, Korea, China, and Japan. | Terminated | November 20, 2012 | 180 | 21892
Z-2021-2012 | 61753 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base, 220/240V. |  |  | June 21, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | TerumoÂż Advanced Perfusion System 1 has displayed a Red X error message on multiple module icons indicating a loss or fluctuation of internal power on one side of the system's base. Terumo Cardiovascular Systems' initial investigation has determined that there may be more than one cause for the malfunction and is investigating further to determine the root cause.Potential risks with the loss or | Process control | Terumo sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 21, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to place the addendum in the Operator's Manual for Terumo System 1, at the beginning of the Central Control Module Section. Additionally, a Customer Response Form was attached for customers to complete and return to the firm. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this notice. |  | 648 | 648 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV. and the countries of ARGENTINA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | March 26, 2013 | 278 | 1640
Z-2031-2012 | 61769 | stains, microbiologic | JTS | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | BD Gram Stain Kit, containing 1- 250 mL bottle of Gram Crystal Violet, and additional reagents: Gram Iodine (Stabilized), Gram Decolorizer, Gram Safranin. |  |  | April 12, 2012 | 2012 | Becton Dickinson &amp; Co. | Potential for increase level of artifacts to appear as large, spherical and easy to distinguish or small &amp; more difficult to distinguish from gram positive cocci. | Material/Component Contamination | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated April 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard the affected product. An Acknowledgement Form was attached for customers to complete and return via fax to 410-316-4258. Contact the BD Customer Service Department for questions regarding this recall. |  | 539 | 539 kits | Nationwide Distribution. | Terminated | December 10, 2012 | 242 | 3084075
Z-2030-2012 | 61691 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMASpectrum Infusion Pump Master Drug Library (MDL) Editor software. |  |  | April 23, 2012 | 2012 | Sigma | Sigma is a safety alert for the Master Drug Library (MDL) Editor software associated with the SIGMA Spectrum Infusion Pump System because it may result in a facility drug library that does not represent common clinical practices. | Software design | Sigma sent an Urgent: Customer Safety Alert dated April 23, 2012 via UPS delivery to all consignees requesting both acknowledgement of notification and confirmation that recommended actions have been taken. Customers are asked to follow the instructions contained in the recall letter. Customers with questions may contact the firm at 800-356-3454 ext 300. |  | 2048 | 2,048 CD's | Worldwide Distribution - US (nationwide) including Puerto Rico, and Canada. | Open, Classified |  |  | 2048
Z-2028-2012 | 62464 | No | KNF | Unipolar endoscopic coagulator-cutter and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Integra(TM) Jarit(R) Take-Apart Endoscopic Instruments. |  |  | May 23, 2012 | 2012 | Integra Limited | The Integra(TM) Jarit(R) Take-Apart Laparoscopic Instrument shafts were released with internal component out of specification. | Nonconforming Material/Component | Integra sent a Urgent Field Corrective Action letter dated May 23, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed customers and sales reps to return the instruments identified on the attached part number list and complete a "Field Corrective Action Acknowledgement and Return Form" to be completed and returned to the firm.We apologize for any inconvenience this issue may cause you. Please call Integra Surgical Customer Service at 800-431-1123 if you have any questions. |  | 108 | 108 | USA (nationwide) including the states of AZ, CA, IL, MA, MN, NJ, OK, TN, TX, VA, WA and WI. | Terminated | March 19, 2013 | 300 | 108
Z-2029-2012 | 62408 | No | FST | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GMF-RC65 Replacement Surgical Headlight Cable for Cogent Microlink. |  |  | April 09, 2012 | 2012 | Gulf Fiberoptics, Inc. | The GMF-RC65 Replacement Surgical Headlight Cable for Cogent Microlink falls outside of the firm's 510k due to light transmitting fiber material difference. | Device Design | The firm, Gulf Fiberoptics, Inc., sent an "RECALL NOTIFICATION-PLEASE DO NOT DISCARD" letter dated April 09, 2012 to its customer via (traceable letter &amp; email). A notification was put on the firm's website. The letter described the product, problem and actions to be taken. The distributors were instructed to determine the status of the product; remove existing product from service immediately and return to their company in a timely manner; send all returned GMF-RC65 cables to Gulf Fiberoptics; forward copies of the recall letters to their customers, ultimately the end user, and complete and return the attached Recall Worksheet Form for each GMF-RC65 cables regardless of its status via fax to: +1 813.855.6627; email to: difigueoa@gulffiberoptics.com or mail to : Gulf Fiberoptics, 448 Commerce Blvd, Oldsmar, FL 34677. The customers were instructed to remove the cables from service immediately, if they think they may have one of the listed cables; contact customer service at +1 813.855.6618 for serial number verification and for instructions to return this cable to Gulf Fiberoptics.If you have any questions regarding this voluntary recall, please contact your Customer Service representative at: +1813.855.6618 or by email at: difigueroa@gulffiberotpics.com. |  | 169 | 169 units | Worldwide distribution: USA (nationwide) including states of: AL, CA, CO, FL, MI, NC, NJ, NY, OH, PA, TN, TX and WI and country of: South Africa. | Terminated | March 07, 2013 | 332 | 169
Z-2027-2012 | 62404 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 8800 Flexiview designed to provide fluoroscopic and spot-film imaging of the patient during diagnostic, surgical and interventional procedures. |  |  | November 19, 2007 | 2007 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC recalled imaging devices OEC 9900, OEC 9800, and OEC 8800 as a result of an FDA inspection identifying that the vertical lift column power supply in the mainframe C-arm is defective and subject to early life failure. | Device Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE Healthcare sent an"URGENT RECALL NOTICE" dated February 22, 2008 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 800-874-7378 for questions concerning this recall. |  | 1349 | 1349 units | Worldwide Distribution. | Terminated | January 28, 2013 | 1897 | 13175
Z-2034-2012 | 61147 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | epoc Blood Analysis BGEM Test Cards for use with the epoc(R) Blood Analysis System |  |  | June 28, 2011 | 2011 | Epocal Inc | Customer reported problem while performing incoming QC of newly received lot of test cards. Liquid quality control results were found to be out of range. Further investigation by Epocal determined that a portion of the test card lot (approximately 3000 cards) could potentially report low glucose results. | Device Design | Epocal Inc. sent a Voluntary Product Recall letter dated June 27, 2011 to the affected consignee. The letter identified the affected products, problem and actions to be taken. The letter ask that all products affected by this recall be sent back to Epocal and a replacement may be obtained. For questions call 1-613-738-6912. |  | 8650 | 8,650 | US nationwide Distribution in the state of CA. | Terminated | December 18, 2013 | 904 | 8650
Z-2052-2012 | 61479 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Sarns Modular Perfusion System 9000, 100V with color screen. |  |  | June 18, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received several reports of malfunctions of the Air Bubble Detection System used with its heart-lung machines over the last three years. The malfunctions were expressed in various failure modes (over sensitivity, false alarms and were found to have various root causes. The malfunctions occurred in systems that were excluded from prior corrective actions or in syst | Nonconforming Material/Component | Terumo CVS sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 30, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers are asked to complete and return a Customer Response Form via fax to 1-800-292-6551. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this recall. |  | 31 | 31 units | Worldwide Distribution-USA (nationwide) and the countries of ARGENTINA, BANGLADESH, BELGIUM, Bolivia, BRAZIL, BRISBANE, CANADA, CHILE, China, COLOMBIA, COLOMBIA, Costa Rica, DENMARK, Dominican Republic, Ecuador, Egypt, FRANCE, GERMANY, Guatemala, HOLLAND, Honduras, Hong Kong, India, INDIA, Indonesia, Israel, ITALY, Jordan, KOREA, Kuwait, Lebanon, LEICESTERSHIRE, Malaysia, Mexico, NETHERLANDS, New Zealand, Nicaragua, NORWAY, Pakistan, Panama, PERU, Philippines, Russia, Saudi Arabia, SINGAPORE , SOUTH AFRICA, South Korea, SPAIN, Sri Lanka, SWEDEN, SWITZERLAND, Taiwan, THAILAND, Turkey, UNITED ARAB EMIRATES (UAE), Uruguay, Venezuela, Vietnam, WEST GERMANY and Yemen. | Terminated | April 12, 2013 | 298 | 4689
Z-2039-2012 | 62293 | No | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Alere Triage Tox+MTD Drug Screen, PN 94400 |  |  | May 14, 2012 | 2012 | Alere San Diego, Inc. | A recall was initiated because Alere San Diego has confirmed that the Triage TOX Drug Screen may have significantly decreased precision and accuracy than claimed. An increased frequency of false positive and/or false negative results has been observed for Triage TOX Drug Screen. | Nonconforming Material/Component | Alere sent an Urgent Medical Device Recall letter dated June 11, 2012 to all affected customers. The letter informed the customers of the affected product, problem identified and actions to be taken. Customers were instructed to discontinue all use of the recalled lots, immediately inform clinicians of this recall at the testing site(s) and complete and fax the enclosed customer verification form and list of affected lots within 10 days to confirm receipt of the letter. Customers with questions about the information contained in the recall letter, can contact Alere San Diego at 9975 Summers Ridge Road, San Diego, CA 92121 or call (877) 308-8287. |  | 9997 | 9,997 kits | US Nationwide Distribution | Terminated | December 04, 2012 | 204 | 9997
Z-2036-2012 | 62284 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCI Concha Therm Neptune Humidifier, Rx Only, |  |  | June 12, 2012 | 2012 | Teleflex Medical | Speaker Failure. If the speaker fails and an event occurs which would generate an audio alarm, the alarm will not sound and could cause a potential delay in patient treatment. | Device Design | Teleflex sent Urgent Medical Device Notification letters dated June 12, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were instructed to complete the enclosed Acknowledgement Form and return per the instructions provided. Once the form is received, Teleflex Medial will schedule for a technician to come to their facility and exchange the affected product. For questions call your local sales representative or Customer Service at 1-866-246-6990. |  | 14794 | 14,794 each | Worldwide Distribution - U.S. (nationwide) and the countries of Argentina, Australia, Brazil, Canada, Chile, China, Europe, India, Japan, Mexico, South Korea, Taiwan and UK. | Terminated | January 28, 2014 | 595 | 14794
Z-2041-2012 | 62420 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian Eclipse Treatment Planning System for radiotherapy; Model Number: H48; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | June 18, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the Eclipse Treatment Planning (for radiotherapy treatments) where changing the prescription in a calculated VMAT or RapidArc plan may lead to Monitor Units which do not reflect the new prescription. If the dose prescription of a VMAT or RapidArc plan is changed after the initial Dose Volume Optimization, the Monitor Units and dose distribution may not reflect | Software design (manufacturing process) | The firm, VARIAN Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION; URGENT FIELD SAFETY NOTICE" dated June 18, 2012, to all affected users. The notice described the product, problem and actions to be taken. Varian is developing a technical correction for this issue. The customers will be contacted by a Varian Customer Service representative when this correction is available to schedule its installation.If you require further clarification, please feel free to contact your local Varian Medical Systems Customer Support District or Regional Manager.In USA and Canada: 1.888.VARIAN5 (888.827.4265)Europe: +41 41 749 8844; email: support-emea@varian.comEmail: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: support-sea@varian.com; China/Asia: support-china@varian.com; Japan: support-japan@varian.com and Latin America: soporte.al@varian.com. Internet: Oncology systems customer site - www.myvarian.com and Varian Medical Systems public site - www.varian.com. |  | 7113 | 7113 units | Worldwide distribution: USA (nationwide) and countries including of: Austria, Australia, Bahrain, Belgium, Brazil, Canada, China, Chile, Colombia, Denmark, Egypt, Finland, France,Germany, Greece, Hungary, Hong Kong, India, Indonesia, Israel, Ireland, Italy, Japan, Latvia, Malaysia, Mexico, Morocco, Nepal, Netherlands, Norway, Panama, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, United Kingdom, Taiwan, Turkey, Ukraine, United Arab Emirates and Venezuela. | Terminated | August 06, 2014 | 779 | 7113
Z-2035-2012 | 62498 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Brilliance 6, 16, 40, 64, Big Bore, iCT and iCT SP are Computed Tomography X-Ray Systems. |  |  | July 02, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare was notified that several of the sixteen bolts securing the brackets which support the carbon table top on various models of their Brilliant CT scanner have been found to be defective and the bolt heads may break off. | Process control | Philips sent a Urgent Medical Device Correction letter dated June 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will visit each site to inspect and correct the issue, and Philips will be tracking the progressof the correction. For further information, Customers were asked to contact their local Philips representative or Philips Healthcare Office. North American and Canadian consignees should contact the Philips Customer Care Solutions Center at 1-800-722-9377 |  | 524 | 524 | Worldwide Distribution -- USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, MJ, NM, NV, NY, OH, OK, OR, PA, PR, TN, TX, VA, VT, and WI. and the countries of Canada, Argentina, Australia, Austria, Azerbaijan, Brazil, Chile, China, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, France, Germany, Greece, Iceland, India, Indonesia, Iraq, Italy, Jordan, Kazakhstan, Kazakhstan, Libyan Arab Jamahiriya, Malaysia, Mexico, Mozambique, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Palestine, Peru, Philipines, Poland, Portugal, Russia, Russia Fed, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Syrian Arab Republic, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Yemen. | Terminated | December 24, 2013 | 540 | 524
Z-2033-2012 | 61723 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker AVS Navigator Trials 6mm-15mm |  |  | January 13, 2012 | 2012 | Stryker Spine | Beginning in February 2011, Stryker Spine began receiving reports related to the shaft of the AVS Navigator trial failing at the junction of the trial head. | Device Design | Stryker sent an "URGENT PRODUCT RECALL" letter to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Customer Response Form was included for customers to complete and return via fax to Regulatory Compliance at 201-760-8370. Contact the firm at 201-760-8298 for questions regarding this recall. |  | 273 | 273 units | Nationwide. | Terminated | July 29, 2013 | 563 | 273
Z-2037-2012 | 62372 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi Scenaria CT System Software |  |  | May 04, 2012 | 2012 | Hitachi Medical Systems America Inc | Hitachi discovered a software error in the Reconstruction Status function that creates a potential risk that a patient study will not be fully reconstructed. When multiple reconstructions are queued for processing, if a queue is suspended and the task order changed, the second queue may not reconstruct the full range of slices. This may result in missing diagnostic data. | Software design | Hitachi sent a "Device Correction Notice" dated May 4, 2012 to their customers. The notice identified the affected product, problem and actions to be taken. Customers were instructed to "not use the Reconstruction Status" function to suspend active reconstructions until the software problem can be corrected. They also informed their customers that a representative will visit their facility to install the software update to permanently resolve the issue at no cost to the facility. For further information regarding this notice call 800-800-3106 x3720. |  | 6 | Six (6) units | US Nationwide Distribution to the following states of: MD, MI, MX, ND, NY, OH and TX. | Terminated | February 19, 2013 | 291 | 6
Z-1993-2012 | 62527 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima Mobile X-ray System. The GE Automatic Mobile X-Ray |  |  | April 04, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue on the interface of the mobile radiographic product, Optima XR200amx, ond Optima XR220amx related to Visual indication as mandated by the US Code of Federal Regulations (21CFR) and IEC 60601-1-3. | Radiation Control for Health and Safety Act | GE Healthcare will, without charge, remedy the defect or bring the product into compliance with each applicable Federal and IEC standard in accordance with a plan to be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication tothe user or through a GE Healthcare field engineer site visit. A GE Healthcare Service Representative will update the firmware on the system to address the issue.Your proposed corrective action plan (CAP) (Field Modification Instructions 10865) describes the rework plan that GEHC Engineers will make to the imaging system to return it to compliance with the applicable performance standard. It appears to adequately address the problem and is hereby approved. |  | 254 | 254 units have been shipped to/installed in the U.S. | Worldwide Distribution (USA) nationwide | Terminated | January 10, 2013 | 281 | 254
Z-2055-2012 | 61242 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base 100/120V |  |  | June 22, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received eight complaints since September 2008 of drive motor failure of the Sarns Centrifugal System of the Terumo Advanced Perfusion System 1. A failure of the drive motor for the Terumo System 1 centrifugal system triggers error messages and alarms; therefore it is easily recognizable by the user.The potential hazards are dependent on the application | Equipment maintenance | TERUMO sent an Urgent Medical Device Recall letter dated June 22, 2012, and an update to the Operator's Manual via Federal Express to inform consignees of the reported malfunction and potential hazards and to raise awareness of the benefits of using a retrograde valve and the importance of backup system availability. Consignees were instructed to review the Medical Device Correction notice, assure that all users are aware of the notice. Place the Addendum at the beginning of the Centrifugal System Section of the Operator's Manual, Confirm receipt of the communication by faxing or emailing a completed Customer Response Form as indicated on the form. For any questions customers should call 1-800-521-2818. |  | 992 | 992 devices | Worldwide Distribution - USA (nationwide) and the countries of Nationwide and worldwide: ARGENTINA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatamala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | March 06, 2013 | 257 | 1632
Z-2057-2012 | 62351 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | OrthoPediatrics PediLoc Locking Plate System, 4.5mm Contour Locking Compression Femur 8 Hole Plate Right. |  |  | June 14, 2012 | 2012 | OrthoPediatrics Corp | The device is a Left 8 hole plate but is incorrectly labeled and etched as a Right device. | Error in labeling | On 15-June-2012, a follow-up, URGENT Recall of MIS-BRANDED PediLoc Locking Plate e-mail was sent to all Distributors and direct representatives to provide them with a written notification of the voluntary recall. The e-mail identified the affected product and the reason for the recall. It also instructed them to quarantine all identified devices from the affected lot number immediately. A certified Voluntary Recall notice was also sent as a follow up along with a return prepaid pouch for the return of the recalled device(s) and Mandatory Reply Form. |  | 15 | 15 devices | Worldwide Distribution -- USA, including states of IN, GA, MI, FL, MO, OH, and SC and country of Australia. | Terminated | April 30, 2013 | 320 | 15
Z-2054-2012 | 62324 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Universal Spinal System Lamina Hooks |  |  | May 31, 2012 | 2012 | Synthes (USA) Products LLC | USS Lamina Hooks, Medium, Right are being recalled because the hooks are etched with the wrong part number - 498.321 (Lamina Hook, Medium, Left). The Lamina Hooks should be etched with part number 498.320 (Lamina Hook, Medium, right). | Employee error | SYNTHES sent a Medical Device Recall letter dated May 31, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the consignee. The letter requested consignees and sales reps remove recalled product from their inventory, call the firm at 1-800-479-6329 to receive an authorization number and complete the enclosed verification form and return to:Credit/Returns, Synthes1101 Synthes AvenueMonument, Colorado 80132.If customers do not have any of the affected product they were instructed to complete the Verification Section at the end of the letter by checking the appropirate box indicating that no affected porudct has been located. Return the document to Synthes by fax to 610-251-9005 or Scan/Email: FieldAction@synthes.com.For any questions please call 610-719-5450. |  | 12 | 12 units | Nationwide Distribution including GA, PA, SC, TX, and UT and Canada. | Terminated | August 31, 2015 | 1187 | 12
Z-2053-2012 | 62114 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Leica M525 F50 Surgical Microscope; Article No. 10338422; Leica Microsystems (Schweiz) Ag, 201 Max Schmidheiny Strasse, Heerbrugg, Switzerland |  |  | June 04, 2012 | 2012 | Leica Microsystems, Inc. | Leica Microsystems evaluated data showing that a malfunction resulting in an illumination reduction can be observed when Leica M525 F50 surgical microscope is used in combination with a remote control connected to a remote video adapter or motorized dual video adapter. | Software design | Leica Microsystems sent an Urgent Medical Device letter dated June 4, 2012 to the affected customers. The letter described the product, problem and actions to be taken. The customers were informed that the Leica surgical microscope M525 F50 required a software upgrade. Customers were instructed to NOT USE a remote video adapter or motorized video adapter until the required software upgrade has been installed by their Leica Microsystems representative. The customers were requested to complete the attached Acknowledgement Form, acknowledging receipt and understanding of the letter, listing the serial numbers of the units at their facilities and to fax it back to Leica Microsystems at (847) 236-3747. |  | 43 | 39 Units - USA , 4 Units - Canada | Worldwide Distribution - US (nationwide) in the states of: California, Colorado, Florida, Georgia, Illinois, Indiana, Massachusetts, Michigan, Minnesota, New York, Oklahoma, Oregon, Pennsylvania, Texas, Utah, West Virginia, including Puerto Rico, and the countriy of Canada. | Terminated | October 29, 2012 | 147 | 43
Z-2058-2012 | 62520 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Oto Ease by Westone, labeled as a bacteria free, greaseless lubricant. Sold in 0.5 oz consumer sized semi transparent flexible plastic bottles. |  |  | July 13, 2012 | 2012 | Westone Laboratories, Inc. | Westone is recalling all Oto Ease ear lubricant in all packaging configurations and sizes due to potential microbial contamination of the product. | Material/Component Contamination | Westone Laboratories sent an Urgent Recall letter and Press Release to distributors and customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine their inventory, quarantine product subject to recall and discard or return the product to Westone. In addition, if the affected product have been further distributed please identify your customers and notify them at once of this product recall. Consumers with questions may contact Westone Laboratories at 1-800-357-3240 between the hours of 8:00am and 5:00pm MST or email at otoeaserecall@westone.com. |  | 115 | Approx 100k sold annually x 15 years. | Worlwide Distribution - US (nationwide) and to countries of: Australia, Belgium, Brazil, Canada, France, Germany, Greece, Hong Kong, Ireland, Japan, Korea, Mexico, Panama, Philippines, Russian Federation, and Singapore, Sweden, Switzerland and Thailand and United Kingdom. | Terminated | December 14, 2012 | 154 | 115
Z-2068-2012 | 61910 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System Software |  |  | May 21, 2012 | 2012 | Ortho-Clinical Diagnostics | software automatically switches the lot when a vancomycin (VANC) or valproic acid (VALP) reagent is depleted or the operator manually changes the lot in use causing lack of verification of proper quality assessment. | Software design | Ortho Clinical Diagnostics notified its foreign direct account consignees via email on May 21, 2012 and their direct account US consignees were sent an Urgent Distributor Safety Notice on May 22, 2012 via US Postal Service Priority Letter. The recall letter notified the consignees of the situation and that Ortho will replace their remaining inventory of non-expired GENs 14 or 15 for VITROS VALP Reagent only. Until then the consignees can continue using the VITROS VALP Reagent, GENs 14 or 15 until the replacement arrives. And to note that VITROS VANC Reagent, GENs 19 &amp;20 are now expired. --- The letter also requested the consignees to complete and return the enclosed Confirmation of Receipt, no later than May 31, 2012. And to review the quality control beginning with the Calibration date of the VITROS VANC Reagent ( GEN 21 or above) or VITROS VALP (GEN &amp; above) so that the QC results can be verified to be within acceptable ranges. --- For customers using VITROS VALP Reagent GEN 14 or 15, they were to respond as to the quantity of the non-expired GENs 14 and 15 remaining inventory; to load and use only one GEN of VITROS VALP Reagent at a time on the system until the replacement arrives. This notification letter should be posted by each VITROS System using the VITROS VALP Reagent or VITROS VANC Reagent or with the user documentation. And in addition, if these products were distributed then this information should be forwarded as well.For questions regarding this recall call 585-453-4110. |  | 1387 | 1387 | Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Brazil, Canada, Chile, Argentina, Paraquay, Ecuador, Colombia, India, Japan, Mexico, Panama, Singapore, Venezuela, UK, Northern Europe, France, Germany, Italy, Spain and Portugal. | Open, Classified |  |  | 3508
Z-2059-2012 | 62307 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | 4.5mm TI Multiloc Screw Length 38MM-Sterile |  |  | May 22, 2012 | 2012 | Synthes USA HQ, Inc. | Two lots of 4.5mm multi-loc screws were recalled due to incorrect labeling of screw size. | Error in labeling | Synthes sent a "NOTICE: MEDICAL DEVICE RECALL" letter dated May 22, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Verification Form was enclosed for customers to complete and return via fax to 610-251-9005. Contact the firm at 610-719-5450 for questions regarding this recall. |  | 20 | 20 | Nationwide Distribution-including the states of AZ, CA, FL, GA, IL, LA, MD, MN, NJ, NY, OH, PA, TN, TX, and VA. | Terminated | September 10, 2015 | 1206 | 20
Z-2024-2012 | 62393 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics Trilogy 100, 200, and 202 Ventilators |  |  | April 27, 2012 | 2012 | Respironics, Inc. | The specified Trilogy units are being recalled because of the potential for a faulty component in the power supply. | Component design/selection | Respironics, Inc. began contacting its affected US customers by phone beginning on Friday, April 27, 2012, and requested that they quarantine any affected devices within their possession and that they retrieve any devices with patients and quarantine those as well. Replacement devices were to be provided to all affected customers. For questions customers should call 724-387-7651.For questions regarding this recall call 877-387-3311.. |  | 173 | 173 | Worldwide Distribution - USA including AZ, CA, IL, IN, LA, MI, MN, MO, NJ, NY, NC, PA, SC, TN, TX, and WI and the countries of Canada, Denmark, Estonia, Finland, France, Great Britian, India, Italy, Japan, Korea, Saudi Arabia, and Spain. | Terminated | November 01, 2012 | 188 | 173
Z-2073-2012 | 62519 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE or ONCOR Digital Linear Accelerators with syngo RT Therapist version 4.2/MOSAIQ OIS Systems and Multiple-X, Manufactured by SIEMENS AG. The ARTISTE and ONCOR family of Linear Accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. |  |  | June 25, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Potential safety issue when using the IMMEDIATE RESUMPTION feature after a Control Console interlock during the delivery of flattening-filter-free (Multiple-X) beams or segments. In a very rare scenario, the RT Therapist will send again the previously completed beam or segment for delivery to the control console, thus repeating the already delivered Monitor Units of the same beam or segment. | Software design | The firm issued "Urgent: Medical Device Correction" letters delivered by the Siemens Customer Service Group or by certified mail to all its direct consignees, informing them of the affected product and providing instructions on the recall. The firm stated that release of this UI TH018/12/S to distribute the notifications began on June 25, 2012. |  | 18 | 18 | Worldwide Distribution, including the state of Florida (USA) and the countries of Germany, India, Poland, Thailand, and the Republic of Korea. | Terminated | October 18, 2012 | 115 | 18
Z-2070-2012 | 61969 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Product is labeled in part: "***OPTETRAK***TIBIAL TRAY***TRAPEZOID CEMENTED***Ti Alloy*** SIZE 1F/1T, 2F/1T***FOR CEMENTED USE ONLY***Exactech, Gainesville, FL 32653-1630.***REF 204-04-21***". |  |  | September 23, 2011 | 2011 | Exactech, Inc. | Exactech, Inc. Gainesville, FL is recalling their Optetrak Cemented Trapezoid Tibial Tray, Size 1F/1T, 2F/1T and the Optetrak Rotating Bearing Knee Tibial Tray, Finned, Cemented Tibial Tray, Size 2F/1T. Both trays have the potential to be mislabeled. | Labeling mix-ups | Exactech, Inc. sent an "IMPORTANT MARKET WITHDRAWAL NOTICE" on September 23, 2011, to all affected customers. The customers were instructed to: 1. Immediately cease distribution or use of the products. 2. Extend the information to accounts that has possession of the products. 3. Verify if they have any of the affected products listed. 4. Fax back the attached form. Customers were instructed to contact Exactech inventory representative to confirm quantities at their location. For questions regarding any inventory restocking, the customers were directed to contact Kaya Davis 1-800-392-2832. |  | 16 | 16 | Worldwide Distribution-USA-including the states of CA, MA, OH, and OK, and the countries of Brazil, Colombia, France, Italy, Japan, Luxembourg, and Spain. | Terminated | June 11, 2014 | 992 | 28
Z-2069-2012 | 62534 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Reporting |  |  | July 10, 2012 | 2012 | AGFA Corp. | When users selected "Left stenosis" in the "Graft Duplex Conclusion" section in IMPAX CV Reporting Non-Invasive(NIV) module, "Right stenosis" appeared in the sentence display within the printed representation of the report. | Error in labeling | AGFA Healthcare sent an Urgent Field Safety Notice dated July 10, 2012, to all affected consignees indicating that a representative from the IMPAX CV product management or professional services team would be contacting them to perform a review of their system. For questions customers were asked to call 401-604-2180.For questions regarding this recall call 864-421-1754. |  | 6 | 6 | Nationwide Distribution including IN, SC, OH, TX | Terminated | October 29, 2013 | 476 | 6
Z-2074-2012 | 62450 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BridgePoint Medical, Stingray Catheter, REF M - 1000, Sterilized with Ethylene Oxide Gas, Prescription Only, Manufacturer BridgePoint Medical Inc. 13355 10th Ave N, Suite 110, Plymouth, MN 55441. |  |  | June 26, 2012 | 2012 | Bridgepoint Medical | BridgePoint Medical Inc. is voluntarily recalling Stingray Catheter, Model M1000 lots BP20121440078 and BP20121560089. BridgePoint Medical, Inc. recently determined the potential of compromised sterility assurance of the product due to a faulty pouch seal. In no case was there any reported adverse clinical event related to the problem. | Process control | The firm, BridgePoint Medical, sent an "Urgent Field Safety Notice" letter dated June 26, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory, quarantine the affected lots and complete and return the enclosed "Verification Form" via fax to 763-225-8718 ,attn: BridgePoint Medical, Inc., Upon receipt of form, BridgePoint will supply the customers with the appropriate materials to return the products.If you have any questions regarding this notice, please call BridgePoint Medical, Vice President of Quality Assurance, at 763-225-8511 or e-mail at cschlawin@bridgepointmedical.com. |  | 95 | 95 (57 USA, 38 OUS) | Worldwide distribution: USA (nationwide) including: AL, AZ, CA, CO, GA, IL, KY, MD, MN, MO, NC, NY, OR, PA, WA, and WI; and countries of: Australia and Sweden. | Terminated | September 06, 2012 | 72 | 95
Z-2040-2012 | 62382 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 6800 Miniview, Image-intensified fluoroscopic x-ray system used during diagnostic, surgical and interventional procedures. |  |  | November 20, 2006 | 2006 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The OEC 6800 Miniview was recalled in November 2006 due to several intermittent potential safety issues including navigation tracking inaccuracy, system Lock-up, incorrect image display, incorrect dose area calculated air kerma in "film mo | Device Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. On 11/20/2006, two written notifications were sent to 1513 identified consignees, addressed to the Facility Administrator and the Radiology Administrator for a combined total of 3026 letters. The Urgent Recall Notice, dated 11/8/2006, advised customers that GE Healthcare had identified several intermittent potential safety issues that may occur with the identified affected product. The notification also provided interim or permanent solutions for each of the problems described. For the issues that did not have permanent solutions at the time, the firm was to contact customers and remedy the issues free of charge. For the issues that did have a permanent fix at the time, the GEHC OEC service team was to contact customers to install the correction. If customers had questions or concerns, they were to contact the service team for further information at 800-874-7378 option 8. On 04/23, 2007, a second written notification was sent via certified mail to the same 1513 consignees providing further clarification on the issue of "incorrect Image Display" and a temporary remedy to the issue. |  | 1933 | 1,933 units | Worldwide Distribution | Terminated | May 13, 2013 | 2366 | 1933
Z-2065-2012 | 62387 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 8800 Flexview, 892.1650 Image-intensified fluoroscopic x-ray system |  |  | November 20, 2006 | 2006 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The OEC 8800 Flexview was recalled in November 2006 due to several intermittent potential safety issues including navigation tracking inaccuracy, system lok-up, incorrect image display, incorrect dose area calculated air kerma in "film mod | Device Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE Healthcare sent two Urgent Recall Notice letters to the Facility Administrator and the Radiology Administrator. The letters identified the affected product, problem and actions to be taken. The letters advised consignees that GE Healthcare is actively working on solutions that will permanently resolve each of the issues. For questions or concerns contact GE Healthcare service team at 800-874-7378 option 8. |  | 1257 | 1257 units (472 domestic systems) | Worldwide Distribution - US Nationwide | Terminated | May 13, 2013 | 2366 | 1257
Z-2038-2012 | 51222 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9800; OEC FluoroTrak 9800 Plus; OEC 9800 Plus; OEC 9800 MD Motorized C-arm System, 892.1650 Image-intensified fluoroscopic x-ray system. |  |  | February 22, 2008 | 2008 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The vertical lift column power supply in the mainframe C-arm of the device is defective and subject to early life failure. | Software design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The firm, GE Healthcare, Surgery, sent an "URGENT RECALL NOTICE" dated November 8, 2006 via certified mail with return receipt request to consignees/customers with specific instructions. Customers were advised "to stop using the system WITH NAVIGATION until this replacement has been successfully completed. OEC@ FluoroTrak 9800 Plus system may be used as a standard C-Arm providing you review the other issues as listed below." GEHC OEC is actively working on solutions that will permanently resolve each of the noted issues.If you have any questions or concerns regarding these issues, please do not hesitate to contact the service team for further information at 800-874-7378 option 8. Information is available at this number 24 hours per day, 7 days a week. |  | 8214 | 8214 units | Worldwide distribution. | Terminated | February 19, 2013 | 1824 | 8214
Z-2077-2012 | 62326 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | .***ANSPACH***Sterile. REF SHD-4820DS***QTY: 1, 4.8 mm x 2mm Twist Drill with Depth Stop. Use with SHORT-HD Attachment ***RX Only |  |  | September 19, 2011 | 2011 | The Anspach Effort, Inc. | Anspach Effort, Inc. Palm Beach Gardens, FL is recalling their Single Use Sterile Bone Cutting Burrs used with XMax, microMax, microMax Plus, eMax, eMax2 and eMax2 Plus High Speed Drill Systems. It is possible for the cutter to come apart. | Device Design | The firm, Anspach, sent an "URGENT: Medical Device Product Removal" letter dated December 15, 2011, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Screen their inventory immediately for any of the batches listed in the letter.2) If they have any of the products listed, contact Anspach Customer Support at (800) 327-6887 to arrange for their immediate return and replacement.3) Complete and return the attached Customer Reply form indicating if the item(s) are being returned and confiming their receipt of this letter, and fax to 1-800-327-6661. If they have distributed any of the products to other services or facilities, please forward this information as appropriate.Should you have any queries, please do not hesitate to contact Anspach Customer Support at (800) 327-6887. |  | 5 | 5 each | US Nationwide Distribution including the states of: Arizona. and California. | Terminated | July 25, 2013 | 675 | 5
Z-2075-2012 | 62582 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Hospira LifeShield Latex-Free 100 mL Burette Set, Convertible Pin, 77 inch with 2 Prepierced Injection Sites and Option-Lok Microdrip Soluset with Calibrated Burette and Precision Drip Chamber; an Rx sterile medical device used to administer fluids; 20 individually wrapped sets per case; Made in Costa Rica, Hospira, Inc., Lake Forest, IL 60045 USA; list 12722-65. |  |  | July 13, 2012 | 2012 | Hospira Inc. | The float valve in the burette sticks to the burette wall and does not open or close properly. | Employee error | The firm, Hospira, sent an "URGENT DEVICE RECALL" letter dated July 13, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to check their inventory and immediately quarantine any affected product; complete and return Reply Form via fax to 1-877-722-7019 or e-mail to: hospira4921@stericycle.com; return affected product to Stericycle using the label provided with the letter; and if they have distributed the product further, notify their accounts that received the product . Please contact Hospira Customer Care at 1-877-946-7747 or your Hospira representative for information regarding product availability. |  | 4420 | 4,420 sets | Worldwide Distribution-USA (nationwide) including the states of California, Colorado, Connecticut, Florida, Illinois, Kentucky, Maryland, Massachusetts, Minnesota, Oklahoma, Texas, Utah, Virginia, Washington, West Virginia and Wyoming, and the countries of Barbados and Canada. | Open, Classified |  |  | 4420
Z-2081-2012 | 54724 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | The Handi-Fil Disposable Syringe Fill Tube is a sterile plastic tube, approximately 10 inches in length, with a "J" shaped curve that projects upward approximately two inches. |  |  | February 01, 2010 | 2010 | Mallinckrodt Inc | A report was received from a Japanese distributor regarding a complaint from a hospital customer stating that ten packages of Handi Fil were not sealed. | Packaging process control | Covidien send an Urgent Device Recall letter dated February 1, 2010, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for affected product and discontinue use of the product immediately. Customers were asked to complete the attached business form and report the quantity of affected product currently in their stock. Once the form is completed customers should fax the form to 314-654-8206. Once the business notification form is received, an RGA number will be sent to the customer for return of the product. Customers will receive credit for all returned product. For questions regarding this recall call the Product Monitoring Department at 800-778-7898. For questions about credit for the product, please call Customer Service at 888-774-1414, menu option 1, then menu option 2.For questions regarding this recall call 513-948-5719. |  | 75 | 75 cases (100 units per case) | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CO, GA, IL, KS, KY, LA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, OR, PA, SC, TN, TX, VA, WA, WI, &amp; WV and countries including: AUSTRALIA, BELGIUM, BRAZIL, CANADA, JAPAN &amp; SINGAPORE. | Terminated | August 03, 2012 | 914 | 75
Z-2084-2012 | 62365 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Tibial Punch Tower Assembly. Non-Sterile. CAT NO. 8000-1089. Howmedica Osteonics Corp., 59 Route 17, Allendale, NJ 07401. |  |  | December 22, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report indicating that a Tibial Keel Punch could not be inserted into the slot of the Tibial Punch Tower. The slot that guides the stem of the keel punch into the channel of the tower was later determined to be undersized. | Nonconforming Material/Component | Stryker Orthopaedics sent an Urgent Product Recall/Recall Acknowledgement letter dated January 10, 2012, to all affected customers. The letter identified the product(s) the problem and the action needed to be taken by the customer.Customers were instructed to contact their Stryker Sales Representative to arrange for return of the affected product. Customers were asked to fax back the attached Product Recall Acknowledgement Form as soon as possible to 201-831-6069. For questions customers should call 201-972-2100.For questions regarding this recall call 201-831-5970. |  | 152 | 152 units | Worldwide Distribution | Terminated | June 13, 2013 | 539 | 152
Z-2078-2012 | 62462 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza. |  |  | February 24, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential malfunction when using syngo.plaza with software version VA20C_HF01. For datasets with distance measurements it can happen, that all images are not loaded to Viewer but no error message is displayed to user. | Software design | Siemens sent a Customer Safety Advisory Notice letter dated February 24, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to ensure that the prefix defined in the viewer settings for "Content Creator's Initials" is not empty. Then all images will be display correctly. This has to be checked for every user. Further questions please call 610-219-6300. |  | 2 | 2 | (USA) nationwide including the states of CA, FL, GA, MA, MN, MO, NJ, NY, OH, WA, and WI. | Terminated | January 14, 2014 | 690 | 2
Z-2083-2012 | 62481 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Slice CT Scanner System is a computed Tomography X-ray system, and has a Part number 989605858501. |  |  | June 08, 2012 | 2012 | Philips And Neusoft Medical Systems Co., Ltd. | One of the four screws that secure the Oil-accumulator to the X-Ray Tube pulled out in the NeuViz 16 CT scanner, which caused the Oil-accumulator to detach. | Device Design | Philips Healthcare North America Company sent an "Urgent Device Correction NeuViz 16 X-Ray Tube Oil-accumulator Detached" letter and attachments of Annex A Method to Identify Affected Products and Annex B Instruction for Mechanic, dated June 14, 2012, to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to follow the instructions on Annex A and Annex B.Philips and Neusoft Medical Systems planned to release the mandatory Field Change Order (FCO) and FCO kits to all customers by the end of June 2012.If you have any questions about this corrective action program, please contact the Service Support Department, Neusoft Medical Systems, Co., LTD by e-mail to NMS-SERVICE@NEUSOFT.COM |  | 15 | 15 units distributed in the US | Nationwide Distribution -- North Carolina, Ohio, and Texas. | Terminated | September 17, 2012 | 101 | 15
Z-2086-2012 | 62117 | No | MKQ |  |  | Pathology | PMA | PrepStain Upgrade Kit: 1.3.0.3/1.2 DiTi Cone. |  |  | December 08, 2011 | 2011 | Tripath Imaging, Inc. | Some Prep Stain Kit, 1.2 DiTi Cones have been found to be leaking and the PrepStain Preparation kits contain incorrect preventive maintenance (PM) instructions. | Nonconforming Material/Component | BD Diagnostics sent a letter dated March 2012 to all affected customers. The letter identifies the product, problem and actions that will be taken by the firm. Contact BD Diagnostics Technical Support at 1-877-822-7771, option 2 for questions regarding this letter. |  | 1274 | 1274 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states AL, AK, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WY, and WV and the countries of Belgium, Brazil, Canada, China, Columbia, Hong Kong, India, Japan, Korea, Peru, Singapore, Taiwan and Victoria (Australia). | Terminated | November 08, 2013 | 701 | 2659
Z-2076-2012 | 62007 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015 |  |  | May 25, 2012 | 2012 | CareFusion 303, Inc. | The recall was initiated because the Alaris PC unit model 8015 has a component on the PC unit power supply board that is causing an error code(120.4630) "System Error" or "Missing Battery Error" at start up. The error code is accompanied by both an audible alarm and visual error messages on the PC unit screen. | Component design/selection | Carefusion sent an Urgent: Medical Device Recall Notification letter dated June 27, 2012 to all their customers who purchased the Alaris PC unit (model 8015). The letter identified the affected units, explanation of the problem, potential risk and actions to be taken. Customers were instructed to not return their devices. If customers observe either of the error messages identified in the letter, then they were instructed to remove the PC unit from use and contact the Carefusion Support Center. Carefusion will contact their facility by phone within 60 days of receiving this notification to schedule a visit to replace their affected power supply board on their PC unit. Customers were instructed to promptly complete and return the enclosed customer response card to expedite the corrective action process. For questions and support contact:CareFusion Support Center at 888-562-6018 for (Recall Related Questions), Customer Advocacy at 800-854-7128, Option 1, Option 1, Option 3 or email customerfeedback@carefusion.com for (Adverse Event Reports), Technical Support at 888-812-3229 for (Technical Questions Regarding the Alaris System). |  | 112912 | 112,912 units total (106,888 units in US) | US Nationwide Distribution | Terminated | May 13, 2014 | 718 | 112912
Z-2082-2012 | 62453 | No | KNC | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Affinity Four Birthing Bed P3700 |  |  | April 02, 2012 | 2012 | Hill-Rom, Inc. | During review of our last braking system via that field corrective action, the system still did not meet our expectations, which is to eliminate the hazards of patient falling. Bed movement immediately prior to, or during the birthing process may still create distractions and delays, resulting in patient or caregiver falls, or contribute to a deviation in aprocedure, (e.g. IV insertion, epidural | Device Design | Hill-Rom sent an URGENT FIELD SAFETY NOTICE letter dated April 2, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were informed that Hill-Rom will visit their facility to modify the affected products, however, until then follow the actions to be taken outlined. For questions concerning this request or the procedure as outlined contact Hill-Rom Technical Support at 800-445-3720 (U.S.). |  | 12726 | 12,726 devices | Worldwide Distribution - US Nationwide and countries including: Argentina , Aruba, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Brazil, Brunei, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guam, Guyana, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Libya, Malaysia, Mauritius, Mexico, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, Turks and Caicos Islands, United Arab Emirates, United Kingdom, United States, US Virgin Islands, Venezuela and Vietnam | Terminated | July 21, 2015 | 1205 | 12726
Z-2079-2012 | 61974 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge brand --- large capacity, floor loading jet spray, washers for mechanical washing, intermediate level, thermal disinfection and drying of moisture and temperature stable Hospital case carts, sterilization containers, material handling carts, stands and utensils. --- Model numbers Getinge 9120, 9122, 9125, and 9128 |  |  | May 25, 2012 | 2012 | Getinge USA Inc | Getinge Disinfection AB is voluntarily correcting Washer Disinfectors, Getinge Electrically Heated 9100-Series manufactured between 2009-05-04 through 09-26-2011. This action is based on a report from the ground fault breaker manufacturer of a production error. | Nonconforming Material/Component | Getinge USA issued a Device Correction Notice dated June 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to perform a trip test to determine if they have a functioning ground fault breaker. Customers will be contacted by a Getinge representative to schedule an appointment to exchange the earth fault breaker. Customers were provided a Customer Response Form to complete and fax back to 585-272-5033. For questions customers should call 800-475-9040, ext. 5270.For questions regarding this recall call 585-475-1400. |  | 8 | 8 washers | Worldwide Distribution - USA including Maryland, Indiana, New Mexico, Massachuesetts, Texas, Kansas and the country of Canada | Open, Classified |  |  | 8
Z-2090-2012 | 62456 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Rejuvenate Modular Stems |  |  | April 30, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has updated the instruction for use (IFU) (product correction)for the ABGII Modular and Rejuvenate Modular Hip Systems. This is based on a reported rate of less than one percent for revisions potentially associated with fretting and or corrosion at or about the modular neck junction. On June 29, 2012, Stryker has decided to remove the affected product from the field since there may be an | Device Design | Stryker sent an "URGENT PRODUCT CORRECTION" letter dated April 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included a Product Correction Acknowledgment Form for customers to complete and return via fax to 201-831-6069. Contact the firm at 201-831-5970 for questions regarding this notice.Stryker sent and "Urgent Update Product Recall" Letter dated January 3, 2013 to all affected customers. The new letter includes updated corrective actions for patients. Another letter was sent out on 7/25/2014. |  | 481092399 | (NLS/NLV) - 23193 units and (SPT) - 20,743 units | Nationwide Distribution | Open, Classified |  |  | 481101402
Z-2100-2012 | 33854 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD Cyclesure Biological Indicator (BI), P/N 14324 are sold in cases, packaged as part of STERRAD System Validation kits and STERRAD test packs. |  |  | June 28, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling certain lots of the STERRAD CYCLESURE 24 Biological Indicator (BI) product because it does not have adequate data to support the entire duration of the product's labeled shelf-life. | Incorrect or no expiration date | ASP has posted the STERRAD CYCLESURE Biological Indicator customer letter on their external facing website at: wwww.aspjj.com/us/news/cyclesureletter on July 3, 2012. The letter informed customers who purchased STERRAD CYCLESURE 24 Biological Indicator (BI) product that Advanced Sterilization Products (ASP) is recalling certain lots because ASP has determined that it does not have adequate data to support the entire duration of the product's labeled shelf-life. The letter also provided the customers with what actions are needed for the recall. Customers were instructed to return all unused STERRAD CYCLESURE 24 BI product that exceeds the newly determined expiration dated using the enclosed prepaid UPS return label to Stericycle, 2670 Executive Drive Suite A, Indiapolis, IN 46241. Customers were instructed to count their inventory of product that exceeds the newly determined expiration date and record the data on the enclosed Business Reply Card and packing slip that are included with the customer recall letter. Customers were instructed to complete the requested information on the enclosed postage paid Business Reply Card and return it to Stericycle. Customers were instructed to contact Stericycle at (877) 907-7030 if they need additional packing labels. If customers further distribute the affected products to other locations, then they were instructed to contact Stericycle to arrange for return of the product. Customers were instructed that new shipments of STERRAD CYCLESURE 24 BI product will be shipped on July 9, 2012 that has accurate expiration dates that do not require conversion. Customers wil be able to differentiate between current product and the newly shipped product by the lot number. Current product has a lot number containing 6 characters, eg. 123456.Newly shipped product will have a lot number containing 8 characters, eg. 123456EE.Customers with any additional questions were instructed to contact (877) 907-7030 or contact their A |  | 38693 | 38,693 units | Worldwide Distribution - USA (nationwide) | Terminated | August 13, 2013 | 411 | 38693
Z-2105-2012 | 62553 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | KIT,DISPOSABLE ACCESSORY,4ARM,IS3000,5 PACK; |  |  | July 13, 2012 | 2012 | Intuitive Surgical, Inc. | Specific lots of the Instrument Arm Drapes were manufactured with a sterile adaptor that may have difficulty engaging an instrument. | Nonconforming Material/Component | Intuitive Surgical sent an Urgent Medical Device Correction letters dated July 16, 2012 via Fed Ex to all affected customers in the US. The customers outside of the US were notified starting July 17, 2012. The letter identified the affected products, problem and actions to be taken. Customers were instructed to Identify affected lots, order replacement products and return both full and partial boxes of the affected products for credit as soon as unaffected products are available. All staff should be made aware of the issue. Customers were instructed to complete and return the acknowledgement form immediately.Questions and concerns should be directed to Intuitive Customer Service at 1-800-876-1310.EXPANSION:Intuitive discovered an additional lot (Lot DA121075) of Instrument Arm Drape IS2000, 20 pack that should have been included in the recall and expanded the recall on September 4, 2012. Letters to customers were sent by Fed Ex on September 5, 2012. |  | 92390 | 92390 total instrument Drapes | Worldwide Distribution - US Nationwide and the countries of: Australia, Spain, India, Italy and Saudi Arabia. | Terminated | June 12, 2013 | 334 | 92390
Z-2091-2012 | 61764 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Nextgen Complete Knee Solution Femoral and Provisional Impactor/ Extractor, REF 00-5901-026-00, Non-Sterile |  |  | April 30, 2012 | 2012 | Zimmer Inc. | The design of the spring clip for the jaw assembly was updated in March 2011 due to fractures of the spring clip. Zimmer has received complaints for the spring clip fracturing on the updated design. The spring clip may fracture or fall off of the instrument during use, and a device fragment may be left in the surgical site. The risk of disassembly or fracture occurs when the instrument is typicall | Component design/selection | Zimmer sent an Urgent Medical Device Recall letter dated April 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Requests for alternative instruments should be requested by calling (574) 372-4470. Fax Back Inventory Return Certification Form is to be sent back to Zimmer at (574) 372-4265. All recalled instruments available at the distributor are to be returned to Zimmer Product Service Department, 1777 West Center Street, Warsaw, IN 46580. The Distributor notice included instructions to notify surgeons immediately of the recall and requested a list of hospitals and surgeons receiving the device be sent back to Zimmer via E-mail. Zimmer is continuing distribution of the replacement jaw ( 005901-026-40) until alternative tools are available. The distributors are asked to identify and deliver preferred alternative instruments to affected accounts. Beginning May 7, 2012 Urgent Recall notices are being sent to hospitals and surgeons identified with instructions to acknowledge the recall communication.For questions regarding this recall call 574-372-4807. |  | 1662 | 1662 units | Worldwide Distribution - USA (nationwide) and the countries of Singapore, China, Hong Kong, Canada, Australia, Japan, Europe, Germany, India, and Finland. | Terminated | June 04, 2014 | 765 | 1727
Z-2108-2012 | 62590 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet Getinge Group Operating Table System Magnus Carbon-Fibre Table Top Version 1180.16 |  |  | April 13, 2012 | 2012 | Maquet Medical Systems, Usa | As a result of a small number of customer complaints, MAQUET conducted an investigation and identified a potential problem with specific MAGNUS Carbon-Fibre Table Tops manufactured between February 23, 2010 and February 28, 2012 and distributed between November 11, 2010 and April 6, 2012. There is a potential issue with the motor brakes for the longitudinal shift drive, and the tilt drive. | Device Design | MAQUET GETINGE Group sent An "Urgent Device Field Correction" notification, dated 13 April 2012 via Federal Express to all affected customers. An acknowledgement form is included with the customer letter which can be mailed or faxed back to Maquet Quality by the notified customers. If the completed Acknowledgement form is not received within 15 days, up to three follow ups will be made and documented via telephone call to each customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and notify their MAQUET Service Representative immediately. The MAQUET Service Representative will contact customers to make arrangements to replace both te logitudinal shift and tilt drives of their table top(s). For questions customers should call 888-627-8383, press option 2, followed by option 1.For questions regarding this recall call 201-995-8968. |  | 191 | 164 Units located in 27 Countries | Worldwide Distribution - USA (nationwide) and the countries of United Arab Emirates, Austria, Australia, Belgium, Brazil, Canada, Switzerland, China, Germany, Denmark, Spain, France, Hong Kong, Ireland, Israel, India, Italy, Japan, South Korea, Netherlands, Poland, Russia, Saudi Arabia, Sweden, Singapore, and Thailand. | Terminated | September 04, 2013 | 509 | 191
Z-2120-2012 | 62566 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Octane-C Cervical, 6Âż Lordotic |  |  | June 15, 2012 | 2012 | Exactech, Inc. | Exactech, Inc. initiated a recall of Equinoxe Replicator Plates, Equinoxe Locking Screws, Equinoxe Compression Screw Kits, Equinoxe Fracture Adapter Tray Screws, and Octane-C Intervertebral fusion devices due to the firm determining that the device labeling identifies the affected parts as having a ten year shelf life; however, the packaging configuration that was used it qualified to a five year | Labeling False and Misleading | Exatech sent an Important Voluntary Recall Notice dated June 18, 2012, to all affected customers. The Notice identified the affected product, the problem, and the action to be taken by the customer. Customers were instructed to immeidately case distribution or use of the affected product. Verify that they have the affected product. Contact their Exatech Inventory Representative for further instruction on how to return their products. The firm sent out recall notification to consignees. Customers were asked to complete and return the Voluntary Recall Fax Notice to 1-352-378-2617. Customers with questions should call 1-800-392-2832.For questions regarding this recall call 352-377-1140. |  | 54877 | 54,877 units for all products in Recall Event | Worldwide Distribution - USA including FL, MO, MA, AL, KY, IN, NC, LA,PA, VA, OH, NY, ME, MI, TX, KS, TN, OK, NJ, WA, CO, GA, CA, MN, IA, OR, and IL and Internationaly to Austria, Luxenbourg, Sweden, Germany, Spain, France, Greece, Italy, Nederland, United Kingdom, Canada, India, Australia, Switzerland, and Sao Palo. | Terminated | July 03, 2013 | 383 | 54877
Z-2113-2012 | 62493 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Padded Swing Away Arm Rest with Receiver; |  |  | June 26, 2012 | 2012 | Sunrise Medical (US) LLC | Reports of injury when the Padded Swing Away Armrest is used for full body weight. Owner's manual states that armrest is only designed for resting of forearm. | Labeling design | Sunrise Medical sent a Safety Notice letter dated June 25, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter request that clinicians/suppliers post this safety notice in a conspicuous location, making customers and therapists aware of the issue. The letter also encourage end users to read the owner's manual and follow the recommenations for proper usage of the recalled product. Customers are instructed to complete and return the Supplier Notification and Response Form. For questions contact Sunrise Medical Regulatory Affairs at 888-208-4901. |  | 14035 | 14, 035 | Worldwide Distribution - US (nationwide) and the countries of Argentina, Brazil, Columbia, Costa Rica, Mexico, Panama, Uruguay and Venezuela. | Terminated | November 01, 2012 | 128 | 14035
Z-2110-2012 | 62526 | No | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | Beekley A-Spots Floral Skin Markers -Order Number: 751 |  |  | July 09, 2012 | 2012 | Beekley Corporation | Digital Skin Mammography Marker separating from the backing of the spot material | Unknown/Undetermined by firm | Beekley Medical issued an Urgent Medical Device Recall letter dated July 9, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were advised to contact Beekley Medical's Customer Service Team at 1-800-233-5539 to return affected lots for replacement. Customers were instructed to complete and return the attached reply form. |  | 32640 | 32,640 units | US Nationwide Distribution | Terminated | July 09, 2013 | 365 | 342910
Z-2114-2012 | 62626 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Aesculap SIBD 4 degree Trial Implant (SJ768T-US) and Aesculap SIBD 9 degree Trial Implant (SJ780T-US) |  |  | July 12, 2012 | 2012 | Aesculap Implant Systems LLC | The 4 degree and 9 degree trial instruments (SJ768T-US/SJ780T-US) are correctly dimensioned, but the part number etched on the instrument may be incorrect. The 4 degree trial instrument (SJ768T-US) may be marked with the incorrect part number of the 9 degree trail instrument (SJ780T-US) and the 9 degree trial instrument (SJ780T-US) may be marked with the incorrect part number of the 4 degree tria | Employee error | Aesculap sent an Important Product Removal Notification letter dated July 11, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to return the affected product by contacting their local sales representative or the customer service department at 1-866-229-3002 for a Returned Goods Request )RGR) number. Customers were asked to complete and return the Distribution Inventory Sheet. For questions call 610-984-9260 or 610-984-9274. |  | 5 | 5 | US Nationwide Distribution - including the states of CO, OH, OK and TX. | Terminated | February 22, 2013 | 225 | 5
Z-2123-2012 | 61655 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | IntegraÂż Advansys Medial Lisfranc Plate (MLP), Integra, 311 Enterprise Drive, Plainsboro, NJ 08536 |  |  | February 09, 2012 | 2012 | Integra LifeSciences Corp. | Incorrect technique applied may cause potential rupture of the tibialis anterior tendon during surgery. | Labeling design | INTEGRA sent an Urgent Medical Device Correction Notice dated February 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to identify members of surgical team that ues the affected product and make them aware of the Medical Device Correction Notice. Customers were asked to sign and return the enclosed Medical Device Correction Notice Acknowledgement and Return Form and to maintain a copy of the notification and signed copy of the acknowledgment form for their records. For any questions regarding this recall call 609-936-6832.. |  | 2437 | 2,437 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NM, NJ, NY, NV, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA &amp; WI and internationally to Belgium; Danemark; France; Germany; IsraÂżl; Italy; Lebanon; Luxemburg; Netherlands; Norway; Portugal; Saudi Arabia; Spain; Switzerland; Turkey &amp; United Kingdom. | Terminated | February 19, 2013 | 376 | 2437
Z-2126-2012 | 62449 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | OEC MiniView 6800 Mobile; MDL number A349855. |  |  | March 27, 2007 | 2007 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC had discovered an issue with the hard disk drives that were installed on OEC device models: UroView 2800; 9800 Plus Digital; MiniView 6800; and FlexiView 8800 Mobile C-Arm. The devices were manufactured between October 15, 2004 and | Software design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC sent an Urgent Safety Advisory Notices dated October 11, 2006 and February 21, 2007 via certified mail delivery letter to affected customers. The notices identified the affected products, problem and actions to be taken. Customers were informed that a GEHC OEC field service engineer would be contacting their facility to arrange to have a replacement hard disk drive installed on the affected system(s). Customers were instructed to ensure that images on the hard disk drive were stored via an alternate permanent media as the hard disk drive approaches the 170 image capacity until a GEHC OEC field service engineer had replaced the drive. For questions regardiug this issues contact the service team at 800-874-7378. |  | 82 | Part of a total of 82 units distributed. | Worldwide Distribution - US (nationwide) and the countries of: Canada, France, Japan, and China. | Terminated | August 02, 2012 | 1955 | 82
Z-2107-2012 | 61247 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, MedfusionÂż Syringe Infusion Pump, Model 4000, Software Version V1.0 &amp; V1.1, Reorder Numbers 4000-0100-50, 4000-0100-95, 4000-0101-50, 4000-0101-51, and 4000-0101-78. |  |  | February 14, 2012 | 2012 | Smiths Medical ASD, Inc. | Smiths Medical has identified a software anomaly in the Medfusion 4000 Syringe Infusion Pump that causes a device history log corruption and triggers a Watchdog Fail-Safe alarm. When the device exhibits this failure mode, visual and audible alarms will sound and the device ceases operation. | Software design | Consignees of MedfusionÂż 4000 pumps with software version 1.0 &amp; 1.1 were visited by a Smiths Medical representative on 2/14/12 informing them that their pumps require a software upgrade. A record of this software upgrade was documented as part of the records for this action. |  | 607 | 607 pumps (257 USA, 350 Canada) | Worldwide Distribution -- USA, including the state of CA, and country of Canada. | Terminated | November 14, 2012 | 274 | 607
Z-2124-2012 | 62555 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IRISpec" CA/CB two-part urine chemistry control twin set, Iris Diagnostics part number 800-0074. |  |  | July 09, 2012 | 2012 | Iris Diagnostics | The firm recalled after it received complaints of IRISpec CA/CB two part Urine Chemistry Control, 101131-12, where the nitrite results read negative for the part B control, which should be positive. | Nonconforming Material/Component | The firm initiated the recall on 07/09/12 by sending out a Medical Device Recall notification to customers via e-mail. The recall notification identified the affected product, reason for recall, and instructions on the recall. Customers were to confirm how many sets of unopened or partially opened IRISpec CA/CB that was in their inventory, dispose of the affected product, and provide written confirmation of the product's disposition. The letter also stated that Iris Diagnostics would replace any unused inventory that was returned at no charge to the customer. Contact information for the recall was also provided. The communications also included a fax response sheet for return/replacement. |  | 862 | 862 | Nationwide Distribution | Terminated | January 27, 2014 | 567 | 862
Z-2122-2012 | 62557 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | The Reliance 130 and 130L Cart Washer/Disinfector is a high capacity mechanical washer intended for use in the washing, low-level disinfecting, and drying of carts, utensils, beds and other miscellaneous reusable items used in the care of patients at healthcare facilities / Crated for shipping. |  |  | July 14, 2012 | 2012 | Steris Corporation | To ensure that users are properly operating the Reliance 130 and 130L Cart Washers, a door close confirmation button and door open delay will be installed on your cart washer. | Labeling design | Steris Corp. sent a Urgent Field Correction Notice dated July 14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.To ensure users are properly operating the Reliance 130 and 130L Cart Washers a 'door close' confirmation button and a door open delay will be installed on the affected cart washer(s). A STERIS Service Representative will contact the affected facilities to arrange for STERIS to upgrade their Reliance 130 and 130L Cart Washer(s). For further information or if affected customers have any questions regarding the STERIS visit to their facility, they can contact STERIS Field Service Dispatch at 1800-333-8828 |  | 367 | 367 | Worldwide Distribution -- USA (nationwide) including the states of : AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY. and the countries of : Australia, South Korea, France, Italy, Lebanon, Russia, Saudi Arabia, Spain, UAE, United Kingdom and Canada. | Terminated | June 06, 2014 | 692 | 367
Z-2121-2012 | 62512 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Perkin Elmer Genetic Screening Processor (GSP) |  |  | June 27, 2012 | 2012 | Perkinelmer | Contamination from GSP Shaker unit results in decreased signal activity of assay measurements for Genetic Screening for Neonatal IRT, TSH and IRT | Component design/selection | PerkinElmer sent an Urgent: Medical Device Recall letter dated June 27, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that replacement units of the recalled product will be installed as soon as they are available. Customers were instructed to complete and return the response form attached. For further information contact your local PerkinElmer representative or GS. Technical.Support@PerkinElmer.com |  | 78 | 78 units | Worldwide Distribution - US (nationwide) and the countries of Australia, Belgium, Brazil, Egypt, France, Germany, Netherlands, Norway, Saudi Arabia,Switzerland, | Terminated | April 18, 2014 | 660 | 78
Z-2080-2012 | 62432 | No | IZH | Mammographic x-ray system. | Radiology | Radiology | 510(k) | Bella Blanket Protective Coverlet, for Mammography, 18cm x 25cm or 24cm x 30cm, Product Code: 311. The purpose of the product is to remove the cold by placing this sheet between the patients breast and the imaging receptor plate. |  |  | July 02, 2012 | 2012 | Beekley Corporation | Imaging artifact can occur on mammogram result | Nonconforming Material/Component | Beekley Medical issued "Urgent: Medical Device Recall" notifications dated June 29, 2012 that were sent via certified mail USPS on July 2, 2012. The product issue was was described and customers were provided the recommended actions. Customer Service number is 1-800-233-5539. |  | 86000 | 86,000 units | International Distribution including Nationwide(USA), and the countries of Japan and Canada. | Terminated | July 09, 2013 | 372 | 86000
Z-2129-2012 | 62538 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Cardiac Catheterization Waste Management System w/1000mL Waste Bag, 2 Fluid Admin Sets |  |  | April 17, 2012 | 2012 | ICU Medical, Inc. | The firm initiated the recall due to a molding defect in a high pressure stopcock where a smallpercentage of the side ports could become totally occluded. | Nonconforming Material/Component | ICU Medical sent an Urgent Medical Device Recall letter dated April 2012 to all affected customers. The letter identified the recalled product, problem and actions to be taken. Customers were instructed to inspect inventory, destroy affected products according to hospital protocol, complete and return verification form and contact ICU Customer Service to order replacement products. The letter states that ICU Medical will credit customer for any product destroyed once the completed verification form is received. This recall should be carried out to the user level and passed on to all those who need to beaware within your organization or any other organization the device may have been transferred to.Should you have any questions or require assistance relating to this recall please contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address: customerservice@icumed.com |  |  |  | Nationwide distribution | Terminated | August 23, 2012 | 128
Z-2132-2012 | 62528 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Extraneal and/or DIANEAL LOW CALCIUM Peritoneal Dialysis Solution in Plastic Container, PL-146 used with Automated Peritoneal Dialysis System/Automated PD Set with Cassette; Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A. The HomeChoice dialysis system is a personal cycler automated peritoneal dialysis system for pediatric and adult renal patients with fill volumes ranging from 60 mL to 3000 mL. |  |  | July 06, 2012 | 2012 | Baxter Healthcare Renal Div | Baxter discovered that a Baxter technical service representative was providing incorrect instructions for inspection of solution bags for a blocked frangible prior to connection to the Automated PD Set with Cassette. | Employee error | Baxter Healthcare Corp. sent "Important Product Information" letters dated July 6, 2012 to the affected home patients and their clinicians, via first class mail on the same date. The letter identified the product, problem and actions to be taken. The customers were instructed to refer to the Extraneal and/or Dineal package insert, Directions For Use - Preparation for Administration and the HomeChoice APD Systems At-Home Guide for proper procedures to connect to a solution line. Customers are instructed to retain a copy of the letter. The customers were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Global Technical Services, Renal Technical Support at 1-800-553-6898, option 1, 24 hours a day.For questions regarding this recall call 1-800-422-9837. |  |  |  | Nationwide Distribution includingArizona, California, Ohio, Pennsylvania and Virginia. | Terminated | October 15, 2012 | 101
Z-2135-2012 | 62504 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | RF Cannula, 100 mm,Curved 10 MM, 20 G, REF 0406-630-125 RX only, Disposable ( 10/pk) |  |  | July 17, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | Stryker has received 6 complaints that the 5mm needle was placed in 10 mm package and 10 mm needle was placed in the 5 mm package. If the clinician does not recognize that the active tip on the cannula is 10 mm instead of 5mm and the cannula is used in a cervical procedure there is a potential for a thermal skin lesion due to a larger lesion creation. If the 10 mm active tip cannula, as opposed | Labeling mix-ups | Stryker sent an Urgent MEDICAL DEVICE Recall Notifications dated July 16, 2012 to all consignees. The letter identified the affected products, problem and actions to be taken. Customers/distributors were instructed to check their inventory and quarantine any affected product found, complete and fax the business reply form to Stryker Instruments Regulatory Department, 866-521-2762 and return all affected product using the pre-paid shipper provided. Upon receipt of the recalled product, a replacement will be issued.For questions regarding this recall please contact Stryker Instruments (Recall Coordinator) Monday-Friday 8am-5pm (EST) at 269-389-4354. |  | 226 | 226 | Worldwide Distribution - US Nationwide and the countries of England, France, Netherlands, and Switzerland. | Terminated | May 02, 2013 | 289 | 436
Z-2133-2012 | 62371 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | I-STAT EG7+ Cartridge Abbott Point of Care, Princeton, NJ. |  |  | January 01, 2011 | 2011 | Abbott Point Of Care Inc. | i-Stat EG7 cartridges may generate falsely depressed PCO2 results for samples above 26 mm hg when stored at room temperature for four weeks or longer. | Nonconforming Material/Component | Abbott Point of Care, Inc sent a Urgent Recall Notice letter dated January 14, 2011 to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to discontinue use of Cartridges from the identified lot. All unused cartridges will be credited and are to be returned as instructed on the enclosed Business Reply Card.If you have any questions regarding this information please contact your Abbott Point of Care Technical Support at 800-366-8020 Option 1, or your Abbott Point of Care representative. |  | 52750 | 52,750 cartridges | USA ( nationwide ) | Terminated | May 22, 2013 | 872 | 52750
Z-2138-2012 | 62514 | No | FIN | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | TANGO3 Water Storage Tank with Ozone Disinfection System, Tango3, LLC, 132 Citizens Blvd., Simpsonville, KY 40067. Models are WOZ-250-PE, WOZ-240-SS, WOZ-180-SS. The TANGO3 Water Storage Tank with Ozone Disinfection System is intended to be used for disinfection of the water distribution system of a dialysis facility. |  |  | March 26, 2012 | 2012 | Tango3 LLC | During an FDA inspection, the firm was notified of Quality System deficiencies, which included lack of procedures for design control and customized modifications that were either not reviewed or received proper clearance from FDA. There was no assurance that all systems in the field were working properly and operating at a minimum of 0.2ppm of ozone. | Process control | On 3/26/2012, the firm sent "Urgent: Customer Notification" letters and Acknowledgement Forms to their customers. The issue was described and recommended actions were provided. Customers were given the number 866-782-7286 to call if questions. |  | 44 | 44 units | Product was shipped to the following states: CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, MA, MD, ME, NH, NJ, NM, OH, PA, RI, TX and VA. | Terminated | March 01, 2013 | 340 | 44
Z-2145-2012 | 62617 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew 5.5 mm Dyonics Bonecutter ,Electroblade |  |  | July 14, 2012 | 2012 | Smith &amp; Nephew, Inc. Endoscopy Division | Device labeled with an expiry period of 48 months instead of 36 months; an additional 12 months for which product and package reliability has not been validated. | Error in labeling | Smith &amp; Nephew sent an Urgent - Product Recall 1st Notification Letter dated July 2012 via Federal Express to all affected customers. The letter identified the affected product, potential risk and actions to be taken. Customers were instructed to examine inventory, quarantine immediately, complete the form attached indicating the lot numbers and quantities. The letter also instructed customers to contact Smith &amp; Nephew Returns Group at 800-343-5717 (option 3) for a Return Authorization (RA) or send an email to endo.andreturns@smith-nephew.com. They will provide customers with instructions for returning affected products and receiving a credit. Each OUS affiliate shall be notified via (email or equivalent) of the recall by Smith &amp; Nephew Inc.; the affiliate must then notify each facility affected by the recall.For question concerning this recall call 405-917-8585 or 978-749-1630. |  | 108 | 108 units | Worldwide Distribution - US Nationwide and the countries of : Australia, Canada, Korea, Turkey, UAE, and Japan. | Open, Classified |  |  | 448
Z-2137-2012 | 62583 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | SPIDER Surgical Instrument: Flexible Clip Applier |  |  | July 02, 2012 | 2012 | Transenterix, Inc | Some clip applier units may be inadequate to achieve clip closure. | Process control | TransEnterix sent an Urgent Device Recall letter date July 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for affected product, remove the affected product from their inventory and either return it or dispose of it to prevent usage. Customers were instructed to complete the attached Inventory Reconciliation Formand return the form even if they no longer have any affected units in their inventory. Customers were instructed to ship the affected product to TransEnterix using the enclosed shipping label. If returning a form only, they should e-mail to adunn@transenterix.com. For questions customers were instructed to call 1-888-879-4111.For questions regarding this recall call 919-765-8420. |  | 89 | 89 boxes(6 units per box) | Worldwide Distribution - USA (nationwide) and the countries of Austria, Germany, Italy, Qatar and Switzerland | Terminated | February 27, 2013 | 240 | 89
Z-2143-2012 | 62023 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab with SA INST versions: 3.1.6.12, 4.0.1.26-4.0.1.32, 4.0.2.21-4.0.2.51, and 4.0.3.5-4.0.3.13. |  |  | October 17, 2011 | 2011 | SCC Soft Computer | On 10/17/2011 SCC Soft Computer, Clearwater, FL initiated a correction on SoftLab with SA INST versions: 3.1.6.12, 4.0.1.26-4.0.1.32, 4.0.2.21-4.0.2.51, and 4.0.3.5-4.0.3.13. A client reported an issue which caused two different sets of results from the Sysmex series of Hematology instruments to display under the same order on the instrument menu. | Software design | SCC Soft Computer's proprietary Task Management System sent a correction communication letter to all consignees with affected software. The letter identified the affected product, problem and corrective actions to be taken. The letter also recommended alternative working solution if the corrective actions cannot be used. Consignees with affected software are scheduling receipt of the software correction. For questions call 727-789-0100. |  | 88 | 88 | US Nationwide Distribution | Terminated | September 03, 2014 | 1052 | 88
Z-2151-2012 | 62665 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PopLok Punch labeled in part: |  |  | July 06, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec recalled one lot of their PopLok bone Punch 3.5MM because the metal shaft diameter is 4.5mm when it should have been 3.5mm. | Process change control | The recall is being conducted to the end user level. ConMed Linvatec Urgent Medical Device Recall Notifications were sent by FedEx next day delivery or equivalent method to notify all direct consignees. The notifications identified the affected product along with the reason for the recall and health risk. All accounts were asked to immediately check their facility for the affected product, stop using any of the product with specified lot number, and to segregate the product. The letter also asked customers to return all affected product to ConMed Linvatec and respond using a Reply Form, even if there is no affected product available. If customers have any additional questions, they should contact the Customer Service department at 800-237-0169. |  | 2 | 2 | Worldwide Distribution -- US and Korea. | Terminated | November 21, 2013 | 503 | 2
Z-2149-2012 | 62563 | Yes | FIR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Plasma Pump Model PP-04a |  |  | July 09, 2012 | 2012 | Apheresis Technologies, Inc. | On 07/09/2012, Apheresis Technologies, Inc. initiated a recall on the Plasma Pump Model PP-04a due to a complaint received regarding a rotor of the PP-04a pump that stopped during a therapeutic plasma exchange treatment causing interruption of flow of both the plasma discard line and replacement fruit line. | Process control | APHERESIS TECHNOLOGIES, INC. sent a letter dated July 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the device for repair. For questions customers should call 727-787-5616. |  | 13 | 13 | Natonwide Distribution | Terminated | December 16, 2013 | 525 | 13
Z-2148-2012 | 62683 | No | FWH | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Chromovision High Definition (HD) Wireless Camera, Pt # CC-9715900. |  |  | July 10, 2012 | 2012 | Berchtold Corp. | Defective camera: There is a possibility that the connection of the wires to the connector may become loose, causing image distortion or loss. | Nonconforming Material/Component | BERCHTOLD Corporation sent a letter dated July 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. BERCHTOLD will replace the customer's cameras at no cost to the customer. Customers with questions should call 843-569-6100, ext 322.For questions regarding this recall call 843-569-6100. |  | 17 | 17 cameras | Nationwide Distribution including FL, MI, MS, and NH. | Terminated | June 06, 2013 | 331 | 17
Z-2146-2012 | 62561 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | djo surgical Screw 6.5 mm, Sz 25mm, low profile. |  |  | July 16, 2012 | 2012 | Encore Medical, Lp | One lot of the 010-55-030, 6.5 mm Cancellous Bone Screw was recalled because a product complaint identified a mislabeling of the device. Although the package label indicates Size 30mm, the screw dimension is actually 25mm. | Packaging | DJO Global sent an Urgent Field Safety Notice dated July 16, 2012, to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the notice to all thoe who need to be aware within their organization or to any other organization where the potentially affected product was transferred, contact customers to place a replacement order and receive an RMA number, and return all affected devices using the RMA number. For questions customers should call 512-834-6255.For questions regarding this recall call 512-832-9500. |  | 18 | 18 devices | Nationwide Distribution including AR, CA, FL, ID, OK, PA, TX, and UT. | Terminated | December 04, 2012 | 141 | 18
Z-2150-2012 | 62556 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Finapres Medical Systems Portapres with NiMh battery pack |  |  | July 06, 2012 | 2012 | FMS, Finapres Medical Systems BV | Battery may become hot and leak | Process control | Finapres Medical Systems BV sent a Field Safety Notice letter dated July 6, 2012 to notified US accounts of the recall. The letter identified the affected product, description of the problem and actions to be taken. The letter states that the recalling firm recommends not using the affected battery packs. A replacement battery was included in shipment. Customers were requested to return recalled battery. For questions call +31 (0)20 609 0974 or email info@finapres.com. |  | 3 | 3 units US | US Nationwide Distribution - including the states of: OH and NY | Terminated | September 12, 2012 | 68 | 3
Z-2147-2012 | 62395 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CrossFT Punch, Product Number PFT-45M. Manufacturer: CONMED LINVATEC. The intended use is to create a hole to facilitate insertion of the CrossFT suture anchor into bone. |  |  | August 12, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec recalled lot number #156920, PFT-45M, CrossFT Punch because there is a possibility that the device was etched incorrectly listing the item number as PFT-00M and includes incorrect depth lines for insertion. | Use error | The firm sent an "URGENT: Medical Device Recall Notification" letter dated August 12, 2011, to its customers after initial phone notifications. The letter describes the product, problem and actions to be taken. The customers were instructed to: 1) Immediately check their facility's inventory for the product number and lot number listed in the letter. If the product is at their facility, they were instructed to discontinue the use of the product. 2) To return the products to ConMed Linvatec for credit and return the products and Reply Form. 3) If the product was transferred to another facility, complete the Reply Form documenting the transfer of the product, and return i by mail or fax. For questions, contact 727-399-5205. |  | 2 | 2 | Nationwide Distribution, including the states of Colorado and Pennsylvania. | Terminated | November 21, 2013 | 832 | 2
Z-2152-2012 | 62535 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model 731 EMV+ Uni-Vent, Impact Instrumentation Inc., 27 Fairfield Place, West Caldwell NJ 07006. Indicated for use in the management of acute or chronic respiratory failure or during resuscitation by providing continuous positive -pressure ventilation. |  |  | March 19, 2012 | 2012 | Impact Instrumentation, Inc. | Ten units of Model EMV+ Devices were shipped out in error before the verification and validation was completed for a new integrated circuit (IC) on the CPU PCB. | Process control | Impact Instrumentation Inc. notified their sole customer via email on 3/15/12 and then emailed an "Urgent Device Product Recall" letter and tracking form dated 3/19/2012. The letter described the reason for the recall and provided recommended actions and the contact number for Customer Service. |  | 10 | Ten Units | US Distribution to the state of Illinois. | Terminated | August 15, 2012 | 149 | 10
Z-2155-2012 | 62648 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LUXOR RETRACTOR BASE; Catalog Number: 48250030; Manufactured by: STRYKER Spine SAS; Z.I MARTICOT-33610-CESTAS-FRANCE; Distributed in the USA by: Stryker Spine. 2 Pearl Ct. Allendale, NJ 07401. |  |  | April 23, 2012 | 2012 | Stryker Spine | Stryker received three reports involving blades of the Luxor Retractor becoming disassembled from the base causing the retractor to collapse insitu. | Nonconforming Material/Component | Stryker sent Urgent Product Recall letters/return response forms dated April 23, 2012 via FedEx to Stryker branches/agencies and hospitals where product might have been distributed. The letter identified the affected product, discussed the issue, potential hazards, risk mitigations, and the return process. Customers are to examine their inventory and reconcile any of the affected product with the information provided. In addition, the Customer Response Form should be completed and faxed back to Regulatory Compliance at 201-760-8370 within 5 days of receipt of the letter. The affected product should also be returned following the instructions provided. Questions should be directed to Dan Roberts, Regulatory Affairs Compliance Specialist at 201-760-8298. |  | 30 | 30 devices | Worldwide Distribution -- US, Switzerland, Netherlands, Brazil, Hong Kong, Germany, and Taiwan. | Terminated | November 13, 2013 | 569 | 30
Z-2154-2012 | 62595 | No | GIM | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | 3mL 9NC Coagulation Sodium Citrate 3.2%, Item #454334, |  |  | July 02, 2012 | 2012 | Greiner Bio-One North America, Inc. | Part of the lot received a lower concentration of citrate solution which may cause falsely lower values for common coagulation tests. | Nonconforming Material/Component | Distributors were notified by letter on 07/02/2012 and customers were sent a letter dated 07/05/2012. The letters identified the affected product and the reason for the recall. Distributors were asked to stop distributing the affected product and to complete the attached Product Disposition form. The form along with the affected product was to be returned to GBO. The firm will provide replacement stock or issue a credit once product return and receipt of the Product Disposition form has been confirmed. Customers were also asked to complete the Product Disposition Site Confirmation form and fax it back to Greiner Bio-One North America, Inc. at the number provided. The letter states that the firm will replace any affected product once the form is received. Questions should be directed to 800-515-8112. |  | 838.9 | 838.9 cases (1,006,700 pieces) | Nationwide Distribution | Terminated | June 11, 2013 | 344 | 838.9
Z-2158-2012 | 62375 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | ***REF IFNT611***Full OSSEOTITE Tapered Certain Implant 6 x 11.5mm***Sterile by radiation***Rx Only***BIOMET 3i Dental Iberica S.L. WTC Almeda Park, Ed, 1, Planta 1 Pl. de la Pau s/n08940 Cornella de Llobregat (Barcelona) Spain. Lot #2011110798. |  |  | March 23, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i is recalling their product Full Osseotite Tapered Certain Implant IFNT611. The depth of the implant's internal hex is too shallow. The condition, if present, may prevent the driver; cover screw or abutment from fully engaging the implant. | Process control | BIOMET 3i sent an Urgent Medical Device Recall letter on March 19, 2012, to all affected customers via email or fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers are to return any unopened and unused product for replacement at no cost. In addition, customers should contact Customer Service at 1-800-342-5454 for assistance if they encounter issues with the mating component, or if they have any questions or concerns. Customers are to respond to the notification regardless of whether they have product to return. |  | 240 | 240 | Worldwide Distribution -- United States, UK, France, Canada, Mexico, Poland, Lebanon and Hong Kong. | Terminated | December 16, 2013 | 633 | 240
Z-2161-2012 | 62575 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Straumann Narrow CrossFit (NC) Screw for CARES Abutment ZrO2. Endosseous dental implant |  |  | July 11, 2012 | 2012 | Straumann USA, LLC | Incorrect Screw (RC) included with NC CARES Abutment Set was not a Narrow CrossFit (NC) Screw for CARES NC Abutment ZrO2 | Packaging process control | Straumann notiifed accounts by telephone on July 10, 2012 and follow-up Field Safety Correction Action letter dated July 11, 2012, via United Parcel Service. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to discontinue use of the affected product and follow the instruction provided until replacement is received. For questions call 978-747-2514. |  | 14 | 14 Distributed in US | US Nationwide Distribution including the countries of FL, IN, LA, NY and TX. | Terminated | September 05, 2012 | 56 | 14
Z-2139-2012 | 62558 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | AUTOMIX High Speed Compounder System, product 2M8077; automated nutrition compounders; |  |  | July 13, 2012 | 2012 | Baxter Healthcare Corp. | After fluid ingress into the keypad of the AUTOMIX Compounder's (AUTOMIX) control module, the AUTOMIX may generate an incorrect device response to an AUTOMIX operator's key press. Also, the AUTOMIX system may experience an intermittent electrical failure, in which unexpected pump motor operations will occur. Also, the AUTOMIX system may intermittently experience an electrical failure where thecompounder may stop compounding without an alarm and not be detected by the user. | Device Design | Baxter Healthcare Corp. sent an Urgent Product Recall letters dated July 13, 2012, via first class mail to all affected customers, informing personnel that Baxter was executing a voluntary controlled product withdrawal of the Automix system from the U.S. market by June 1, 2013. Until their units can be replaced, customers were advised to exercise care when entering data, and verify the accuracy of the keypad response; to inspect the umbilical cable of the Automix for any visible cuts, nicks deformation, or physical damage to the cable, discontinuing use of the device if such damage is found; and to return the device to Baxter if the device stops compounding without an alarm and without a visible green complete light, discarding the solution bag. The customers were requested to convey this information to all users of the Automix compounders, and a wall chart was provided to hang over the device to remind users of the mitigations to take. The customers were also requested to complete the attached customer reply form listing the model and serial number of each of the Automix compounders at the facility and confirming their receipt of the letter, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to their local Baxter representative, and technical questions were directed to Baxter's Global Technical Services at 1-800-626-2667. |  | 270 | 270 units | Worldwide Distribution - USA - Nationwide including Puerto Rico, and internationally to Brazil, Canada, Colombia, Costa Rica, Czech Republic, El Salvador, Egypt, Germany, Guatemala, Honduras, Luxembourg, Mexico, Panama, Saudi Arabia, Singapore, Switzerland, Taiwan, Trinidad, Turkey, United Arab Emirates and United Kingdom. | Terminated | August 11, 2017 | 1855 | 5166
Z-2163-2012 | 62503 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The ABL software version 6.12 for the ABL800 and ABL700 analyzer. The analyzers are intended for the in Vitro testing of: |  |  | July 05, 2012 | 2012 | Radiometer America Inc | A customer has reported an error on an ABL8xx configured to a Hospital Information System using the data transfer protocol HL7 version 2.5. If the instrument transmits results to a HIS system and the connection is broken the results are queued for later transmission. If the ABL and HIS system are configured with HL7 2.5, any results in the queue will be mixed up when the instrument transmits t | Software design | The firm, Radiometer, sent two "Priority Level: Urgent" letters one dated July 9, 2012 and the other dated July 24, 2012 to its affected customers via FedEx with tracking. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to complete and return the attached Recall Response Fax Form via fax to 800-736-0601. The letter informed the customers that a local service representative will be contacting them shortly to install a software update that will prevent this from occuring (in the event that you utilize HL7 version 2.5 or intend to at some future date.) If you have any questions regarding this letter or the Recall, please contact Radiometer America Technical Support at 1-800-736-0600 opt. 4. |  | 6861 | 6854 ABL 7xx units | Worldwide distribution: USA (nationwide) including the states of AK, AL, AR, AZ, BM, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and the countries of Canada, Algeria, Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cuba, Cyprus, Czech republic, Denmark, Ecuador, Egypt, Eritrea, Estonia, Finland, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libya, Lithuania, Malaysia, Malta, Mauritius, Mexico, Moldova, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Peru, Philippines, Poland, Puerto Rico, Qatar, Republic of Korea, Romania, Russia, Saudi Arabia, Switzerland, Singapore, Slovenia, South Africa, Spain, Sweden Syria, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, Yugoslavia, and Zimbabwe. | Terminated | August 20, 2013 | 411 | 20315
Z-2157-2012 | 62536 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | I-STAT PT/INR cartridge Abbott Point of Care Inc. Abbott Park, IL 60064 USA |  |  | March 01, 2012 | 2012 | Abbott Point Of Care Inc. | The i-STAT PT/INR cartridges have the potential to exhibit incorrectly elevated results. | Process control | Abbott Point of Care Inc, sent a URGENT RECALL NOTICE dated March 2012 , to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were provided in the letter with instructions in the letter on how to handle the affected product. Please complete and return the enclosed Business Reply card to acknowledge receipt of this communication and installation of the interium software CLEW and also 1) Update all handhelds 2) download appropriate assignment sheets. For further questions please call (609) 454-9000. |  | 3106704 | 2,950,296 cartridges US; 156,408 cartridges OUS | Worldwide Distribution -- (USA) nationwide | Terminated | February 08, 2013 | 344 | 3106704
Z-2159-2012 | 62500 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas KRAS Mutation Test for In Vitro Diagnostic Use |  |  | February 24, 2012 | 2012 | Roche Molecular Systems, Inc. | During the real-time stability testing for the cobas KRAS Mutation Test on the commercial lot PO6778, it was identified that the kit was just within specifications for the codon 61 peak height ratio, and that the kit has trended towards the upper specification limit at the 6 month time point but did passed the 6 month stability testing. The kit is labeled for 9 months shelf life stability. The d | Nonconforming Material/Component | Roche issued a Safety Board Notice (Prooduct Support Notice 2012-01) and Product Bulletin 2012/03 dated February 24, 2012 to all affected customers. The notice idenfied the affected product, description of situation and actions to be taken. Customers were instructed to discontinue use of lot P06778 by March 4, 2012 and to quarantine and discard any remaining kits from lot P06778. The notice stated that a new lot will be made available as a replacement. |  | 3432 | 264 kits to 13 EU countries; 2 kits to Canada and 33 kits to 11 countries in the Rest of the World | Worldwide Distribution - including the countries of Austria, Belgium, Czech Republic, Denmark, Germany, Italy, Luxemburg, Netherlands, Poland, Slovakia, Slovenia, Spain, United Kingdom, Canada, Chile, Greenland, Hong Kong, Israel, Malaysia, New Zealand, Philippines, S. Africa, Switzerland, Taiwan and Turkey. | Terminated | July 16, 2013 | 508 | 3432
Z-2153-2012 | 62562 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian ARIA Oncology Information System, Import Export |  |  | July 02, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Import Export Application version 11.0 whereMLC leaf positions for plans using BrainLab m3 and mMLC are incorrectly imported. | Software design | Varian Medical Systems sent a " Urgent Medical Device Correction Urgent Field Safety Notice" dated July 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed until a correction is deployed all treatment plans using the BrainLab m3 mMLC must be imported using the Import Wizard only.Varian Medical Systems is developing a technical correction for this issue. You will be contacted by a Customer Service Representative when this correction is available to schedule its installation on your system.For further questions please call 1.888.827.4265. |  | 2 | 2 Units. | Nationwide Distribution including the states of New York and Ohio. | Terminated | July 01, 2013 | 364 | 2
Z-2160-2012 | 62645 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | The finger cots are made out of natural rubber. They are disposable products used to cover a fingertip bandage when working with liquids and are intended to keep the fingertip bandage dry. Cots are packed in clear plastic boxes, with 60 cots in a box |  |  | February 08, 2012 | 2012 | Afassco Inc | The firm is initiating a field correction because the product(s) were missing the Latex Warning Label. Afassco Inc. is implementing a voluntary field correction for MULTI-COLORED FINGER COTS because product lacked the warning label:"CAUTION: This product contains Natural Rubber Latex which may cause allergic reactions." | Packaging process control | Afassco sent a Field Correction Notice to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Labels have been included with the notices that the firm will be sending to their distributors. The distributors were asked to examine their inventory and affix the caution label to the boxes of affected product. If additional warning labels are needed, they may contact Afassco (800-441-6774) for additional labels.If they have further distributed these products, they should notify consignees of the field correction. Their notification to their customers may be enhanced by including a copy of the notice sent to them. The recall should be carried out to the user level. For any questions customers may call 1-880-441-6774.For questions regarding this recall call 775-783-3555. |  | 1956 | 1956 Boxes | Nationwide Distribution including CA, FL. GA. HI, IL, IN, MI, NH, NV, TN, TX, and WV. | Terminated | January 14, 2014 | 706 | 1956
Z-2156-2012 | 62551 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva with Flat Detector (MD Eleva FD)Model: 708034 |  |  | July 11, 2012 | 2012 | Philips Healthcare Inc. | When doing a 3 Dimensional Rotational Angiography (3DRA) the 3D reconstruction does not always start, even when all images have been acquired. | Software design | Philips issued an Urgent Field Safety Notice dated July 2, 2012 to all affected customers. The letter identified the affected product, problem description, actions to be taken by the customer/user in order to prevent risks for patients or users and actions planned by Philips to correct the problem. For question or support concerning this issue contact your local Philips representative. |  | 41 | 41 units | Worldwide Distribution - US (nationwide) including the states of CA, CT, MI, MN, MN, NC, NY, NY, VA and the countries of Belgium, Columbia, Czech Republic and Denmark.FranceGermanyItalyLatviaLuxembourgNetherlandsRussian Fed.SpainSwedenSwitzerlandThailandUnited Kingdom | Terminated | March 26, 2013 | 258 | 41
Z-2136-2012 | 62414 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Radiological Image Processing System |  |  | April 02, 2007 | 2007 | GE OEC Medical Systems, Inc | The FluoroTrak Spinal Navigation Application on the OEC 9900 EliteNAV could result in an incorrect position of the navigated instrument(s) versus the displayed reference image. | Software design | GE Healthcare sent an Urgent Recall Notice dated April 2, 2007, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were asked to call 800-874-7378. For questions regarding this recall call 801-536-4952, |  | 17 | 17 | Worldwide Distribution - USA including MO, IL, LA, FL, NH, TX, IN, PA, NJ and the countries of France and China | Terminated | September 28, 2012 | 2006 | 17
Z-2181-2012 | 62596 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Kimberly Clark Patient Warming System - Model 1000 Control Unit. |  |  | April 27, 2012 | 2012 | Kimberly-Clark Corporation | A defect of the "fuse holder" can lead to a Control Unit failure that may cause the device to lose its ability to warm and maintain set temperature and may also result in excessive heat with the fuse assembly that could lead to the generation of smoke. | Device Design | Kimberly-Clark sent an "URGENT: FIELD CORRECTIVE ACTION NOTICE" dated April 27, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Corrective Action Response Form was attached for customers to complete and return. Contact the Field Corrective Action Coordinator at 770-587-7400 for questions regarding this notice. |  | 129 | 129 units | Nationwide Distribution-including the states of AZ, CA, CO, DE, FL, GA, IA, ID, IN, IL, KS, KY, LA, MA, MD, ME, MO, MS, NE, NM, OH, OR, PA, RI, SC, TN, TX, VA, WI and WV. | Terminated | September 05, 2012 | 131 | 129
Z-2172-2012 | 62448 | No | NSX |  |  | General Hospital | Contact ODE | EasyLink(TM) Informatics System, Part Number 1000034805. EasyLink(TM) Informatics System is classified as a data calculator and processing module for use with In-Vitro diagnostics devices. |  |  | June 07, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | There is a potential for the EasyLink system to display and transmit to the Laboratory Information System networks an incorrect result following the EasyLink systems application of the configured result number format. The problem is encountered when the result number format configured on the Sysmex CA-7000 or CA-1500 coagulation analyzer does not match the EasyLink system settings on the following | Software design | The firm initiated this recall on June 7, 2012 by issuing an "Urgent Field Safety Notice" to all customers who are known to have the EasyLink system and who also own and operate Sysmex CA-1500 and CA-7000 instruments. The Notice described the problem and provided recommended actions. Customers can contact 800-441-9250 for questions. |  | 237 | 237 | Worldwide distribution, including the countries of Australia, Austria, Belgium, Canada, Czech Republic, France, Germany, Italy, Netherlands, New Zealand, Slovakia, Spain, Switzerland and the United States, with nationwide distribution. Center Recall Depth is Retail. | Terminated | August 26, 2014 | 810 | 237
Z-2179-2012 | 62747 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medstream 81" (206 cm) 20 Drop Standard bore Universal Administration Set with Y-Site, Sterile |  |  | August 03, 2012 | 2012 | Medical Specialties Distributors, Inc. | Devices leaking at the Y-site | Process control | Medical Specialties Distributors LLC sent an Urgent Medical Recall letter dated August 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and remove all affected product from use immediately. Further distribution of the affected product should be discontinued immediately. Customers were instructed to contact their customers and inform them of the recall; complete and return the acknowledgement form and fax to 781-344-8320 to arrange for a return and to contact the Customer Service Department at (800) 967-6400 with questions. For questions regarding this recall call 781-344-6000. |  | 472 | 472 cases (50/case) | Nationwide Distribution | Terminated | January 23, 2015 | 903 | 472
Z-2165-2012 | 62525 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew 5.5 mm TWINFIX Ultra Ti Suture Anchor with two #2 ULTRABRAID sutures (blue, blue-COBRAID) with Needles |  |  | July 03, 2012 | 2012 | Smith &amp; Nephew, Inc. Endoscopy Division | Distal part of the anchor may break on insertion into bone during surgery | Device Design | Smith &amp; Nephew sent an Urgent Product Recall 1st Notification letters dated June 2012 via Federal Express to all affected customers. The letter identified the affected product, potential risk and actions to be taken. Customers were instructed to inspect inventory and quarantine them immediately, complete and return the attached form indicating the Lot# and quantities that need to be returned and contact Smith &amp; Nephew Returns Group at 800-343-5717 (option 3) or email to endo.andreturns@smith-nephew.com. for instructions for returning affected products and receiving a credit. For questions call 1-508-261-3731. |  | 913 | 913 units | Worldwide Distribution - US Nationwide and the countries of: Argentina, AUSTRALIA, AUSTRIA, Belgium, CANADA,CHILECOLOMBIA, Denmark, DUBAI, FINLAND, FRANCEGDC BAAR Distribution,GERMANY, GREECE, INDIAITALY, KOREA, MALAYSIA, MEXICO, NETHERLAND, NORWAYPORTUGAL, PUERTO RICO, SHANGHAI, SOUTH AFRICASPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | April 18, 2016 | 1385 | 12189
Z-2176-2012 | 62415 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Reliance AL Implant Blocker (Various sizes) |  |  | March 21, 2011 | 2011 | Stryker Spine | Stryker Spine has received six reports regarding the breakage of one or more prongs at the extremity of the instrument. | Component design/selection | Stryker sent an Urgent Product Recall letters/return response forms dated April 5, 2012, via FedEX to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and hospital location to identify the product. Reconcile any SLIDE Implant Blockers with catalog and lot numbers referenced above. Retrieve and return the affected product using the enclosed pre-paid mailing label to Stryker Spine 59 Route 17 South, Allendale, New Jersey, 07401. For questions customers were instructed to call 201-760-8298. |  | 109 | 109 Lots | Nationwide Distribution including NC, AR, TN, MA, IL, GA, NY, LA, MN, MO, OH, KS, NM, TX, KS,and IN. | Terminated | March 21, 2013 | 731 | 109
Z-2175-2012 | 62609 | No | LEH | Vancomycin test system. | Toxicology | Toxicology | 510(k) | Dimension Vista(R) Vancomycin Flex(R) Reagent Cartridge (VANC, Catalog # K4086). |  |  | July 17, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has received complaints of customers experiencing slopes &gt;1.05 when calibrating Flex(R) reagent cartridge lot 12095AD. This results in a slope failure and prevents the VANC assay from being run on the Dimension Vista(R) System. | Nonconforming Material/Component | The firm initiated their recall of this product on July 17, 2012 by sending an Urgent Field Safety Notification via FedEx next day delivery to all consignees. The Notice informed consignees of the high calibration slopes with the affected device. Affected consignees were requested to immediately discontinue use and discard any remaining Dimension Vista(R) VANC Flex(R) reagent cartridges, lot 12095AD. Affected customers were provided with replacement cartridges. Customers were to also complete the attached form and fax it back to the number provided. The notification letter should be forwarded to anyone, whom customers may have further distributed the affected lot. Questions and concerns should be directed to Siemens Technical Solutions Center at 800-441-9250. |  | 1894 | 1,894 cartons | Worldwide Distribution -- US, including Washington, DC and the states of AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WV, and the countries of Canada, New Zealand, Germany, and Puerto Rico. | Terminated | January 14, 2014 | 546 | 1894
Z-2182-2012 | 62611 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) Reagent Management System (RMS) - Power Cord. |  |  | July 17, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed that partial insertion of the power cord on the Dimension Reagent Management System (RMS) power input module can cause the plug to overheat and potentially create a fire hazard. RMS without a power cord retention clip may be affected. | Device Design | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated June 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to complete and return an attached Field Correction Effectiveness Check Form and fax it to (302) 631-8467. Contact Siemens Technical Solutions Center at 800-441-9250 for questions regarding this notice. |  | 2554 | 2,554 | Distributed in the US including Washington, DC, Puerto Rico and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WI, WV, and WY and the Virgin and Pacific Islands. | Terminated | April 03, 2014 | 625 | 2554
Z-2180-2012 | 62341 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics Angiographic catheter --- Soft-Vu (JB-1) Non-braided. --- Size: 5F; Length: 100cm; designed to accept guidewire: .035 in. --- Catalog No. [REF]: 10734201. Packaged in boxes of 5 catheters each, with manufacturer listed as AngioDynamics, Inc., Queensbury, NY USA. |  |  | June 15, 2012 | 2012 | Angiodynamics Worldwide Headquarters | AngioDynamics is recalling one lot of Soft-Vu Angiographic catheters because one of the catheter pouches within the outer box may be mislabeled. | Labeling mix-ups | The firm, AngioDynamics, Inc., sent an "Urgent Medical Device Recall - Immediate Action Required letter dated June 15, 2012 to all consignees/customers via FedEx. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to carefully read this recall notification in its entirety; IMMEDIATELY DISCONTINUE USE OF AND SEGREGATE RECALLED PRODUCT in a secure location for return to AngioDynamics, Inc.; immediately remove recalled product from their inventory; immediately forward a copy of this recall notification to all sites to which they have distributed affected product; and promptly complete, sign and return the Reply Verification Tracking Form (even if you do not have any product to return) via fax to Soft-Vu Angiographic Catheter Recall Coordinator at 1-518-798-1360 or email rdenino@agniodynamics.com. If affected product is located in your institution, please call AngioDynamics, Inc. Customer Service at 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday - Friday: Eastern Standard Time) to obtain a replacement. The consignees/customers were also instructed to package and return the recalled product, write the RMA number on the box and send to AngioDynamics, Inc., 603 Queensbury Avenue, Queensbury, NY 12804; Attn: Soft-Vu Angiographic Catheter Recall Coordinator.If you have any questions about this recall action, contact your local Sales Representative or AngioDynamics, Inc. Customer Service Manager at 800-772-6446 or by email at customerservice@angiodynamics.com |  | 245 | 7 boxes/35 units | Worldwide distribution: USA (nationwide) including states of: CA, KS, MO, and NC and countries of: Czech Republic and Switzerland. | Terminated | February 10, 2016 | 1335 | 245
Z-2171-2012 | 62533 | No | MYT | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-Stat CK-MP cartridges |  |  | February 01, 2012 | 2012 | Abbott Point Of Care Inc. | The frequency of suppressed results for i-Stat cTnl, BNP and CK-MP cartridges is affected by atmospheric pressure. | Device Design | Abbott Point of Care (APOC) sent an URGENT RECALL NOTICE dated February 2012 to all affected customers via e-mail and also contacted by phone. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that if their facility was performing testing at more than approximately 7500 feet above sea level, they should have an alternate test method available should they experience an issue. If they have forwarded any cartridges to another facility, they should provide a copy of the Urgent Recall Notice to them. For questions customers were instructed to call 800-366-8020, Option 1 or contact their Abbott Point of Care representative.For questions regarding this recall call 609-454-9000. |  | 400874 | 368,149 cartridges US; 32,725 OUS | Worldwide Distribution - USA including AZ and CO and internationally to Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Chile, China, Colombia, Costa Rica, Dominican Republic, El Salvador, Finland, France, Georgia, Germany, Greece, Honduras, Hong Kong, India, Ireland, Italy, Korea, Kuwait, Mauritius, Mexico, Netherland (Antilles), Nigeria, Norway, Oman, Paraguay, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, R.O.C., Thailand, UAE, United Kingdom, and Yemen. | Terminated | February 08, 2013 | 373 | 9289349
Z-2177-2012 | 62444 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 base 100/120V |  |  | July 26, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems sent a Safety Advisory (AA-2010-007-C) on December 2, 2010 (Recall# Z-0882/0883-2011 RES 57416) regarding malfunctions of the Occluder and Air Bubble Detection (ABD) Modules to a confined number of Terumo Advanced Perfusion System 1 users. The investigation of the malfunction found a faulty component on the modules' application board caused the failures. Terumo CVS re | Component design/selection | The firm, Terumo CVS, sent an "URGENT MEDICAL DEVICE RECALL" letter dated July 26, 2012 to its customers. The letter described the product, problem and actions to be taken. To customers were instructed to review this Medical Device Recall notice and addendum; assure that all users are aware of this notice and addendum; place the addendum in the front of Section 7: Modules (p 7.1) and confirm receipt of this communication by faxing the attached Customer Response Form to the fax number/email address indicated on form. Terumo CVS is issuing a Safety Advisory to alert all Terumo System 1 users of the potential for failure and providing an addendum for the Operator Manual with instruction on how to respond to the failure.If you have any question or concerns contact Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours are Monday-Friday, 8AM - 6PM ET. |  | 992 | 992 units | Worldwide distribution: USA (nationwide) and countries of: ARGENTINA, AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Japan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | November 16, 2012 | 113 | 1641
Z-2173-2012 | 62403 | No | NSX |  |  | General Hospital | Contact ODE | Siemens syngo(R) Lab Data Manager. |  |  | June 07, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | The recalling firm has discovered that when using a Sysmex(R) CA-7000 or CA-1500 coagulation analyzer with the syngo Lab Data Manager (LDM), there is a potential for the syngo LDM system, to display and transmit to the LIS an incorrect result following the syngo LDM's application of the configured result number format. The problem is encountered when the result number format configured on the Sys | Software design (manufacturing process) | Siemens Healtcare Diagnostics sent a Urgent Field Safety Notice dated June 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised to discuss the contents of this letter with your Laboratory Director regarding the need to review any previous test results.If you have any technical questions or concerns, please contact the Seimens Technical Solutions Center at 1-800-441-9250, for further assistance. |  | 15 | 15 | USA (nationwide) including the states of AZ, CA, MS, MO, NJ, NC, OH, OR, PA, and TX. | Terminated | April 01, 2014 | 663 | 15
Z-2187-2012 | 62499 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | The NCB Polyaxial Locking Plate System NCB screws 26mm: 02.03155.026, non-sterile |  |  | July 13, 2012 | 2012 | Zimmer, Inc. | Zimmer investigation following a complaint found one lot of NCB Polyaxial Locking Plate System NCB screws, 26mm, distributed in the United States contained incorrect 32 mm screws. | Labeling mix-ups | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE RECALL" letter dated July 13, 2012, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to locate the implicated lot in their inventory and user accounts; return all packaged units of this lot; and provide Zimmer with the name, complete mailing address, and phone number of each user facility and surgeon that have used the product via E-mail to dawn.kindle@zimmer.com. Zimmer will send a follow-up Urgent Recall notification to these sub-accounts with instructions to verify the screw length of any screws already in the implant rack inventory prior to use in surgery. All implicated product is to be returned. The customers were also instructed to complete and return the Inventory Return Certification form to Zimmer via fax at +1-574-372-4265. If you have any questions about this correction and/or product removal, contact Zimmer at +1-877-946-2761. |  | 31 | 24 to US Distributor with 7 further distributed | Nationwide distribution: USA including states of: AZ, IL, IN, ME and NJ. | Terminated | September 24, 2013 | 438 | 31
Z-2186-2012 | 62598 | No | GBW | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TAUT, Cholangiogram Catheter Intraducer, Peritoneal Catheter, Teleflex Medical, Research Triangle Park, NC. |  |  | July 20, 2012 | 2012 | Teleflex Medical | Product is missing the label on the internal packaging. | Packaging process control | Consignees were notified by letter on 07/20/2012 |  | 611 | 611 units | AK, AL, AZ, CA, CO, FL, GA, IA, IL, IN, KS, LA, MA, MI, MO, MS, NC, NY, OH, PA, SD, TX, UT, VA, WI, WV. | Terminated | February 12, 2015 | 937 | 611
Z-2184-2012 | 62646 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Cardiovascular (CV) Reporting |  |  | July 27, 2012 | 2012 | AGFA Corp. | Content entered into the "Conclusions" free text box on the Report Writer screen was not represented on the final printed report. | Software change control | AGFA Healthcare sent an URGENT FIELD SAFETY NOTICE dated July 27, 2012, to all affected customers. The letter described the safety alert and mitigation. Customer were asked to complete the Acknowledgment form and fax back to 864-421-1664 that the information was received and understood.For questions customers were instructed tol call 401-604-2180.For questions regarding this recall call 864-421-1754. |  | 280 | 280 | Worldwide Distribution - USA (nationwide) and Canada | Terminated | January 13, 2015 | 900 | 280
Z-2174-2012 | 62559 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Shiley Reusable Cannula Low Pressure Cuffed Tracheostomy Tubes, model numbers (Worldwide) 8LPC, 8FEN; (Outside U.S.) 8FENJ, 8FENJ-S, 8LPC-S, 8FEN-S. |  |  | July 18, 2012 | 2012 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is recalling certain cuffed Shiley tracheostomy tubes due to consumer complaints related to the connections and functions of the inner and outer cannula which were observed during mechanical ventilation of patients. | Device Design | Covidien sent an Urgent Medical Device Recall letter dated July 18, 2012 to all affected consignees via mail (FedEx) notification. The customer communication identified the affected product , problem and actions to be taken. Customers were instructed to examine their inventory, segregate product with the identified affected product codes and lot numbers and complete the enclosed verification form and fax to 925-463-4600. Customers were asked to return unused products from the affected product codes and lots. The firm also recommended the replacement of affected tubes that were already in use as soon as clinically appropriate. The letter stated to return the affected product for credit or question contact Technical Services group at 800-635-5267 option 3, option 1 and reference the Return Goods Authorization number. |  | 337227 | 337,227 units | Worldwide Distribution - US Nationwide | Terminated | January 22, 2015 | 918 | 337227
Z-2183-2012 | 61855 | [6 empty] | BeautyAngel Health Fitness Natural Collagen Stimulation Anti Aging |  |  | May 07, 2012 | 2012 | JK Products &amp; Services, Inc | The Beauty Angel was marketed without a 510k. | No Marketing Application | JK Products sent a Mandatory Field Correction Notice - Immediate Action Required letter dated February 10, 2012, to all affected customers. The letter identified the product the problem, and the action to be taken by the customer. The firm instructed customers that they would be replacing the operating instructions for the affected product with new ones that meet the requirements defined by the regulatory authorities. Customers were instructed not to make claims that indicate that using the affected product will result in medical or health benefits, or that directly or indirectly imply that the affected product is intended to cure, treat, diagnose or pevent an illness, or disease. Customers were asked to complete the new owner information request on the return post care enclosed and return to the firm. If product was further distributed those accounts should be contacted immediately and should be advised of this Mandatory Field Correction Notice. For questions customers should call 870-935-1130. For questions regarding this recall call 870-268-2702. |  | 598 | 598 units | Worldwide Distribution - USA (nationwide) and Canada | Terminated | January 07, 2015 | 975 | 598
Z-2188-2012 | 62278 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Swan-Ganz Thermodilution VIP Catheters Model 131F7J. |  |  | May 11, 2012 | 2012 | Edwards Lifesciences, LLC | The device is being recalled because the firm received multiple complaints related to passage of a J-tip guidewire when inserted through the distal lumen hub of the catheter. | Process control | Edwards Lifesciences sent a "PRODUCT RECALL" letter dated May 8, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Confirmation Form was attached for customers to complete and return via fax to 1-800-422-9329. Contact Customer Service at 1-800-424-3278 for questions regarding this notice. |  | 32145 | 32,145 units for all products in Recall Event | Worldwide Distribution-USA (nationwide) and the countries of Canada, Japan, EMEA (Europe, Middle East and Africa), Asia Pacific, (Hong Kong, Indonesia, Malaysia, Australia, and Taiwan, Thailand), Latin America, (Chile, Colombia, Panama, Mexico, and Honduras). | Terminated | June 17, 2014 | 767 | 32145
Z-2202-2012 | 62310 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-1 System Battery --- Made in U.S.A. for Carestream Health, Inc., Carestream Health, Inc., 150 Verona St, Rochester, NY 14608. --- The battery may be sold with the DRX-1 imaging device or as a stand-alone accessory. |  |  | June 12, 2012 | 2012 | Carestream Health, Inc. | Carestream DRX-1 System batteries manufactured before June 2012 due to two occurrences of lithium-ion battery overheating, which may cause burns. | Nonconforming Material/Component | Firm sent consignees "Urgent: Customer Notification and Field Corrective Action" letters on June 8, 2012. The letters described the issue and provided recommended actions. Service dealers will contact all affected consignees to evaluate the battery. U.S. customers with questions can contact Carestream Customer Care at 800-328-2910. Customers outside the United States should contact their local Carestream service number. |  | 9971 | 9971 units | Worldwide Distribution, including Nationwide (USA). | Open, Classified |  |  | 9971
Z-2200-2012 | 62588 | No | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | 13.5 mm x 50 mm BrainPath Sheath |  |  | July 17, 2012 | 2012 | Nico Corp. | BrainPath sheaths were mislabeled. Sheaths which measure 50mm were labeledas 60mm. The length callout on the label does not match the actual length of the sheath within thepackaging resulting in an improper fit as further explained below. This mislabeling wasdiscovered on 29-June-2012 during a clinical case. | Labeling mix-ups | NICO Neuro and Spine sent an Urgent Recall letter to all affected customers on July 18, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the recall response form.Affected sheaths distributed within the US will be obtained and discarded by NICO. Affected sheaths shipped to Canada will be obtained and discarded by NICOs distributor, Galen Medical.For questions customers should call 888-632-7071. |  | 100 | 20 devices (4 boxes, 5 devices per box) | Worldwide Distribution - USA including MI and Canada | Terminated | November 07, 2012 | 113 | 175
Z-2203-2012 | 61614 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare Advantage Workstations, Advantage Worksation version 4.2 and 4.3 with Volume Viewer 2 (version 6.4-6.8) provided with Advance Vessel Analysis (AVA) and Advantage Workstation 4.4 with Volume Viewer 2 (version 6.8 or before) or Volume Viewer 3 (version 7.2 -7.3) provided with Advanced Vessel Analysis (AVA), Model 2378698-2 |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare is informing you of incorrect measurements when using the Aorta analysis protocol or customized protocols. Until a new software version is installed they are asking you not to use the aorta analysis protocol or the customized protocols for an aneurysm.A second issue was also discovered internally. When loading PET data into the Volume Viewer, ifslices are missing in a PET serie | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 23, 2012 to consignees. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers &amp; Radiologists. The letter described the affected product details, safety issue, safety instructions and product correction. For more information on this topic, contact a GE Application Specialist or your GE Field Service representative.For questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative |  | 1055 | 1055 | Worldwide Distribution - US (nationwide) including the states of : AL, AZ, AR, CA, CO, DE, DC, FL, GA, GU, HI, IK, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO,MT, NE,NH,NJ, NM, NY, NC, ND, OH, OK, PA, PR, SC, SD, TN, TX, VA, WA, WV, WI and the countries of YUGOSLAVIA, VIETNAM, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, UKRAINE, TURKEY, TUNISIA, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAÂżDE ARABIA, RUSSIAN FEDERATION, ROMANIA, REPUBLIC OF MALTA, PORTUGAL, POLAND, PHILIPPINES, NORWAY, NIGERIA, NEW ZEALAND, MEXICO, KOREA, KENYA, JAPAN, ITALY, IRELAND, INDONESIA, INDIA, HUNGRY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, EGYPT, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BULGARIA, BRAZIL, BELGIUM, BELARUS, AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA. | Terminated | December 11, 2012 | 201 | 1055
Z-2199-2012 | 62589 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | SenTec, V-Sign"2 Sensors |  |  | July 09, 2012 | 2012 | Sentec AG | SenTec is recalling a number of V-Sign"2 Sensors due to a problem regarding configuration parameter writing during manufacturing. | Device Design | SenTec sent a Field Safety Notice dated July 12, 2012, via email to all affected customers. It identified the product the problem and the action needed to be taken by the customer. Customers were instructed the affected Sensors in the field must be replaced.Please locate the affected Sensors and return them to SenTec for analysis immediately. SenTec will provide you with replacement Sensors within the next three months.This notice needs to be passed on to all those who need to be aware within your organization. Please transfer this notice to other organizations on which this action has an impact.Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. Further questions please call (262 ) 594-2114. |  | 8 | 8 | Worldwide Distribution -- Nationwide including the state of Missouri,. and the countries of Germany, Spain, France, Netherlands, Australia, Belgium, Thailand. | Terminated | October 11, 2012 | 94 | 8
Z-2185-2012 | 55121 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump Module model 8100. |  |  | June 29, 2012 | 2012 | CareFusion 303, Inc. | The recall was initiated because Carefusion identified a potential risk associated with the Alaris Pump module model 8100. Carefusion has received reports of customers experiencing motor stalls during infusion with Alaris Pump Module (model 8100) manufactured between August 2010 and July 2011. | Under Investigation by firm | CareFusion sent an Urgent Medical Device Recall Notification letter dated July 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Required action for users instructed high risk patients undergoing infusions at high rates, consider having additional devices as back up. Clinicians should weigh the risk/benefit to the patients before continuing to use the device. Customers with recall related questions were instructed to contact Carefusion support center at (888) 562-6018. Customers with adverse event report questions were instructed to contact customer advocacy at (800) 854-7128, option 1, option 1, option 3 or email at customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris system were instructed to contact Technical support at (888) 812-3229.For questions regarding this recall call 858-617-4000. |  | 102248 | 102,248 units | Worldwide Distribution - USA (nationwide) and the countries of: South Africa and Canada.Additional Countries: Netherlands. | Terminated | August 26, 2015 | 1153 | 102248
Z-2209-2012 | 62586 | No | JJF | Micro chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COBAS AmpliPrep Sample Tube Input Barcode Clips, catalog number 03137040001. Input S-tube barcode clips hold sample input tubes in the COBAS AmpliPrep Instrument (CAP) sample racks and have a barcode to identify the sample. |  |  | February 01, 2012 | 2012 | Roche Molecular Systems, Inc. | Certain lots of COBAS AmpliPrep (CAP) S-tube Input (Box of 12 bags 12x24 S-tube and 12x24 S-clip) can cause the sample input tubes to stick in the barcode clips. As a consequence, the CAP instrument will lift both the sample input tube together with the barcode clip from the sample rack, resulting in an aborted run. No erroneous results would be produced if the issue occurs, but samples in proce | Nonconforming Material/Component | On 1/17/2012 an "Analyzer Bulletin" notification was sent to Roche affiliates and on 2/1/2012 an "Analyzer Bulletin" notification was sent to all affected US Consignees. The notification described the product issue and provided recommended actions. Customers were referred to Roche Technical Support at 800-526-1247 for questions. |  | 2132 | 2,132 boxes | Nationwide Distribution (USA). | Terminated | July 16, 2013 | 531 | 2132
Z-2211-2012 | 62612 | No | LFM | Digitoxin test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products DGXN Slides --- [REF] 834 3386. Product is packaged in a box with multiple slides. The lot and expiration codes are found on the side of the package. |  |  | June 28, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics revised the VITROS Chemistry Products DGXN Slides Instructions for Use (IFU) and removed heparin plasma as a recommended specimen type for VITROS DGXN Slides | Process control | Ortho Clinical Diagnostics sent an Urgent Field Safety Notice dated June 28, 2012, via US Postal Service Priority Mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: Immediately discontinue using heparin plasma samples and use only serum samples when processing VITROS DGXN Slides; Notify their laboratory and medical staff of this change and forward the information if the product have been distributed outside of your facility; Replace the pages of their VITROS Instructions for Use Manual with the enclosed Instructions for Use and Table of Contents; and Discontinue using any copies of the VITROSÂż DGXN Slide IFU (Version 7.0).--- On 28 June 2012 an Urgent Field Safety Notice was sent via US Postal Service Priority Mail to all affected Distributors. They were advised of the issue and instructed to discontinue distributing and discard their remaining inventory of VITROS DGXN Slides. --- Foreign affiliates were informed by email on 28 June 2012 of the issue and instructed to notify their consignees of the issue and actions. The letters included a Confirmation of Receipt form to be completed and return. For questions contact Customer Technical Servives at 1-800-421-3311 |  | 13338 | 13338 sale units | Worldwide Distribution - US (nationwide) | Open, Classified |  |  | 13338
Z-2061-2012 | 61747 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Neptune 1 Silver Rover, Part 0700-003-000 Stryker Instruments Neptune Rover |  |  | June 08, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | Stryker has received two reports of serious injury as a result of tissue damage associated with the use of the Neptune 2, including an event in which one customer connected the Neptune 2 System to a passive chest drainage tube post operatively, resulting in a fatality. The current IFU did not specifically warn against connecting all Neptune Waste Management devices, which is a high vacuum/high flow device, to a passive drainage tube. June 12, 2012, customers were instructed to review the revised IFU, distribute to affected departments, and educate users of the Neptune on this warning. Warning Labels, which are pending distribution, are required on all Neptune devices. Further investigation found the Neptune 1 Silver, Neptune 2 Ultra (120V) and Neptune 2 Ultra (230V) lacked FDA 510(k) clearance and therefore are not legally marketed without determination of their safety and effectiveness. FDA advises Neptune 1 Silver, Neptune 2 Ultra (120V) and Neptune 2 Ultra (230V) not be used. However, customers who do not have an alternative device to use should weigh the risks and benefits associated with continued use of these devices. If customers choose to continue use of the Neptune Silver, Neptune 2 Ultra (120V) or Neptune 2 (230V), they must complete a Certificate of Medical Necessity and return it to Stryker by October 12, 2012. | Labeling False and Misleading | On June 5, 2012, Stryker notified customers that it was recalling the IFUs for the above products. The current IFU did not specifically warn against connecting the Neptune Rover, which is a high vacuum/high flow device, to a passive drainage tube. Customers were instructed to review the revised IFU, distribute to affected departments, and educate users of the Neptune on this warning. Customers must confirm with Stryker via business reply form that they have completed these actions. Customers who have the Neptune 1 Gold, Neptune 1 Gold International or Neptune 2 Bronze will receive a follow up mailing containing warning labels for the device and instructions detailing how to apply them. Customers may continue to use the Neptune 1 Gold, Neptune 1 Gold International, and the Neptune 1 Bronze without any further actions.On September 18, 2012, Stryker notified customers via letter delivered by FedEx overnight delivery that it is expanding the recall on the Neptune 1 Silver, Neptune 2 Ultra (120V) and Neptune 2 Ultra (230V) because FDA has also advised Stryker that these devices require, but do not currently have, 510(k) clearance. FDA is therefore unable to determine whether these devices are as safe and effective as their legally marketed predicate, the Neptune 1 (Gold) Waste Management System (510(k) K012992). As such, Stryker has ceased distribution of the Neptune Silver, Neptune 2 Ultra (120V) and Neptune 2 Ultra (230V) devices until FDA clears these devices.At this time, FDA does not consider the Neptune Silver, the Neptune 2 Ultra (120V) or the Neptune 2 Ultra (230V) to be legally marketed devices because their safety and effectiveness have not yet been determined. As such, FDA advises that the devices not be used. However, customers who do not have an alternative device to use should weigh the risks and benefits associated with continued use of these devices. If customers choose to continue use of the Neptune Silver, Neptune 2 Ultra (120V) or Nept |  | 2657 | 2657 | Worldwide Distribution -- USA (nationwide) and country of: Canada.. | Terminated | August 23, 2016 | 1537 | 20329
Z-2210-2012 | 62471 | No | NCP | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet Getinge Group Pediatric Arterial Cannula |  |  | May 29, 2012 | 2012 | Maquet Inc. | Certain pediatric arterial cannulas may have limited flexibility due to an increased length of the "hardened" cannula tip. | Device Design | Maquet issued Urgent Device 2, Recall letters/Urgent Product Recall Forms dated May 29, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their current stock and identify any affected product and remove from their inventory. Customers were instructed to contact the customer service department at 800-777-4222 (option #2, followed by option #2). Customers were also instructed to completed the enclsoed Recall Form to acknowledge receipt of the notification and to document the current status of the devices shipped to their facility. For questions regarding this recall call 973-709-7660. |  | 142 | 20 devices US; 122 devices foreign | Worldwide Distribution - USA including MI, MN and Internationally | Terminated | April 03, 2014 | 674 | 142
Z-2208-2012 | 62568 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Philips Easy Diagnost Classic-Stationary x-ray system |  |  | July 11, 2012 | 2012 | Philips Healthcare Inc. | Inclination switch in the stand may have been mounted incorrectly, as a result, the switch could not stop the tilting movement of the table base due tol hardware or software errors. | Equipment maintenance | Philips issued an Urgent Field Safety Notice dated July 9, 2012, to notify users that the inclination swtich may be mounted incorrectly. A Field Safety Engineer will be dispatched to verify if a switch was installed properly, and in the event it was not installed properly they will re-install it properly. The Field Service Engineer will add a label as described in the spare part replacement manual, which is part of the FCO kit.. Customers were instructed review the information contained in the Field Safety Notice and to retain a copy of the notice with the equipment Instruction for Use. For further questions customers were instructed to call their Phillips representataive Technical Support Line at 1-800-722-9377. For questions regarding this recall call 978-687-1501 |  | 12 | 12 units | Worldwide Distribution - USA including IN, MA, MI, MO, NC, NY PA, SC, and WI and to the countries of Canada, HongKong, New Zealand. | Open, Classified |  |  | 12
Z-2207-2012 | 62650 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Essenta DR, X-Ray System, Diagnostic, general-purpose |  |  | July 25, 2012 | 2012 | Philips Healthcare Inc. | Essenta DR C-Arm may fall | Equipment maintenance | Philips Medical sent an "URGENT FIELD SAFETY NOTICE" dated July 09, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers are asked to inspect the C-Arm unit and discontinue use if not parallel to the column, do not use the system and call for a Service Technician. A Field Change Order (FCO) is being released that requires an inspection of C-ARM connection and, if necessary, repair of systems. The bolts will be disconnected and thread inserts will be checked. If not the thread inserts are no longer secured, additional holes will be drilled to use in place of the original ones. For further information or support concerning this issue, contact your local Philips representative at 1-800-722-9377. |  | 7 | 7 units | Worldwide Distribution-USA including the states of CA, NC, and TN and the countries of China, Germany, and Hungary. | Terminated | March 01, 2013 | 219 | 7
Z-2215-2012 | 62540 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | 00-4309-029-01 - Trabecular Metal Reverse Shoulder Liner 36 mm 65ÂżÂż Neck Angle |  |  | July 18, 2012 | 2012 | Zimmer, Inc. | Lot specific Trabecular Metal (TM) Reverse Shoulder Liner Impactors are being recalled due to reports of alignment peg fractures. All lots of TM Reverse Shoulder Liner Impactors manufactured from 455 stainless steel are being retrieved, as all reported fractured instruments were manufactured with thismaterial. The TM Reverse Shoulder Liner Impactor units that were manufactured with 13-8stain | Device Design | Zimmmer sent an Urgent Medical Device Recall letter dated July 18, 2012, to all affected customers. The letter identified the product, the problem, and thea action to be taken by the customer. Customers were instructed to inspect all TM Reverse instrument sets for affected lot(s)(instruments. If affected product was found customers were instructed to immediately place an order to backfill the identified product with Customer Service. If an instrument was was further distributed customers were to provide a copy of the Notice to ensure their customers were aware of the removal. Customers were also instructed to provide a list of all hospitals that currently use the affected product. For questions customers were instructed to call 800-348-2759.For questions regarding this recall call 574-371-1689. |  | 212 | 212 (179 USA) | Worldwide Distribution - USA (nationwide) and internationally to Germany, Belgium, Denmark, Finland, France, Italy, Netherlands, Spain, Sweden, Switzerland, and the United Kingdom. | Terminated | April 17, 2014 | 638 | 724
Z-2217-2012 | 62766 | [6 empty] | DePuy Custom Implant Devices |  |  | July 31, 2012 | 2012 | DePuy Orthopaedics, Inc. | In connection with a Warning Letter received from the U.S. Food &amp; Drug Administration on December 9, 2011, DePuy Orthopaedics is taking steps to retrieve custom device implants that surgeons may have ordered from DePuy Orthopaedics, but not implanted. DePuy Orthopaedics received a Warning Letter from the FDA on December 9, 2011 regarding custom device implants. DePuy Orthopaedics has discontinue | PMA | DePuy sent a written communication letter dated July 31, 2012 to all consignees. The communication to the hospital customers who received the devices includes a "Surgeon Letter" with instructions to provide the surgeon letter to the Surgeon users. All DePuy Distributors will be notified via email. The sales representatives notified the hospital and surgeons either by mail or in person with a written communication. The letters identified the affected product, problem and actions to be taken. Customers are instructed to complete the attached Roponse form and return to fax number 574-371-4939 or email: kseppa@its.jnj.com. The sales representatives are expected to aid customers in returning of unused Custom Device implants, if needed. The devices can be returned through the normal DePuy Returns process, to attention of "Returns" and marked "Custom Returns" on the outside of the box.Questions Regarding the RemovalQuestions regarding the removal may be directed to the following DePuy representatives -Questions about removal information provided: Katie Seppa, Manager of Post-Market Surveillance, 574-372-7333 (M-F; 8 am  5 pm EDT)-Clinical related questions from surgeons: Direct to DePuys Scientific Information Office at 1-888-554-2482 (M-F; 8 am-5 pm EDT) |  | 8338 | 8338 | US Nationwide Distribution - including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, HI, IA, ID, IL, IN, KS, KY, LA, MA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | January 13, 2014 | 531 | 8338
Z-2216-2012 | 62344 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory SMART &amp; Sunquest Laboratory SMART Select |  |  | April 28, 2011 | 2011 | Sunquest Information Systems, Inc. | In Online Entry (OEx) there are two scenarios where a Test result from one container may file to a Test on another container. In order for this to occur:1) Site parameter SIT 1,2,1,26 (3,217) [Restrict resulting of tests to owner container ( | Software design | Sunquest sent an Urgent Product Safety Notice dated April 27, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. For the 103 sites that still have this issue Sunquest distribute a new maintenance release. Customers were instructed to use LabAccess Results Workstation (LARS) instead of Online Entry when entering results for instrument method codes that are defined to use replicate tests that are typically ordered on the same accession. Customers were also instructed to share this notification with appropriate laboratory personnel. Customers were questions should call 520-570-2000. For questions regarding this recall call 520-570-2252. |  | 103 | 103 sites | Nationwide Distribution | Terminated | September 04, 2012 | 495 | 103
Z-2218-2012 | 62732 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 8000 Base |  |  | June 20, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During a retrospective review of complaint events, TCVS identified on 06/15/12 that three (3)Sarns" System 8000 bases were manufactured and distributed from 11/13/06-11/29/06 with anincorrect incorrect amperage rating identified on the name plate label. The mislabeling wasidentified in January 2007 a set up process of the automark label printer. All units were correctedin the field in Marc | Labeling mix-ups | Terumo field correction/service records were completed in 2007 and all affected units were corrected in the field in March 2007 by replacing the incorrect label at the customer location. For questions about this report call 734-741-6056. |  | 3 | 3 | US Nationwide Distribution - including the state of New York | Terminated | November 07, 2012 | 140 | 3
Z-2206-2012 | 62592 | No | MTJ |  |  | General &amp; Plastic Surgery | 510(k) | Synthes(R) Hemostatic Bone Putty, Part numbers 08.901.001.97S, 08.901.001.98S, 08.901.001.99S, 08.901.001D, and VB1025.10S. |  |  | July 05, 2012 | 2012 | Synthes USA HQ, Inc. | There is the potential for Hemostatic Bone Putty to ignite if contacted with electrosurgical cautery systems under certain conditions during surgery. | Labeling design | The firm initiated their recall of this product on July 5, 2012 by sending a "Notice: Medical Device Recall" letter to their consignees. The letter identified the affected product along with the reason for recall; and, requested that consignees immediately cease using the product, examine their inventory, and remove the product from inventory. The letter further instructed consignees to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, and to complete and return the verification form that was included with the letter. Questions should be directed to 610-719-5450. |  | 8853 | 8,853 packs | Worldwide Distribution -- US, including Washington, DC and states of AK, AZ, CA, CT, FL, GA, ID, KS, LA, MD, MA, MI, MN, MO, NE, NH, NY, NC, OH, OR, PA, SC, TX, VA, WA, and WI, and the country of Canada. | Terminated | November 06, 2015 | 1219 | 8853
Z-2220-2012 | 62649 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Altrx" Altralinked" Polyethylene liners |  |  | July 24, 2012 | 2012 | DePuy Orthopaedics, Inc. | In the recalled lot, the products were not machined according to the production router, a size 52 cup (122136052) was machined instead. This resulted in a size 52 liner being etched and labeled as a size 54 liner. | Employee error | Depuy sent an URGENT MEDICAL DEVICE RECALL notification dated July 24, 2012, to all affected consignees. The recall is extended to the Depuy Distributor level and hospital level. All DePuy Distributors were notified via email. The sales representatives notify the hospital either by mail or in person with a written communication.Distributors and customers were instructed to cease use or distribution of recalled products and to return them to DePuy in Warsaw, IN. The sales representatives are expected to aid customer in returning of recalled products. Distributors and sales representatives are to return the recalled products through the normal DePuy Returns process, to attention of Returns and marking H12-16 on the outside of the box. Reconciliation forms should be returned to the DePuy Sales Representative or faxed to 574-372-7567. For clinical questions from Surgeons contact DePuy's Scientific Information Office at 1-888-554-2482. For questions regarding recall information call 574-372-7333. |  | 20 | 20 liners | Worldwide Distribution - USA including PA and the country of Canada | Terminated | April 30, 2013 | 280 | 20
Z-2219-2012 | 62581 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Version 9.5 |  |  | July 13, 2012 | 2012 | Siemens Medical Solutions, USA, Inc | If multiple uncompressed images/clips are sent to the syngo Dynamics server, around the same time, it is possible for an image from a patient to be placed in another patient's study. | Software design | On 07/13/2012, Siemens issued a "Customer Safety Advisory Notice" to affected consignees according to Update Instruction SY062/121S. The notice informs the customers of the potential issue with the identified affected product and provides details regarding the resolution. A software update to relieve the symptom of the issue will be distributed and installed by a Siemens Service Engineer on all affected systems under Update Instruction SY063/121S. A software update to correct this issue will be distributed and installed by a Siemens Service Engineer on all affected systems under Update Instruction SY064/121S. |  | 30 | 30 | Nationwide Distribution -- MA, GA, MO, IA, CA, PA, OR, CO, NY, NJ, WA, UT, OH, WI, MI, KY, and NC. | Terminated | July 15, 2013 | 367 | 30
Z-2222-2012 | 62576 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Conformis iTotal CR-Cruciate Retaining Knee Replacement System |  |  | July 17, 2012 | 2012 | Conformis Inc | Addendum Labeling iTotal CR Surgical Technique with Detailed Instructions of Total CR Poly Insertion Technique | Device Design | ConforMIS sent an Urgent Field Safety Notice letter dated July 5, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. An Addendum is included that provided updates to the Surgical Technique Guide. The updated technique guide will be delivered to all iTotal surgeons along with an on-site vist from ConforMIS sales representative. Customers were instructed to place the Surgical Technique Guide Addendum alongside the surgical plan, it puts the information in a highly visible location for the surgeon to review prior to the surgical procedure. Customers were asked to sign the communication verification form attached. For any questions regarding this notice, please contact your local ConforMIS representative or MDSS GmbH, Tel.: +49-511-62628630, info@mdss.com. |  | 1427 | 1427 units | Worldwide Distribution - US Nationwide and the country of Germany | Terminated | February 27, 2013 | 225 | 1427
Z-2205-2012 | 62642 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump Module Model 8100 (formerly Medley Pump Module) |  |  | June 15, 2012 | 2012 | CareFusion 303, Inc. | The recall was initiated because Carefusion has identified a potential risk associated with the Alaris Pump module model 8100. The pump module door keypad overlay may separate from the keypad assembly | Device Design | CareFusion sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated July 20, 2012. The letter identified the product, problem, and actions to be taken by the customers. The potential risk identified in the letter was that when the pump module door keypad overlay has delaminated (loose, peeled away or separated from the door assembly) this could cause a potential for fluid ingress which could lead to a keypad malfunction causing termination of infusion with alarm. Termination of infusion could result in serious injury or death. Customers were informed that CareFusion will contact their facility with in 60 days of receiving the notification to schedule a visit to replace the door assembly of their affected pump module. Customers with recall related questions were instructed to contact CareFusion Support Center at (888) 562-6018. Customers with adverse event reports questions were instructed to contact Customer advocacy at (800) 854-7128, option 1, option 1, option 3 or e-mail at customer feedback @carefusion.com. Customers with technical questions were instructed to contact technical support at (888) 812-3229. Customers were instructed to promptly complete and return the enclosed Customer Response Card to expedite the corrective action process. |  | 43862 | 43,862 units total (40,780 units in the US) | Worldwide Distribution-USA (nationwide) and the countries of Australia, European Union, and South Africa. | Terminated | May 13, 2014 | 697 | 43862
Z-2223-2012 | 62641 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance 64 and Ingenuity CT are whole body Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | July 18, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | 1. The current system design of software version 4.0.0xx379 does not prevent the system from scanning if the dose level exceeds a set upper dose limit for CT Brain Perfusion scanning, which previous versions did. However, the current version does have the Dose Check (NEMA XR-25) feature, which provides dose notifications and alerts, but it is possible for the user to disable Dose Check after ente | Software design | Philips Healthcare sent an URGENT - Medical Device Correction letter date July 18, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Phillips will be installing a software update version 4.0.0xx516. A Philips Service Engineer will contact the customer for implementation of the software update on the affected systems free of charge.For further information or support concerning this issue, customers may contact their local Philips representatives or Philips Healthcare office. For North America and Canada, they may also contact the Customer Care Solutions Center (1-800-722-9377, option 5)For questions regarding this recall call 440-483-7000. |  | 8 | 8 | Worldwide Distribution - USA including IN and OH, also, the countries of Australia, Belgium, China, Israel, and Sweden | Terminated | February 22, 2013 | 219 | 8
Z-2224-2012 | 62745 | device, prosthesis alignment | IQO | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | Oxford Gap Gauge SML 3/4mm. |  |  | July 27, 2012 | 2012 | Biomet U.K., Ltd. | A drafting error during an unrelated design change has led to 4 misleading dimensions on the drawing. The CAD modelremained the same thickness but the problem dimensions were accidentally struck from an incorrect point result in anincrease of 1.0mm in the value of each dimension | Device Design | Biomet sent a "FIELD SAFETY CORRECTIVE ACTION" letter dated July 24, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Response Form was attached for customers to complete and return. Contact the firm for questions regarding this notice. |  | 119 | 119 units | Worldwide Distribution-USA-including the states of IN, SD, and TX and the countries of Austria, China, Japan, Netherlands, Sweden, Turkey, and UK. | Terminated | February 15, 2013 | 203 | 119
Z-2221-2012 | 62798 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GEHC OEC 8800 Mobile Fluoroscopic X-ray System, GE Healthcare, Surgery, Salt Lake City, UT. |  |  | October 19, 2007 | 2007 | GE OEC Medical Systems, Inc | X-ray units did not meet manufacturer's specification for linearity. | Nonconforming Material/Component | GE updated their user manual and changed their internal specifications to match those of the CFR In addition, a software code correction was to be made to forward productions in order to correct the reported defect.*** This recall issued was previously classified under an incorrect recall number. It is being reclassified with a new number. *** |  | 175 | 175 units | Worldwide Distribution - USA - Australia, Brazil, Cameroon, China, Colombia, Cyprus, The Czech Republic, Finland, France, Germany, Greece, Italy, Japan, Korea, Malaysia, The Netherlands, New Zealand, Norway, Russian Federation, Spain, Sweden, Taiwan, Turkey, UAE, UK, and Canada. | Terminated | August 17, 2012 | 1764 | 175
Z-2225-2012 | 62755 | No | CKG | Adrenocorticotropic hormone (ACTH) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCel DxH Slidemaker Stainer Instrument System, PN 775222. |  |  | December 13, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the DxH Slidemaker Stainer Floor Stand shipped with Slidemaker Stainer instruments were missing vinyl caps, exposing sharpt edges on the DxH Slidemaker Stainer Reagent Cabinet (Floor Stand) drawer slides. The sharp edges on the drawer slides may expose operators to physical injury when replacing reagents and waste containers whic | Process change control | Beckman Coulter initiated a "PRODUCT CORRECTIVE ACTION (PCA)" which was communicated via phone calls starting the week of December 19, 2011 to all customers who purchased UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System. The phone calls provided the customers with an explanation of the problem identified, potential impact should the problem occur and an action to be taken. Customers were instructed to share the PCA information with their laboratory staff and if they had any questions, they should contact Beckman Coulter at (800) 526-7694. |  | 41 | 41 units (15 units in the US) | Worldwide Distribution-USA (nationwide) including the states GA, NY, IL, TX, WA, PA, IN, MS, NC, and MO and the countries of Belgium, France, Hong Kong, Italy, Spain, and Switzerland. | Terminated | August 20, 2012 | 251 | 41
Z-2227-2012 | 62791 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Total T3 Reagent, PN: 33830 |  |  | October 11, 2011 | 2011 | Beckman Coulter Inc. | A recall was initiated because Beckman Coulter has identified a negative bias in test results when comparing elevated Total T3 patient samples diluted with Sample Diluent A to samples diluted with Access Total T3 Calibrator S0. | Device Design | Beckman Coulter sent a Product Correction letter dated October 11, 2011, to all affected customers. The PCA letter informed the customers of the problem identified and the actions to be taken. Customers were informed to stop SDA usage as an alternative dilutent and that the S0 dilutent is to be used for manual dilution of the patient sample for Total T3. They will also be informed to no longer use the On Board Dilution (OBD) option for Total T3. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions regarding this notification were instructed to contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter Representative.For questions regarding this recall call 714-961-4941. |  | 117385 | 117,385 units total (33,918 units in US) | Worldwide Distribution - USA (nationwide), Canada, and internationally. | Terminated | March 14, 2013 | 520 | 117385
Z-2228-2012 | 62640 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | Siemens ADVIA CentaurÂż Syphilis (SYPH) Assay , 200 ReadyPack |  |  | July 17, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Syphilis Assay Interference with HBsAg Assay | Component design/selection | Siemens Healthcare sent Urgent Device Recall Notice letters dated July 2012 via e-mail to all affected customers inside and outside the US. The letters identifed the affected product, reason for voluntary recall and risk to health. The letters provided actions to be taken by the customers and if there are any remaining inventory of the affected product, follow the guidelines outlined. Customers were instructed to complete and return the Field Correction Effectiveness Check form. For questions contact your local Technical Support Provider or Distributor. |  | 2050 | 2050 Kits | Worldwide Distribution - US Nationwide and the countries of:Foreign: Brazil, China, Denmark, Hong Kong, Malaysia, and New ZealandViet Nam | Terminated | November 22, 2013 | 493 | 2050
Z-2230-2012 | 62684 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Serial interface module RS-232 |  |  | July 26, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received reports of malfunctions involving the loss of data transfer capabilities using the data transfer capabilities using modules on the Terumo Advanced Perfusion System 1. Some reports stated that the odor of smoke was as associated with the failures. The malfunctions occurred during set-up, during cardiopulmonary bypass, and during routine testing of the modu | Device Design | Terumo sent an URGENT MEDICAL DEVICE Recall letter dated July 24, 2012 to all affected customers. The letter identified the affected product, reason for correction, potential hazard and actions to be taken. The letter stated that Terumo CVS will replace all affected modules. Customers were informed that a Terumo VS Representative will contact the user to schedule a service call for the replacement upon receipt of the Response Form attached. For questions contact Terumo CVS Customer Service at 800-521-2818. |  | 579 | 579 units | Worldwide Distribution - US Nationwide including the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI, and WV and the countries of: AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Germany, Hong Kong, Indonesia, Japan, Jordan, Malaysia, Singapore, South Africa, South Korea, Taiwan, Thailand, UNITED ARAB EMIRATES (UAE), and Vietnam | Terminated | June 27, 2014 | 701 | 1427
Z-2234-2012 | 62134 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OPTETRAK LOGIC Proximal Tibial Spacer &amp; Screws (2) 2.5 LT BLUE Size 2.5, 8mm***STERILE***Match insert/spacer/tray color. Use with LOGIC Tibial Components. REF 02--012-42-2508. |  |  | October 17, 2011 | 2011 | Exactech, Inc. | Exactech, Inc. of Gainesville, FL recalled their Optetrak Proximal Tibial Spacer after the devices were determined to be dimensionally incompatible to mate as intended with the Logic Tibial Tray. | Device Design | The firm, Exactech, Inc., sent a "IMPORTANT PRODUCT RECALL NOTICE" date October 17, 2011 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately cease distribution or use of these products; extend this information to their accounts that may have possession of this product; verify if they have any of the subject Logic size 2.5 Proximal Tibial Spacer products, complete and return the attached PRODUCT RECALL FAX NOTICE form within five (5) working days via fax to: 352-337-3915, and to contact the firm inventory representative to confirm quantities at the location. For questions regarding inventory restocking, please call 1-800-392-2832. |  | 70 | 70 | Nationwide distribution: USA including states of: AL, AZ, CO, FL, GA, IL, LA, MD, MN, NC, NJ, NY, OH, OK, PA, PR, TN, VA, and WA. | Terminated | November 07, 2013 | 752 | 70
Z-2229-2012 | 62452 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory SMART &amp; Sunquest Laboratory SMART Select |  |  | August 12, 2011 | 2011 | Sunquest Information Systems, Inc. | In GUI inquiry, searching by ID (CID) may return the wrong patient demographic data if the accession number associated to the CID is no longer inactive. | Process design | Sunquest sent an Urgent - Product Safety Notice letter dated August 12, 2011 to all affected customers. The letter identified the affected product, summary of the problem, actions to eliminate the problem and actions to be taken by clients. Sunquest will distribute a new maintenance release.Client Support can be reached at 1-877-239-0337 for U.S and Canadian sites and at +44 (0) 101-335-0562 for international sites. |  | 47 | 47 sites | US Nationwide Distribution | Terminated | September 04, 2012 | 389 | 47
Z-2244-2012 | 62571 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-armÂż Imaging System. Product Numbers: BI-700-00027-100, BI-700-00027-100R, BI-700-00027-120, BI-700-00027-120R, BI-700-00027-230, BI-700-00027-230R. |  |  | July 11, 2012 | 2012 | Medtronic Navigation, Inc. | A fault in a high voltage relay in the 0-armÂż Imaging System might cause the relay to overheat and burn. | Component change control | Medtronic Navigation issued an "Urgent Field Safety Notice" to consignees dated 7/11/12. The problem was described and recommended actions were provided. For each affected O-armÂż Imaging System, a Medtronic service representative will visit the site and replace the 500VDC relay on the system and test the O-armÂż Imaging System. |  | 17 | 17 units | Worldwide Distribution - USA, including the states of AL, CO, FL, IA, MI, IA, NC, NE, NY and the countries of South Africa, Italy, and Spain. | Completed |  |  | 17
Z-2238-2012 | 62567 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging system. |  |  | May 29, 2012 | 2012 | Mckesson Medical Imaging Group | There is a potential to save reports in RadReport with an Incorrect Final Status. | Software design | McKesson sent an Urgent Field Safety Notice letter dated May 29, 2012 to all affected consignees. The letter identified the affected product, circumstances under which the issue occurs and immediate actions to be taken. Customers were advised to notify the Customer Support department that they have read and understood this Advisory Notice at 1-800-663-2533. For questions contact the Support Manager at the Customer Support department. |  | 516 | 516 units | Worldwide Distribution - U.S. Nationwide and the countries of: Canada, France, Ireland and United Kingdom. | Terminated | September 12, 2013 | 471 | 516
Z-2239-2012 | 62892 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Positron Emission Tomography (PET) and X-ray Computed Tomography, GEMINI TF 16 PET/CT, GEMINI GXL 16 PET/CT, GEMINI TF Big Bore PET/CT, Philips Medical Systems, Cleveland, OH. |  |  | August 08, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips became aware of an issue relating to a UPS battery cabinet that appeared to have burned a hole through the casing on one of the batteries. Supplier of the UPS, Chloride/Emerson, determined the batteries, housed within the UPS battery cabinet, entered thermal runaway causing the battery to overheat. | Material/Component Contamination | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated August 8, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Field Service Engineers will be implementing a Field Change Order (FCO) to replace the batteries of the affected UPS units. Contact the Philips Healthcare Customer Care Solutions Center at 1-800-722-9377 for questions regarding this notice.. |  | 7 | 7 units | Nationwide Distribution-including the states of CO, FL, MD, MN, NE and PA. | Terminated | April 28, 2014 | 628 | 7
Z-2246-2012 | 61671 | No | ILK | Powered patient transport. | Physical Medicine | Physical Medicine | 510(k) | Powered Mobility Offers Versatility (MÂżV)TM Chair |  |  | June 27, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Paramed Corporation (d/b/a, Descent Control Systems), a manufacturer and distributor of patient transport devices, merged with an affiliate of Stryker Corporation in August 2009. As part of the due diligence process, Paramed's powered Mobility Offers Versatility (MÂżV) chair was reviewed to identify potential compliance issues in relation to product documentation and classification. | PMA | Stryker sent a Urgent Medical Device Recall notice dated June 18, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to " Return the enclosed post card to confirm receipt of this notification." Locate the chairs listed in this notice." Call 269-389-6750 to arrange shipment of the chairs to Stryker Medical." If you have loaned or sold any of the chairs listed in this letter, please, forward a copy of this notice to the new users and advise us of their new location." If you have disposed of any of the affected chairs and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers.We apologize for any disruption that this advisory notice may cause. We strive to make products that meet our customer expectations for quality and durability.For further questions please call (269) 389-6927 |  | 159 | 159 units | Worldwide Distribution -- USA (nationwide) including the states of AK, AL, AZ, CA, FL, ID, IN, KS, KY, LA, MA, MD, MN, MO, NC, NJ, NY, OR, PA, TN, TX, UT, UT, VA, WA, WI, and WV, and the country of Canada. | Terminated | December 30, 2013 | 551 | 159
Z-2240-2012 | 62691 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | 892.1650 Image-intensified fluoroscopic x-ray system; the device is designed to provide fluoroscopic and spot film imaging of the patient during diagnostic surgical and interventional procedures. The systems include features specifically designed for use in diagnostic and interventional cardiac imaging procedures, and are also intended for cholangiography, endoscopic, urologic, orthopedic, neurologic, vascular, critical care and emergency room procedures. The systems may be used for other imaging applications at the physician's discretion. |  |  | November 06, 2009 | 2009 | GE OEC Medical Systems, Inc | GE OEC is recalling certain 9800 Image Intensified Flouroscopic x-ray systems due to customer complaint analysis showing that a replacement Hitachi 160 GB Hard Drive, when installed on certain models is susceptible of causing unanticipated system shut downs, no boots, data loss or data mix, or unexpected hard drive corruption. | Device Design | GE OEC sent a Urgent Recall Notice dated November 6, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The notification advised them that GE Healthcare had identified a potential safety issue associated with a "Significant increase on Data Loss, Data Mix, and System file corruption of the Hard Drives" on the 9800 Systems.The field correction consisted of installing a "Smart Power Switch" between the power control PCB and the system power switch on 9800 Systems with a Celeron single board computer and software version 29. This will remedy the potential for data loss, data mix, and system file corruption of the hard drives.For further questions please call 800-874-7378 Option 8. |  | 3595 | 3,595 units | Worldwide Distribution -- USA (nationwide) | Terminated | January 28, 2013 | 1179 | 3595
Z-2237-2012 | 62410 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT cTnI cartridges |  |  | March 07, 2012 | 2012 | Abbott Point Of Care Inc. | Abbott Point of Care has determined that i-Stat CTNL cartridges have the potential to exhibit incorrectly elevated cardiac troponin I results. This recall was expanded on July 19 2012 to include additional information regarding user documentation. | Device Design | Abbott Point of Care (APOC) sent an Urgent Recall Notice letters, via Fed Ex on March 07, 2012 to all affected customers. APOC sent an important product information update on July 19, 2012 to all of their affected customers. The letter identified the affected product, problem and actions to be taken. Customers are asked to review the information included as it may pertain to their cartridges. Customers were instructed that if the affected product have been forwarded to another facility, provide a copy of this letter to them. For additional information, contact APOC at 800-366-8020 option 1 or your Abbott Point of Care representative. |  | 7830075 | 7,057675 cartridges US; 772,400 cartridges OUS | Worldwide Distribution - US Nationwide, International | Terminated | February 19, 2013 | 349 | 7830075
Z-2241-2012 | 62729 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE HC OEC 9900 Elite and GE OEC 9800 Fluoroscopes Snubber Boards are a PCB with a 70 amp fuse to provide fused protection of internal system components from the high voltage system components, such as the X-ray tube or High Voltage Transformer. System model number A349855, Snubber Board number 00-880405-03 rev 5. This product is used for image intensified fluoroscopic x-rays for interventional radiologic procedures. |  |  | April 26, 2010 | 2010 | GE OEC Medical Systems, Inc | GE OEC recalled certain 9800 Fluoroscope &amp; 9900 Elite Models because the Snubber Board did not conform to the IEC 60601-1 standard for clearance and creepage of voltage isolation. | Component change control | No customer communication took place. On April 26, 2010, GE began opening service cases for field engineers to begin replacing the snubber boards. This action has been completed as of October 29, 2010. |  | 845 | 845 units | Worldwide and Nationwide (USA) distribution. | Terminated | August 22, 2012 | 849 | 845
Z-2243-2012 | 62751 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Encompass |  |  | April 13, 2009 | 2009 | Sunquest Information Systems, Inc. | Sunquest is recalling the Sunquest Encompass versions 2.4 and later because the specimen collect time is displaying and printing one hour later than the actual collect time only when Daylight Saving Time is active. | Software design | The firm, Sunquest, sent a "Urgent-Product Safety Notice: PSN-09-01" letter dated April 13, 2009 all their customers who purchase the Sunquest Encompass versions 2.4 and later. The letter informed the customers of the problem identified and the action to be taken. Customers were instructed to request software correction and to create a service request via client support webpage at www.sunquestinfo.com/supportweb. Then choose Submit a Service Request. Select Issue Type and select Software Request. When Software Request is selected, the Work Queue field automatically fills in with "Software Request." Promptly advise appropriate personnel.Customers with questions can contact Client Support and Services at (877) 239-6337 for US and Canadian sites and at +44(0)0800-960-972 for international sites. |  | 23 | 23 clients | Worldwide distribution: USA (nationwide) and country of: Bahamas. | Terminated | August 22, 2012 | 1227 | 23
Z-2242-2012 | 62460 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatic System is included with the Dimension Vista System. The Dimension Vista Integrated System is an in vitro diagnostic device intended tot duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. |  |  | May 29, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | When using the EasyLink Quality Control Package the system may not perform as intended for Multi-Rule QC violations [e.g., 2(2S), 4(1s) and 10x], unless configured using the instructions send with the Urgent Field Safety Notice. When this situation occurs, QC status may not be flagged appropriately. If auto-verification rules are defined to hold patient results when QC status is out, the patient | Software design | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated May 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Correction Effectiveness Check Form was attached for customers to complete and return via fax to 302-631-8467. Contact Siemens Technical Solutions Center at 800-441-9250 for questions regarding this notice. |  | 637 | 637 | Worldwide Distribution-USA (nationwide) and the countries of Canada, France, Germany, Spain, Portugal, Italy, Japan, Norway, Australia, Austria, New Zealand, Belgium, Switzerland, Netherlands, Malaysia, Slovenia, South Korea, Denmark, Saudi Arabia, Slovakia, Taiwan, Philippines, Czech Republic, and Singapore. | Terminated | April 10, 2014 | 681 | 637
Z-2252-2012 | 62796 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | AGFA IMPAC CV DICOMStore with Media Purge Daemon (MPD) and IMPAX CV DICOMStore with Cardiovascular Purge Services (CPS), Medical Image Storage Device |  |  | August 09, 2012 | 2012 | AGFA Corp. | Loss of patient data can occur under certain circumstances due to misconfiguration of DICOMstore configured with Agfa's Medical Image Storage Devices - Media Purge Daemon (MPD) and Cardiovascular Purge Service (CPS). | Process design | On 8/9/2012 AGFA Healthcare issued an "URGENT FIELD SAFETY NOTICE" letter to consignees. This letter describes the problem and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood was also required from the consignees |  | 535 | 535 | USA Nationwide, Virgin Islands and Canada | Terminated | December 08, 2015 | 1216 | 535
Z-2245-2012 | 62790 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance 6, 10, 16, 40, 64 and CT Big Bore are Computed Tomography X-Ray Systems These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | August 03, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The patient support may move in an unintended manner if the foot switch cover becomes damaged and impinges on the foot switch. | Device Design | Philips Medical Systems (Cleveland), Inc. sent an Urgent - Medical Device Correction letter dated August 1, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taked. The letter contained important information for the continued safe and proper use of your equipment. The letter stated that, a Field Service Engineer will visit each affected site to install the hardware upgrade. For further information or support concerning this issue contact your local Philips representative or local Philips Health office. For North America and Canada contact Customer Care Solution Center (1-800-722-9377), option 5: Enter Site ID or Follow the prompts). |  | 126 | 126 units | Worldwide Distribution - US Nationwide including the states of: AL, AR, AZ, CA, DC, FL, GA, KY, MD, MI, MN, MS, ND, NM, NY, OH, OK, SC, TN, TX, VT, WA, and WI and the countries of: Austria, China, Denmark, France, Germany, Iceland, Japan, Netherlands, South Africa, Sweden, Switzerland, United Kingdom. There were two government consignees. There were no Canadian consignees. | Terminated | July 23, 2015 | 1084 | 126
Z-2235-2012 | 62663 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | The Instructions For Use for the following: |  |  | April 25, 2012 | 2012 | Stryker Spine | The front cover of Instructions For Use (IFU) #NOLI135B03Rev02 was improperly marked with a "Sterile" logo in the lower right corner. The products these IFUs were packed with are sold as "Non-Sterile" and require sterilization before use in surgery. | Labeling mix-ups | Stryker sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter to affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Customer Response Form was attached for customers to complete and return via fax to 201-760-8370. Contact the firm at 201-760-8298 for questions regarding this notice. |  | 40 | 40 units. | Distributed in Atlanta, Georgia. | Terminated | February 27, 2013 | 308 | 40
Z-2236-2012 | 62744 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | OEC InstaTrak 3500, picture archiving and communications system intended as an aid to locate anatomical structures during open percutaneous surgical procedures. It is indicated for use in medical conditions that may benefit from the use of stereotactic surgical technique. The system provides a reference to rigid anatomical structures such as sinus, skull, long bone, or vertebra, which are visible on medical images such as CT, MRI, or X-ray. |  |  | October 11, 2006 | 2006 | GE OEC Medical Systems, Inc | GE Healthcare had recalled certain OEC InstaTrak 3500 Carts due to the potential for the cart to tip over when the arm of the imaging device is extended during use. | Component design/selection | GE Healthcare sent an "URGENT SAFETY ADVISORY NOTICE" dated October 11, 2006 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter stated that GEHC OEC Field Service Engineers would be contacting their facility to place a caution label on the articulating monitor arm of the cart and to provide a revised operators manual that will contain additional information on the safe usage of the cart. Until a GEHC Engineer had taken the actions described, they advised their consignees to exercise caution when fully extending the articulating monitor arm on the device to avoid potentially tipping over the cart. Contact the firm at 800-874-7378 for questions regarding this notice. |  | 99 | 99 units | Nationwide Distribution. | Terminated | August 22, 2012 | 2142 | 99
Z-2259-2012 | 62706 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC 660i, 680i, 860i, 880i SYNCHRON Access Clinical Systems, |  |  | October 13, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has received seven reports of a software anomaly for the UniCel Dxl Immunoassay Systems. If undetected, this anomaly may cause erroneous test results. | Software design | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated October 13, 2011 with attached PCA Response Form to all customers who purchased the UniCel Dxl Immunoassay Systems and DxC Clinical Systems which includes all software versions. The letter describes the product, problem and actions to be taken. The customers were instructed verify the system temperatures immediately following each initialization; do not begin processing samples until this verification is complete; refer to the appendix included with the letter for detailed instructions about verifying system temperatures; review their quality control (QC) performance history; and share this information with their staff and retain this information as part of their quality system documentation, and to complete and return the enclosed response form within 10 days to Beckman Coulter, Inc., 1000 Lake Hazeltine Drive, Chaska, MN 55318; Attn: Regulatory Affairs or fax to: 1-786-639-4000. Customers with any questions regarding the notice, were instructed to contact Customer Support Center via our website, http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter Representatives. |  | 2442 | 2442 units total (784 units in US) | Worldwide distribution: USA (nationwide) and countries of: Algeria, Angola, Argentina, Australia, Austria, Bangladesh, Belgium, Botswana, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Korea,Republic of; Kuwait, Liechtenstein, Malaysia, Mayotte, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Philippines, Poland, Puerto Rico, Romania, Russian Federation, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Viet Nam. | Terminated | August 23, 2012 | 315 | 2442
Z-2254-2012 | 62870 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Echelon MRI System/Oasis MRI, Software. These are software controlled MRI devices. |  |  | June 15, 2012 | 2012 | Hitachi Medical Systems America Inc | Hitachi discovered a software error with a feature called Multi-Planar Reconstruction(MPR). MPR takes an original volume dataset reconstructed in any plane view and creates a new image set from another view (i.e. axial to coronal, coronal to sagittal, etc.). The error causes the first image in the dataset to have a right/left reversed orientation. The software error can occur randomly when execu | Software design | Hitachi sent a Device Correction Notice letter dated June 18, 2012 to all affected customers. All letters are recorded by tracking number and date received by the consignee. The letter identified the affect product, problem and actions to be taken. All sites will be visited by Hitachi Service to install the software update to permanently resolve the issue at no cost to the facility. Software correction is targeted to start by September 1, 2012. For question contact the Manager of Regulatory Affairs at 800-800-3106 x3720 |  | 165 | 165 units | US Nationwide Distribution - including the states of:: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY and Mexico. | Terminated | February 19, 2013 | 249 | 165
Z-2261-2012 | 62871 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Visual-ICEÂż Cryoablation System, Model FPRCH 6000 |  |  | July 20, 2012 | 2012 | Galil Medical, Inc. | Galil Medical is recalling a number of Visual Ice Cryoablation Systems because the gas regulators may fail to regulate the gas pressure, rendering the devices inoperable. | Device Design | Galil Medical sent a "Voluntary Medical Device Field Correction" letter dated July 20, 2012, via UPS to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that a Galil Field Service team will be contacting customers to schedule a service visit the week of July 23rd. Customers were instructed to forward and communicate this field correction notice and instructions to individuals in their organization and/or to other organizations affected by this action. For questions call Galil Medical Customer Service. |  | 4 | 4 | Worldwide Distribution - US nationwide including the state of MN and the countries of United Kingdom and Netherlands | Terminated | September 13, 2012 | 55 | 4
Z-2262-2012 | 62746 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GEHC OEC Injection Cable for MedRad Provis V Contrast Injector. |  |  | October 10, 2008 | 2008 | GE OEC Medical Systems, Inc | GE initiated a correction to the instructions/use due to the potential for erroneous auto injection of contrast dye. | Device Design | An Urgent Safety Advisory Notice was sent to consignees on October 10, 2008. The letter identified the affected product and notified customers of the situation, along with instructing them to discontinue use until proper validation and instruction update for cable use was completed. GEHC OEC was to have contacted customers when the solution became available and resolved it at no cost to the customer. |  | 149 | 149 units | Nationwide distribution. | Terminated | January 10, 2013 | 1553 | 149
Z-2266-2012 | 62851 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba CT systems, Asteion and Aquilion, that are equipped with a PC based computer, or to those older systems (SGI based) that received an improved table by an upgrade purchase. |  |  | February 27, 2008 | 2008 | Toshiba American Medical Systems Inc | The systems all have a similar table control board. It hasbeen found that this board may cause the table position to be incorrect if the table cradle is movedmanually. This is referred to as "Free" mode in the operation manuals. If this error does occur it ispossible that the patient could be scanned at an incorrect position. | Process design | Toshiba sent a "URGENT: MEDICAL DEVICE CORRECTION" letter to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to please avoid moving the table in the "Free" or manual mode as much as possible. If it is necessary to move the table in this manner please verify the table position prior to initiating the scan sequence.If it is noted that the position is not correct, exit the eXam plan and start a new one beginning with the scanoscopy. Please share this information with all users of the equipment and your clinical or biomedical engineering departments.Please find the attached customer reply form. The purpose of this form is for FDA tracking. Please take a moment to complete the form and fax it to the toll free number on the form.". The contact information for this recall is:"(800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968.". |  | 757 | 757 | USA ( Nationwide Distribution ) | Terminated | September 04, 2012 | 1651 | 757
Z-2268-2012 | 62881 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Mini-Infuser Model 150XL Single Speed Infusion Pump, an Rx battery operated syringe infusion pump; |  |  | March 04, 2011 | 2011 | Baxter Healthcare Corp. | Should a battery with the minimum contact diameter (.51 inches) be placed on the battery spring (.53 inches), an electrical circuit will not be completed and the Mini-Infuser will not operate. The reason being that the Mini-Infuser Battery Spring has a flat contact that measures .53 inches Âż.01 inches on the inner diameter of the spring coil. The IEC 60086-2 and ANSI 08 standards specify a minimum | Nonconforming Material/Component | Baxter sent a Safety Alert letter on January 11, 2011 for the InfusO.R. Pump. Baxter expanded the scope of the recall to include the Mini-Infusers System and sent an Urgent Device Correction letter dated March 4, 2011, via First Class mail to all affected customers. The letters informed the customers that the Mini-Infusers use the same size battery and incorporate the same battery contact interface as the InfusOR pump, and that the same feature that inhibits electrical contact in over-the counter Energizer C size batteries may also be incorporated by other battery manufacturers. The customers were requested to verify that the InfusOR and Mini-Infuser will power-on before use; if the device does not power on, replace the batteries with a different brand of size C batteries and verify that the device will power-on before use; and to place a copy of the letter with each operator's manual for the InfusOR and Mini-Infuser System for future reference. For Technical questions call Baxter Global Technical Services at 1-800-843-7867. For question regarding this communication contact the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 26485 | 26,485 units | 45,865 units | Terminated | August 27, 2012 | 542 | 45865
Z-2270-2012 | 62708 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory Versions 6.4 and later using Microbiology Result Entry |  |  | December 05, 2011 | 2011 | Sunquest Information Systems, Inc. | The recall was initiated because Sunquest has confirmed that the Sunquest Laboratory versions 6.4 and later while using Microbiology Result Entry (MRE), the specimen ribbon containing the Accession Number (AN) and the Battery Name may be missing or contain the wrong information. | Software design | Sunquest sent a recall letter dated December 5, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Your reply to this email message serves as verification you have received, read, understood and/or followed instructions outlined in the attached Product Safety Notification. Your reply assists Sunquest Information Systems to satisfy our obligation to Regulatory Agencies.Please note in your email reply if your site will not be requesting the software correction. If your site plans to request the software correction, please follow the directions noted on the attached PSN.Further questions please call (520) 570-2347. |  | 127 | 127 sites | Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | August 27, 2012 | 266 | 127
Z-2267-2012 | 62864 | No | LRJ | General purpose disinfectants. | General Hospital | General Hospital | 510(K) Exempt | CaviCide, Part Numbers: 13-1025, 13-1000, 13-1024, MC-1000. |  |  | February 06, 2012 | 2012 | Metrex Research, LLC. | Metrex Research is recalling Cavicide containing the Spring Fresh Fragrance because the affected lot of product containing this fragance has caused an increase in reported health reactions in some individuals such as headache, nausea, and respiratory discomfort. | Component change control | Metrex sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 16, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Mandatory Recall Response Form was attached for customers to complete and return via fax to 734-947-6725.Contact Customer Care Center at (800) 841-1428 for questions regarding this recall. |  | 17690 | 17,690 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, AE, and YE. | Terminated | August 24, 2012 | 200 | 17690
Z-2271-2012 | 62578 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Low Profile Abutment, Non-Hexed Castable Cylinder Model LPCCC1. Product is labeled in part: "***REF LPCCC1***Low profile Abutment Hexed Castable Cylinder***Cylinder; Cylindre; Zylinder; Cilindro; Cilindro; Cilindro***RX only***NON STERILE***Caution: Law prohibits dispensing without prescription***Consult instructions for use*** |  |  | May 31, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i conducted a recall on the device Low Profile Abutment, Non-Hexed Castable Cylinder due to some packages labeled LPCCC1, Hexed Castable Cylinder possibly containing LPCCC2, Non-Hexed Castable Cylinders. | Labeling mix-ups | Biomet 3i sent a "MEDICAL DEVICE ADVISORY NOTICE" dated May 31, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-342-5454 for questions regarding this recall. |  | 611 | 611 units | Worldwide Distribution-Distributed in the USA and the countries to Spain, France, Germany, Switzerland, and Australia. | Terminated | November 21, 2013 | 539 | 611
Z-2272-2012 | 62714 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory LabAccess Results Workstation (LARS) |  |  | October 05, 2010 | 2010 | Sunquest Information Systems, Inc. | The recall was initiated because Sunquest has confirmed that the Sunquest Laboratory LabAccess Results Workstation (LARS) may file results to the wrong patient by incorrectly combining results. | Software design | An "Urgent - Product safety notice " was issued on 11/4/10 to all customers who purchased the Sunquest Laboratory LabAccess Results Workstation (LARS). The safety notice informed the customers of the problems identified and the actions to be taken. US customers may contact 1-877-239-6337. |  | 151 | 151 sites | Worldwide Distribution, including Nationwide (USA) and the countries of Canada, Denmark, Ireland, the United Arab Emirates and the United Kingdom. | Terminated | August 24, 2012 | 689 | 151
Z-2263-2012 | 62770 | No | NSX |  |  | General Hospital | Contact ODE | Cisco brand HealthPresence, version 2.1; Model/Catalog Number: CHP-ENTPRSYR2.I-K9; Product is manufactured and distributed by Cisco Systems, Inc. San Jose, CA |  |  | July 23, 2012 | 2012 | Cisco Systems Inc | Software anomaly may result in incorrect values and interpretations. The initial vital signs reading will display accurately on the Cisco Health Presence Vitals tab and the OnePlace tab. However, subsequent vital signs readings will only update on the Cisco HealthPresence Vitals tab and within the OnePlace database, but will not update under the OnePlace tab, and thus, will not provide current vit | Software design | Cisco notified their one customer by phone on July 25, 2012, with a follow up email and instructed the customer to stop using the OnePlace vitals tab and to use the CHP only. On July 27th, 2012, the vendor of the OnePlace application, fixed the bug remotely.On July 31, 2012, Cisco's staff physician contacted the customer to confirm that the customer was aware of the fix. Validation of the correction was completed by Cisco, along with customer on August 9, 2012. Until validation was satisfactorily completed, the integrated system lie dormant at customer's site. |  | 1 | 1 unit | Nationwide Distribution including the state of TN. | Terminated | August 24, 2012 | 32 | 1
Z-2265-2012 | 62828 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Canon Solid State X-ray Imager (Flat Panel/Digital Imager), Digital Radiography, Models CXDI-55C, CXDI-55G and CXDI-70C. These devices are intended to replace radiographic film/screen systems in all general purpose diagnostic procedures. |  |  | November 01, 2010 | 2010 | Canon Inc. | Screws holding the 55REF circuit board are loosening and roaming inside of the sealed sensor. | Nonconforming Material/Component | Canon Inc. notified sales companies including Canon USA on 9/15/2010 with Service Bulletin 10-017, and on 8/25/2011 with Service Bulletin 11-011, and on 1/20/2011 with Service Bulletin 11-001. Sales representatives were to visit customers and check and replace the screws. |  | 468 | 468 units | Nationwide Distribution (USA) | Terminated | January 07, 2014 | 1163 | 468
Z-2264-2012 | 62678 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS. |  |  | July 18, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction when using syngo Imaging XS, version VA70A or higher. In some situations it can happen that a miscalculation is done for the grey scale values in the functions "Regions of Interest", "Pixel Lens", "Edge Enhancement" and "Histograms". | Software design | Siemens sent a "CUSTOMER SAFETY ADVISORY NOTICE" dated July 12, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. For questions concerning this notice contact your local Siemens Service Rep.. |  | 52 | 52 | Nationwide Distribution-CA, CO, CT, FL, GA, IL, IN, IA, KS, MD, MN, MO, NE, NV, NJ, NY, NC, OH, PA, TN, TX, and WI. | Terminated | November 06, 2014 | 841 | 52
Z-2273-2012 | 51437 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TRESTLE LUXEÂż Anterior Cervical Plating System, Product code KWQ. 510(k) K102820. |  |  | May 29, 2012 | 2012 | Alphatec Spine, Inc. | On May 9th 2012 Alphatec Spine received a complaint which reported that the distal tipof the Trestle Luxe Slide Alignment Tool (Part Number 71721), does not properly fit intothe hexalobe screw on the Trestle Luxe Plate. | Device Design | Alphatec Spine sent an "URGENT-MEDICAL DEVICE RECALL NOTIFICATION" letter dated May 29, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A form was attached for customers to complete and return via fax to 760-431-0289. Contact Alphatec Spine's Customer Service Department at customerservice@alphatecspine.comfor questions regarding this recall.) |  | 38 | 38 | Nationwide Distribution. | Terminated | September 04, 2012 | 98 | 38
Z-2274-2012 | 62564 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Apelo Modular Screwdriver labeled in part: "APELO PEDICLE SCREW SYSTEM***Modular Screwdriver***Atlas Spine, Inc.***www.AtlasSpine.com*** 1555 Jupiter Park Drive #4***Jupiter, Florida 33458***561-741-1108***Tissue Sleeve***MFG Date: XX/XX/XXXX***Lot #: XXXX***Weldment***MFG Date: 03/23/2009***Sleeve Assembly***MFG Date: XX/XX/XXXX***Lot #: XXXX***Qty: 1 Rx Only***Non-Sterile. Sterilize Prior to Use*** Rev. A***. The Modular Screwdriver is used to insert modular pedicle screw implants which are part of the Apelo System. |  |  | June 04, 2009 | 2009 | Atlas Spine, Inc. | Atlas Spine, Inc. recalled their Apelo Modular Screwdriver, a component of the Apelo Pedicle Screw System Instrumentation, because of complaints received about the tip breaking. If the tip is fractured inside of the modular screw, the tip could become implanted in the body or if the entire screw is removed, it could cause a delay in the surgery. | Component design/selection | Consignees were notified by Telephone and visit. All products will be returned, quarantined and then destroyed. |  | 4 | 4 pieces | USA Distribution to the state of Florida only. | Terminated | July 16, 2014 | 1868 | 4
Z-2276-2012 | 62919 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | CIDEX OPA, High Level Disinfectant Solution, Glutaraldehyde-free (0.55% ortho-phthaladehyde) solution. The CIDEX OPA solution is intended for automated and manual cleaning and disinfection. |  |  | June 01, 2012 | 2012 | Advanced Sterilization Products | The recall was initiated because Advanced Sterilization Products has discovered instances of missing labels on bottles of CIDEX OPA Solution, Part # 20390 manufactured between May 2011 and April 2012. | Process control | An "Urgent Medical Device Recall" letter dated August 10, 2012 was sent to all customers who purchased the CIDEX OPA solution, part #20390 (manufactured between May 2011 and April 2012). The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to notify their customers if they have distributed bottles of CIDEX OPA Solution to other locations. Customers were instructed to call Stericycle at (866) 629-6182 to arrange for return of the product. |  | 1310844 | 1,310,844 units | Nationwide (USA) and Puerto Rico. | Terminated | September 11, 2013 | 467 | 1310844
Z-2277-2012 | 62780 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 16 Fr HemoStar XK Long-Term Hemodialysis Catheter, implanted Standard kit containing an AirGuard introducer. The device is used for patients requiring hemodialysis. |  |  | August 07, 2012 | 2012 | Bard Access Systems | Bard Access Systems is recalling one lot of 16 Fr HemoStar XK Long Term Hemodialysis Catheters because they were identified with an incorrect expiration date, 03/2015 rather than 10/2013. | Incorrect or no expiration date | Bard Access Systems sent an "URGENT: *16 FR HEMOSTAR* LONG-TERM HEMODIALYSIS CATHETER RECALL" letter dated August 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Reply Form and Inventory Reconciliation Form was enclosed for customers to complete and return via fax to 1-801-522-5674. Contact BAS Customer Service at 1-800-290-1689 for questions regarding this recall. |  | 120 | 120 units | National US distribution only: MO, KS, AZ, OK, OR, AL, MA, TX, FL, MI, OH, LA, NJ, NC, WA, AR. | Terminated | January 07, 2013 | 153 | 120
Z-2275-2012 | 62679 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo.plaza, Model number 10592457. |  |  | July 18, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction in which a miscalculation occurs for images with pixel depth values greater that 12 bit in certain functions. | Software design | The firm initiated their recall/correction of this device on July 18, 2012 by sending a "Customer Safety Advisory Notice" letter, dated July 12, 2012, to their consignees. The letter identified the affected product along with the malfunction. It also discussed the steps the user can take to avoid the issue. Users were asked to not use the mentioned functions for images from those modalities until the issue was resolved and the system was updated with the software version containing the solution. Users were to contact their local Siemens Service in order to plan for the updated software version that will resolve the issue. Additionally, users should place the safety advisory notice in the system's instructions for use, and personnel should be instructed accordingly and maintain awareness. If the device was sold and is no longer in the users' possession, the notice should be forwarded to the new owner and the firm should be notified of the new owner. |  | 14 | 14 | Nationwide Distribution -- CA, FL, GA, MA, MN, MO, NJ, NY, OH, WA, and WI. | Terminated | January 24, 2014 | 555 | 14
Z-2286-2012 | 61982 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Product is labeled in part: "***MEDLINE***Extended PTFE Needle 6.5" (16.51cm) Length***Reorder: ES0016***Contents: 1 Each***SINGLE USE ONLY**latexfree***RX Only***STERILE EO***www.medline.com Manufactured for Medline Industries, Inc. MEDLINE is a registered trademark of Medline Industries, Inc. 1-800-MEDLINE***Mundelein, IL 60060 USA Made in USA.***" |  |  | June 16, 2010 | 2010 | Unimed Surgical Products, Inc. | On 6/16/2010 Unimed Surgical Products, Inc., 10401 Belcher Road, Largo, FL 33777 initiated a recall of the Medline Industries, Inc. Electrosurgical electrodes used for cutting and coagulation of soft tissue due to concerns of sterility. | Process control | Unimed Surgical Products, Inc. issued a Medical Device Removal letter dated June 4, 2010 to Medline Industries, Inc. The letter identified the affected product, problem and actions to be taken. Medline was instructed to quarantine any undistributed product and notify its customers of the recall. For further information, contact Unimed Surgical Product, Inc. by phone 800-531-2627 or FAX 800-886-1225. |  | 1440 | 1440 ea | US Nationwide Distribution - including the state of IL. | Terminated | August 30, 2012 | 806 | 25440
Z-2281-2012 | 61171 | No | LGW |  |  | Neurology | PMA | The Eon Mini Neurostimulation (IPG) System (Model 3788) is a 16-channel, rechargeable, multi-program system that allows the connection of one or two leads with 4, 8, and 16 electrode leads. |  |  | December 19, 2011 | 2011 | Advanced Neuromodulation Systems Inc. | As of 11/30/2011, St. Jude Medical has received 110 patient complaints of warmth or heating at the implantable pulse generators (IPGs) implant site during charging for the Eon IPG and 116 reports of similar symptoms for the Eon Mini IPG, respectively. | Labeling design | ST. Jude Medical sent an Important Medical Device Correction letter dated December 19, 2011 to all affected customers. The letter identified the affected product, problem and recommendations to be followed. For questions contact your St. Jude Medical Neuromodulation Division Representative. |  | 30560 | 30,560 units | Nationwide within the US, and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Isreal, Italy, Japan, Mexico, Netherlands, Puerto Rico, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | July 24, 2015 | 1313 | 56948
Z-2283-2012 | 62712 | No | DAP | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Siemens Stratus(R) CS Acute Care(TM) D-dimer Calibrator (DDM CalPak) lots 502058002 and 501318002 |  |  | June 07, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the Stratus(R) CS Acute Care(TM) D-Dimer CalPak calibrator lots 502058002 and 501318002 may exhibit SLOPE or CALCV error messages (instrument errors), which indicate calibration failures. | Nonconforming Material/Component | SIEMENS sent an Urgent Field Safety Notice dated June 7, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included a response form to be faxed to the recalling firm at 302-631-8467 by the consignee. If customers experienced calibration failures, they were instructed to contact the Siemens Technical Solutions Center at 800-405-6473 for assistance. Customers were also asked to forward the notification to anyone to whom they may have distributed the affected product. For questions regarding this recall call 302-631-6311. |  | 786 | 786 cartons | Nationwide Distribution - USA including AK, AL, AR, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, WA, WI, WV, and WY. | Terminated | January 16, 2014 | 588 | 786
Z-2251-2012 | 62715 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | OEC IT3000; IT2500 (EnTrak); IT2500 Plus (EnTrak Plus); IT3500 (InstaTrak); IT3500 Plus (InstaTrak Plus); ConneCTstat; and ConneCTstat Plus picture archiving and communications systems. The systems are an aid to locate anatomical structures during open or percutaneous surgical procedures. It is intended for use in medical conditions that may benefit from the use of stereotactic surgical technique. The system provides a reference to rigid anatomical structures such as sinus, skull, long bone, or vertebra, which are visible on medical images such as CT, MRI, or X-ray. |  |  | June 22, 2009 | 2009 | GE OEC Medical Systems, Inc | GE Healthcare Surgery had discovered that using the Inverted Headset Placement is not a validated configuration with the IT 3000, 2500, 2500 plus, 3500, 3500 plus, ConneCTstat, and ConneCTstat Plus Surgical Navigation equipment. | Labeling False and Misleading | GE Healthcare sent an Urgent Safety Advisory Notice dated June 22, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised not to use the ENT registration Headset in the inverted position for any CT scans or surgical procedures and to remove all references in their Naviation System Operator's Manual that discuss an inverted headset placement, and replace it with the enclosed updated instructions for the Navigation System Operator's Manual. For any questions customers should contact the service team for further information at 800-874-7378, option #8.For questions regarding this recall call 801-536-4516. |  | 768 | 768 units | Worldwide Distribution - USA (nationwide) and the countries of India, Japan, Canada, Switzerland, Italy, United Kingdom, Germany, Austria, Hong Kong, France, and Denmark. | Terminated | September 28, 2012 | 1194 | 768
Z-2253-2012 | 62769 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | ON-Q pump with ONDEMAND bolus button, ON-Q C-bloc with ONDEMAND; ON-Q C-bloc with Select a Flow and ONDEMAND Model/Reference #: 7000180, 7000181, 7000182, 7000367, 7000397, CB003, CB005, CB006, P400X1-7+5/60, P400X2-14+5/30, P400X5+5/60, and PMB01. |  |  | May 08, 2012 | 2012 | I-Flow Corporation | It was determined that in a small quantity of ON-Q pumps with ONDEMAND, the bolus button may not lock in the down position when depressed and/or the orange bolus refill indicator may stay in the lowermost position. As a result, the patient may receive continuous infusion at a rate greater than expected. | Device Design | An Important Voluntary Recall Notice letter dated 5/8/12 was sent to all their customers who purchased the ON-Q pump with ONDEMAND bolus button. The letter provided customers with the problem indicated and the actions to be taken. Customers were instructed to complete and sign the attached Recall Response Form and fax to (920)969-4565 or scan and email to IflowProductInquiry@kcc.com.Once I-Flow receives the completed FAX BACK form, an I-Flow Customer Service Representative will immediately contact the customers to provide instructions on product return and credit. Customers with questions regarding the recall letter, were instructed to contact I-Flow Customer Service between 8am to 5pm (PST) at (800) 443-3569. Or customers can contact I-Flow 24 hour Technical Support line at (800) 444-2728. |  | 80646 | 80,646 units total (65,926 units in US) | Worldwide Distribution -- USA, Canada, Australia, the EU (Germany), United Arab Emirates, and Cyprus. | Terminated | April 25, 2013 | 352 | 80646
Z-2279-2012 | 62880 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | V Series Patient Monitors; Mindray DS USA, Inc. |  |  | May 07, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified two software anomalies contained in the V Series Monitoring System. There have been no reports of injuries associated with these anomalies. These anomalies affect V Series Systems distributed between September 31, 2010 and January 9, 2012. | Software design | Mindray sent an Urgent: V Series Monitoring System Correction letters dated May 7, 2012 via certified mail to all affect customer. The letter identified the affected product, problem and actions to be taken. Customers are directed that VPS modules should not be used between systems until a software upgrade has been completed. This work will be done at no cost to the customer. For questions contact your Mindray Serviice Representative at 1-800-288-2121. |  | 979 | 979 units | Worldwide Distribution - US Nationwide including the states of: NY, SD, NE, IL, MO, CA, OH, OK, WA, IN, WY, CO, TX, FL, NC, AL and Puerto Rico. and the countries of: Australia, Canada, Columbia, New Zealand, Pakistan, Russia, Saudi Arabia, United Kingdom and Venezuela. | Terminated | July 17, 2013 | 436 | 979
Z-2285-2012 | 62572 | No | MID | Multiple autoantibodies immunological test system. | Immunology | Immunology | 510(k) | PHADIA Varelisa Cardiolipin IgM |  |  | August 02, 2012 | 2012 | Phadia US Inc | Customers were not following the labeled assay procedure. This was due to a letter issued to them in 2004 from Phadia GmbH, which indicated that customers could omit the pre-wash step. The omission of the pre-wash step, when combined with a change in material source in kit lot #0067, caused elevated results. | Labeling Change Control | On 8/2/2012 an URGENT PRODUCT CORRECTION INFORMATION FOR PHADIA VARELISA CARDIOLIPIN IGM, Article Number 15696 letter was sent to all consignees the received the Direction for Use modification letter in 2004. |  | 67254624 | 837 kits (96 assays per kit), total of 80,352 individual assays | USA Nationwide. | Terminated | December 12, 2012 | 132 | 67254624
Z-2284-2012 | 62776 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | QKC1692E |  |  | August 10, 2012 | 2012 | Steris Corporation | STERIS has been actively seeking devices for quick connect revalidation projects. STERIS was unable to obtain two (2) of the needed devices; these two devices represent the design of seven (7) models in one quick connect. Due to the unavailability of these two validation devices, STERIS will remove 7 models from the labeling of QKC1692E and will re-label all QKC1692Es in the field.. | Device Design | The firm, Steris, sent a "Voluntary Field Correction: Re-labeling of QKC1692E" letter on August 10, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the steps outlined in the letter to re-label the affected QKC1692E(s) and return the completed form by September 30. 2012 to: Regulatory Affairs Specialist via Kate_Beveridge@STERIS.com or via fax at (440) 350-7082.If you have any questions, contact Director, Regulatory Affairs at 440-392-7519. |  | 64 | 64 units | Nationwide distribution: USA including states of: AR, AK, CA, CO, DE, FL, IL, MA, MD, NC, NE, NH, NY, OH, PA, SD, TN, TX, WI, and WV. | Terminated | July 05, 2013 | 329 | 64
Z-2282-2012 | 53055 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Whole Body X--ray Scanner |  |  | July 13, 2009 | 2009 | Toshiba American Medical Systems Inc | 1. It has been found that in raw data processing for Go&Return; helical scanning, when the raw data acquired in the first scan (Go) and that acquired in the second scan (Return) are selected and reconstructed at one time, incorrect couch positions and scan times are displayed on the reconstructed images for the data acquired by Return scanning, while all related information is displayed on the reco | Software design | The firm, Toshiba America Medical Systems, Inc., sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated July 30, 2009 to its Consignees/customers. The letter described the product, problem and actions to be taken. The letter informed them of the above mentioned reason for recall and telling them that a corrective measure will be installed to prevent the occurrence of the above problems. They were told that their Toshiba service representative will contact them to install the measures.Furthermore Toshiba instructed them to adopt the following temporary measures to avoid these phenomena until corrective measures have been taken.1. When performing raw data processing for Go&Return; helical scan, select the raw data acquired for Go and that acquired for Return separately in order to perform reconstruction2. When using the Sure Exposure 3D function, confirm that the error message shown in the preceding page is not displayed on the monitor at the time of setup (until the Scan Start button lights after the Confirm button is clicked).3. When executing an eXam Plan in which Sure Exposure 3D is selected as the actual scan for Sure Start, confirm that the error message show in the preceding page is not displayed on the monitor when monitoring scan is switched to the actual scan. If the error message is displayed, click [OK] in the error message dialog to continue processing.The letter should be shared with facility's Biomedical or Clinical Engineering Department. Also the customers were instructed to complete and return the Customer Reply Form via fax to: Regulatory Affairs at 877-349-3054.If you have any questions regarding this letter, contact Director, Regulatory Affairs at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968 with any questions regarding the letter. |  | 24 | 24 Systems | Nationwide distribution: USA including states of: AL, AR, CA, CO, FL, GA, KY, MA, MD, MT, ND, NJ, NV, NY, OH, PA and Puerto Rico. | Terminated | September 04, 2012 | 1149 | 24
Z-2278-2012 | 62853 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | EXCELART Vantage, Atlas, and Titan MRI Systems, MRT-1503 and MRT-1504. |  |  | May 13, 2008 | 2008 | Toshiba American Medical Systems Inc | The status display in the sequence queue window may not change from "Current" to "Done". | Process design | A recall notification letter titled: "URGENT: MEDICAL DEVICE CORRECTION," was sent to customers. The letter stated the problem, corrective action, and (instructions) request to customers. Customers were instructed to perform the temporary recovery procedures if the problem occurs. Additionally, the acknowledgement sheet should be completed and faxed back. The information should be shared with all users, including the bio-medical or clinical engineering departments. Questions pertaining to the information for this recall were to be directed to (800) 421-1968. |  | 54 | 54 | Nationwide Distribution | Terminated | September 04, 2012 | 1575 | 54
Z-2300-2012 | 49060 | No | FTN | Plastic surgery kit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Various SRI Surgical Disposable Accessory Packs. |  |  | July 16, 2008 | 2008 | SRI Surgical | Distribution of a product that did not meet specifications. (labeling error) | Component change control | SRI sent a "MEDICAL DEVICE CORRECTION" letter dated July 16, 2008 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 813-891-9550 ext. 3124 for questions regarding this recall. |  | 802135 | 802,135 | Nationwide Distribution-including the states of CA, FL, GA, KY, MD, MI, NC, NJ, OK, PA, SC, TN, TX, and UT. | Terminated | September 10, 2012 | 1517 | 802135
Z-2301-2012 | 62705 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DGPHP RFA High-Power Single Use Grounding Pads and Cool-Tip RFA Electrode Kits (containing DGPHP RFA High-Power single use grounding pads) |  |  | August 01, 2012 | 2012 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is conducting a recall of various production lots of DGHP RFA High Power single use grounding pads and Cool-tip RFA Electrode Kits due to the potential for foil degradation on the DGP-HP single use grounding pad which may result in an electrical burn to the patient at the application site. | Process control | The firm, Covidien, sent an "URGENT MEDICAL DEVICE RECALL" letter dated August 6, 2012, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine and discontinue use of the device; immediately advise all DGPHP RFA High-Power single use grounding pads and Cool-tip RF Electrode kit users of this recall; complete and return the attached Recalled Product Return Form in its entirety via fax to (800) 895-6140; return affected product to Covidien Attn: Field Returns Department, 195 McDermott Road, North Haven, CT 06473; if purchased directly from Covidien include completed form via fax or by email to : feedback.customerservice@covidien.com; if purchased from distributor complete form and contact your Distributor directly (completed form and all affected units must be returned through the Distributor); and reply to Covidien WHETHER OR NOT you have affect product at your site. If you have any questions or concerns, please do not hesitate to contact your Covidien representative or Quality Assurance, Monday through Friday, 8:30-5:00, directly at 203-492-5232. |  | 184934 | 184,934 units | Worldwide distribution: USA (nationwide) and countries including: Australia, Brazil, Canada, Columbia, Deutchland, France, Japan, Mexico, Panama, Puerto Rico, Singapore, and Uruguay. | Terminated | September 06, 2013 | 401 | 184934
Z-2304-2012 | 60514 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL80 FLEX (with FLEX software) and ABL80 FLEX CO-OX (with CO-OX software) - pH Blood Gas and Electrolytes Analyzer System. |  |  | June 06, 2012 | 2012 | Sendx Medical Inc | The firm recalled because the glucose measurements from a patient sample that the customer felt was too low and did not reflect the patient's clinical condition. | Software design | Radiometer sent a Field Safety Notice letter dated June 06, 2012 via email to all affected customers. The letter identified the affected product, problem and actions to be taken. Also included is a Note to Users to be placed at the front of customers existing Operator's Manual. Customers were instructed to provide confirmation to Radiometer that all customers have received the letter and the Note to Users. For questions contact your Radiometer representative. |  | 3599 | ABL80 FLEX is 1680, ABL80 CO-OX is 1759 | Worldwide Distribution - US Nationwide including the state of OH. | Terminated | September 04, 2012 | 90 | 3599
Z-2302-2012 | 62801 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKnee e+ Tibial Insert - HXLPE-VE, Sterile, Material: Highly cross linked vitamin E PE, Sz 2/9mm. |  |  | August 09, 2012 | 2012 | Encore Medical, Lp | The firm is recalling one lot of 392-09-706, 3DKnee e+ Tibial insert, 9 mm, Size 6, Right. The shelf life labeling of specific lots for a component were found to be incorrect and were found to be expired. | Incorrect or no expiration date | DJO Global sent an Urgent Field Safety Notice dated August 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the Notice to all those who need to be aware within their organization or to any organization where the potentially affected product has been transferred. Customers were instructed to contact the firm to place a replacement order and receive an RMA number. For questions customers were asked to call 512-834-6255.For questions regarding this recall call 512-832-9500. |  | 7 | 7 units | Worldwide Distribution including Venezuela. | Terminated | January 25, 2013 | 169 | 7
Z-2303-2012 | 62865 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Philips SureSigns VSi/VS2+ patient monitors |  |  | August 15, 2012 | 2012 | Philips Healthcare Inc. | Philips SureSigns VSi/VS2+ patient monitors may not be properly secured and may fall from the wall mount | Process design | Philips sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated August 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers using the VSi/VS2+ with the GCX wall mount are instructed to complete and fax a Customer Reply Wall Mount Order Card to the Business Unit or Key Market indicating the number of wall mounts in use. Upon receipt of the Customer Reply Wall Mount Order Card, Philips will send a replacement of the mounting hardware. Customers should contact their local Philips representative for questions regarding this recall. |  | 2314 | 2314 units | Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, DENMARK, EL SALVADOR, FRANCE, GERMANY, ICELAND, INDIA, IRELAND ISRAEL, ITALY, JORDAN, KENYA, LEBANON, MALAYSIA,NAMIBIA NETHERLANDS, NORWAY, POLAND, PORTUGAL, RWANDA SAUDI ARABIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | March 13, 2014 | 575 | 2314
Z-2305-2012 | 62709 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | syngo(R) Lab Data Manager and syngo(R) Lab Process Manager, all versions. |  |  | May 29, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | When using the syngo Quality Control package the system may not perform as intended for Multi-Rule QC violation [e.g., 2(2S), 4(1s) and 10x], unless configured using the instructions on page 2 of the recall notification. QC status may not be flagged appropriately. If auto-verification rules are defined to hold patient results when QC status is out, the patient results will not be held as expecte | Software design | Siemens sent a Urgent Field Safety Notice letter dated May, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed the customers to re-set the Multi-Rule option and instructions are provided on how to do so. This recall notification also included a Field Safety Effectiveness check form to please fax to the Siemens Technical Solutions Center at 302-631-8467 BY June 15th. |  | 44 | 44 | USA ( nationwide ) including the states of AZ, CA, FL, IL, KY, MD, MA, MN, MS, MO, NE, NH, NJ, NY, NC, OH, PA, TX, VA and WV. | Terminated | April 10, 2014 | 681 | 44
Z-2307-2012 | 62537 | No | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Poly stat Flu A &amp; B Test |  |  | March 09, 2012 | 2012 | Princeton Biomeditech Corp | There have been reports of dual positive (Flu A + Flu B positive) results being encountered. | Nonconforming Material/Component | OraSure Technologies, Inc. sent an "IMPORTANT NOTICE FOR ORASURE QUICKFLU RAPID FLU A + B TEST" letter dated March 13, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 610-882-1820 for questions concerning this recall. |  | 103 | 103 kits (22 tests per kit) | Nationwide Distribution-including the states of NJ, NY, and PA. | Terminated | October 14, 2014 | 949 | 563
Z-2260-2012 | 62848 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | V60 Ventilator, Continuous Ventilator, Model# V60, PN 1056490 |  |  | July 31, 2012 | 2012 | Respironics California Inc | Respironics is recalling the V60 ventilator because certain blower motor assemblies in the V60 ventilators may not meet specifications. Specifically, the impeller may not be properly seated on the blower motor shaft. This could subsequently cause the blower to cease functioning which would result in the failure of the V60 ventilator to deliver therapy to the patient due to the loss of ventilatio | Nonconforming Material/Component | Respironics California a division of Philips Healthcare sent a Medical Device Correction letter dated July 31, 2012 to all affected customers. The letter identified the affected product, problem identified and actions to be taken. The letter states that a Phillips Respironics Field Service Engineer or distributor will be contacting customers to schedule a no-cost replacement of the blower motor assembly on all affected V60 ventilators. For information or support concerning the recall issue, contact Philips Respironics at (800) 722-9377. |  | 197 | 197 units total (116 in the US) | Worldwide Distribution - US (Nationwide) and the countries of: Australia, Belgium, Canada, China, Germany, Italy, Japan, Kuwait, Netherlands, Norway. Qatar, Saudi Arabia, Turkey, and United Kingdom. | Terminated | March 22, 2013 | 234 | 197
Z-2309-2012 | 61451 | Yes | NIO |  |  | Cardiovascular | PMA | Zilver 518 Vascular Self Expanding Stent |  |  | January 05, 2012 | 2012 | Cook Ireland Ltd. | Cook Ireland received a product complaint from Japan on December 22nd, 2011. The complainant reported that a 5 French Zilver Vascular Stent (part number ZIV5-18-125-8-80 from Lot# CF694429) was inserted into a sheath over a wire guide where it became lodged in the sheath hub. After further examination, the user determined that the product was a 6 French Zilver Vascular Stent (part number ZIV6-35 | Packaging process control | The firm, Cook Medical, sent a letter dated February 13, 2012 to foreign consignee. The letter described the product, problem and actions taken. The firm stated that re-training of the relevant personnel has been carried out; they will continue to monitor to identify potential emerging trends and further corrective and preventive action will be considered in an effort to avoid the recurrence of similar events.If you have any question contact in US: Customer Relations Manager at 800.346.2686.ext 2222 or email: rita.harden@cookmedical.com; OUS: Customer Quality Supervisor at +353 61 239343. |  | 15 | 15 units of ZIV5-18-125-8-80 LOT# CF694429 | Worldwide distribution: USA (nationwide) and country of: Japan. | Terminated | February 04, 2013 | 396 | 15
Z-2250-2012 | 62856 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT70 and HT70 Plus Ventilators, Model Number: HT70 and HT70 Plus. |  |  | August 03, 2012 | 2012 | Newport Medical Instruments Inc | May emit a continuous high priority alarm and the ventilator may stop ventilating, due to a component failure on the control board. | Nonconforming Material/Component | Newport Medical sent an Urgent Field Correction Notice letter dated August 3, 2012 to all affected customers, via DHL and FedEx courier services. The letter identified the affected product, reason for recall and actions to be taken. Customers were instructed to immediately stop using the affected ventilator and retrieve the replacement kit provided free of charge. In addition, the letter states that if the product was further distributed, to contact their customers and notify them of the recall. Customers were asked to complete and return the enclosed Replacement form as replacements are completed. For questions contact Newport Medical Technical Support Department at 800-451-3111, extension 282, for U.S. customers; 714-427-5811, extension 500, for international customers.". |  | 72 | 72 | Worldwide distribution: USA (nationwide) and countries of: Brazil, Chile, Egypt, France, Hong Kong, India, Japan, Norway, Philippines and South Africa. | Terminated | April 14, 2014 | 619 | 72
Z-2310-2012 | 61984 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OPTETRAK LOGIC TIBIAL INSERT POSTERIOR STABILIZED***Match Component Colors 6 YELLOW, 11mm, Use with LOGIC Femoral and Tibial Components***REF 02-012-35-6011. |  |  | August 19, 2011 | 2011 | Exactech, Inc. | Exactech, Inc. of Gainesville, FL is recalling their Optetrak Logic Posterior Stabilized Tibial Insert, Size 6 after the device was determined to be out of dimensional specification. | Device Design | Exactech, Inc., sent an "IMPORTANT PRODUCT MARKET WITHDRAWAL NOTICE" letter dated August 19, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to: 1. Immediately cease distribution or use of the products. 2. Extend the information to accounts that are in possession of the products. 3. Verify if they had any of the affected products on the list. 4. Fax back the attached form. In addition, they were instructed to contact the Exactech inventory representative to confirm quantities at their location. If you have any questions regarding inventory restocking issue, please call -800-392-2832. |  | 193 | 193 | USA ( nationwide) WA, OH, OK, TX, NC, NY, CO, VA, MA, IL, TN, MN, NJ and FL. | Terminated | November 07, 2013 | 811 | 193
Z-2312-2012 | 62982 | No | IZD | Nuclear uptake probe. | Radiology | Radiology | 510(K) Exempt | Capintec Captus 3000 Thyroid Uptake System |  |  | July 24, 2012 | 2012 | Capintec Inc | Customers who received a Capintec Captus 3000 Thyroid Uptake System with Software Revisions 1.22, 1.23, 1.24, 1.25, and 1.26 may encounter an error while using the thyroid uptake protocol ?Measure Liquid and Subtract Residual." This protocol was a new feature added to Software Revision 1.22 in January 2011. The error occurs only if the user does NOT measure the residual dose immediately after me | Software design | The firm, CAPINTEC, INC., sent a "SOFTWARE ERROR ALERT" letter dated July 20, 2012 to all affected consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the four measuring steps outlined in the letter; DO NOT PRINT, EXIT THE THYROID UPTAKE MODULE, OR MEASURE ANOTHER PATIENT BEFORE YOU HAVE COMPLETED THESE FOUR MEASUREMENT STEPS; and follow the attached addendum for step-by step instructions. Capintec will send corrected software to all affected customers in approximately 2-3 weeks. The customers are instructed to contact Capintec Customer Support Representative if you have any questions or concerns regarding this notification, or if you do not receive your replacement software. Capintec Customer Support can be reached by phone at 1-800-631-3826 or by email at getinfo@capintec.com or by visiting our website at www.capintec.com. |  | 256 | 256 | Worldwide distribution: USA (nationwide) and countries of: Argentina, Belgium, Brazil, Canada, China, Greece, Hong Kong, India, Israel, Malaysia, Philippines, Spain, South Korea, Thailand, and Turkey. | Terminated | December 10, 2013 | 504 | 256
Z-2311-2012 | 62593 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Acrobat 3000 (AC3000) Spring Arm as support arm for iLED operating lights, TRUMPF Medizinsysteme GmbH + Co KG 82178 Puchheim. |  |  | May 29, 2012 | 2012 | Trumpf Medical Systems, Inc. | Cracks can arise on the welded seam on the rear joint. | Process control | TRUMPF sent an Urgent Safety Information letter dated February 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that within their organization, all users of the affected product be informed about the Urgent Safety Information. If affected product was further distributed, customers should forward a copy of the information or inform the necessary parties. Customers were asked to confirm receipt of the Urgent Safety Information by completing and returning the confirmation of receipt.For questions regarding this recall call 888-474-9359. |  | 402 | 402 units | Worldwide Distribution - USA including CA, OH, OK, NY, PA, TX, WY and the country of Canada | Terminated | April 11, 2013 | 317 | 402
Z-2315-2012 | 62928 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Toshiba Ultra-Short Magnetic Resonance Imaging System, EXCELART Vantage. MRT-1503/P5-Vantage MRI System. |  |  | February 21, 2012 | 2012 | Toshiba American Medical Systems Inc | The firm recalled due to a potentail problem with the position of the hold in the lug teminals on the power cables connected to the terminal blocks on the gradient coil. | Component change control | Toshiba America Medical Systems, Inc, sent a Urgent Medical Device Correction letter dated February 21, 2012, via USPS return receipt mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed if you subsequently find that images are abnormally noisy compared to images acquired under the same conditions in the system before this modification was performed, contact your Toshiba Service Representative.Please share this information with all users and reviewing radiologist as well as clinical engineering or biomedical group at your facility. Please complete and return the attached form and fax it to the toll free number at the top of the form. This form can also be sent via email toraffairs@tams.com.Contact information was also provided (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 3 | 3 units | US (nationwide) distribution including the states of Ohio and Texas. | Terminated | September 07, 2012 | 199 | 3
Z-2316-2012 | 62505 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD GeneOhm MRSA ACP Assay Catalog #441637, box 48 tests labeled in part***GeneOhm Sciences Canada, Inc., 2555 boul, du Parc-Technologique, Quebec, QC, Canada G1P 4S5*** |  |  | May 22, 2012 | 2012 | Becton Dickinson &amp; Co. | Leakage in Cepheid SmartCycler reaction tubes. | Nonconforming Material/Component | BD Diagnostic Systems issued an Urgent Product Recall letter via UPS overnight delivery on May 22, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions:Discontinue use and discard any affected product. BD would send replacement tubes for the discarded material.Complete the attached form whether or not they have any inventory remaining so that the firm may acknowledge their receipt of the notification.Fax the form to:Attention: Regulatory ComplianceFAX: 410-316-4258Email: RegulatoryComplianceFax@bd.comIf further assistance is needed, customers should contact BD Technical Services Department at 1-800-838-8663.For questions regarding this recall call 410-316-4054. |  | 38 | 38 kits | Nationwide Distribution including CA, DE, FL, GA, IL, IN, KS, KY, ME, MT, NC, NJ, NM, NY, OH, OK, PA, SC, TX, VA, WA, WI, and WV. | Terminated | November 02, 2012 | 164 | 200
Z-2321-2012 | 62891 | No | FIP | Water purification system for hemodialysis. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AmeriWater MRO Dialysis RO+ System, AmeriWater 1303 Stanley Avenue, Dayton, OH 45404 (800) 535-5585 www.AMERIWATER.com. |  |  | August 09, 2012 | 2012 | AmeriWater Inc | The firm was notified that a malfunction of the device was caused by an incorrect fuse 32 volt 15 amp fuses instead of 250 volt 15 amp fuses, that was installed in the water system. The incorrect fuse may result in excessive heat in the fuse holder eventually resulting in failure of the device to operate. | Nonconforming Material/Component | The firm, AmeriWater, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 9, 2012, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the product subject to recall; remove affected devices from service until the correction is completed to prevent malfunction of the device during patient treatment; and complete and return the attached response form to AmeriWater within 30 days via fax to 937/461-1988 Attn: Quality Assurance. AmeriWater dealer will be contacting their customers to schedule replacement of the fuses.If you have any questions, contact your local AmeriWater dealer or AmeriWater customer service at 1-800-535-5585. |  | 80 | 80 Units | Worldwide distribution: USA (nationwide) including states of: AL, AR, CO, CT, FL, GA, IL, IN, KS, MA, MD, MO, MS, NC, NY, OH, OK, SC, TN, TX and VA; and country of: Canada. | Terminated | April 17, 2014 | 616 | 80
Z-2322-2012 | 63042 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilett Mira. |  |  | August 21, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During regular product monitoring, Siemens has become aware of a potential issue with the Mobilett Mira system. During normal clinical operation the switch on the motor drive handle of the unit may become loose and slide horizontally to the right and become trapped under the cover. This may cause the system to continuously move. In response, Siemens has issued Update Instruction XP038/12/S to r | Device Design | Siemens initiated their correction of the MOBILETT Mira on August 21, 2012 by sending an "UPDATE INSTRUCTIONS XP038/12/S to all affected customers. The Instruction identifies the product, problem, and work steps for customers to follow. A XP038/12/S Completion Protocol letter was included for customers to complete and file with the instruction once the update is completed. |  | 20 | 20 | Nationwide Distribution- including the states of GA, ID, IL, IA, KS, NC, OH, SD, TX, and WI. | Terminated | April 03, 2014 | 590 | 20
Z-2326-2012 | 58247 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Polar Dressing, Knee/MU/Shldr Intelli-Flo-Breg, PN 10630 |  |  | August 11, 2010 | 2010 | Breg Inc, An Orthofix Company | This recall has been initiated due to confirmation that the Polar Dressing, Knee/MU/Shldr Intelli-Flo-Breg does not meet the required thickness as specified. The dressing in the sterile pouch measures a thickness of 0.010" instead of 0.020". This change in thickness makes the dressingthinner, less durable and subject to tearing. The product is mislabeled. | Nonconforming Material/Component | The firm, Breg Orthofix, called their customers on August 11, 2010 sent a letter dated August 12, 2010 and emails to all customers. The call and letter described the product, problem and actions to be taken. The customers were instructed to take immediate action to identify and quarantine the product at their facility. Breg established a return materials authorization and is contacting their customers to provide a Return Authorization to return the dressings. Breg is also providing replacement dressings. If you have any questions regarding communication or the Intelli-Flo Sterile Dressing, contact Breg Customer Relations at (800) 321-0607. |  | 502 | 502 units | Worldwide distribution: USA (nationwide) and country of: Nova Scotia. | Terminated | September 06, 2012 | 757 | 502
Z-2313-2012 | 62702 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Multi-Diluent 1 - Used for automated dilutions for the ADVIA Centaur BNP and TSH3 Ultra assays when on-board the ADVIA Centaur or ADVIA Centaur XP systems. |  |  | July 25, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Decrease in onboard dilution recovery when using Multi-Diluent 1 that has been stored on board the ADVIA Centaur and ADVIA Centaur XP systems for for BNP and TSH3 Ultra assays | Nonconforming Material/Component | Siemens sent an "URGENT DEVICE RECALL NOTICE" dated July 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Correction Effectiveness Check Response Form was included for customers to complete and return via fax to (302) 631-7597. Contact your local Sales Representative for questions regarding this notice. |  | 7414 | 2 Packs = 15879 and 6 Packs = 7414 | Worldwide Distribution-USA (nationwide) and the countries of Canada, Argentina, Australia, Brazil, Chile, China, Croatia, Denmark, Ecuduar, Egypt,Guatemala, Hong Kong,India,Ireland, Israel,India, Japan,Korea, Mexico, Malyasia, Peru, Pakistan, Saudi Arabia,Singapore, Thailand,Uruguay, Viet Nam, and South Africa. | Terminated | November 04, 2015 | 1197 | 7414
Z-2314-2012 | 62949 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | TxRx Knee 15 Coil 3T. |  |  | July 27, 2012 | 2012 | Quality Electrodynamics LLC | On July 27, 2012, it was reported to the firm that the Siemens 3T 15ch Knee Coil was not functional because the coil identification code for the device was incorrect. | Error in labeling | Since all of the devices were located at various Siemens distribution locations, all communications were internal to Siemens concerning bringing the coils from one warehouse to another. |  | 17 | 17 units | Worldwide distribution: USA (nationwide): Tennessee only and country of: Germany | Terminated | February 19, 2013 | 207 | 17
Z-2323-2012 | 62757 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | VersaTREK Windows Software, Version 5.4.3 is a component of the VersaTrek System. The Software CD will be individually packaged using a standard CD plastic case and placed into a standard CD mailer |  |  | July 23, 2012 | 2012 | Trek Diagnostic Systems | Review of instrument software databases and investigation of two customer reports determined that on rare occasions (less than 0.3%), following the removal of a positive signaling bottle from the VersaTREK instrument, the positive test result inadvertently reverted to a negative result in the instrument database. | Software design | ThermoFisher Scientific sent an Important Medical Device Product Correction Notice dated July 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact the Technical Service Department at 800-642-7029 to schedule their upgrade and to complete and return the Correction Notice Customer Acknowledgement by fax or in the enclosed return envelope. Customers were also asked to return the checklist. For questions customers were instructed to call 800-642-7029.For questions regarding this recall call 913-985-4185. |  | 333 | 333 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO , CT, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, RI, SD, TN, TX, VA, VT, WA, WI, WV, and WY. Internationally to the countries of Bahamas, Bangladesh, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, India, Ireland, Italy, Japan, Jordan, Kenya, Kuwait, Malaysia, Mexico, Pakistan, Peru, Philippines, Portugal, Puerto Rico, Russian Federation, South Africa, Spain, Thailand, Turkey, and United Kingdom | Terminated | June 18, 2013 | 330 | 333
Z-2325-2012 | 62852 | No | LRT | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | 100 ft roll of tubing inside a cardboard box. Medical devices are sealed into the tubes prior to ethylene oxide sterilization. Tube maintains sterility of the enclosed medical devices until opened. |  |  | August 02, 2012 | 2012 | Steris Corporation | On 7/20/2012 the firm became aware that a lot of tubing was distributed by the supplier after it was rejected by Steris due to failing process indicators. | Nonconforming Material/Component | Steris sent a Urgent Voluntary Recall Notice letter dated August 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Steris has learned that the VIS-U-All tubing is not performing to product specifications. Upon receipt of replacement product we ask that you destroy any remaining inventory of the affected lot in your possession. In the interim period, should you receive a failing result following a sterilization cycle, the results must be handled in accordance with your internal hospital procedures.We apologize for any inconvenience this matter may cause you. For further questions please call 1-(800) 548-4873. |  | 77 | 77 packs | Worldwide Distribution--USA (nationwide) including the states of AL, AZ, CA, FL, IL, IN, KY, LA, MO, MS, NJ, OH, PA, SD, TX, and VA., and the countries of Japan, Peru, and Suriname. | Terminated | September 04, 2013 | 398 | 77
Z-2324-2012 | 62967 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Brand name: RayStation. Generic name: Treatment Planning System. Software version and build numbers are 2.5.0.144 and 2.5.1.89. Product lifetime is 5 years. |  |  | May 09, 2012 | 2012 | RAYSEARCH LABORATORIES AB | The firm decided to recall when an issue was found with the dose calculation in RayStation 2.5, Raystation 2.5.1. The problem has not caused any patient mistreatment or other incidents. However, the user must be aware of this information to avoid incorrect dose calculations during treatment planning. The dose calculation is not correctly normalized in the unusual case of preventing the optimizatio | Device Design | The firm, RaySearch Laboratories, sent an "Urgent Field Safety Notice, Medical Device Correction" letter dated May 9, 2012 to its customers. The letter describes the product, problem and actions to be taken. The letter informed the customers that the "workaround is to not use segment MU optimization" and that there "is in general no need to manually adjust the segment shapes to such an extent there is need for a separate optimization of the segment MU." Additionally, the customers were instructed -Do not uncheck the "Segment shapes" check box; educate planning staff and all users about this workaround; inspect your product and identify all installed units with the noted software version numbers; and complete and return the Reply Form via fax to: RaySearch Americas Inc. at 888-501-7195; email: freddie.cardel@raysearchlabs.com and/or email to: Raysearch Laboratories AB at niclas.borglund@raysearchlabs.com.For further information and return of the Reply Form contact For the Americas: Director of Customer Support, RaySearch Americas at +1877 778 3849 or freddie.cardel@raysearchlabs.com; and For Europe, Asia and rest of the world: Director of Service, RaySearch Laboratories AB at +46 8 5450 6130 or niclas.borglund@raysearchlabs.com. |  | 19 | 19 units | Worldwide distribution: USA (nationwide) including states of: FL, MA, NJ, and NY; and countries of: Canada Germany, Netherlands and Switzerland. | Terminated | January 25, 2013 | 261 | 19
Z-2338-2012 | 62925 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Cassette, Batch # 12A078. |  |  | June 19, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling (removing) the Sterrad 100S and Sterrad NX System cassettes because the camera used to verify printing on the Sterrad System cassettes was inadvertently disabled. The STERRAD 100S and STERRAD NX System cases were shipped without verification of proper printing. | Error in labeling | Advanced Sterilization Products sent an "URGENT MEDICAL DEVICE RECALL (REMOVAL) letter dated July 16, 2012 to all their customers who purchased the STERRAD 100S and STERRAD NX System cassettes. The letter informed the customers of the problem identified and the actions to be taken. For the STERRAD NX System cassettes, there should be no impact to the functionality of these affected cassettes. Although printing may be smeared, illegible, or missing, the packaging has multiple references to the information, and the cassettes should be usable despite any printing error.For the STERRAD 100S System cassettes, there is a chance that the barcode will be unreadable and the STERRAD 100S Sterilizer will not accept the cassette.Customers were instructed to complete and return the business reply card and packing slip that are included in the letter. Customers were instructed to contact Stericycle at (888) 406-9303 for additional labels. Customers with questions about the recall letter were instructed to contact Stericycle at (888) 406-9303. |  | 7550 | 7,550 units | Nationwide Distribution. | Terminated | November 08, 2013 | 507 | 15045
Z-2335-2012 | 62895 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | The SYSTEM 1E Liquid Sterilant Processing System. |  |  | August 01, 2012 | 2012 | Steris Corporation | The firm initiated a recall after complaint investigations revealed that when the units are left on and not in use at night and/or on weekends the hose disconnects/separates resulting in water leakage that can cause damage to property. | Nonconforming Material/Component | STERIS sent a Urgent Voluntary Field Correction Notice letter dated August 1, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.STERIS has received no reports of injuries associated with the disconnection of a SYSTEM IE water hose.STERIS Corporation will instill new hoses and connections on your SYSTEM IE Liquid Chemical Sterilant Processing System. For further information or if you have any questions regarding our visit to your facility, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 6051 | 6,051 units | Worldwide Distribution -- USA (nationwide) and the countries of Bahrain and Hong Kong. | Terminated | June 06, 2014 | 674 | 6051
Z-2334-2012 | 61869 | No | LEJ | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SUN¦DASH RADIUS 252 tanning system; 52 lamps, Ultraviolet Radiation. |  |  | July 08, 2005 | 2005 | JK Products &amp; Services, Inc | The firm has tracked a systematic failure in the door cord on the Sun¦Dash Radius 252 product line. In some instances when the door cord is installed with little or no slack outside the upper decor piece, the cord has been known to fail causing a short in the wiring at the pivot point in the cord. | Equipment maintenance | JK Global Services sent out Mandatory Inspection/Modification letters to their Distribution Service Teams dated 07/08/2005, 01/12/2006, 09/15/2010, 10/04/2010 and 03/01/2012. A letter dated September 15, 2010, was sent to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to contact a JK representative or their Distributor as soon as possible to schedule an inspection to the affected product in their possession. For any questions customers were instructed to call 1-800-445-0624 or send an e-mail to support@jkamerica.com. |  | 1948 | 1948 units | Worldwide Distribution - USA (nationwide) and Canada | Terminated | January 13, 2017 | 4207 | 1948
Z-2332-2012 | 62768 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory |  |  | March 15, 2012 | 2012 | Sunquest Information Systems, Inc. | Sunquest is recalling Sunquest Laboratory versions 6.1.0125 through 6.1.0139, 6.2.1, 6.2.2, 7.0, and 7.0.1 because when the Windows-based merge application is used and the "Event merge across patients" option is used, a receiving system (for example a HIS or CoPath system) may interpret the event merge as a full patient merge. This occurs because a trigger from the Laboratory system for the event | Software design | The firm, Sunquest, sent an "Urgent- Product Safety Notice: PSN-12-02" dated March 15, 2012 to all their customers who purchased the Sunquest Laboratory versions 6.1.0125 through 6.1.0139, 6.2.1, 6.2.2, 7.0, and 7.0.1. The notice described the product, the problems and the actions to be taken. Customers were instructed to request the software correction by creating a Case via the Client Support webpage at: www.sunquestinfo.com/supportweb. Customers with questions were instructed to contact Client support at (877) 239-6337 for US and Canadian Sites and at +44 (0) 161 335 0562 for international sites. |  | 126 | 126 sites | Worldwide distribution: USA (nationwide) and countries of: England, United Kingdom and United Arab Emirates. | Terminated | October 02, 2012 | 201 | 126
Z-2336-2012 | 62970 | No | EWL | Laryngeal prosthesis (Taub design). | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | BLOM-SINGER indwelling TEP Occluder, catalog IN2004-TO |  |  | March 08, 2012 | 2012 | Helix Medical LLC | Helix Medical LLC is recalling the BLOM-SINGER indwelling TEP Occluder due to an invalid expiration date. | Labeling mix-ups | A recall letter dated March 8, 2012, was sent to customers who purchased the Blom-Singer Rapid Response Voice Prosthesis TEP Occluder 4mm 20Fr. (IN2004-TO). The letter informed the customers of the problem identified and the action to be taken. Customers were instructed to contact Customer Service Center at (800) 477-5969 if they have products to be returned. Customers were instructed to return Blom-Singer Rapid Response Voice Prosthesis TEP Occluder 4mm 20Fr. (IN2004-TO) to:Helix Medical, LLCAttn to: Customer Service-IN2004-TO.1110 Mark AvenueCarpinteria, CA 93013For questions regarding this recall call 805-684-3304. |  | 2 | 2 devices | Worldwide distribution: USA (nationwide) including state of: Ohio and country of United Kingdom | Terminated | April 02, 2013 | 390 | 2
Z-2331-2012 | 62875 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CMS-2981-R3 Arthrogram Tray containing 2 oz. bottle PVP Prep Solution |  |  | September 30, 2011 | 2011 | Custom Medical Specialties, Inc. | The firm saw on the FDA web site a recall by the manufacturer due to no microbial testing in place at the manufacturer and product did not meet proper finished goods specifications. | Nonconforming Material/Component | Custom Medical Specialties sent an Urgent Medical Device Recall letter dated August 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their inventory for the affected product and to discontinue the use of any affected product. Customer were instructedto complete and return the Return Response form, acknowledging they have received the letter and checked their inventory. Any product still in the customers inventory should be returned to the firm. Customers should call (919) 202-8462, ext 205 to obtain a Return Authorization Number and the shipping number.For questions regarding this recall call 919-202-8462. |  | 15 | 15 cases | Nationwide Distribution including GA, MI, NY, and PA. | Terminated | September 10, 2012 | 346 | 38
Z-2333-2012 | 62842 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Aesculap S4 Element 4.5mm Polyaxial Screws (ST240T, ST241T, ST242T, ST243T, ST244T, and ST245T). |  |  | July 25, 2012 | 2012 | Aesculap Implant Systems LLC | There is a possibility that the screw channel that holds the rod in place may be out of tolerance. When the depth of the channel is not deep enough, the rod may not be able to provide ample compression on the insert preventing the screw head from locking down to the bone screw. | Nonconforming Material/Component | Aesculap sent an "IMPORTANT PRODUCT REMOVAL NOTIFICATION" letter dated July 25, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Distribution Inventory Sheet was attached for customers to complete and return with the affected product. Contact the firm at (610) 984-9275 for questions regarding this recall. |  | 5 | 5 | Nationwide Distribution-including the states of IL, MN, NV, OK, and TX. | Terminated | March 18, 2014 | 601 | 5
Z-2327-2012 | 62774 | No | EZC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Foley procedural trays packaged in breather bags. The product is packaged ten units in each shipper box, and are sold sterile for single use. |  |  | June 29, 2012 | 2012 | C.R. Bard, Inc., Urological Division | Incomplete seal on sterile packages of product. | Process control | C. R. Bard Inc, sent a Urgent Product Recall Notification letter dated June 27. 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This notification is being sent because there is a potential for a limited number of units within the subject lots (noted above) that may contain an open seal of the sterile packages. The issue is visually obvious even before the package is opened. No patient injury or adverse events have been reported. Do not use or further distribute any of the affected products.Once the product affected by tlle recall has been removed from your inventory 1.Fill out the Recall &amp; Effectiveness Check Form. Be sure to state the quantities affected of the recalled product that you have in stock. It is extremely important that we receive this information. 2. Fax the Recall &amp; Effectiveness Check Form to BMD at 1-770-784-6469. The Recall Coordinator will issue you with a Return Authorization (RCL) to facilitate the expedient return and replacement of the product. 3. A mailing label is enclosed for your convenience to return the affected product. Please mark the outside package as "RECALLED PRODUCT" and include the RCL number. All products should be retumed to the following shipping address that is in the letter. |  | 56694 | 56,694 units | Nationwide distribution: USA | Terminated | February 07, 2013 | 223 | 56694
Z-2346-2012 | 62909 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Sterile Empty Vial and Injector (30mL PCA vial), Rx only, One/List No. 6021-03, Hospira, Inc., Lake Forest, IL 60045. |  |  | August 02, 2012 | 2012 | Hospira, Inc. | Reports of leaking during filling and administration. | Nonconforming Material/Component | Hospira sent an Urgent Device Field Correction letter dated August 3, 2012, to all affected customers. The letter identifed the product, the probem, and the action to be taken by the customer. Customers were instructed to first consider the use of Prefilled Morphine PCA Vials wherever possible. In some instances, the treating physician may require an alternative to morphine.Customers may continue using the affected product by following the instructions provided in the Correction Letter. Customers were asked to complete the attached Reply Form and fax it to 1-888-879-8144.For questions regarding this recall call 877-946-7747. |  |  | unknown | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NY, NC, NK, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY and the countries of Canada, Costa Rica and the Virgin Islands. | Terminated | May 08, 2014 | 644
Z-2340-2012 | 62915 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Multileaf Collimator |  |  | July 27, 2012 | 2012 | Elekta, Inc. | Information released with the product Agility did not originally contain information on compatibility with the add-on Dynamic Multileaf Collimator Apex. | Error in labeling | ELEKTA sent an Important Field Safety Notice, FCO Ref: 20001405000 dated July 7, 2012, to current Agility users and a separate notice FCO Ref: 20001407026 for current Apex users who may be in the process of considering an upgrade to Agility. Both notices made it clear that Apex has not been integrated or classified as compatible for use with Agility. Users are advised not to use Apex with Agility.For questions regarding this recall call 770-300-9725. |  | 2 | 2 | Nationwide Distribution including DC &amp; MD | Terminated | October 06, 2014 | 801 | 2
Z-2347-2012 | 62328 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx monitor/defibrillators |  |  | June 20, 2012 | 2012 | Philips Healthcare Inc. | Therapy Cable Connection wear may lead to a malfunction in detection of defibrillation pads/paddles therapy Cables | Device Design | Philips notified customers on June 20, 2012, by an Urgent Medical Device Correction notification/Field Safety Notice and a revised Instructions for Use addendum for each HeartStart MRx unit. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of units affected and and identifies the action. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:Carefully read the attached HeartStart MRx Instructions for Use Addendum. Immediately implement ongoing therapy connection inspection on all of your MRx devices to detect wear from higher than expected levels of stress. The HeartStart MRx Instructions for Use Addendum describes how a user can identify wear. If wear is detected, remove affected devices immediately from use and contact Philips to arrange for service.WARNING: The service life of your therapy cables/external paddles is up to three years. To maintain reliable performance and reduce the possibility of failure during patient use, replace them every three years from the time they were initially placed into service or if they fail the inspection criteria in the attached Instructions for Use AddendumThe addendum provides the customer with complete user information regarding ongoing inspections of port. In addition, service support will be provided for Therapy Port and Cables requirement replacements.Further information or support concerning this issue,contact your local Philips representative at 1-800-722-9377.For questions regarding this recall calll 978-687-1501. |  | 58701 | 58,701 units | Worldwide Distribution - USA (nationwide) and the countries of : AFGHANISTANALBANIAARGENTINAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHRAINBANGLADESHBELGIUMBOLIVIABOSNIA AND HERZEGOWINABRAZILBRUNEI DARUSSALAMBULGARIACAMBODIACANADACHILECHINACOLOMBIACOSTA RICACROATIA (local name: Hrvatska)CYPRUSCZECH REPUBLICDENMARKECUADOREGYPTEL SALVADORESTONIAETHIOPIAFINLANDFRANCEGERMANYGREECEGUADELOUPEHONG KONGICELANDINDIAINDONESIAIRAQIRELANDISRAELITALYJAPANJORDANKENYAKOREA, REPUBLIC OFKUWAITLAOSLATVIALEBANONLIBYAN ARAB JAMAHIRIYALITHUANIALUXEMBOURGMACAUMACEDONIAMALAYSIAMALTAMEXICOMOROCCONAMIBIANETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANIGERIANORWAYOMANPAKISTANPANAMAPAPUA NEW GUINEAPARAGUAYPERUPHILIPPINESPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIAN FEDERATIONSAUDI ARABIASINGAPORESLOVENIASOUTH AFRICASPAINSRI LANKASWEDENSWITZERLANDTAIWAN, PROVINCE OF CHINATANZANIA, UNITED REPUBLIC OFTHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYUGANDAUKRAINEUNITED ARAB EMIRATESUNITED KINGDOMUZBEKISTANVENEZUELAVIET NAMYEMEN | Open, Classified |  |  | 58701
Z-2344-2012 | 63008 | No | GWF | Evoked response electrical stimulator. | Neurology | Neurology | 510(k) | Interfaces subject to recall: |  |  | September 27, 2010 | 2010 | Medtronic Xomed, Inc. | In January 2010, Medtronic issued a recall of Medtronic NIM 3.0 Patient Interfaces after Medtronic identified a trend of open channel conditions (signal loss from the electrode). When the NIM 3.0 detects the signal loss from the electrode (i.e. open channel), it fails the electrode check and provides an audible and visual warning to the user which could also create a delay in the surgical procedu | Device Design | The firm, Medtronic USA, Inc., sent an "URGENT Product Recall Notification" letter dated September 27, 2010, to its U.S. and International customers. The letter described the product, problem and actions to be taken. The customers were instructed to verify their receipt of this notice and complete and sign the Product Replacement checklist form indicating that their sales representative has completed the replacement of their devices and fax to 1-651-367-8055.U.S. Customer Units - replacement kits with their contents will be provided to each Area Sales Manager (ASM) for each customer along with documentation of serial number units to be replaced. Each ASM will be responsible for delivering a replacement unit(s) and picking up and returning each of their customer's units using the pre-labeled replacement kit boxes.International Customer Units- International units, depending on the geography will either be replaced or repaired and returned. If replacement is planned, the International Business Manager (IBM) will receive replacement kits as described for the U.S. In those geographies where devices will be repaired and returned, kits with the following contents will be sent to IBM's along with documentation of serial number units to be repaired.Foreign distribution: In the larger territories (i.e. Australia, Japan, Canada, EMEA-Europe, Middle East, Africa), repairs/rework were handled at the Medtronic facilities in those geographic areas. In the smaller geographic areas (Singapore, Thailand, Hong Kong, Korea, Latin America ), where there were limited sales, devices were returned to the U.S. In all cases (large and small geographic areas), devices were returned via a Medtronic entity (sales office or distribution center).If you have any questions regarding the subject Actron or the content of this letter, contact the Distinguished Regulatory Affairs Advisor at 904-279-7532. |  | 7 | 7 | Worldwide distribution: USA (nationwide) and countries of: Australia, Austria, Belgium, Canada, Columbia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, Guatemala, Hong Kong, India, Iran, Ireland, Israel, Italy, Japan, Lebanon, Libya, Luxembourg, Netherlands, New Zealand, Norway, Portugal, Saudia Arabia, Singapore, Slovak, Spain, SRI Lanka, Sweden, Turkey, United Arab Emirates, and United Kingdom. | Terminated | November 26, 2013 | 1156 | 524
Z-2345-2012 | 63102 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano brand Intravascular Ultrasound, Volcano s5/s5i Family Software Version 3.2.1 or 3.2.2, Models/Part Numbers: 435-0607.01, 435-0607.04, 435-0602.01, 435-0602.02; Product is manufactured and distributed by Volcano Corporation, Rancho Cordova, CA. |  |  | August 24, 2012 | 2012 | Volcano Corporation | The display on the system has the potential to change. In house testing eventually identified the unique sequence of events that must occur in order for this issue to manifest itself. | Software design | Volcano Corporation sent an "IMPORTANT FIELD CORRECTIVE ACTION" letter dated August 24, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact your local Volcano sales representative or Volcano Customer Service for questions regarding this recall. |  | 1020 | 1020 units | Worldwide Distribution. | Terminated | November 25, 2014 | 823 | 1020
Z-2339-2012 | 62807 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Tip Cover for 8mm Monopolar Curved Scissors (Hot Shears)Accessory, Disposable, 8mm MCS, Box of 10. |  |  | August 09, 2012 | 2012 | Intuitive Surgical, Inc. | There is potential for the sterility of the product to be compromised. | Employee error | Intuitive Surgical sent a "MEDICAL DEVICE CORRECTION" letter dated August 9, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. An Acknowledgement Form was included for customers to complete and return via fax to (408) 716-3040. Contact Customer Service at 1-800-876-1310 for questions regarding this recall. |  | 11121 | 11,121 Boxes | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Canada, and Lebanon | Terminated | December 31, 2012 | 144 | 11121
Z-2341-2012 | 62988 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Uroskop Access. |  |  | April 30, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens has discovered a possible hazard to patients or hospital personnel when tilting the Uroskop Access system starting with serial number 6076. When tilting the table on the Uroskop Access table, the motor may become disconnected and the entire table assembly may fall on the floor. | Nonconforming Material/Component | Siemens sent a Update instructions letter dated April 2012, to all affected customers. The safety instruction letter identified the product the problem and the action needed to be taken by the customer.The update has been completed successfully. The customer has been informed of the benefits of improvements resulting from this update.Further questions please call (610) 219-4834. |  | 5 | 5 | Nationwide distribution: USA including states of: IN, MI, MS, TN and TX. | Terminated | September 20, 2012 | 143 | 5
Z-2367-2012 | 62989 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilett Mira mobile x-ray system |  |  | April 30, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During regular product monitoring, firm became aware of a potential issue with the Mobilett Mira system, (serial numbers 1034, 1038, 1041, 1042, 1043, 1050, 1053, and 1054. During normal clinical operation the center screw on the driving wheel may become loose resulting in system immobilization. There is no risk of the system tipping over or collapsing on the floor, however the system may tilt t | Device Design | Siemens has issued an Update Instruction XP018/12S in April 2012 to all affected consignees notifying them of the problem. The update instruction provides a correction of the wheel fixation. For questions call Siemens at 610-219-4834 or 610-448-4634. |  | 9 | 9 | US Nationwide Distribution including the states of: GA, IA, KS, NC and TX. | Terminated | September 25, 2012 | 148 | 9
Z-2362-2012 | 62140 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF 101-05-30***NOVATION MODULAR DRILL BIT LENGTH: 30mm, DIAMETER 3.2mm***Use with 30mm Screw***STERILE***EXACTECH 2320 NW 56th Court, Gainesville, FL 32653***. |  |  | January 24, 2012 | 2012 | Exactech, Inc. | Exactech Inc. of Gainesville, FL is recalling their Novation Modular Drill Bits after the devices were reported to have been used by end user while in a non-sterile condition. | Finished device change control | The firm, Exactech, Inc., called the lone customer on January 24, 2012 to discuss the product, problem and actions to be taken. The customer was instructed to return all unused devices and associated packaging. |  | 2 | 2 | Nationwide distribution: USA state of: Arkansas only | Terminated | September 26, 2012 | 246 | 2
Z-2368-2012 | 62926 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Enduramesh Corpectomy Spacer System |  |  | July 15, 2012 | 2012 | Lucero Medical LLC | The firm voluntarily recalled the device after they discovered that the sterilization recommendations on the product insert did not pass sterilization validation. | Error in labeling | Lucero Medical sent a Field Correction Notice letter dated July 15, 2012 to it's sole consignee. The notice identified the affected product., problem and actions to be taken. The customer was instructed to complete the response card and submit back to their local Lucero Medical Representative. |  | 169 | 169 | US Nationwide Distribution | Terminated | February 20, 2013 | 220 | 169
Z-2348-2012 | 62321 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Safety monitor included with System 8000 Base |  |  | August 12, 2009 | 2009 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received complaints from our Affiliate, TerumoDubai that the safety monitors were tripping off. Terumo CVS' investigation determined that 8monitors were manufactured with an underrated circuit breaker due to an error where an associatepulled an underrated circuit breaker from an incorrect inventory location. Since this time, inventorylocations hav | Mixed-up of materials/components | TERUMO Cardiovascular Systems communicated with their affiliate, Terumo Dubai, in 2009 at the time the complaints were reported. No further communication was warranted.For questions regarding this recall call 734-741-6173. |  | 8 | 8 | Worldwide distribution: International only-country of: Terumo Dubai, UAE | Terminated | January 29, 2013 | 1266 | 8
Z-2369-2012 | 54659 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 V1.05(FG-5400-00). |  |  | November 17, 2009 | 2009 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster has become aware of information regarding a potential hazard that may occur during the operation of a Bloom EP Stimulator device in conjunction with RF generators and other electrophysiology (EP) equipment, including the CARTO 3 Electropysiology Mapping System. | Labeling mix-ups | Biosense Webster sent an " URGENT FIELD SAFETY NOTICE" dated November 20, 2009 to all affected customers who purchased the Carto 3 Electrophysiology Mapping System. The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to continue use of the CARTO 3 System and that it was safe and no further action is needed. Customers with questions about the notice were instructed to contact Biosense Webster representatives.On January 12, 2010 another letter was sent from Biosense to all their customers recommending to avoid any connection of an EP stimulator's pacing pins to the direct ports on the CARTO 3 System when applying RF energy, unless the set up is indicated by the stimulator's instructions for use. The letter also informed the customer that the user manual for the CARTO 3 System has been formally revised to include a warning. Customers were instructed to complete an acknowledgement form to confirm receipt of the revised user manual. |  | 71 | 71 units total (30 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, Denmark, France, Germany, Israel, Italy, Netherlands, and United Kingdom. | Terminated | September 12, 2012 | 1030 | 71
Z-2376-2012 | 62760 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2121IQ, Innova 3131IQ Cardiovascular X-ray imaging systems. |  |  | August 16, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the Innova 2121IQ &amp; Innova 3131IQ Systems. While performing a fluoroscopic acquisition with the Innova 2121IQ or Innova 3131IQ systems, there is a potential for loss of usable imaging mode which requires the system to restart in order to recover. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated August 16, 2012, to all affected consignees. The letter described the Safety Issue, safety Instructions, Affected Product Details, Product Correction, and Contact Information. Consignees were instructed that hospital staff should wait for Photo or Image to appear on the display (when store photo function is performed) before start of the next short duration fluoroscopy acquisition. For questions customers were instructed to call:United States (800) 437-1171Japan: 0120 - 055 - 919For other countries consignees should contact their local GE Healthcare Service Representative.For questions regarding this recall call 262-513-4122. |  | 240 | 240 (133 US and 107 OUS) | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, HI, IK, IL IN, IA, KS, LA, MD, MA, MI, MN,MO, NV, NH, NH, NM, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, WV, and WI,.and internationally to:AUSTRALIA, AZERBAIJIN, BAHRAIN, BELARUS, BULGARIA, CANADA, CHILE, CHINA, CZECH REPUBLIC, EGYPT, FRANCE, GERMANY, INDIA, INDONESIA, IRAQ, IRELAND, ITALY, JAPAN, KUWAIT, LEBANON, LIBYA, MALAYSIA, MEXICO, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAÂżDI ARABIA, SINGAPORE, SOUTH KOREA, SPAIN, TAIWAN, THAILAND, UNITED ARAB EMIRATES, UNITED KINGDOM, and YEMEN. | Terminated | October 20, 2015 | 1160 | 240
Z-2372-2012 | 62912 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Omnilife Science Apex Knee System Modular Tibia Augment, Size 3 x 4mm |  |  | August 09, 2012 | 2012 | Omni Life Science | The Apex Knee System Modular Tibia Augment may breach the inner and outer sterile pouches, compromising the sterility | Packaging | Omnilife Science contacted affected customers by phone and sent an "URGENT MEDICAL DEVICE RECALL" e-mail on August 9, 2012. The e-mail identifies the product, problem, and actions to be taken by the customers. Customers with questions regarding product return should contact Customer Service at 800-448-6664. Customers with questions regarding the recall should contact the firm at 774-226-1845. |  | 310 | 310 units | Worldwide Distribution-USA (nationwide including the states of CO, NY, OK, PA, UT, VA, and WI, and the countries of Spain and the Netherlands. | Terminated | January 10, 2013 | 154 | 1396
Z-2256-2012 | 62946 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Optima XR200amx. Mobile general purpose radiographic imaging of the human head and body. |  |  | July 10, 2012 | 2012 | GE Healthcare, LLC | Software issue: a violation of 21 CFR 1020.31(a). After an image is taken, when the user selects and clears an additionally displayed screen the system inhibit is also cleared, allowing the user to take an exposure without the technique factors being displayed. | Radiation Control for Health and Safety Act | The firm will issue Important Electronic Product Radiation Warning letters that identify the affected product along with a description of the defect and related hazards. Users are instructed to ensure that the acquisition mode is always selected prior to attempting an exposure to confirm the techniques selected. The firm will remedy the issue or bring the device into compliance free of charge. A GE Healthcare Service Representative will update the software on the system to address the issue. Field Modification Instructions (FMI) 10865(2) describes the rework plan the GEHC Engineers will make to the imaging system to return it to compliance with the applicable performance standard. It appears to adequately address the problem and is hereby approved. Users are to contact 800-437-1171 in the US, if they have any questions or concerns. |  | 254 | 254 total in the US | Worldwide Distribution -- USA (nationwide) Distribution | Terminated | January 10, 2013 | 184 | 254
Z-2377-2012 | 62844 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Capsule Neuron Docking Station, a component in the DataCaptor Connectivity System |  |  | August 10, 2012 | 2012 | Capsule Tech Inc. | Capsule Neuron Docking Station may experience a lost connection, resulting in loss of data or loss of power to the Caspsule Neuron | Device Design | The firm, Capsule Tech, sent a recall letter dated August 10,2012 to its customers. The letter described the product, problem and actions to be taken. Capsule is sending field correction kits to customers with instructions. Capsule's Professional Services team will contact the customers to schedule shipment of their kit. Final shipments of the field correction kits will be completed by Oct 2012. If you have any questions or concerns, contact your professional services representative and/or call in US: 978-482-2340 or France: +33 153-34-14-00. |  | 12448 | 12,448 units | Worldwide distribution: USA (nationwide) and countries of: Belgium, Canada, Denmark, France, Germany, Netherlands and UK. | Terminated | April 22, 2015 | 985 | 12448
Z-2382-2012 | 63007 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL 91493 Integrated Module Housing. |  |  | August 27, 2012 | 2012 | Spacelabs Healthcare, Llc | There is a potential failure of the Integrated Module Housing causing the patient module parameters to drop off the monitor display or the patient module will fail to sign on when inserted into the housing. | Mixed-up of materials/components | Spacelabs Healthcare sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated August 27, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-522-7025 and select 2 for assistance regarding this recall. |  | 527 | 527 units total; 350 US and 177 International | Worldwide Distribution. | Terminated | April 15, 2013 | 231 | 527
Z-2384-2012 | 62884 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DigitalDiagnost Systems with M Cabinet |  |  | August 14, 2012 | 2012 | Philips Healthcare Inc. | Creepage Distance between the two primary phases in the GEO Transformer is 4.5 mm.does not meet (IEC) International standard of 5.5 mm. | Device Design | Philips Healthcare Imaging issued an Urgent - Field Safety Notice dated August 8, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The unit will be exchanged with a new unit that complies with the IEC requirements. Please reference Field Change Order (FCO) 71200081 when contacting I your local Philips representative. 1-800-722-9377.Customer were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the letter and to maintain a copy of the letter with the equipment instruction for Use. For further assistance, customers should contact their local Phillips representative: Technical Support Line 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 4 | 4 systems | Worldwide Distribution - USA including OH and WA and the country of Canada | Open, Classified |  |  | 4
Z-2387-2012 | 62894 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | Instrumentation Laboratory ACL TOP (Base) |  |  | August 03, 2012 | 2012 | Instrumentation Laboratory Co. | Low Abnormal Fibrinogen Sample Reports as Normal and High Fibrinogen Sample Reports as Low | Software design | The firm, Instrumentation Laboratory, sent a "Mandatory Test Parameter Upgrade" letter dated June 7, 2012 with upgrade kits P-14.8 via USPS on August 3 thru August 10 to US and Canadian customers and notified International Affiliates via e-mail on August 2, 2012. The customers were ask to complete and return the Mandatory Test Parameter Upgrade Tracking Form via fax: 781-861-4207 or email: ra-usa@ilww.com. If you have any questions, call 781-861-4204. |  | 1074 | 1074 units | Worldwide distribution: USA (nationwide) and countries of: ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CHILE, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, REPUBLIC OF LUXEMBOURG, MACAU, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PANAMA, PERU, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SLOVAKIA (Slovak Republic), SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TURKEY, UNITED KINGDOM, URUGUAY and VIETNAM. | Terminated | October 02, 2015 | 1155 | 4297
Z-2378-2012 | 63110 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The product affected by this notice is sold under the trade name TomoTherapy SharePlan 1.1 . To determine if the version you are using is affected, bring up the About box in the SharePlan application and check if the build number reported there is 1.3.1.10 . |  |  | May 30, 2011 | 2011 | RAYSEARCH LABORATORIES AB | This field notice concerns two behaviors of the SharePlan (Multiple Fraction Groups and Depth Offset) that may be unexpected and are not clearly described in the user manual. None of these behaviors have caused patient mistreatment or other incidents. However, the described user actions must be adhered to for best agreement between treatment plan evaluation and the clinical outcome of the delivere | Labeling False and Misleading | RaySearch Laboratories sent an Urgent Field Safety Notice Medical Device Correction dated May 30, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consigness were instructed to completely avoid affecting the erroneous behaviour, verify in TomoTherapy Hi-Art that the plans sent to TomoTherapy SharePlan 1.1 only have one single fraction group with a single fraction size,used for all fractions, and do not attempt to import treatment plans having more than one fraction size into the product. To completed avoid affecting the erroneous behaviour the user can either offset the measured curves before importing them into beam commissioning or add the depth offset tothe calibration depth.For further information consigness should contact RaySearch Laboratories AB Quality and Regulatory Affairs at +46 722 366 110 or the Tomo Therapy support at support@tomotherapy.com.North America 1 866 368 4807Belgium 0800 38783France 0805 631 565Germany 0800 000 +541Italy 800-986 399Netherlands 0800 0201364Spain 800 3000049Switzerland 0800 001927UK 0868 238 6035China/Northern 10 800 712 1701China/Southern 10800 120 1701Hong Kong 800 967912Japan 0044 22 132374Singapore 800 1204683South Korea 0079 81 4800 7204All other locations + 1 608 824 2900 or +32 2 40 44 44 |  | 15 | 15 units | Worldwide Distribution - USA including CA, KY, MI, NY, OK, and OR. Internationally to Canada, Belgium, United Kingdom, Australia, France, Germany, and Poland. | Terminated | November 16, 2012 | 536 | 15
Z-2381-2012 | 62728 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The 2.7mm Drilling Guide is a reusable instrument used in conjunction with a drill to prepare oriented holes in bone for insertion of a Standard SurfixÂż Locking Screw with plates for the Tiblaxis, Uni-CP, Advansys MLP and Advansys DLP compression plate systems. |  |  | June 14, 2012 | 2012 | Integra Life Sci. | Integra' s quality system noted an adverse trend in complaints for breakage of the drilling guides during use. None of the complaints resulted in patient injuries but in some cases resulted in an increase in the time to complete the surgical procedure. A physician was consulted to assess patient risk and concluded that any potential for risk would be transient, self-limiting or minor, or result | Device Design | Integra sent a "SUBJECT: MEDICAL DEVICE RECALL" letter dated June 14, 2012 to all affected customers. The letter identified the product, problem, and action to be taken by the customers. Customers were asked to complete a Recall Acknowledgement and Return Form whether or not they were affected by the recall, and return all their affected inventory to Integra LifeSciences.Customers were asked to contact their regional managers or the firm at 609-936-6832 for questions regarding the recall. |  | 2389 | 2,389 units, Catalogue No. 219635ND; 2.7mm Drilling Guides | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NV, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI and the countries of Austria, Czech republic, Denmark, France, Germany, Ireland, Israel, Italy, Kingdom of Saudi Arabia, Lebanon, Norway, Portugal, Switzerland, South Africa, Spain, Sweden, and United Kingdom. | Terminated | February 20, 2013 | 251 | 2389
Z-2379-2012 | 62867 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HALO Flex Energy Generators |  |  | August 03, 2012 | 2012 | Covidien LLC | Potential failure of HALO Energy Generator to enter "stand by" mode when initially powered on; and therefore can not be used for patient care. | Software design | COVIDIEN sent an Urgent Medical Device Correction letter dated August 3, 2012, to all affected customers via Federal Express. Additionally, Field Service Representatives will be contacting their accounts to notify them of the correction.. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps prior to initiating patient treatment and specifically, prior to administering conscious sedation and performing upper endoscopy:1. Power-on the HALO Flex Energy Generator and confirm that the generator passes self-test, enters standby mode and displays "Ready, Connect Catheter."2. Confirm that the HALO Flex Energy Generator is functioning properly at power-up and enters the standby mode then the procedure can begin.3. Do not use the HALO Flex Energy Generator if the generator does not pass the power on self-test. In this situation, please call Covidien, GI Solutions customer service at (888) 662-2779 or your local territory manager (sales representative) so that the firm can provide them with a replacemenent generator.4. Customers were asked to reply to Covidien, GI Solutions (formerly BARRX Medical) by signing and returning the attached Product Correction Form via fax to the number provided on the form. Customers with questions were instructed to call 888-662-2779.For questions regarding this recall call 408-328-7300. |  | 617 | 617 generators | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Canada, Chile, Colombia, Czech Republic, Denmark, France, Germany, Greece, Hungary, India, Israel, Jordan, Netherlands, Norway, Poland, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Venezuela and UK. | Terminated | September 30, 2013 | 423 | 617
Z-2386-2012 | 63014 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice CT Scanner System, models 728131 and 728132. The MX 16-slice CT Scanner System can be used as a whole body computed tomography X-ray system. |  |  | August 24, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | On August 24, 2012 the firm recalled the MX 16-slice after discovering that a bolt which attaches the oil accumulator within the gantry may fail, resulting in the oil accumulator detaching from the tube housing assembly within the gantry. | Process design | An "Urgent Medical Device Correction" Field Safety Notice dated August 24, 2012 was sent to all US consignees via certified mail. The problem was described and recommended actions were provided. Philips Key Market representatives distributed the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Orders. US Customer contact number: 1 800-722-9377. |  | 328 | 328 units | Worldwide distribution, Nationwide (US) including the states AL, AR, CA, FL, GA, IL, IN, MS, NJ, NV, NY, OH, PA, and VA and the countries of Albania, Argentina, Australia, Belgium, Brazil, Chile, China, Colombia, Dominican Republic, Ecuador, France, Germany, Greece, India, Indonesia, Ireland, Italy, Jordan, Lebanon, Malawi, Mexico, Nepal, Netherlands, Oman, Palestine, Philippines, Poland, Romania, Russia, Saudi Arabia, South Africa, South Korea, Spain, Switzerland, Turkey, Uganda, and Venezuela. | Terminated | August 28, 2013 | 369 | 328
Z-2383-2012 | 63033 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac, Trilogy, Novalis Tx, Unique, linear accelerators; Model Number: H14, H18, H29, Reference/FSCA Identifier: CP-08881; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA. |  |  | August 14, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Some X-jaw (lower collimator jaw) carriers were made using the incorrect metal alloy, which can cause the jaw carrier to crack. If both carriers on the same X-jaw crack, the jaw would be able to move freely in the closed direction. The position readout interlock circuitry will not detect this jaw position deviation. This may lead to treatment with an under-dose of the intended target volume. | Nonconforming Material/Component | Varian sent an "URGENT MEDICAL DEVICE CORRECTION URGENT FIELD SAFETY NOTICE" dated August 14, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-888-827-4265 for questions regarding this recall. |  | 221 | 221 Units | Worldwide Distribution. | Terminated | December 04, 2013 | 477 | 221
Z-2385-2012 | 62899 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Sterile Distraction Screws of 3 different sizes: 12mm, 14mm and 16mm sold under TSI, 3COR, DB Surgical, and American Surgical Products brands and Sterile XL Screws of 7 different sizes: 90mm, 100mm, 110mm, 120mm, 130mm, 140mm, 150mm sold under TSI brand |  |  | June 11, 2012 | 2012 | Tedan Surgical Innovations Llc | This letter is to notify you that several of our sterile screw products are being voluntarily recalled. Product sterility may be compromised due to a mechanical failure of the protective packaging tube. Tube end plugs may dislodge creating the potential for the screw to come out of the protective tube and then the sharp end of the screw pierce the outer sterile barrier pouch. | Packaging | The firm, TeDan Surgical Innovations (TSI), sent an "URGENT: Medical Device Recall" letter dated June 11, 2012 to their consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine all products subject to recall and return product to TSI. In addition, if they have further distributed this product, identify all customers and notify them of this product recall at once in writing; and complete and return the enclosed Mandatory Response Form as soon as possible and no later than June 25, 2012 via mail: 12675 W. Airport Blvd., Suite 200, Sugar Land, TX 77478; fax: (713) 726-0846 or email: thernandez@tedansurgical.com.If you have any questions, contact the Business Manager at 713-726-0886, Monday thru Friday, from 8am - 5pm or via email at THERNANDEZ@TEDANSURGICAL.COM. |  | 12065 | 12,065 Sterile Distraction Screws (2413 boxes); 384 Sterile XL Screws (128 boxes) in distribution | Worldwide distribution: USA (nationwide) and countries of : Australia and Taiwan. | Terminated | July 31, 2013 | 415 | 12065
Z-2380-2012 | 62939 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Ultraview DM3 Monitor |  |  | June 06, 2012 | 2012 | Zoe Medical Incorporated | Unit fails to power up, resulting in an equipment alarm | Software design | The firm, Zoe Medical, coordinated with the Distributor via email on June 7, 2012. The email described the product, problem and actions to be taken. Zoe will remove product the site and return the units for correction.If you have any questions, contact Manager, Global Technical Support at 425-657-7200 ext 5508. |  | 35 | 35 units | Nationwide distribution: WA only. | Terminated | September 18, 2012 | 104 | 35
Z-2394-2012 | 63105 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, Pre-Filled Flush Solution, 10 mL in 12 mL Single Use Syringe, Rx only, List no. 1078-20, made in Nashville, TN, manufactured for Hospira, Inc., Lake Forest, IL 60045 USA. |  |  | August 15, 2012 | 2012 | Amsino Medical USA | One lot of 0.9% Sodium Chloride Flush syringe, 3 mL in 12 mL may be mislabeled as 0.9% Sodium Chloride Flush syringe, 10 mL in 12 mL. | Employee error | The firm, Amusa, notified their sole consignee of the mislabeled product via e-mail and telephone on August 15, 2012. An "URGENT-CLASS II MEDICAL DEVICE RECALL" letter dated August 22, 2012 was sent to the customer. The letter identified the product, problem, and actions to be taken by the customer. A Response Form was included for the customer to complete and return. Contact the firm at 615-521-2723 for questions regarding this recall. |  |  | unknown | Distributed to one customer in Illinois. | Terminated | August 13, 2013 | 363
Z-2395-2012 | 63078 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Empty Evacuated Container 250 mL and 1000 mL. |  |  | August 28, 2012 | 2012 | Hospira Inc. | Embedded iron oxide glass defect which may have the potential to break off into the solution. | Material/Component Contamination | Hospira sent an "URGENT DEVICE RECALL" dated August 28, 2012 to all affected customers. The letter identified the product, problem,and actions to be taken by the customers. The letter included an Urgent Device Recall Reply Form for customers to complete and return via fax to 1-888-237-5582. Contact Hospira Customer Care at 1-877-946-7747 for questions regarding this recall. |  | 191094 | 191,094 units | Worldwide Distribution-USA (nationwide) and the countries of Saudi Arabia, Canada and Barbados. | Terminated | September 18, 2013 | 386 | 191094
Z-2402-2012 | 54705 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | A mobile, electro-hydraulically operated, fluoroscopic/C-Arm compatible general surgical table. |  |  | February 17, 2010 | 2010 | Steris Corporation | Customers may be storing objects on the base and/or around the column of the surgical table. This prohibited practice and misuse known by the customers can damage the override switch assembly located at the top of the column cover. Depending upon the nature and severity of the damage to the electronics of the table, it is possible that the table could move without prompting by an operator . | Employee error | Steris sent an Urgent Field Correction Notice letter dated February 17, 2010 to their customers. The letter identified the affected product, description of the problem and actions to be taken. Customers were instructed to examine their units for any structural damage. If found they are to discontinue use and arrange for a field representative to visit the facility to assess the damages and arrange for repair of the unit. For information or questions regarding STERIS visit to your facility contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 11044 | 11,044 Units | Worldwide Distribution - US Nationwide including District of Columbia and Puerto Rico and the following countries: American Samoa, Argentina, Australia, Bahrain, Bangladesh, Bermuda, Brazil, Canada, China, Colombia, Costa Rica, Dominican Republic, Ecuador, Egypt, Germany, Greece, Guam, Hong Kong, India, Indonesia, Israel, Italy, Japa, Jordan, Kazakhstan, Korea, Kuwait, Lebanon , Libya , Malaysia , Mexico, Morocco, New Zealand, Panama, Peru, Philippines , Poland, Portugal, Qatar, Russia, Saudi Arabia , Singapore, South Africa, Spain, Sri Lanka, Taiwan, Thailand, Trinidad and Tobago, Turkey &amp; United Kingdom. | Terminated | March 13, 2013 | 1120 | 11044
Z-2401-2012 | 63004 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll X Series Defibrillator/Pacemaker/Monitor |  |  | August 16, 2012 | 2012 | ZOLL Medical Corporation | Shipped with incorrect software | Process control | Zoll Medical telephoned affected customers on August 16, 2012. Customers were informed of the affected product, problem and actions to be taken. Customers were instructed to remove and return the affected product for replacement. For question contact ZOLL's Technical Support. |  | 11 | 11 units | Worldwide Distribution - US Nationwide including the states of: LA, MA, ME, NC, ND, NE, OH and the country of Germany. | Terminated | September 19, 2012 | 34 | 11
Z-2396-2012 | 63047 | [6 empty] | InfinityÂż sampling device, Cytology Brush, United States Endoscopy Group, Inc., 5976 Heisley Road, Mentor, OH 44060 |  |  | December 22, 2011 | 2011 | US Endoscopy Group Inc | US Endoscopy received five complaints which noted difficulty in deploying the cytology brush through the catheter. | Nonconforming Material/Component | US Endoscopy sent an Addendum and Correction to the Instructions for Use letter dated December 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to rinse the brush head with sterile water immediately prior to use (reference the attached Addendum. Customers were also instructed to retract and deploy the brush to ensure that is is functioning properly prior to use. Customers were asked to review the instruction for use for this product along with the Addendum to ensure proper usage and performance.For questions customers should call 1-410-639-4494 or their local Endoscopic Product Specialist.For questions regarding this recall call 440-639-4494, ext 378. |  | 393 | 393 Devices | Worldwide Distribution - USA including CA, FL, GA, IA, IL, MA, ME, MN, MO, NC, NH, OH, OK, PA, TN, TX, VA, WA &amp; WV and the countries of Australia, Canada, Columbia, Finland, Germany, India, Israel, Italy, Netherlands, New Zealand &amp; United Kingdom. | Terminated | March 22, 2013 | 456 | 393
Z-2397-2012 | 62494 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACS 7-color Setup beads |  |  | June 29, 2012 | 2012 | BD Biosciences, Systems &amp; Reagents | Incorrect version of the CD was included with BD FACS 7 Color Setup Beads. | Employee error | BD Biosciences sent a Important Product Recall Information letter beginning July 3, 2012, to all affected consignees via Fed Express. The letter identified the product the problem and the action needed to be taken by the customer.The customers were informed that the CD enclosed in their recent purchase of BD FACS 7-Color Setup Beads, lot number 33412, contain CD version 2.0 instead of the newer version 3.0. The letter goes on to explain the the potential risk and include a table of listed range values to determine if a potential false pass or false fail for the listed detectors has occurred. Customers are asked to check the CD with the kit to ensure that it is version 3.0 (the kit lot number ID is on the CD label). If it is not the correct version, discard version 2.0 and do not use it with this bead kit. Call BAD customers service for a replacement.If the 2.0 kit was already uses, letters request that clinicians inspect the FITC and APC-CY7 stains in the plots of the laboratory reports for the clinical applications in BD FACSCanto clinical software for compromised resolution between negative and dim populations.The verification form/effectiveness check should be completed and returned as requested. Questions or concerns may be directed to BD Customer Support at 1-855-236-2772. |  | 96 | 96 kits | Worldwide Distribution -- USA (nationwide) including the countries of Australia, Belgium, Canada, Chile, Korea, Mexico, Singapore and Taiwan. | Terminated | March 25, 2013 | 269 | 96
Z-2399-2012 | 62391 | No | LRY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | JUGGERKNOT INÂżGUIDE PUNCH, 2.9MM. |  |  | August 14, 2012 | 2012 | Biomet, Inc. | Investigation found that the current design of the product allows the press-fit of the knob/shaft interface to be overcome by the impaction force of a mallet, causing the knob to potentially unfasten and slide down the shaft. | Device Design | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated August 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Response Form was included for customers to complete and return via fax to 574-372-1683. Questions related to the recall should be directed to 574-371-3758, Monday through Friday, 8 a.m. to 5 p.m. |  | 62 | 62 | Nationwide Distribution-including the states of WA, IN, AK, PA, GA, SD, KS, NY, VA, AR, AL, OH, WI, MA, and CA. | Terminated | April 18, 2013 | 247 | 62
Z-2400-2012 | 63098 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Guardian II Hemostasis Valve, Model 8210 and Guardian II NC Hemostasis Valve, Model FH102. |  |  | August 30, 2012 | 2012 | Vascular Solutions, Inc. | Vascular Solutions, Inc., became aware of a potential problem involving pouches missing a seal from Guardian II (Model 8210) and Guardian II NC (Model FH102) Hemostasis Valve which poses a risk of product contamination. | Process control | The firm, Vascular Solutions, sent an "URGENT - Medical Device Field Safety Notice" dated August 24, 2012 and a follow-up "Urgent Medical Device Removal - Lot Specific" letter dated August 30, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check their inventory and to follow steps outlined in the letter which included to complete and return the "VSI Account Inventory Form" via email to: customerservice@vasc.com or fax: to (763) 656-4251.If you have specific questions or concerns regarding this matter, please contact your local Vascular Solutions Account Manager; call (763) 656 4300 or email: Senior Regulatory Operations Associate at SOuellette@vasc.com and/or email: www.vascularsolutions.com. |  | 3500 | 3500 units | Worldwide distribution: USA (nationwide) including states of: AZ, AR, CA, CT, DE, FL, GA, ID, IL IN , KY, LA,MD, MN, MO, NV, NM, NY, OH, PA, SC, TN, TX, VA, WA, and WV; and countries of: Italy and Korea.Italy, Korea | Terminated | December 04, 2012 | 96 | 3500
Z-2398-2012 | 62850 | No | LOM |  |  | Microbiology | PMA | The Bio-Rad MONOLISA Anti-HBs EIA is for In Vitro Diagnostic Use and it has the catalog number 25220. It is a qualitative and quantitative enzyme immunoassay for the detection of antibody for hepatitis B surface antigen in human serum and EDTA or citrated plasma. |  |  | July 27, 2012 | 2012 | Bio-Rad Laboratories | Prolonged exposure of the MONOLISA Anti-HBs EIA Conjugate Diluent component to light may result in elevated Optical Densities (OD) in some samples such that a negative sample could elevate to a grey zone or low positive result. | Packaging | The firm, Bio-Rad Laboratories, Inc., sent an "Important Product Information Bio-Rad MONOLISA Anti-HBs EIA Catalog #25220", letter dated 27 July 2012, to their domestic consignees/customers via FedEx. The International subsidiaries will be notified by e-mail. The consignees/customers were informed of the important storage and handling information regarding the MONOLISA Anti-HBs EIA Conjugate Diluent component . The consignees/customers were instructed to follow the storage and handling instructions as stated in the Package Insert, and compare their laboratories practices to the storage and handling instructions stated by Bio-Rad.Customers who have questions can call Bio Rad Technical Support at 1-800-224-6723, option 2, then option 3. |  | 7400 | 7,400 (approximately)- All lots since distribution date May 9, 2007 | Worldwide distribution: USA (nationwide) and countries of: Canada, South Africa, and Zimbabwe. | Terminated | December 14, 2012 | 140 | 7400
Z-2403-2012 | 62738 | No | OYT |  |  | Hematology | 510(K) Exempt | Amplilink (AL) Software 3.2.3 for use with COBAS AMPLICOR, CAP/CA, COBAS TaqMan,and COBAS AmpliPrep/COBAS TaqMan. The software can be used with a Laboratory Information System. |  |  | April 23, 2012 | 2012 | Roche Molecular Systems, Inc. | When using a Laboratory Information system (LIS) with automated systems utilizing AMPLILINK software v3.2 series (AL v3.2) there is a potential for wrong results to be assigned to an order. | Device Design | Roche Molecular Systems Inc. sent "Urgent Field Safety Notice" letters and Faxback forms on 4/23/2012 via UPS to all affected customers. The issue was described and recommended actions were provided. A Product Advisory Notice and an Updated Field Safety Notice were issued as well. |  | 1292 | 1292 instruments | Nationwide (USA) Distribution | Terminated | February 12, 2014 | 660 | 1292
Z-2405-2012 | 62917 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Precise Digital Accelerator |  |  | August 06, 2012 | 2012 | Elekta, Inc. | There have been a small number of instances in hospitals where the joint between the drive gear and the gantry base drive wheel has become loose, leading to fatigue failure of some of the securing bolts. | Nonconforming Material/Component | Elekta sent an Important Field Safety Notice dated August, 2012 to inform users that inspection and possibly bolt replacement needs to be carried out. Users should return the signed acknowledgement to Elketa, and an inspection should be carried out under FCO20002103053.For questions customers should call +44(0)1293 654200.For questions regarding this recall call 770-300-9725. |  | 110 | 110 | Nationwide Distribution including AL, AZ, CA, CO, CT, FL, GA, IL, IA, MD, MA, MI, MN, NE, NV, NJ, NM, NY, ND, OH, OK, OR, PA, VA, SC, TN TX, WA, WV, WI and Puerto | Terminated | October 08, 2014 | 793 | 110
Z-2410-2012 | 62897 | [6 empty] | ADVIA Centaur FSH Assay: ADVIA Centaur Calibrator B |  |  | August 07, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | ADVIA Centaur Prolactin Assay and FSH, Calibrator B fails the analytical sensitivity stated in the IFU for lots below 40 | Nonconforming Material/Component | Siemens Healthcare issued an Urgent Field Safety Notice letter dated August 2012 by e-mail to all affected Siemens Healthcare Diagnostics Branch offices on August 7, 2012. Customers were instructed to discontinue using the ADVIA Centaur Calibrator B lots 40 and below following their obtaining Calibrator 8 lot 41 or higher. Customers were asked to complete and fax-back the response form attached. |  | 15445 | 15,445 kits | Worldwide Distribution - US Nationwide and the countries of: Canada, Argentina, Australia, Bangaldesh,Brazil, Chile, China, Columbia, Croatia, Ecuduar, Egypt, Hong Kong, India, Israel,Japan, Korea, Maylasia, New Zealand, Mexico, Peru, Philipines, Pakistan, Saudia Arabia, Singapore, Thailand, Taiwan, Viet Nam, South Africa and UK. | Terminated | February 26, 2014 | 568 | 30900
Z-2404-2012 | 63006 | [6 empty] | The Brilliance CT 16 System |  |  | August 28, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips was notified that the system logout in software version 2.3.6 is now longer and may sometimes fail to logout, requiring a forced system restart or hard shutdown with the power button. | Software change control | Philips sent an Urgent - Medical Device Correction letter dated August 24, 2012 via certified mail to all US consignees. Philips Key Market representatives distributed the letter internationally. The letter identified the affected product, problem and actions to be taken by customer/user. Philips is implementing a software update to correct the issue free of charge. A Philips representative will contact affect sites to schedule the software installation. For further information or support contact your Philips representative or local Philips Healthcare office. |  | 93 | 93 units | Worldwide Distribution - US Nationwide including the states of: AL, CA, GA, MA, MD, NC, OH, PA, TX, and WA; and to the following countries of: Algeria, Oman, Australia, Brazil, China, Saudi Arabia, United Arab Emirates, and India. | Terminated | December 24, 2013 | 483 | 93
Z-2415-2012 | 63001 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Kimberly-Clark MIC Safety PEG Kit - Pull Type |  |  | June 22, 2010 | 2010 | Kimberly-Clark Corporation | The placement guidewire diameter is too large and does not pass-through the introducer needle during the placement procedure of the PEG Feeding Tube. | Nonconforming Material/Component | Kimberly-Clark Health Care sent an Urgent Voluntary Medical Device Withdrawal letter dated June 22, 2010, to all affected customers. The letter included a Product Withdrawal Response Sheet which was to be returned to Kimberly-Clark.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine and destroy all affected products and complete the attached Product Withdrawal Response Sheet and Fax the form to 920-225-4182.For questions customers should call 770-587-7362.For questions regarding this recall call 770-587-8393. |  | 277 | 277 units | Nationwide Distribution including AL, AK, AZ, CA, CO, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, ME, MA, MI, MN, MS, MO, NE, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, and WI. | Terminated | February 07, 2013 | 961 | 1043
Z-2423-2012 | 63146 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Asahi KASEI, REXEED(TM)- 15 LX, HEMODIALYZER |  |  | September 21, 2012 | 2012 | Asahi Kasei Medical Co., Ltd., Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products, though there is no sufficient information to date to conclusively link these adverse events to the products. | Process change control | AsahiKASEI Medical Co., LTD sent a Medical DeviceRecall letter on September 21, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to confirm that they have no inventory and to forward the letter to all clinics who may have received the product requesting they check their stock and arrange for the return of any recalled product. Non-responding consignees will be contacted by telephone.For questions regarding this recall call Asaki Kasei Medical America, Inc. at 888-362-6105. |  | 24 | 24 units | REXEED SX: Worldwide Distribution - USA including TN and the countries of Japan and GermanyREXEED LX: Worldwide Distribution - USA including TN and Germany | Terminated | February 05, 2013 | 137 | 325008
Z-2417-2012 | 63129 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | VerifyÂż Bowie-Dick Test Card, Model #S3098, Steris Biological Operations, 9325 Pinecone Drive, Mentor, OH 44060 |  |  | July 03, 2012 | 2012 | Steris Corporation | STERIS has learned that the Verify Bowie Dick Test Cards are not performing to product specifications. Specifically, test cards have resulted in false fail results when sterilizer performance is within acceptable ranges. | Nonconforming Material/Component | Steris Corporation sent a Urgent Voluntary Recall Notice dated July 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to examine their inventory and compare the lot numbers printed on the box with the listed lots. If they have products from the affected lot numbers to contact STERIS Customer Service for account credit. After notifying STERIS of the amount of recalled product on hand they are to immediately destroy the product. Distributors are to also notify their customers of the recall. They are to send a copy of the Customer Recall Notice. They are to inventory their stock and notify STERIS of any remaining recalled product for replacement product. All recalled product should be destroyed.Please contact your STERIS Account Manager or STERIS Customer Service at 1-800-548-4873. |  | 629 | 629 cards | Worldwide Distribution--USA (nationwide) including the states of AL, CA, CO, CT, FL, GA, IA, IN, IL, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK,OR, PA, RI, SC, TX, UT, VA, WA, WI, WV and WY. and the country of Canada. | Terminated | March 27, 2013 | 267 | 629
Z-2413-2012 | 63069 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bard Ventralight ST Mesh. - Ellipse 7" x 9" (17.8 cm x 22.9 cm) |  |  | September 04, 2012 | 2012 | Davol, Inc., Subs. C. R. Bard, Inc. | Mislabeled: The 8" x 10" package contains 7" x 9" VENTRALIGHT ST Mesh product and the 7" x 9" package contains 8" x 10" VENTRALIGHT ST Mesh | Mixed-up of materials/components | Davol issued an "URGENT: MEDICAL DEVICE RECALL" letter dated September 4, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter advises customers to examine their inventory and quarantine affected product for return to the firm. Customers were asked to complete and return a Recall Effectiveness Check Form. Contact Davol Customer Service Department at 1-800-556-6275 or C.R. Bard's Medical Services &amp; Support Department at 1-800-562-0027 if you have any questions. |  | 49 | 49 units | Nationwide Distribution. | Terminated | September 03, 2013 | 364 | 49
Z-2430-2012 | 62736 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Model 324JL On-Board Suction System |  |  | March 30, 2012 | 2012 | Impact Instrumentation, Inc. | A shipment mix-up took place. | Process control | The firm, Impact Instrumentation Inc., contacted their customers via phone and email notification on April 9, 2012. Customers were provided the correct packing slip and asked to acknowledge the correction.If you have any questions, contact the Regulatory Affairs Manager at 973-882-1212 or email: aglordano@impactii.com. |  | 2 | 2 devices | Nationwide distribution: USA including states of: FL and OH. | Terminated | July 21, 2015 | 1208 | 2
Z-2425-2012 | 62577 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ULTRA Fabric Reinforced Surgical Gown, X-Large, Sterile, Test File. |  |  | December 31, 2010 | 2010 | Kimberly-Clark Corporation | The impacted product was not intended for commercial distribution to end users. | Packaging process control | Kimberly-Clark sent a Urgent Medical Recall Notice dated on December 31, 2010. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to: Evaluate their Distribution Facilities Inventory of Ultra Fabric Reinforced surgical Gown to determine if any of impacted product remained within their facility. If present, to discontinue distribution of the impacted product and place in a quarantined location. Complete the attached Distributor Product Withdrawal Response Form and fax it to KC within five days. Destroy the impacted product according to their facility procedures. And to review their distribution records to identify all customers who were shipped any of the impacted products. Notify each of their customers who were shipped any of the impacted products by sending them a copy of the attached Customer Letter with the Customer Withdrawal Response Form. A minimum of three (3) attempts to contact each non-responding customer within interval of approximately four to six weeks will be performed. Once all customers who were notified have responded or after approximately four weeks after the third attempt to contact the last non-responder, a recall Termination will be initiatedIf your Kimberly-Clark Representative has not already contacted you, please contact the Ultra Fabric Reinforce surgical Gown Coordinator at (770) - 587-7700. |  | 32 | 32 cases | US (nationwide) including the states of CA, CO, LA, NM, OK, TX and WA. | Terminated | November 01, 2012 | 671 | 32
Z-2411-2012 | 63011 | No | HPT | Perimeter. | Ophthalmic | Ophthalmic | 510(K) Exempt | Octopus 101 Perimeter; PeriTrend Software Analysis |  |  | August 17, 2012 | 2012 | Haag-Streit USA Inc | The firm recalled the device after learning of a possible deviation in the stimulus luminance of the Perimeter OCTOPUS 101. | Equipment maintenance | The firm, HAAG-STREIT USA, Inc., sent a "Field Safety Notice 2012-01/01" dated August 10, 2012 via certified mail or onsite visit to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to make all users of the product and relevant personnel aware of this Field Safety Notice; forward a copy to third parties, if affected product has been passed on; keep this notice at least until the completion of the corrective measure, and complete and return the enclosed Confirmation of Receipt Form via fax to: 513-229-3867 or scanned and emailed to : fieldsafetynotice@haag-streit-usa.com.If you have any questions, call 513-336-7255. |  | 64 | 64 units | Worldwide distribution: USA (nationwide) including states of: AR, CA, CT, FL, GA, IA, IL, LA, MA, MI, MN, MS, NJ, NV, NY, OH, OR, PA, SC, SD, WA, WI, and WV; and to the countries of Jamaica, Puerto Rico, and Mexico. | Terminated | May 29, 2015 | 1015 | 64
Z-2427-2012 | 63089 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 12/14 TAPER HIGH OFFSET NECK USE WITH SMITH &amp; NEPHEW STEMS AND 12/14 FEMORAL COMPONENTS ONLY CO-CR, REF 71352112, Smith &amp; Nephew, Inc. Memphis, TN 38116 USA. |  |  | August 22, 2012 | 2012 | Smith &amp; Nephew Inc | Modular necks were manufactured out of specification by a vendor; the taper provide is undersized. | Employee error | Smith &amp; Nephew sent an "URGENT:CLASS II RECALL-HIGH OFFSET MODULAR NECK R1213 letter dated August 22, 2012 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. Customers were instructed to cease use and distribution of the product and quarantine for return. An Inventory Return Certification Form was included for customers to complete and return via fax to 901-566-7975. |  | 70 | 70 units | Worldwide Distribution-USA (nationwide) including the states of CA, FL, GA, IL, IN, MA, NC, NJ, NY, OH, OK, PA, TN, TX, and WA, and the country of Canada. | Terminated | February 27, 2013 | 189 | 70
Z-2431-2012 | 63016 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Sekisui Spectrolyse PAI-1, In-Vitro Diagnostic for Quantitative determination of Plasminogen Activator Inhibitor Type- 1 (PAI-1) activity in human plasma. |  |  | August 28, 2012 | 2012 | Sekisui Diagnostics Llc | SPECTROLYSE PAI-1 activity assay producing lower than expected absorbance values with "0" standard. This will reduce the slope of the standard curve requiring the user to make more frequent sample dilutions. | Process control | The firm, Sekisui Diagnostics, LLC, sent an "URGENT PRODUCTS RECALL NOTICE" dated August 28, 2012 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to immediately stop using the product; to discard any remaining inventory and complete and return the Customer Acknowledgement Form by September 4, 2012 documenting the # of kits destroyed via FAX to Quality Assurance at 203-602-5553. The customers were also instructed to return the form even if they do not have the affected product currently in inventory. Sekisui will replace kits.If you have any questions contact, Customer Service Department at 800-637-3375, ext 10, 12, or 31, or email to recall coordinator at MKHAN@AMDIAG.COM. |  | 699 | 699 kits | Nationwide distribution: USA including states of: FL, IN, IL, and WA. | Terminated |  |  | 699
Z-2360-2012 | 62132 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | BR980-9600 Hysteroscopy Sterile Procedure Kit containing 2 oz. bottle PVP Prep Solution |  |  | September 30, 2011 | 2011 | Custom Medical Specialties, Inc. | The firm saw on the FDA web site a recall by the manufacturer of the Povidone Iodine Prep Solution due to no microbial testing and product did not meet proper finished goods specifications. | Material/Component Contamination | Custom Medical Specialties sent an Urgent Medical Device Recall letter dated September 27, 2011 and a follow-up letter dated September 4, 2012 to all affected customers. The letter identified the affected product, problem, and actions to be taken. The letter requested that customers inspect inventory, discontinue use and return any remaining affected products. A Return Authorization Number and shipping number will be given to customer by calling 919-202-8462 ext 205. Customers were instructed to complete and return the attached Acknowledgement and Return Product form. |  | 135 | 75 cases &amp; 60 Bottles for reps | US Nationwide Distribution including the states of CO, FL, KS, MA, MI, NJ and PA. | Terminated | April 29, 2013 | 577 | 994
Z-2429-2012 | 62978 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Integra HRAIM Head Ring Assembly with Intubation Mounts |  |  | August 09, 2012 | 2012 | Integra LifeSciences Corporation | T-handle screw is used on a complementary product to the CRW, the Integra HRAIM, which is the CT-compatible intubation Head Ring Assembly. When the T-handle on the HRAIM intubation hoop is tightened, the T-handle could stop in a vertical position that prevents the CRWPRECISE and certain models of the CRW-ASL from seating properly on the head ring although it appears to be fully seated | Device Design | The firm, Integra, sent an "URGENT: VOLUNTARY MEDICAL DEVICE RECALL NOTIFICATION" dated August 10, 2012 to US &amp; Canada consignees/customers via traceable courier service and Non-US &amp; Non-Canada consignees/customers - through written notification delivered either by traceable courier, email, facsimile, or through a directed visit by an Integra Sales Representative. The notification describes the product, problem and actions to be taken. The Integra sales representative will replace the screws according to the attached rework instructions. If for some reason it is not possible to replace the screws during the visit, the sales representative will leave the screws and customers can then replace them by following the attached rework instructions and discard the old screws immediately after replacement. The customers were instructed to complete and return the Product Recall Acknowledgement Form via the sales representatives visit; fax at 1-609-750-7999 or email to: csnirecalls@integralife.com and include the number of devices to be identified and replaced . Note:The T -Handle screws that are replaced will be disposed of at the consignee's/customers location.Please feel free to contact our service hotline at 888-772-7378 should you have any additional questions. |  | 188 | 188 units | Worldwide distribution: USA (Nationwide) and countries of: Argentina, Australia, Belgium, Canada, China, Czec, Denmark, France, Hong Kong, Hungary, Italy, Japan, Malaysia, Philippines, Singapore, Spain, South Africa, and UK. | Terminated | March 26, 2014 | 594 | 363
Z-2426-2012 | 63142 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STERIS 5085 and 5085 SRT Surgical Tables, |  |  | April 06, 2012 | 2012 | Steris Corporation | STERIS has learned from Customer feedback and field service experience that damage from external impact to the table hydraulic column can cause the stainless steel shroud sections covering the table hydraulic column to become misaligned; if the damage is severe, the up/down movement of the table may be impeded. In addition, the table hand control may indicate that the table floor locks are unlock | Device Design | Steris sent an Urgent Voluntary Field Correction Notice dated April 6, 2012 to all affected customer. The letter identified the affected product, description of the problem and actions to be taken. Customers were informed that a STERIS Service Representative will contact affected facility to arrange for the replacement of the stainless steel shroud covering the hydraulic column and relocation of the floor lock pressure switches on their 5085 and/or 5085SRT table(s). For further information or if you have questions regarding the visit to your facility, contact STERIS Field Service Dispatch at 1-800-333-8828.If you have questions regarding this matter, call the Director of Marketing Tables and Case Goods at 440-392-7148 or STERIS Customer Service at 1-800-548-4873. |  | 379 | 379 tables | Worldwide Distribution - US Nationwide including the states of: AL, AR, CA, CO, CT, FL, HI, ID, IL, IN, IA, KY, LA, MA, MD, ME, MI, MN, MO, MS, NV, NYH, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WI &amp; WV and countries of: Australia, Canada, China, India, Korea, Mexico, Panama and Thailand. | Terminated | September 04, 2013 | 516 | 379
Z-2414-2012 | 62765 | No | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | S-RomÂż Total Hip System Poly Liner |  |  | August 22, 2012 | 2012 | DePuy Orthopaedics, Inc. | Orthopaedics, Inc. (DePuy) is issuing a voluntary recall of four lots of the S-ROM" liner because the Carton label has an Incorrect Color Code Series Designation. We have received one complaint regarding this issue. | Labeling Change Control | DePuy sent an "URGENT INFORMATION-MEDICAL DEVICE RECALL NOTICE" dated August 22, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Reconciliation Form was included in the letter for customers to complete and return via fax to 574-371-4939. Contact your local Depuy Orthopaedics sales representative for questions related to this recall. |  | 12 | 12 liners | Nationwide Distribution including the states of MA, TX, OH, CA, and NC. | Terminated | October 09, 2013 | 413 | 12
Z-2440-2012 | 62606 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | BP7561 OPTI CCA-TS B-Lac Cassette (25 per box) |  |  | July 03, 2012 | 2012 | OPTI Medical Systems, Inc | The values reported for PCO2 by BP7561 B-Lac cassettes are biased at high concentrations when used at altitudes above 2800 feet (853 meters). | Process design | OPTI Medical sent a Urgent Field Correction Notice dated July 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to apply correlation factors to correct the PCO2 measurement reported by the device. They were instructed to determine the elevation (altitude above sea level) for their location and to use the tables provided in the recall notification to determine the correct correlation factor. A website was provided if the consignee did not know their elevation. Steps were provided to enter the correlation factor into their OPTI CCA TS analyzer.Please contact OPTI Medical Technical Support at 1-770-510-4444 or 1-800-490-6784. |  | 895 | 895 boxes (25 cassettes per box) | Worldwide Distribution -- US (nationwide) including the states of AZ, ID, TX and UT., and the countries of Argentina, Bolivia, China, Ecuador, France, Germany Greece, India, Indonesia, Iran, Ireland, Japan, Kuwait, Lithuania, Luxembourg, Malaysia, Mexico, Morocco, Nigeria, Poland, Romania, Russia, Saudi Arabia, Singapore, South Africa, Sudan Switzerland, Turkey, United kingdom, Venezuela, and Viet Nam. | Terminated | April 11, 2013 | 282 | 895
Z-2434-2012 | 63015 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Duet TRS 45 4.8 mm Universal Straight Single Use Loading Unit. |  |  | August 21, 2012 | 2012 | Covidien LP | Duet TRS may have the potential to injure adjacent anatomical structures which may result in life threatening post-operative complications. | Device Design | Covidien issued an "URGENT MEDICAL DEVICE RECALL" letter dated August 21, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The firm sent a revised customer recall notification letter via FedEx overnight delivery to all customers who purchased Duet TRS" devices from August 2009- August 2012. This letter went to all customers whether they had been sent the August 21st letter or not. Users were advised that Covidien was recalling all lots of the Duet TRS Universal Straight and Articulating Single Use Loading Units (SULU). On August 23rd 2012, Covidien notified, via e-mail, perspective surgeons that were planning to participate in the "Duet TRS Patient Registry" of the recall and that as a result the Duet TRS registry program was discontinued. Customers were asked to complete and return a Recalled Product Form via fax to (800) 895-6140. Contact Covidien Customer Service at (800) 722-8772, option 1 for questions regarding this recall. |  | 4867 | 4,867 units | Worldwide Distribution. | Open, Classified |  |  | 1078939
Z-2442-2012 | 62927 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Sleek ***OTW ***Ref ***426-1501X***Inflated Balloon Dia/mm 1.5 Inflated Balloon Length/mm 15 Usable Catheter Length/cm 150***Lot 50026688***2013-09***Cordis***PTA Dilatation Catheter***Cordis***Johnson Johnson Company***Sterile***EO***Method of Sterilization: ETO***ClearStream***Clearstream Technologies, Ltd.***Moyne Upper, Enniscorthy Co Wexford, Ireland***Phone: +353-53-92-37111***Fax: +353-53-92-37-100***Made in Ireland***Distributed by: Cordis Cashel***Cahir Road, Cashel, Co. Tipperary, Ireland***". |  |  | January 31, 2011 | 2011 | Cordis Corporation | On 1//20/2011 Cordis Corporation , Miami Lakes, FL initiated a recall of their Sleek OTW PTA Dilatation Catheter Model #4261201X Lot #50027149 and Model #4261501X Lot #50026688. Product is incorrectly labeled. | Error in labeling | Cordis (Johnson &amp; Johnson Company) sent an Urgent Medical Device Field Corrective Action Notice dated January 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review, sign, and return the enclosed Acknowledgment Form directly to Cordis at the fax number on the form, or to their local sales representative. Pass the revised labeling notice to all who need to be aware within their facility or organization. Also, customers were asked to maintain awareness of this notice until all affected product in their facility has been consumed. In addition to their local Cordis representative, they may contact the normal HCS Customer Service number at 1-800-327-7714, option 1, option 3.For questions regarding this recall call 786-313-2365. |  | 1 | 1 unit | Worldwide Distribution - USA and Germany | Terminated | November 26, 2013 | 1030 | 6
Z-2445-2012 | 62440 | No | GQX | Varicella-zoster virus serological reagents. | Microbiology | Microbiology | 510(k) | Trinity Biotech, Captia VZV IgG. Product code 2325600: 96 Tests; product code: 2325601: 480 tests. |  |  | May 29, 2012 | 2012 | Clark Laboratories, Inc. (dba,Trinity Biotech USA) | The Positive Control for the Captia Varicella-Zoster Virus (VZV) IgG Kit has been reported as running high out of its assigned ISR range. | Nonconforming Material/Component | Trinity Biotech sent a "PRODUCT RECALL LETTER" dated May 29, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Product Recall Fax Back Form was included in the letter for customers to complete and return via fax to (716) 488-1990. Customers were instructed to discard any of the affected product. Contact the Trinity Biotech Help Desk at 800-325-3424, option 2 for questions regarding this recall. |  | 810 | 810 kits | Nationwide Distribution-including Puerto Rico and the states of AL, CA, IL, IN, MA, MD, NJ, OH, PA, TN, TX. | Open, Classified |  |  | 810
Z-2444-2012 | 63186 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SharePlan 1.1, aka t-RayAutoplan version 1.2.1 (build number 1.3.1.10). Manufacturing name RayAutoplan, aka t-RayAutoplan, commercial name (TomoTherapy) SharePlan |  |  | October 10, 2011 | 2011 | RAYSEARCH LABORATORIES AB | This notice concerns behavior of the SharePlan that may be unexpected and not clearly described in the user manual. This behavior has not caused patient mistreatment or other incidents. However, the described user actions must be adhered to for best agreement between treatment plan evaluation and the clinical outcome of the delivered treatment. The behavior appears, when importing RFA-files (.asc) | Labeling mix-ups | RaySearch Laboratories sent an Urgent Field Safety Notice Medical Device Correction letter dated October 10, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. To completely avoid effects of the erroneous behaviour, RFA-file has to be modified to have the crossline profiles sgtored in the X column of the file and the inline profiles in the Y column. Upon request, RaySearch Laboratories will provide help to the customers by transferring the RFA-files to the correct format through a scrip. Customers were instructed to pass on the notice to all those who need to be aware within their organization and to maintain awareness on this notice until the new labeling version has been received to ensure effectiveness of the corrective action.For questions regarding this recall call +46 8 54 50 61 42. |  | 20 | 20 units | Worldwide Distribution - USA including CA, KY, MI, NY, OK, OR, SD, and WI. Internationalyl to Canada, Sweden, France, Germany, Australia, Poland, Belgium, and United Kingdom. | Terminated | January 25, 2013 | 473 | 20
Z-2443-2012 | 62313 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 ASSEMBLY FLOW POD |  |  | September 10, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (CVS) received two reports of the false back flow alarms for the TerumoÂż Advanced Perfusion System 1 when there was no tubing in the flow sensor. Terumo CVS' investigation found that the voltage level on a circuit board was improperly set in a population of flow modules for Terumo System 1, the sensor operates correctly if the tubing is installed and it is highly unli | Release of Material/Component prior to receiving test results | Terumo CVS issued an URGENT MEDICAL DEVICE Recall letter dated September 10, 2012 to alert all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that a Terumo CVS representative will contact all affected US account.to schedule an appointment to replace the affected product. The affected product will be return to Terumo for destruction. Non-US Terumo Affiliates will be instructed to remove the devices from the field and provide a record of destruction. For questions call the Terumo CVS Customer Service at 1-800-521-2818 |  | 150 | 150 units | Worldwide and US Nationwide Distribution | Terminated | January 13, 2014 | 490 | 150
Z-0573-2 | 3529 | [6 empty] | Brand name: Olympus Bronchoscopes |  |  | November 30, 2001 | 2001 | Olympus America Inc. | Loosened bronchoscope biopsy channel port &amp; microbial contamination of the port. |  | Please be aware that this is not a new recall. The firm took action in 2001; but, due to changes in FDA's recall reporting system this recall is now being posted by the Agency. |  | 5135 | 5135 units (revised final #): 4663 units (revised final #) in USA, 472 units to Canada &amp; South America. | Medical facilities nationwide &amp; Puerto Rico. The devices were also distributed to U.S. Government/military accounts and 53 foreign accounts. Distribution was provided to CDRH &amp; OE on 3/4/02.Foreign distribution: Canada, Mexico, Dominican Republic, Argentina, Brazil, Panama, Chile, Peru, Ecuador, Colombia, Venezuela, Paraguay, Uruguay, Costa Rica, El Salvador. | Completed |  |  | 5135
Z-2451-2012 | 63044 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray Systems, (Brilliance 6, Brilliance 16, MX8000 Dual v. Exp, GEMINI Dual), Philips Healthcare Systems, Cleveland, OH. |  |  | August 21, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | A bolt which attaches the oil accumulator within the gantry may fail, resulting in the oil accumulator detaching from the tube housing assembly within the gantry. | Nonconforming Material/Component | Philips Healthcare sent an Urgent Medical Device Correction letter dated August 22, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that no action needed to be taken at that time b the customer or user. A Philips Field Service Engineer would schule a time to inspect and repair all the affected systems within approximately 6 months. For further information customers were instructed to contact their local Philips representative or contact the Philips Healthcare Customer Care Solution Center at 1-800-722-9377 and follow the recorded menu options toreach a Customer Solutions Engineer.For questions regarding this recall call 440-483-7000. |  | 112 | 112 Units | Worldwide Distribution - USA including AL, AZ, CA, CO, FL, GA, IL, KS, KY, MI, MS, NJ, NM, NV, NY, OH, TX, WA &amp; WI and internationally to: Australia, Belgium, Brazil, Canada, China, Denmark, France, Germany, India, Italy, Japan, Jordan, Mexico, Saudi Arabia, South Korea, Spain, Thailand, Turkey &amp; United Kingdom. | Terminated | December 24, 2013 | 490 | 112
Z-2449-2012 | 62322 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Flow Module |  |  | September 20, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (CVS) received one report of a flow module for the Terumo Advanced Perfusion System 1 not displaying blood flow information on the Central Control Monitor or the Centrifugal Control Unit. The tricolor LED on the flow module was not illuminated. Terumo CVS' investigation found that the malfunction was caused by an improper component on a circuit board used in a populat | Process control | TERUMO sent an Urgent Medical Device Recall letter dated September 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the Medical Device Recall notice; assure that all users are aware of the notice; and confirm receipt of the communication by faxing the attached Customerr Resonse Form to 1-734-741-6149. Terumo CVS will send a representative to each affected US account to replace the affected devices and return the affected devices to Terumo for destruction. Non-US Terumo Affiliates will be instructed to remove the devices from the field and provide a record of destruction.For question customers should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 72 | 72 units | Worldwide Distribution - USA (nationwide) and the countries of AUSTRALIA, Singapore, Jordan, Japan, Vietnam, Germany, Russia, Saudi Arabia, BELGIUM, Mexico, and CANADA. | Terminated | January 13, 2014 | 480 | 72
Z-2453-2012 | 63180 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Hologic Fluoroscan Mini C-Arm Model with InSight- FD Software Version 5.0.1and 5.0.2 |  |  | September 14, 2012 | 2012 | Hologic, Inc. | Fluoroscan Mini C-Arm models -InSight2, version 4.0 and 4.0.1 - DICOM images stored in PACS may be flipped 180Âżaround vertical axis.InSight FD, version 5.0.1 and 5.0.2 - DICOM images stored in PACS may be lipped 180Âż around vertical axis as well as 180Âż around a horizontal axis. | Software design | Hologic issued an Urgent: Medical Device recall letter dated September 14, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Halogen corrected the anomaly and they provided the customer with the revised version of the software. Customers were asked to update their systems in accordance with the instructions provided with the software included with the letter. Customers were also asked to complete the attached acknowledgement form and fax or mail it to the firm at their earliest convenience. For any questions customers were instructed to call Hologic's Help Desk at 1-800-321-4659.For questions regarding this recall call 781-999-7300. |  | 46 | 46 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia,Belgium, Canada, China, Germany, Ecuador, Spain, France, Great Britain, Greece, Hong Kong, Italy, Japan Korea, Lithuvania, Neherlands, qatar, Russia, Saudi Arabia, Sweden, Singapore, Slovenia, Thailand, Taiwan, South Africa, and Switzerland. | Terminated | January 10, 2014 | 483 | 460
Z-2456-2012 | 63009 | No | GCT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | ***REF LS7700***Qty 1*** EC REP*** ConMed Linvatec Xenon Universal Light Source |  |  | May 20, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | Linvatec Corporation dba ConMed Linvatec, Largo, FL initiated a recall of their products LS7700 Xenon Universal Light Source and LS7600 Xenon Light Source are intended to be used with an endoscope to provide illumination during endoscopic procedures. | Nonconforming Material/Component | Linvatec Corporation sent an "URGENT Medical Device Recall Notification letter dated May 20, 2011 to its customers via FedEx service. Firm advised their suppliers of the Light Sources potentially has unapproved components and does not meet the finished product requirement. The distributors and end user level are advised if any affected product is in their possession can be used until returned for advance exchange. The customers were instructed to: 1) locate the identified product by serial number at your facility and complete the Reply Form attached; 2) Contact ConMed Linvatec Customer Service (800-237-0169) to receive SR Number Advance Exchange affected product only. Once Customer Service authorizes the SR Number your Advance Exchange Light Source will be shipped to your facility; 3) Once replacement unit is received, utilize the packaging material to return the recalled Light Source to ConMed Linvatec using the enclosed FedEx Airway Bill. Mail or FAX copy of Reply Form to ConMed Linvatec; 4) Complete and return a copy of the Reply Form. For questions contact your local ConMed Linvatec office or email Custseryl@linvatec.com, Customer Service at 800-237-0169 or fax to 727-319-5701. |  | 57 | 57 | Worldwide Distribution - US Nationwide and the countries of Canada, DK, HK, IN, IT, KR, MY, RU, CO, VN, ZA. | Terminated | September 26, 2012 | 495 | 58
Z-2446-2012 | 63100 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Embolectomy Catheters (Models A44XX, A4FXX, CE0XXXST, and CE0XXX): |  |  | August 09, 2012 | 2012 | Applied Medical Resources Corp | Applied Medical is recalling specific lot numbers of its Vascular Catheters because they may have a potential for packaging particulate matter to reside on the product. | Material/Component Contamination | Applied Medical sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to complete the attached Recall Notification Confirmation Form and to return it to the firm via e-mail to recall60476587@appliedmedical.com or fax to (949) 713-8832. For product return questions, customers were instructed to contact Sales Operations Administrator at (949) 713-8652.For questions regarding this recall contact the firm at (949) 713-8041. |  | 40238 | 40,238 units for all products in Recall Event | Worldwide Distribution-USA (nationwide) and the countries of United Arab Emirates, Austria, Australia, Belgium, Canada, Germany, Ecuador, Estonia, Spain, France, United Kingdom, Hong Kong, Ireland, Israel, Italy, Republic of Korea, Libyan Arab Jamahiriya, Malaysia, Netherlands, New Zealand, Pakistan, Poland, Portugal, Qatar, Serbia, Saudi Arabia, Taiwan, and South Africa. | Terminated | March 22, 2013 | 225 | 40238
Z-2450-2012 | 63092 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Armada 35/ Armada 35 LL PTA Catheters. The device is intended for dilatation of lesions in the vasculature. |  |  | August 20, 2012 | 2012 | Abbott Vascular | Abbott Vascular is recalling the Armada 35 and Armada 35 LL PTA Catheters because they have discovered that some devices may exhibit difficulty inflating and/or deflating. | Process design | An "Urgent Device Recall" letter dated 8/20/2012 was sent to all customers who purchased the Armada 35 and Armada 35 LL PTA Catheters. The letter informed the customers of the problem identified and the action to be taken. A list of lot numbers and an effectiveness check form was sent to customers with the customer notification letter. Customers were instructed to call Abbott Vascular Customer Service at (800) 227-9902. |  | 61799 | 61,799 units | Nationwide(US) Distribution and Puerto Rico. | Terminated | July 15, 2013 | 329 | 61799
Z-2454-2012 | 63156 | No | NBE |  |  | General &amp; Plastic Surgery | PMA | Progel Pleural Air Leak Sealant, Model #: PGPS002 |  |  | August 13, 2012 | 2012 | Neomend Inc | Product was distributed with an incorrect expiration date. Product was labeled with anexpiration date of 2013-04 and should have been labeled with 2012-09. | Incorrect or no expiration date | Neomend sent a Product Recall Notice letter dated August 13, 2012 to all affected customers. The notification letter identified the product affected, problem and actions to be taken. Customers were instructed to check their stock supply and ask the Central Supply Department to determine whether they have any unused product from the affected lot. If you do have unused units of this product follow the instructions provided. If you do not have any product with the affected lot number complete and return the enclosedTrackingVerification Form. For questions contact your Neomend sales representative. |  | 95 | 95 units | US Nationwide Distribution | Terminated | September 25, 2012 | 43 | 95
Z-2460-2012 | 63091 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Alere Triage TOX Drug Screen 9 Panel, P/N: 94402. |  |  | July 12, 2012 | 2012 | Alere San Diego, Inc. | Alere San Diego is recalling Alere Triage TOX Drug Screen because there is a possibility that after installing the code chip provided with the affected lots: AMP, BAR, and PCP, the threshold ranges may not be consistently set correctly in all meters. | Nonconforming Material/Component | Alere sent a 'NOTICE OF CORRECTION" letter dated July 12, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers with questions were instructed to contact Technical Services at (877) 308-8287. Customers were instructed to complete the Customer/Distributor Verification Form Notice of Correction even if they do not have any involved product and fax back to Technical Services at (858) 805-8457 or scan and email to responses.ts@alere.com. |  | 591 | 591 kits | Nationwide Distribution. | Terminated | February 01, 2013 | 204 | 591
Z-2464-2012 | 63197 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens Magnetom Skyra, Verio, Spectra and Biograph MMR |  |  | September 12, 2012 | 2012 | Siemens Medical Solutions USA, Inc | When positioning the Flex Large 4 Coil (part No. 8625761) off-center in left-right direction as it is used for example for hip and shoulder imaging and, in addition at an off-center position with respect to the head-foot direction, the coil can heat up in the area of the electronic housing (white plastic lids). | Device Design | Siemens Healthcare sent a Customer Safety Advisory Notice dated September 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included updated instruction to all customers.For questions regarding this recall call 610-219-6300. |  | 370 | 370 | Nationwide Distribution - USA including Washington, DC and AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI and Puerto Rico. | Terminated | January 14, 2014 | 489 | 370
Z-2458-2012 | 62711 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | The MAYFIELDÂż Horseshoe Gel Pads are used in conjunction with the MAYFIELDÂż Base Units and MAYFIELDÂż Swivel Adaptors for neurosurgeries in the prone or supine positions. Those devices provide cranial stability and immobility for various craniotomies and spinal procedures. |  |  | July 27, 2012 | 2012 | Integra Limited | An investigation of customer complaints identified that for two product lots, Left Horseshoe Gel Pads were packaged as a Right Horseshoe Gel Pads and vice versa. None of the complaints resulted in a report of patient injuries or an increase in the time to complete the surgical procedure. A physician was consulted to assess patient risk and concluded that there is no potential harm to patients, onl | Packaging process control | Integra Lifesciences sent an "URGENTURGENT: VOLUNTARY MEDICAL DEVICE RECALL NOTIFICATION" letter dated July 27, 2012 to all affected customers. The letter identifies the affected products, problem and actions to be taken by the customers. Customers were instructed to check their inventory for affected products and quarantine them for shipment back to Integra. Customers were asked to complete the Recall Acknowledgement and Return Forms ASAP whether or not they intended on returning any recalled product. Contact Integra Customer Service at 1-877-444-1114, Option 3 for Technical Support. |  | 94002 | 18 from Lot 093984 for Right Gel Pads | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DC, GA, IA, IL, IN, NC, and NY and the countries of: Canada and Belgium. | Terminated | February 20, 2013 | 208 | 1033003
Z-2462-2012 | 63093 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Eclipse Treatment Planning System, Model Number: H48; Eclipse 8.9 and 10.0 |  |  | August 17, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the Eclipse Treatment Planning where an outdated assigned Hounsfield Unit [HU] value can be used for dose calculation when the clock on the Eclipse Client Workstation is not synchronized with the Database/System Server. Treatment of the patient using these values can therefore lead to under or over-dose. | Software design | VARIAN Medical Systems sent an Urgent Medical Device Correction Urgent Field Safety Notice on August 17, 2012, to all affected customers with a description of the problem and user corrective action steps. Customers were instructed to ensure that the workstation clock be synchronized with the DatabaseSystem server clock. Customers were also instructed to contact Varian service for assistance with time synchronization if required. Customers were to complete the attached Proof of Notification Form once they have read the document and return it to Varian Medical Systems. For questions customers should call the following numbers:USA and Canada 1-888-827-4265.Europe +41 41 749 8844.For questions regarding this recall call 650-424-5731. |  | 6087 | 6087 | Worldwide Distribution - USA (nationwide) and Internationally... | Terminated | August 06, 2014 | 719 | 6087
Z-2459-2012 | 63141 | No | DRA | Steerable catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Z Flex 270, Steerable Sheath, Sterile EO, Rx only, Distributed by Boston Scientific Corporation, One Boston Scientific Place, Natick, MA 01760-1537, Manufactured by Greathbatch Medical, 2300 Berkshire Lane North, Minneapolis, MN 55441. |  |  | September 06, 2012 | 2012 | Greatbatch Medical | During the investigation of four returned Z Flex 270 Steerable Sheaths that were used with cardiac cryoablation balloon catheters, the internal PTFE sheath liners were found to display varying degrees of damage that could pose an embolic risk to a patient. | Nonconforming Material/Component | Greatbatch Medical sent a "SAFETY ALERT-Z-FLEX-270 STEERABLE SHEATH" letter dated September 6, 2012 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. The letter advised the customer to communicate with their customers to assure the that the recalled sheath model is not used in conjunction with cryoablation catheters. Contact the firm at 763-951-8214 for questions regarding this issue. |  | 787 | 787 | Distributed in MA. | Terminated | January 22, 2015 | 868 | 787
Z-2463-2012 | 62830 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DEXIS Imaging Suite - catalog number 1.008.7932, version numbers 10.0.0 through 10.0.3 |  |  | July 30, 2012 | 2012 | Imaging Sciences International, LLC | A condition can occur during renumbering of the patient database using the DEXCopy feature. When renumbering is being performed using the DEXCopy feature and multiple patients with an identical name are present, the software will assign all subsequent records to the first patient record, resulting in mixed patient records. In order for this problem to occur several conditions need to be met: 1) | Software design | DEXIS sent a Medical Device Corrective Action letter dated July 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. An Acknowledgement Form was included with the notification that is to be returned by the consignee to the recalling firm. Customers were instructed to call DEXIS Technical Support at 888-883-3947 as soon as possible to have Technical Support review and potentially correct any issues that may have resulted from its use. To request a free-of-charge upgrade kit, customers should email a request to orders@dexis.com.For questions regarding this recall call 215-997-5666, |  | 366 | 366 | Worldwide Distribution - USA including Washington, DC, AL, AZ, AR, CA, CO, CT, DE, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and the country of Canada. | Completed |  |  | 366
Z-2461-2012 | 62965 | chair and table, medical | KMN | Medical chair and table. | General Hospital | General Hospital | 510(K) Exempt | InvacareÂż I-Fit" Shower Chair (Models 9780, 9781, 9781-1), Invacare Corporation, Elyria, OH 44035. |  |  | August 06, 2012 | 2012 | Invacare Corporation | On August 6, 2012, Invacare voluntarily recalled the InvacareÂż I-Fit" Shower Chair (Models 9780, 9781, 9781-1) based on the return, quality and adverse event data received. | Device Design | Invacare sent an "URGENT: RECALL INFORMATION" letter dated August 31, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use of any affected product and to call the Recall Support Line at 877-496-5041 for assistance. |  | 21168 | 21,168 chairs | Worldwide Distribution-USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of AUSTRALIA, CANADA, GUAM, ISRAEL, NEW ZEALAND, THAILAND, TRINIDAD AND TOBAGO. | Terminated | February 11, 2014 | 554 | 21168
Z-2468-2012 | 63166 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301A/2 Aquilion One CT System. |  |  | May 13, 2008 | 2008 | Toshiba American Medical Systems Inc | The SureStart function may not operate in the usual manner, and you may obtainabnormal analysis results when using the cerebral blood flow analysis function. | Software design | Toshiba American Medical Systems, Inc. sent an "URGENT: MEDICAL DEVICE CORRECTION" letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.customers: The firm advised customers to take the following temporary measures to avoid issues until a corrective action has been taken:I. In SureStart:" Perform scanning with Observation Direction set to VFF when the PatientInsertion Direction and Observation direction are FF." Perform scanning with the Observation direction set to VFH when the PatientInsertion Direction and Observation Direction are HF.2. When acquiring images for 4D perfusion analysis:" Perform scanning with the Scan Direction set to OUT when the Patient Insertion Direction and Scan Direction are HF.Perform scanning with the scan direction set to IN when the Patient InsertionDirection and Scan Direction are FF.Contact the firm at (800) 421-1968 for questions regarding this issue. |  | 3 | 3 units | Nationwide distribution: USA including states of: MA, MD, and NV. | Terminated | September 27, 2012 | 1598 | 3
Z-2472-2012 | 63188 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Marathon Flow Directed Micro Catheter, Model #105-5055. |  |  | September 23, 2011 | 2011 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | The recall was initiated because ev3 Neurovascular has identified one lot of UltraFlow Flow Directed Micro Catheters and Marathon Flow Directed Micro Catheters may have been mislabeled. | Error in labeling | Micro Therapeutics Inc, Dba Ev3 Neurovascular sent a recall notification letter dated September 23, 2011 to all their affected customers. The letter identified the affected product, problem, and actions to be taken. The notification letter also provides the customer (physician) with appropriate information regarding the affected products. The letter instructs customer to stop using the affected product, segregate from inventory and return for replacement with conforming product. The notice requests the customer to complete and return the reply form indicating use or return of devices in the affected lot(s). Customers with questions regarding the recall letter were instructed to contact Kim Salceda at (949) 680-1305. |  | 23 | 23 units total | Worldwide Distribution - US Nationwide including the states: MN, FL, OR, TX and the countries of: Brazil, Belgium, Czech Republic, Finlind, France, Germany, Netherlands, Poland, and United Kingdom. | Terminated | September 28, 2012 | 371 | 59
Z-2467-2012 | 62948 | No | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | KY LIQUIBEADS Vaginal Moisturizer; Distributed By: McNeil-PPC, Inc. Skillman, NJ 08558-9418; UPC 380041087704; 380040087958. |  |  | July 02, 2012 | 2012 | Mc-NEIL-PPC, Inc. | Johnson &amp; Johnson Healthcare Products (JJHP) is voluntarily recalling this product to address FDA concerns that the product should have been cleared under a separate medical device application. This voluntary recall is not being conducted on the basis of any consumer safety or product performance issues related to K-Y LIQUIBEADS. JJHP received a Warning Letter on May 23, 2012 indicating that thi | No Marketing Application | Johnson &amp; Johnson sent Recall Letters/Response forms dated June 29, 2012 on July 2, 2012. The recall notification included affected lots and instructions for wholesale distributors and retailers that handle the product. The recalled product is being handled by Stericycle. The notification included phone and fax numbers in the event assistance is required. Customers were asked to complete the enclosed business Reply Form and return immediately; identify all warehouse/wholesale inventories of products and return immediately per the attached "Return Authorization Form." For any questions regarding the form, customers were asked to call Stericycle at 888-406-9309. Customers could also fax questions to 877-877-0318. Once customers have collected all returns, they should send back immediately per the enclosed instructions.For questions regarding this recall call 908-874-1000. |  | 818712 | 818,712 | Nationwide Distribution | Terminated | September 27, 2013 | 452 | 818712
Z-2470-2012 | 62419 | No | NGY | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | This voluntary medical device correction is related to Alcon products that present a 23G nonvalved trocar cannula and infusion cannula together. These items are available as components within Accurus 23G TotalPLUS Pak, Accurus 23G Combined Procedure Packs (0.9mm and 1.1mm), Constellation 23G Total PLUS Pak, Constellation 23G Combined Pack (0.9 mm and 1.1 mm) and 23G Small Parts Kit. |  |  | May 29, 2012 | 2012 | Alcon Research, LTD. | This medical device correction is related to Alcon products that present a 23G non-valved trocar cannula and infusion cannula together because of an issue regarding the loose fit between the 23G non-valved trocar cannula and infusion cannula that could potentially cause a detachment between the non-valved trocar cannula and infusion cannula during surgery. | Device Design | Alcon sent an Urgent a Voluntary Medical Device Correction letter on May 9, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed NOT to use the affected product, instead customers should use a 1-Count Valved Trocar Cannula and to complete and immediately return the Medical Device Correction Response Form to Alcon via fax at: (817) 916-9087 or email to: ReplacementCannulas@alconlabs.com. For questions customers should call 1-800-862-5266. |  | 93319 | 93,319 (30, 413 distributed domestically &amp; 62,906 internationally) | Worldwide Distribution - USA (nationwide) and the countries of Saudi Arabia, United Kingdom, Greece, Singapore, Germany, Switzerland, Kuwait, United Arabian Emirates, and Austria. | Terminated | February 12, 2013 | 259 | 93319
Z-2465-2012 | 63119 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost-Mobile Detector Holder for Digital Diagnost (Single Detector, Dual Detector, Release 3) |  |  | September 06, 2012 | 2012 | Philips Healthcare Inc. | The Wireless Portable Detector (WPD) is not fixed securely inside the Mobile Detector Holder and may falldown | Device Design | Philips Healthcare sent a Urgent Field Safety Notice dated August 13, 2012, to all affected customers. Notifications began the week of September 6, 2012, for US and Worldwide Customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips field service engineer will Secure the 4 countersunk screws, which are used to attach the lock mechanism, with thread locker (loctite). Further information or support concerning this issue, contact your local Philips representative. 1-800-722-9377. |  | 93 | 93 Units | Worldwide Distribution--USA (nationwide) and the countries of Australia, Austria, Canada, Chile, Finland, France, Germany, Hungary. Italy, Japan, Latvia, Netherlands, New Zealand. Oman, South Korea, Switzerland and United Kingdom. | Open, Classified |  |  | 93
Z-2466-2012 | 62979 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | BariAir Therapy System |  |  | September 17, 2012 | 2012 | KCI USA, Inc. | KCI has issued a medical device correction for the BariAir Therapy System Hand Pendent. The BariAir Therapy System is not designed to allow the patient to place the bed into the patient exit position. KCI has received complaints indicating that patients used the hand pendent to inadvertently place the BariAir Therapy System into a pseudo-exit position. In a small number of instances, this allow | Device Design | KCI sent an Urgent - Voluntary Medical Device Correction letter dated September 13, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers would be contacted by KCI to coordinate a shipment of a replacement product. For questions customers would instructed to call 1-800-275-4524, Select Option 3 followed by Option 2.For questions regarding this recall call 1-800-275-4524. |  | 575 | 575 | Worldwide Distribution - USA (nationwide) and the countries of Saudi Arabia, United Kingdom, Singapore, Germany, Switzerland, Kuwait, United Arabian Emirates, and Austria. | Terminated | July 29, 2013 | 315 | 575
Z-2469-2012 | 63107 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Catheter Irrigation Set labeled in part: ***24K Arthroscopy Outflow/Suction Tubing Set*** |  |  | July 01, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec recalled their 24k Arthroscopy Outflow/Suction Tubing Set (Product Number 24k100)because there is a possibility that the product may have a breach in the Tyvek seal that could potentially compromise the sterility of the contents. | Employee error | ConMed Linvatec sent a Urgent Medical Device Recall Notification via FedEx to all affected customers. This notice should be received July 5, 2011. The letter identified the product the problem and the action needed to be taken by the customer. All returned product under this removal action will be identified, segregated, and quarantined per ConMed Utica, N.Y. procedures. Returned product will be dispositioned in accordance with ConMed Utica, N.Y. procedures. This can include repackaged and resterilization with new lot number or product scrap. ConMed Utica, N.Y. will maintain control of all individual lot numbers returned and dispositioned.Firm is striving for 1005 effectiveness. If a Reply form is not received within approximately 30 days of the first mailing, ConMed Linvatec will make two additional contacts by mail approximately 30-35 days apart.Should you have any additional questions please contact our Customer Service at 800-237-0169 or fax at 727-319-5701. |  | 30064 | 30064 units | Worldwide Distribution--USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, WA ,WI and Hawaii., and the countries of Australia Canada, China, Switzerland, Germany, Czech Republic, Denmark, Egypt, Spain, France, South Africa, United Kingdom, Israel, Italy, Netherlands, New Zealand, Poland, Portugal, Russia and Saudi Arabia. | Terminated | July 25, 2013 | 755 | 30064
Z-2072-2012 | 62621 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | EOS X-ray Medical Diagnostic , X-ray Equipment System |  |  | June 19, 2012 | 2012 | Eos Imaging Inc | It was discovered of several generator failures of the EOS System due to failure of its anode controller board that drives the rotation of the X-ray tube anode. | Radiation Control for Health and Safety Act | EOS Imaging notified users by letter on 6/19/12 and that an on site correction would be performed by an EOS Imaging Field Service engineer.EOS indicated in its notification that although this specific defect has not been observed in EOS systems installed in the U.S, they will replace the anode driver boards with new boards equipped with high current capacitors. Similarly, the generator firmware will be upgraded with the software patch that improves the cooling fans' controls. This will be replaced at no charge to the customers, during regularly scheduled preventative maintenance of the user's EOS system. This should be completed by the end of calendar year 2012. Your proposed CAP appears to adequately address the issues you raised in your notification, and is hereby approved.Your proposed customer notification letter as revised, appears to adequately provide instructions with respect to the use of the product pending the correction of the defect and it appears to provide a clear evaluation in nontechnical terms of the potential hazards related to the defect. Your proposed notification letter is hereby approved for distribution to all affected parties. Further questions please call 678-564-5400. |  | 12 | 12 EOS Systems installed in the US | Nationwide distribution | Terminated | September 16, 2013 | 454 | 12
Z-2494-2012 | 63165 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost with patient carriage |  |  | September 12, 2012 | 2012 | Philips Healthcare Inc. | Hook does not securely hold the footplate in vertical position, causing the footplate to disengage and may fall on patient or operator | Device Design | Philips Healthcare sent an Urgent Field Safety Notice letter dated August 15, 2012 to all affected customers. The letter identified the affected product, problem and the actions to be taken to avoid or minimize the occurrence of the issue. Customers were advised to follow the instruction for Use and that Philips plans to replace the old hook with one that has a new design. For questions contact your local Philips representative 1-800-722-9377. |  | 536 | 536 units | Worldwide Distribution - US Nationwide and the countries of :Australia, Austria, Bahrain, Belgium, Canada, Chile, China, Czech Republic, Denmark, Finland, , France, Germany, Ireland, Italy, Japan, Netherlands, New Zealand, Norway, Poland, Qatar, Romania, Russia, Saudi Arabia, Sweden, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan. Thailand, Turkey,United Arab Emirates and United Kingdom. | Terminated | July 26, 2013 | 317 | 536
Z-2493-2012 | 63125 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Cemented Tibial Drill |  |  | September 07, 2012 | 2012 | Zimmer, Inc. | Zimmer is initiating a recall of the Cemented Tibial Drill for the Zimmer Persona Personalized Knee System due tothree reports being received as of August 29. 2012 for tibial cortex perforation during preparation of the tibial bone.The investigation into the reports that have been received found that the perforations are occurring while using theCemented Tibial Drill. The perforations have o | Device Design | Zimmer sent an Urgent Medical Device Recall dated August 2012 to distributors via electronic mail and Hospital risk managers and surgeons had notifications hand delivered by sales representatives. The recall notice describes the reason for the recall, emphasizes surgical technique to be used, includes instructions for delivering notifications to surgeons and risk managers, instructions for replacing the affected product, and instructions for returning the affected product. Customers were instructed to locate the affected product and contact their sales representative for replacement. Consignees are asked to fill out the Certificate of Delivery, scan it and send back per the instruction provided. |  | 60 | 60 | Worldwide Distribution - US Nationwide including the states of: AZ, CA, CO,FL, MN, IL, NJ, NY, OH, OR and the countries of Austrailia, Utd. Arab Emir., France, Germany, India, and Italy. | Terminated | April 17, 2014 | 587 | 60
Z-2492-2012 | 61796 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Tabs Professional Voice + Monitor, with wire bracket, Model 25223. Stanley Security Solutions, Inc., Lincoln, NE. |  |  | May 04, 2012 | 2012 | Stanley Security Solutions, Inc. | The monitor may fail to sound an alarm | Device Design | Stanley Security Solutions, Inc. sent an Important Recall Notice dated May 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease using any of the affected product. The customers were informed, however, that they may continue to use the units as long as they were powered with a 9-volt alkaline battery in accordance with current instructions because the issue only occurs when using the 12V DC adaptor. The letter also informs the consignee to expect a follow-up notice no later than 6/1/12 that will provide additional instructions.The recalling firm issued additional letters dated 6/5/12 to distributors and end users with updated instructions. The letter informed them further testing revealed the possibility that a subset of affected monitors may also fail to alarm when powered by a 9V battery if weight is rapidly applied and removed from the pressure pad. Therefore, the recalling firm decided to recall and repair this subset of affected monitors and provided the serial number ranges that must be returned for repair. The letter also provided a list of serial number ranges that could be corrected in the field by permanently blocking access to the 12V DC jack on the monitor. Instructions were enclosed to show the customer how to disable to the jack. In addition the distributor letters request the customer to provide the recall information to their downstream customers. A response form was included to report the amount of affected product in the customers' inventory.For questions customers were instructed to call 888-224-2460.For questions regarding this recall call 407-271-2097. |  | 12111 | 12,111 units | Worldwide Distribution - USA (nationwide) and the countries of Canada and Great Britain. | Terminated | September 24, 2013 | 508 | 23938
Z-2486-2012 | 63073 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Optilock T15 AO Driver (3.5mm). |  |  | August 23, 2012 | 2012 | Biomet, Inc. | Field complaints indicate that recent shipments of the 14-400795 and 36505 drivers do not properly mate("lock") into some of the specified screw heads. Drivers near MMC (Maximum Material Condition) or that are misshapen may not properly mate with screws at MMC. Design dimensions and tolerances of the drivers are not properly specified to prevent an incorrect fit between the drivers and the screw | Mixed-up of materials/components | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE'' dated August 23, 2012 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. A Fax Back Response Form was included for customers to complete and return via fax to 574-372-1683. Contact the firm at 574-372-1570 for questions related to this recall. |  | 363 | 363 units | Worldwide Distribution-USA (nationwide) including the states of AL, GA, MT, TA, AZ, WA, CA, TA, IN, NY, NJ, VA, MA, MI, MD, PA, KY, NC, UT, NH, FL, OK, CO, LA, OH, MO and SD, and the countries of Japan, Netherlands, Australia, and Costa Rica. | Terminated | August 27, 2013 | 369 | 363
Z-2487-2012 | 62594 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Restoration Modular Cone Body Trials |  |  | April 02, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware of a burr, or a metal shaving, observed in the internal threaded portion of the Cone Body Trial. The typical size of the shaving observed was 1-2 mm in length and 0.037 mm squared in cross section. | Device Design | Stryker Orthopaedics sent a Urgent Product Field Action - Product Inspection Branch Acknowledgement Form on April 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.All inspections must be completed immediately as per the provided Product Correction Bulletin. Please document the results of your inspection on the Tracking Log, indicating conforming devices as "Not Affected" and non-conforming devices as "To be returned." If the devices listed on your tracking log have previously been returned or are otherwise not on hand, indicate the appropriate disposition following our usual process.Please assist us in meeting our regulatory obligation by faxing back the attached Product Acknowledgment Form and Tracking Log within 10 business days to 855-251-3635. Any further questions please call 201-831-5523. |  | 10983 | 10, 983 units | Nationwide distribution | Terminated | June 13, 2013 | 437 | 10983
Z-2480-2012 | 62110 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | 00780302601 MIS 12/14 Slotted Femoral Head Provisional -3.5X26 |  |  | April 01, 2011 | 2011 | Zimmer, Inc. | Complaints have been received reporting sub-optimal fit between provisional femoral heads (VerSys and MIS) and the mating components (ie stems and provisional femoral necks), in some cases alleging the fit was too loose and in other cases, alleging the fit was too tight. Intraoperative dislodging of the provisional head and the provisional head being left in-vivo after separating from mating comp | Device Design | Zimmer, Inc. sent an E-mail letter dated April 21, 2011, to all affected customers in the US. A June 2011 E-mail was sent to all customers outside the US. The letters requested Inventory &amp; Operations Managers to check product for proper fit /function, complete the return form with worn provisional heads to Zimmer Inc, Attention Product Service, 1777 West Center St. Warsaw, IN 46580.For questions regarding this recall call 574-372-4753. |  | 231641 | 231,641 | Worldwide Distribution - USA (nationwide) and internationally to Canada, Mexico, South America, Europe, Puerto Rico, China, and Korea. | Terminated | May 16, 2014 | 1141 | 231641
Z-0001-2013 | 62773 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | MIRS for MATRIX Spine System. Minimally Invasive Reduction Screw (MIRS) for percutaneous instrumentation of the thoracolumbar spine. |  |  | June 18, 2012 | 2012 | Synthes USA HQ, Inc. | Synthes is initiating a Medical Device Labeling Correction following a detailed review of the current labeling in response to reported complaints. Synthes has identified technique measures and has updated the labeling associated with this system. The labeling reflects updated technique recommendations and emphasizes the importance of avoiding rod mis-alignment. | Labeling design | SYNTHES sent an Urgent Medical Device Labeling Correction letter dated June 18, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to replace the previous version of the labeling with the new labeling provided. In addition, they can access an electronic version of the labeling by accessing Product Literature at the following link: http://us.synthes.com. The new version can be identified by code J11031-B located at the bottom right corner of the back cover. To confirm their receipt and understanding of the information contained in the letter, the recall package also included a response form to be completed by consignee and returned to Synthes.by Fax: 610-251-9005 or Scan/email: FieldAction@Synthes.com.Customers with questions were asked to call 1-800-620-7025, ext 5403 or contact their Synthes Spine Sales Consultant. For questions regarding this recall call 610-719-5377.The firm initiated this recall action on June 18, 2012 by sending an Urgent Medical Device Labeling Correction letter to all accounts. The updated labeling was Included with the notification letter. |  | 1602 | 1602 | Nationwide Distribution including CA, FL, ID, IN, MD, MI, MT, LA, NY, OH, PA, TX, VA and WA. | Terminated | September 10, 2015 | 1179 | 1602
Z-0003-2013 | 63152 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent, PN 447070 |  |  | January 27, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports of QC shifts, lot-to-lot variation, and an increase in the percent of patient samples recovering &gt;20 IU/mL using the Lots# M012376 and M101865 of IMMAGE RF reagent. The lots identified for the RF reagent might produce false abnormal results. | Nonconforming Material/Component | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated January 27, 2012 with an attached PCA Response Form to all customers who purchased the IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue use of IMMAGE RF reagent lots M012376 and M101865; call Order Assistance at (800) 526-3821 in the United States, (800) 463-7828 in Canada or contact your local Beckman Coulter representative for replacements; and complete and return the enclosed FAX BACK RESPONSE FORM within 10 days via mail at: Beckman Coulter, Inc., 250 S. Kraemer Blvd, Mail Stop A2.0246, Brea, CA 92822-8000; Attn: Regulatory Affairs or fax to: (714) 961-4234. Note: Beckman Coulter has added additional release testing that will assure future lots RF will meet the reference interval claim of the IMMAGE RF Instructions for Use (IFU).If you have any questions concerning this letter, contact Beckman Coulter Support Center (call Center) on the Beckman Coulter Website at: http://www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter Representative. |  | 925 | 925 units | Worldwide distribution: USA (nationwide) and country of: Canada. | Terminated | May 30, 2013 | 489 | 925
Z-0002-2013 | 63134 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COULTER PrepPlus, Part Number: 286600 |  |  | January 20, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Coulter PrepPlus may experience intermittant reagent level sense failures where the instrument may not produce a warning that a reagent vial is low or out of reagent. | Device Design | Beckman Coulter sent an Urgent Product Correction letter dated January 20, 2012 to all their customers who purchased the affected product. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were instructed to complete and return the Response Form with in 10 days. Customers with questions regarding the PCA letter were instructed to contact Beckman Coulter Service at (800) 526-7694 in the US and Canada or contact their local Beckman Coulter Representative. |  | 36 | 36 units total (15 units in the US) | Worldwide Distribution - US Nationwide and the countries of: Australia, Botswana, Canada, France, Germany, Japan, Malawi and Sweden. | Terminated | October 02, 2012 | 256 | 36
Z-0004-2013 | 63241 | No | EZD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Male, 16Fr, Straight-tip Intermittent Urinary Catheters M16. The product is distributed individually or in intermediate boxes containing 30 individual, sterile packaged catheters (PA/PE film and dialyzing paper pouches). |  |  | September 19, 2012 | 2012 | Cure Medical LLC | The firm initiated the product recall because labeling on some of the intermediate boxes of catheters identify the contents to be M16C, a coudÂż-tip catheters while the actual contents are M16Fr male, 16Fr, straight-tip catheters with the pouch and the case box labeled as M16Fr male, 16Fr, straight-tip catheters. | Labeling mix-ups | Cure Medical LLC sent an Urgent Medical Device Correction and Removal Notification letter dated September 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter instructed the customers to:a) Check their shelves for M16C with lot number 120425 and quarantine all affected product.b) Verify the contents (pouch label) are M16C and correspond with the intermediate box label.c) If all match, have the dealer sign and return a supplied verification record.d) If any mismatch is found, contact Cure Medical for assignment of a return goods authorization.e) Cure Medical will replace any incorrect product along with processing costs for this situation to the dealer.If affected product was further distributed customers were asked to identify those customers who received them and notify them of the problem. For questions customers were instructed to call 1-800-570-1778. |  | 4200 | 420cases (= 4,200 boxes) | Nationwide Distribution | Terminated | February 07, 2013 | 141 | 4200
Z-0005-2013 | 61780 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | IR1200, IR1250(IR1200 Plus), 2020 and One Touch(R) Ping(R) Insulin Pump glucose management system. |  |  | July 16, 2012 | 2012 | Animas Corporation | Wear of the keypad that may lead to damage such as tearing or peeling of the edge of the keypad from the pump. | Nonconforming Material/Component | The firm, Animas Corporation, sent an "IMPORTANT PATIENT NOTIFICATION" letter dated July 16, 2012 to its consignees/customers. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to examine their insulin pump for wear or damage on a regular basis, as part of their routine process of inspecting their pump. If keypad is worn or damaged, the buttons may not respond as intended. If consignees/customers see any damage or have issues with the keypad, they are to contact Animas through their website at www.animas.com/contact-animas or telephone Animas Customer Technical Support at 1-855-230-7574, or the phone number provided on the back of their insulin pump. Animas carefully investigates all complaints related to the insulin pump keypad.If you have any questions contact Animas Customer Technical Support at 1-855-230-7574. |  | 236470 | 236,470 | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, V, VT, WA, WI, WV, WY and Washington, DC and countries of: Austria, Australia, Canada, Czech Republic, France, Finland, Germany, Hungary, Israel, Italy, Mexico, New Zealand, Puerto Rico, Spain, and United Kingdom. | Terminated | January 15, 2014 | 548 | 236470
Z-0007-2013 | 62966 | No | JOL | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | BriteField illuminated Suction Device; |  |  | May 12, 2012 | 2012 | Invuity, Inc. | Invuity is notifying customers to return all obsolete BriteField Illuminated Suction Devices from the field. It has not been distributed since March 2011 as the next generation product was introduced at that time. | Other | Invuity sent E-mails to those locations thought to still be in possession of the old model of the device in July 2012. Distributors had been e-mailed in May 2012 to make sure that they had returned any of the device that was in stock.For questions regarding this recall call 415-655-2100. |  | 197 | 197 devices | Nationwide Distribution including AR, GA, SC, CA, OR, UT, and WA. | Terminated | April 23, 2013 | 346 | 197
Z-0008-2013 | 63253 | No | HGI | Gynecologic electrocautery and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sterile, single-use, disposable electrosurgical electrode holder. It connects the electrode to the foot activated electric generator. Label reads in part "***GYNEX*** Electrode Handpiece, Foot Activated*** Manufactured for GYNEX***" |  |  | September 06, 2012 | 2012 | Gynex Corporation | The wrong electrical connector pin was attached to the cord during assembly by the manufacturer. this electrical connector is too big to fit into the port on the foot activated generator. this makes the device unusable. | Nonconforming Material/Component | Consignees were first notified by phone with a follow up e-mail or fax. The consignees were asked to complete the response section regarding the destruction/return of the device.For questions regarding this recall call 425-882-1179. |  | 228 | 19 boxes (containing 12 electorde hand pieces each) | Nationwide Distribution | Terminated | October 22, 2012 | 46 | 228
Z-0006-2013 | 63196 | No | FGO | Ureteral stone dislodger. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Endoscopic retrieval device. |  |  | November 14, 2011 | 2011 | US Endoscopy Group Inc | When the outer catheter is either coiled or in a contorted configuration outside of the endoscope and the patient, and excessive force is applied to the handle, the jaws may become inoperable and may become stuck in the open condition. | Device Design | US Endoscopy initiated a Voluntary Product Recall letter initiated November 11, 2011. The letter identified the product the problem and the action needed to be taken by the customer.US Endoscopy will provide a free replacement or grant your facility a credit toward a future purchase for every device that is returned to US Endoscopy or verified as having been used or destroyed.For further questions please call 440-639-4494 ext. 378. |  | 1153 | 1153 | Worldwide Distribution -- USA (nationwide) including the states of AL, AR, AZ, CA, FL, GA, HI, IA, IL, KY, LA, MD, ME, MI, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, and WI; and to the following countries: Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Caribbean, Chile, Columbia, Croatia, Czech Republic, Denmark, Ecuador, El Salvador, Estonia, Finland, France, Germany, Great Britain, Greece, India, Ireland, Israel, Italy, Jordan, Lithuania, Malta, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, and United Arab Emirates. | Terminated | July 02, 2013 | 596 | 1153
Z-0009-2013 | 63064 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Wound Pro Apex Negative Pressure Wound Therapy Powered Suction Pump; Accuro Medical Products LLC, 2854 Hitchcock Ave, Downers Grove, IL 60515; part number 84635400182, sold inside the Apex Negative Wound Therapy system, part number 84635400181 |  |  | August 31, 2012 | 2012 | Accuro Medical Product LLC | Accuro has received reports of some Apex Pumps either failing to operate or malfunctioning, at times without notifying the user of such failure or malfunction on the unit's display screen. | Device Design | The firm, Accuro, sent an "Urgent: Medical Device Recall" letter dated August 30, 2012 to the customers via e-mail on August 31, 2012. The letter describes the product, problem and actions to be taken. The customers were instructed to remove all devices from use immediately; to immediately complete and return the enclosed Accuro response form via email to; (customer.service@accuromedical.com) or fax to: (1-800-866-0545) . Note: the form still needs to be returned even if you no longer have or use the device. Upon receipt, Customer Service will call the account with instructions to arrange for all devices to be returned to Accuro. General questions were directed to Customer Service at 1-800-669-4757. |  | 919 | 919 units | Worldwide distribution: USA (nationwide) and country of: Canada. | Open, Classified |  |  | 919
Z-0011-2013 | 63116 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL M4735A Defibrillator/Monitors, used with the M3516A batteries |  |  | September 11, 2012 | 2012 | Philips Healthcare Inc. | HeartStart XL Defibrillator/Monitor (M4735A) used on a patient when AC power is not in use, with a battery that may fail, the device may shut down unexpectedly without providing the user with a low battery warning message or audible tone. | Nonconforming Material/Component | The firm, Philips Healthcare, sent an "URGENT-MEDICAL DEVICE RECALL/Field Safety Notice" letter dated September 6, 2012 to its customers. The letter describes the product, problem and actions to be taken by the customer and also identifies what action Philips plans to take to remedy the issue. The customers were instructed to operate the XL and AC power instead of battery power alone; use a battery from a different date of manufacture (must not have the R-2011-12 date of manufacture code) if they need to operate the XL on battery power alone; and use an alternative source for treatment or therapy (i.e., backup defibrillator/monitor) if unable to operate with AC power alone or locate a battery with a different date of manufacture than that identified. Phillips will provide a replacement battery(s) to customers as soon as new production is available.If you need any further information or support concerning this issue, contact your local Philips representative or call Quality Specialists,Recalls at 978-687-1501 or email: debra.levasseur@philips.com. |  | 2500 | 2500 units | Worldwide distribution: USA (Nationwide) and countries of: AUSTRALIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, FINLAND, FRANCE, GERMANY HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA, REPUBLIC OF KUWAIT LEBANON, MACAU, MALAYSIA, MEXICO, NAMIBIA, NETHERLANDS, NORWAY, OMAN, PERU, PHILIPPINES, POLAND, PORTUGAL, ROMANIA RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VIET NAM. | Open, Classified |  |  | 2500
Z-0010-2013 | 63136 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 8 Fr. Bardex Pediatric All-Silicone Foley Catheter with a 3cc balloon. The product is packaged individually with 12 units per corrugated case. The product is sold sterile for single use and is used to provide continuous urinary drainage from the bladder. |  |  | September 07, 2012 | 2012 | C.R. Bard, Inc., Urological Division | The 8 Fr. Bardex All-Silicone Foley Catheter with 3cc balloon had an incorrect inflation volume (10 ml instead of 5 ml) printed on the valve cap. | Employee error | An "Urgent: Product Recall Notification" was sent to all customers and distributors in September 2012 via Federal Express. The letter described the product issue and provided recommended actions. Customers can contact Bard at (770) 784-6471 if questions. |  | 4032 | 4,032 units | Nationwide distribution: USA and Puerto Rico | Terminated | March 21, 2013 | 195 | 4032
Z-0016-2013 | 63140 | No | CIB | Angiotensin I and renin test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active Renin IRMA Kit, Part Number DSL25100 |  |  | January 20, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Active Renin IRMA Test Kit may exhibit a decrease in stability of the coated tubes. As a result, quality control and patient sample results may be falsely elevated by up to 72%. | Nonconforming Material/Component | Beckman Coulter sent an Urgent Production Correction letter dated January 10, 2012 to all their customers who purchased the Manual Immunoassay Active Renin IRMA Test Kit. The letter provides the customers with an explanation of the problem and actions to be taken. Customers were instructed to discontinue use and discard all remaining affected inventory. Customers were asked to complete and return the enclosed Response Form with in 10 days. Customers with questions contact Customer Service Support at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633. |  | 216 | 216 units total (6 units in the US) | Worldwide Distribution - US Natiowide including the state of: OH, CT, CA and the state of Argentina, Austria, Belgium, Brazil, Canada, China, Columbia, Czech Republic, France, Germany, Greece, Hungary, Italy, Kuwait, Netherlands, Poland, Portugal, Slovakia, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | October 05, 2012 | 259 | 216
Z-0013-2013 | 63059 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Buretrol Solution Set, 117" (3.0 m), with 150 mL Burette (Ball-Valve Drip Chamber) (Non-DEHP); a sterile Rx IV fluid pathway device; |  |  | September 07, 2012 | 2012 | Baxter Healthcare Corp. | Baxter has determined that the ball-valve feature of the Buretrol Solution Sets may not function as expected, allowing air to flow past the valve and into the tubing at the completion of dose within the burette. | Device Design | Baxter Healthcare Corp. sent an Urgent Product Recall letter dated September 7, 2012, via first class mail to all affected customers, informing them of the problem with the ball-valve, and instructing them to discontinue use of the affected product and to contact Baxter for instructions to return the affected product.. The customers were also requested to complete the attached customer reply form, confirming their receipt of the letter, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Medical Information Services at Baxter at 1-800-933-0303. |  | 1995040 | 1,995,040 units | Worldwide Distribution - US Nationwide and Internationally to Canada, Costa Rica, Australia, New Zealand, Hong Kong, and the United Kingdom. | Open, Classified |  |  | 5027417
Z-0025-2013 | 63154 | No | LGW |  |  | Neurology | PMA | The Eon Mini Implantable Pulse Generator (IPG) is a 16-channel, rechargeable, multi-program system that allows the connection of one or two leads with four, eight, and 16 electrode leads. It is powered by a hermetically sealed battery within a titanium case. It uses microelectronic circuitry to generate constant current electrical stimulation. Stimulation programs can be delivered as either single stimulation or MultiStim programs depending on patients need. |  |  | September 04, 2012 | 2012 | Advanced Neuromodulation Systems Inc. | Eon Mini IPGs manufactured in April 2012 could potentially exhibit a sudden, brief surge in stimulation that would be felt by the patient as uncomfortable or painful. SJMs routine testing, monitoring, and analysis of manufacturing processes identified three lots of Eon Mini IPGs that the internal battery has the potential to come into contact with the internal microcontroller board. | Nonconforming Material/Component | ST. Jude Medical sent an important Medical Device Advisory letter dated September 4, 2012 to all affected customers. The letter identified the affected product, problem and recommended actions to be taken. For questions contact your St. Jude Medical Neuromodulation Division Representative or call the Technical Services Department at 1-800-727-7846. |  | 56 | 56 units | US Nationwide Distribution | Terminated | May 12, 2014 | 615 | 56
Z-0024-2013 | 63211 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Reference materials "RETINAL OCT Analysis and Interpretation Method" shipped with iVue System (K091404) and the "Retina Poster" shipped with iVue System (K091404) and RTVue Systems (K101505); Models: iVue-100, RTVue-100; |  |  | September 07, 2012 | 2012 | Optovue Inc. | Devices labeled for an intended use not included in the existing 510(k) or PMA. Specifically, "The Retina" poster provided with the systems, and the book "RETINAL OCT Analysis and Interpretation Method" by Bruno Lumbroso and Marco Rispoli which are not labeled within 51O(k) cleared product description. | No Marketing Application | Optovue Inc., sent a Voluntary Field Correction and Removal letter dated September 6, 2012, to all affected customers. The letter identified the product the problem and the action to be taken by the customer .Please sign and return the attached form indicating your acknowledgement that that the Affected Materials have been returned or destroyed. If you have any further questions please contact Optovue, at 1-866-344-8948. |  | 795 | 795 units | Worldwide Distribution -- USA (nationwide) and Internationally. | Terminated | March 04, 2013 | 178 | 795
Z-0020-2013 | 63174 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Newport Portable Package, Model #50000, Serial # |  |  | April 29, 2012 | 2012 | Accutron Inc | Accutron is recalling the Flowmeter (Analgesia Gas Machine) because when it flows Nitrous Gas without Oxygen gas being turned on, it functions in an incorrect manner. | Device Design | Accutron sent an Urgent Advisory Notice letter dated April 6, 2012 to all their customers who purchased the Ultra PC% Cabinet Mount Flowmeter. The letter informed the customers of the problem identified and actions to be taken. Customers were instructed to discontinue use or to continue use of the device only with extreme care until a replacement can be installed. Extreme care includes the use of the titration method for nitrous oxide administration coupled with close monitoring of both the patient and the actual gas flow through the flowtubes. For question contact Accutron at (800) 531-2221 or by email at jheadrick@accutron-inc.com. |  | 220 | 220 units total | US Nationwide Distribution | Terminated | February 01, 2013 | 278 | 220
Z-0021-2013 | 63158 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Unicel DxH 800 Coulter Cellular Analysis System, Part Number 629029. |  |  | April 03, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is recalling the UniCel DxH 800 Coulter Cellular Analysis System because the NRBC, Differential or Reticulocyte mix chamber drain ports may become plugged by rubber pieces from specimen tube caps, and result in biohazard or chemical exposure and/or the potential for generating incorrect results. | Device Design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated April 3, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return an enclosed Response Form. Customers were notified that Beckman Coulter is in the process of developing and testing a design change that will prevent the mix chamber waste port from plugging. Customers with questions were instructed to contact Beckman Coulter Customer service at (800) 526-7694 in the US and Canada or contact their local Beckman Coulter representative. |  | 1281 | 1281 units total (693 units in the US) | Worldwide Distribution-USA (nationwide) and the counties of Australia, Belgium, Canada, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Republic of Kuwait, Macao, Malaysia, Mexico, Netherlands, Norway, Oman, Panama, Philippines, Puerto Rico, Qatar, Romania, Russian, Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | June 20, 2013 | 443 | 1281
Z-0023-2013 | 63019 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PREVI Isola System, Catalog numbers 29500 and 29500R, bioMerieux, Inc., Durham, NC. |  |  | August 28, 2012 | 2012 | Biomerieux Inc | The instrument may not dispense the specimen onto the pre-poured media resulting in a false negative result. | Labeling design | bioMerieux sent an "URGENT PRODUCT CORRECTION NOTICE" letter dated August 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. An Acknowledgement Form was included for customers to complete and return. Contact Customer Service at 1-800-682-2666 for questions regarding this notice. |  | 350 | 350 systems | Worldwide Distribution-USA (nationwide)-including DC and the states of CA, CO, CT, FL, GA, IL, IN, KY, LA, MO, NC, NJ, NY, OH, OR, PA, RI, SC, TX, and WI, and the country of Canada. | Terminated | September 04, 2014 | 737 | 350
Z-0026-2013 | 54725 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | OrthoHelix Part #HCS-056-32 3.2MM Cannulated Hex Driver NON STERILE MAT: 465 SS |  |  | January 29, 2010 | 2010 | Orthohelix Surgical Designs Inc | The tips of the drivers are breaking while being used to drive the screws into patients. | Software design (manufacturing process) | The firm, OrthoHelix Surgical Designs, Inc., sent an "IMPORTANT RECALL NOTICE: Mini MaxTorque HCS-056-32 &amp; HCS-056-38 Drivers" dated January 29, 2010 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue the use of these drivers immediately and return them to OrthoHelix Surgical Designs Inc, 1065 Medina Rd., Suite 500, Medina, OH 44256; ATTN: Driver Recall.Should you have any questions about system, call 330-247-1145 or the Director of Regulatory Affairs &amp; Quality Assurance at 330-247-1444. |  | 47 | 47 Pieces | Nationwide distribution: USA including the states of CA, CO, FL, GA, IL, MO, NC, NH, NJ, NM, NY, OH, OR, PA, SD, TN, TX, VA, WA, and WI. | Terminated | February 20, 2013 | 1118 | 167
Z-0022-2013 | 63214 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | 3M Steri-Vac Gas Sterilizer/Aerator models 5XL and 8XL. |  |  | September 13, 2012 | 2012 | 3M Company - Health Care Business | 3M determined that the electrical grounding mechanism for the Steri-Vac 5XL and 8XL gas Sterilizer/Aerator did not meet all of the requirements of the standard in that the electrical wiring, which is insulated and contained within the sealed door should have a stronger connection to the protective electrical grounding mechansim. | Component design/selection | 3m Healthcare Service Center sent a Urgent Medical Device Field Correction letter dated September 19. 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.3M will contact you to schedule an appointment for a field correction on any unit in the range of serial numbers provided in the table. 3Ms goal is to service the affected units with minimal disruption to your facility.We apologize for any inconvenience this field correction may cause. Please be assured that 3M is committed to the highest level of safety and customer service and we are proud of our longstanding safety record for 3M" Steri-Vac" Gas Sterilizers.For Further Information If you have any questions, please contact the 3M Health Care Service Line at 1-800-688-5888 or your local 3M representative.. |  | 707 | 707 units (61 US, 646 OUS) | Worldwide Distribution--USA (nationwide) including the states of AL, CO, CT, GA, IN, KY, MA, MI, MN, MO, MS, NJ, OH, PA, SC, TN, VA, WA and WI., and the countries of Argentina, Barbados, Chile, China, Colombia, Costa Rica, Ecuador, Germany, India, Indonesia, Japan, Korea, Kuwait, Lebanon, Malaysia, Mexico, Morocco, Pakistan, Paraguay, Peru, Philippines, Poland, Republic of Guinea, Russia, South Africa, Taiwan, Thailand, Trinidad, UAE, Venezuela, Vietnam and Zimbabwe. | Terminated | January 29, 2013 | 138 | 707
Z-0028-2013 | 63070 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Automatic Cutting Needles, Soft Tissue Biopsy Needle, NAC-1820M, Sterile, Remington Medical, Inc., 5830 Meadowridge Court, Alpharetta, GA 30005. |  |  | August 29, 2012 | 2012 | Remington Medical Inc. | Sterility of the product may be compromised. | Material/Component Contamination | The firm, Remington Medical, Inc., sent an "URGENT - MEDICAL DEVICE RECALL" letter dated August 29, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the products; complete and return the FaxBack form via Fax to: Quality Engineer Remington Medical, Inc., at 770-887-0732 or email: tracyw@remmed.com; and call the Quality Engineer or customer service at (800) 989-0057 ext 233 to obtain a RGA number and have a call tag issued.Should you have any questions, please do not hesitate to call the Quality Engineer at (800) 989-0057 ext 233. |  | 800 | 800 needles | Nationwide distribution: USA including states of: AL, AZ, CA, FL, IN, MA, NC, NY, OH, PA and TX. | Terminated | January 15, 2013 | 139 | 800
Z-0042-2013 | 62968 | No | NEO | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IDS-iSYS N-MidÂż Osteocalcin, Catalog Number IS-2900. Packaged in cardboard Box containing plastic bottles. |  |  | August 10, 2012 | 2012 | Immunodiagnostics Systems Ltd | It has been noted by staff at IDS Inc. that stock of the affected products held at IDS Inc. are labeled with IVD labels although they are not registered for sale as IVDs in the United States. The products should have been labeled as Research Use Only or Investigational use only dependent on their usage. | Labeling False and Misleading | Immunodiagnostics Systems sent Customer notification letters via email and notified customers of the recall. The notification included the reason for recall, product information with codes, lot number, manufacturing/distribution date, and exp date. The notification included instructions to customers: Actions to be taken by the customer:Customers should complete the attached Recall and Acknowledgement Form and take the following actions:1) Check your stock for the products included within the scope of the recall. Cease use and distribution and quarantine all affected product lots immediately.2) Complete the enclosed Recall &amp; Acknowledgement Form and immediately fax back to Immunodiagnostic Systems. This will allow us to document your receipt of this letter and the amount of product you have on hand for return.3) Return any affected product freight collect, along with the original completed Recall &amp; Acknowledgement Form to your local IDS office.Contact information was also provided.For questions regarding this recall call +44 (191) 5195212. |  | 5 | 5 | Nationwide Distribution including NY, CA, ME and CT. | Terminated | December 17, 2012 | 129 | 56
Z-0037-2013 | 63157 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ATLANTIS(R) Anterior Cervical Plate System FIXED ANGLE SCREW MODULE, REF 176-504, QTY: 1 EA, Rx only, Medtronic Sofamore Danek USA, Inc., 1800 Pyramid Place, Memphis, Tennessee 38132. |  |  | August 30, 2012 | 2012 | Medtronic Sofamor Danek USA Inc | Lines on the screw gauge may not placed on the caddy properly which could result in improper measurement of screws. An incorrect gauge could measure a screw to be approximately 1 mm shorter than its actual length. | Nonconforming Material/Component | Medtronic conducted a conference call on August 30, 2012 and sent an "URGENT VOLUNTARY RECALL" notice dated September 3, 2012 to all affected customers. The notice identified the product, problem, and actions to be taken by the customers. A Recall Questionnaire Form was attached to the letter for customers to complete and return. Contact your Medtronic Sales representative or the firm's Global Quality Department at 800-876-3133, ext. 3197 for questions regarding this notice. |  | 1014 | 1014 units | Worldwide Distribution-USA (nationwide) and the countries of India, Netherlands, Hong Kong, S. L Korea, Singapore, Canada, Australia, Japan, and Mexico. | Terminated | August 15, 2013 | 350 | 2308
Z-0031-2013 | 63085 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Caremor Cliner"***Manufactured By Winco***Ocala, FL 34474 www.wincomfg.com***Model No. 536 Serial No. 536A-101475***" |  |  | July 10, 2012 | 2012 | Winco Mfg., LLC | Winco Mfg., LLC in Ocala, FL is recalling the Caremor Cliner model numbers 5351 and 5361. The recall was initiated due to a report that a back frame did not "snap" onto a chair frame properly. | Process control | Winco notified consignees via phone except for the one located in Saudi Arabia, who was contacted by e-mail. All consignees have been sent a follow-up e-mail. The notification identified the product, the problem, and the action to be taken to be taken by the customer. Customers should remove the chair from use and the upholstered back will not be required to be returned. Winco will send out new fully assembled upholstered backs. For questions customers should call 352-854-2929, ext 123. |  | 23 | 23 chairs | Worldwide Distribution - USA including Arkansas, California, Maryland, Mississippi, New York, Pennsylvania and the country of Saudi Arabia | Terminated | May 23, 2013 | 317 | 43
Z-0047-2013 | 63169 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Folate WHO Calibrators, Part Number: A98033. |  |  | July 24, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access Folate Calibrators because it may produce falsely elevated results that could mask a folate deficiency. | Nonconforming Material/Component | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated July 24, 2012 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. Customers were instructed to complete and return the enclosed Response Forms within 10 days. Customers with questions were instructed to call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter representative. |  | 5266 | 5,266 units total (3,214 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Austria, Australia, Bangladesh, Belgium, Botswana, Brazil, Canada, China, Colombia, Czech Republic, Croatia, Ecuador, Ethiopia, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Malaysia, Mexico, Morocco, Namibia, Netherlands, New Zealand, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Switzerland, Sweden, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, and Zimbabwe. | Terminated | October 17, 2013 | 450 | 5266
Z-0033-2013 | 63161 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CYTO-STAT tetraCHROME CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5, Part Number 6607013. |  |  | May 22, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has determined that the specimen and prepared sample stability claims stated in the product labeling for tetraCXP software and CYTO-STAT reagents (CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5 and CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5) need to be updated for alignment with current standards and clinical decision-making associated with lymphocyte immunophenotyping | Process control | Beckman Coulter sent a Product Corrective Action (PCA) letter dated May 22, 2012, with attached PCA Response Form to all customers who purchased the tetraCXP Software, CYTO-STAT tetraCHROME CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5, CYTO-STAT tetraCHROME CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were informed that the tetraCXP SYSTEM Guide and CYTO-STAT tetraCHROME CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5 &amp; CYTO-STAT tetraCHROME CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5 product labeling will be revised to reflect the updated claims. Customers were instructed to complete and return the enclosed Response form within 10 days. Customers with questions were instructed to contact Beckman Coulter Customer Service at (800) 526-7694 in the US and Canada or contact their local Beckman Coulter representative.For questions regarding this recall call 714-961-4483. |  | 137445 | 137,445 units total (26,155 units in US) | Worldwide Distribution - USA (nationwide) and the following countries:AlbaniaAustraliaAustriaBahrainBangladeshBarbadosBelarusBelgiumBosnia and HerzegovinaBulgariaBurundiCanadaChileChinaColombiaCroatiaCzech RepublicDenmarkEcuadorEl SalvadorFinlandFranceFrench GuianaFrench PolynesiaGermanyGreeceGuadeloupeHong KongIndiaIndonesiaIraqIrelandIsraelItalyJapanKenyaKorea, Republic ofKuwaitLibyan Arab JamahiriyaLithuaniaMacaoMalaysiaMexicoMonacoMoroccoNamibiaNetherlandsNew ZealandNigeriaNorwayOmanPakistanPanamaPolandPuerto RicoQatarReunionRomaniaRussian FederationSaudi ArabiaSingaporeSlovakiaSloveniaSouth AfricaSpainSudanSwedenSwitzerlandTaiwanTurkeyUnited Arab EmiratesUnited KingdomUruguay | Terminated | October 10, 2012 | 141 | 137445
Z-0030-2013 | 62832 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Emit(R) 2000 Phenobarbital Assay (OSR4D229) and Syva(R) Emit(R) Phenobarbital Assay (4D019UL) |  |  | July 30, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed highter imprecision with the recalled lots on Beckman Coulter AU(R) Systems. Also, if the Syva(R) Emit(R) 2000 Phenobartbital Assay (4D019UL) lots D4, D5, D6, and/or D7 as an application on AU(R) Systems is used, the issue may be experienced. Firm has found that within-run imprecison of the recalled lots when run on an AU(R) System may be higher that what would normally be obs | Nonconforming Material/Component | Siemens sent an Urgent Field Safety notice letter dated July 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and discard any remaining inventory of the affected product. Customers were asked to share the content of the letter with laboratory director regarding the need to review previous Phenobarbital results, conduct patient follow up and/or repeat testing. For a no charge replacement of remaining inventory customers were instructed to contact Siemens Customer Service Center at 888-588-3916. For questions contact Siemens Technical Solutions Center at 800-227-8994. |  | 79744 | 79,744 cartons | United States Nationwide Distribution in the states of: CA, FL, KS, MA, MN, MS, NM, NY, NC, OR, PA, TN, TX, UT, VA, and WA. | Terminated | August 27, 2014 | 758 | 79744
Z-0029-2013 | 63175 | No | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Triage TOX Drug Screen Panel, Model# 94400. |  |  | May 28, 2009 | 2009 | Alere San Diego, Inc. | Biosite is recalling the Triage TOX Drug Screen kits where the box containing test devices because it was mislabeled. | Error in labeling | Biosite sent a Notice of Correction letter dated May 28, 2009 to all their customers who purchased the Triage TOX Drug Screen. The letter informed the customers of the problem identified and actions to be taken. Customers are instructed that performance of Triage TOX Drug Screen test devices is not affected in any way by the misprinted box flap, provided the test is used with a urine sample as per the instructions for use. Customers are instructed to complete and fax the enclosed verification form within 10 days. In the US, customers were instructed to contact Biosite at (877) 441-7440 or email productadvisory@invmed.com. Customers in UK are instructed to contact European Representative, Lesley Paice, lesley.paice@invmed.com. |  | 45638 | 45,638 units | Worldwide Distribution - US Nationwide and the countries of Australia, Belgium, Brazil, Canada, Colombia, Czech Republic, Finlind, France, Germany, Great Britian, Greece, Guatemala, Hong Kong, India, Indonesia, Italy, Lebanon, Mexico, Netherlands, Norway, Puerto Rico, Romania, Saudi Arabia, South Africa, South Korea, Spain, Switzerland, Taiwan, Tunisia, Turkey, United Arab Emirates, and Uruguay. | Terminated | October 10, 2012 | 1231 | 45638
Z-0048-2013 | 63045 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza radiological image processing system. |  |  | June 06, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens became aware of an unintended behavior when using syngo.plaza. In syngo.plaza within the online "Short Time Storage" (STS), some studies are marked with an archived icon, however, the entry in the Archive Database referring to the physical file may be missing. | Software design | The firm initiated this correction on June 6, 2012 by issuing update instructions SY129/12/S to consignees. A script will be distributed to all potential affected customer systems which will prevent data loss by setting the delete protection on all affected studies. |  | 11 | 11 | Nationwide Distribution including the states of MO, FL, MA, NY, WI, GA, CA and WA. | Terminated | April 13, 2013 | 311 | 11
Z-0049-2013 | 63079 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Battery, 100/115V (Includes charger and connecting cable) |  |  | September 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received multiple reports since 2004 of variousmalfunctions of the battery backup system for the Sarns" Centrifugal System.The reports can be broadly categorized as follows:" Failure to charge" Failure to hold charge" Failure to operateIn the most severe case, a battery malfunction could result in an unexpected loss of arterial flow. In | Component design/selection | TERUMO sent an Urgent Medical Device Recall letter dated June 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. On 09/25/2012 an URGENT MEDICAL DEVICE RECALL notification and addendum was sent to all consignees. TCS issued this notification as a Safety Advisory to alert all users of the Sarns centrifugal system of the potential for malfunction and remind them to review the appropriate sections of the Operators Manual for the Sarns Delphin Battery. Providing an addendum to the Operators Manual for the Sarns Delphin Battery with additional instruction on the use and charging of a discharged battery.Customers were instructed to review the Medical Device Recall notice; assure that all users are aware of the notice; place the addendum in the Operator's Manual; and confirm receipt of the communication by faxing or emailing the attached Customer Response Form.For questions customers were instructed to call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 1076 | 1076 | Worldwide distribution: USA including: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, UT, VA, WA, WI, and WV. Puerto Rico and internationally to:ALEXANDRIA, ARGENTINA, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CENTRO AMERICA, CHILE, COLOMBIA, Dominican Republic, Ecuador, FRANCE, GERMANY, GREECE, Guatamala, HOLLAND, HONG KONG, India, ITALY, Japan, KL SEGUNDO, KOREA, LEICESTERSHIRE, MALAYSA, Malaysia, Mexico, Mexico, Mexico, NETHERLANDS, Panama, Peru, Philippines, RALEIGH, Republic of Georgia, SANTIAGO CHILE, Saudi Arabia, SEOUL KOREA, SHANGHAI, CHINA, Singapore, South Africa, SPAIN, SWITZERLAND, TAIWAN, Thailand, UNITED ARAB EMIRATES (UAE), Uruguay, USA, and Venezuela, | Terminated | January 15, 2014 | 477 | 1365
Z-0055-2013 | 63362 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | 3 PORT "OFF" MANIFOLD w/WASTE BAG, SYRINGE and |  |  | September 28, 2012 | 2012 | ICU Medical, Inc. | ICU Medical, Inc. has identified a potential comingling of components where a 60 drop micro dripchamber may have been included on a 20 drop macro drip set. | Mixed-up of materials/components | ICU Medical sent an "URGENT:MEDICAL DEVICE RECALL" letter dated September 28, 2012 to its customer. The letter identified the product, problem, and actions to be taken by the customer. The customer has been instructed to inspect their inventory for the specific lot numbers, examine for the presence of a metal cannula, and certify destruction of any affected devices. The customer letter included a Verification Form to be completed and returned. Contact ICUCustomer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address: customerservice@icumed.com" for questions regarding this notice. |  | 60 | 60 | Nationwide distribution: Virginia only. | Terminated | February 19, 2013 | 144 | 60
Z-0051-2013 | 62059 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Centrifugal System Control module, 100/115V |  |  | September 24, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received one report of loss ofpower to the control module and subsequently the drive motor for the Sarnscentrifugal system during cardiopulmonary bypass. | Component design/selection | TERUMO sent an Urgent Medical Device Recall dated April 13, 2012, to all affected consignees. An Addendum was sent on September 25, 2012. All US Consignees will receive a Safety Advisory and an update to the Operator's Manual via Federal Express to inform them of the reported malfunction and potential hazards and to raise awareness of the benefits of using a retrograde valve and the importance of back-up system availability. TERUMO will include a customer response form and follow up with all affected users to obtain a response to the recall.For questions customers should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 847 | 847 units | Worldwide Distribution - USA (nationwide) and the countries of Bangladesh, Seoul Korea, Chile, Brazil, Venesuela, Thailand, Mexico, India, Taiwan, Canada, Dominican Republic, Guatemala, Ecuador, Germany, Republic of Georgia, Hong Kong, Japan, United Arab Emirates (UAE), Japan, Belgium, Phillipines, Singapore, and South Africa. | Terminated | November 13, 2013 | 415 | 1186
Z-0054-2013 | 63111 | No | HXS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | TORQ Sternal Closure Device. The common name is TWISTER, WIRE. The model number is TQ01A. |  |  | August 28, 2012 | 2012 | Kardium | The TORQ Sternal Closure Device, Lot 062711, is being recalled due to probability one of the metal parts may be dislodged during use. | Employee error | Kardium Inc. sent an "URGENT: TORQ STERNAL CLOSURE DEVICE RECALL" e-mailed dated August 30, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use of the TORQ Sternal Closure Device lot 062711 immediately. Contact the firm at 604-248-8891, x285 for any questions about this recall. |  | 480 | 480 units | Nationwide Distribution-including the states of FL, GA, IL, and NH. | Terminated | February 05, 2013 | 161 | 480
Z-0058-2013 | 62703 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Belmont Rapid Infuser, a Fluid Management System, |  |  | September 25, 2012 | 2012 | Belmont Instrument Corporation | Leak from a saline bag caused saline to drip along the back of the Rapid infuser close to the Power Entry Module causing an electrical flash at the power entry area. | Device Design | Belmont Instrument initiated recall on 9/25/12. A "Field Correction" notice was provided that described the product issue and recommended actions. US-based sales representatives visited all facilities to provide the modified power cord/shield. For units sold outside the United States, the modified power cords were sent to distributors who would ensure the new power cords were provided to all facilities within their accounts. |  | 2913 | 2,913 units | Worldwide Distribution including Nationwide (USA) and the countries of Australia, Austria, Belgium, Brazil, Canada, China, Denmark, England, Estonia, France, Germany, Hong Kong, India, Israel, Italy, Jordan, Mongolia, Netherlands, New Zealand, Norway, Pakistan, Peru, Puerto Rico, S. Korea, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Turkey, UK, United Arab Emirates, and Vietnam. | Terminated | February 28, 2014 | 521 | 2913
Z-0056-2013 | 63155 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate 1250 and AutoMate 2550, Part Numbers ODL25125 and ODL25255. |  |  | April 11, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports of secondary tubes being labeled with two different barcode labels on the AutoMate 2500 Family (previously referred to as OLA2500 Systems) units which can cause misidentification of a sample when software versions lower than version 4.1 is used. | Labeling False and Misleading | Beckman Coulter initiated an Urgent Product Correction letter dated April 11, 2012, with attached PCA Response Form to all customers who purchased the affected product. . The letter identified the product, the problem, and the action taken by the customer. Customers were instructed to complete and return the enclosed Response Form within 10 days. Customers with questions were instructed to contact their local Beckman Coulter representative. Customers with questions should call 714-993-5321.For questions regarding this recall call 714-961-4941. |  | 468 | 468 units total (51 units in US) | Worldwide Distribution - USA (nationwide) and the following countries:ArgentinaAustraliaAustriaBelarusBelgiumChinaCroatiaCzech RepublicDenmarkFinlandFranceGermanyHungaryIndiaIrelandIsraelItalyJapanKorea, Republic ofKuwaitMacaoMalaysiaNetherlandsNew ZealandNorwayPanamaRussian FederationSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTurkeyUnited Kingdom | Terminated | February 11, 2013 | 306 | 468
Z-0059-2013 | 63191 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Unicel Closed Tube Aliquotter (UCTA) for UniCel DxC 660i, UniCel DxC 680i, UniCel DxC 860i, UniCel DxC 880i Synchron Access Clinical Systems. Part Number A45853. |  |  | July 06, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the current sample syringe pump drive assemblies installed on DxC Integrated ystem with UniCel Closted Tube Aliquotter (UCTA) serial numbers between 547 and 574 may cause a hardware motion error (ID0x08000501). | Nonconforming Material/Component | The recall notification was initiated on July 6, 2012 with Beckman Coulter forwarding a Product Corrective Action (PCA) letter with attached PCA Response Form to all customers who purchased the UniCel DxC 880i, UniCel DxC 860i, UniCel 680i, UniCel DxC 660i Clinical System Integrated Workstations UCTA Sample Syringe Pump Drive Assembly. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were notified that Beckman Coulter will schedule a Service Call to replace the affected sample syringe pump drive assemblies. Customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with any questions regarding the PCA, were instructed to call Customer Support at (800) 854-3633 in the US and Canada. Outside the US and Canada, customers were instructed to contact Beckman Coulter Representative. |  | 18 | 18 units | Worldwide Distribution and US (nationwide) and the countries of China, France, Italy and Taiwan. | Terminated | November 15, 2012 | 132 | 18
Z-0057-2013 | 62879 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hemorrhoidal Circular Stapler and Accessories 33mm |  |  | August 03, 2012 | 2012 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated this voluntary global recall for specific production lots of PROXIMATEÂż PPH Hemorrhoidal Circular Stapler and Accessories 33mm (PPH01 and PPH03) and TranstarÂż Circular Stapler Procedure Set (STR10) due to difficulty firing the device which may result in incomplete firing stroke that may result in an incomplete staple formation. | Process change control | Ethicon Endo-Surgery sent an Urgent Device Recall letter dated August 3, 2012 via overnight, Federal Express UPS to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately, discontinue use, quarantine the material and return per the instructions provided. Customers were advised to complete the Business Reply Form (BRF) and fax it to Stericycle at 1-888-628-0733 within 3 business days, even if you do not have affected product. The letter states that Ethicon Endo-Surgery will issue credit for recalled products returned by November 5, 2012. Any product returned after November 5, 2012 will not be eligible for credit. For clinical or product support call 1-800-873-3636, option 6 or contact your local sales representative and reference Event 7283. |  | 157881 | 157,881 devices | Worldwide Distribution - US (nationwide) in the states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and the countries of: ARGENTINA, AUSTRALIA, BAHRAIN, BELGIUM, BRAZIL, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, GREECE, GUATEMALA, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, MALAYSIA, MEXICO, NEPAL, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, PUERTO RICO, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIETNAM, and YEMEN. There was one Canadian consignee and 43 U.S. Government consignees. | Terminated | July 12, 2013 | 343 | 157881
Z-0061-2013 | 63124 | [6 empty] | CPT Hip System Femoral Stem Petite, Item No. 00-8011-000-05. |  |  | September 13, 2012 | 2012 | Zimmer, Inc. | Zimmer Inc. is initiating a removal of the CPT Modular Stem (item 00-8011-001-00), CPT Stem Petite (item 00-8011-000-05; ) and CPT Femoral Stem 12/14 Neck Taper Extended Offset,(item 00-8114-001-10 ) because devices were packaged in incorrect cartons and distributed internationally and domestically. Zimmer received five complaints of this packaging issue. | Packaging change control | Zimmer Inc. sent an "URGENT MEDICAL DEVICE RECALL-LOT SPECIFIC" letter dated September 13, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Consignees were asked to return all recalled product with a completed Inventory Return Certification Form. Contact the firm at 1-800-348-2759 for questions regarding this notice. |  | 20 | 20 | Worldwide Distribution-USA (nationwide) including the states of Indiana and Oklahoma and the countries of Australia, Thailand, Malaysia, Germany, India, Bulgaria, Egypt, UK, Russian FED, Slovak, and Brazil. | Terminated | January 15, 2014 | 489 | 138
Z-0066-2013 | 63115 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | "***PressFT***2.1 w/ Two #0 (3.5 metric) Hi-Fi***Sutures***NP212***2009-02***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA*** |  |  | June 26, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | Linvatec Corporation dba ConMed Linvatec in Largo, FL is recalling PressFT HI-FI Sutures (items NP211, NP212, NP261, NP262, NP211H, NP261H-all lot numbers) due to the potential for the distal tip of the driver to become separated from the driver shaft after successful insertion of the implant. | Process control | ConMed Linvatec sent an Urgent Medical Device Recall Notification letter on June 26, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their facility for the recalled products, segregate them for return to ConMed Linvatec. The customers were also told to return the REPLY FORM to the Regulatory Department at the firm. For questions customers were instructed to call 800-237-0169 or email at Custserv1@linvatec.com.For questions regarding this recall call 727-392-6464. |  | 879 | 879 units for all products in recall event | Worldwide Distribution - USA including AZ, CA, CO, FL, GA, LA, MA, ME, MI, MN, NJ, NM, NY, PA, TX, UT, and WA. Internationally to BE, CA, ES, MY, NZ and PL | Terminated | October 03, 2014 | 829 | 879
Z-0071-2013 | 63225 | No | KRI | Cardiopulmonary bypass accessory equipment. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Backplate Replacement CXFVR 1500 Holder |  |  | October 02, 2009 | 2009 | Terumo Cardiovascular Systems Corporation | Backplate reservoir holders were manufactured with an incorrect configuration of connector supports. The devices were manufactured and configured with two 3/8 inch and one Âż inch connector supports instead of one 3/8 inch and two Âż inch connector supports (See location of supports as # 7 and # 8 on backplate attachment). The backplates were manufactured as described in manufacturing instructions a | Labeling design | The firm, Terumo, interviewed the Sr. Product Manager from the Marketing Department of the company on August 29, 2012. The interview included the description of the product, problem and action taken. Replacement connectors were provided to users and all units were corrected.If you have any questions, call 734-663-4145 or Manager, Regulatory Affairs at 734-741-5576. |  | 33 | 33 units | Worldwide distribution: USA (nationwide) including states of: CA, MI, NC, NJ, PA, and WI;and country of: EUROPE. | Terminated | January 29, 2013 | 1215 | 33
Z-0114-2013 | 63027 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare innova 3100-IQ, Mobile X-ray system |  |  | December 20, 2011 | 2011 | GE Healthcare, LLC | It was discovered of a potential failure of the spectral filter mechanism inside the collimator of the innova 3100, innova 3100-IQ and innova 3131-IQ systems. | Radiation Control for Health and Safety Act | GE Healthcare will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services. A GE Healthcare Service Representative will replace the collimator to address the issue.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification. Further questions please call 262-513-4122. |  | 199 | 199 installed in the US | Nationwide distribution | Terminated | September 18, 2013 | 638 | 199
Z-0072-2013 | 63359 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ***This recall is being conducted due to an incomplete recall for RES 57829*** Curlin Infusion Administration Set, REF 340-4165, Sterile EO, |  |  | September 26, 2012 | 2012 | MOOG Medical Devices Group | This recall is the same issue for the FDA processed recall initiated by Moog Medical Devices on 1/25/11 which was related to customers reporting unclearable Air-In-Line alarm with specific lot numbers of Curlin Administration Sets used with Curlin Ambulatory Infusion Pumps. The lots associated with this recall were inadvertently omitted from the original recall which has since been terminated by | Process control | MOOG sent a Medical Device Recall Notification letter dated September 26, 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove product from their inventory and return to the firm. Customers were also instructed to contact their distributor to arrange for replacement product. For questions customers were instructed to call 1-800-970-2337, prompt #7.For questions regarding this recall call 801-264-1001. |  | 12000 | 12,000 sets | Nationwide Distribution including California | Terminated | April 05, 2013 | 191 | 12000
Z-0073-2013 | 62516 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory : |  |  | November 15, 2011 | 2011 | Sunquest Information Systems, Inc. | Online calculations may not be performed as expected or may yield results that are different than expected when the calculation module is called from the Online Entry (OEM) application for a subtest whose subcolumn number on the instrument online string is divisible by 10. | Software design | Urgent - Product Safety Notice: PSN-11-06,dated, 15 November 2011, was sent to customers. Costumers should follow instructions in the notification letter. To request the software correction, create a Case via the Client Support webpage at: www.sunquestinfo.com/supportweb.Return receipt acknowledging PSN-11-06 is required .Client Support can be reached at 1-877-239-6337 for U.S. and Canadian sites and at +44 (0) 161 335 0562 for International sites. |  | 127 | 127 sites | Nationwide distribution | Terminated | October 19, 2012 | 339 | 127
Z-0079-2013 | 63336 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LeVeen Standard Needle Electrode System, Sterile (3.5/12), (3.5/15), (4.0/15), (3.5/25), (5.0/15), (5.0/25), and (4.0/25). |  |  | September 17, 2012 | 2012 | Boston Scientific Corporation | Leveen Standard and Leveen Coaccess Needle Electrode System, Directions for Use(DFU) in Portuguese, are incorrect. all other languages in the DFU are correct. Consignees in Brazil, Portugal,and Hong-Kong are being notified. | Labeling Change Control | Boston Scientific sent notifications, dated September 17, 2012, to each account in Brazil, Portugal and Hong Kong affected by the recall. The affected accounts were provided the corrected pages of the Portuguese DFU as part of the communication to ensure that those users that rely on the Portuguese DFU have the correct information. |  | 19 | 19, 716 total all units | Distributed in Brazil, Portugal and Hong Kong. | Terminated | July 17, 2013 | 303 | 19
Z-0087-2013 | 62569 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Grafton DBM Crunch |  |  | April 20, 2012 | 2012 | Osteotech Inc | Multiple lots of GRAFTON Âż and XPANSE Âż, which are used for bonegrafting procedures are being recalled because there is a possibility that sterility of the outer surface of the inner pouch may becomecompromised. | Device Design | Osteotech sent out recall letters/questionairres dated April 20, 2012, to Risk Managers/service representatives. The recall was expanded on August 1 2012. The letter identified the product, the problem, and the action to be taken by the customer. Sales Representatives would facillitate the return of any affected product they may have or provide cautionary labels for any affected product and/or support in the completion of the attached questionnaire. Customers were asked to disseminate this information to additional personnel within their facility as appropriate and return the questionnaire to Medtronic.. Customers with questions should contact their Medtronic Sales Representative at 800-876-3133, extension 3197.For questions regarding this recall call 732-542-2800.Osteotech sent an additional letter stating that The OPTIUMTM (DBM) product line was added in error. |  | 3541 | 3,541 units | Worldwide Distribution - USA (nationwide) and Internationally. | Terminated | January 17, 2014 | 637 | 79207
Z-0078-2013 | 63365 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hem-o-lok Large Polymer Ligating Clips, REF #544240, Rx only, Sterile, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | September 25, 2012 | 2012 | Teleflex Medical | Product is being recalled due to the possibility that the tray may contain pinholes, compromising product sterility. | Packaging process control | Teleflex sent a Urgent Medical Device Recall Notification dated September 25, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to check their inventory and quarantine any affected inventory found. They were to complete the enclosed Field Corrective Action Acknowledgement Form and return it to the Teleflex Medical sales representative who will then remove the affected product from the facility and ensure its return to Teleflex Medical.We sincerely apologize for any inconvenience this action may cause your operations. If you have any questions please contact your local sales representative or Customer Service at 1-866-246-6990. |  | 126 | 126 units | Nationwide Distribution including the state of Nevada. | Terminated | November 08, 2012 | 44 | 126
Z-0082-2013 | 63223 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Columbus 4-IN-1 Femoral Cutting Guide M3 Instrument (NQ083R) |  |  | May 09, 2012 | 2012 | Aesculap, Inc. | The manufacturer informed the recalling firm of a potential laser marking error in size identification on two Columbus 4-IN-1 Femoral Cutting Guide Instruments. The guides are correctly dimensioned, but the laser markings have been incorrectly etched. | Process change control | Both consignees were notified verbally on May 15, 2012 and asked to inspect their inventory for the recalled items, remove the mislabeled instruments and return them to the recalling firm. Both mislabeled instruments were returned to the firm. Customer questions were directed to (610) 984-9074. |  | 2 | 2 | Within the US, product was distributed to PA and TN. | Terminated | January 15, 2014 | 616 | 2
Z-0074-2013 | 63086 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | "***Electric Systems Foot Control****Rx Only***" |  |  | January 09, 2012 | 2012 | The Anspach Effort, Inc. | The Anspach Effort, Inc. in Palm Beach Gardens, FL is recalling the foot control/foot pedal due to the Foot Control Cable lacking the shield connection from the cable to the connector body. | Nonconforming Material/Component | Anspach sent an Urgent Product Removal letter dated January 9, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to contact their sales representative to facilitate the removal action by providing a replacement for the affected product. For question contact your local sales representative or the Product Support department. |  | 16 | 16 units | Worldwide Distribution - US (nationwide) and the countries of: Italy, Japan, Switzerland, Turkey and Taiwan. | Terminated | September 16, 2014 | 981 | 67
Z-0077-2013 | 63255 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Helion and Xenion surgical lights, 501/01 series KLM lights. |  |  | September 19, 2012 | 2012 | Trumpf Medical Systems, Inc. | Incidents regarding fractures of the front joint of the spring arm have been reported. | Device Design | Trumpf sent an "URGENT RECALL NOTICE" dated September 19, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Confirmation of Receipt Form was attached for customers to complete and return. Contact the firm at 888-474-9359 for questions regarding this notice. |  | 275 | 275 units | Nationwide Distribution-USA (nationwide) including the states of AR, CA, CO, CT, FL, GA, IL, KS, MA, MI, MO, NC, NJ, NY, OH, SC, TX, UT, VT and WI. | Terminated | January 21, 2014 | 489 | 275
Z-0070-2013 | 63056 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq Two Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids; Hospira, Inc., Lake Forest, IL 60045 USA, Made in Costa Rica; List Number 16027-01 |  |  | August 29, 2012 | 2012 | Hospira Inc. | The Symbiq pump touchscreen may not respond to user selection, may experience a delayed response, or may register a different value from the value selected by the user. Failure of the touchscreen to respond to user input could result in a delay in therapy or result in over delivery or under delivery of medication if the user does not confirm the programmed values on the pump's confirmation screen prior to starting the infusion. | Device Design | Hospira sent an Urgent Device Correction letter dated August 29, 2012, to all affected customers. Customers have been informed of reports that the Symbiq pump touchscreen may not respond to user selection, may experience a delayed response, or may register a different value from the value selected by the user. Failure of the touchscreen to respond to user input could result in a delay in therapy or result in over delivery or under delivery of medication if the user does not confirm the programmed values on the pump's confirmation screen prior to starting the infusion. The users were instructed to complete and fax the enclosed reply card to 1-888-943-5180 or email it to Hospira194@stericycle.com indicating the number of infusers at the facility and if they notified their users throughout the facility. Wholesalers were asked to notify their customers. If they have further distributed the affected product, they should notify their accounts and ask them to contact Stericycle at 1-877-272-6141 to receive a reply form. For questions regarding this recall call 224-212-2000. |  | 14067 | 14,067 units | Worldwide Distribution - USA (nationwide) Puerto Rico and Internationally to Canada | Terminated | March 15, 2017 | 1659 | 33681
Z-0095-2013 | 61603 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Seno Advantage 1.x, Seno Advantage 2.0, Seno Advantage 2.1 |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of an issue associated with the magnification images acquired on non-GE digital mammography systems and your Seno Advantage Workstation that may impact patient safety. The measurement values provided by the Seno Advantage Workstation may be incorrect when applied to magnification images acquired on non-GE digital mammography systems. If not noticed by the caregiv | Software design | GE Healthcare sent a Urgent Medical Device Correction letter dated May 23, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.GE Healthcare will install new software to correct this issue with the release of Field Modification Instruction (FMI) 12129 in the near future. This correction will not affect the performance or any other functionality of your current equipment. A GE Healthcare service representative will contact you to arrange for this modification. This activity will be performed at no cost to you.If you have any questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800 437 1171, Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. |  | 188 | 188 | Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, DC, FL, GA, KY,LA, MA, MO, NV, NJ, NY, NC, OH, OK, PA, SD, TN, TX, VA and WI., and the countries of UNITED KINGDOM, TUKEY, TAIWAN, SPAIN, SINGAPORE, SAUDI ARABIA, RUSSIA, PORTUGAL, NEW ZEALAND, NETHERLANDS, MEXICO, MALAYSIA, KOREA, JAPAN, ITALY, HUNGARY, GREECE, GERMANY, FRANCE, CHINA, CANADA, BRAZIL, AUSTRIA and ALGERIA. | Terminated | March 26, 2013 | 306 | 188
Z-0108-2013 | 63276 | No | CBT | Arterial blood sampling kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Pulset Arterial Blood Gas (ABG) sampling kit. |  |  | January 18, 2011 | 2011 | Westmed Inc | Westmed is recalling Westmed ABG convenience kits containing TRIAD brand alcohol pads because TRIAD is recalling TRIAD brand alcohol pads due to potential contaminants. | Material/Component Contamination | Westmed sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 18, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were also given the TRIAD GROUP: "Urgent Drug Recall" and "Facts: Alcohol Recall" both dated January 3, 2011. Westmed customers were instructed to remove the TRIAD alcohol pad from the kit and dispose of the alcohol pad only. Customers were instructed to complete and return the enclosed response form as soon as possible. Customers with questions were instructed to contact the firm at (520) 264-7987 x1233. |  | 1259200 | 1,259,200 units total | Worldwide Distribution-USA (nationwide) and the countries of Canada, El Salvador, United Arab Emirates, Holland, Costa Rica and Lebanon. | Terminated | October 22, 2012 | 643 | 1259200
Z-0110-2013 | 63130 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron Systems Synchron Calibrator Multi Calibrator Diskettes, Part Number 442600. |  |  | December 15, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed a mislabeling error, some Synchron Systems Multi Calibrator Diskettes (REF 442600, lot M002642G) may actually be Thyroxine (T4) Reagent Calibrator Diskettes. | Labeling mix-ups | Beckman Coulter sent an Urgent Product Correction letter dated December 15, 2011, withan attached Fax Back Response Form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. If customers are currently using Multi Calibrator lot M002642, then they were instructed to verified diskette and discard all Multi Calibrator diskettes lot M002642 and M002642G without a dark green dot. Customers were instructed to complete and return the enclosed Response Form within 10 days. Customers were instructed to contact Beckman Coulter Support Center at http:www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and Canada. Outside of the US or Canada, customers were instructed to contact their local Beckman Coulter Representatives.For questions regarding this recall call 714-961-4483. |  | 5746 | 5746 units total (2729 units in US) | Worldwide distribution: USA (Nationwide) and Internationally to: Afghanistan Angola Australia Bangladesh Belgium Bhutan Brazil Brunei Darussalam Bulgaria Canada China Colombia Costa Rica Cote d'Ivoire Czech Republic Denmark Ecuador Egypt El Salvador Eritrea Ethiopia Finland France French Polynesia Germany Ghana Greece Guatemala Hong Kong India Indonesia Ireland Israel Italy Japan Jordan Kenya Kuwait Lebanon Lesotho Libyan Arab Jamahiriya Lithuania Macao Malawi Malaysia Mexico Morocco Namibia Netherlands New Zealand Nigeria Oman Pakistan Panama Paraguay Philippines Poland Puerto Rico Qatar Reunion Romania Russian Federation Rwanda Saudi Arabia Serbia Singapore Slovakia South Africa Spain Swaziland Sweden Switzerland Taiwan Thailand Tunisia Turkey Uganda United Arab Emirates United Kingdom Viet Nam Zambia Zimbabwe | Terminated | December 28, 2012 | 379 | 5746
Z-0112-2013 | 63215 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx GNR Traffic Light PNA FISH, In Vitro Diagnostic for Culture Identification Kit for Escherichia coli,and/or Klebsiella pneumoniae and/or Pseudomonas aeruginosa |  |  | September 14, 2012 | 2012 | Advandx Inc | GNR Traffic Light PNA Fish may have false positive results with Pseudomonas putida and Pseudomonas fulva | Device Design | AdvanDx sent a "GNR Traffic Light PNA FISH (KT011) Package Insert Replacement" letter dated September 14, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. All customers received an Acknowledgement Form to complete and return to AdvanDx so that the effectiveness of the replacement can be evaluated. Contact AdvanDx Technical Support at 1-866-376-0009 for questions regarding this notice. |  | 187 | 187 kits | Worldwide Distribution-USA (nationwide) and the countries of Austria, Denmark, Italy, Germany, Finland, and UK. | Terminated | March 05, 2013 | 172 | 187
Z-0117-2013 | 63353 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | HiRes 90K devices with Helix Electrodes Model CI 1400-02H |  |  | September 14, 2007 | 2007 | Advanced Bionics Corporation | Advanced Bionics is recalling the the HiRes 90K cochlear implant with HiFocus Helix Electrode because of loose-fitting insertion tool assemblies. | Nonconforming Material/Component | The firm, Advanced Bionics, sent a " RECALL NOTIFICATION" letter dated September 14, 2007 to all their customers who purchased the HiRes 90K devices with Helix Electrodes Model CI 1400-02H. The letter identified the product, problem, and actions to be taken. The customers were instructed to immediately locate and return the affected product to Advanced Bionics. The firm's Service representative will contact the customers within 48 hours to retrieve the affected product and send a replacement as needed. The customers were also instructed to complete and return the enclosed Notification Acknowledgement form via fax to: (661) 362-1511.If you have any questions regarding this letter, call Advanced Bionics at (877) 829-0026 between the hours of 5 AM and 5 PM Pacific Time, Monday through Friday. |  | 17 | 17 units | Worldwide distribution: USA (nationwide) including state: IL and country of: Europe. | Terminated | October 22, 2012 | 1865 | 17
Z-0111-2013 | 63254 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | EVICEL Fibrin Sealant (Human) Rx Only. |  |  | July 02, 2012 | 2012 | Ethicon, Inc. | Ethicon and Omrix Biopharmaceuticals have received reports of air or gas embolism occuring in association with the use of spray devices employing a pressure regulator to administer EVICEL; these events appear to be related to the use of the spray device at higher-than-recommended pressures and/or in close proximity to the tissue surface. The field correction initiated is to provide the recommend | Employee error | Ethicon sent a notification letter "Urgent Product Information EVICEL Fibrin Sealant (Human) All Lots and Product Codes" dated June 27, 2012, to all affected customers. Notifications were sent via UPS on July 2, 2012. The notification informed customers of the identified risks and included instructions for use. The notification included a business reply card consignees were requested to fill out and return.Customer were notified by a notification titled "Urgent Product Information Omrix Pressure Regulator All Product Codes" that was sent by Ethicon on August 1, 2012, as a follow up to the letter dated June 27, 2012. The notification informed consignees of the identified risks and included instructions for use. Customers were notified that a company representative would visit them to confirm receipt of the notification and to ensure the enclosed label has been affixed correctly on all Omrix Pressure Regulators.For further questions please call 1-877-384-4266. |  | 6300 | 4800 pressure regulators distributed worldwide; 1500 distributed within the US | Worldwide Distribution -- USA (nationwide) and the countries of Canada, European Union, Singapore, Israel, Mexico and Brazil. | Terminated | September 15, 2015 | 1170 | 6300
Z-0116-2013 | 63199 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx E. coli/P. aeruginosa PNA FISH, In Vitro Diagnostic Culture Identification Kit for Escherichia coli and Pseudomonas aeruginosa |  |  | September 14, 2012 | 2012 | Advandx Inc | Some strains of Pseudomonas putida cross-react with the P. aeruginosa PNA probe used in E. coli/P. aeruginosa PNA FISH (KT007) due to a sequence similarity resulting in a cross-reaction that may produce a red-positive signal | Process design | AdvanDx notified customers on September 14, 2012 by email along with a PDF file of the new package insert to each contact followed by written notification and a paper copy of the new package insert via FedEx courier. The notification described the product, problem and actions to be taken. The customers were instructed to destroy the package insert in any existing inventory of the kit and replace it with the enclosed package insert. "Note: All shipments of E. coli/P.aeruginosa PNA FISH as of September 14, 2012, will contain this new package insert." The customers were also instructed to complete and return the Acknowledgement of Receipt form to AdvanDX: Regulatory Affairs; 400 TradeCenter, Suite 6990, Woburn, MA 01801-7476; Fax to: +1 (781) 376-0111; bsc@advandx.com. If you have any questions or requests for further information, please contact AdvanDx Technical Support: US (Toll-Free): 1-866-376-0009, techsupport@advandx.com. |  | 33 | 33 units | Worldwide distribution: USA (nationwide) including: CA, FL, IL, MI, and NJ; and country of: Great Britain. | Terminated | March 05, 2013 | 172 | 33
Z-0119-2013 | 63313 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Osseotite Tapered Certain Implant 5 X 10mm Model INT510 |  |  | August 03, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i recalled their Osseotite Tapered Certain Implant, Model # INT510 due to a small percentage of the devices are missing the internal threading. | Process control | Biomet 3 i sent an Urgent Medical Device Recall letter dated August 3, 2012, to all affected customers via e-mail, facsimile, telephone, or postal mail. Each customer was instructed to check their inventory for the affected product and provide feedback to Biomet 3i. Product is to be returned to:Biomet 3iP.M. Regulatory Compliance4555 Riverside DrivePalm Beach Gardens, Florida 33410For questions customers were instructed to call 1-800-342-5454. For questions regarding this recall call 561-776-6700. |  | 227 | 227 | Worldwide Distribution - USA including PA, GA, OR, KS, NY, VA, CT, NH, IL, OK, and MA and internationally to Australia, Brazil, Columbia, France, Germany, Ireland, Italy, Japan, Paraguay, Portugal, and Spain. | Terminated | December 20, 2012 | 139 | 227
Z-0121-2013 | 63424 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Discovery NM/CT670; GE Discovery NM 630 |  |  | October 12, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue associated with thetable pallet of your Discovery NM/CT 670 and/or Discovery NM630 system which may impactoperation safety.At scan setup, during gantry rotation towards scanning position, a potential contact between the table-pallet-edge and the detector was identified. Such occurrence may damage the pallet front-edge and | Software design | GE Healthcare, LLC sent an " Urgent Medical Device Correction" letter dated October 12, 2012 to all affected customers. The letter was addressed to Managers of Nuclear Medicine, Hospital Administrators / Risk Mangers, and Mangers of Radiology / Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details and Product Correction. For questions or concerns contact your local GE Healthcare Service Representative. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other countries: 262-896-2890. |  | 210 | 210 | Worldwide distribution: USA (Nationwide) in the states of: AL, CA, FL, GA, IL, IN, KS, LA, MD, MA, MN, MS, MO, NJ, NY, NC, OH, OK, PA, SC, TN, TX, VA, WI and WV. AUSTRALIA, BELGIUM, CANADA, CHINA, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GUADELOUPE, HONG KONG, INDIA, IRELAND, ISRAEL,ITALY, JAPAN, REPUBLIC OF KOREA, NETHERLANDS,NORWAY, POLAND, RUSSIAN FEDERATION, SPAIN, SWITZERLAND, TAIWAN,THAILAND, TURKEY, UNITED KINGDOM. | Terminated | November 14, 2013 | 398 | 210
Z-0118-2013 | 63405 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncentra External Beam 4.1. |  |  | August 21, 2012 | 2012 | Elekta, Inc. | During the planning process the beam weights can be changed by the customers. When the beam weights are changed and renumbered and this happens during the dose calculation it is possible to save and reopen the plan without any error messages. | Software design | Nucletron sent a Customer Information Bulletin 555.00248 to all affected customers. The bulletin identified the product, problem, and actions to be taken by the customers. The bulletin contained information on how an incorrect dose association can be prevented by not renumbering beams during recalculation. A Read Confirmation Form was included for customers to complete and return. Contact your local Nucleton Service Representative for questions regarding this notice. |  | 52 | 52 units | Nationwide Distribution-including the states of AL, CA, DC, KS, LA, MO, NB, NY, NC, ND, OH, OR, PA, RI, SC, TX, and WA. | Terminated | November 03, 2014 | 804 | 52
Z-0120-2013 | 63296 | No | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | The brand name of the device is Damon 3MX Bicuspid Bracket, an orthodontic bracket. Part Number 494-4491, Description -Damon 3MX Bicuspid Bracket, Lot Number - 011265039 |  |  | April 24, 2012 | 2012 | Ormco/Sybronendo | The firm initiated the recall as a result of a misidentification of the brackets. Damon 3MX Bicuspid Brackets are produced with a purple colored ID dot to identify them as maxillary bicuspid brackets; however the affected lot of product listed below was produced with dark blue colored ID dots,which identified them as maxillary central incisor brackets. | Labeling mix-ups | Ormco/Sybronendo sent an Urgent Medical Device Recall letter via USPS 1~1 class mail, on April 24, 2012 to all consignees. Consignees were sent an updated customer notification via USPS lSI class mail on June 28, 2012. The letters identified the affected product, problem and actions to be taken. Consignees were instructed to complete the Acknowledgement/Return Form and to return any affected product in their inventory. Customers were instructed to contact Ormco Customer Care at 1-800-854-1741 to receive an RMA number. The RMA will allow for a quick return and replacement or credit. |  | 148 | 148 | Worldwide Distribution - US Nationwide and the countries of Canada, China, Costa Rica, Ecuador, French Polynesia, India, Kuwait, Malaysia, Philippines, Singapore, South Korea, Thailand and UAE. | Terminated | December 12, 2012 | 232 | 148
Z-0125-2013 | 61612 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage 2.2 workstations with software version 22_01 and 22_02 |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of inaccurate Scale factor annotation provided by SenoAdvantage 2.2 workstations that may impact patient safety. This issue may result in a potential miscalculation of size. | Software design | Consignees were sent on 5/24/12 a GE Healthcare "Urgent Medical Device Correction" letter dated May 23, 2012. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, and Product Correction. Customers should verify if their workstation is affected. GE Healthcare will correct all affected workstations by providing a software upgrade and Operator Manual update. A GE Healthcare service representative will contact you to arrange this correction. This activity will be performed at no cost to you. If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. |  | 12 | 12 | Worldwide Distribution -- USA (nationwide) including states of: CA, CT, NY, TX, and MS and countries of TAIWAN, ITALY, CANADA, and BAHRAIN | Terminated | December 11, 2012 | 201 | 12
Z-0123-2013 | 63143 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCelÂż DxC 600/800 with |  |  | September 03, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter has confirmed that the Shuttle to Barcode Reader and Cap Piercer alignmentprocedure cannot be performed on certain DxC configurations with v5.0.11 software installed. If the Shuttle to Barcode Reader becomes misaligned, the Operator will not be able to perform thisalignment procedure on the DxC systems listed above. Software v5.0 .11 does not allow access to theShuttle to Bar | Software change control | Beckman Coulter sent an Urgent Field Safety Notice dated August 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that their Beckman Coulter Service Representative would contac them to reinstall the software. Customers were asked to complete and return the enclosed Response Form. For questions customers were instructed to call their Customers Support Center on the Beckman Coulter website at www.BeckmanCoulter.com/CustomerSupport/SupportRequest.For questions regarding this recall call 714-993-5321. |  | 7 | 7 units total (both US and International) | Worldwide Distribution - USA including TX and OH. Interntionally to China and Turkey | Terminated | February 14, 2014 | 529 | 7
Z-0122-2013 | 63081 | No | ONN | Stereotaxic instrument. | Neurology | Neurology | 510(k) | "****ORTHO SENSOR**12-123-A Knee Trial for Trithion***1560 Saw grass Corporate Pkwy***4Th floor***Sunrise, FL 33323***www.orthosensor.com***" |  |  | February 04, 2012 | 2012 | Orthosensor | Ortho sensor in Sunrise, FL is recalling the Graphics User Interface Software 1.8 (which was updated to version 1.82, validated on 02/13/12). Recall was initiated due to customer feedback which included the line of site between the knee ba lancer and the Link station was obstructed by OR metal stands or trays or the knee balancer device was placed atop of or near the metal stands or trays, interru | Software design | Orthosensor did not provide a written customer notification. As the recall was reported to the FDA well after all nine of the products were corrected (the recall stemmed from an initial inspection of the firm), then no customer notification was requested. The removal and replacement of the previous version of the Graphic User Interface Software is being handled internally by members of Operations and the Sales force. The items the firm is recalling is software, and the software cannot be returned; the old version (version 1.8) would be erased. The new version 1.82 would be installed and downloaded into the equipment at the site during servicing. This does not involve the customer of which none are consignees. Effectiveness checks associated with the update of the software is being handled internally by the Operations and the Sales force.Further questions please call (813) 645-2855. |  | 9 | 9 | Nationwide Distribution including the states of CA, FL, MI and NY. | Terminated | December 12, 2012 | 312 | 9
Z-0131-2013 | 62315 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Perfusion System 8000, base 5 pump, 100V safety monitor included |  |  | September 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received multiple reports since 2007 of various malfunctions of the power distribution and battery back-up systems for the Sarns Modular Perfusion System 8000. Very few of the reports occurred during cardiopulmonary bypass; the majority of the malfunctions occurred during set-up or were reported during routine preventative maintenance. The malfunction | Labeling design | TERUMO sent an URGENT MEDICAL DEVICE letter dated September 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed review the Medical Device Recall notice, assure that all users are aware of the notice and place the addendum in the Operator's Manual before page 1.13. Customers were also instructed to confirm receipt of the Urgent Medical Device letter by faxing, the attached Customer Response Form to 1-800-292-6551. For questions customers were instructed to call 1-800-521-2818. |  | 81 | 81 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV, and WY and the countries of ARGENTINA, AUSTRALIA, BANGLADESH, BELGIUM, Bolivia, BRAZIL, CALLO PERU, CANADA, CAROLINA PR, CHILE, China, COLOMBIA, Costa Rica, DEKALB, DENMARK, Dominican Republic, Ecuador, Egypt, FRANCE, Germany, GUAYQUIL ECUADOR, Hong Kong, India, Indonesia, ITALY, Japan, , JARAMA MADRID, Lebanon, LEICESTERSHIRE, MADRID SPAIN, MADRID SPAIN, Malaysia, MEDLEY, Mexico, D.F., MIAMI, MUMBAI 26 INDIA, MUMBAI, New Zealand, Nicaragua, Pakistan, Panama, Philippines, REPUBLICA, SANTIAGO, Saudi Arabia, SELANGOR, MALAYSA, SEOUL, KOREA, Singapore, SOUTH AFRICA, South Korea, Sri Lanka, SWITZERLAND, TAIPEI, TAIWAN, Thailand, TUJUNGA, UNITED ARAB EMIRATES (UAE), URUGUAY, Venezuela, Vietnam, VIGNATE ITALY, WOODDALE, and Yemen. | Terminated | August 21, 2015 | 1060 | 2256
Z-0132-2013 | 63316 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster MobiCath Bi-Directional Guiding Sheath. Small Curve catalog no. D140010, Large Curve catalog no. D140011. The MobiCath Bi-Directional Guiding Sheath is a flexible tipped percutaneous catheter introducer (or guiding sheath) designed for gaining access to the coronary systems. The device features adjustable tip geometry through use of a rotating handle to deflect the catheter. This catheter is assembled with a hemostasis valve and a sideport infusion line. |  |  | September 24, 2012 | 2012 | Greatbatch Medical | We are notifying you of a potential interaction between therapeutic catheters and MobiCath Bi-Directional Guiding Sheaths. We became aware of these interactions while investigating recent complaints regarding catheters "self-advancing" while being used in conjunction with the MobiCath Sheaths. | Device Design | The firm, Greatbatch, Ltd., sent an "URGENT CUSTOMER NOTIFICATION Medical Device - Voluntary Field Safety Notice" dated Monday, September 24, 2012 to its customer. The letter described the product, problem and actions to be taken. The customer was instructed to read the "Description of the Problem" and "Recommendations for Clinical Use" sections carefully; complete and return the attached Field Action Acknowledgement Form via mail to: Biosense Webster, Incl, 15715 Arrow Highway, Irwindale, CA 91706; Attn: Recall Coordinator, or Fax to: (909) 839-8514; pass on this notice to anyone in your facility that needs to be informed; maintain awareness of this communication until information has been incorporated into the device labeling, and to avoid the catheter inside the MobiCath Bi-Directional Guiding Sheath from "self-advancing", physicians should reposition the catheter handle when manipulating the catheter in and out of the sheath. Greatbatch, Ltd. is updating the MobiCath Bi-Directional Guiding Sheath Instructions For Use (IFU) accordingly. Supplemental information will be provided with future shipments of these devices, until the revised IFU is available.For questions related to the issue please contact your Biosense Webster representative or call (800) 729-9010, Monday through Friday from 9:00 AM to 6:00 PM EST. For questions related to the Acknowledgment Form and its return, please contact Field Action Coordinator at (909) 839-8468. |  | 9040 | 9040 | Nationwide distribution: CA only. | Terminated | May 15, 2014 | 598 | 9040
Z-0133-2013 | 63283 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DePuy Mitek VAPR Hook Electrode 3.5mm Hook Electrode with Integrated Handpiece |  |  | March 12, 2012 | 2012 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Product may have wrong large product identification label with symbols on the box indicating a S90 VAPR Electrode instead of the VAPR 3.5 hook electrode | Error in labeling | The firm, Depuy Mitek, sent a "Medical Device Labeling Alert" letter dated March 14, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect your inventory to determine if any inventory of product code 227305 lot with the number M1108086 that is incorrectly labeled; review the following label information and the attached examples of correct and incorrect labels; and complete and return the Business Reply Form via fax to: 1-508-828-3750. To return the product call 1-877-379-4871 for an RGA and replacement product and return to DePuy Mitek, Attn: Recall Coordinator, 50 Scotland Boulevard, Bridgewater, MA 02324.Contact your local DePuy Mitek Representative or DePuy Mitek Customer Service directly at 1-877-379-48 |  | 376 | 376 units | Nationwide distribution: USA including states of: AL, CA, CT, FL, GA, IA, ID, IL, IN, KY, MD, MI, MN, NJ, NY, OH, OR, TN, TX, UT, VA, and WA. | Terminated | June 05, 2015 | 1180 | 376
Z-0137-2013 | 63302 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | AquaLiner Hydrophilic Guidewire 0.035" X 260cm, Straight/ Stiff. |  |  | July 27, 2012 | 2012 | Nipro Medical Corporation | Nipro Medical Corporation, Miami, FL recalled their AquaLiner Hydrophilic Guidewire because the wires contain insufficient lubricity which causes them to get stuck in the housing and become nonfunctional. | Equipment maintenance | Nipro Medical Corp. sent notifications to their consignee via e-mail and fax on July 27, 2012 and September 13, 2012 to their affected customer. The letters identified the product, problem, and actions to be taken by the customer. The letter instructs the customer to place all affected product in stock on hold. A Aqualiner-Recall Response form was attached for the customer to complete and return. Contact the firm at 305-599-7174 ext. 249 for questions regarding this notice. |  | 300 | 300 | Worldwide Distribution-Distributed in the state of New York and the countries of Spain, Germany, Argentina, Panama, Mexico, Chile, Colombia, Ecuador, Peru, and Brazil. | Terminated | May 25, 2013 | 302 | 1500
Z-0155-2013 | 63121 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | SpiraLok 6.5 w/out Needles with violet and blue Orthocord |  |  | September 06, 2012 | 2012 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Fracture of SPIRALOK Anchors post op, requiring patient revision | Device Design | DePuy Mitek sent an Urgent Voluntary Product Recall letter dated September 6, 2012, to all affected customers. Dr Doctor letters were included with the recall letter. Recall is coordinated with Stericycle , contact 1-888-202-3694 The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the enclosed Business Reply Form, even if they have no product remaining in their inventory. Customers were instructed to isolate all inventory of the affected product and return to Stericycle at the following address or fax to 1-888-912-2189 with the completed Business Reply Form:Stericycle2670 Executive Drive, Suite AIndianapolis, IN 46241Attn: Event #8874Customers with questions were instructed to call 1-888-202-3694.For questions regarding this recall call 508-880-8100. |  | 4321 | 4,321 units | Worldwide Distribution - USA (nationwide) and the countries of AE, AR, AT, AU, BE, BR, CA,, CH,CL .CN, CO, CZ,DE, DK, EC, ES, FI, FRGB, GP, GR, IE, IL, IN, IT, KR, LU, LY, MX, NC NL, NO, NZ, PE, PLPT, RE, RU, SG, SI, TH, TN, TR, UG ,UY , VE, and ZA. | Terminated | June 03, 2015 | 1000 | 78594
Z-0158-2013 | 63361 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Gish 3/8 x 1/2" Reducer Connector with Luer fitting, Item Numbers: EC2180S, Lot Numbers: 1108900099, 1119900091, 1203700129, 1208600013, 1219200107, Exp. Dates: 3/31/14, 7/31/14, 2/28/15, 3/31/15, 7/31/15. |  |  | September 14, 2012 | 2012 | Sorin Group USA, Inc. | Due to a molding irregularity, certain lots of connectors have the potential to leak fluid or entrain air at the interface of the luer fitting and mating cap. Entrained air could result in patient injury. | Process control | Sorin Group USA, Inc. sent a Product Recall letter to all affected customers by certified mail on September 14, 2012. The letter identified the affected products, problem and actions to be taken. Customers were instructed to discontinue use of the affected product, return all unused product to the firm and to complete and return a Response Form via fax or email. For further questions, contact the firm at 1-800-650-2623. |  | 1380 | 1,380 units | US Nationwide Distribution in the states of: AL, AR, CA, CO, CT, FL, IL, MI, MN, MO, MS, NC, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA and WI. | Terminated | April 23, 2013 | 221 | 3020
Z-0156-2013 | 63379 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Repeater Pump, Class II device used for facilitating repeatable drug dosage distribution and reconstitution in hospital pharmacies. The device are Fluid Transfer Tube Sets, Sterile/EO, Std Volume Trifurcated, BAXA Corporation, Englewood, CO. |  |  | October 03, 2012 | 2012 | Baxa Corporation | The firm is issuing a field correction to the operators manual included with the 099 Repeater Pump because the Technical Manual was distributed with the device as opposed to the Operator Manual and the instruction for use contradict each other with regards to multiple source container use. | Packaging process control | Baxa Corporation sent a "REPEATER PUMP PRODUCT NOTICE" to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. All customers received a CD containing a copy of the correct Operator Mannual, 5300-0754. Contact Technical Support at 800-678-2292 for questions regarding this notice. |  |  |  | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Brazil, Canada, Columbia, Guam, Japan, Malaysia, Mexico, Panama, Puerto Rico, Taiwan, UK, and Venezuela. | Terminated | February 15, 2013 | 135
Z-0139-2013 | 63321 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 6.5mm Cancellous Bone Screw |  |  | September 28, 2012 | 2012 | Encore Medical, Lp | A product complaint was received where a size 20 screw was labeled as size 30. After receipt, measurement of the screw determined it to be a size 25. Further investigation also determined that there were other screws labeled as size 25, but a size 30 screw was inside of the package. It was thought that all of these screws were found in-house and were contained, but it was determined later that on | Labeling mix-ups | Encore Medical sent a Urgent Field Safety Notice dated September 28, 2012, to all affected customers.This notice needs to be passed on all those who need to be aware withinyour organization or to any organization where the potentially affecteddevices have been transferred.Please contact Customer Service to place a replacement order and receive an RMA number. Return all affected devices using the RMA number.For further questions please call (512) 834-6255 |  | 1 | 1 unit | Nationwide Distribution only to CA. | Terminated | January 25, 2013 | 119 | 1
Z-0160-2013 | 62947 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | QUADROX iD HMOD 2030 Diffusion Membrane Oxygenator; QUADROX-i ADULT HMO 70000 / HMO 71000; QUADROX-iD Pediatric HMOD 30000 |  |  | June 12, 2012 | 2012 | Maquet Cardiovascular Us Sales, Llc | It has come to firm's attention that, in some rare cases, the blood outlet connector on some models of MAQUET Oxygenators may detach from the oxygenator. This event has occurred in about 0.01% (1:10,000) of units shipped. | Nonconforming Material/Component | Maquet sent a Field Safety Notices and Confirmation Forms dated June 13, 2012 to all affected customers. A Follow-up communication dated July 20, 2012 was mailed to all affected US customers. The Field Safety Notice identifies the problem, affected product, risk factors and actions to be taken by the customer. Customers were instructed to complete the Confirmation form and return in the self-addressed prepaid envelop. For questions call 201-995-8968.Maquet issued an updated Field Safety Notice/Field Safety Notice Confirmation Form dated 2/14/13 via Fed Ex on 2/18/2013.Maquet sent out an Updated Field Safety Notice/field Safety Notice Confirmation Form dated 4/4/2013 in order to notify customers they would be providing safety clamps. |  | 23203 | 23,203 units | US Nationwide Distribution | Open, Classified |  |  | 23203
Z-0159-2013 | 62386 | No | GDI | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Progressive Medical PMI Endoscopic Dissector with Kittner Tip (5mm x 40cm), Reorder Number: PMILK40, Product Sterile in Unopened/Undamaged pouch, Cotton Tip - X-Ray Detectable, 1 Pc/Pouch, 10 Pouches/Box, 20 boxes/case. |  |  | October 28, 2011 | 2011 | Progressive Medical Inc | The tip of the product could come loose or unravel. | Packaging | The recalling firm sent a letter dated October 28, 2011 to the customer who reported the problem and calls were initiated 6/26/12 to the remainder of the customers. The reason for recall was explained and the customer were asked if any of the lot remained in their inventory. If inventory remained, the recalling firm sent them a return goods authorization number and informed them the product would be replaced. For questions call 1-800-969-6331. |  | 124 | 124 boxes | US Nationwide Distribution in the states of: CO, IL, IN, KS, LA, MA, MN, MO, NE, OH, OK, TN, TX, VA, and WI. | Terminated | October 26, 2012 | 364 | 124
Z-0124-2013 | 63139 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartSine Samaritan(R) 300/300P PAD devices with serial numbers 0400000501 to 0700032917; 08A00035000 to 10A0070753; 10C00200000 to 10C00210106 |  |  | September 13, 2012 | 2012 | Heartsine Technologies, Limited | Certain Samaritan 300/300P PAD devices may experience one or both of the following conditions that could affect the ability to deliver therapy to a patient in an SCA event if needed: Issue 1 (On/Off Issue): the device may turn itself on and off without input from the user. When this occures, the normal sequence of audible instruction prompts that occurs when the device is turned on will be emit | Device Design | HeartSine sent an Urgent: Medical Device Recall dated September 2012 to all affected customers. The letter identifed the affected product, problem and actions to be taken. A response form was included with the letter to be completed by customers and returned to the recalling firm's representative. In addition, HeartSine is providing a software upgrade to bring all users up to a more recent version of the software. For questions call 1-877-877-0147 or your service representative.The company issued anadditional press release on September 13, 2013 in a further attempt to reach out to end customers thathad not yet responded to the companys efforts. |  | 81982 | 81,982 in total | Worldwide Distribution and US (nationwide) in the following states: AK, AL, AP, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, WA, and WI. | Terminated | July 10, 2015 | 1030 | 81982
Z-0162-2013 | 63167 | No | MOI | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Siemens ADVIA Centaur BR Assay: Catalog Number: |  |  | September 11, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | BR Assay for CA 27.29 Calibration Interval Change due quality control (QC) results outside of the expected ranges. | Under Investigation by firm | Siemens sent two "URGENT FIELD SAFETY NOTICE" letters dated August and September 2012 to all affected customers. The letters identified the product, problem, and actions to be taken by the customers. A Field Correction Effectiveness Check Form was included for customers to complete and return. |  | 15455 | 15,455 kits | Worldwide Distribution-USA (nationwide) and the countries of Australia, Brazil, Canada, Columbia, Japan, Korea, Mexico, New Zealand, Peru, Saudi Arabia, and Spain. | Open, Classified |  |  | 15455
Z-0165-2013 | 63112 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Myocardial Heart Wires |  |  | July 25, 2012 | 2012 | Oscor, Inc. | Oscor, Inc. in Palm Harbor, FL is recalling myocardial heart wires, models TME60S, TME64S, TME64C, TME64S-3, TME65S, and TME66S. The recall was initiated due to the spacing between anchoring zig-zag to electrode is out of specification. | Process control | Oscor sent a customer notification letter dated July 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to send the listed models and lots in their inventory back to the recalling firm for immediate replacement. They were also instructed to call Oscor Customer Service at 727-937-2511 to obtain a Return Goods Product (RGA) number.For questions regarding this recall call 727-937-2511, ext. 133. |  | 1425 | 1425 | Worldwide Distribution - USA including OR, FL, Al, TX, NM, CA, AL, PA, KY, IN, OH, and IL and internationally to Canada, Russia, and New Zealand | Terminated | December 17, 2013 | 510 | 1425
Z-0164-2013 | 63090 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***All in one CrossFT***Punch***PFT-00M***4.5mm to 6.5mm***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA***" |  |  | April 09, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec is recalling the product due to the punch being mis-assembled with the incorrect handled. The correct handle is a clear anodized silver colored handle. | Process control | Linvatec Corp sent a Urgent Medical Device Recall Notification letter dated April 9, 2012, to all affected customers. The letter was sent via FEDEx and by e-mail. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to check their inventory for the recalled product and segregate them for return to the recalling firm. Furthermore, return the products along with the "REPLY FORM".Should you have any further questions please contact our Customer Service department at (800) 237-0169. |  | 2 | 2 units | Worldwide Distribution--US (nationwide) including state of Minnesota and the country of Thailand | Terminated | November 21, 2013 | 591 | 2
Z-0168-2013 | 63385 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | ILLlCOÂż MIS Posterior Fixation System, Bone-screw Internal Spinal Fixation System, |  |  | September 25, 2012 | 2012 | Alphatec Spine, Inc. | The firm initiated the recall because the stainless steel guide-wire of theIllico MIS Fixation System did not completely fit/pass through the tip of certain Illico TiCannulated Polyaxial Screws, therefore rendering such screws nonfunctional. | Nonconforming Material/Component | Alphatec Spine sent an Urgent Medical Device Recall Notification letter dated September 25, 2012, to all affected customers. The letter identifed the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Alphatec Spine's Customers Service Department (customerservice@alphatecspine.com) immediately for instructions on how to return the affected product. Customers were asked to fill out the lower portion of the letter and fax to (760) 431-0289 to confirm that they have read the notification and will take all necessary removal actions. For questions customers were instructed to call 1-800-922-1356. For questions regarding this recall call 760-431-9286. |  | 209 | 209 screws | Nationwide Distribution including AZ, IL, CA, SC, TN, TX, GA, CT, NY, FL, OK, LA, MO, NV,and UT. | Terminated | March 29, 2013 | 185 | 209
Z-0166-2013 | 63381 | Yes | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | SCP Pump Control Panel, Product Code: 60-02-15, Serial Numbers: 60S09121, 60S09123-60S09129, 60S09136-60S09139. |  |  | September 20, 2012 | 2012 | Sorin Group USA, Inc. | Sorin Group USA, Inc. has issued an Urgent Field Safety Notice regarding certain serial numbers of the Stockert Centrifugal Pump (SCP) Pump Control Panel . An out of specification circuit board component within the Panel may lead to a malfunction consisting of a blank or flashing display. The flow and pressure display information may also be incorrect after the malfunction. | Mixed-up of materials/components | The firm, Sorin Group USA, Inc., sent an "URGENT FIELD SAFETY NOTICE" dated September 20, 2012 to all affected customers by certified mail. The notice described the product, problem and actions to be taken. Customers were instructed to discontinue use of and segregate affected product for return and replacement to the firm; to complete and return the attached Customer Response Form via fax to: 303-467-6502 or email to: yvonne.feyerherm@sorin.com; and if they have transferred the affected products to a third party, to pass this information to the Sorin Group.Further questions may be directed to Sorin Group Customer Service at 1-800-650-2623. |  | 12 | 12 units (3 units USA and 9 units OUS) | Worldwide distribution: USA (nationwide) including states of : MI and NY; and countries of: Belgium, Canada, Japan, India, Russia, United Arab Emirates. | Terminated | January 28, 2013 | 130 | 20
Z-0163-2013 | 63272 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Aviator Drill Bit (10 mm, 12 mm, 14 mm, 16 mm, |  |  | July 05, 2012 | 2012 | Stryker Spine | Aviator drill bits are 10 mm longer than the Reflex Hybrid drill bits and are not interchangeable which could lead to over-drilling by as much as 10 mm. | Device Design | Stryker Spine sent an Urgent Medical Device Field Correction Notification letter beginning July 5, 2012, to all affected customers. Teh letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check the drill bit length in each respective drill guide prior to surgery to ensure the correct length protrudes through the guide. Customers with questions should call 201-760-8298. Customers were asked to report any adverse events or product quality problems to 1-888-457-7463. |  | 6154 | 6154 units | Nationwide distribution. | Terminated | January 28, 2013 | 207 | 6154
Z-0169-2013 | 63335 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ |  |  | September 18, 2012 | 2012 | Elekta, Inc. | A patient was seriously mistreated after a user made and saved an inadvertent change to the definition of a treatment filed in MOSAIQ. | Device Design | The firm, MOSAIQ, sent an "IMPORTANT SAFETY NOTICE- LINMSQ0003" was sent to consignees/customers on September 18, 2012. The notice described the product, problem and actions to be taken. The customers were instructed to check their Department Setup configuration settings, and re-enable the Field Edit Delta Display feature if they have disabled it. Note: sites can take this action without an upgrade/patch. The customers were also instructed to complete and return the CONFIRMATION OF RECEIPT form via fax to: 702-992-5002 in US or +44 (0)1293 654401 in OUS, Attn: Elekta OIS Support; scan and email to: support@impac.com or europe.support@impac.com; Attn: Elekta OIS Support and/or mail to: C/O Elekta OIS Support, 2310 Corporate Circle, Suite 275, Henderson, NV, USA; or C/O Elekta OIS Support, Elekta Limited, Linac House, Fleming Way, Crawley, West Sussex, RH10 9RR, UK.(Note: Product upgrade/patch is not mandatory or even necessary to reduce the related risk.)For questions, concerns, and requests for upgrade, email: support@impac.com (North America and Rest of World) or europe.support@impac.com (Europe) or call Vice President of RA/QA RNA at 770-670-2548. |  | 300 | 300 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | June 06, 2013 | 261 | 300
Z-0161-2013 | 60030 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Medpro Elastomeric Infusion Pump. |  |  | July 09, 2011 | 2011 | First Medical Source LLC | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being reclassified by the Agency as a Class I. The recall was initiated because First Medical Source has confirmed that these lots may have a higher flow rate than specified. The use of this product may lead to over-administration of drug solutions to the patients. The pr | Nonconforming Material/Component | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being reclassified by the Agency as a Class I. First Medical Source received an email from Medpro International with the following lnstruction to Customer - Informed customers to remove product from sales/discontinue use, cease distribution, remove the products from inventory and quarantine, inform their customers to retum the unused products to them, and destroy the product. Distributors' customer service will contact the customer to coordinate the return of the goods in question where required.First Medical Source emailed to their customer, Williams Medical Technologies, to informed them of the recall of the Medpro Elastomeric infusion pumps. Customer was instructed to inform them of the date and method of disposal of recalled products. |  | 500 | 500 units | US Nationwide distribution in the state of IL | Terminated | September 22, 2014 | 1171 | 500
Z-0182-2013 | 63387 | No | NOP | West Nile virus serological reagents. | Microbiology | Microbiology | 510(k) | West Nile Virus IgG DxSelect" kit Catalog No. EL0300G |  |  | October 03, 2012 | 2012 | Focus Diagnostics Inc | The firm recalled due to higher reactivity with some samples in the effected lots, which may reflect ahigher false positive rate. | Nonconforming Material/Component | Focus Diagnostics Inc., sent a Urgent Medical Device Recall letter dated October 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you received at least one of the effected lots in the month of August 2012. Please document the number of kits in your inventory on the table below, even if zero (0), and fax to 562-240-6526.For further questions please call (714) 220-1900. |  | 425 | 425 units | Worldwide Distribution--US (nationwide) and the country of: Canada, Deuche, France, Greece, Mexico, Israel, and Japan. | Terminated | December 04, 2012 | 62 | 425
Z-0184-2013 | 63338 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Knee Components: |  |  | September 05, 2012 | 2012 | Wright Medical Technology Inc | There may be inaccuracies on the labeled shelf life on all products packaged at the Toulon, France facility. | Packaging process control | Wright sent an Urgent Product Recall Notification letter dated September 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for affected product and immediately return to Wright Medical Technology, Inc. Customers were asked to complete and return the enclosed notification form by fax to 901-867-7401. For questions customers should call 901-867-4324.For questions regarding this recall call 901-867-4655. |  | 729 | 729 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Austria, Belgium, Bulgaria, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Italy, Latvia, Lithuania, Norway, Poland, Portugal, Serbia, Slovakia, Slovenia, Spain, The Netherlands, Argentina, Brazil, Chile, China, Japan, Taiwan, Turkey, United Arab Emirates, Australia, South Africa, Russia, and United Kingdom. | Terminated | January 07, 2014 | 489 | 33135
Z-0175-2013 | 63347 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | AirFlow Resuscitator, Adult Resuscitator, AF1000 Series, Single use, Manufactured by: Ventlab Corporation, 155 Boyce Drive, Mockville, NC 27028 |  |  | July 11, 2012 | 2012 | Ventlab Corporation | Possible volume leakage through the inlet valve during compressions of the resuscitator bag. | Component design/selection | VentLab Corporation sent an Attention: Recall Notification letter dated July 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product, complete and return the enclosed response form by fax to 336-753-5002, indicating whether or not their facility has any of the affected product. Customers can also scan and email the response form to CSR@Ventlab.com. For questions regarding this recall call 336-753-5000. |  | 14602 | 14,602 total for all units | Nationwide Distribution including IL, OH, NY and TN. | Terminated | June 30, 2013 | 354 | 14602
Z-0186-2013 | 63104 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Software for Vitrea, Vitrea fX, and Vitrea Enterprise Suite (VES): |  |  | August 13, 2012 | 2012 | Vital Images, Inc. | Vital Images has found two potential errors in calcium score values in restored snapshots of studies generated through use of Vitrea or Vitrea Enterprise Suite. This issue affects Version 3.1 of Vitrea fX; Version 5.2 of Vitrea; Versions 6.0, 6.1 and 6.2 of Vitrea and Vitrea fX; and Versions 1.3, 6.0, 6.1 and 6.2 of Vitrea Enterprise Suite; and all updates to those versions. | Software design | Vital Images sent an "URGENT DEVICE CORRECTION" letter dated August 13, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter gave advise on how to resolve the two issues and advised consignees to immediately inform all potential users. Contact Customer Support at 1-800-208-3005 for questions regarding the notice. |  | 2088 | 2088 (1133 USA) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, and the countries of ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BRAZIL, BULGARIA, CANADA CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, FINLAND, FRANCE, GREAT BRITAIN, GEORGIA, GERMANY, GHANA, GREECE, GUAM, HONDURAS, HONG KONG, INDIA, INDONESIA, IRAQ, IRELAND, ITALY, IVORY COAST, JAPAN, JORDAN KAZAKHSTAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MALI, MOSCOW, MYANMAR, NAMIBIA, NETHERLANDS, NEW ZEALAND, NIGERIA, NORTHERN IRELAND, NORWAY, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA, SERBIA AND MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE NETHERLANDS, TIBET, TURKEY, UKRAINE, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, AND YEMEN | Terminated | March 25, 2015 | 954 | 2088
Z-0187-2013 | 63471 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA Radiation Oncology, Version 11. Varian Medical Systems, Palo Alto, CA. The ARIA Radiation Oncology product is a treatment plan and image management application. |  |  | September 24, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the ARIA Oncology Information System Version 11 where Monitor Units of one fraction may be incorrect, when partial treatment is being continued. | Software design | "Urgent Medical Device Correction/ Urgent Field Safety Notice" letters were sent on September 24, 2012 by Certified Mail with return receipt to all affected users. The problem was described and recommended user actions were provided. The notification letters were to also be distributed to Varian Sales, Marketing, and Service organizations. US customers may contact the Varian help desk at 1-(888) 827-4265. |  | 205 | 205 | Worldwide Distribution - Nationwide(USA) and including the countries of Australia, Denmark and the United Kingdom. | Terminated | December 05, 2013 | 437 | 205
Z-0188-2013 | 63310 | No | OWD |  |  | Molecular Genetics | PMA | cobas 4800 BRAF V600 Mutation Test for In Vitro Diagnostic Use Roche Molecular Systems, Inc., |  |  | July 31, 2012 | 2012 | Roche Molecular Systems, Inc. | It was identified that BRAF controls dispositioned as "Internal Use Only" were incorrectly packaged into cobas 4800 BRAF V600 Mutation Test CE-IVD kit lot R03607 and released for distribution. Two control batches manufactured during two process validation studies are involved:PV7711: BRAF Mutant (MUT) ControlPV0853: BRAF Wild Type (WT) Control | Mixed-up of materials/components | Roche sent a Safety Board Notice - Product Bulletin 2012-013 on 31-Jul-2012 to affected customers. The notice identifed the affected product, description of situation and actios to be taken. Customers were instructed to discontinue use, quarantine and discard any remaining kits of the affected product in inventory. Customers were advised to contact their affiliate sales office for replacement. The notice states a new cobas 4800 BRAF V600 Mutation Test lot (R07749) is anticipated to be available on 03-Aug-2012. For questions call +01.908.253.7569. |  | 44 | 44 kits | Worldwide Distribution - Australia, Austria, Belgium, China, Hungary, Turkey and United Kingdom. | Terminated | September 11, 2013 | 407 | 44
Z-0190-2013 | 63096 | No | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | Boule Con-Diff Hematology Control |  |  | March 24, 2011 | 2011 | Clinical Diagnostic Solutions | Clinical Diagnostic Solutions, Inc. in Plantation, FL recalled Boule Con Diff US Tri Pack (Product Part 501-605, Lot #1102-575 and 1102-576; Product Part 501-607, Lot #1102-576; and Product Part 502-012, Lot #1102-574) which is used for in vitro diagnostic use as a control to monitor the performance of multi-parameter hematology instruments. The recall was initiated due to the receipt of a complai | Nonconforming Material/Component | Clinical Diagnostic Solutions, Inc. sent a letter dated March 24, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The firm istructed the accounts to replace the assay sheet they were currently using with the new assay sheets enclosed. For an unused product they would need to discard. To correct the ranges in their instrument, they would need to scan in the new bar codes specific for their instrument type. For any questions customers were instructed to contact CDE Technical Support atl 800-453-3328.For questions regarding this recall call 954-791-1773. |  | 901 | 901 vials | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MS, ND, ND, NE, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WI, WV,and WY. Foreign: Brazil and Colombia | Terminated | August 13, 2014 | 1238 | 1721
Z-0192-2013 | 63309 | No | JJT | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cliniqa Liquid QCTM Cardiac Marker Control for i-STAT Level 1 (List Number 06F12-09; Lot Number 071039) |  |  | July 19, 2012 | 2012 | Abbott Point Of Care Inc. | Abbott Point of Care (APOC) has determined that for the Cliniqa Liquid QC Cardiac Marker Control for i-STAT Level 1 (List #06F12-09; Lot Number 071039); i-STAT Level 2 (List #06F12-10; Lot Number 081039); and i-STAT Level 3 (List #06F12-11; Lot Number 091039) results may be generated below the lower value assignment range for some vials. | Material/Component Contamination | The firm, Abbott Point of Care, Inc. (APOC), sent an "URGENT RECALL NOTICE" dated July 2012 via FED-EX on July 19, 2012. The recall notice identified the product, problem and the action needed to be taken by the customer. The customers were instructed to discontinue use of the control lot numbers; return all unused controls on the enclosed Customer Reply Form; If another institution was provided the controls, provide a copy of this notice to them. Note: if any used vials gave acceptable product performance, a review of the cartridge performance and associated patient results is not necessary)Note: APOC is shipping replacement products to customers facilities to minimize any disruption this might cause.If you have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020, option 1 or your Abbott Point of Care representative. |  | 154 | 154 (06F12-09 (74) boxes; 06F12-10 (35) boxes and 06F12-11 (45) boxes.) 1 box contains six, 3ml plastic bottles. | Worldwide distribution: USA (nationwide) including states of: AZ, GA, ID, IL, LA, MD, MO, MN, NC, NV, OK, PA, TN, TX, VA and WI; and country of: Italy. | Terminated | June 20, 2013 | 336 | 154
Z-0206-2013 | 63430 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 VCC P/N 4598 001 41341 REV A, Philips Medical Systems (Cleveland), Inc., Fitchburg, WI. |  |  | October 05, 2012 | 2012 | Philips Medical Systems (Cleveland), Inc. | Philips Medical Systems Cleveland), Inc. is recalling Pinnacle3 Virtual Client Connection software version 3.1 due to an improper release and subsequent distribution. | Process control | The firm, Philips Healthcare, sent a "Customer Information Letter" dated 2012 Oct 05. The letter listed the affected products, problem description, how to identify affected products, advise on actions by customer / user, actions planned by Philips and further information and support. The letter stated that no action is needed on the part of the customer. The field service engineer will be scheduling a site visit to complete the replacement. Philips Healthcare will be delivering the correct VCC kit.For customers in North America if you need any further information or support concerning this issue, please contact our Customer Care Solutions Center at 1-800-722-9377. Select option 5 for "All Imaging Systems". Enter your site ID # (If you do not have a site ID#, simply pause for a moment.). Select option 5 for "Nuclear Medicine" and finally select option 3 for "Pinnacle" support. |  | 11 | 11 | Nationwide distribution: USA including states of: ME, MI, NY, and WI. | Terminated | February 21, 2013 | 139 | 11
Z-0203-2013 | 62945 | No | GES | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | "***Sharpoint IQ Geometry Slit Knife Angled, Bevel Up***Surgical Specialties Corporation Reading, PA***REF 72-3040***" |  |  | August 15, 2012 | 2012 | Surgical Specialties Puerto Rico Inc. dba Angiotech PR Inc. | Potential for bent tips. | Process control | Angiotech sent an Urgent Product Recall Notification letter dated August 15, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the current inventory, segregate the affected lot numbers, and to return the identified product. Customers were asked to complete the attached Customer Acknowledgement Form and fax to the Quality Assurance department at 787-658-1073. If product was further distributed customers should notify their customers to whom they sold the affected product. Customers with questions should call 787-658-1800 option 2 ext. 1775 or 787-658-1800 option 2 ext 1807.For questions regarding this recall call 787-658-1800. |  | 117948 | 117,948 units (total) | Worldwide Distribution - USA including DE,IL, NC, CA, CO, GA, KY, PA, MD, MN, NY, and MA and internationally to Austria, UAE, China, Argentina, Italy, Taiwan, Czech Republic, Australia, UK, India, Germany, Spain, France, South Africa, Malaysia, Japan, Lebanon, Korea, Greece, Saudi Arabia, Switzerland, Pakistan, Tunisia, Serbia, Morocco, Netherlands, Portugal, Truk, Belgium, and New Zealand | Terminated | January 13, 2014 | 516 | 117948
Z-0205-2013 | 63412 | No | DZN | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Vector TAS Modular Driver, a manual driver for Vector TAS dental screws. |  |  | December 12, 2011 | 2011 | Ormco/Sybronendo | The Instructions For Use (IFUs) for Vector TAS Modular Driver for some foreign language sections incorrectly lists the sterilization temperature as 130ÂżC. The correct sterilization temperature should be listed as 132ÂżC and/or 270ÂżF. | Error in labeling | "Urgent - Medical Device Correction" letters dated 12/12/2011 were sent via USPS 1st class mail to all customers and again on 2/21/2012 for recall expansion. The device issue was described and recommended actions were provided. Customers were instructed to acknowledge and complete the Return Form. Questions were referred to (800) 854-1741. |  | 4543 | 4,543 | Worldwide distribution, including Nationwide (USA) and the countries of Canada, Austria, Belgium, Bulgaria, Switzerland, Germany, Denmark, France, Estonia, Spain, Finland, Iceland, United Kingdom, Greece, Croatia, Hungary, Ireland, Italy, Luxembourg, Netherlands, Monaco, Poland, Romania, Serbia, Slovenia, Turkey, Jersey, and Martinique. | Terminated | November 05, 2012 | 329 | 4543
Z-0195-2013 | 63368 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL80 FLEX CO-OX analyzer, model #393-841 (all software versions prior to 1.35). |  |  | January 06, 2010 | 2010 | Sendx Medical Inc | The recall was initiated because SenDx Medical Inc. became aware of a software issue in the ABL80 FLEX CO-OX analyzer related to the reporting of the hemoglobin fractional parameters. | Software design | SenDx Medical sent a recall letter to all their foreign customers on January 6, 2010, to informed them of the ABL80 FLEX CO-OX recall via the Radiometer Field Action Notification (FAN) system. The Field Action Notes informed the customers of the problem identified and the action to be taken. Customers were instructed to submit confirmation fax#1 to confirm that the customer letter has been distributed to all ABL80 CO-OX customers by January 31, 2010. Customers were instructed to submit confirmation fax #2 to confirm that all ABL80 CO-OX analyzers have been upgraded to version 1.35 by 6/30/10.In the U.S., the customer was provided with a software upgrade CD along with a letter advising them to the issues associated with the recall. Customers with questions were instructed to contact TechLine at (800) 736-0600 opt 2. |  | 767 | 767 units | Worldwide Distribution-- US Distribution including the state of Ohio and the countries of Switzerland, Germany, France, Poland, Austria, Italy, Japan, Australia, Denmark, China, Sweden, Norway, Spain, United Kingdom, Netherlands, and Czech Republic. | Terminated | November 05, 2012 | 1034 | 767
Z-0207-2013 | 63396 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Trochar Scalpel Instrument (Surgical Punch) |  |  | September 13, 2012 | 2012 | OrthoPediatrics Corp | A complaint received from sales representative that a replenishment instrument was too short. Upon investigation, it has been determined that the device was not manufactured to design specifications and will not function as intended. | Reprocessing Controls | OrthoPediatrics Corp sent an Urgent Recall e-mail dated September 14, 2012, to all affected customers. The e-mail identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine all identified devices from the affected lot numbers immediately. Customers were also asked to complete the Mandatory Reply Form. Customers with questions should call 574-268-6379. |  | 78 | 78 | Nationwide Distribution including AL, IL, NY, FL, TX, MI, KY, NC, CO, and GA. | Terminated | September 18, 2013 | 370 | 78
Z-0193-2013 | 63094 | No | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | Boule Con Diff US Tri Pack, 6 Vials. |  |  | August 04, 2010 | 2010 | Clinical Diagnostic Solutions | Clinical Diagnostic Solutions, Inc. in Plantation, FL recalled the Boule Con Diff US Tri Pack (Product part 501-605, Lot #1005-549 and Product part 501-607, Lot #1005-549) which is used for in vitro diagnostic use as a control to monitor the performance of multi-parameter hematology instruments. The recall was imitated due to receipt of customers stating the MCV values and RDW% recovered for Boule | Process control | Clinical Diagnostic Solutions, Inc. sent a "PRODUCT NOTIFICATION:BOULE CON DIFF TRI-LEVEL LOT 1005-549" letter dated August 4, 2010 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For questions regarding this notice, contact CDS Customer Service at 1-800-453-3328. |  | 532 | 532 packs | Nationwide Distribution. | Terminated | August 14, 2014 | 1471 | 1222
Z-0196-2013 | 63506 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The touchscreen display (model 94267) designed for use with the XPREZZON bedside monitor (model 91393) is width-452 mm (17.2 inches) and height- 368 mm (14.5 inches). The depth without the pedestal --- 67.7 mm (2.67 inches). The display is an integrated LCD 19-Inch. |  |  | October 05, 2012 | 2012 | Spacelabs Healthcare, Llc | Potential failure of Spacelabs Healthcare Touchscreen Display, Model 94267 where the alarm lights failed to illuminate upon installation. | Vendor change control | Spacelabs Healthcare sent a URGENT - MEDICAL DEVICE CORRECTION letter dated October 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are advised to advise their staff of this potential failure and advise them to ensure the alarm lights are functional. The firm will contact the consignee at the earliest possible date to schedule a convenient time for the hospital to upgrade the affected product at no cost.Customers can call Technical Support at 1-800-522-7025 and select 2 for any questions about this recall. |  | 169 | 169 units distributed (98 in US and 71 Internationally) | Worldwide Distribution--US (nationwide) including states of Alaska, California, Florida, Illinois, Michigan, South Carolina, and Texas., and the countries of Canada, Chile, France, Libya, Panama, Qatar, Saudi Arabia, Switzerland and United Kingdom | Terminated | March 25, 2013 | 171 | 169
Z-0253-2013 | 62833 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension Vista(R) Homocysteine Flex(R) Reagent Cartridge |  |  | August 03, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Cartridges are leaking. The leaking is observed when opening the wrapper that holds the reagent cartridge. | Packaging | Siemens sent a Urgent Field Safety Notice letter dated July 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.At this time Siemens Healthcare Diagnostics is conducting a voluntary recall for Dimension VistaÂż HCYS FlexÂż reagent cartridge lots 12062MA and 12066MA. Please immediately discontinue use and discard any remaining inventory of Dimension VistaÂż HCYS (K7044) lots 12062MA and 12066MA. Please complete the attached form and fax it to (302)631-8467 to indicate that you have received this information. We ask that you forward this notification to anyone to whom you may have distributed this lot.Contact your Siemens Customer Service Center to report your no-charge replacement needs at 888-588-3916.If you have technical questions or concerns, please contact the Siemens Technical Solutions Center at 800-441-9250 for further assistance.We apologize for the inconvenience that this situation has caused. Thank you for your patience and continued support. |  | 583 | 583 cartons | USA (nationwide) including the states of AL, CA, CO, CT, DC, DE, FL, IA, IL, IN, LA, MA, MD, MI, MO, NC, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA and WI. | Terminated | January 15, 2014 | 530 | 583
Z-0250-2013 | 63369 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory, version 7.0 |  |  | November 03, 2011 | 2011 | Sunquest Information Systems, Inc. | Sunquest is recalling Sunquest Laboratory version 7.0 because when saving an Accession to a non-current Episode/Event, a User-Defined field can be incorrectly replaced by the current Episode/Event User-Defined data. | Software design | Sunquest Information Systems, Inc. sent a recall letter dated November 3, 2011, to all affected customers. Customers were informed of the problem identified and the actions to be taken. Customers were informed that a change request CR11-00174907 has been created to address this issue. This will be available as part of Sunquest Laboratory version 7.0.1. To request the software correction, customers were instructed to create a Case via the Client Support webpage at www.sunquestinfo.com/supportweb. Then choose a new case, and finally select work queue and select software request. Customers with questions were instructed to call (877) 239-6337 for US and Canadian sites and at +44 (0) 161 335 0562 for international sites.For questions regarding this recall call 520-570-2252. |  | 14 | 14 sites | Worldwide Distribution - USA (nationwide) and Internationally to Ireland | Terminated | November 06, 2012 | 369 | 14
Z-0252-2013 | 63270 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | STA Satellite Automated Multi-Parameter Analyzer |  |  | June 15, 2012 | 2012 | Diagnostica Stago, Inc. | Following a single customer complaint that incorrect INR values were occasionally printed for PT tests, even though the primary units (seconds) were correct, the firm investigated and confirmed the potential for printing incorrect secondary units when using an external printer and the customer printout mode on the STA SatelliteÂż. | Software design | Diagnostica Stago sent an Urgent Field Safety Notice dated June 12, 2012, to all affected customers via email return receipt requested on June 15, 2012. Customers were instructed to follow the instructions in the notification and complete the Customer Verification form and return to the Hotline by fax to 973-644-9348. For additional information, customers were instructed to contact the Hotline at 800-725-0607.For questions regarding this recall call 973-631-1200, ext 2044. |  | 237 | 232 units US; 5 units Canada | Worldwide Distribution - USA (nationwide) and Canada | Terminated | July 19, 2013 | 399 | 237
Z-0251-2013 | 63380 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Application Interfacing Outbound Phlebotomy Point of Care Interface, v3.0.0_07222004 and later. |  |  | November 19, 2009 | 2009 | Sunquest Information Systems, Inc. | The recall was initiated because Sunquest Information System has confirmed reports of an Outbound Phlebotomy Point of Care (POC) interface is combining patient orders inappropriately. | Nonconforming Material/Component | Sunquest sent an "URGENT-PRODUCT SAFETY NOTICE: PSN-09-06" letter dated November 19, 2009 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were informed that Sunquest Application Interfacing would be contacting their facility for scheduling to correct the issue. |  | 9 | 9 units | Nationwide Distribution. | Terminated | November 06, 2012 | 1083 | 9
Z-0243-2013 | 63103 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | NCBÂż cancellous screw |  |  | September 04, 2012 | 2012 | Zimmer, Inc. | Zimmer Inc., is initiating a correction to the patient labels of products manufactured before March 20I0 by the Zimmer GmbH production site in Winterthur, Switzerland, after receiving four complaints of a discrepancy between product patient labels and product labels with respect to the reference and lot numbers of the product. While the information printed on the product label is correct, the info | Error in labeling | Zimmer, Inc. sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated September 4, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Certificate of Acknowledgement Form was attached for customers to complete and return via fax to Zimmer, Inc. at 574 372-4265. |  | 2 | 2 | Nationwide Distribution. | Terminated | April 16, 2014 | 589 | 23110
Z-0255-2013 | 62580 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Sysmex XE-5000 Automated Hematology System; an automated differential cell counter; Sysmex Corporation, Kobe, Japan; Sysmex America, One Nelson C. White Parkway, Mundelein, IL 60060; catalog #063-7591-7 |  |  | June 01, 2012 | 2012 | Sysmex America, Inc. | There is a variation of reticulocyte counts between XE Series (XE-2100, XE-5000) and XT-Series (XT-2000i, XT-4000i) hematology analyzers that may result in a retic counts that are an average of 21% lower on the XE-Series because of a low bias. | Device Design | Sysmex America, Inc. sent a "PRODUCT NOTIFICATION" letter dated May 2012 to all affected customers on June 1, 2012. The letter identifies the product, problem, and actions to be taken. The letter instructs customers to review patient results to determine if their Reference Intervals need adjustment, and to review QC results via Insight and document any change that may be linked to the date of calibration if appropriate. Any questions concerning the information contained in the notification were directed to the Sysmex Technical Assistance Center at 1-888-879-7639. |  | 763 | 763 units | Nationwide distribution: USA including Puerto Rico. | Open, Classified |  |  | 1931
Z-0209-2013 | 52796 | Yes | CBF | Hyperbaric chamber. | Anesthesiology | Anesthesiology | 510(k) | Sechrist Monoplace Hyperbaric Chamber with Electronic Control Module Model 3600E/ER |  |  | August 28, 2008 | 2008 | Sechrist Industries Inc | Sechrist conducted a field upgrade on the hyperbaulic chambers equipped with Electronic Control Modules in order to correct a software bug that could potentially affect the performance of the chamber. | Software design | Sechrist Industries contacted consignees by phone to notify them of a software upgarde and to schedule appointments to perform the required field upgrades. An approximate summary of the phone notification is provided below."I am contacting you to notify you of a field upgrade we are performing on ourhyperbaric chambers equipped with Electronic Control Modules in order to correct a software bug that could potentially affect the performance of your chamber. The upgrade includes the installation of a new motherboard with revised software, a reset switch, and the exchange of other control panel components. The process will take approximately 2.5 hours and will be performed at no charge to you."For questions regarding this recall call 714-579-8400. |  | 65 | 65 units total (61 units in the US) | Worldwide Distribution - USA (nationwide) and the countries of Bulgaria, and Canada | Terminated | November 06, 2012 | 1531 | 138
Z-0259-2013 | 63540 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AMS AdVance" Male Sling System |  |  | October 19, 2012 | 2012 | American Medical Systems, Inc. | American Medical SystemsÂż (AMS) is issuing a voluntary recall on a number of lots of the AdVanceÂż Male Sling System (AdVance).AMS has recently received a small number of complaints regarding the resorbable suture in the AdVance sling not lasting for the entire 15 month shelf life of the product. An internal investigation has been conducted and confirmed that some sutures degraded prior to the | Device Design | American Medical Systems sent an "URGENT MEDICAL DEVICE RECALL" dated October 18, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to return any existing stock of the recall product and to complete and return the Product Recall Acknowledgment Form. Contact AMS Customer Service at 1-800-328-3881 for additional information or questions.. |  | 4872 | 4872 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the countries of Austria, Australia, Brazil, Canada, Germany, Estonia, Spain, Finland, France, Great Britain, Hong Kong, Croatia, India, Morocco, Mexico, Malaysia, Netherlands, New Zealand, Panama, Poland, Russian Federation, Solomon Islands, Singapore, Thailand, Turkey, Venezuela, South Africa, Chile, Colombia, Costa Rica, Slovenia, Argentina, China, and Taiwan. | Terminated | April 11, 2013 | 174 | 4872
Z-0258-2013 | 63598 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MicroCool Breathable Gown with Towels, X-Larg, X-Long, Sterile |  |  | July 05, 2012 | 2012 | Kimberly-Clark Corporation | Package seam of some units may not remain properly sealed to ensure the contents might not be sterile as labeled. | Packaging process control | Kimberly Clark sent an Urgent Voluntary Product Recall letter dated July 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to determine if they have any affected product in their inventory. Any affected product should be quarantined and distribution discontinued. Distributors were also asked to conduct a sub-recall for any affected product further distributed. Customers were asked to complete the attached Distributor Product Recall Response Form and fax to 1-920-225-3522. Upon completion of the response form customers were asked to destroy any affected product.For questions regarding this recall call 770-587-7835. |  | 933 | 933 ((30 units per case) | Nationwide Distribution | Terminated | June 13, 2013 | 343 | 90789
Z-0270-2013 | 63083 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | AC Powered adjustable hospital bed, Prime Bed TL500, Distributed by: Primus Medical LLC, 8401 Southern Blvd, Boardman, OH 44512. |  |  | February 02, 2011 | 2011 | Primus Medical LLC | The "Retaining washers &amp; Clevis Pins" provided with the AC -powered adjustable hospital beds may fail under certain circumstances. The pins at the top of the leg were disengaging from the channel. This may cause the bed frame to come off tracks or collapse. | Equipment maintenance | The firm, Primus Medical LLC, sent an "Important Notification Relating to Primus Medical Product TL500" letter dated March 29, 2011 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to replace the "Retaining washers &amp; Clevis Pins" with the "Clevis Pins &amp; Bow Tie Cotter Pins" in the TL500-001 Service Pack provided with the letter in order to avoid any potential failure and/or bed collapse and to return the" Retaining washers &amp; Clevis Pins" after the replacement procedure is performed. If you have any questions, or need assistance in performing the replacement, call 330-965-9681. |  | 342 | 342 Beds | Nationwide distribution: USA including states of: AZ, IL, KY, MD, MI, OH, OR, PA, TX and VA. | Terminated | May 14, 2014 | 1197 | 342
Z-0268-2013 | 63297 | No | ELW | Impression material. | Dental | Dental | 510(k) | The brand name of the device is AlgiNot, a dental impression material. |  |  | February 21, 2012 | 2012 | Kerr Corporation | The firm recalled the product because it may cure harder than expected within theshelf life, which could result in a material that may be more difficult to remove. Use of this product onpatients with loose teeth may result in possible tooth fracture or tooth extraction. | Nonconforming Material/Component | Kerr Corporation sent "URGENT: MEDICAL DEVICE RECALL" letters to all affected customers. The letters identified the product, problem, and actions to be taken by the customers. US, Canadian, and Rest of World consignees were notified via USPS 1st class mail on June 19, 2012, and Australianconsignees on July 5, 2012. Consignees were instructed to complete and return the Acknowledgement Form. Notification of the recall expansion was sent, via USPS 1st class mail, on February 21, 2012 to US, Canadian, Australia and European, consignees; however, once further investigation determined that all unexpired lots of the product were affected, the Alginot recall was expanded a second recall letter was sent to clarify that all lots of AlgiNot FS Volume recall and AlgiNot recall expansion are affected by this expandedrecall. Letters were mailed to US, Canadian and Rest of World (ROW) customers on June 19, 2012; Australian and European consignees were notified of the expansion on July 5, 2012.All consignees were instructed to complete the Return Form to acknowledge the updated instructions for use and the added cautionary statement. Consignees not responding are to be contacted at least three times via telephone/fax/e-mail as follow up on the recall. |  | 132 | 45 kits and 87 refills | Worldwide Distribution-USA (nationwide) and the countries of Canada, Europe, and Australia. | Terminated | November 15, 2012 | 268 | 132
Z-0260-2013 | 62271 | No | IZW | Diagnostic x-ray beam-limiting device. | Radiology | Radiology | 510(K) Exempt | R-30H , Diagnostic X-Ray Beam-Limiting Device. |  |  | October 28, 2011 | 2011 | Shimadzu Medical Systems | Shimadzu is recalling Diagnostic X-ray Beam-Limiting Device, Mobile x-ray system and Diagnostic x-ray tube assembly because they do not comply with the Product Performance Standards, 21CFR 1020.30 (m) Beam quality --(1) Half-value layer (HVL), which was revised in 2006. Certain devices might not comply with the requirements for them in some x-ray generating condition after the revision in 2006. | Nonconforming Material/Component | Shimadzu sent an Urgent Voluntary Medical Device Recall - Field Safety Notice letters dated September 19, 2012 to all the affected customers. The letter identifed the description of non-compliance, how to identify the affected products, corrective action plan by Shimadzu and actions to be taken customers. Customers were instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at Shimadzu Medical Systems National Technical Support (800) 228-1429. |  | 665 | 665 units | US Nationwide Distribution | Open, Classified |  |  | 1402
Z-0274-2013 | 63383 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5 Mast Roller Pump 150, Item Number: 10-88-00 |  |  | September 19, 2012 | 2012 | Sorin Group USA, Inc. | Sorin Group USA, Inc. has issued an Urgent Field Safety Notice regarding certain serial numbers of the Sorin Group S5 Perfusion System due to reports of a "Fault in motor controller" alarm message displayed on the system panel. This could lead to unintended pump stoppage which may result in permanent impairment or life-threatening injury. An out of specification electrical component was determin | Mixed-up of materials/components | The firm, Sorin Group, sent an "URGENT FIELD SAFETY NOTICE" dated September 19, 2012 to all affected customers by certified mail. The notice described the product, problem and actions to be taken. The customers were instructed to follow the provided instructions for the continued safe use of the Perfusion systems until they have been serviced or replaced by Sorin Group. The Sorin Group Service Team will contact the customer to schedule servicing and replacement of the affected product. The customers were also instructed to complete and return the Customer Response Form via fax to 303-467-6502 or by email to yvonne.feyerherm@sorin.com; assure this notice is distributed to all personnel who need to be aware of this notice; and if they have transferred the affected products to a third party, pass this information with them as well as with Sorin Group Customer Service at 1-800-650-2623. For questions regarding this notice, contact Sorin Group USA Customer Service at 1-800-650-2623. |  | 2 | 2 units | Worldwide distribution: USA (nationwide) including states of: AZ, GA, MI, MS, NC, NY, PA, SC, and TX; and countries of: Austria, Australia, Belgium, Canada, People's Republic of China, Columbia, Denmark, Finland, Germany, India, Iran-Islamic Republic of, Ireland, Japan, Lebanon, Morocco, Pakistan, Poland, Russian Federation, Saudia Arabia, Spain, South Africa, Thailand, Turkey and United Kingdom. | Terminated | January 18, 2013 | 121 | 47
Z-0266-2013 | 63273 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Hoffman LRF Telescopic Strut Short |  |  | July 03, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Three complaints have been filed where the Hoffman LRF Telescopic Struts have broken during load bearing application by patients having a body weight greater than 250 LBS. | Device Design | Stryker Orthopaedics sent an "URGENT MEDICAL DEVICE FIELD REMOVAL" letter dated July 2, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Business Reply Form was included for customers to complete and return to via fax to 1-865-252-3635. Contact the firm at 1-866-626-7747 for assistance. |  | 165 | 165 units | Nationwide Distribution. | Terminated | April 17, 2014 | 653 | 165
Z-0269-2013 | 63599 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TI Standard Polyaxial Screw Assembly 6.5MM x 55MM. Part # 62065-55. |  |  | December 09, 2010 | 2010 | Alphatec Spine, Inc. | The recall was initiated because Alphatec Spine confirmed that the TI Standard Polyaxial Screw Assembly was laser marked with the incorrect length of 45 MM. | Error in labeling | Alphatec Spine distributors who received the TI Standard Polyaxial Screw Assembly were contacted by phone and by email on December 9, 2010. A phone script was used as a recall notification to contact their three distributors who received the recalled products. The recall notification informed the customers of the problem identified and the action to be taken. Customers were informed that they will receive replacements within 2-3 days. Customers with questions were instructed to call (800) 922-1356.For questions regarding this recall call 760-431-9286. |  | 11 | 11 units | Nationwide Distribution including TN, WI, and FL. | Terminated | November 09, 2012 | 701 | 11
Z-0267-2013 | 53337 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Quadra S Offset Broach Handle |  |  | October 13, 2008 | 2008 | Medacta Usa Inc | Medacta USA Inc has initiated a recall on the Quadra S Offset Broach Handle because of the potential of the rivet that locks the metallic wire to the lever may become loose during a surgery and make it impossible to lock/unlock the broaches. | Nonconforming Material/Component | Medacta USA, sent a Urgent Safety Notification letter dated October 13, 2008, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.The letter informed the customers of the lot numbers on recall and the reference numbers. The recall was initiated because the offset broach handles contain a lever mechanism that is connected to the locking axis by means of a metallic wire. The rivet that locks the metallic wire to the lever may become loose during a surgery and make it impossible to lock/unlock the broaches. For further questions please call (858) 705-0350. |  | 12 | 12 units | Worldwide Distribution -- US including the states of Texas, Nevada and Colorado., and the country of Belgium. | Terminated | November 09, 2012 | 1488 | 12
Z-0276-2013 | 63445 | No | GDI | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Progressive Medical PMI Endoscopic Dissector with Kittner Tip (5mm x 40cm), Reorder Number: PMILK40, Product Sterile in Unopened/Undamaged pouch, Cotton Tip - X-Ray Detectable, 1 Pc/Pouch, 10 Pouches/Box, 20 boxes/case. Exclusively Distributed in the U.S. Market by: Progressive Medical, Inc., St. Louis, MO, Made in China. |  |  | December 09, 2011 | 2011 | Progressive Medical Inc | Some of the sterile packages had broken seals. | Packaging | Progressive Medical Inc contacted their customers via telephone on December 9, 2011, through December 15, 2011.Customers were informed the reason for recall was explained and the customer was asked if any of the lot remained in their inventory. If inventory remained, the recalling firm sent them a return goods authorization number and informed them the product would be replaced. For further questions please call (314) 961-5786. |  | 70 | 70 boxes | USA (nationwide) including the states of CO, IN, KS, MA, MN, MO, NE, OH, SC, TN, TX, VA, and WI. | Terminated | November 14, 2012 | 341 | 70
Z-0287-2013 | 53336 | bath, flotation, tissue | IDY | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | PT Link Waterbath, Model#PT100. |  |  | May 01, 2008 | 2008 | Dako North America Inc. | Dako is recalling the PT Link Waterbath because of a possible electrical hazard. Specifically, if the heater pad under the reagent tank becomes wet, it may present an electrical shock hazard when the lid is open. | Device Design | The firm, Dako, sent a "DAKO PT LINK ADVISORY NOTICE" dated May 1, 2008 to customers who purchased the PT Waterbath along with a customer acknowledgement document informing the users of the potential electrical hazard. The notice described the product, problem and actions to be taken. The customers were instructed to follow the attached Safety Instructions; place the enclosed warning labels directly on the instrument as shown in the attached instructions and adhere to the warning labels as described. Customers with questions were instructed to call Dako Products' Technical Support at (800) 424-0021. |  | 136 | 136 units | Nationwide distribution: USA including states of: AK, CA, CO, CT, DC, FL, GA, HI, IA, IL IN, KS, KY, MA, MD, MI, MN, MO, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, and WA. | Terminated | November 13, 2012 | 1657 | 136
Z-0288-2013 | 63586 | No | MIH |  |  | Cardiovascular | PMA | TriVascular brand Fill Polymer Kit, Endovascular Graft, Aneurysm Treatment System, European Model Number: TVFP14-B, (in European distributor inventory only); |  |  | October 19, 2012 | 2012 | Trivascular, Inc | Selected lots of the Fill Polymer Kit (European Model Number: TVFP14-B) in European distributor inventory only that potentially contain leaking stopcocks due topossible cracks in the stopcock component. TriVascular indicated that these Fill Polymer Kits are for use with the Ovation or Ovation Prime Abdominal Stent Graft Systems that may display leaks upon initiation of mix but prior to use in | Nonconforming Material/Component | Trivascular sent an Urgent Field Safety Notice letter dated October 19, 2012 to all affect consignees. A verbal (telephone) communication was also initiated. The forms communication identified the affected product, problem and actions to be taken. Consignees were instructed to immediately quarantine and return affected products. A return response form was included for consignees to complete and returned via the e-mail address or fax number provided. Consignees were instructed to further notify its (sub-level) consignees, and have them return affected product directly to the recalling firm. For questions call 1-707-543-8732. |  | 41 | 41 units | Worldwide Distribution - US Nationwide and the countries of Belgium, Cyprus, Italy, Spain, Switzerland and Turkey. TriVascular stated that no purchased devices are in hospital inventory in Europe or the U.S. | Terminated | December 17, 2012 | 59 | 41
Z-0282-2013 | 62733 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Perfusion System 8000, base 4 pump, 220V/240V safety monitor included |  |  | September 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received multiple reports since 2007 of various malfunctions of the arterial and cardioplegia monitors for the Sarns ModularPerfusion System 8000. The malfunctions can be broadly categorized as follows monitor power malfunctions, inability to measure pressure and inaccurate and/or false pressurealarms.If the malfunction of the arterial or cardio | Device Design | The firm, Terumo CVS, sent an "URGENT MEDICAL DEVICE RECALL" notice/Safety Advisory Alert dated September 24, 2012 and an addendum to the Operator's Manual to alert all affected Customers. The notice described the product, problem and actions to be taken. The customers were instructed to review this medical device recall notice; assure that all users are aware to this notice; place the addendum in the Operator's Manual just before page 4.1 and complete and return the attached Customer Response Form via Fax to: 1-734-741-6149 or email to:www.termuo-cvs.com.If you have any questions or concerns contact Terumo CVS Customer Service at 1-800-521-1818; Monday - Friday, 8 AM - 6 PM ET. |  | 466 | 466 | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, and WY; and countries of: ARGENTINA, AUSTRALIA, BANGLADESH, BRAZIL, CALI COLOMBIA, CALLO PERU, GERMANY, GUAYQUIL ECUADOR, HONG KONG, INDIA, INDONESIA, JAPAN, LEBANON, MADRID SPAIN, MEXICO, NEW ZEALAND, SANTIAGO, SANTIAGO CHILE, SINGAPORE, SOUTH AFRICA, SOUTH KOREA, SRI LANKA, TAIWAN, THAILAND, and UNITED ARAB EMIRATES (UAE). | Terminated | July 11, 2013 | 289 | 2256
Z-0277-2013 | 63446 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes External Fixation System Tube to Tube Clamp, Part Number 390.007. |  |  | September 28, 2012 | 2012 | Synthes USA HQ, Inc. | The insert provided with the product for the specific lot number indicates that this product is MR safe. This product is MR Conditional and is etched as MR Conditional. The insert is incorrect. | Labeling mix-ups | Synthes sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated September 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return a Verification Form via fax to 610-251-9005. Contact the firm at 610-719-5450 for questions regarding this notice. |  | 8 | 8 | Nationwide Distribution-including the states of CA, MN, NC, and TN. | Terminated | September 21, 2015 | 1088 | 8
Z-0278-2013 | 63344 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | HX2 Temp Manage SYS 120V |  |  | October 16, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Based on a retrospective review of quality data, AC Control boards used in the manufacture of theHX2 Temperature Management System were identified as being manufactured using the incorrectflux material. Use of the incorrect flux may contribute to premature board failure. A search ofTCVS complaints revealed 3 complaints received related to the HX2 failure. Each failure occurredduring set up | Nonconforming Material/Component | TERUMO sent an Urgent Medical Device Recall letter dated October 15, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Terumo has determined that a field correction is required to notify affected customers that the boards may fail, instruct the users that in a case of failure to swich to the second channel on the device or use a back-up unit, and that TCVS will replace the affected boards. Customers with questions were instructed to call 1-800-621-2818. For questions regarding this recall call 734-741-6173. |  | 19 | 19 units | Nationwide Distribution including CA, HI, KS, LA, MI, MO, OR, PA, UT, and WI | Terminated | July 11, 2013 | 268 | 34
Z-0286-2013 | 63275 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Special Orthopaedic Solution |  |  | July 16, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has become aware that insufficient taper torsional strength may result when a PAR extension piece is used with a forged CoCr stem. | Device Design | Stryker Orthopaedics sent an Urgent Product Recall letter dated July 16, 2012 via Fed Ex with return receipt to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter requested customers to contact their Stryker Sales Representative to arrange for return of the affected product. Customers were instructed to fax back the attached Product Recall Acknowledgement Form to 201-821-6069. For questions call 201-972-2100. |  | 301 | 301 units | Worldwide Distribution - US Nationwide | Terminated | September 09, 2015 | 1150 | 301
Z-0289-2013 | 63536 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABX PENTRA Creatinine 120 CP, Model Number Creatinine Part # A11A01868, Lot Numbers: 203273706, 203273707, 203273808, 203273709. |  |  | September 14, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | This recall was initiated because some customers were experiencing an alarm during Creatinine calibration "REAG RANGE HIGH". This alarm caused the calibration to fail and prevented the user from using the Creatinine assay. | Nonconforming Material/Component | HORIBA Medical sent customer communication letters dated October 24, 2011, to all affected customers via Fax with the exception of two communications that were emailed. The response or confirmation of receipt was conducted via telephone and also documented on attached Notification List. Customers were given a list of procedures to follow and asked to complete and fax the Acknowledgement of Receipt Form to +33 (0)4 67 14 15 17.For questions regarding this recall call 949-453-0500, ext 208. |  | 8503 | 8,503 units | Nationwide Distribution | Terminated | November 29, 2012 | 76 | 8503
Z-0295-2013 | 63603 | No | LYZ | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Select Medical Products Vinyl Powder-Free Exam Gloves, Medium Model # 6370 and Large Model # 6371, 100 per box; 10 boxes per case. |  |  | June 20, 2012 | 2012 | Gulf South Medical Supply Inc | Gulf South Medical Supply recalled their Select Medium and Large size Vinyl Exam Gloves because they were inadvertently shipped prior to U.S. Customs clearing them for sale. | Employee error | The firm, PSS World Medical, sent an "URGENT RECALL NOTICE" dated June 20, 2012 via USPS to all customers that received the product. The notice described the product, problem and actions to be taken. The customers were instructed to examine their inventory, quarantine the affected lot numbers listed in this recall and return the product; If the product was further distributed, to inform their customers of the recall and request they return the affected product directly to them; and contact their Gulf South Customer Care team via telephone at 1-800-347-2456 or email to customercare@gsms.com to request a return goods authorization.If you have any questions, contact the Product Integrity Leader Quality Assurance at 904-380-4501. |  | 292 | 275 cases and 17 boxes | Nationwide distribution: USA including states of: DE, KY, LA, MD, NJ, NY, PA, TX, VA, and WV. | Terminated | August 19, 2014 | 790 | 292
Z-0296-2013 | 63458 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ATS 3000 Automatic Tourniquet System |  |  | September 05, 2012 | 2012 | Zimmer, Inc. | Zimmer is conducting a voluntary removal of some Zimmer A.T.S. 3000 tourniquets after receiving complaints of both out-of-box failures and failures during surgical procedures on some of the affected units. During failure the unit will alarm, lock into a non-operational mode, and the display screen will go blank. Based on investigation, this problem appears to be due to a counterfeit chip on the | Counterfeit | Zimmer sent an Urgent Device Removal Notice letter dated September 5, 2012, to all affected customers. Zimmer Distributors and Sales Leadership were notified via E-mail of the pending recallthe product under recall beginnig September 11, 2012. The notification listed the reason for the recall, risks involved, and requested return of the unit for repair To: Zimmer Surgical, Attn: QAlRA Dept. - Recall, 200 West Ohio Avenue, Dover, Ohio 44622 USA The notice asked to ensure all users be made aware of the recall, remove the affected ATS 3000 tourniquest sytem for repair and return the fax back form/acknowledgement certification to 1-888-429-7380 or email it to zimmer3611@stericycle.com. Questions should be directed to 1-888Âż943Âż5167 between 8 a.m. and 5 p.m. EDT.For questions regarding this recall call 574-372-4807. |  | 204 | 204 | Worldwide Distribution - USA (nationwide) and Internationally to Belgium, Czech Republic, Egypt, France, Germany, Morocco, Netherlands, Romania, Saudi Arabia, Spain, Switzerland, Turkey, and UAE. | Terminated | April 17, 2014 | 589 | 204
Z-0294-2013 | 63660 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | EnteraLite Infinity Enteral Feeding Pump identified with the following list numbers: INFKIT2 and INFORNG. The device is used with disposable PVC administration tubing sets to provide a means for the volumetric delivery of fluids used in enteral applications. |  |  | November 08, 2012 | 2012 | MOOG Medical Devices Group | MOOG Medical Devices Group is conducting a voluntary for a limited set of serial numbers of the EnteraLite Infinity and Infinity Orange Small Volume Enteral Feeding Pumps because certain diodes on the PCB assembly may result in insufficient voltage to a portion of the pump resulting in an error 13 message and shutting the pump down entirely. | Component design/selection | MOOG Medical Devices Group sent a Urgent Medical Device Recall letter dated November 8, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to contact Moog Medical Devices Group to discussthe process for product replacement at 1-800-970-2337, prompt #7.See Attachment A to determine if you have a serial number that is related to a suspect pump. Determine the number of affected pumps. Notify all of your affected customers of the recall and the protocol for getting a replacement pump. Provide serial numbers, number of product(s) to be returned and a shipping address for replacement product. Execute the pump replacement process discussed in the conversation with Moog Customer Advocacy. Package and return product via shipping label. |  | 8736 | 8,736 | Worldwide Distribution - USA (nationwide) including the countries of Canada and Lithuania. | Terminated | January 30, 2014 | 448 | 8736
Z-0198-2013 | 63597 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima XR 220amx |  |  | September 20, 2012 | 2012 | GE Healthcare, LLC | It was discovered GE Healthcare has become aware of a firmware issue on the mobile radiographic product, Brivo XR285amx, Optima XR200amx, Optima XR22Oamx and Optima XR200amx with Digital Upgrade, related to Technique Accuracy as mandated by the US Code of Federal Regulations (21CFR) and IEC 60601-2-54. | Radiation Control for Health and Safety Act | GE Healthcare will remedy the defect or bring the product into compliance with each applicable Federal and IEC standard. A Service Representative will update the firmware on the system to address the issue.Field Modification Instructions (FMI) 10870, 10872, and 10873 are describing the rework plan the GEHC Engineers will make to the imaging systems compliant with the applicable performance standard.Further questions please call (262) 513-4122. |  | 364 | 364 total units installed in US | Nationwide Distribution | Terminated | December 05, 2013 | 441 | 364
Z-0292-2013 | 63322 | No | LMF | Absorbable hemostatic agent and dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | ETHICON SURGIFLO HEMOSTATIC MATRIX KIT Plus FlexTip with Thrombin distributed by ETHICON ETHICON Inc. Somerville, NJ 08876 USA |  |  | August 21, 2012 | 2012 | Ethicon, Inc. | SURGIFLO Hemostatic Matrix Kit &amp; SURGIFLO Hemostatic Matrix Kit with Thrombin have an issue within the packing process where a cut could potentially breach the double Tyvek pouch of the packaging. | Packaging | Ethicon Inc. sent an Urgent Product Information letter dated August 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory to determine if they have any of the affected product. Affected product should be returned for replacement to Stericycle using the enclosed pre-paid shipping label. Customers with questions should call 1-877-384-4266 or contact their Ethicon Biosurgery Sales Representative.For questions regarding this recall call 877-384-4266. |  | 28276 | 28,276 | Worldwide Distribution - USA (nationwide) and the countries of Russia, China, and the European Union. | Terminated | September 11, 2014 | 751 | 73476
Z-0299-2013 | 63467 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | bioMerieux BacT/ALERT Control Module, catalog number 210147, Clinical instrument |  |  | October 16, 2012 | 2012 | Biomerieux Inc | The product may mislabel patient information on culture specimen bottles due to a software miscommunication. | Software design | BIOMERIEUX sent an Urgent Product Correction Notice to all affected customers starting on October 16, 2012, The letter identified the affected product, the problem, and the action to be taken by hte customer. The firm provided modified instructions for use to clarify how to properly load bottles and to provide appropriate cautions and warnings concerning bottle loading. Customers were instructed to ensure that the letter and guidance on how to properly scan and load bottles is distributed, and training is provided to all appropriate personnel within their organization. Check for lost bottles in their system; fill out and return the Acknowledgement Form in Attachment A by fax to confirm receipt of the notice. Customers were instructed to contact their local bioMerieux Customers Service Representative with any questions.For questions regarding this recall call 919-620-2682. |  | 70 | 70 modules | Worldwide Distribution - USA including AR, AZ, CA, FL, KS, LA, MD, MN, MT, NC, NV, NY, OR, PA, RI, SC, TX, VA and WA. Internationally to Argentina, Armenia, Austria, Bangladesh, Bosnia and Herzegowina, Brazil, Bulgaria, Burkina FASO, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, COTE DIVOIRE, CROATIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LITHUANIA, MALTA, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, PANAMA, PERU, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UNITED STATES, URUGUAY, VENEZUELA, VIETNAM, PALESTINE, AND SERBIA. | Terminated | October 28, 2013 | 377 | 6400
Z-0331-2013 | 63581 | No | LMF | Absorbable hemostatic agent and dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | Superstat Modified Collagen Hemostatic Sponge.115mmx115mm Pledget, 1.5%, Part#9500-15. |  |  | October 24, 2012 | 2012 | Superstat Corp | The recall was initiated because Superstat Corporation has confirmed that no documentation are available to verify that quarterly dose sterility audits were performed for the Superstat Modified Collagen Hemostatic Sponge. | Process control | Superstat initially contacted customers by telephone and sent an Urgent - Device Recall Notification Letter on October 24, 2012, to all affected customers. The recall letter provided the reason for recall was due to inadequate documentation of quarterly dose sterility audits. The recall includes all lots and part numbers of Superstat produced between 2002 and 2012. Customers are instructed to complete and return the enclosed response form as soon as possible. Customers with questions are instructed to call customer service at (800) 487-3786.For questions regarding this recall call 310-605-1655. |  | 26422 | 26,422 units | Worldwide Distribution - USA (nationwide) and the countries of Saudi Arabia, Thailand, Mexico, Canada, Pakistan, Israel, Greece, Serbia, Canda, Isle of Man, Nova Scotia, and Singapore. | Terminated | February 06, 2013 | 105 | 26422
Z-0330-2013 | 63481 | No | HGP | Fetal scalp circular (spiral) electrode and applicator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Corometrics Qwik Connect Plus Spiral Electrode- Model #7000AA0 |  |  | October 16, 2012 | 2012 | Vital Signs Devices, a GE Healthcare Company | Thirty fetal scalp electrodes were found to have an incomplete package seal and lacked the printed lot number stamped on during sealing. | Employee error | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter on October 16, 2012 via Fed Ex to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: DO NOT use Qwik Connect Plus Spiral Electrodes part number 7000AAO with lot number S03123 or product with missing lot numbers; isolate all affected product; complete and return the Urgent Medical Device Correction Confirmation form via fax at 800-535-7923, return form even if no recall product is in inventory; to return only unused product, and if they have forwarded any affected lot numbers of this product to any other healthcare institutions, forward a copy of this letter to those institutions.If you have any questions or concerns regarding this notification, contact GEMSIT Customer Service at +1-800-588-7044 (Domestic) or Vital Signs Customer Service at +1-800-932-0760 (International). Hours of operation: 8:00 am EST to 6:00 pm EST. |  | 126 | 126 cases (50 pouches per case) | Worldwide distribution: USA (nationwide) including states: AZ, CA, CT, FL, GA, IL, KS, KY, MA, MD, MN, MO, NJ, NM, NY, OH, PA, SC,TX, WA, and WY; and countries of: Australia and Canada. | Terminated | April 16, 2014 | 547 | 126
Z-0297-2013 | 63618 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Product is Full OSSEOTITE 2 Tapered Certain Implant, Model Number XIFNT413, and size 4 x 13 mm. The product is sterile. |  |  | September 28, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i recalled Full OSSEOTITE 2 Tapered Certain Implant and Full OSSEOTITE Parallel Walled Certain Implant due to a manufacturing condition where the depth of the implant's internal hex is too shallow. This condition may prevent the driver or abutment from fully engaging in the implant. | Nonconforming Material/Component | Biomet 3i sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 28, 2012. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-342-5454 for questions regarding this notice. |  | 191 | 191 devices | Worldwide Distribution-USA (nationwide) including the states of VA, SC, HI, LA, OH, FL, IL, NJ, NC, CA, CT, MI, TN, AK, NH, NY, MA, GA, PA, and KY and the countries of United Kingdom, Spain, Germany, Canada, Japan, Singapore, Taiwan, Israel, Poland, Hong Kong, Ireland, The Netherlands, Austria, and France. | Terminated | August 06, 2015 | 1042 | 267
Z-0291-2013 | 63314 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Mindray A5 Anesthesia Delivery System |  |  | August 13, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified the potential for a system leak resulting an issue related to the carbon dioxide absorber canister gasket on A3/A5 Anesthesia Delivery Systems. A potential system leak resulting from the canister gasket will likely present itself during the Automatic Circuit Leak and Compliance Test performed at startup and the Manual Leak Test recommended to be performed before each case. | Device Design | The firm, Mindray DS USA Inc., sent an "URGENT: A3/A5 ANESTHESIA DELIVERY SYSTEM CORRECTIVE ACTION" letter dated August 13, 2012 to their customers via Certified Mail return receipt requested. The letter described the product, problem and actions to be taken. The customers were instructed to route this letter to all potential uses of the product within their facility; increase fresh gas flow, should the leak occur during use, to compensate for any pressure loss resulting from the leak; alternatively, the absorber canister may be unlocked and locked again to reseat the gasket and verify that the canister and gasket are fully seated, a leak test is performed after replacement of the Pre-Pak or loose-fill absorbent and also be performed before each patient use until the canister gaskets are replaced on systems in their facility. A Mindray Representative will replace the canister gasket to correct the issue.If you have any questions or need to arrange for a replacement of the absorber canister gasket on the A3/A5 System(s) in your facility, contact your Mindray Service Representative, at 1-800-288-2121 (Monday-Friday 8:30am-5:30pm).A press release was issued by the firm on 11/14/2012. |  | 346 | 346 units (311 US - 35 Foreign) | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Australia, Chile, Colombia, and Venezuela. | Terminated | June 04, 2013 | 295 | 415
Z-0344-2013 | 62916 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | IO Fix IntraOsseous Fixation System; Extremity Medical. www.extremitymedical.com; Customer Service: 888.499.0079. |  |  | April 13, 2011 | 2011 | Extremity Medical LLC | Product Event Reports revealed that lag screw passes through washer (e.g. X-Post) intra-operatively. | Employee error | Extremity Medical sent an e-mail notification (Extremity Medical- Sales Bulletin- IO Fix) dated April 13, 2011 to distributors notifying them of a change made to the 4.0 mm Lag Screw. |  | 120 | 120 Units | Worldwide Distribution-USA (nationwide) including the states FL, MA, NC, CA, MD, VA, IN, AL, NJ, AZ, OK, MT, NY, OH, IA, MI, TX, WI, GA, IL, TN and NE, and the country of Switzerland. | Terminated | March 21, 2013 | 708 | 120
Z-0368-2013 | 52634 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 4.5mm x 35mm Noncannulated Screw Shank |  |  | January 05, 2009 | 2009 | NuVasive Inc | The tulip portion of the screw assembly can disengage from the shank after implantation. | Device Design | Telephone communications to the sales representatives of surgeons actively using the Posted Pedicle Screw System were made on 01/05/2009. The calls were unscripted and requested that all inventory be forwarded to NuVasive's San Diego office for quarantine. Surgeons who were actively using the system were also contacted and notified of the situation. Follow-up communications were sent via mail on 01/19/2009 to all surgeons who participated in the Alpha release. The letter informed the customers of the problems identified. Customers were instructed that the recalling firm does not believe any further action is required and they were instructed to monitor the progress of their patients and notify NuVasive if any issues are detected.Customers with questions or concerns are instructed to contact the firm at (858) 909-1964. |  | 240 | 240 units total | Nationwide Distribution-including the states of OR, MO, FL, WI, CA, and NC. | Terminated | November 19, 2012 | 1414 | 240
Z-0379-2013 | 62987 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Thermo Scientific MAS Omni IMMUNE and Omni IMMUNE PRO; |  |  | August 20, 2012 | 2012 | Microgenics Corporation | In house testing confirmed that CK-MS and Insulin claimed in both Omni IMMUNE and Omni IMMUNE PRO products are displaying reduced stability. | Nonconforming Material/Component | Microgenics Corporation sent a Recall Notification letter dated August 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed that immediate attention is required and identify the two affected constituents. The letter contains a list of the affected product with their catalog number, lot number and expiration date and request that the CK-MB and Insulin use be discontinued. The remaining components of the product may be used. Evaluate quality controls according to your laboratory procedures.Retain a copy of the letter for your laboratory records. If the product was forwarded to another laboratory, provide them with a copy of the letter.Complete the attached Recall Response form and return within 5 days.Questions may be directed to 1-800-232-3342. |  | 2076 | 2076 total in US , 596 total Foreign | Worldwide Distribution--USA (nationwide) and the countries of Germany, Canada and Australia. | Terminated | March 04, 2013 | 196 | 2076
Z-0345-2013 | 63505 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel Dxl 800 Access Immunoassay System, Part Numbers: 973100, A71456, and A71457. |  |  | August 01, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the UniCel Dxl and DxC Access Immunoassay Systems had field modifications to the aspirate and dispense probe assemblies. | Process design | Beckman Coulter sent a field modification was released in August 2011 which included:- Installing a new aspirate and dispense tubing routing bracket.- Replacing existing aspirate and dispense probe tubing with modified tubing to help eliminatemisrouting of aspirate and dispense probe tubing.During installation the FSE informed the customer:-That aspirate and dispense tubing are different lengths.- That probes are not interchangeable with other instruments once tubing has been modified.- To route probe tubing by referring to the labels on the tubing.An updated service document, Routine Maintenance Changes, A82294B was provided to every customer.For questions regarding this recall call 714-961-4483. |  | 1756 | 1,756 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, AE, AU, BG, BW, CH, CN, CZ, DE, ES, FR, GB, HK, HR, HU, IN, IT, JO, JP, LB, MA, MX, NL, PL, QA, RO, RU, SK, TN, TR, ZA, AO, AR, AT, BE, CO, DZ, FI, IE, IL, KR, KW, LI, NA, NO, NZ, PF, PH, PR, SE, SG, SZ, TH, TW, VE, VN, and YT. | Terminated | November 19, 2012 | 476 | 2188
Z-0383-2013 | 63077 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***SENTINEL***Cannulated Drill Bit***S8580***8 X 229mm (9")***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA***LOT***2012-03***". |  |  | April 10, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | Linvatec Corporation dba ConMed Linvatec in Largo, FL is recalling the S8580, 8.0MM SENTINEL Cannulated Drill Bit, lot numbers 322225 and 322226; and the S8585, 8.5MM SENTINEL Cannulated Drill Bit, lot number 322287. Recall was initiated due to the incorrect labeling on package. | Labeling Change Control | ConMed/Linvatec sent an Urgent Medical Device Recall Notificaiton letter via phone and FedEx on April 9, 2012 and again on May 11, 2012, to all affected customers. They also sent the same Notification on May 14, 2012, to international accounts. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete a reply formand fax it to 727-319-5701. Customers were also instructed to return all affected product to the firm. All returned product under this removal action will be identified, segregated and quarantined per firm procedures, etc.Customers with questions should contact Customers Service at Custserv@linvatec.com.For questions regarding this recall call 727-392-6464. |  | 70 | 70 total | Worldwide Distribution - USA including NJ, UT, GA, NY, PA, IL, LA, OH, CA, MS and Internationally to Canada, Japan, and Romania. | Terminated | November 21, 2013 | 590 | 70
Z-0385-2013 | 63607 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | The RIO (TGS 2.), Model # MAKO TGS 2.0 (Part No. 204000). |  |  | September 12, 2012 | 2012 | Mako Surgical Corporation | MAKO Surgical Group recalled their RIO System software, version 2.4 and is implementing software version 2.5 to address a software functional issue with the existing version of the system's software. Loss of tactile feedback constraining the cutting burr has been reported. | Software design | Mako Surgical Corporation sent an Important Medical Device Correction letter dated August 31, 2012, to all affected customers. The letter included specifics regarding the software issue and how the correction will take place. The letter informed the customers that their MAKOplasty Sales Specialist will coordinate with them to establish the appropriate time for the upgrade .Customers were asked to return by fax an Acknowledgement Form to 954-927-0446, ATTN: Quality Assurance. For questions customers were instructed to call 954-628-0447.For questions regarding this recall call 954-927-2044, ext 605. |  | 137 | 137 devices | Worldwide Distribution - USA including IL, OR, IA, AZ, CA, CO, FL, MS, OK, PA, WV, NC, TX, NY, GA, NJ, OH, WA, VA, MA, MD, WI, MI, NV, MO, AL, UT, RI, AR, LA, ID, IN, and TN and Internationally to Japan, Scotland, Italy, Singapore, Turkey, Hong Kong, and Korea. | Terminated | May 22, 2013 | 252 | 137
Z-0387-2013 | 63358 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics |  |  | August 17, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has received reports from the field indicating that certain lots of 6.5mm Cancellous Bone Screws are either associated with a product mix or potentially associated with a product mix. Packaging associated with a 35 mm screw may contain a 25mm screw. Conversely, packaging associated with a 25mm screw may contain a 35mm screw. | Mixed-up of materials/components | The firm, Stryker Orthopaedics, notified Domestic locations of this action by e-mail on August 6, 2012 and foreign locations of this action by e-mail on August 7, 2012. Stryker follow-up by sending an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" dated August 17, 2012 to their consignees/customers via Fed Ex with return receipt. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine and return affected product to Attn: Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, 07430 (attach fluorescent orange sticker and mark box with words "Product Recall"), or contact your Stryker Sales Representative to arrange for return of the product ,and complete and return the attached Product Recall Acknowledgment Form within 5 days via fax to 855-251-3635.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100. |  | 72 | Approximately 72 units | Worldwide distribution: USA (nationwide) and countries including: Canada, Belgium, Finland, Germany, Greece, Hong Kong, Ireland, Italy, Japan, Poland, Portugal, Spain, Turkey and UK. | Terminated | April 10, 2014 | 601 | 430
Z-0382-2013 | 63386 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SPECIALTY AUDIBLE TORQUE WRENCH, XIA |  |  | August 09, 2012 | 2012 | Stryker Spine | Multiple complaints have reported the hex tip of the torque wrench fracturing during final tightening. | Nonconforming Material/Component | Stryker Spine sent an Urgent Medical Device Removal Notification letter/Product Accountability Forms dated August 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.For questions regarding this recall call 201-760-8150. |  | 28 | 28 units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Canada, France, Germany, Netherlands, Japan, Spain, Denmark, Switzerland, and the United Kingdom. | Terminated | September 30, 2015 | 1147 | 101
Z-0389-2013 | 63066 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Product 100789502202 VerSys Femoral Head Provisional 22mm +0 |  |  | October 10, 2010 | 2010 | Zimmer, Inc. | Surgical technique instructions for the Versys and Slotted Provisional Femoral Heads have been enhanced to inform users that the head may become difficult to visually locate if it becomes displaced deep or medial to the iliopsoas tendon and provide non-invasive imaging methods that can be used to locate the head. To reduce incidence of sub-optimal fit between provisional femoral heads and the mat | Device Design | The firm, Zimmer, sent notification by email to consignees beginning October 10, 2012 with a follow-up letter titled "URGENT MEDICAL DEVICE CORRECTION" dated October 16, 2012 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to complete a spreadsheet of current users and contact information and immediately provide this information via email to: CorporateQuality.Postmarket@zimmer.com, and review and become familiar with the content and recommendations of the attached letter sent to the customers. For questions call 1-877-946-2761. |  | 128946 | 128,946 all lots, all sizes | Worldwide distribution: USA (nationwide) and countries of: Asia, Barbados, Canada, Europe, Mexico, Puerto Rico, San Juan, and South America. | Terminated | April 30, 2014 | 1298 | 128946
Z-0386-2013 | 63126 | No | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trilogy II Constrained Liner |  |  | August 27, 2012 | 2012 | Zimmer, Inc. | Trilogy Longevity Constrained Liner devices were packaged in incorrect cartons. The product is packaged in the incorrect carton. Product labels and patient labels are correct, matching the product inside the carton. | Labeling Change Control | Zimmer contacted their customer beginning the week of August 27, 2012, via telephone. Zimmer informed the customer the reason of the recall. The customer was instructed to quarantine and return the device to Zimmer Product Service, 1777 W. Center St. Warsaw, IN 46580. The customer was asked to provide information if the product was further distributed. For further questions please call 574-372-4316. |  | 1 | 1 | USA (nationwide) including the states of Ohio, Carolinas and New England. | Terminated | August 27, 2013 | 365 | 1
Z-0406-2013 | 63596 | No | JAD | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | INTRABEAM Balloon Applicator Set; |  |  | October 25, 2012 | 2012 | Carl Zeiss Meditec, Inc. | A manufacturing defect has been identified which could results in the small tube becoming dislodged between the luer fitting and the check valve. | Nonconforming Material/Component | Carl Zeiss Meditec sent an Important Field Corrective Action letter dated October 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately return the affected product to Carl Zeiss Meditec, Inc. A Carl Zeiss Meditec representative will contact the customer to schedule a visit to support the removal of the affected product. For questions customers were instructed to call 925-580-5371 or by email at v.venturo@meditec.zeiss.com.For questions regarding this recall call 925-557-4616. |  | 36 | 36 total, all lots | Nationwide Distribution including IL and MD | Terminated | November 20, 2012 | 26 | 36
Z-0409-2013 | 62805 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | RotoRest Delta Kinetic Therapy System |  |  | June 27, 2012 | 2012 | KCI USA, Inc. | The firm is initiating a voluntary medical device correction for the RotoRest Delta Therapy System head pack. KCI has received complaints alleging patients experiencing skin breakdown on the head and face due to the positioning of the head pack. In a few instances, the patient received injuries ranging from Stage III to Stage IV pressure ulcers. | Component design/selection | KCI sent an Urgent - Voluntary Medical Device Correction letter dated June 27, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that KCI would contact them soon to coorinate a shipment of a replacement of the affected product. For any questions customers were instructed to call 1-800-275-4524, Selection Option 3 followed by Option 2.For questions regarding this recall call 1-800-275-4524. |  | 282 | 282 units | Nationwide Distribution including AL, CA, CO, OK and GA. | Terminated | April 22, 2013 | 299 | 282
Z-0401-2013 | 63557 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | BioPlant, a dental bone grafting material. BioPlant |  |  | October 25, 2012 | 2012 | Kerr Corporation | The firm initiated the recall becaue disturbing the BioPlant product soon after it is placed may halt or delay the ossification process. | Labeling design | Kerr Corporation sent an Urgent Medical Device Correction letter and acknowledgement form dated October 25, 2012, to all affected customers via USPS First Class. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to: "...review the table with product description , part numbers and lot numbers to determine if they have any of the affected product in their inventory and confirm that they were aware of the additional placement instructions necessary to achieve desired ossification results detailed in the paragraph in the letter. Customers were instructed to contact Kerr Customer Care at 1Âż800Âż537Âż7123 if they have any questions. Customers were also instructed to complete and fax back the enclosed Acknowledgement Form to 714-288-4609 in order to confirm their receipt of this correction regardless of whether they have any affected product in their inventory.For questions regarding this recall call 714-516-7400. |  | 22410 | 22,410 units | Worldwide Distribution - USA (nationwide) Canada, Europe, Australia and ROW consignees. | Terminated | September 23, 2013 | 333 | 22410
Z-0403-2013 | 63504 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | qube Compact Monitor (Patient Monitor) 91390. The monitor is 26.2 cm x 31.5 cm x 13.2 cm and weighs 4.1 kg (10.3 inches x 12.4 inches x 5.2 inches - 9 lbs). It has a 26.2 cm (12.1 inches) LCD display with 1024 x 768 resolution. |  |  | October 19, 2012 | 2012 | Spacelabs Healthcare, Llc | Spacelabs Healthcare has learned through severalreports that the input circuit may fail and the monitorwill go on battery mode. The monitor will functionnormally until the batteries become depleted and thenthe monitor will shut down. In addition the monitor mayfail to turn on due to an internal short. | Nonconforming Material/Component | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated October 19, 2012. The letter identified the product, problem, and actions to be taken by the customers. Contact Spacelabs Healthcare at 1-800522-7025 option 2 for questions regarding this notice. |  | 742 | 742 total devices distributed; 530 with the U.S. and 212 Internationally. | Worldwide Distribution.On 4/16/13 the firm advised FDA that they added serial number 1390-000055 and 1390-000913. The two additional serial numbers were added because they were provided as replacements to a French customer whose devices had been returned to support the firm's investigation. The customer had already received the Urgent Medical Device correction, customer letter and required Spacelabs to provide these devices with the understanding that they would require the field corrective action | Terminated | May 22, 2013 | 215 | 742
Z-0407-2013 | 63532 | No | KAK | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Vital Signs Disposable General Purpose 9 French Temperature Probe |  |  | August 17, 2012 | 2012 | Vital Signs Colorado Inc. | There is a potential for epistaxis during use of the Disposable General Purpose 9 French Temperature Probe. | Device Design | The firm,GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter and Confirmation form on August 17, 2012 via Fed Ex to its consignees/customers. The letter describes the product, problem and action to be taken. The customers were instructed to do the following: Do not use the Disposable General Purpose 9 French Temperature Probe P/N 1016 or 1016EU; return all unopened product(s) to Vital Signs Devices (VSD) 11039 East Lansing Circle, Englewood, CO 80112-note: replacements and/or refunds discussed in letter are available only in exchange for the return of any unopened product(s); complete and return the URGENT MEDICAL DEVICE CORRECTION CONFIRMATION via fax to: 800-535-7923; forward a copy of the letter to any institutions they may have forwarded the product and for customers using the GXXXXX64 kit that includes the Disposable General Purpose 9 French Temperature Probe follow the directions listed in the letter.Note: VSD is removing p/n 1016 and 1016EU, Disposable General Purpose 9 French Temperature Probe, from the market and offering a replacement. If you have any questions or concerns regarding this notification, contact Customer Service at +1-800-932-0760 (Domestic &amp; International). Hours of Operation 8:00 am EST to 6:00 pm EST. |  | 243685 | 243,685 units | Worldwide distribution: USA (nationwide) including states of: AR, CA, FL,GA, IL, LA, MA, MI, MO, NC, NH, NY, OH, OK, PA, RI, SE,TX, and WA; and countries of: AUSTRALIA, BARCELONA, GERMANY, NEW ZEALAND, NETHERLANDS, SWITZERLAND, and UNITED KINGDOM. | Terminated | June 14, 2013 | 301 | 243685
Z-0402-2013 | 63512 | No | EBA | Dental operating light. | Dental | Dental | 510(K) Exempt | Freedom Cordless LED Light System. Freedom System, Standard Length, Part Number 910863-1, Serial Numbers 77001000-77003470 and Freedom System, Long Length, Part Number 910863-2, Serial Numbers 77001000-77003470. |  |  | November 17, 2011 | 2011 | Kerr Corporation | The firm recalled the Freedom Cordless LED Headlight because some units were assembled with an incorrect component in thecontrol board electronics that may cause some Freedom Cordless LED Headlight units to become warm to the touch; however, no skin damage will occur. | Labeling mix-ups | Orascoptic sent an "URGENT: MEDICAL DEVICE RECALL" letter dated November 17, 2011. The letter identified the product, problem, and actions to be taken by the customers. An Acknowledgement Form was included for customers to complete and return to the firm. Contact Orascoptic Customer Care at 608-831-2555 for assistance regarding this notice. |  | 1330 | 1330 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, Australia, and Europe. | Terminated | November 20, 2012 | 369 | 1330
Z-0410-2013 | 63122 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | MultipolarÂż Bipolar Cup Liner 22 mm LD.; |  |  | August 31, 2012 | 2012 | Zimmer, Inc. | Zimmer Inc. is initiating a removal of the MultipolarÂż Bipolar Cup Liner 22 mm J.D. (item 00-5001-040-22) due to reports that the split poly ring was incorrectly assembled. | Process design | Zimmer sent an Urgent Medical Device Recall letter to all affected consignees during the week of September 3, 2012. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately, quarantine the recalled product and return remaining inventory. Customers were advised to complete Inventory Return Certification form attached. |  | 195 | 195 | Worldwide Distribution - USA nationwide including the states of: AK, CA, IL, FL, HI, MI, MO, NC, NJ, NY, OH, PA, TX and the countries of Germany, Brazil, Canada, Japan, Taiwan, Thailand, India and Korea. | Terminated | September 23, 2013 | 388 | 195
Z-0411-2013 | 63629 | No | ETD | Tympanostomy tube. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | PARPARELLA-TYPE VENT TUBE, 1.02 MM, SILICONE, REF 240044, QTY 6, Rx Only, STERILE EO, GYRUS ACMI INC., |  |  | October 17, 2012 | 2012 | Gyrus Acmi, Incorporated | One lot of tympanostomy tubes were manufactured with silicon that was found to be contaminated with ethylene vinyl acetate copolymer. | Material/Component Contamination | The firm, Olympus, sent an "URGENT: Medical Device Recall" letter dated October 15, 2012 to their consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to inspect their stock to confirm if any of the affected lot of product remains in their possession; immediately cease any further use of any affected product; remove it from their stockroom and quarantine it until it is shipped back to Gyrus ACMI, Inc.; call their Gyrus ACMI customer service representative (800-773-4301) to obtain a Returned Goods Authorization and promptly return any affected product to Gyrus ACMI, Inc., 2925 Appling Rd, Bartlett, TN 38133., and complete and return the RECALL REPLY FORM to Attn: Regulatory Affairs via fax to: (508) 804-2624, even if there is no affected stock left at your facility.If you have any questions, contact the customer service at 800-773-4301. |  | 600 | 10 boxes/60 units | Worldwide distribution: USA state of Kansas and country of Korea. | Terminated | April 09, 2013 | 174 | 600
Z-0408-2013 | 62835 | No | KLS | Theophylline test system. | Toxicology | Toxicology | 510(k) | Dimension Vista(R) Systems Vista(R) Theophylline THEO Flex(R) Regent Cartridge (K4071) |  |  | August 09, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has received complaints from customers of "Abnormal Reaction" errors occurring on calibrations, QC and patient samples. The Abnormal Reaction flag is being triggered by an internal check in the method for atypical particle reagent aggregation. | Nonconforming Material/Component | Siemens Healthcare Diagnostic sent a Urgent Field Safety notice dated July, 20. 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please immediately discontinue use and discard any remaining inventory of Dimension VistaÂż Theophylline FlexÂż reagent cartridge (THEO, Catalog # K4071) lot 12135AE and/or 12171BB.Please complete the attached form and fax it to (302)631-8467 to indicate that you have received this information and to facilitate your additional no charge replacement.We ask that you forward this notification to anyone to whom you may have distributed this lot.If you have technical questions or concerns, please contact the Siemens Technical Solutions Center at 800-441-9250 for further assistance. For questions concerning the status of your no charge replacement order please contact your Siemens Customer Service Center at 888-588-3916.We apologize for the inconvenience that this situation has caused. Thank you for your patience and continued support. |  | 959 | 959 cartons | USA (nationwide) including the states of AL, AZ, CA, CO, DC, DE, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, MS, MT, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, VA, VT, WA and WI. | Terminated | January 08, 2014 | 517 | 959
Z-0422-2013 | 63580 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | KING LAD Standard ClearSeal SINGLE USE, King Systems ETO Sterilie10 devices/case |  |  | November 02, 2012 | 2012 | King Systems Corp. | The firm received two complaints from a user facility in New Jersey regarding color-coded EtO sterilization indicators on the primary device packaging that did not undergo the expected color change. The complainant reported devices that had pink color indicators when the product packaging specified that the color should change from pink to yellow (or brown) if the product had subjected to an EtO | Nonconforming Material/Component | King Systems sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated October 31, 2012 to all affected consignees via email. The letter identified affected product, problem and actions to be taken. The letter states that further use and/or distribution of any remaining product affected by this recall should cease immediately. The letter instructs customers to destroy all of the affected products and indicate the destruction by signing and returning the attached forms using the enclosed self-addressed envelope. For questions contact your King Systems customer service representative at 800-642-5464. |  | 263 | 263 cases of product. (2,630 devices) | US Nationwide Distribution - WI, TX, NY, MO, IN, KY, FL, NJ, PA, TN, and NC | Terminated | May 28, 2013 | 207 | 263
Z-0412-2013 | 63144 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Spine Frontier Inspan compressors |  |  | August 28, 2012 | 2012 | SpineFrontier, Inc. | Inspan Compressors may break at weld causing a loss of compression | Device Design | Spine Frontier issued e-mail notices to customers on May 14, 2012 advising of the potential for breakage. Units were replaced with weld free design. |  | 17 | 17 units | Worldwide Distribution-USA (nationwide) including PR and the states of FL, MA, MO, OK, and TX, and the country of Mexico. | Terminated | June 11, 2015 | 1017 | 17
Z-0413-2013 | 63633 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | CapLOX II Final Torque Driver, Indicator is a Pedicle screw spinal system. |  |  | September 20, 2012 | 2012 | Captiva Spine, Inc | Captiva Spine recalled their CapLOX II Final Torque Driver, Indicator due to the device prematurely breaking | Nonconforming Material/Component | Captiva Spine contacted the customer by phone to have them return the device.For questions regarding this recall call 561-277-9480. |  | 10 | 10 | Nationwide Distribution (Distributed to one customer located in MN) | Terminated | January 11, 2013 | 113 | 10
Z-0416-2013 | 63541 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Trilogy Acetabular System Shell with Cluster Holes Porous 56 mm O.D. Sterile |  |  | September 20, 2012 | 2012 | Zimmer, Inc. | Zimmer TrilogyÂż Acetabular System Shell with Cluster Holes, Porous,56 MM O.D., sterile; is missing polar hole threads. Zimmer received one complaint in which no injury was reported. | Process control | Zimmer sent a Urgent Medical Device Recall Notice dated September 20, 2012, to all affected customers. The letter identified the products the problem and the action needed to be taken by the customer. Customers were instructed to locate quarantine and return affected product to Zimmer Product Service Department, 1777 West center St. Warsaw, IN 46580 for product in the USA. Customers were asked to complete the Inventory Return Certification Form and follow the instructions where to return this form.For further questions please call (574) 372-4807. |  | 95 | 24 USA; 71 worldwide | Worldwide Distribution--USA (nationwide) including the states of AL, CA, CT, KS, MA, MD, MN, MO, NY, SD, TX and WI and the countries of Bulgaria,Canada, Germany, Italy, Netherlands, Russian Fed, Sweden and UK | Terminated | April 18, 2014 | 575 | 95
Z-0428-2013 | 63500 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | IMRIS ORT100 (Part Number 110470-000) and ORT200 (Part Number 112216-00), are tables used in both the Neuro II-SE and Neuro III-SV Magnetic Resonance Imaging Systems. |  |  | October 19, 2012 | 2012 | Innovative Magnetic Resonance Imaging Systems Inc. | IMRIS Inc. initiated a recall of their OR Tables ORT100 and ORT200. When proper procedure is not followed, the foot end of the IMRIS intra-operative OR table may be attracted to the MRI magnet as the magnet approaches the table. This poses a potential risk of injury to both the patient and operators standing near the table. | Labeling design | Consignees were sent on 10/19/12 a IMRIS "Urgent Medical Device Correction" letter dated October 16. 2012. The letter described the product and the problem involved in the recall. It also provided changes to the Operator's Manual and labeling which addresses the issue. Customers may contact 1-866-475-0525. |  | 23 | 23 | Nationwide Distribution, including the states of FL, GA, IN, MD,MA, MI, MN, MO, TN, UT, and VA. | Terminated | December 11, 2012 | 53 | 23
Z-0429-2013 | 63539 | No | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | Minotrol 16, Model Numbers: 2043748, 2043752, 2043756, Lot Number: MX090(Levels: Low, Normal, and High). |  |  | September 14, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical Irvine Technical Support received reports from some customers that have experienced high recover values for Hemoglobin when using Minotrol 16 lot #MX090 (Level High). Customer communication was sent 10/01/2010. | Nonconforming Material/Component | HORIBA Medical sent a Product Information Notice letter dated November 5, 2010 via Fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the instructions included with the Product Information Notice. For questions customers were instructed to call the Technical Support Hotline at 1-888-903-5001.For questions regarding this recall call 949-453-0500, ext 208. |  | 22954 | 22,954 | Nationwide Distribution | Terminated | April 22, 2013 | 220 | 22954
Z-0414-2013 | 63516 | No | FSP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The brand name of the device is Laser Loupes, a dentalloupe |  |  | April 10, 2012 | 2012 | Kerr Corporation | The firm initiated the recall for Laser Loupes because the ink used to mark which laser frequency each laser Loupe protectsagainst may wipe off. The loupes work according to specification, but the marking may not permanently adhere tothe device. | Finished device change control | Customers were notified via USPS 1st class mail on April 10, 2012 to US and Canadian, and Rest of World (ROW) consignees, and on April 27, 2012 to European consignees. The notification letter identifies the product, reason for recall, and instructions to customers:"Please review the table below to verify that the laser loupes being used in your office has the correct filter color. ORASCOPTIC KINDLY REQUESTS YOUR COOPERATION IN COMPLETING AND FAXING BACK THE ENClOSED ACKNOWLEDGEMENT FORM IN ORDER TO CONFIRM YOUR RECEIPT OF THIS FIELD CORRECTION.If you are an authorized Orascoptic distributor, please inform your affected customers of this notification within forty-eight (48) hours of receipt.". Customer contact information was also provided:"Please contact Orascoptic Customer Care at 608.831.2555 or toll free in the United States and Canada at 1-800369-3698 for more information.". |  | 639 | 639 units | Worldwide Distribution--USA (nationwide) and the countries of Canada and EU. | Terminated | November 21, 2012 | 225 | 639
Z-0415-2013 | 63613 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | Ranawat/Burnstein Acetabular Shell 54mm/Limited Holes/Porous Coated is a metallic, full-hemisphere, acetabular shell component. |  |  | October 04, 2012 | 2012 | Biomet, Inc. | The color code on label conflicts with the liner size (liner size 24 should be green, not yellow) | Labeling Change Control | BIOMET called distributors who received the affected product to return it to Biomet with follow-up Urgent Recall Notices sent via E-mail October 22, 2012.. The notice included the identity of the recalled device, reason for the recall, and directions to complete the FAX back form and return it to 574-372-1683. Product is to be returned to Return Goods, Recalls, Biomet, Inc.56 East Bell Drive, Warsaw, IN 46580. Questions related to the recall should be directed to Field Action Coordinator (574) 372-1570, Monday through Friday. 8 a.m. to 5 p.m.For questions regarding this recall call 574-267-6639, ext 1676. |  | 16 | 16 (5 USA) | Nationwide Distribution including CA, IN, NC, OH, and UT. | Terminated | March 18, 2013 | 165 | 16
Z-0417-2013 | 63075 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | S-Lift Implant Guides (also known as S-LIF) |  |  | August 22, 2012 | 2012 | SpineFrontier, Inc. | When using the optional S-LIFT (also known as S-LIF) Implant Guides without stops, there is a known risk of moving the Implant Guides, and possibly the implant, too far into the disc space creating the potential for injury to the patient. The instructions for Use for the S-Lift Implant Guides will be amended to add the following warning: "Care is to be taken to not over insert the Implant Guides | Labeling design | Spine Frontier issued a Correction Notice Letter August 22, 2012 to all affected customers to update the Instructions for Use. The letter identified the affected product, problem and actions to be taken. Customers were Instructed to remove the Instructions for Use they currently have and replace them with the update. The letter instructs customers to complete and return the Response Form to 978-232-3990. |  | 19 | 19 units | US Nationwide including the states of: MA and TX. | Terminated | June 11, 2015 | 1023 | 19
Z-0430-2013 | 63479 | [6 empty] | 0.9% Sodium Chloride Injections, USP, 1000ml, VWR/Baxter |  |  | October 17, 2012 | 2012 | Biomerieux Inc | The firm inventoried and distributed 0.9% Sodium Chloride Saline bags as 0.45% Sodium Chloride Saline bags for use with the VITEK 2 system. | Employee error | Biomerieux sent an Urgent Product Removal Notice letter dated October 16, 2012 to all affected customers via certified mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the affected product and immediately discard any remaining inventory. The letter included instructions on receiving product replacement or credit. For questions or concerns contact your local BioMerieux representative. |  | 199 | 199 cases/14-1000mL bags | Worldwide Distribution - USA (nationwide) including the states of: AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, PA, RI, TN, TX, VA and WI. Distribution was also made to Canada, military facilities and other foriegn countries. | Terminated | December 05, 2013 | 414 | 199
Z-0432-2013 | 63554 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 4.5mm PLEO Osteotomy Guide. |  |  | October 18, 2012 | 2012 | OrthoPediatrics Corp | Complaint received from sales representative that the drill guide tubes would not screw into replenishment instrument. Upon investigation, it has been determined that the device was not manufactured to design specifications and will not function as intended. | Nonconforming Material/Component | OrthoPediatrics sent an "URGENT RECALL" e-mail of the Trochar Scalpel Instrument to all affected customers. The e-mail identifies the product, problem, and actions to be taken by the customers. The e-mail states that the instrument has been manufactured incorrectly and will not work (function as intended). The threads in the instrument are faulty and the guide tubes will not thread in correctly. Customers will receive follow up information. |  | 24 | 24 | Nationwide Distribution-including the states of OH, NC, CO, FL, IN, MI, and TX. | Terminated | September 18, 2013 | 335 | 24
Z-0434-2013 | 63631 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | SenTec V-Sign Membrane Changer Insert (VS-MCI), 5 pieces. The SenTec V-Sign Membrane Changer Insert is indicated for use with the V-Sign Membrane Changer only. |  |  | November 05, 2012 | 2012 | Sentec AG | A manufacturing defect was observed for the membrane changer insert that does not allow electrolyte to be present under the new membrane. The PCO2 performance of a sensor without electrolyte will be deteriorated. | Component design/selection | SenTec notified the one US distributor via e-mail on November 5, 2012. The distributor was advised to stop distribution of affected lot numbers. The communication to the distributor described the issue, identified affected product, and stated that SenTec will replace current stock at the end-customer level. The end user can discard the affected products. The distributor is asked to respond back with an e-mail to SenTec that they took appropriate action. |  | 265 | 265 boxes (5 per box) | Distributed only in MO. | Terminated | December 18, 2012 | 43 | 460
Z-0431-2013 | 62834 | No | LMF | Absorbable hemostatic agent and dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | Styker Orthobiologics Vita Plasma Separator, a component of Vitagel Surgical Hemostat. |  |  | July 26, 2012 | 2012 | Stryker Orthobiologics, Inc. | When preparing plasma using VitaPrep, it was reported that a small amount of blood collected in the bottom of the centrifuge cup following the centrifugation step. Firm has since discovered that if blood is present in the Luer connection of the VitaPrep following the blood draw step, this blood could be displaced to the outside of the syringe when the VitaPrep cap is reattached to the syringe. T | Device Design | Stryker sent a "MEDICAL DEVICE FIELD CORRECTION" dated July 26, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter included a Business Reply Form for customers to complete and return via fax to 1-484-323-8803. Contact the firm at 1-888-774-8870 for questions regarding this notice. |  | 3268 | 3,268 | Nationwide Distribution-including the states of AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, PA, SC, SD, TN, TX, VA, WA, WI, and WV | Terminated | October 29, 2014 | 825 | 3268
Z-0437-2013 | 63544 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen CompleteKnee Solution Stemmed Nonaugmentable Tibial component |  |  | September 21, 2012 | 2012 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the NexGenÂż Stemmed Nonaugmentable Tibial Component due to the devices being processed through a manufacturing cleaning operation that was operating outside of the validated parameters.As a result, the devices may contain residual particulate from the manufacturing process.Risks include: Allergic reaction due to the foreign material or late onset | Process change control | Zimmer sent an Urgent Medical Device Recall letter dated September 21, 2012, to all affected customers. The letter identifiedthe product, the problem, and the action to be taken by the customer. Customers were instructed to locate and quarantine the affected product. Customers were to return the affected product along with the completed Inventory Return Certificate Form. If product was further distributed the Recall Letter should be forwarded to their customers. Product was to be returned to USA accounts - Zimmer Product Service Department, 1777 West Center St. Warsaw, IN 46580. If product has been implanted, affected surgeons will be notified. For questions regarding this recall call 574-372-4807. |  | 123 | 123 worldwide (33 USA) | Worldwide Distribution - USA including AZ,CA, FL, IA, IL, IN, MO, NE, NJ, NY, OH, PA, WA. Internationally to Thailand, Canada, Germany, France, UK, Hungary, Netherlands, Russian Fed, and Sweden, | Terminated | April 17, 2014 | 573 | 123
Z-0440-2013 | 63681 | No | LKF | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical Advincula Arch , Reusable Uterine Manipulator Handle |  |  | October 31, 2012 | 2012 | Cooper Surgical, Inc. | Excessive pressure exerted on the uterine tip when attached to the Arch may, under certain circumstances, cause the internal steel rod of the tip to protrude through the cleaning hole on the shaft of the handle | Device Design | CooperSurgical issued notification on October 31, 2012, to all affected customers Fed X . The notification identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return units for replacement. Customers with questions should contact Customer Serivce at 1-203-601-9818 or 800-243-2974. Customers were asked to fill out the attached Replacement of Recalled Product Form and fax it to 1-800-262-0105.For questions regarding this recall call 203-601-9818. |  | 1215 | 1215 total | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Canada, Czech Republic, Germany, Italy, Spain, France, United Kingdom, Japan, New Zealand, South Korea, Saudi Arabia, Sweden, Singapore, Slovenia, and Taiwan. | Terminated | March 04, 2014 | 489 | 1215
Z-0438-2013 | 62918 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | IO FiX X-Post Reamer, Extremity Medical, www.extremitymedical.com; Customer Service: 888.499.0079 |  |  | April 25, 2011 | 2011 | Extremity Medical LLC | Product Event Reports revealed a broken reamer experienced intra-operatively. Pre-drilling was not performed prior to reaming. | Employee error | Extremity Medical sent a notification by e-mail that describes the "Slight modification to the Surgical Technique" on April 25, 2011. For questions regarading this recall call 973-588-8980. |  | 156 | 156 units | Worldwide Distribution - USA including FL, MA, NC, CA, MD, VA, IN, AL, NJ, AZ, OK, MT, NY, OH, IA, MI, TX, WI, GA, IL, TN and NE. Internationally to Switzerland. | Terminated | March 21, 2013 | 696 | 156
Z-0436-2013 | 63543 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | 220/240V AC, Advanced Perfusion System Platform (APS) |  |  | November 14, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular System (TCVS) has received reports of a situation where usersexperienced a total loss of functionality for some System 1 units. The reports indicate that theunits went blank and shut down with no sign of power and battery backup did not initiate. Theresult is all pumps stop, with no safety system functionality, and the battery would not beactivated. The user would be | Device Design | Terumo Cardiovascular Systems sent a Urgent Medical Device Recall Correction letter dated November 14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Terumo CVS is alerting all users of Terumo System 1 of the reports of malfunction. Customers were instructed to: 1. Review this Medical Device Safety Advisory.2. Assure that all users are aware of this notice.3. Confirm receipt of this communication by faxing, or emailing the attached Customer Response Form to the fax number/email address indicated on the form.We encourage you to contact us with any questions or concerns:Terumo CVS Customer Service 1-800-521-2818, Recall Fax 1-734-741-6149Customer Service Hours: Monday thru Friday, 8 AM - 6 PM ET |  | 1647 | 1647 units total | Worldwide Distribution--USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | August 30, 2013 | 289 | 1647
Z-0444-2013 | 63583 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | Pentron Clinical Metrex Research; Pinnacle Cure Sleeve, a protective barrier |  |  | September 20, 2011 | 2011 | Metrex Research, LLC. | The firm initiated the recall due to a labeling error. The boxes labeled as Pinnacle Cure Sleeve 4511, which fits 11mm light guides, actually contain Pinnacle Cure Sleeves that fit 7mm or 8mm light guides (part number 4500). | Error in labeling | The firm, Sybron Dental Specialties, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 20, 2011 to its consignees/customers. The consignees/customers were instructed to determine if they have any of the affected product in their inventory; if they have the affected product to contact Metrex Research Customer Care at 1-800-841-1428 to receive an RMA number; return any affected product in their inventory; and complete and return the Recall Response Form via mail to: Metrex Research, 29210 Wick Road, Romulus, Michigan 48174, Attn: QA Lab or fax to: 734-947-6725.If you have any questions regarding this recall, you can contact our Customer Care Center at 1-800-841-1428. |  | 3348 | 3,348 units | Worldwide distribution: USA (nationwide) and countries of: Australia, Canada, and Europe. | Terminated | November 29, 2012 | 436 | 3348
Z-0442-2013 | 63711 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | Echo Bi-Metric Press-Fit Stems/Echo Bi-Metric 11mm x 135mm Collarless Porous Stem (Orthopedic Implant) |  |  | November 06, 2012 | 2012 | Biomet, Inc. | It was reported that the surgeon noticed that the Echo 11mm hip stem had an etching indicating it was 13mm stem. | Labeling Change Control | The firm, Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated November 8, 2012 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to take the following actions:Âż Immediately locate and remove the identified device(s) listed below from circulation.Âż Carefully follow the instructions on the enclosed FAX Back Response Form .Âż Fax a copy of the Response Form to 574-372-1683 prior to return of product.Âż Use priority carrier for your shipment.Questions related to notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 4 | 4 | Nationwide distribution: USA including states of: AZ, MD, and KS. | Terminated | August 28, 2013 | 295 | 4
Z-0446-2013 | 63675 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The brand name is MRidium Series 1000 MR Infusion Sets, Type 1058 MR IV Extension Set, Lot Code LBG001. The generic name is infusion administration set. |  |  | August 24, 2012 | 2012 | Iradimed Corporation | Iradimed Corporation is recalling MRidium Series 1000 MR Infusion Sets, Type 1058 MR IV Extension Set, Lot Code LBG001 because a section of the device does not meet specifications. | Nonconforming Material/Component | Iradimed Corporation contacted all affected customers directly by phone, and followed up with an email notification. Customers were instructed to return the product to Iradimed Corporation. For questions regarding this recall call 407-677-8022. |  | 719 | 719 sets | Nationwide Distribution including AZ, CA, FL, KY, MD, MA, NJ, NY, PA, and TX. | Terminated | July 19, 2013 | 329 | 719
Z-0449-2013 | 63212 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima XR 220amx 15kw Mobile X-Ray System |  |  | August 01, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue on the interface of the mobile radiographic product, Optima XR22Oamx and Optima XR200amx with Digital Upgrade. | Radiation Control for Health and Safety Act | GE Healthcare will without charge remedy the defect or bring the product into compliance with each applicable Federal and IEC standard in accordance with a plan to be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication through a GE Healthcare field engineer site visit. Field Modification Instructions (FMI) 10868 describes the rework plan the GEHC Engineers will make to the imaging system compliant with the applicable performance standard.This CAP appears to adequately address the problem and is hereby approved.Call Center phone number: 800-437-1171 (United States). For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. |  | 217 | 217 total Installed in the US | Nationwide distribution | Terminated | January 10, 2013 | 162 | 217
Z-0451-2013 | 63649 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Kimberly-Clark Patient Warming System - Model 1000 Control Unit |  |  | September 16, 2010 | 2010 | Kimberly-Clark Corporation | Error 85 alarm on the Model 1000 Control Unit was occasionally triggered inappropriately. | Component design/selection | Kimberly Clark sent an Urgent Field Corrective Action letter dated September 16, 2010, to all affected customers. The letter identified the product, the problem, and the aciton to be taken by the customer. Customers were asked to confirm whether their facility as the affected product by faxing the completed Field Corrective Action Response Form to 920-380-6682. Kimberly-Clark will arrange to perform the upgrade at the customer's facility. Customers with questions should call the Field Corrective Action Coordinator at 770-587-7400.For questions regarding this recall call 770-587-8393. |  | 400 | 400 units | Nationwide Distribution including AZ, AK, CA, CO, DE, DC, FL, GA. ID, IL, IN, IA, KS, KY, LA, MD, ME MA, MS, MO, MT, NE, NJ, NM, NV, NY, NC, OH, OR, PA, RI, SC, TN, TX, VA, WV, and WI | Terminated | December 20, 2012 | 826 | 400
Z-0441-2013 | 63722 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | The Vertex Max Drill Tube is a component within the Navigated Vertex Max Drill Guide Set (item 9733370). The Navigated Vertex Max Drill Guide Set is intended to be used with a Medtronic Computer Assisted Surgery System to help assist with screw placement during spine surgical procedures. The device is reusable and requires sterilization according to product labeling. |  |  | November 13, 2012 | 2012 | Medtronic Navigation, Inc. | Medtronic is recalling certain Vertex Max Drill Tubes which are used with the StealthStation System due to a defect related to the internal threading for component set up. | Process control | Medtronic Navigation sent a Urgent Medical Device Recall letter dated November 13 -14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed do not use the Vertex Max Cannulated Drill Tube clinically. Medtronic Navigation will remove the Vertex Max Cannulated Drill Tube. Medtronic Navigation will provide a replacement Vertex.Max Cannulated Drill Tube manufactured from a different lot which was inspected to verify the defect has been corrected.If you have any questions about this notification, please contact a Sales Representative in your area. For further questions please call (720) 890-3200 Ext. 3303 . |  | 12 | 12 Units | Worldwide Distribution--US (nationwide) including the states of CA, FL, MA, LA, and MI and the countries of Switzerland, Saudi Arabia, Lebanon, Canada and Korea. | Terminated | January 04, 2013 | 52 | 12
Z-0447-2013 | 61835 | No | MXM |  |  | Neurology | PMA | The product has the Catalogue/Part number 60010. |  |  | May 09, 2012 | 2012 | Natus Medical Incorporated | The Olympic Cool-Cap's Control Module has experienced a frozen screen during use. When this occurs, the on-screen information remains on display, but the system is no longer providing cooling treatment to the infants. The visible clock displayed at the upper right hand corner of the screen stops advancing. | Nonconforming Material/Component | The firm, Natus Medical Incorporated, sent its customers an "URGENT: MEDICAL DEVICE" letter dated May 16, 2012 with an attached Field Safety Notice. The letter describes the product, problem and actions taken. Customers can call Natus Technical Service at 888-496-2887 for any questions about this recall. |  | 324 | 324 units | Worldwide distribution: USA (nationwide) including states of: AK, AR, AZ, CA, CO, DE, FL, GA, HI, ID, IL, IN, KY, MA, MI, MN, MS, ND, NE, NJ, NV, NY, OH, OK, PA, SC, SD, TN, TX, UT, WA, Puerto Rico, and countries of: Argentina, Australia, Austria, Bahrain, Canada, Chile, China, Colombia, Ecuador, Egypt, Germany, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Italy, Kuwait, Latvia, Malaysia, Mexico, Morocco, Panama, Puerto Rico, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Terminated | January 27, 2015 | 993 | 324
Z-0443-2013 | 63235 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | Devices includes a Multi Port Inflator or without Multi Port Inflator (MPI). All the devices are vascular blood flow machines. Devices are distributed with pressure cuffs of varying sizes as accessories. |  |  | August 08, 2012 | 2012 | Parks Medical Electronics, Inc | Parks Medical Electronics, Inc. was made aware of two incidents where an air supply tube from their PARKS PVR unit was inadvertently connected to an IV in the patient's arm instead of into the blood pressure cuff. | Employee error | The firm, Parks Medical Electronics, Inc., sent an "IMPORTANT SAFETY ISSUE SERIOUS HEALTH CONSEQUENCE" letter, dated September 10 2012 to their consignees/customers. The letter described the product, problem and actions to be taken. The customers are advised to order LUER UPGRADE KIT WITH MPI if the Parks device includes a Multi Port Inflator (MPI), and/or order the LUER UPGRADE KIT WITHOUT MPI the Parks device does not have a Multi Port Inflator. While waiting for the kit or its installation, consignees/customers should use extreme caution when hooking up the air supply tubes to the cuff's Luer fittings. For ordering your kit, please call 800-547-6427 option 0.For responses, questions or concerns, customer should contact Parks Medical Electronics at 888-356-9522 (Phone), 503-591-9753 (Fax) or send email to support@parksmed.com. |  | 349 | 349 units (Domestic: 270 units - Foreign: 79 units) | Worldwide distribution: USA (nationwide) and countries of: Australia, Brazil, Canada, Chile, Egypt, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Korea, Mexico, Peru, Republica Argentina, Spain, Switzerland, and Venezuela. | Open, Classified |  |  | 349
Z-0445-2013 | 63410 | No | LLL | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | The Is-anti-Sm test kits****Catalog No: 720-240*** |  |  | August 06, 2012 | 2012 | Diamedix Corporation | Diamedix Corporation recalled their Is-anti-Sm Test Kit due to the Positive Control recovering high out of specification. | Nonconforming Material/Component | Diamedix sent a Customer Notification letter to all affected customers via Certified Mail. The letter identifed the affected product. problem and actions to be taken. The letter instructed customers to contact Diamedix Technical Service Department at 1-800-327-4565 or 305-324-2300 to arrange for the return and replacement of any remaining affected product. For questions call 305-324-2314. |  | 65 | 65 Test Kits | Worldwide Distribution - USA Nationwide including the states: CA, ID, NJ, NY, OH, PA, PR, RI, TN, and TX, FL, and the countries of Italy, Chile and Ecuador | Terminated | August 08, 2014 | 732 | 65
Z-0454-2013 | 63634 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | NeptuneÂż 2 Rover - ULTRA ~120V ~ 60 Hz 12A |  |  | May 24, 2011 | 2011 | Stryker Instruments Div. of Stryker Corporation | Beginning May 24, 2011 Stryker Neptune 2 (N2) Waste Management Systems castors were replaced with a new castor design because investigation found the old design had the potential for the casters to bind compromising the Rover's ability to roll. This can result in additional strain on the operator, due to multiple exposures resulting in back pain. | Component design/selection | Stryker Instruments sent a "STRYKER NEPTUNE 2 ULTRA 120V &amp; 230V WASTE MANAGEMENT SYSTEM" letter dated June 22, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return a Reply Form via fax to Stryker at 269-906-9121 or email to sales.service@stryker.com. Technical questions should be addressed to Neptune Field Service at 800-800-4236 ext. 5052. |  | 5865 | 5865 | Worldwide Distribution-USA (nationwide) and Pacific countries, and the countries of EEMEA, Canada, Latin America, and Japan. | Terminated | May 28, 2014 | 1100 | 6019
Z-0464-2013 | 62749 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom Procedural Trays: Merit's Custom Total Hip Pack, K12T-05537, Merit's Total Knee Pack, K12T05538 |  |  | June 08, 2012 | 2012 | Merit Medical Systems, Inc. | Custom Procedural trays contain Stryker Hytrel Togas which are being recalled because the clear tape that is intended to aid in the prevention of patient's fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga, was not applied during the manufacturing process. | Process control | The firm, Merit Medical, sent an "URGENT PRODUCT RECALL NOTICE" letter dated June 8, 2012 to its customers by letter sent by US Postal Service - Certified Mail Receipt describing the problem, instructions as to actions that need to be taken, and requesting the customer to complete and sign the Customer Response Form and immediately fax a copy to this form to 1-804-416-1031. Customers were instructed to mail the completed original Customer Response Form using the provided postage-paid, self-addressed envelope to Merit. Should the customer choose to return the product for rework, they are to return the affected products by shipping them back to Merit Medical Systems via Federal Express 2nd Day Air (Acct #1129-5275-6) to: RGA Department RGA #NC-53417, Merit Medical Systems, Inc., Attn: Manager of Quality Assurance and Regulatory Affairs, 12701 Kingston Ave., Chester, VA 23836. For questions customers were instructed to contact Manager of Quality Assurance and Regulatory Affairs at 1-804-416-1048 or email: jbraxton@merit.com. |  | 83 | 83 packs | Nationwide distribution: US including states of: NE and SD. | Terminated | December 10, 2012 | 185 | 83
Z-0459-2013 | 63407 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | 18 Fr 6mm Sft Flow Angled Wire |  |  | November 21, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | The presence of plastic flash was identified at the tip of certain lots of Sarns Soft-Flow Aortic Cannula. The plastic flash has the potential to detach and contribute to an adverse patient effect. While this potential failure has not been reported from the field as an observed defect, it was determined that the product does not meet specifications and this condition could result in an unacceptabl | Process design | TERUMO sent an URGENT MEDICAL DEVICE RECALL letter dated September 20, 2012, to all affected customers. The letter identified the product, the problem, and the aciton to be taken by the customer. Customers were instructed to remove this cannulae from circulation and return them for credit. Customers with questions should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 5760 | 5760 total | Worldwide Distribution - USA including AL , CA , CO , DC , DE , FL , GA , IL , IN , KS , KY , LA , MA , MI , MO , MS , MT , NC , ND , NE , NH , NY , OH , PA , TN , TX , VA , and WI. Internationally to Australia, United Arab Emirates (UAE), Singapore, Malaysia, BELGIUM, Japan, and CANADA. | Terminated | April 23, 2013 | 153 | 5760
Z-0463-2013 | 63615 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Sekisui Diagnostics, Spectrolyse PAI-1 In-Vitro Diagnostic Use for the quantitative determination of Plasminogen Activator Inhibitor Type 1 in human plasma. |  |  | October 24, 2012 | 2012 | Sekisui Diagnostics Llc | Spectrolyse PAI-1 activity assay, producing lower than expected absorbance values 0 IU/ml standard reducing the slope of the standard curve requiring more frequent sample dilutions. | Nonconforming Material/Component | Sekisui Diagnostics issued an "URGENT PRODUCT RECALL NOTICE" dated October 24, 2012 to all consignees. The letter instructed consignees to immediately stop using the specified lot of product, discard any remaining inventory and return the enclosed response form. Contact the Customer Service Department at 800-637-3375, ext. 10, 12, or 31 for assistance with this recall. |  | 63 | 63 kits | Nationwide Distribution-including the states of FL, IN, NC, OH, UT, and WA. | Terminated | January 08, 2013 | 76 | 63
Z-0465-2013 | 63521 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza Radiological Image Processing System |  |  | June 20, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Firm has become aware of an unintended behavior when using syngo.plaza and issued an Update Instruction. The Update Instruction SY050/12/S provides a software update and addresses the following issues: - Patient Rename messages for the same patient are being continually processed for several days. This fills up the 'pcvdbsrv' log and potentially affects performance of the system. - Data could | Software design | The firm, Siemens, sent "Update Instruction SY050/12/S" on June 20, 2012 to its consignees/customers. The Update Instructions described the product, problem and actions to be taken. It also provides a software update.If you have any questions call 610-219-4834 or email: anastasia.mason@siemens.com. |  | 15 | 15 | Nationwide distribution: USA including states of: CA, FL, GA, MA, MN, MO, NJ, NY, OH, and WI. | Terminated | January 14, 2014 | 573 | 15
Z-0453-2013 | 63642 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | OmniGuide BP-ROBOTIC ( BeamPath Robotic), Single Use |  |  | November 05, 2012 | 2012 | OmniGuide, Inc. | Failure of glue joint, allowed a portion of the single use device to shift inside, may separate and fall into the patient | Process control | OmniGuide issued letter notifiction on November 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that an OmniGuide sales representative would contact them immediately to recover the affected product. For any questions customers were instructed to call 617-551-8561. For questions regarding this recall call 617-551-8410. |  | 45 | 45 units | Nationwide Distribution including AZ, FL, MA, NY, OH, and OR. | Terminated | October 16, 2014 | 710 | 45
Z-0452-2013 | 63429 | No | JIY | Iron (non-heme) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista(R) IRON Flex(R) reagent cartridge (K3085) |  |  | October 08, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed the potential for falsely elevated results with IRON Flex(R) reagent cartridges when tests are processed from a well set contaminated with iron from environmental sources. All lots of IRON Flex(R) reagent cartridges have the potential for this issue. If a well set is contaminated with trace amounts of iron, the entire well set is impacted, and the mag | Material/Component Contamination | The firm intiated their recall of the specified lots on October 8, 2012 by issuing a letter to all consignees informing them of the problems with the reagent cartridges. Firm recommended batching the processing of the tests on one instrument and provided instructions to disable normal flow of IRON and to run the QC on each IRON Flex(R) reagent cartridge well set. In addition, a a reply form and Frequently Asked Questions attachment were enclosed with the communication.For questions regarding this recall call 302-631-6311. |  | 19310 | 19,310 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, Washington DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and Internationally to Canada, AU, NZ, JP, DE, SA, and MY | Terminated | January 14, 2014 | 463 | 19310
Z-0470-2013 | 63442 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics |  |  | July 27, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Offset Bushings associated with the lots identified were mismarked: the rotational reference numbers on the Tibial Offset Bushing were marked counterclockwise instead of clockwise; the Femoral Offset Bushing did not include rotational markings. | Error in labeling | Stryker Orthopaedics sent an "URGENT PRODUCT RECALL" letter dated July 27, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Product Recall Acknowledgement Form was attached for customers to complete and return via fax to 201-831-6069. Contact the firm at 201-972-2100 for assistance with this recall. |  | 8 | 8 | Worldwide Distribution. | Terminated | September 30, 2015 | 1160 | 15
Z-0472-2013 | 63612 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | OSS Reamer Sleeve REF132-472662 KNEE REAMER SLEEVE 13.0 thru 20.0 STAINLESS STEEL |  |  | November 01, 2012 | 2012 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that the thin shaft reamer may get stuck and not slide into the sleeve. | Nonconforming Material/Component | The firm, Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated November 1, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately locate, discontinue use of and remove the product from circulation; and return recalled devices with directions to distribute notices, if product was further distributed; and to complete and return the enclosed FAX Back Response Form within three (3) business days to FAX # 574-372-1683. Distributors were charged with locating and returning products to Biomet ATTN: Return Goods-Building B RECALLS RGA#: _, Biomet, Inc., 56 East Bell Drive, Warsaw, IN 46580. Questions should be directed to the Field Action Coordinator at (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m |  | 36 | 36 | Worldwide distribution: USA (nationwide) and countries of: Japan and Netherlands. | Terminated | April 29, 2013 | 179 | 82
Z-0467-2013 | 63646 | No | ISW | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | BiOM Rechargeable Li Ion 22.2 V Battery and BiOM Lithium Battery Charger sold as an accessory to the BiOM lower leg system: |  |  | October 26, 2012 | 2012 | iwalk inc | Battery may over-heat during charging and smoke | Unknown/Undetermined by firm | iWalk notified customers directly by sales representatives by telephone on October 26, 2012, and instructedthem to stop use of the battery and charger and return them to iWalk. Accounts provided with Return Material Authorizations (RMA).For questions regarding this recall call 781-761-1560. |  | 30 | 21 Batteries; 9 Chargers | Nationwide Distribution including OH, NC, and NY. | Terminated | September 26, 2014 | 700 | 30
Z-0468-2013 | 63643 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Corporation Expression MRI Patient Monitoring System. |  |  | October 15, 2012 | 2012 | Invivo Corporation | Invivo Corporation located at 12151 Research Parkway, Orlando, FL 32826 is recalling the Expression MRI Patient Monitoring System because during transport of the systems cart with a docked Display Controller Unit (DCU), a caster (wheel) may become detached. If this occurs, the cart can become unbalanced and may fall. | Nonconforming Material/Component | Philips Healthcare sent an "URGENT-MEDICAL DEVICE RECALL" letter dated October 15, 2012 to all affected customers. The letter identified the product, problem, and actons to be taken by the customers. Customers were instructed to inspect all casters of the unit to ensure that they are all secured. If a caster is loose, customers were told to lock the caster in place, limit movement of the cart and contact their local Phillips Invivo Representative. Nevertheless, a Philips Invivo representative will contact the customer regarding their affected device. All affected devices will have new casters installed in order to correct the problem. Contact your local Philips Invivo Representative at 1-800-722-9377 for further information and support. |  | 543 | 543 units | Worldwide Distribution-USA (nationwide) including the states of MO; MS; MA; IL; OH; NY; KS; PA; WI; VA; SD; AZ; FL; OR; TX; MA; MO; FL; MA; MI; CT; CA; NJ; NC; ME; IL; MA; MA; WV; MA; MN; CA; AL; AL; WI; DC; OH; WV; CA; NC; NY; MT; NJ; CA; WA; IN; MI; PA; TX; NC; TN; TX; ID; NM; IL; PA; IN; IL; WA; CO; MN; CA; IN; MD; MI; DC; GA; DE; CO; CT; WA; NJ; MD; TX; TN; MN; WA; CA; FL; OH; PA; NY; IN; OK; ID; IA; NY; WI; VA; CA; HI; MS; TX; HI; CT; TN; OR; WA; PA; NH; CA; MA; LA; IN; PA; AZ; GA; UT; ND; GA; TX; TX; IN; CA; MI; CT; CA; NC; MS; VA; IN; MI; NE; NV; MI; WA; IL; KY; LA; OR; OK; LA; TN; KY; TX; CA; PA; NM; NC; CO; DC; AL; MI; TX; SD; PA; OR; ID; NM; SD; CO; MN; NJ; MI; ND; MN; MT; LA; CT; WA; TX; IA; NJ; WI; SD; and UT, and the countries of Canada, Italy, Thailand, Russian Federation, Saudi Arabia, Slovenia, Singapore, Republic of Korea, Uzbekistan, Iceland, Israel, Austria, United Kingdom, Afganhistan, Japan, South Africa, Germany, Qatar, Netherlands, Switzerland, Denmark, and Sweden. | Terminated | February 13, 2015 | 851 | 543
Z-0466-2013 | 63553 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Sliced CT Scanner System, part number 989605858501. NeuViz Dual series CT Scanner System, part number 989605651321 |  |  | October 18, 2012 | 2012 | Philips And Neusoft Medical Systems Co., Ltd. | The four screws that secure the patient table top to the carrier pulled out and caused the table top to detach. This issue may affect NeuViz Dual and NeuViz 16 scanners. | Device Design | Philips and Neusoft Medical Systems Co., Ltd. sent an Urgent Device Correction NeuViz Dual, NeuViz 16 letter, dated Oct. 18, 2012, to all affected US customers and distributors. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to:1. Have a mechanic check the patient table top screws fixing. Detailedinstructions are attached as Annex with the letter.2. If it is found that table top screws are improperly assembled, stop usingthe product and contact local Service personnel for repair.Philips and Neusoft Medical Systems Co., Ltd is in the process of upgrading devices with aforementioned kits and will contact customers to arrange for an update to affected systems. If you have any questions about this corrective action program, please contact the Service Support Department, Neusoft Medical Systems Co. Ltd. at nms-service@neusoft.comFor questions regarding this recall call 425-487-7665. |  | 320 | 20 units of NeuViz 16 CT Scanner Systems and 25 units of NeuViz Dual series | Nationwide Distribution including North Carolina, Ohio, Texas, and Puerto Rico. | Terminated | July 08, 2013 | 263 | 320
Z-0473-2013 | 63577 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas 8000 Modular Analyzer Series, Model/Catalog/Part Number: 05641489001. The Cobas 8000 is a fully automated system for clinical chemistry analysis. |  |  | November 01, 2012 | 2012 | Roche Diagnostics Operations, Inc. | On cobas c 701 and cobas c 702 modules of the Cobas 8000 series analyzer, a mismatch between the reagent probe positioning and the aspirated reagent volumes under certain conditions can lead to insufficient reagent pipetting. This may affect patient sample and controls recovery for the following tests: Alkaline Phosphatase (ALP2), Alanine Aminotransferase (ALT), Bicarbonate (CO2-L), Creatinine (C | Software design | On 11/01/2012 an "Urgent Medical Device Correction" notification was sent via UPS Ground (receipt required) to consignees. The device issue was described and recommended actions were provided. A Roche Field Application Specialist will install the new e-barcode during a scheduled visit. Customers also will have an option to install the new e-barcodes themselves, following the instructions provided in the notification. A Software update will be released when available. Customers may contact Technical Support at 1-800-428-2336 for more information. |  | 78 | 78 analyzers | Nationwide Distribution, including the states of AL, AZ, CA, FL, GA, IL, IN, KS, KY, MA, MD, MI, MN, NC, NJ, NY, OH, VA, WA, and WI. | Terminated | January 08, 2014 | 433 | 78
Z-0475-2013 | 63619 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Software Version R2.0.2 |  |  | October 30, 2012 | 2012 | Philips Healthcare Inc. | With Digital Diagnost R2.0.2 a mirrored "R in a circle" is burnt into the image (if mirrored) in the lower right corner of the image. When such an image is mirrored back inside the PACS system, an unmirrored "R in a circle" appears in the lower left corner. This can be mistaken for a "Right patient side" marker, although this can appear on the left patient side. | Software design | Philips Healthcare initiated a recall on October 25, 2012 via Urgent - Field Safety Notice for the Digital Diagnost X-ray System. The letter identified the affected product, problem and actions to be taken. The letter informs customers that Philips will supply a software update kit to version 2.1.4. For questions please contact your local Philips representative 1-800-722-9377. |  | 623 | 623 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, India, Indonesia, Iraq, Italy, Japan, Latvia, Lithuania, Libya, Malta, Mexico, Netherlands, New Zealand, Norway, Oman, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE and UK. | Open, Classified |  |  | 623
Z-0474-2013 | 63708 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bio-Modular Humeral Head 48mm x 24mm Ti 6AI4V/Ion Implanted (Orthopedic Implant) |  |  | November 01, 2012 | 2012 | Biomet, Inc. | The items in this lot are missing the ion implantation feature. The surface of the head may be softer and more prone to scratching, which in tum could potentially cause higher wear of the poly bearing. Increased wear could potentially lead to a shorter implant life. | Process control | Biomet, Inc. sent an URGENT MEDICAL DEVICE RECALL NOTICE dated November 1, 2012 to all affected consignees. The letter identified the affected product, problem and actions to be taken. The notification letter instructed the surgeon to montior the patient as their is a potential for increased wear which have an unknown effect on implant life. For questions call 574-372-1570. |  | 4 | 4 | USA Nationwide Distribution in the state of WI. | Terminated | August 29, 2013 | 301 | 4
Z-0480-2013 | 63495 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilette Mira, intended use as Mobile x-ray system. |  |  | October 10, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a risk of lost images during examination. | Software design | The firm initiated their recall/correction of this product on September 9, 2012 by sending a Field Safety Notification letter with Update Instruction XP054/12/S to all consignees. The letter described the steps to be taken to avoid the potential risk and included a response form to be returned to the recalling firm. |  | 21 | 21 | Nationwide distribution: USA including states of: GA, ID, IL, IA, NE, NC, OH, SD, TX, and WI. | Terminated | January 15, 2014 | 462 | 21
Z-0484-2013 | 63312 | No | NSX |  |  | General Hospital | Contact ODE | Data Station AMPLILINK software is an instrument user interface/ data management software used with COBASÂż AmpliPrep instrument, COBASÂż TaqManÂż analyzer, COBASÂż TaqManÂż 48 analyzer, COBASÂż AMPLICORÂż analyzer, and Cobas p 630 instrument for sample preparation, amplification and detection of PCR (Polymerase Chain Reaction) molecular diagnostic analytes. |  |  | July 16, 2012 | 2012 | Roche Molecular Systems, Inc. | Roche Molecular Systems, Inc. has become aware of a device-associated diagnostic software issue. | Software design | Roche sent an "Urgent Medical Device Correction" (UMDC) notification letter dated July 16, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Roche Molecular Systems, Inc. has disabled this software on affected Roche platforms and discontinued the distribution of the affected software. Contact Roche Diagnostic Technical Support, 24 hours a day, seven days a week at phone number : 1-800-428-2336 if you have questions regarding this notification. For questions regarding this recall call 908-253-7569. |  | 884 | 884 units | Nationwide Distribution | Terminated | July 22, 2015 | 1101 | 9250
Z-0478-2013 | 63635 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | The GlideScope Video Laryngoscope (GVL) is a video laryngoscope that incorporates a high-resolution color camera, an LED light source, a rechargeable lithium battery, and video output to a remote monitor. It consists of a video laryngoscope and a portable video monitor. |  |  | October 25, 2012 | 2012 | Verathon, Inc. | Re-usable video laryngoscope blades of the GlideScope GVL3, GVL4, and GVL5 were recalled due to potential cracking. | Device Design | The firm, Verathon, sent an "URGENT - Medical Device Recall" letter dated October 25, 2012, to their consignees/cutomers. The letter described the product, problem, and actions to be taken. The customers were instructed to return the recalled blades back to Verathon and complete and return the Device Recall Notification Reply Form via fax to (425) 883-2896 or email: cservice@verathon.com. All returned products will be quarantined and will be destroyed later. Verathon will replace the affected video laryngoscope blade at no cost to customers. Customers in Canada and US with questions can call 800-331-2313. If you are located outside of the US or Canada, please call 425-867-1348. |  | 2710 | total 2710 units | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NM, NV, NJ, NY, OK, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY; and countries of: Argentina, Australia, Canada, China, France, Germany, Indonesia, Japan, The Netherlands, Pakistan, Panama, Peru, Philippines, Singapore, South Korea, Taiwan, Thailand, United Kingdom, and Venezuela. | Terminated | September 27, 2013 | 337 | 2710
Z-0481-2013 | 63518 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Single Channel Infusion Pumps with Plug-and-Play Modules that provide wired Ethernet and wireless local area networking; Hospira, Inc., Lake Forest, IL 60045; pump list number 11971, module list numbers 12101, 12102 and 12393; pump list number 11973, module list number 12380; pump list number 12391, module list numbers 12097 and 12680 |  |  | October 26, 2012 | 2012 | Hospira Inc. | The volume control knob on some Plum A+ single channel infusers (located on the back of the infuser) may not function as described in the System Operating Manual - the direction for Loud and Quiet may be reversed. The manual indicates a clockwise rotation of the knob decreases the volume; however, on some of the devices the alarm volume increases when the knob is turned clockwise. | Device Design | The firm, Hospira Inc. sent "URGENT DEVICE FIELD CORRECTION" letters dated October 26, 2012 to all affected customers. The letters identified the products, problem and action(s) to be taken. The customers were instructed to read the information provided carefully, complete the attached Reply Form and return it to the firm via fax at 1-877-650-8667. The letter states that the System Operating Manual will be updated to clarify the instructions for setting the volume, and that the updated manuals will be sent as soon as they are available. Customers were instructed to notify their consignees if they have further distributed the affected product. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for additional information and technical assistance. |  |  |  | Worldwide Distribution - USA (nationwide) including Puerto Rico, and the countries of Australia, Bahrain, Belgium, Brazil, Canada, Chile, Colombia, Denmark, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Lavenia, South Africa, Spain, Switzerland, Taiwan, Turkey, United Kingdom, United Arab Emirates and Vietnam. | Open, Classified
Z-0504-2013 | 63728 | No | EOQ | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | WANG Transbronchial Aspiration Needles Cat. No. SW-121, WANG Cytology Needles, Central, Carinal and Mediastinal Regions, 21 g x 15 mm long needle. --- Common Name: Bronchoscope (flexible or rigid) and accessories. The ConMed WANG Transbronchial Aspiration Needles are designed for use with a bronchoscope. They are used to puncture the tracheobronchial wall and to aspirate sufficient tissue and/or cell specimens to stage bronchogenic carcinoma. These are disposable sterile, single-use devices. |  |  | November 20, 2012 | 2012 | ConMed Corporation | ConMed received complaints from a user facility reporting that the WANG Transbronchial Aspiration Needles would not insert properly through the instrument channel of a 2.0 mm bronchoscope. If the device will not fit through the instrument channel of a 2.0 mm bronchoscope, it must be replaced with an alternate device and could result in minor inconvenience to the user. | Component design/selection | ConMed Corporation sent an Urgent Field Safety Notice letter dated November 30, 2012 with Response Form , to all affected customers via USPS Priority Mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete Attachment II and return it with any unused affected product to:ConMed Corporation525 French Road, Utica, NY 13502United Sates of AmericaFor information regarding this recall call 315-624-3225. |  | 452 | Total 452 | Worldwide Distribution - USA including CA , FL, IA, IN, MA, MD, ME, MO, MS, NC, NM, OH, PA, TX. Foreign distribution to Spain, Portugal, Japan and Hungary. | Terminated | August 25, 2015 | 1008 | 452
Z-0495-2013 | 63709 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Information Center/Client and Philips IntelliVue Information Center iX |  |  | November 09, 2012 | 2012 | Philips Healthcare Inc. | Internal audio cable may have diminished retention force due to damaged contacts, which could lead to intermittent or loss of audible alarms. | Component design/selection | Philips sent an URGENT - Medical Device Correction letter dated November 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Philips is conducting a voluntary correction to replace the audio cable on affected devices. Customer were instructed to follow the "Action to be taken by CustomerUser" section of the instructions. Customers with wquestions were instructed to contact their Phililps representataive or contact the Customer Care Solutions Center at (800) 722-9377.For questions regarding this recall call 978-687-1501. |  | 3165 | 3165 | Worldwide Distribution - USA (nationwide) and Canada. | Terminated |  |  | 3165
Z-0501-2013 | 61209 | No | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | Empi STIMC RE Self Adhesive electrodes. For use on TENS, NMES, FES. a DJO brand |  |  | February 09, 2012 | 2012 | Covidien LLC dba Uni-Patch | Covidien is conducting a voluntary recall regarding specific lots of TENS electrodes because it was determined that during the manufacturing process the surface of the wires in these lots may have been exposed to an employee who had a minor cut. | Employee error | Covidien sent An Urgent Product Recall letter, dated February 9, 2012, to all affected consignees. The letter identified the affected product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine the affected product if found. The letter asked customers to fill in teh quantity remaining in their inventory and fax it to 651-565-5888. A Covidien Customer Service rep would contact the consignee for returning product. Consignees were also requested to notify their customers of the recall.For questions regarding this recall call 508-261-8532. |  | 28138 | 28,138 pouches | Nationwide Distribution including CA, IL, MA, MN, and SD. | Terminated | January 03, 2013 | 329 | 37036
Z-0508-2013 | 63682 | No | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Siemens Healthcare ADVIA Centaur CA 19-9 |  |  | November 14, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Operating range of the ADVIA Centaur systems is 18C to 30ÂżC (64.4ÂżF to 86.0ÂżF), a change in room temperature may increase or decrease assay results, depending on the specific assay and ADVIA Centaur system used. | Device Design | Siemens Healthcare issued An Urgent Medical Device Correction Notice (for US customers) and an Urgent Field Safety Notice (for customers outside the US) dated November 5, 2012, to all affected customers. These notices inform customers who used the ADVIA Centaur BNP, BR, CA 19-9, Digitoxin, Folate, and Theophylline 2 assays to reduce their instrument operating temperaturerange to 20 -25 degrees C and for the DHEAS assay to reduce theinstrument operating temperature range to 18 -26 degrees C when theseassays are being used. These notices also direct customers to review thecontents of the correction notice with their Laboratory Director or MedicalDirector. A confirmation fax-back form was included to ensure that allcustomers worldwide have been notified and have complied with themandatory action.Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local technical support representative.For questions regarding this recall call 508-668-5000. |  | 50 | 250 test=525 and 50 test=12243 | Worldwide Distribution - USA (nationwide) and Internationally to the following countries: ArgentinaAustraliaAustriaAzerbaijanBelgiumBosnia Herzeg.BrazilCanadaChileCOLOMBIACroatiaCzech RepublicDenmarkEgyptEstoniaFinlandFranceGermanyGreeceHong KongHungaryIcelandIndiaIndonesiaIranIrelandItalyJapanKazakhstanLatviaLithuaniaMalaysiaMexicoNetherlandsNew CaledoniaNorwayOmanP.R. ChinaPhilippinesPolandPortugalQatarRep. of YemenRepublic KoreaReunionRomaniaRussian Fed.Saudi ArabiaSerbiaSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTaiwanThailandTurkeyU.A.E.United KingdomUruguayVietnam | Terminated | September 25, 2015 | 1045 | 25245
Z-0513-2013 | 63614 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | DPM 6/7 Patient Monitor |  |  | August 11, 2010 | 2010 | Mindray DS USA, Inc. d.b.a. Mindray North America | There is an issue with the DPM 6 and DPM 7 monitors where the following features may be unavailable : Full Disclosure, Drug, Hemodynamic Calculations, Renal Calculations, Oxygenation Calculations, and Ventilation Calculations. | Device Design | Mindray sent a "PRODUCT CORRECTIVE ACTION" letter dated August 11, 2010 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers should contact their Mindray Service Representative at 1-800-288-2121 for assistance with this issue. |  | 124 | 124 units | Nationwide Distribution-USA (nationwide) including the states of WA, IL, FL, MN, SC, NV, MI, TN, PA, LA, NE, TX, and IA. | Terminated | August 02, 2013 | 1087 | 124
Z-0515-2013 | 63496 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MAVIG PORTEGRA 2 system operated in conjuction with a Siemens AX system |  |  | October 10, 2012 | 2012 | Siemens Medical Solutions USA, Inc | A particular component may not have been fitted during the installation of some MAVIG PORTEGRA 2 systems. | Equipment maintenance | Siemens sent a Customer Safety Advisory Notice on October 3, 2012, to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were reminded that they should contact the service department in the event of damage or other irregularity. Customers were instructed to forward the information contained in the notice to their employees responsible for the operation of the affected product. Customers were also asked to forward the safety information to other organization affected by the product and to any users of the product.For questions regarding this recall call 610-219-4834. |  | 1315 | 1315 | Nationwide Distribution | Terminated | November 06, 2014 | 757 | 1315
Z-0520-2013 | 63758 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD CYCLESURE 24 Biological Indicator (BI) product, P/N 14324. |  |  | October 03, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling all lots of Sterrad Cyclesure 24 Biological Indicator (BI) manufactured between February 2008 and December 2011 because it may not have adequate data to support the entire duration of the labeled shelf-life. | Nonconforming Material/Component | Advanced Sterilization Products (ASP) sent an "URGENT MEDICAL DEVICE RECALL (REMOVAL) letter dated November 2, 2012 to all customers who received the Sterrad Cyclesure 24 Biological Indicator (BI) product P/N 14324 manufactured between February 2008 and December 2011. The letter informed the customers of the problem identified and the actions to be taken. ASP determined that the 24 BI may not have adequate data to support the entire duration of the labeled shelf life of affected product manufactured during the above timeframe. The letter also informed customers that this is a different recall from the previous July 3, 2012 STERRAD CYCLESURE 24 BI recall. Customers with questions or need additional packing labels are instructed to call Stericycle at (877) 650-7692. |  | 38532 | 38,532 units | Worldwide Distribution-USA (nationwide) and the countries of AR, AU, BE, BR, CA, CL, CN, CO, EC, GT, HK, ID, IN, JP, KR, MX, MY, PA, PE, PH, SG, TH, VE, and RU. | Terminated | July 19, 2013 | 289 | 38532
Z-0519-2013 | 63530 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | Quasar Power Pack contains a Baby Quasar and a Baby Blue. The Baby Quasar is an infrared LED lamp. The Baby Blue uses blue LED light to be used in combination with the Baby Quasar to treat acne. |  |  | October 25, 2012 | 2012 | Quasar Bio-Tech, Inc. dba Silver Bay LLC | Quasar Bio-Tech is recalling their Baby Quasar, Baby Quasar Pink, Quasar Power Pack, and Quasar MD because the device is unapproved. | No Marketing Application | The firm, Quasar Bio-Tech, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated November 16, 2012 to its customers. All domestic customers were notified by Certified mail. The letter described the product, problem and actions to be taken. The customers were instructed to cease distribution and remove the products from sale; to complete and return the enclosed Response Form via e-mail to DeNette@quasarbiotech.com or fax to: Quasar Bio-Tech, Attn: Quality/Recall Coordinator: at 941-306-5816, and return all remaining stock to Quasar Bio-Tech at Quasar Bio-Tech Inc., Attn: Quality/Recall Coordinator, 1431 Tallevast Road, Sarasota FL 34243.If you have any questions, contact the Quality/Recall Coordinator at 941-306-5812, Monday-Friday from 9:00 am to 5:00 pm Eastern Standard time. |  | 7734 | 7,734 devices total (included in batches identified for the Baby Quasar ) | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CO, FL, GA, HI, ID, IL, KS, MD, MI, MN, MO, NC, NJ, NY, NV, OH, OR, TN, TX, UT, VA, and WA; and countries of: Aruba, Canada, Cayman Islands, China, Estonia, Ireland, Malaysia, New Zealand, Singapore, Sweden, Taiwan, and United Kingdom. | Terminated | May 21, 2013 | 208 | 7734
Z-0522-2013 | 63680 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cobas b 123 POC system |  |  | November 14, 2012 | 2012 | Roche Diagnostics Operations, Inc. | It has been determined that cobas b 123 Fluid Pack COOX cuvettes for lot numbers 21426121 and 21426171, exhibit a significant increase in the out-of-box failure rate caused by a change in the thickness in the cuvette layer. Internal investigations found the altered thickness of the cuvette was due to production equipment settings. Cuvettes were produced within specifications, but within the upper | Software design (manufacturing process) | Roche sent an Urgent Medical Device Removal (UMDR12-268) notification and faxback form on November 14, 2012, to all affected consignees. Customers were requested to discontinue the use of cobas b 123 Fluid Pack COOX, 200 lots 21426121 and21426171 and discard the product per local regulations. Customers were instructed to use the enclosed faxback form to order replacement product at 1-800-722-7222. For questions customers were instructed to contact Roche Diagnostics Technical Support at 1-800-526-2272.For questions regarding this recall call 317-521-3966. |  | 4600 | 23 units of Fluid Pack COOX, 200 were distributed | Nationwide Distribution including NE and AZ. | Terminated | August 07, 2013 | 266 | 4600
Z-0523-2013 | 63723 | No | GRO | Haemophilus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD BBL Taxo XV Factor Strips, Catalog number 231104, in shelf pack vials, fifty strips/vial labeled in part ***Becton, Dickinson and Company, 7 Loveton Circle, Sparks, MD 21152 USA 800-638-8663, www.bd.com/ds*** |  |  | October 22, 2012 | 2012 | Becton Dickinson &amp; Co. | In vitro diagnostic test kit was not manufactured according to specifications and may cause false negatives or misidentification of Hemophilus species present in patient specimens. | Employee error | BD Diagnostic Systems sent an Urgent Product Recall letter dated October 22, 2012, to all affected customers.. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to discontinue distribution and discard affected inventory for credit. . Customers were requested to return a response form by fax to 410-316-4258. For questions customers were instructed to contact BD Customer Service Department at 1-800-638-8663. For assistance with replacements, customers were instructed to call 1-800-675-0908.For questions regarding this recall call 410-316-4258. |  | 650 | 650 vials | Worldwide Distribution - USA (nationwide) to international customers in Belgium, India, Japan, South Korea, Thailand and Mexico | Terminated | January 24, 2013 | 94 | 650
Z-0497-2013 | 63712 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II, Model 8637, (The SynchroMedÂż II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The implantable components of the SynchroMed II Infusion System include the pump, catheter, and catheter accessories. |  |  | November 09, 2012 | 2012 | Medtronic Neuromodulation | Medtronic notified Healthcare Professionals of the impact of unapproved drugs on the performance of the SynchroMed infusion pump system. Use of unapproved drugs with SynchroMed pumps can result in an increased risk of permanent motor stall and cessation of drug infusion. | Nonconforming Material/Component | The firm Medtronic, sent an "Urgent medical Device Safety Notification" dated November 2012, to its customers beginning November 09, 2012. The letter described the product, problem and actions to be taken. The customers were instructed to follow the recommendations provided in the letter and complete and return the PHYSICIAN REPLY FORM via mail at Medtronic Neuromodulation Quality Assurance, 7000 Central Ave NE, Mailstop RCC150, Minneapolis, MN 55432; fax to: 800-897-3899 or email to: neuro.quality@medtronic.com.If you have questions, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am - 6pm CST. |  | 174700 | 174,700 devices total | Worldwide distribution: USA (nationwide) and countries including: Canada and Mexico. | Completed |  |  | 174700
Z-0521-2013 | 63694 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Ready-to-Use Mouse Monoclonal Antibody Terminal Deoxynuceotidyl Transferase Staining Reagent; an in vitro staining reagent for the qualitative identification by light microscopy of Terminal Deoxynucleotidyl Transferase molecules in paraffin sections; Catalog No. RTU-TdT-339; packaged in 7 mL vials; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom. |  |  | November 14, 2012 | 2012 | Leica Microsystems, Inc. | RTU-TdT-339 is not stable up to the expiry date on the product labeling, affecting the staining intensity. | Nonconforming Material/Component | Leica Microsystems, Inc. sent a Product Recall Notification letter dated November 14, 2012, via first class mail on the same date, to all affected customers. The letter identified the product ,the problem, and the actions to be taken by the customers. Customers were instructed to discontinue use of the product and to return any unused/part used reagents to the Richmond, IL, location for replacement. The accounts were asked to complete the attached Acknowledgement Form, acknowledging receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-236-3747. Additional questions should be addressed to the account's Leica Representative. |  | 322 | 322 vials | Worldwide Distribution -- USA including Arkansas, California, Florida, Illinois, Maryland, Michigan, Minnesota, Missouri, North Carolina, Oregon, Rhode Island, Tennessee, Texas and Wisconsin, and the country of Canada. | Open, Classified |  |  | 322
Z-0514-2013 | 63710 | No | LSZ |  |  | Anesthesiology | PMA | Patient Circuit used with the Life Pulse High Frequency Ventilator. The Patient Circuit is a required disposable component of the Life Pulse High Frequency Ventilator. |  |  | November 19, 2012 | 2012 | Bunnell, Inc. | Customer complaints received indicate the heater wire insulation can melt, causing a short which creates sparking and smoke, in the circuit close to the humidifier cartridge. Bunnell is recalling all lots of circuits distributed between March 12, 2012 and November 30, 2012. | Component design/selection | On 12/5/12, the firm, Brunnell, Inc., issued a public press notification to their consignees. A. On or before 12-12-2012 a certified letter will be sent to all customer listed in the customer database. Specifically each customer will receive two certified letters; one directed to Director of Respiratory Therapy and a second to the Director of Biomedical Engineering. The certified letter will contain a copy of the Recall Notice and a copy of the Certificate of Medical Necessity. The Certified Letter receipts will be maintained. On the mailing date a copy of the Recall Notice will be placed on the Bunnell website. B. An Excel database will be maintained for each customer/lot #. The database will contain the current status for each. Included will be: 1. Customer Information 2. Date Certified Letter Sent 3. Date receipt received and filed 4. Date product returned or that no stock remains 5. Lot number &amp; Quantity C. On or before 02-01-2013, a second letter will be sent to those customers who have not responded to the original certified letter. D. Data will be maintained in the database E. On or before 04-01-2013, a phone call will be placed to those customers who have not responded to either of the letters. F. Data will be maintained in the database G. On or before 04-15-2013, a second phone call will be placed to those customers who have not responded to the letters or previous phone call. H. Data will be maintained in the database I. On or before 06-07-2013, a Final Recall Notification certified letter will be sent to those customers who have not responded to either of the previous letters or phone calls. J. The Certified Letter receipt will be maintained K. Data will be maintained in the database L. When PMA for alternate heater wire is approved, steps A-K will be repeated. Dates will be determined. M. Following completion of Step L, a final review of the database will occur. Final tabulation of the status for each custo |  | 5743 | 5743 patient circuits distributed in commerce | Worldwide Distribution-USA (nationwide) and the countries of Canada, Australia, Malaysia, and Uruguay (animal study only). | Terminated | February 21, 2014 | 459 | 5743
Z-0524-2013 | 63655 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Carescape Patient Data Module with v2.0 software used with Solar 8000M/l, Transport Pro or Carescape B650/B850 host monitor. The Patient Data Module (PDM) is intended to provide physiologic parameter data on adult, pediatric and neonatal patients during bedside and transport patient care episodes. |  |  | February 17, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare became aware of two issues with the Patient Data Module using Version 2.0 software that may impact patient care. | Software design | Consignees were sent on 2/17/2012 three GE Healthcare "Important Product Information" letters dated 10/17/ 2011. The letters were addressed to Healthcare Administrator/Risk Manager, Director of Biomedical Engineering, and Chief of Nursing. The letters described the Issue, Instructions, Affected Product Details, Product Correction and Contact Information. |  | 10250 | 10,250 | Worldwide distribution, including the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, LA, KS, KY, LA, ME,MD, MA, MI, MN, MS,MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK OR, PA, SC, TN,TX, UT,VA, WA, WT, WY and the countries of the UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVAK IA, SINGAPORE, SAUDI ARABIA, QATAR, POLAND, PERU, PANAMA, NORWAY, NEW ZEALAND, NETHERLAND, MEXICO, MALAYSIA, LEBANON, KUWAIT, KOREA, JAPAN, ITALY, INDONESIA, INDIA, IRELAND, GERMANY FRANCE, FINLAND, EGYPT, CZECH REPUBLIC, COSTA RICA, COLOMBIA, CHILE CANADA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, ARGENTINA, and ALGERIA. | Terminated | December 03, 2013 | 655 | 10250
Z-0494-2013 | 63650 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash 3000/4000/5000. |  |  | June 26, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential issue due to an unexpected loss of CustomAutomatic NBP Measurement cycling associated with the Dash 3000/4000/5000 patient monitors. | Software design | GE Healthcare sent a Important Product Information Letter dated June 21, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If your facility is using the Custom Auto NBP Measurement feature in conjunction with Automatic View On Alarm and/or View Other Beds options, GE Healthcare recommends that you stop using the Custom Automatic NBP Measurement feature and use the Auto NBP function instead.All Dash3000/4000/5000 patient monitors with software version 6.0 and later with the networking option enabled may be affected. You can determine the software version of the Dash monitors by selecting MORE MENUS &gt; MONITOR SETUP &gt; SOFTWARE CONFIGURATION. Product CorrectionGE Healthcare is working on a software correction that will be provided to you at no charge once it is available. We will contact you to schedule the correction.Contact InformationIf you have any questions regarding this notification, please contact Technical Support at 1-800-558-7044 or your local Service Representative.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above.. |  | 119260 | 119,260 devices | Worldwide Distribution -- USA (nationwide) including the countries of AUSTRALIA, ANGOLA,ARGENTINA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELGIUM, BOSNIA HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GHANA, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNG ARIA, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAHSTAN, KENYA, KOREA, KUWAIT, LATIVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, MALTA, MOROCCO, NEPAL, NETHERLAND, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PALESTINIAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN THAILAND, TRINIDAD AND TOBAGO,K TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARB EMIRATES , UNITED KINGDOM and URUGUAY | Terminated | June 23, 2017 | 1823 | 119260
Z-0535-2013 | 63549 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Titan Spine Endoskeleton TA Implant products. |  |  | October 19, 2012 | 2012 | TITAN SPINE, LLC | Titan Spine LLC, is conducting a recall on the Endoskeleton TA VBR products. The labeling accompanying this device contains incorrect indications for use. | Labeling False and Misleading | Titan sent an Urgent Medical Device Recall letter dated October 19, 2012, to all affected customers. The letter described the problem , the product affected and the corrective measures to be taken. Customers were instructed to review any surgeries that were performed at the L5-51 level with the affected device. Customers were asked to communicate any complications associated with those cases to Titan Spine. Customers with questions should call 1-262-242-7801.For questions regarding this recall call 262-242-7801. |  | 92 | 92 pcs. | Nationwide Distribution including NV | Terminated | January 14, 2013 | 87 | 92
Z-0536-2013 | 63653 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Polaris 5.5 Button Lock Screw Inserter. Catalog number 14-500178 |  |  | September 29, 2012 | 2012 | Ebi, Llc | Broken screw insertor hardness was measured and found to be below the acceptable limits. | Process control | Biomet mailed Urgent Medical Device Recall Notice/Fax Back Response Form dated 9/25/2012 to the distributors on 9/29/2012. |  | 15 | 15 units | Nationwide Distribution, including the states of MO, TX, and IN. | Terminated | December 07, 2015 | 1164 | 15
Z-0527-2013 | 63725 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | EBI XFIX DFS Rail System (Adult Rail Gradual Swivel Clamp). The Biomet Rail System consists of a unilateral external fixator rail, modular bone screw clamps, and compression and distraction units. The device is used in the treatment of bone fractures. |  |  | November 09, 2012 | 2012 | Biomet, Inc. | It was reported that the number scale/graduation markings were missing on 14 pieces of the swivel clamp. | Process design | On 11/9/2012 an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter was sent to all consignees (distributors and Doctor notices) informing them of the issue and providing recommended actions. Customer contact number (574) 372-1570 Monday through Friday from 8 am to 5 pm. |  | 49 | 49 | Worldwide Distribution, including the states of TX, TN, PA, CA, NY, FL, and MI and the countries of Japan and Columbia. | Terminated | August 28, 2013 | 292 | 49
Z-0526-2013 | 63824 | No | HCC | Biofeedback device. | Neurology | Neurology | 510(K) Exempt | EEG NeuroAmp, Model Number: CS 10090; CS 10137. |  |  | November 15, 2012 | 2012 | EEG Info | EEG Info, Inc. is recalling the EEG Neuroamp device because they have identified a potential risk associated with the mislabeling of certain devices. | Labeling False and Misleading | EEGInfo sent a recall notification letterdated November 29, 2012, to all affected customers. The letter informed the customers they have not received approval from the FDA to market and distribute a device for therapeutic applications other than that of its FDA intended use-biofeedback relaxation. Customers are informed that the safety and effectiveness for other indications has not been established. The risk or likelihood of serious adverse health consequences is remote. Customers were instructed to forward this notice if affected product was further distributed. Customers with questions about the notification letter are instructed to contact EEG Info's customer support line during normal business hours, Monday-Friday, from 8am-5pm PST. For questions regarding this recall call 818-456=5965. |  | 1308 | 1308 units | Nationwide Distribution | Terminated | June 11, 2013 | 208 | 1308
Z-0532-2013 | 63520 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Medline Total Knee CDS; a single patient prescription procedure pack and disposables, including a back table cover; |  |  | September 20, 2012 | 2012 | Medline Industries Inc | The Back Table Cover component in the custom pack was placed in the non-sterile portion of the pack when it should have been in the sterile portion. | Packaging | Medline Industries Inc notified the sole direct account on September 20, 2012 via telephone that the affected custom pack were recalled. Instruction was provided to the staff at the facility not to use the pack until it had been reworked with a sterile back table cover. |  | 18 | 18 packs | US (nationwide) in the state of West Virginia | Terminated | December 17, 2012 | 88 | 18
Z-0531-2013 | 63713 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand ARIA Radiation Oncology, ARIA Oncology Information System Radiation Oncology, Import Export Application, Model Number: HIT, Affected Versions: ARIA version 11.0 below 11.0.55, Build 11.0.28 and 11.0.34, Reference/FSCA Identifier: CP-09798; |  |  | November 05, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian Medical Systems has identified an anomaly with the ARIA Oncology Information System Radiation Oncology [ARIA RO] DICOM import/export functionality where a 360 degree arc field may convert to static field during DICOM export or import. | Software design | Varian sent an Urgent Medical Device Correction Urgent Field Safety Notice dated November 6, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that the plan was not converted unintentionally during import/export. Customers would be contacted by a Customer Service Representative when the correction was available to schedule installation on their system. Customers with questions were instructed to contact their local Varian Medical Systems Customer Support District or Regional Manager.US and Canada 1-888-827-4265.For questions regarding this recall call 650-424-5731. |  | 264 | 264 units | Worldwide Distribution | Terminated | March 03, 2014 | 483 | 264
Z-0529-2013 | 63732 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Mitek 11 MM Fully Fluted Reamer, Sterile |  |  | November 16, 2012 | 2012 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Specific units of the DePuy Mitek Sterile Reamers are not labeled correctly, 7 units were labeled as 11MM while the box actually contained a 10 MM reamer | Packaging process control | DePuy Mitek sent an Urgent Voluntary Product Recall letter dated November 16, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to refer to the attached instructions to report their inventory status and for returning the affected product. Customers with questions should call 1-508-880-8477. For questions regarding this recall call 508-880-8100. |  | 20 | 20 | Nationwide Distribution including IL and NJ | Terminated | December 21, 2012 | 35 | 20
Z-0533-2013 | 63781 | No | MBV | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Offset Tibial Tray 2.5 mm Adaptor, REF 141490 |  |  | November 15, 2012 | 2012 | Biomet, Inc. | Biomet investigation identified that the square end of the offset tibial tray adaptor locking insert is oversized and may not engage insert tool screwdriver. This could cause a possible delay in surgery greater than 30 minutes, potentially exposing the patient to the increased risks from being under anesthesia for a longer duration. | Employee error | Biomet Recall sent an Urgent Medical Device Recall Notices letter dated November 15, 2012 to distributors, hospital operating room(OR)managers and Surgeons receiving the affected product. The letter identified the affected product, the reason for the recall and actions to be take, including directions to immediately locate and discontinue use of the affected product. Customers were asked to confirm receipt of the notice by calling 1-888-348-9500 extension 1570. Customers with affected product to be returned are asked to obtain RGA # by calling 574-372-6677 or email: rgarequest@biomet.com. A FAX Back Response Form is to be filled out and sent within 3 days of receipt to 574-372-1683. Return product to the Attention of Audrey Daenzer, Field Action Coordinator, Biomet, Attention Return Goods, Building B Recalls, 56 East Bell Drive, Warsaw, IN 46580. Questions related to the notice should be directed to 474-372-3983 or 574-371-3009, Monday- Friday 8 a.m. to 5 p.m. |  | 200 | 200 | Worldwide distribution - USA (nationwide) and the countries of Australia, Canada, Chile, Europe, Japan, Korea, Panama, Puerto Rico and Thailand. | Terminated | November 08, 2013 | 358 | 200
Z-0534-2013 | 63808 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Arobella Medical 6mm Qurette used with the Arobella Medical AR1000 Wound Therapy System. |  |  | November 05, 2012 | 2012 | Arobella Medical, LLC | This recall has been initiated because the Arobella Medical 6mm Qurette Q07 Lot Numbers 0001 and 0002 can fracture during use and may release a fragment or fragments. The 6 mm Qurette is a component of the AR1000 Quostic Wound Therapy System. | Nonconforming Material/Component | Arobella Medical sent an " Urgent: Medical Device Recall" letter dated November 2, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine their inventory, quarantine and return any product subject to the recall. If the affected product have been further distributed, customers were advised to notify their customers of the product recall. For questions call 952-288-2810. |  | 99 | 99 | Worldwide distribution - US (nationwide) and the states of AL, CA, ID, IN, IA, MI, MN, MS, NJ, NM, OH, PA. SC, TX, UT, VA, WI. including the country of SAUDI ARABIA. | Terminated | March 04, 2015 | 849 | 99
Z-0528-2013 | 63652 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray V12 Size: 320 mm X 320 mm X 450 mm N.W.: 3 kg G.W.: 6 kg Qty:1 Manufactured in China |  |  | September 24, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified an issue with the V Series Monitor where the monitors touch screen may stop responding to touch. | Nonconforming Material/Component | Mindray issued an "Urgent: V Series Monitoring System Corrective Action" letter dated 9/24/2012 via certified mail with return receipt verification to all hospitals.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Mindray Service Representative at 1-800-288-2121 to arrange for relplacement of the affected product. For questions regarding this recall call 201-995-8407. |  | 206 | 206 units (180 units domestic, 26 units foreign) | Worldwide Distribution - USA (nationwide) and internationally to Great Britian, Colombia, and Australia. | Terminated | August 15, 2013 | 325 | 206
Z-0538-2013 | 63641 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Coonrad/Morrey New Ulnar Revision Kit for Use with Non-Interchangeable Regular Humerals STERILE Rx, Tivanium QTY/PKG:1 |  |  | November 19, 2012 | 2012 | Zimmer, Inc. | Zimmer is initiating a lot specific recall because affected lots could include an extra Small Inner Pin instead of the required Regular Inner Pin in the Coonrad/Morrey Elbow Replacement Set and the New Ulnar Revision Kit . | Process change control | Zimmer sent "URGENT MEDICAL DEVICE RECALL" letter dated November 19, 2012 to Distributors, Risk Managers/ Implanting Surgeons, and direct accounts. The notice describes the product, problem, risks and actions to be taken. Consignees were instructed to immediately locate, quarantine and return affected product with the completed Inventory Return Certification form to Zimmer Product Service Department, 1777 West center St. Warsaw, IN 46580 (US accounts) or Zimmer International Logistics GmbH, ATTN: Fao Tanja Herold (Recall Warsaw), Max-Immelmann-Allee 12 79427 Eschbach Germany for international accounts. If an implant is removed from customer inventory, they were directed to provide a copy of the Risk Manager letter upon retrieval of that implant(s) to ensure all facilities are aware of the removal. Surgeons are asked to consider risk information in the letter should a patient present with unexplained symtoms. Hospitals that received direct shipments from the Warsaw Distribution Center will be sent a copy of the Risk Manager letter directly. Questions and concerns along with information on any implanted product is to be addressed to CorporateQuality.PostMarket@zimmer.com using the provided spreadsheet or contact Zimmer at 1-800-447-5633. |  | 17 | 17 | Worldwide distribution - USA nationwide and the countries of: Canada, Germany, Finland, Japan, Italy, San Juan, Switzerland, and UK. | Terminated | April 17, 2014 | 514 | 40
Z-0546-2013 | 63690 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona, The Personalized Knee System, Tibial Drill Guide, Cemented, 15.7 mm Diameter, Nonsterile, Nonsterile |  |  | November 14, 2012 | 2012 | Zimmer, Inc. | All lots of the Persona Cemented Tibial Drill, 5 Degree Cemented Stemmed Natural Tibia implants, and associated tools including the Tibial Drill Guide, Broach, and Sizing Plate are being removed and replaced with enhanced products and instruments to reduce risks of tibial perforation which could result in tibial bone fracture due to the stress riser created by a perforation; tibial plate loosening | Device Design | Zimmer sent distributors beginning the week of November 14, 2012, E-mail information requesting all equipment be collected in distributors office before shipment back to Zimmer and replacement. A telephone conference discussing the recall and amended surgical technique was held the week of November 14, 2012, and Urgent Recall notices were sent to Distributors describing the devices; risks; updated surgical technique with directions to hand deliver surgeon/risk manager letters to all affected surgeons and risk managers in hospitals receiving the product. The notices included directions to remove all affected product and ensure risk manager letter is provided. The letter stated a Zimmer Inc. employee will conduct onsite reconciliation and asked return of distribution Spreadsheet to CorporateQuality.Postmarket@zimmer.com within 24 hours of receiving notification. Product is to be returned with certificate of sterilization to Zimmer Product Service Department 1777 West Center St. Warsaw, IN 46580 or Zimmer International Logistics, Eschbach Germany for International Accounts.For questions regarding this recall call 574-372-4807. |  | 94 | 94 | Worldwide Distribution - USA (nationwide) and internationally to Germany and Europe. | Terminated | April 16, 2014 | 518 | 4035
Z-0556-2013 | 63775 | cane, safety walk | KHY | Cane. | Physical Medicine | Physical Medicine | 510(K) Exempt | Invacare Height Adjustable (walking) Canes, Model Numbers: 8916, 8917, and 40918-6. The device is packaged six (6) canes per shipping carton. |  |  | November 06, 2012 | 2012 | Invacare Corporation | Following their receipt of several customer complaints, Invacare recalled their height adjustable walking canes. The device was recalled from distribution based on reports indicating that there is a possibility that the button which holds the cane at the desired height, may suddenly and unexpectedly compress, causing the cane to collapse and cause possible injury to the user. | Device Design | The firm, Invacare, sent an "URGENT MEDICAL DEVICE RECALL (Removal)" letter dated November 6, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to examine their existing stock using the enclosed detailed 'Product Tracking Sheet' and quarantine any affected canes; Acknowledge receipt and understanding of the urgent recall notice and indicate the number of canes remaining in inventory on the enclosed, 'Provider Response Card;'. Fax the completed Provider Response Card to Invacare Regulatory Affairs at: 1- 440-326-3544; disassemble each cane and discard both pieces in the trash; (Disassembly instructions are also provided); and lastly, contact Invacare Customer Service at 1-800-668-2337, Monday - Friday, 8 a.m. - 8 p.m. EST, to order replacement merchandise at no charge.If you have any questions concerning these instructions, contact Invacare Customer Service at 1-800-668-2337, Monday - Friday, 8 a.m. - 6 p.m. ET. |  | 38280 | 38,280 canes | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN KS, KY, LA,MA, MD ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH,OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries of: Australia, Puerto Rico, and Taiwan Province of China. | Terminated | August 20, 2013 | 287 | 38280
Z-0557-2013 | 62826 | No | ILK | Powered patient transport. | Physical Medicine | Physical Medicine | 510(k) | ThyssenKrupp Accessibility LEVANT Stairlift. |  |  | August 06, 2012 | 2012 | ThyssenKrupp Access Corp | The seat could break as a result of damage resulting from a seat belt being trapped underneath the foldable part of the seat when the seat belt is not buckled. | Device Design | The firm, ThyssenKrupp Access sent an "IMPORTANT SAFETY AND PRODUCT USE INFORMATION" letter dated August 6, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.Contact the firm at 1-800-925-3100 ext. 5484 for assistance regarding this recall. |  | 5413 | 5,413 stairlifts | Worldwide Distribution-USA (nationwide) including the states of IL, IN, MI, OH, WI, IA, KS, MN, MO, NE, SD, AR, LA,OK, TX. AZ, CO, ID, MT, NM, UT, MT, AK, CA, HI, OR, WA, AL, FL, GA, KY, NC, SC, TN, MD, PA, VA, WV, CT, MA, ME, NH, NJ, NY, RI and VT and the countries of Canada, Nova Scotia and British Columbia | Terminated | January 29, 2014 | 541 | 5413
Z-0568-2013 | 63739 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Medline Angio Drape Pack Radiology-LF; a sterile custom single patient prescription procedure pack, including a waste bag, for use by the Interventional Radiology Department; Packaged for Medline Industries, Inc.; Reorder #DYNJ35993D |  |  | November 08, 2012 | 2012 | Medline Industries Inc | The non-sterile Waste Bag component in the custom pack was attached to the outside of the pack in its own packaging. The product was not labeled or identified as sterile, but was assumed to be sterile by the customer. | Packaging | The firm, Medline Industries, Inc., telephoned the consignee/customer on November 8, 2012. The phone call informed the customer of the product, problem and actions to be taken. The customer was instructed to inform the staff at the facility not to use the pack until it had been reworked to replace the non-sterile waste bag with the sterile equivalent. It was also noted that the product does not need to be returned as it is being replaced directly at the facility.If you have any question call 847-837-2759 or email lsimmons@medline.com. |  | 105 | 105 packs | Nationwide distribution: New York only. | Terminated | December 20, 2012 | 42 | 105
Z-0561-2013 | 63744 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | AutoTac Delivery Handle, REF 400-200, Rx Only, non-sterile, BIOHORIZONS, BIRMINGHAM, AL 35244. |  |  | November 19, 2012 | 2012 | BioHorizons Implant Systems Inc | The cleaning instructions provided may not be sufficient to remove residual contaminants after surgical use. | Component design/selection | BioHorizons sent an "URGENT DEVICE CORRECTION" letter on November 19, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers should contact the Customer Care Department at (888) 246-8338 for questions regarding this recall. |  | 174 | 174 units | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Brazil, Bulgaria, Canada, Cayman Islands, Chile, Colombia, Costa Rica, Ecuador, Egypt, Fiji, France, Germany, Greece, Hong Kong, India, Indonesia, Iran, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Panama, Peru, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Vietnam. | Terminated | September 24, 2013 | 309 | 1420
Z-0566-2013 | 63604 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | The Stryker Model 6390, Power-LOAD, is a power-loading cot fastener system designed to lift, lower, or steer ambulance cots into and out of the ambulance. |  |  | October 11, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Model 6390 Power-LOAD systems built between September 12, 2012 and September 27, 2012 may have been built with a load switch slide that does not meet specifications causing the lifting ability of the arms to be negatively impacted. | Component design/selection | All in process and Load Switch Slides within Stryker control were quarantined on 09/27/2012. Direct accounts were telephoned beginning 10/26/2012 with "Urgent Medical Device Correction" notices sent beginning on 11/30/2012. The notices explained the reason for the recall, hazards involved, and recommended actions. Stryker field service was to contact consignees to schedule replacement of the affected load switch slide. Consignees were asked to forward a copy of his notice to the new users and advise Stryker of their new location. Questions or concerns should be directed to (269) 389-6927 or (269) 389-6954. |  | 24 | 24 | USA, Nationwide including the states of PA, NM, MI, IA, CA, CO, WA, VA, IL, WI, NY, WY, MO, TX, and CT. | Terminated | August 30, 2013 | 323 | 24
Z-0569-2013 | 62764 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Pinnacle Cancellous Screw |  |  | November 28, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a voluntary recall of ten lots of the PinnacleÂż Cancellous Screw due to an incorrect description on the label. The label description reads Pinnacle Cannulated Screw and should read Pinnacle Cancellous Screw. | Labeling Change Control | The firm, DePuy Orthopedics, Inc., sent an "URGENT INFORMATION - MEDICAL DEVICE RECALL NOTICE" dated July 2012 to its customers via email and mail. The notice described the product, problem and actions to be taken. The customers were instructed to determine whether they would like to retain and use the product or return affected product to DePuy in Warsaw, IN, and to complete and return the Reconciliation forms to your DePuy Sales Representative or by fax to: 574-371-4939. DePuy sales representatives are expected to aid customers in returning of recalled products.For questions about recall information provided, please contact the Manager, Post-Market Surveillance, 574-372-7333 (M-F; 8 am - 5 pm EDT.) |  | 118 | 118 units | Nationwide distribution: USA states of: GA, MA, MN, SC and VA. | Terminated | March 20, 2013 | 112 | 118
Z-0560-2013 | 61593 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Advantage Workstation Volume Share 2 (versionAW4.4_04) with Volume Viewer 3 and 3.1 (version7.0 to version7.5 and version 8.1) and only when or more of these applications is installed: |  |  | May 29, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has became aware of two issues with the Advantage Workstation.1) A possible mismatch between the label of the tracked vessel and the underlying image associated with the cardio- vascular applications of the Advantage workstation that may impact patient safety.2) A refresh problem of the stenosis / aneurysmmeasurement tools associated with the cardio-vascular applications of | Software design | GE Healthcare sent two Urgent Medical Device Correction letters dated May 23, 2012, to all affected customers. Each letter addressed a separate issue. The letters described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. Customers were instructed to use the following precautions until corrected software was installed:1. Please avoid the use of leading or trailing spaces in Vessel Names2. Be sure to press "Show Tracking" after depositing new bifurcation points3. When eating a multi-point vessel tracking, remove and replace the entire branch, not only the final end-point.4. When analyzing multi-phase cardiac images, do not delete vessel branches from a phase other than the one in which it was created.Customers were instructed that the new application version that addresses this issue will soon be installed on all systems. Customers with questions or concerns were instructed to call the following:United States 1-800-437-1171.Japan 0120-055-919For other countries they were instructed to contact their GE Healthcare Service Representative.For questions regarding this recalla call 262-513-4122 |  | 548 | 548 (118 USA; 430 OUS) | Worldwide Distribution - USA including AL, CA, CO, FL, GA, IL, IN, IA, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX , UT, VA, WI and Internationally to the following countries:VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, TAIWAN, SWITZERLAND, SPAIN, SOUTH AFRICA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIA, ROMANIA, REUNION, PORTUGAL, POLAND, PHILIPPINES, PERU, PANAMA, NORWAY, NIGERIA, NEW CALEDONIA, NEPAL, MEXICO, MALAYSIA, KOREA, JAPAN, ITALY, IRELAND, INDONESIA, INDIA, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, ECUADOR, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BULGARIA, BRAZIL, BELGIUM, BELARUS, AUSTRIA, AUSTRALIA, ARGENTINA ALGERIA. | Terminated | January 17, 2013 | 233 | 548
Z-0563-2013 | 63807 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | CoverStainer Slide Rack |  |  | November 26, 2012 | 2012 | Dako Denmark A/S | The expected life time of the CoverStainer slide racks (CS 10330) does not meet established performance requirements, originating from variations in the manufacturing process. | Device Design | Dako sent a Voluntary Product Recall Notification dated November 13, 2012, to all affected customers. This voluntary recall notification was initiated because the expected life time of the CoverStainer slide racks (CS 10330) does not meet established performance requirements, originating from variations in the manufacturing process. The result can be that slides are falling out of the rack during the staining process. While this has been observed at only a few customer sites, the corrective action includes all CoverStainer installations. Customers were notified that Dako would correct the rack situation by using a new manufacturer and has imroved the rack design. Customers with questions were instructed to contact their local Dako representative.For questions regarding this recall call 805-566-3036. |  | 2650 | 25 Units US (50-100 racks per unit) 106 total | Worldwide Distribution - USA (nationwide) and the following countries:AustraliaFrance Rep. of KoreaBelgium Germany RussiaBrazil Israel Saudi ArabiaCanada Italy SpainChile Japan SwedenChina Mexico TaiwanDenmark Netherlands TurkeyEcuador Norway UKFinland Poland | Terminated | May 14, 2013 | 169 | 2650
Z-0567-2013 | 63777 | No | MHY |  |  | Neurology | PMA | Medtronic, DBS Extension Kit for Deep Brain Stimulation, Catalog # 7482, 37085, 37086. Rx Only, Manufactured in: Medtronic, Inc. Villalba, Puerto Rico. |  |  | December 03, 2012 | 2012 | Medtronic Neuromodulation | Medtronic has identified 140 Restore DBS Extension Kits (model numbers 7482, 37085, and 37086), which may contain a damaged winged connector boot. The connector boot may not have been molded properly and may be cracked. Approximately 14% of the 140 affected extension kits may contain a damaged radiopaque winged connector boot. | Nonconforming Material/Component | Medtronic beginning November 15, 2012 , initiated a retrieval of unused inventory in the field. In the US, three customers were identified. Medtronic field representatives visited these US accounts to facilitate product identification and removal. If product was removed a letter was provided to the facility describing the recall. The letter dated November 2012, "Medtronic Medical Device recall" was addressed to the Hospital. Beginning on December 3, 2012, Medtronic mailed healthcare professional letters to physician who have used the affected product. The Medtronic " Medical Device Recall" letter was dated November 2012. The letter identified the problem, the product involved in the recall, and provided advise if the product was used. A reply card was provided to each physician to confirm they have received and understood the communication. For further questions please call (763) 514-2000. |  | 140 | 140 units | Worldwide Distribution -- USA (nationwide) including the states of CT, MI, NH, NY, OH, PA and VT., and the countries of AUSTRALIA, JAPAN, BELGIUM, FRANCE, GERMANY, ITALY, NETHERLANDS, POLAND, SPAIN, SWITZERLAND, TURKEY and UNITED KINGDOM. | Terminated | November 14, 2013 | 346 | 140
Z-0558-2013 | 63776 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | MAYFIELD Composite Series Base Unit, Model A3101, a Swivel Adaptor (A3018) is an integral component of the Base Unit, Model A3101. |  |  | November 05, 2012 | 2012 | Integra LifeSciences Corporation | An investigation of an adverse trend of complaints conducted by Integra LifeSciences identified various product performance issues which could impact the functionality of the firm's Mayfield Composite Series Base Units and Mayfield Skull Clamps. In order to prevent additional complaint reports, the firm decided to recall and replace certain lot codes of the Base Units and Skull Clamps that were d | Under Investigation by firm | Integra issued an 'Urgent: Voluntary Medical Device Recall Notification with 'Recall Acknowledgment and Return Form' to all Integra Sales Representatives and Customers via visit and E-mail on November 5, 2012. The recall notice informs the customers of the firm's receipt of complaints pin-pointing quality issues with the MAYFIELD Base Units and MAYFIELD Skull Clamps which may impair their functionality. The customers are asked to sign and return the enclosed Recall Acknowledgement and Return Form so as to document their receipt of the recall notification and replacement products. The replacement products are to be delivered by the Integra NeuroSpecialist who is also responsible for returning the signed forms to Integra as documentation that the recalled devices have been replaced. Should the customer(s) have questions or concerns, they are directed to contact Integra Customer Service at: 1-877-444-1114,Option 3 for Technical Support. |  | 5 | 5 base units were distributed | Nationwide Distribution including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KY, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OK, OR, PA TN, TX, UT, VA, WA and WI. | Terminated | February 21, 2014 | 473 | 13
Z-0570-2013 | 63737 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | BD Phoenix PMIC-108 Panels, Catalog number 448418, labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, USA. |  |  | November 06, 2012 | 2012 | Becton Dickinson &amp; Co. | An antibiotic susceptibility test kit may produce false intermediate or false resistant results with clinically significant bacteria. | Nonconforming Material/Component | The firm, BD, sent an "URGENT PRODUCT RECALL" letter dated November 2012 via the telephone and fax to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and discard any remaining packages for replacement and to complete and return the Acknowledgement Receipt Form, whether or not you have any inventory remaining, via fax to: BD Regulatory Compliance at 410-316-4258. BD will issue replacements for the discarded materials.If further assistance is needed regarding replacements please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. |  | 12250 | 12,250 tests | Nationwide distribution: USA including states of: AR, CA, KY, OK, RI, WI, and WY. | Terminated | January 09, 2013 | 64 | 12250
Z-0565-2013 | 63714 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | BleaseFocus Anesthesia Workstation and 700/900 Series Ventilators are anesthesia gas machine. |  |  | November 02, 2012 | 2012 | Del Mar Reynolds Medical, Ltd. | The BleaseFocus Anesthesia Workstation and 700/900 Series Ventilators were shipped with manuals that describe a secondary O2 Supply Failure alarm which has not been incorporated into the system. The manuals were prematurely updated for a feature that is not yet available. | Labeling design | Spacelabs has sent the Urgent -Medical Device Correction-Spacelabs Healthcare BleaseFocus Anesthesia Workstation Blease 700/900 Series Ventilators letter, dated 16 November 2012, to US customers.On November 26, 2012, a customer letter (translated as necessary) wasemailed to all international subsidiaries and distributors of record.Customers are advised that the primary alarm will sound only in the event of an O2 supply failure and they will be provided with the corrected user and service manuals that do not reference a secondary O2 supply failure alarm.Customers with questions can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 3124 | 3 in US and 3121 Outside US | Worldwide Distribution--USA (nationwide) including states of North Carolina and Oklahoma and the countries of Algeria, Argentina, Bahrain, Bangladesh, Bolivia, Chile, China, Columbia, Dominican Republic, Ecuador, Egypt, Finland, Ghana, Great Britian, Guatemala, Honduras, Hong Kong, India, Indonesia, Iraq, Jamaica, Jordan, Lebanon, Libya, Malaysia, Mauritius, Mexico, Moldova, Nepal, Netherlands, New Zealand, Nicaragua, Nigeria, Pakistan, Palestine, Paraguay, Peru, Philippines, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sudan, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates , Uruguay, Venezuela, and Vietnam. | Terminated | July 11, 2013 | 251 | 3124
Z-0564-2013 | 62394 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Case Carton: |  |  | October 02, 2012 | 2012 | Becton Dickinson &amp; Company | BD received a complaint report for mixed syringe tips (oral and non-oral syringe tips) within the same lot. | Mixed-up of materials/components | Becton Dickinson sent a Urgent Product Recall letter/Customer Packing list dated October 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to review inventory and determine if they have any of the affected product. If so, it should be removed from inventory. The enclosed customer list should be completed as well. Any questions please call 1-201-847-4267. |  | 2786000 | 2,786,000 units | (USA) nationwide | Open, Classified |  |  | 2786000
Z-0593-2013 | 63057 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Bausch + Lomb Stellaris PC Vision Enhancement System, Posterior Fluidics Module (PFM). The system is intended for the emulsification and removal of cataracts, anterior, and posterior segment vitrectomy. |  |  | August 31, 2012 | 2012 | Bausch &amp; Lomb Inc | Posterior fluidics modules installed in the system may need to be recalibrated. | Process control | "Urgent Product Recall" notification letters dated August 31, 2012 were issued via first class mail, e-mail, and/or phone calls to affected customers. The product issue and the field correction was provided. Customers in the US may contact 1-800 338-2020 for questions. |  | 124 | 124 units | Worldwide Distribution was made in the U.S., Canada, Australia, Singapore, Korea, Switzerland, France, Germany, United Kingdom, Portugal, Italy, Spain, Ireland, Africa, UAE, Jordan, Turkey, Ukraine, and Lebanon.There was no government or military distribution. | Terminated | July 11, 2013 | 314 | 124
Z-0572-2013 | 63784 | No | CEP | Luteinizing hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Access hLH assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of luteinizing hormone (LH) levels in human serum and plasma using the Access Immunoassay Systems. |  |  | November 05, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter Inc. (BEC) has demonstrated that real time stability testing of hLH Calibrator of certain lots have failed the 13-month and 14-month time points. Per BEC's real time stability procedure, in order to verify expiration dating, one additional month is needed beyond the claimed expiration period. Therefore, these lots did not meet BEC's procedural requirements for a 12-month shelf-lif | Nonconforming Material/Component | Beckman Coulter Inc. (BEC) sent an "URGENT FIELD SAFETY NOTICE" dated November 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Per BEC's real time stability procedure, in order to verify expiration dating, one additional month is needed beyond the claimed expiration period. Therefore, these the affected lots did not meet BEC's procedural requirements for a 12-month shelf-life claim.Contact the firm at 952-448-4848 for assistance with this recall. |  | 14409 | 14409 | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia and Herzegovina, Botswana, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Djibouti, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Georgia, Germany, Ghana, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Macao, Malaysia, Mexico, Moldova Republic, Morocco, Myanmar, Namibia, Netherlands, Netherlands Antilles, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Rwanda, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, Yemen, Zambia, and Zimbabwe | Terminated | May 14, 2014 | 555 | 14409
Z-0607-2013 | 63667 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Segmented Cervix Applicator Set, part number GM11004310, a component used with the tube for segments and cervical sleeves |  |  | September 28, 2012 | 2012 | Varian Medical Systems, Inc. | The guide tube which is part of the Segmented Cervix Applicator Set is difficult to disconnect from cervical sleeve. | Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION - URGENT FIELD SAFETY NOTICE" letter dated September 28, 2012 to all affected consignees/Customers. The letter identified the affected product, description of problem and recommended user actions to be taken. The consignees/customers were instructed to immediately remove all affected guiding tubes from clinical use, and sequester the components for pending return to Varian Brachy Therapy. Consignees/Customers were requested to complete the attached proof of notification form and return it to Varian Medical Systems.For questions contact-AT&T; international direct dialing network (http://www.business.att.com/bt/dial_guide.isp. Then after hearing the AT&T; signature noise, and. "Please enter the number you are calling", dial 800-360-7909; UK:+44 1293-601-327 or US:+1-434-979-1540. To submit a support request or general inquiry, compose an email with a blank subject line to: brachyhelp@varian.com. |  | 15 | 15 units | Worldwide Distribution - U.S. state of Georgia and the countries of Canada, Germany, India, Poland, and Thailand. | Terminated | April 02, 2013 | 186 | 15
Z-0577-2013 | 63545 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SOL SYS 8IN CALC 1.5/18.0 |  |  | December 03, 2012 | 2012 | DePuy Orthopaedics, Inc. | An Engineering Analysis has determined the Solution Calcar Stem has the potential to result in impingement with a variety of heads. This impingement may not allow the stem and head to lock as intended. The amount of possible impingement varies depending on the femoral head. Worst case condition: Calcar stem with -2mm offset articuleze maximum of 0.937mm radial impingement. It is evident that a | Device Design | Deputh Orthopaedics sent an Urgent Information - Medical Device Recall Notice dated December , 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The notification letter instructed consignees to please cease further distribution or use of products immediately. The purpose of the communication was to inform them of this recall and request acknowledgement of receipt of this letter by their facility. Reconciliation forms should be returned to your DePuy Sales Representative or faxed to the following Fax Number: 574-371-4939.For questions regarding this recall call 574-267-8143. |  | 1732 | 1732 total product codes | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY and GUAYNABO, PR | Terminated | November 13, 2013 | 345 | 1732
Z-0571-2013 | 63788 | No | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | SpineSix consists of a motorized treatment table that is controlled by an integral computer. The system is operated by means of 12-inch Color LCD touch screen that displays treatment time, positioning and magnitude and direction of motion. |  |  | November 12, 2012 | 2012 | BioMotion Medical Systems, LLC | BioMotion is issuing a field correction for all models of the SpineSix (SpineSix01 and SpineSix02; serial numbers 1001 through 1054) to rectify all regulatory deficiencies involving the former owners of the SpineSix product and concurrently preparing the submission of a 510(k) application in order to receive premarket clearance for SpineSix. | No Marketing Application | BioMotion issued a field correction letter dated October 18, 2012 for all models of the SpineSix (SpineSix01 and SpineSix02; serial numbers 1001 through 1054) to all affected customers. The letter identified the product, problem, and actions to be taken. BioMotion is preparing the submission of a 510(k) application in order to receive premarket clearance for SpineSix. Contact the firm at 772-405-7899 for questions concerning this letter. |  | 28 | 28 units | Distributed in Florida, Georgia, Tennessee, New Jersey, New York, Pennsylvania, and Wisconsin. | Terminated | November 05, 2013 | 358 | 28
Z-0596-2013 | 61734 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medrad(R) Continuum MR Infusion System Standard Administration Kit |  |  | April 18, 2012 | 2012 | Medrad Inc | Bayer HealthCare is recalling all MEDRAD(R) Continuum MF Infusion System tubing sets (macrobore and microbore), catalog numbers MIK200A, MIL 200B, 3015153, 3015155, 3015156, 3015157, 3015158, 3015159, and 3015161. The firm is initiating this recall because this tubing exhibits a variation that, when used in conjunction with any Continuum pump, may result in flow rate accuracy performance issues. | Nonconforming Material/Component | The firm, Bayer HealthCare, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 17, 2012 via the USPS certified mail to its consignees/customers. The firm expanded the recall on May 11, 2012 and sent a second letter entitled "EXPANDED URGENT MEDICAL DEVICE RECALL" to its affected consignees/customers notifying them of the additional issue. The letters described the product, problem and actions to be taken. The customers were instructed to return pump now; contact customer support at 1-877-229-3767 for a return goods authorization (RGA); complete and return Response Form, include RGA numbers, via Fax to Bayer Compliance at 412-4060942, and/or wait to return the pump when they begin contacting customers with this subset of pumps in August to make arrangements for a replacement pump.If you have questions, please contact our customer support team at 1-877-229-3767 or visit www.medrad.com for more information. |  | 1229 | 1,229 Pumps total | Worldwide distribution: USA (nationwide) and countries of: Australia, Bermuda, Brazil, Canada, China, Chile, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Great Britain, Guam, Hong Kong, India, Israel, Italy, Japan, Korea, Lebanon, Mexico, Malaysia, Netherland, Poland, Portugal, Puerto Rico, Sweden, Singapore, Taiwan, Thailand, and Venezula. | Terminated | August 17, 2015 | 1216 | 1229
Z-0608-2013 | 37530 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Qualicheck 1+, Aqueous Quality Control, Level 2, S7140, REF 944-042. |  |  | June 06, 2006 | 2006 | Radiometer America Inc | The barcode for the ABL77 on the insert incorrectly enters a control range for pH of 7.90-7.95 into the analyzer. The insert correctly states that the control range is 7.40-7.45. | Process control | Radiometer sent Field Action Notes regarding the affected product. The note identified the product, the problem and the action to be taken by the customer. Instead of scanning the barcorde on the insert to enter the QC data into the anyalyzer, customers were instructed to manually key in the QUALICHECK type, lot number, and correct control range.For questions regarding this recall call 440-871-8900. |  | 840 | 840 Ampules | Distributed to KY, NC, and NY. | Terminated | February 05, 2013 | 2436 | 840
Z-0525-2013 | 63729 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Praxair Vantage Grab n Go - Gas Cylinder pressure regulator |  |  | November 16, 2012 | 2012 | Praxair Inc. | Isolated incidents of ignition inside Grab n' Go cylinders that had been knocked over or otherwise subject to significant physical impact. | Device Design | The firm, PRAXAIR, sent a "MEDICAL DEVICE CORRECTION" letter dated November 26, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the Grab'n Go operating instructions provided in the letter which include: storing the Grab'n Go systems indoors in a safe area; never store a system in an elevated location; handle systems safely and avoid jarring or dropping the system and locate at the following web link: www.praxair.com/GNGinstructions. The customers were also instructed to ensure that all of the Grab'n Go units that were sent by Praxair are sent to Praxair for servicing and refilling after use, in accordance with their normal process.If you have any questions, contact your sales representative at 630-320-4452 or email: Thelma_Brantley@Praxair.com. |  | 320451 | 320,451 units (255,690 US ; 55,884 Canada; 8877 OUS) | Worldwide distribution: USA (nationwide) and countries of: Brazil, Canada, China, Columbia, Costa Rica, India, Paraguay and Uruguay. | Terminated | March 24, 2014 | 493 | 320451
Z-0479-2013 | 63394 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System, Extended Brilliance Workspace (EBW), Model #728260 Philips Healthcare, Highland Heights, OH |  |  | September 28, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips is issued a Field Change Order to update Extended Brilliance Workspace (EBW) systems to software version 4.5.5. The update will provide a number of improvements to address nonhazardous reliability issues which will improve the overall quality and functionality of the EBW. | Software design | Philips sent a Urgent Medical Device Correction letter dated September 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Philips is issuing a system update to all of the Extended Brilliance Workspace (EBW) systems currently running software version 4.5.4. The update will provide a number of improvements to the overall quality and functionality of the EBW systems.If you need further information or support concerning this issue, please contact your local Philips representative. In the United States and Canada, please call the Philips Healthcare Customer Care Solutions Center at 1-800-722-9377, and follow the recorded menu options to reach a Customer Solutions Engineer; in all other countries please dial your local Philips Healthcare office. |  | 265 | 265 units | Worldwide Distribution--US (nationwide) including the states of AR, OH, OR, PA, TX VA &amp; WI., and the countries of Algeria, Australia, Brazil, China, Ecuador, Finland, France, Germany, Hong Kong, India, Italy, Lebanon, Libyan Arab Jamahiriya, Russia, South Korea &amp; United Kingdom | Terminated | December 24, 2013 | 452 | 265
Z-0611-2013 | 63676 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster MobiCath Bi-Directional Guiding Sheath. Small Curve catalog no. D140010, Large Curve catalog no. D140011. The MobiCath Bi-Directional Guiding Sheath is a flexible tipped percutaneous catheter introducer (or guiding sheath) designed for gaining access to the coronary systems. The device features adjustable tip geometry through use of a rotating handle to deflect the catheter. This catheter is assembled with a hemostasis valve and a sideport infusion line. |  |  | November 12, 2012 | 2012 | Greatbatch Medical | Greatbatch Medical observed some anomalies of the inner lumen of the MobiCath Bi-Directional Guiding Sheaths upon implementation of a new visual inspection tool. The anomalies may include loose, string-like liner material, liner abrasion (fraying), and/or attached string-like liner material (from scratches or damage on the inner wall of sheath lumen). | Nonconforming Material/Component | Biosense Webster sent an "Urgent Field Safety Notice" dated November 12, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Voluntary Field Removal Certification Form was attached for customers to complete and return. Contact the firm at 800-729-9010 for questions related to this issue. |  | 10201 | 10,201 units | Worldwide Distribution-USA including the states of AL, AK, AZ, ALR, CA, CO, CT, DC, FL, GA, HI, IL IN, IA, KY, MD, MA, MI, MN, MO, NE, NV,NM, NJ, NH, NY, NC, OH, OK OR, PA, RI,SC,SD, TN, TX, UT,VT, VA, WA, and WI, and the countries of AUSTRIA, BELGIUM, SWITZERLAND, GERMANY, DENMARK, SPAIN, FINLAND, FRANCE, UNITED KINGDOM, GREECE, HUNGARY, IRELAND, ISRAEL, ITALY, THE NETHERLAND, NORWAY, SWEDEN, AUSTRALIA, and CANADA | Terminated | May 15, 2014 | 549 | 10201
Z-0610-2013 | 63457 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Howmedica Osteonics; Stryker T2 Humerus System |  |  | August 06, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | During the manufacturing of the affected lot a manufacturing error has led to embrittlement of the tube material of the T2 Humerus Teflon Tube. | Nonconforming Material/Component | The firm, Stryker Orthopaedics, sent an "URGENT MEDICAL DEVICE REMOVAL" letter &amp; Acknowledgement form dated August 7, 2012 via FedEx Return Receipt to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to inform their users of this Medical Device Removal and pass this notice to all those individuals who need to be aware within their organization; return all affected products available at their location to: Stryker Osteosynthesis, 325 Corporate Drive, Mahwah, NJ 07430 REF: PFA#2012-124 or contact Stryker customer service; complete and sign the enclosed Product Recall Acknowledgement Form and Fax to 1-855-251-3635 or email to Recall Coordinator at janice.pevide@stryker.com and keep a copy of their records.If you have any questions, contact the Manager, Regulatory Compliance, at 1-201-972-2100 or email: colleen.omeara@stryker.com. |  | 100 | 100 units (73 units US; 27 units OUS) | Worldwide distribution: USA (nationwide) including states of: AL, CO, CT, IL, IN, MO, MS, NC, NJ, NY, OR,and TX; and countries of: Australia, Canada, France, Spain and United Kingdom. | Terminated | September 11, 2013 | 401 | 100
Z-0609-2013 | 63803 | [6 empty] | Diamigo i-Phone App |  |  | September 21, 2012 | 2012 | sanofi-aventis US, Inc. | Diamigo, a software application, was released through the i-Phone Global Store which inadvertently allowed global access to the application as opposed to the original intended restriction of access to Brazil. | No Marketing Application | Apple sent a message ("push notification") on September 21, 2012 to its customers. The message describes the product, problem and actions to be taken. The customers were instructed not to use the app to calculate insulin does and should delete if from their device(s). If you have additional questions about the app or this message, contact Sanofi's offices in your country. If you are in the U.S. contact Sanofi US Medical Information at 1-800-633-1610, for Brazil please contact 0800 703 0080. |  | 1600 | 1600 downloads (140 US; 1460 Foreign) | Worldwide distribution. | Terminated | July 22, 2015 | 1034 | 1600
Z-0613-2013 | 61212 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Total T4 Calibrators, Part Number: 33805. |  |  | July 27, 2011 | 2011 | Beckman Coulter Inc. | A recall communication was initiated because Beckman has identified that Access Total T4 calibrator lots 019818, 021654, and 024072 do not meet their 12 month expiration date claim. | Nonconforming Material/Component | A recall communication was initiated on July 27, 2011 with Beckman forwarding a Product Corrective Action (PCA) letter with attached PCA response form to all their customers who purchased the Access Immunoassay Systems Total T4 Calibrators. The letter provides the customers with an explanation of the problems identified and actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 days. Customers can call for product replacement in the US, by contacting Client Services at (800) 526-3821, option 1. Customers in Canada can contact Customer service at (800) 463-7828. Customers with questions regarding this notice can contact Customer Support Center at http:www.beckmancoulter.com or call (800) 854-3633 in the US and Canada. |  | 6634 | 6,634 units total (2,742 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Egypt, El Salvador, Ethiopia, France, Germany, Greece, India, Indonesia, Ireland, Italy, Japan, Jordan, Kazakhstan, Lebanon, Libyan Arab Jamahiriya, Liechtenstein, Macao, Malaysia, Mexico, Netherlands, Nigeria, Pakistan, Panama, Philippines, Puerto Rico, Romania, Russian Federation, Serbia, Singapore, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, and Zambia. | Terminated | December 26, 2012 | 518 | 6634
Z-0594-2013 | 63683 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Zimmer Spine Ardis Inserter, Rx, non-sterile |  |  | November 12, 2012 | 2012 | Zimmer, Inc. | Zimmer Spine has received reports of the PEEK Ardis Interbody Spacer breaking into fragments when too much lateral force is applied to the Ardis Inserter. If the PEEK Ardis Interbody Spacer breaks during implantation, surgical delays of up to 60 minutes could occur. Sharp edges of the broken implant may contact and damage adjacent structures, requiring surgical intervention. Observed health risks | Device Design | Zimmer Spine, sent a "MEDICAL DEVICE CORRECTION" memo dated November 12, 2012 to all consignees / customers describing product, problem and actions to be taken including revised surgical technique. Beginning the week of December 17, 2012, Zimmer began the immediate retrieval of the affected units as a Class I recall. December 20, 2012 nation wide press outlining the recall was issued. Customers may report adverse events or quality problems experienced with the use of these products to Zimmer Spine by calling 866-774-6368 between the hours of 8AM and 5PM CST or by email at qualitycompliance@zimmer.com. |  | 315 | 315 (172 in USA) | Worldwide distribution: USA (nationwide) including states of: AR, AZ, CA, CO, CT, DC, FL, GA, ID, LA, MD, MI, MN, MO, NC, NV, NY, OH, OK, PA, TN, TX, WI, and WV; and countries of: Australia, El Salvador, France and Singapore. | Terminated | February 25, 2014 | 470 | 315
Z-0612-2013 | 63632 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay System, |  |  | November 01, 2012 | 2012 | Beckman Coulter Inc. | It has been determined that, in some cases,pulley drive components have broken parts or have become loose, mispositioned, ordisconnected from their drive shafts. When these defects occur, the pulley may not turn at all, ormay turn at an incorrect speed, resulting in inadequate washing or mixing of the contents of thereaction vessel. | Device Design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 1, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return an enclosed Response Form. Contact Technical Support for questions regarding this notice. |  | 6858 | 6,858 units | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | December 23, 2013 | 417 | 6858
Z-0614-2013 | 63814 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COBAS AmpliPrep Instrument (CAP) with AMPLILINK (AL) Software v.3.3.5 |  |  | October 11, 2012 | 2012 | Roche Molecular Systems, Inc. | Certain COBAS AmpliPrep instruments using AMPILINK software v.3.3.4 had the sample clot detection flagging feature disabled during quality control release resulting in samples that may not be flagged as having clots. | Device Design | Roche sent an Urgent Medical Device Correction notification/Fax Back forms dated 10/11/12 to potentially affected customers. The letter identified the affected product, the issue, clinical significance, workaround, and required actions, along with contact information. Customers are to use the workaround until their Roche FSE has performed a Master Initialization on their affected instruments. Also, the letter should be filed for future reference. Questions should be directed to Roche Molecular Diagnostics Technical Support at 1-800-526-1247. |  | 7 | 7 units | Distributed in the states of GA, CA, NC, and MN. | Terminated | September 11, 2013 | 335 | 7
Z-0606-2013 | 63123 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Giraffe and Panda T-Piece Resuscitation System labeled ***T-Piece Neonatal Patient Circuit Kit Disposable with Size 1 Mask***M1091335 QTY: 10 T-Piece Disposable Circuit, QTY: 10, M1091316 T-Piece Disposable Circuit Kit with Mask Size O; M1091365 QTY: 10 T-Piece Disposable Circuit Kit with Mask, Size 1, Lot 1981713 *** Ohmeda Medical, A Division of Datex-Ohmeda, Inc. A General Electric Company Laurel, MD 20723 USA MADE IN USA*** |  |  | June 26, 2012 | 2012 | GE Healthcare, LLC | Disposable T-Piece circuits used with resuscitation systems are unable to achieve maximum Peak Inspiratory Pressure (PIP) during the pre-use check out procedure. | Nonconforming Material/Component | GE Healthcare sent an Urgent Medical Device Correction letter, dated June 26, 2012, with a return receipt to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer along with the product correction. Required Action for Users: Inspect existing stock of T-piece circuits to look for the Red inspection stamp, the T-pieces with the blue stamp are acceptable for clinical use. Separate affected circuits (with Red inspection stamp) from unaffected stock (with Blue inspection stamp). Destroy affected circuits that have the Red inspection stamp. Fill out the customer acknowledgement form and fax or forward to the GE contact address on the form. Further, it is recommended that healthcare professionals continue to emphasize the recommended pre-use checkout practices, and do not use a T-piece patient circuit that does not reach the maximum PIP level of 45 +/- 5 cmH20. Clinicians should always use the built-in Airway Pressure Manometer to verify the PIP value. For questions or concerns regarding this notification, customers should call the following phone number: 1-800-345-2700. |  | 209804 | 209,804 circuits | Worldwide Distribution - USA, including DC and the states of AL, AK, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IN, IL, KY, KS, LA, MA, MD, ME, MI, MN, MO,MS, MT, NC, NE, NJ, NM, NY, NV, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Albania, Australia, Austria, Bahrain, Bangladesh, Belgium, Canada, China, Czech Rep., Egypt, Finland, France, Germany, Great Britain, Greece, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, S. Korea, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, and Arab Emirates, | Terminated | April 14, 2014 | 657 | 209804
Z-0616-2013 | 63497 | [6 empty] | Brookstone Heated Body Bean, electrically heated |  |  | November 02, 2012 | 2012 | Brookstone Company, Inc. | Additional Warning Statement to prevent burns to the skin | Labeling design | Brookstone sent a "MEDICAL DEVICE CORRECTION" letter dated December 10, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Recall Information is on the website link from the homepage, www. Brookstone.com. Recall posted: http://www.brookstone.com/recall-information. Questions call Brookstone Customer Care at 800.846.3000 or visit any Brookstone Store. |  | 109956 | 109,956 units | Nationwide Distribution. | Terminated | June 13, 2013 | 223 | 109956
Z-0615-2013 | 63857 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | Flowtron Trio DVT Pump; an Rx intermittent pneumatic compression deep vein thrombosis pump; ArjoHuntleigh AB, Eslov, Sweden; Model 512003; a non-invasive prophylaxis system for reducing the incidence of deep vein thrombosis. |  |  | December 10, 2012 | 2012 | Arjo, Inc. dba ArjoHuntleigh | The Flowtron Trio DVT Pump has been marketed without proper FDA Approval. | No Marketing Application | ArjoHuntleigh sent an Urgent Field Correction Recall letters dated December 10, 2012, to all affected customers, informing the accounts that the pumps were marketed without proper FDA approval, and that ArjoHuntleigh is removing the pumps from the market and replacing them with a suitable alternative. The accounts were requested to disseminate the information to all organizations where the devices may have been transferred, and to complete the Customer Response Form and return it to ArjoHuntleigh by mail, fax or e-mail. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 2859 | 2,859 unts | USA (nationwide) including the states of Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Florida, Georgia, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Maryland, Maine, Michigan, Minnesota, Missouri, Mississippi, Montana, North Dakota, New Jersey, New York, Ohio, Oklahoma, Oregon, Pennsylvania, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin and Wyoming. | Open, Classified |  |  | 2859
Z-0619-2013 | 63825 | No | JIQ | Urinary protein or albumin (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Analyzing system consists of a small and portable analyzer and plastic microcuvettes. The microcuvette contains reagents deposited on its inner walls. The urine sample is drawn into the cavity by capillary action. The filled cuvette is inserted into the device where the contents of the cuvette are mixed through vibration. Within 90 seconds, the immunochemical reaction is completed and the turbidity is measured photometrically at 610 nm. The albumin concentration is proportional to the turbidity. When the end point is reached, the result is displayed in mg/L. |  |  | November 07, 2012 | 2012 | HemoCue AB | Incorrect version of Quick Reference Guide (QRG) was delivered with products in shipment made from April1 st 2011 and onwards. In addition, a warranty statement was missing in the analyzer package. | Labeling mix-ups | HemoCue sent an "IMPORTANT MEDICAL DEVICE CORRECTION NOTIFICATION" letter dated November 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Correction Notification Verification Form was included for customers to complete and return. Contact the firm at 562-668-5741 for questions regarding this notification. |  | 48 | 48 affected analyzers | Distributed in the states of FL, MA, CA, OR, NY, NC, MO, ID, IN, KY, PA, TX, OK, IL, GA, KS, and UT. | Terminated | March 14, 2013 | 127 | 48
Z-0618-2013 | 63769 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Custom Tubing Pack with or without coating ( Carmeda BioActive Surfact, Carmeda BioActive Surfact and Trillium Biosurface,or Balance Biosurface) with the following Model numbers: |  |  | November 19, 2012 | 2012 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is initiating an Urgent Medical Device Customer Notification. We have confirmed that certain Medtronic Intersept Tubing Packs contain packaging trays that are susceptible to damage that can compromise product sterility. | Package design/selection | Medtronic sent a "Urgent Medical Device Customer Notification" letter dated November 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately identify and quarantine all affected product and to return affected product, but if they have concerns about replacement than they recommend to inspect and return any damaged product..For further questions please contact your local Medtronic sales representative or Lifeline Technical Services at 1-877-526-7890. Outside the US, please call1+763-526-7890. We apologize for any inconveniences this issue may have caused you or your institution. |  | 1552 | 1552 | Worldwide Distribution--USA (nationwide) including the states of AZ, CA, CT, FL, GA, IL IA, MD, MA, MI, MS, NJ, NY, OH, OK, SC, TN, TX and WI. and the countries of CANADA and MALAYSIA. | Terminated | April 04, 2013 | 136 | 1552
Z-0620-2013 | 63915 | No | DSZ | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular | PMA | Companion External Battery, Part Number: 293001-001 |  |  | December 05, 2012 | 2012 | SynCardia Systems Inc. | Syncardia Systems is recalling the Companion External Batteries used with the Companion 2 Driver System because they may potentially be at risk of thermal events and/or other physical damage such as melting, charring and/or burning. | Process control | Syncardia Systems sent a" Urgent Field Safety Notice" dated December 10, 2012, to all affected customers. The letter identified the product the problem and what action the customer needs to take.. The letter informed the customers that there were 3 reported thermal battery events (2 of which occurred on-site at either SynCardia or at the battery manufacturer, Dow Kokam, and one incident that occurred in the field) which caused physical damage to the batteries including burning and/or charring. The batteries were not in use with a patient at time of the reported issue, and there were no reports of any adverse impact to any patient or user. Customers are instructed to follow the directions on the attached Return page, and return their customer response form to SynCardia as soon as possible.Customers with questions are instructed to call (520) 545-1234 or (520) 440-7593. |  | 81 | 81 units | Worldwide Distribution -- US (nationwide) | Terminated | June 07, 2013 | 184 | 81
Z-0617-2013 | 63855 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Fast-Cath Transseptal Guiding Introducer, 90, 8F, REF 406586, Length 60 cm, Sterile EO, Rx only, St. Jude Medical, 14901 DeVeau Place, Minnetonka, MN 55343 |  |  | March 28, 2012 | 2012 | St Jude Medical Inc | St. Jude Medical is recalling a single batch of Fast Cath Transseptal Guiding Introducers which was packaged as a 180 degree curve style introducer instead of a 90 degree curve style introducer as indicated on the labeling. | Process control | St Jude Medical hand delivered the" Product Recall" letter dated March 27, 2012, to the affected customers on April 5, 2012, and March 28, 2012. The letter identified the product, the problem and the action needed to be taken by the customer.The returned devices and those not distributed were quarantined at the SJM Minnetonka distribution center until their destruction.Customers were requested to to complete and return the reconciliation form.Further questions please call (651) 756-6526 |  | 26 | 26 | US Distribution only including the states of FL, KY and OK. | Terminated | July 12, 2013 | 471 | 26
Z-0623-2013 | 63862 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Space Pump Pole Clamps for the Infusomat Space Pump, model number 8713130. |  |  | March 23, 2012 | 2012 | B. Braun Medical, Inc. | Firm has become aware of the potential for breakage of the pole clamp locking grids, or rasters, in the Space Pole Clamps. | Device Design | B. Braun Medical, Inc sent an Urgent Medical Device letter dated March 23, 2012 , to all affected customers. The letter identified the product the problem and the action needed by the customer to take.B. Braun has implemented an improvement to the pump locking mechanism of t he older style pole clamps and will be upgrading any older style pole clamps that may remain within your facility.Please return the completed form from attachment 1 in the self-addressed envelope provided or fax the form to (610) 849-1197 within two (2) weeks of receipt\.If you require any additional information or have any questions, comments or concerns please contact your B. Braun Medical Customer Service at 800-627-7867. |  | 13078 | 13,078 | Worldwide Distribution--US (nationwide) including the states of AL, AZ, NC and TN. | Terminated | February 02, 2015 | 1046 | 13078
Z-0626-2013 | 63772 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Hyperbaric Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; list number 11005 Product Usage: The Plum A+ Hyperbaric Infusion System is exclusively designed for hyperbaric therapy in monospace and multiplace chambers. |  |  | December 07, 2012 | 2012 | Hospira Inc. | The door roller assembly on the Plum A+/A+3 infuser family of infusers has the potential to break which can lead to possible unrestricted flow and/or over-delivery during the removal of the IV administration set's cassette from the pump. | Device Design | Hospira Inc. sent "URGENT DEVICE FIELD CORRECTION" letters, dated December 7, 2012, to all affected customers. The letters identified the products, problem and action(s) to be taken until doors have been replaced. The customers were instructed to read the information provided carefully, including instructions to inspect the door roller assembly prior to loading a cassette for signs of damage. Customers are to complete the attached Reply Form and return it to the firm via fax at 1-877-650-8359. Customers were instructed to notify their consignees if they have further distributed the affected product. If customers require additional assistance, they are to contact Stericycle at 1-877-650-0362.. |  | 180 | 180 units | Worldwide Distribution -- USA, including Puerto Rico, and the countries of Austria, Australia, Argentina, Bahrain, Belgium, Brazil, Canada, Chile, Colombia, Denmark, Finland, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, United Arab Emirates, and Vietnam. | Open, Classified |  |  | 326612
Z-0622-2013 | 62413 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Standard Slingbar 450. |  |  | December 06, 2012 | 2012 | Hill-Rom, Inc. | Liko/Hill-Rom has received reports from facilities which indicate that serious injuries have been received related to the Standard Sling bar 450. There have been 6 complaints of injury related to the product over the past 32 years (6 out of 838,687,856 estimated patient uses). One of the injuries was associated with a patient during use, four injuries occurred when the slingbar was not in use wher | Device Design | Hill-rom sent an URGENT MEDICAL DEVICE RECALL notification letter dated December 6, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.. Customers were requested to please read, fill in completely, and return the Response Form to Hill-Rom immediately. This information will be compiled and reported to the applicable regulatory governing bodies. If we do not receive your response we will be required to contact you asking for your response.Continue to follow safe lifting practices, and do not allow unauthorized persons, especially children, to play around or operate the lift.If you have any questions concerning this Customer Safety Notification, please contact Hill-Rom Technical Support at 800-445-3720. |  | 261 | 261 units | Worldwide Distribution -- US (nationwide) including the states of AR, AZ, CA, CO, FL, GA, HI, IA, IL, IN, KS, KY, MA, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OK, OR, PA, TN, TX, UT, WA, WV, and WY., and the country Canada,We have not yet defined our European or International facilities at this time. | Terminated | November 06, 2014 | 700 | 261