## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities
Z-0577-2016 | 72727 | joint, knee, external brace | ITQ | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | ORTHO SPEC GRP CONTENDER POST-OP-KNEE BRACE LITE |  |  | November 12, 2015 | 2015 | Corflex | Knee brace could fail to maintain the locked position and a loss of balance or undesired flexion/extension may occur at the affected knee joint, | Nonconforming Material/Component | Corflex notified customers on November 12, 2015, initially via phone. On December 9, 2015, a follow-up letter dated December 8, 2015, was issued via email with read receipt. Customers have been instructed to return all product to Corflexs Manchester, NH manufacturing facility. This removal includes product that is new, in use and no longer in service. Customers have been provided instructions on how to contact our Customer Service department to initiate the return process and generate a Returns Material Authorization. UPS Call Tags are also being issued to facilitate the return of braces at no cost to the customer.Contact Corflex at service@corflex.com or via phone at (800) 426-7353, Monday through Friday, 8:00 am to 4:30 pm, EST |  | 42 | 42 units | Worldwide Distribution - US (nationwide) and Internationally to Canada , Taiwan , Singapore , and England. | Open, Classified |  |  | 1557
Z-0572-2016 | 72710 | No | DSB | Impedance plethysmograph. | Cardiovascular | Cardiovascular | 510(k) | CNSystems TASK FORCE MONITOR 3040i, plethysmograph, impedance. Noninvasively measure and display patient's hemodynamic parameters time synchronized using Impedance Cardiography (ICG), Electrocardiography (ECG), oscillometric Blood Pressure (oscBP) and continuous Blood Pressure (contBP). Furthermore Sympathic Tone and Parasympathic Tone are calculated by power spectral analysis of ECG (heart rate variability) and continuous blood pressure (blood pressure variability) as well as the Baroreceptor Reflex Sensitivity (BRRS). For use by medically trained personnel. It has been designed for diagnoses aiding and must not be used for vital sign monitoring of critically ill patients under absence of additional, suitable self-monitoring devices or medical trained personnel. The device measures continuously the subject's hemodynamic parameters without reporting any diagnosis. It is the physician's responsibility to make proper judgment based on these parameters. |  |  | December 01, 2015 | 2015 | Partners In Medicine Llc | Device non-conformity which may lead to an incorrect oscillometric measurement of upper arm blood pressure values in some subjects - mainly in tall and slim persons (overestimation of systolic blood pressure, underestimation of diastolic blood pressure). | Mixed-up of materials/components | CNSystems sent an Urgent Medical Device Recall Safety Notice dated December 1, 2015, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. CNSystems decided to perform a voluntary recall of the affected devices and to return the devices after repair (free of charge). For questions regarding this recall call 314-324-2470. |  | 30 | 28 worldwide 2 monitors distributed in the U.S. | Nationwide Distribution including MI and CA. | Open, Classified |  |  | 30
Z-0581-2016 | 72763 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Natus Quantum System with NeuroWorks Software. Catalog /Part Numbers: 013926. |  |  | November 11, 2015 | 2015 | Natus Neurology DBA Excel Tech., Ltd. (XLTEK) | During an internal testing, Natus Medical Incorporated, identified that in Neuro Works 8.1 with Quantum hardware, the incoming signal is displayed with reversed polarity. | Software Manufacturing/Software Deployment | The firm, Natus Neurology, sent an "Urgent Medical Device Correction Notification" letter on 11/11/2015. The letter described the the problem, the product involved in the recall and actions to be taken. Natus will provide further information once an updated software is released. The consignees were instructed to complete and return the confirmation form by mail to Senior Manager Quality and Regulatory Affairs, Excel-Tech, Ltd (XLTEK) Division, Natus Medical Incorporated, 2568 Bristol Circle, Oakville, ON L6H 5S1 or by email to :ots@natus.com.If you have any questions, contact Senior Manager Quality and Regulatory Affairs at 905-287-5055 or email: sanjay.mehta@natus.com. |  | 56 | 56 (43 US, 13 OUS) | Worldwide Distribution: US Distribution including states of: AL, AZ, AR, CO, FL, GA, MD, MA, MN, MS, NJ, NY, OH, SC, TN, TX and UT; and countries of: INDIA, KOREA, SPAIN, SWITZERLAND, THAILAND, UNITED KINGDOM, and CANADA.. | Open, Classified |  |  | 56
Z-0573-2016 | 72866 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew GENESIS(R) II/LEGION(TM) PATELLAR COMPONENT, RESURFACING ROUND, 32MM STANDARD, FOR CEMENTED USE ONLY WITH JOURNEY PEG, STERILE EO, Rx only, REF Number: 71934490 |  |  | November 09, 2015 | 2015 | Smith &amp; Nephew, Inc. | One lot of 35 mm patella implants is labeled as 32 mm. These implants have pegs which are spaced proportionally; therefore the mismatch prevents the pegs from properly aligning with pre-drilled holes during implantation. | Process design | The firm, Smith &amp; Nephew Inc, notified the distributors by Fed- Ex on 10/20/2015 and sent an "Urgent - Product Recall 1st Notification - Urgent R-2015-22" letter dated 11/9/2015 to their customers. The letter described the product, problem, and actions to be taken. The customers were instructed to inspect your inventory and locate any unused devices, quarantine them immediately, complete and return the Inventory Return Certification Form by fax to: +1-901-566-7975 or email to: FieldActions@smith-nephew.com, even if you do not have the product, contact Smith &amp; Nephew's Field Actions Department via email or fax (above) to obtain a return authorization (RA) number, and return any affected product to Smith &amp; Nephew, Attn: Global Field Actions, Global Distribution Center, 3303 E Holmes Road, Memphis, TN 38118.If you have any questions, call Director, Regulatory Affairs 901-399-5520 or email: jason.sells@smith-nephew.com. |  | 20 | 20 units | US Distribution to states of:FL, MO, IN and TX. | Terminated | March 17, 2016 | 129 | 20
Z-0478-2016 | 72715 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | The Fuhrman Pleural/Pneumopericardial Drainage Set |  |  | November 17, 2015 | 2015 | Stryker Sustainability Solutions | Stryker Sustainability Solutions is recalling the Fuhrman Pleural &amp; Pneumopericardial Drainage Set due to receiving complaints that it broke off in the pleural cavity while inserting the device into the patient. | Device Design | The firm, Stryker Sustainability Solutions, sent a "URGENT MEDICAL DEVICE RECALL" letter dated 11/17/2015 to its customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to discontinue the use of the product, and complete the Recall Effectiveness Check Form even if no product has been found in inventory. Customers should return the completed Recall Effectiveness Check Form to their local Stryker Sustainability Sales Representative, email to ssspfa@stryker.com, or mail to: Stryker Sustainability Solutions, 1810 West Drake Drive,Tempe, AZ 85283.If the firm indicates that affected devices remain in inventory, a prepaid shipping label will be issued for the return of the product. Customers will receive credit for all affected devices returned.Adverse reactions or quality problems experienced with the use of this productmay be reported to:- Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433x5555- http://www.stryker.com/productexperience/- The FDAs MedWatch Adverse Event Reporting program either online, byregular mail, or by fax.If you have any questions, contact Senior Regulatory Affairs Specialist at (888) 888-3433x5506 or email: chelsea.cullen@stryker.com. |  | 34 | 34 devices | US Distribution in the states of: AK, CA, WI, IN, FL, OR, NY, and NC. | Completed |  |  | 34
Z-0582-2016 | 72794 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Remote Patient Monitoring System, MDP 2000 Series, Part Numbers: MDP2040-0100, PGM340R3, or PGM340R4. |  |  | November 08, 2015 | 2015 | Vidco, Inc. | Testing at customer site showed unit Remote Patient Monitoring System MDP2040-0100 in a continuous trap condition, not allowing system to reset and reboot. Two customers recently complained of the system freezing and it could only be restarted if the user re-applied power. | Device Design | The firm began notifying consignees on Nov. 8, 2015 and followed up with the Advisory Notice letter, dated November 10, 2015.VIDCO identified the root cause and has corrected software that is currently going through validation. This software will be made available for updating the MDP2040-0100 as soon as it is released. The units may be updated on-site with the software program that will become available on the File Transfer Protocol (FTP) site. If you are unable to perform the procedure yourself, VIDCO will assist you in this effort.The current software version can still be used for monitoring as long as the user confirms the displays are still dynamically writing waves to the screen (i.e. waves will be changing in real-time).Consignees with questions can call (503) 641-1804 or 800-638-4326, or send email to sales@vidcoinc.com |  | 59 | 59 units | US Nationwide distribution in the states of AZ, CA, MD, NM, NJ, and OH. | Open, Classified |  |  | 59
Z-0583-2016 | 72814 | No | DZN | Dental hand instrument. | Dental | Dental | 510(K) Exempt | InterGuard 5.5mm Econo Refill, sizes, 4.0mm for short anatomical crowns and 5.5mm, part # 4012 |  |  | December 22, 2015 | 2015 | Ultradent Products, Inc. | Ultradent is recalling InterGuard 5.5mm Econo Refill, lot#s BC2GN and BC4L7 because product was received from the supplier with the floss tether hole not completely punched through and were distributed. | Nonconforming Material/Component | Ultradent Products sent an Urgent Medical Device Correction letter dated December 21, 2015 to affected customers via FedEx. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discard remaining inventory. Replacement product will be sent to each of the identified customers upon completion. For questions regarding the recall call 801-553-4366 or 801-553-4220. |  | 157 | 157 | Worldwide Distribution - US Nationwide and the countries of Austria, Canada, Colombia, Denmark, France, Germany, Great Britain, Iceland, and Portugal. | Open, Classified |  |  | 157
Z-0584-2016 | 72719 | No | KWL | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | BIRMINGHAM HIP RESURFACING MODULAR HEAD, REF numbers: |  |  | November 16, 2015 | 2015 | Smith &amp; Nephew, Inc. | The 16 devices distributed in the United States were incorrectly labeled as the monoblock Birmingham Hip Modular Head, with labeling not cleared for marketing in the US. The actual devices were monoblock Modular Femoral Head for hemiarthroplasty use. | Labeling design | The firm sent their field safety notices on 11/16/2015 by letter delivered by Fed Ex (overnight) and email. Non- responding consignees will be contacted within 4 weeks of the initial notice |  | 16 | 16 units in US | Worldwide distribution. US nationwide, France, UAE, Austria, Australia, Belgium, CANADA, Switzerland, Germany, Denmark, Spain, Finland, UK, Hong Kong, ISRAEL, India, Italy, Japan, South Korea, Netherlands, Norway, Poland, PORTUGAL, Sweden, SINGAPORE, Thailand, Turkey, and SOUTH AFRICA. | Open, Classified |  |  | 16
Z-0590-2016 | 72750 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGEN System REF 5980-37-01 Tibial Component Size Stemmed 3 Precoat, and REF 5986-37-01 Nonaugmentable Tibial Component Option CR/PS/LPS Size Stemmed 3. Tibial component of implanted knee prosthesis. |  |  | December 03, 2015 | 2015 | Zimmer Manufacturing B.V. | A product complaint was received stating that a carton labeled as a Precoat Stemmed Tibia Size 3 was opened for use and the contents were a non-augmented Stemmed Tibia Plate size 3. | Labeling mix-ups | Zimmer Biomet sent an "Urgent Medical Device Recall-Lot Specific" letter to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to do the following:Review the notification and ensure affected personnel are aware of the contents.2. Locate all affected product identified above and quarantine them immediately.3. Carry out a physical count of all affected product in your territory and complete the Inventory Return Certification Form (Attachment 1). Email a completed copy of Attachment 1 to corporatequality.postmarket@zimmer.com.4. Return the recalled product along with the completed Inventory Return Certification Form (Attachment 1). Clearly mark the outside carton of each product return shipment made as Recall. 5. Please notify Zimmer Biomet of any hospitals to which you have further distributed the affected product. Supply the information for any hospitals that you have identified, as well as the affected surgeons, using the provided spreadsheet template. The template will be emailed to you for completion. Please return to corporate quality.postmarket@zimmerbiomet.com.6. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-800-348-2759 between 8:00 am and 8:00 pm EST. |  | 768 | 384 distribution events; 1 unit was returned leaving 383 units . | Worldwide Distribution to Germany, Belgium, Denmark, France, UK, Greece, Ireland, Netherlands, Norway, Russian Fed., Saudi Arabia, Sweden, and Thailand. | Open, Classified |  |  | 768
Z-0589-2016 | 72652 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Maquet HCU 30 Coolant Type R134A Filling weight 250 g 120 V 50-60 Hz 1920 W |  |  | November 13, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Failure of the HEATER-COOLER UNIT (HCU 30) Type 2 power supply board. A drop of more than 10% voltage can cause overheating and subsequent failing of the R56 thermistor of the power supply board. Board failure typically occurs when the compressor is switched on (prior to patient use); however, it can occur during patient use. This can shut down the compressor or the entire HCU 30 Type 2 unit. | Device Design | MAQUET sent "Product Recall Urgent Medical Device Field Correction" notices and "Medical Device Field Correction Response Forms" dated 11/13/2015 via UPS to affected customers. The notification identified the issue and how to identify the affected unit. The notification instructed customers to always keep a replacement heater-cooler unit on standby in the event of a unit failure; complete the attached Field Correction Response form (fax to 1.973.629.1518 or email to HCU30@maquet.com; and that they would be contacted by a representative of the Maquet Service Team to schedule an onsite service to determine if the power supply board in their HCU 30 Type 2 unit is affected adn replace the power supply board as necessary. If the customer had any questions they are instructed to contact their Maquet Service Representative or Maquet Technical Support at 1.888.627.8383 (Press option 3, followed by Option 1 and then option 1 again), Monday through Friday between the hours of 8:00 am and 5:00 pm. |  | 100 | 100 Units (distributed in United States) | United States Nationwide Distribution | Open, Classified |  |  | 100
Z-0585-2016 | 72774 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SYNTHES; 13.5MM Medullary Reamer Head; 352.135 NON STERILE. Part of the Flexible Reamer Set for orthopedic surgical procedures. |  |  | December 01, 2015 | 2015 | Synthes (USA) Products LLC | The 13.5mm Medullary Reamer head, lot #F-17180, is etched as dimension 13.5mm while the actual dimension is 14mm. | Process design | On Dec 1, 2015, Synthes issued a recall notification letter and requested that all consignees review inventory and contact Synthes for a return authorization number. The verification form is to be completed and sent to Stericycle. |  | 65 | 65 | US nationwide distibution. | Open, Classified |  |  | 65
Z-0588-2016 | 72739 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific ACUITY Pro Coronary Sinus Guide Catheter, 9 Fr. Model 8104, 8105, 8106, 8107, 8108, 8109, 8110, 8111, 8112, 8113, 8114, 8115, 8116, 8117, 8118, and 8119. |  |  | November 25, 2015 | 2015 | Boston Scientific Corporation | Potential for separation of the proximal handle from the shaft. | Process control | Boston Scientific personnel began retrieving product from hospital shelves on November 25, 2015. The hospital was given a Medical Device Retrieval letter. The letter provided the reason for product retrieval, identified affected product, asked for product to be segregated to ensure that it will not be used, and to contact Boston Scientific for a return authorization number to ship product back to Boston Scientific. A Certificate of Removal was asked to be completed and emailed to BSCFieldActionCenter@bsci.com Questions can be addressed to Technical Services at 1-800-227-3422. |  | 900 | 900 units | Worldwide Distribution - US (nationwide) and Internationally to Argentina, Australia, Austria, Belgium, Colombia, Cyprus, Czech Republic, Egypt, Finland, France, Germany, Great Britain, Greece, Hong Kong, Ireland, Italy, Japan, Malaysia, Netherlands, Poland, Portugal, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, and Thailand. | Terminated | April 08, 2016 | 135 | 900
Z-0602-2016 | 72840 | No | PFJ | Hygroscopic Laminaria cervical dilator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Gynecare Thermachoice III Uterine Balloon Therapy System Ethicon, Inc. |  |  | December 03, 2015 | 2015 | Ethicon, Inc. | Stability data does not substantiate the labeled two-year shelf life of affected product. | Device Design | Ethicon, Inc. sent "Urgent: Medical Device Recall (REMOVAL)" notifications and Business Reply Forms dated 12/3/2015 on same date via UPS Next Day Mail to its customers. The notification informed the customers of the issue with the product; how to identify affected product and action required. The customers were instructed to examine their inventory immediately to determine if they have affected product on hand and quarantine the affected product(s); remove the affected product and communicate the issue to relevant operating room or materials management personnel, or anyone else in their facility who needs to be informed; if any affected product has been forwarded to another facility, contact that facility to arrange return; complete and return the enclosed Business Reply Form (BRF) confirming receipt of this notice within three (3) business days to Stericycle by fax at 1-866-792-5453 or by email at Ethicon7427@stericycle.com, return BRF even if you do not have affected product. Stericycle is handling returns. If the customer requires any assistance with returning product or have any questions, they were instructed to contact the Customer Support Center at 1-877-ETHICON (1-877-384-4266).Ethicon, Inc. issued an update to their 12/3/2015 letter in a notification dated 12/23/2015. The 12/23/2015 notification clarified the lots of the product subject to their recall previously communicated in their 12/3/2015 letter. There are three (3) lots of product subject to the recall that were inadvertently omitted from the original lot listing. The notification stated that "All THERMACHOICEÂż Catheter lots with an expiration date prior to December 10, 2017 are subject to this recall and are required to be returned. The lot number listing of all product subject to the recall is contained in Attachment 1. The expiration date and lot number can be determined by using the Product Identification Tool attached hereto at Attachment 2." The attachments included the updated pr |  | 96124 | 96,124 units | Worldwide Distribution: US (nationwide) including Puerto Rico and countries of: India, Belgium, Argentina, Mexico, Brazil, Canada, Colombia, Ecuador, Singapore, Aruba and Venezuela. | Terminated | May 30, 2017 | 544 | 96124
Z-0601-2016 | 72780 | protector, skin pressure | FMP | Skin pressure protectors. | General Hospital | General Hospital | 510(K) Exempt | 3M Coban Self-Adherent Wrap Shipper carton label identified the product as 3M Coban LF Latex Free Catalog #2081, lot number 2018-10AM. |  |  | December 01, 2015 | 2015 | 3M Company - Health Care Business | Product mislabeled as latex free. Actual Catalog number 1581, lot 2020-10AN. contains latex. This could pose a health risk to users with sensitivity or allergic reaction to latex, which could range from mild to severe. | Process control | Consignees were sent via certified mail on 12/1/2015 a 3M letter dated November 25, 2015 and an emal on 12/5/2015 for the updated 3M "Urgent Medical Device" letter dated November 25, 2015. The letter was addressed to 3M Health Care Distributor. The letter described the problem, device involved in the recall and the action required. Advised consignees to quarantine the product, notify their customers and to complete and return the "Product Recall Form". Distributors were also provided with a "Customer" letter for distribution. For questions contact 3M Customer Helpline 1-800-228-3957. |  | 167 | 167 cases (30 rolls/case) | US nationwide distribution. | Open, Classified |  |  | 167
Z-0599-2016 | 72707 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | NIPRO SafeTouch II Safety AVF Needle, 17G x 1 Âż inch. Item code FS+173230BC. |  |  | April 01, 2014 | 2014 | Nipro Medical Corporation | Device has the potential to have torn wings. | Process control | The only one Distributor was notified (2014) by e-mail by the firm of the quality deficiency before the product arrived at their facility. It was immediately returned to Nipro Medical when it arrived. |  | 160000 | 160,000 pieces. | US distribution including TN. | Completed |  |  | 160000
Z-0603-2016 | 72825 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | AMSCOÂż Warming Cabinet. Manufactured by: STERIS Corporation |  |  | November 20, 2015 | 2015 | Steris Corporation | The firm discovered that information printed on the quick reference magnet is incorrect; specifically, Fahrenheit to Celsius temperature conversions and the default Warming Cabinet temperature. | Process change control | The firm, Steris, sent an "URGENT VOLUNTARY FIELD CORRECTION NOTICE" letter dated 12/14/2015 to its customers. The letter described the product, problem and actions to be taken. The firm included an updated quick reference magnet for each affected unit with the letter. The customers were instructed to replace your existing magnet with the enclosed magnet. If you have questions regarding this matter, please contact Product Manager, at 440-392-7705 or STERIS Customer Service at 1-800-548-4873. |  | 2319 | 2,319 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Austria, Bahrain, Canada, China, Colombia, Costa Rica, France, Hong Kong, India, Indonesia, Israel, Kuwait, Lebanon, Mexico, Panama, Peru, Republic of Korea, Russian Federation, Saudi Arabia, Singapore, Spain, Taiwan, Thailand, Turkey, Uganda, United Arab Emirates, Viet Nam &amp; Virgin Islands. | Open, Classified |  |  | 2319
Z-0596-2016 | 72879 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Lumify Diagnostic Ultrasound, Catalogue Number: 795216 |  |  | December 09, 2015 | 2015 | Philips Ultrasound, Inc. | Color Flow direction is displayed incorrectly in Lumify 1.0. The system displayed Color Flow direction does not correctly represent the annotated Color Bar or Velocity Markers. (e.g.) When the Color Bar conveys that Red is to be displayed for color flow toward the Transducer, the system displays Blue for color flow toward the Transducer. This could result in misdiagnosis in some studies. | Software design | Philips began mailing the URGENT - Medical Device CorrectionPhilips Model Lumify Ultrasound System, FSN 79500372, dated 2015 DEC 11, to consignees.Consignees are instructed to do the following:1. Check the version of software on your system using the instructions provided above on How to Identify Affected Products. Philips has corrected this issue in version 1.0.1 and higher.2. Connect the tablet to the internet.3. The system will alert you to download the upgrade from Google Play.4. Follow all on screen to upgrade the software version.If consignees cannot promptly connect to the internet to download the upgraded software, Philips recommends that you do not use Color Flow. In accordance with the subscription terms and conditions, you will be required to reconnect the system to the internet within no more than 30 days. If consignees do not connect their device to the internet after 30 days, they will be unable to use the system.Consignees with questions should contact your local Philips representative or Philips Customer Service at 1-800-722-9377. |  | 21 | 21 units in the US only | distributed in CA, CT, ND, NV, OR, RI, TN, and WA. | Open, Classified |  |  | 21
Z-0600-2016 | 72753 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ActiV.A.C. 300ml Canister (with Gel) for ActiV.A.C. Therapy System for the application of V.A.C. Therapy in the acute extended and home care settings. The software-controlled therapy unit applies negative pressure wound therapy. The dressing, to which the therapy unit is connected, enables distribution of the negative pressure across the surface of the wound, while the tubing transfers accumulated fluids to the canister. |  |  | November 25, 2015 | 2015 | KCI USA, INC. | There is potential for small holes, the size of a pin, to develop in the pouches of the ActiV.A.C. Sterile Canisters. | Package design/selection | Customer notification letters were sent out the week of November, 23, 2015. The international distributors will be responsible for notification to their sub-consignees. |  | 12728 | 12,728 units | Japan and Australia | Open, Classified |  |  | 12728
Z-0594-2016 | 72446 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | BIOMET 3i, PROVIDE IMPRESSION COPING KITS, Catalog # PRK484, PRK485, PRK654, and PRK655 |  |  | September 21, 2015 | 2015 | Biomet 3i, LLC | Inadequate biocompatibility testing. | Process control | The firm, Biomet 3i, sent an "URGENT DEVICE RECALL NOTICE" letter dated September 16, 2015 to both Domestic and Foreign Customers via traceable courier. The letter described the product, problem and actions to be taken. The customers were instructed to review the notice and check your inventory for the affected units; immediately quarantine and remove all affected products from service; if you have product to return, call Biomet 3i Complaints Department at 1-800-443-8166 or 1-561-776-6700 to obtain a CMP# and return affected product to Biomet 3i Returns, 4555 Riverside Dr., Palm Beach Gardens, FL 34410; complete and return the attached Business Reply Form via fax to: +1-561-514-6316 or email to: postmarket@biomet.com, and maintain a copy of the notice for your records.For assistance or any other questions that you may have, please contact Biomet 3i Complaints Department at 1-800-443-8166 or 1-561-776-6700. The Complaints Department is available 8:00am to 6:00pm (Eastern), Monday through Friday. |  | 30111 | 30,111 devices. | Worldwide Distribution: US (nationwide) and countries of: Canada, Argentina, Austria, Australia, Brazil, Chile, China, Colombia, Costa Rica, Germany, Dominican Republic, Spain, France, Great Britain, Hong Kong, Ireland, Israel, India, Korea, Kuwait, Lithuania, Mexico, Netherlands, Portugal, Saudi Arabia, Sweden, Singapore, Taiwan, Venezuela, and South Africa. | Open, Classified |  |  | 30111
Z-0597-2016 | 72751 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Software version VD10E for Syngo X-Workplace; Picture archiving and communication system. |  |  | November 18, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Potential post-processing software issue when using tabcard "4D" on X-Workplace with software version VD10E. A too small measurement in the MPRs of a volume when images are acquired with a CT scanner with a tilted gantry. In a 2x2 layout, the 4D tabcard will show an incorrect, too small length measurement. This could result in selecting a device of the wrong size, which then needs to be exchanged. | Software design | Siemens sent an Important Customer Safety Notice dated November 17, 2015. The Notice identified the product, the problem, and the action to be taken by the customer. The potential issue will be remedied with a software update via Update Instructions AX067/15/S. Siemens service organization will contact customers to arrange a date for the installation of the software update. Customers with questions were advised to call 800-888-7436. |  | 73 | 73 units | Nationwide Distribution including:AL, AZ, CA, CO, FL, IA, ID, IL, IN, MA, MD, MI, MN, MO, MT, NC, NY, OH, OK, PA, SC, TN, TX, VA, and WI. | Open, Classified |  |  | 73
Z-0595-2016 | 72851 | No | FPD | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MED-RX Single Use REF 54-3680R, 8 FR. x 36" (91CM) Polyurethane Feeding Tube For Pediatric Use with Orange Radiopaque Stripe and Enteral Connector, Made in Canada, Benton, Oakville, Ontario, Canada |  |  | December 11, 2015 | 2015 | Benlan, Inc. | The product may be packaged with a 5 or 6.5FR tube instead of an 8FR size. | Under Investigation by firm | Benlan, Inc. sent a Product Recall Notification letter dated December 11, 2015, to all affected customers. Customers were asked to examine their stock and to discontinue/use of the product and return product. The letter included a response form which was to be returned via fax to 905-825-9600. Customers with questions were instructed to call 905-825-9300. |  | 120 | 120 units | Worldwide Distribution - US including GA &amp; Internationally to Canada. | Open, Classified |  |  | 120
Z-0604-2016 | 72869 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncentra External Beam - VMAT - Radiation therapy planning system |  |  | December 18, 2015 | 2015 | Elekta, Inc. | When using the option "Tumor Overlap Fraction" in VMAT planning it has been observed that in rare cases the system does include an organ at risk as target volume. This could result in open MLC, and open jaws in areas away from the target volume. | Software design | Elekta sent an Important Field Safety Notice 806-01-ETP-001 to all affected customers on December 18, 2015. The notice informs users of the issue and that should the issue occur it could result in overdosage of organs at risk outside the intended treatment volume. Users are strongly advised to performproper Quality Assurance for all treatment plans before delivery of the first fraction to the patient. To prevent the issue from occurring, the workaround is to not use the Tumor Overlap Fraction for VMAT planning.and not to use the Tumor Overlap Fraction for VMAT planning. The issue will be solved in Oncentra External Beam version 4.5.2. The notice included an acknowledgement form which is to be returned to Elekta. For questions contact your local Elekta representative. |  | 154 | 154 | Worldwide Distribution - US Nationwide in the states of CA, NY, PA, and the countries of Austria, Belarus, Canada, China, Denmark, Finland, France, Germany, Greece, India, Italy, Japan, Netherlands, Norway, Poland, South Korea, Spain, Sweden and United Kingdom. | Open, Classified |  |  | 154
Z-0611-2016 | 72570 | No | DAP | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Siemens Acute Care DDMR TestPak |  |  | November 11, 2015 | 2015 | Siemens Healthcare Diagnostics Inc | Stratus CS Cartridge Barcode labels may not scan on the Stratus CS systems | Labeling design | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on November 11, 2015 both in the United States and internationally for communication with affected customers. This notice informs customers of the issue and provides mitigation instructions. Customers with questions were advised to contact their Siemens Customer Care Center or their local Siemens technical support representative. For questions regarding call 781-269-3000. |  | 2647 | 2647 | Worldwide Distribution - US (nationwide), Algeria, Armenia, Australia, Austria, Bangladesh, Bosnia Herzegovina, Canada, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Libya, Malaysia, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., and United Kingdom. | Open, Classified |  |  | 12288
Z-0606-2016 | 72787 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter INTERLINK System, Y-Type Solution Set, 78 (2.0 m), 2 Injection Sites, Male Luer Lock Adapter, 1C8581 s, |  |  | November 30, 2015 | 2015 | Baxter Healthcare Corp. | Potential for dark loose particulate matter (400 microns or smaller) on the spike of the spike adapter or solution set. | Under Investigation by firm | Baxter sent an Urgent Product Recall letter by letter on 11/30/2015 to affected customers. The recall was extended to the consumer/user level. The letter instructed customers to locate and remove all affected product, contact Baxter Healthcare Center for Service to arrange for return of affected products and credit and complete the enclosed Baxter Customer Reply Form and return it to Baxter by either fax or scanned email. |  | 9600 | 9600 units | US Nationwide Distribution and the country of Canada | Open, Classified |  |  | 96000
Z-0624-2016 | 72701 | Yes | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NIPRO SET / SLIMLINE BLOOD TUBING SET for Hemodialysis with Transducer Protectors and Priming Set |  |  | March 30, 2015 | 2015 | Nipro Medical Corporation | Blood foaming was found during evaluation of the product. | Device Design | NIPRO contacted their consignees by email 9/25/2013 of the affected product. Customers were instructed to return product back to Nipro Medical Warehouse or discard at point of use. Please contact Director of QA/RA at Nipro Medical Corporation for any additional information or return correspondence to JessicaO@nipromed.com / 305.599.7174x249. |  | 1484 | 1,484 pieces. | US Distribution to states of: TN and NJ. | Terminated | February 04, 2016 | 311 | 1484
Z-0622-2016 | 72962 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | UniAxial Pedicle Screw &amp; UniAxial Reduction Pedicle Screw |  |  | November 10, 2015 | 2015 | OrthoPediatrics Corp | Response Spine UniAxial Pedicle &amp; UniAxial Reduction Pedicle Screws may have been manufactured incorrectly, thus preventing the driver from fully engaging into the hexagonal drive feature of the screw. | Process change control | The firm, OrthoPediatrics Corp., sent an "URGENT MEDICAL DEVICE RECALL" letter dated 11/13/2015 to affected customers via email. The letter included a description of the reason for the recall, affected product, consignee responsibilities, and actions to be taken. The customers were instructed to take immediate possession of all the identified devices; quarantine them and immediately return them to OrthoPediatrics, and complete and return the Mandatory Reply Form via fax to (574) 269-3692 or email: MF0x@OrthoPediatrics.com.If you have any questions, please contact VP-Regulatory Affairs (574) 268-6379 or email: mfox@orthopediatrics.com. |  | 102 | 102 | US Distribution to states of: AL, AZ, CA, DE, FL, GA, HI, KY, LA, MI, MO, MN, MS, NC, NJ, NY, OH, OR, TN, and TX. | Completed |  |  | 102
Z-0623-2016 | 72789 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Cavity 8 Gauge Spine Wand indicated for resection, ablation, and coagulation of soft tissues and hemostasis of blood vessels in spinal procedures. |  |  | October 26, 2015 | 2015 | ArthroCare Corporation | There is cracking in the tray when the device is snapped into place. The crack is through the full thickness of the tray, compromising the sterile barrier. | Under Investigation by firm | Affected consignees will be notified via letter beginning 11/2/15. |  | 262 | 262 units | Worldwide distribution. US nationwide including OH, MO, AZ, AL, CA, FL, SC, GA, and PA; Spain, France, United Kingdom, Germany, Switzerland, Singapore, Russia, Turkey, and Ecuador. | Open, Classified |  |  | 262
Z-0625-2016 | 72832 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Zimmer, Virage OCT Spinal Fixation System, Item Number: 07.01738.001, a posterior system intended to achieve an implant and construct of the Occipital-Cervical-Thoracic spine (Occiput-T3). |  |  | November 17, 2015 | 2015 | Zimmer Biomet, Inc. | Zimmer Biomet is initiating a recall of the Virage Occipital Eyelet, as it was found through product testing that the staking of the set screw was out of specification and can interfere with the insertion of the occipital rod when the set screw is staked too low. | Process control | Consignees were sent a Zimmer Biomet "Urgent Medical Device Recall" letter dated November 17, 2015. One letter was addressed to Risk Manager and Surgeons and a separate letter was sent to Distributors, Sales Representatives, and Distributor Operations Mangers. The letters described the problem and the product involved in the recall. The letter described the "Risk" and the responsibilities of consignees and requested consignees to quarantine the product. Risk Managers and Surgeons were requested to complete and return the Certificate of Acknowledgment form. Distributors, Sales Representatives, and Distributor Operations Mangers were requested to complete and return the Inventory Return certification Form.For questions consignees can call the customer call center at 1-866-774-6368 between 8:00 am and 5:00 pm EST. |  | 372 | 372 | US: AL, AZ, CA, CO, FL, GA, ID, MI, MN, MO, NV, NJ, OK, OR, PA, SD, TN, TX, WA, WI.OUS: AUSTRALIA, GERMANY, SWITZERLAND, SWEDEN. | Open, Classified |  |  | 372
Z-0626-2016 | 72809 | No | PHX | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AltiVate Reverse Humeral Broach, Size 10 Catalog number 804-06-040, Size 14 Catalog number 804-06-044, and Size 16 Catalog number 804-06-046. Orthopedic surgical instrument for use with Encore Shoulder prosthesis. |  |  | December 08, 2015 | 2015 | Encore Medical, Lp | The trial shell mating feature on the broach is undersized. | Nonconforming Material/Component | The recalling firm notified customers of the recall via letter on 12/8/15. |  | 23 | 23 units | Throughout the US to the states: FL, PA, ID, CA, OH, AZ, TN, NV, WA, MS, GA, MA, KS, TX | Terminated | March 01, 2016 | 84 | 23
Z-0586-2016 | 72867 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | MOVES System, an emergency ventilator with suction, oxygen concentrator and multi-parameter patient monitoring capabilities. |  |  | November 19, 2015 | 2015 | Thornhill Research Inc | Batteries found unresponsive after storage, not recognized by main device, and are perpetually enabled. | Component design/selection | Thornhill Research sent a Recall Notification letter dated December 15, 2015 to affected customers. The letter identified the affected products, problem and actions to be taken. Customers were instructed to update their device software.For questions customers were instructed to call 416-597-1325. |  | 83 | 83 units | Worldwide Distribution: US distribution in state of: Virginia and to the country of: Singapore. | Open, Classified |  |  | 261
Z-0628-2016 | 72885 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | FMC7 FITC, Catalog No. 340918 Analyte Specific Reagent. |  |  | December 14, 2015 | 2015 | Becton, Dickinson and Company, BD Biosciences | Becton, Dickinson and Company (BD) has determined that the FMC7 FITC (ASR) is exhibiting functional degradation. The functional degradation has the effect of reducing the percent positive from 100% down to approximately 87%. | Under Investigation by firm | BD Biosciences sent an Urgent Product Recall letter dated December 2015, to all affected customers to inform them that Becton, Dickinson and Company (BD) has determined that the FMC7 FITC (ASR) is exhibiting functional degradation. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to return the Tracking/Verification form to BD Biosciences. Customers are instructed to contact BD Customer Support at (877) 232-8995 (prompt 3, 3) in the US. Customers outside of the US are instructed to contact their local BD Biosciences or distributor. |  | 1388 | 1,388 units | Nationwide Distribution | Open, Classified |  |  | 1388
Z-0627-2016 | 72817 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medrad Source Administration Sets, Catalog number INT CSS. For use with MEDRAD Intego PET Infusion pump Systems, in nuclear medicine diagnostic procedures. |  |  | November 30, 2015 | 2015 | Bayer Healthcare | The firm's investigation of the Medrad Intego indicates that the inner diameter of the saline tubing is out ofspecification and may result in the system failing to position the dose correctlyduring extraction, potentially leading to a low measurement of the extracted dose orto a higher dose being extracted. The occluded SAS tubing may impair the primingfunction and prompt an RP prime fail | Nonconforming Material/Component | Bayer Healthcare sent an "Urgent Medical Device Field Corrective Action" letter dated November 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please immediately take the following steps.Please review your current inventory. Do not use any Source Administration Sets (SAS) with the impacted batch numbers that are included in the index below.If you have unused affected product, immediately quarantine the product and please call Bayer Customer Service at 1-800-633-7231 to receive a Returned Goods Authorization (RGA) number and return the product immediately to Bayer accordingly.Please complete the response form and fax it to Bayer at 1-412-406-0942 or e-mail to randiproductrecalls@bayer.com, whether or not you haveimpacted product.To provide you with new product, please submit a new product order.Please note that due to limited new inventory, there may be a delay in fulfillment.New batches of SAS are currently being shipped and lead times will be dependent on available supply until inventory returns to normal levels. We appreciate your cooperation and sincerely regret any inconvenience. We are committed to providing effective products and service to support your patient care. If you have questions, please contact our customer support team at 1-800-633-7231. |  | 198598 | 198,598 | Nationwide Distribution to the states of : AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MI, MN, MO, NC, NE, NH, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT and WI. | Open, Classified |  |  | 198598
Z-0639-2016 | 72540 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Mercury Medical Neo-Tee T-Piece Resuscitator and T-Piece Circuit. |  |  | October 22, 2015 | 2015 | Mercury Enterprises, Inc. dba Mercury Medical | Incorrect use of T-Piece resuscitators will result in complete occlusion of the patient circuit and consequent failure to ventilate the patient. | Device Design | Mercury Medical sent notification letters dated October 22, 2015, to all affected Domestic Consignees by courier and for the foreign consignees the notifications were sent via e-mail.The letter identified the product, the problem, and the action to be taken by the consignee. Customer Immediate Actions - Immediately notify all personnel subject to using these devices to remove protective cap before connecting a mask. Further this filed safety advisory notice should be read in conjuction with the manufacturer supplied instructions for use.In addition, if you further distributed this product, please identify your customers and notify them at once of this product correction. Notificatin to your customers should include a copy of this letter. For questions contact Mercury Medical at 787-573-0088, 727-573-4966 fax or via e-mail at regulatoryaffairs@mercurymed.com |  | 753245 | 753,245 devices | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Bahrain, Canada, Chile, China, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Guam, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Netherlands, Norway, Poland, Portugal, Puerto Rico, Qatar, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, United Arab Emirate, and United Kingdom.. | Open, Classified |  |  | 753245
Z-0633-2016 | 72808 | No | JJJ | Beta or gamma counter for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Perkin Elmer WIZARD2 2-detector, 550 samples Product Code: 2470-0020 |  |  | December 08, 2015 | 2015 | Perkinelmer | The content of the second #023 barcode ID label in the Barcode ID Label Binder is incorrect. The second barcode ID label #023 symbol incorrectly identifies as #024 when scanned. If measurement protocols have been programmed for both ID #023 and #024, the protocol associated with barcode ID #024 is executed. If the error is undetected, the Gamma Counter may produce erroneous results. | Nonconforming Material/Component | Perkin Elmer sent an Urgent Medical Device Correction letter dated December 7, 2015 to affected customers. The letter identified the affected products, problems and actions to be taken. Customers were instructed to destroy all affected products in their possession. Customers were asked to complete theenclosed Correction Response Form and return it by scan and e-mail to QMresponse.SG@perkinelmer.com . Corrected barcode labels ID #023 will be provided upon request and return of the completed Correction Response Form to PerkinElmer. For further information contact your local PerkinElmer representatives.. |  | 60 | 60 units | Worldwide Distribution - US Nationwide and the countries of ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, GREECE, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, MYANMAR, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, and THAILAND. | Terminated | August 04, 2017 | 605 | 430
Z-0629-2016 | 72880 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Reprocessed Stryker Color Cuff and Color Cuff II Adult Tourniquet Cuffs |  |  | December 17, 2015 | 2015 | Stryker Sustainability Solutions | Reprocessed Stryker Pressure Tourniquet Cuffs were mislabeled, including some labeled as single port when they contained dual port, and vice versa. | Under Investigation by firm | The firm sent out customer notification letters on 12/17/15. The firm requests that customers discontinue the use of the affected lots of the recall product, and other Pressure Tourniquet Cuffs reprocessed by Stryker Sustainability Solutions should be considered as alternatives. A Recall Effectiveness Check Form including the quantity of affected devices shipped to the customer is included with the notification.The firm states to complete the Recall Effectiveness Check Form and indicate if any devices from the affected lots remain in inventory. The letter states the form needs to be completed even if no product is found.If the form indicates that affected devices remain in inventory, a prepaid shippinglabel will be issued for the return of the product. Customers will receive a credit forall affected devices returned.Adverse reactions or quality problems experienced with the use of this productmay be reported to:- Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433X5555- http://www.stryker.com/productexperience/- The FDAs MedWatch Adverse Event Reporting program either online, by regular mail, or by fax. |  | 450 | 450 devices | Nationwide Distribution including the states of AK, NV, VT, FL, CA, CO, WI, IA, MN, OK, PA, MI, OR, KY, and MA. | Open, Classified |  |  | 450
Z-0649-2016 | 72925 | No | JIL | Urinary glucose (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | iChem 10 SG Urine Chemistry Strips, PN 800-7005. |  |  | December 08, 2014 | 2014 | Beckman Coulter Inc. | Iris International is recalling the iChem 10 SG Urine Chemistry Strip because of false positive leukocyte result for samples that are truly negative. | Nonconforming Material/Component | An "Important Product Notice," dated 12/8/14 was sent to customers to inform them of the recall by Iris International. The recall letter informs customers that The identified lots of iChem 10 SG Urine Chemistry Strips may have leukocyte pads that are more sensitive at the first reporting range of 1+, corresponding to 25 white blood cells (WBCs) /microliter. The letter provides the impact of the recall and the actions to taken by the customers. Customers are instructed to discard any remaining strips from the affected lots and use another lot number. Customers can either contact their local Customer Service Support or complete and return the Replacement Order form, if replacement product is needed. The enclosed response form should be completed and returned within 10 days. Customers with questions are instructed to refer to http://www.beckmancoulter.com; (800) 854-3633 in US and Canada. Customers outside the US and Canada are instructed to contact their local support representative. |  | 16206 | 16,206 vials total (7,827 vials in US) | Worldwide Distribution -- US, Brazil, Canada, Chile, Ecuador, France, Germany, Greece, Italy, Kuwait, Luxembourg, Malaysia, Philippines, Portugal, Puerto Rico, Singapore, Spain, Switzerland, Turkey, and United Arab Emirates. | Terminated | January 13, 2016 | 401 | 16206
Z-0645-2016 | 72698 | No | LYJ |  |  | Neurology | PMA | Cyberonics Vagus Nerve Stimulator VNS Therapy AspireSR Generator Model 106 |  |  | November 18, 2015 | 2015 | Cyberonics, Inc | Certain Model 106 Pulse Generators demonstrate delays in sensing during use of the 'Verify Heartbeat Detection' feature and exhibit the potential for decreased battery longevity. | Manufacturing material removal | Cyberonics sent out written recall notification beginning October 14, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. For further questions, please call 866-882-8804 |  | 29 | 29 units | US Distributions to the states of : IL, NC, KS, and TX., and Internationally to the country of: United Kingdom. | Open, Classified |  |  | 29
Z-0646-2016 | 72938 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Zimmer Periarticular Locking Plate System 3.5mm Locking Screw 58 mm Length; Zimmer Periarticular Locking Plate System 2.4mm Locking Screw 14 mm Length. |  |  | December 09, 2015 | 2015 | Zimmer Biomet, Inc. | Product on the package label did not match what was inside the package. The label was for part number 00-2359-058-35 (Zimmer Periarticular Locking Plate System 3.5mm Locking Screw 58mm Length) lot number 62918019 while the product inside the package was part number 00-2359-014-24 (Zimmer Periarticular Locking Plate System 2.4mm Locking Screw 14mm Length) lot number 62896038. | Labeling Change Control | On 12/9/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital risk managers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761. Hours of operation are Monday through Friday, 8 a.m. through 5 p.m. EST. |  | 400 | 400 | Worldwide Distribution-US including the states of CA, NY, PA, AL, FL, GA, ID, WI, MO, SC, NC, and ME, and the countries of Brazil, Dominican Republic, Taiwan, China, Uruguay, Australia, Japan, Malaysia, India, Belgium, France, Italy, Netherlands, Poland, United Arab Emirates, and Canada. | Open, Classified |  |  | 400
Z-0650-2016 | 72521 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Integral Centralizer Hip System, Lateralized Femoral Stem (Co-Cr-Mo), |  |  | November 17, 2015 | 2015 | Biomet, Inc. | PMMA is listed as a material on the label but the product does not contain PMMA. | Labeling Change Control | On 11/17/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 21 | 21 | Worldwide distribution. US nationwide including IA, UT, MN, PA, MO, OK, IL, OH, and WA; CANADA, and NETHERLANDS | Open, Classified |  |  | 21
Z-0642-2016 | 72597 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 4600 Chemistry System, Catalog Number 6802445, Unique Device Identifier Number 10758750012343, Software Version 3.2 and below; IVD. |  |  | November 03, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly may allow VITROS Systems to sample and process assays while the fluids and reagent temperatures are not in the required temperature range for optimal processing. When this intermittent issue occurs, the VITROS Systems continue to operate and process results without notification to the User, possibly producing erroneous results. | Software design | Ortho-Clinical Diagnostics sent an Urgent Product Correction Letters (Ref. CL2015-211, dated November 3, 2015) via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected devices (VITROS 3600 and 5600 Systems). The notification identified the affected products, problem and actions to be taken. Foreign affiliates were informed via e-mail on 11/03/2015 of the issue and instructed to notify their consignees of the issue and required actions.Ortho-Clinical Diagnostics sent an Urgent Product Correction Letters (Ref. CL2015-221, dated November 11, 2015) via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected devices (VITROS 5,1 FS and 4600 users). The notification identified the affected products, problem and actions to be taken. Foreign affiliates were informed via e-mail on 11/11/2015 of the issue and instructed to notify their consignees of the issue and required actions. |  | 427 | Domestic: 157, Foreign: 270 | Worldwide Distribution - US Nationwide and the countries of Puerto Rico, and foreign distribution to: Argentina, Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Israel, Italy, Japan, Kenya, Mexico, Panama, Philippines, Poland, Romania, Russia, Singapore, Spain, Taiwan, Turkey, and Venezuela. | Open, Classified |  |  | 4399
Z-0644-2016 | 72915 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Model NuCARTGN. The NuCART is a mobile display and image management platform used in Minimally Invasive Surgery (MIS) environments to facilitate the routing and ergonomic display of video signals from source devices to flat-panel monitors. It is tailored for use with an OEC C-arm and Endo-Cam in MIS surgery suites, Ambulatory Surgery Centers, Urology, GI and other suites. |  |  | December 04, 2015 | 2015 | CompView Medical, LLC | The attachment of the NuCART Spring Arm to the NuCART horizontal boom may come apart if the securing ring (snap ring) was not installed correctly causing the spring arm to release and detach. | Nonconforming Material/Component | The firm began inspecting devices between 12/4/15 and 12/16/15, and sent the Nucart Product Investigation letter to consignees between 12/11/15 and 12/15/15. Consignees with questions can call the firm at 503-601-5603. |  | 10 | 10 units | US nationwide distribution including AZ, CA, CO, GA, OH, and TN. | Terminated | March 31, 2016 | 118 | 10
Z-0647-2016 | 72804 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RT Elements are applications for radiation treatment planning for use in stereotactic, conformal, computer planned, Linac based radiation treatment of cranial, head and neck, and extracranial lesions. The simulated plan is intended for treatment evaluation for example in tumor board meetings or operating rooms |  |  | November 05, 2015 | 2015 | Brainlab AG | Large objects with fine resolution are potentially displayed cropped when imported into Brainlab Brain Metastases 1.0.0. | Software design | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 5, 2015, to all direct accounts informing the of the recall and actions to be taken.User Corrective Action:Users of Brain Metastases Version 1.0.0 or Adaptive Hybrid Surgery Analysis 1.0.0 shall adhere to the following:After import always carefully review and verify all objects for correctness and validity.For reviewing an object, select the image set which was used to outline or to modify the objects in the Data menu. Also compare the contours of an object in Brain Metastases or Adaptive Hybrid Surgery Analysis software to the contours in SmartBrush. Please note that the error will not be visible in SmartBrush. Always verify the correct shape of the objects in Brain Metastases or Adaptive Hybrid Surgery Analysis software. If possible do not use image data sets with large slice distance and/or with a low amount of slices. To retrospectively identify if cropped objects have been used for treatment planning, the object contours of a specific treatment plan can be reviewed in Brainlab Dose Review, where the error will also be visible.Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact affected customers tentatively starting February 2016 to schedule the update. The request that users advise the appropriate personnel working in their department of the content of the Notification letter. If the user facilities require further clarification, thay can contact their local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers)E-mail: support@brainlab.com (for US customers: us.support@brainlab.com)Fax: Brainlab AG: + 49 89 99 15 68 33 |  | 20 | 20 systems | Worldwide Distribution - US (nationwide) and Internationally to Australia, Belgium, Brazil, Canada, China, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Mexico, Portugal, South Korea, Spain, Switzerland, Turkey and the United Kingdom. | Open, Classified |  |  | 68
Z-0654-2016 | 72944 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Zimmer Universal Locking System 3.5mm Locking Reconstruction Plate |  |  | December 09, 2015 | 2015 | Zimmer Biomet, Inc. | The product on the package label did not match what was inside the package. The label was for part number 00-4936-011-13 lot number 62968956 while the product inside the package was part number 00-4936-008-07 lot number 62968974. | Labeling Change Control | On 12/9/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital risk managers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Specifically, customers were asked to review the notification and ensure all affected personnel are aware of the information; assist the Zimmer Biomet sales representative with the quarantine and removal of any of the identified affected product; and complete and return the Certificate of Acknowledgement form. If after reviewing this notification customers have further questions or concerns please call the customer call center at 1-877-946-2761. Hours of operation are Monday through Friday, 8 a.m. through 5 p.m. EST. |  | 52 | 52 | Worldwide Distribution -- USA, including the states of MN, NY, MS, AL, OH, KY, ME, and NM; and, the countries of Taiwan, Korea, Singapore, El Salvador, Uruguay, Czech Republic, France, Italy, Poland, United Kingdom, South Africa, and Slovakia. | Open, Classified |  |  | 52
Z-0655-2016 | 72897 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ergon 2 Articulating Equipment Boom Monitor Brackets |  |  | December 02, 2015 | 2015 | Skytron, Div. The KMW Group, Inc | The hardware that holds the VESA plate and the display interface / monitor to the bracket may fail causing the display interface to fall from the bracket. Over time and with use the threaded rod hardware material can weaken and detach from the bracket. A display interface falling from its mounting location may result in injury to the patient or provider. | Device Design | Skytron sent an Urgent Medical Device Recall letter dated December 02. 2015, to all affected customers via certified mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers with questions were instructed to contact at 616-656-2900 with any questions. For after-hours support, please call 1-800-SKYTRON. |  | 286 | 286 | Worldwide Distribution - US including WI, TX, FL, NC, MD, OH, AZ, CA, UT, MT, OK, HI, CO, WY, MO, NJ, IL, NY, MA, IA, DE, PA, MI, IN, and WV and Internationally to Canada, and Japan. | Open, Classified |  |  | 286
Z-0653-2016 | 72776 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 30 deg, 8mm Endoscope, IS4000: |  |  | November 23, 2015 | 2015 | Intuitive Surgical, Inc. | This correction to the da Vinci Xi I&A; User Manual communicates proper use of the da Vinci Xi 8 mm Endoscopes in combination with the da Vinci 8 mm and 12 mm cannulae, cannula reducers, and monopolar energy instruments. It informs customers of scenarios that may result in unintended patient harm due to capacitive coupling from monopolar energy sources. | Device Design | Intuitive Surgical sent an Urgent Medical Device Correction letter to affected customers on December 4, 2015,. Letters explained the issue and provided action to follow to prevent the issue. Customers were instructed to do the following:Distribute a copy of this letter to all da Vinci Xi users at your facility.3. Place a copy of this letter with your user manual.4. Complete the attached Acknowledgement Form and return it to Intuitive Surgicalas instructed.5. Retain a copy of this letter and the Acknowledgement Form for your files.If you need further information or support concerning this Medical Device Notification ,please contact your Clinical Sales Representative or contact Intuitive Surgical CustomerService at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (6 AM to 5 PM PST) ormail: customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8AM to 6 PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 003-5575-1362 (9 AM to 6 PM JST) |  | 2624 | 2624 total all models | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belgium, Cyprus, Czech Republic, Denmark, France, Germany, Hong Kong, India, Israel, Italy, Japan, Monaco, Netherlands, Norway, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, South Korea, Spain,Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela.. | Terminated | April 05, 2016 | 134 | 2624
Z-0659-2016 | 72896 | No | LYJ |  |  | Neurology | PMA | VNS Therapy AspireSR Generator Model 106. |  |  | December 18, 2015 | 2015 | Cyberonics, Inc | Recall being initiated in response to three reports of "Burst Watchdog Timeout" events occurring with the Model 106 AspireSR Generator, resulting in a device reset condition where stimulation output is disabled. | Device Design | The firm notified consignees of the issue via letter on 12/18/15. The letter identified the affected device, the issue involved, and actions to be taken. Physicians are to contact Clinical Technical Support at 866-882-8804 to report if a patient's generator has been disabled due to the issue identified. Users are to complete and return the effectiveness card as soon as possible. If further information is needed, customers can contact Clinical Technical Support at 866-882-8804 or via e-mail at cservices@livanova.com. |  | 4935 | 4,935 units | Worldwide Distribution -- United States, Austria, Belgium, Croatia, Czech Republic, Finland, France, Germany, Iceland, Italy, Netherlands, Norway, Poland, Slovakia, Spain, Sweden, Switzerland, United Kingdom, Cyprus, Israel, Kuwait, Lebanon, Oman, Qatar, Saudi Arabia, and United Arab Emirates. | Open, Classified |  |  | 4935
Z-0657-2016 | 72697 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | FMS neXtraÂż 3.5mm Aggressive Cutter Shaver Blades; Product Code: 283315 |  |  | November 20, 2015 | 2015 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Product is incorrectly labeled. The blade configuration in the affected lots does not match the description on the label. | Process control | On November 20, 2015, customers were notified of the recall via letter. They were asked to take the following actions: Customer immediate actions:1. Review inventory located within your facility and quarantine any products listed in this product recall notification. Please follow the instructions on the Business Reply Form to return any affected inventory located within your facility; 2. Review, complete, sign and return the attached business reply form to Elizabeth Messana (emessan1@its.jnj.com) in accordance with the directions on the business reply form within 5 business days of receipt of this notification; 3. Forward this notice to anyone in your facility that needs to be informed; 4. If any potentially affected product has been forwarded to another facility, contact Elizabeth Messana at 508-828-3150; 5. Keep a copy of this notice. |  | 45 | 45 units | Worldwide Distribution. US states of TX, OH, MI, and ME; and the countries of FR, DE, SK, IT, NL, CH, and CZ. | Terminated | April 12, 2017 | 509 | 210
Z-0630-2016 | 72860 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | arkray SPOTCHEM II Glucose Reagent test strips, Item No. 77101, MEDTEST DX Catalog # 77424. |  |  | December 18, 2015 | 2015 | Arkray Factory USA, Inc. | One lot each of SPOTCHEM II Basic PANEL-1 Reagent Test Strip and SPOTCHEM II Glucose Reagent Test Strip are being recalled due to a lowered reaction rate in the reagent pad for glucose. The affected lots will display a result of approximately 300 to 400 mg/dL for any glucose value that exceeds the upper limit of 416 mg/dL. | Component design/selection | The firm sent an Urgent Medical Device Recall letter to customers beginning 12/18/2015. The letter identified affected product, stated the reason for recall and asked for unused product to be returned. For instructions on returning unused product contact Lesty Jaranilla 1-877-538-8872 x3243. |  | 88 | 88 boxes (25 foil packaged test strips per box) | Distributed in the states of FL, IL, KY, MI, NC, NY, OH, and TN. | Terminated | January 04, 2017 | 383 | 99
Z-0661-2016 | 72953 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Fukuda Denshi patient monitor model DS-8100M and DS-8100N. |  |  | December 21, 2015 | 2015 | Fukuda Denshi USA, Inc. | A possible malfunction where a Short-term Battery Error message has appeared and has been traced back to a bad (damaged) Lithium-Ion capacitor (Short-term Backup battery). | Process control | Fukuda Denshi USA, Inc. of Redmond, WA plans to perform 100% field correction of units distributed in the US.The Recall/Field Correction Notice letter (FCN-A01-201512), dated December 21, 2015, and the Customer Confirmation Sheet will be sent to consignees using a clearly red-marked URGENT via Express Mail.The consignees are requested to fill out the Customer Confirmation Sheet and return via fax to 888-224-7090, email to DBlakely@fukuda.com, or send via postal service back to Fukuda Denshi USA, Inc. in a pre-stamped envelope.Consignee can call Doug Blakey directly at 210-240-4521 with any questions or concerns. |  | 95 | 95 units distributed in the US | Distributed in the states of CO, ID, KS, MI, MO, OR, PA, TX, UT, and WI. | Open, Classified |  |  | 95
Z-0598-2016 | 71924 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Brainlab Cranial Navigation System: An Image Guided Surgery System / Stereotactic. Radiology Departments. |  |  | April 22, 2013 | 2013 | Brainlab AG | Software Error: The effect of setup on overall navigation accuracy could potentially intensify small inaccuracies arising from individual steps of a complex navigation procedure that may cause an inaccurate display of instruments by the navigation system compared to the actual patient anatomy. | Software in the Use Environment | The firm, Brainlab AG, sent out a "Field Safety Notice / Product Notification" CAPA-20130417-000315 Brainlab Cranial Navigation System: Effect of setup on overall navigation accuracy for all versions of software on 4/22/2013 and an update on 5/29/2015.- From April 22, 2013 has been distributed via email to the consignees.- UPDATE from May 29, 2015 was distributed via email to the medical users.An active reply by the consignees will be followed, if necessary via further different communication media.FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter describes the product, problem and actions to be taken. The customers were instructed to adhere to the measures described in document Measures to Improve Cranial Navigation Accuracy in the Product Notification letter when using the Brainlab Cranial Navigation System.Brainlab is requesting that 1. Existing potentially affected customers receive this Product Notification Letter.2. These customers receive the attached supplement to the Instructions for Use regarding the Brainlab Cranial Navigation System in hardcopy version as an amendment to their current User Manuals. Tentative planned timeline for availability: June 2013Additionally to advise the appropriate personnel working in your department of the content of Product Notification letter.If the medical users have the need for further clarification, they are asked to contact their local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers) or byE-mail: support@brainlab.com. Fax Brainlab AG: + 49 89 99 15 68 33 Address: Brainlab AG (headquarters), Kapellenstrasse 12, 85622 Feldkirchen, Germany. |  | 1430 | 1,430 units | Worldwide Distribution: US (Nationwide) including Puerto Rico; and countries of: Algeria, Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Latvia, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Open, Classified |  |  | 1430
Z-0662-2016 | 72878 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Excella II Rod Cat. No.: SR 55100 5.5 x 100 mm LOT 1502 and Excella II Rod Cat. No.: SR 55110 5.5 x 110 mm LOT 1502 |  |  | December 07, 2015 | 2015 | Innovasis, Inc | Innovasis recalled two lots of Excella II titanium rods which were packaged using each other's labels. | Labeling mix-ups | Innovasis sent a Voluntary Product Recall letter, dated December 7, 2015 to customers regarding two lots of Excella II titanium rods which were packaged using each other's labels. Customers were asked to check their inventory and either remove the rod from the mislabeled package and request additional rods, or return the unopened packages immediately. The firm will replace all returned inventory upon receipt. |  | 1519 | 17 units Labeled SR55100 Lot 1502 with SR55110 rods in package | Distributed in US -- CA; KS; OH; &amp; TX. | Terminated | April 07, 2016 | 122 | 1519
Z-0660-2016 | 72569 | Yes | LWS |  |  | Cardiovascular | PMA | Optisure Dual Coil Defibrillation Leads, Model No. LDA220, LDA220Q, LDA230Q, LDP220Q. St. Jude Medical. Designed for long term attachment to an Implantable Cardioverter Defibrillator (ICD). |  |  | November 03, 2015 | 2015 | St Jude Medical Cardiac Rhythm Management Division | Leads may have been compromised during the manufacturing process, with cuts in the insulation layer of the SVC shock coil. | Process control | An Important Medical Device Advisory letter was sent to doctors on 11/3/15 to inform them of the recall by St. Jude Medical. The letter informs the doctors that St. Jude Medical identified that during the manufacturing process of a limited number of Optisure leads, a trim technique to remove excess medical adhesive around the SVC shock coil may have introduced damage to the leads insulation. The letter provides the doctors with the recommendations for patients with a potentially impacted Optisure lead connected to a device with Dynamic TX technology and for patients implanted with a potentially impacted Optisure lead connected to a device without DynamicTx technology. Doctors with questions about the advisory are instructed to contact their local Sales Representative or St. Jude Medical Technical Services at (800) 722-3774. |  | 447 | 447 units total (281 units in US) | Worldwide Distribution US (nationwide), Belgium, Switzerland, Germany, France, United Kingdom, India, Italy, Japan, Luxembourg, The Netherlands, Saudi Arabia, and Sweden. | Open, Classified |  |  | 447
Z-0686-2016 | 72836 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Wilson Penile Implantation System Retractor Frame, Retractor Stays &amp; Penile Elevation Strap, SINGLE USE ONLY Rx ONLY. |  |  | November 25, 2015 | 2015 | Abeon Medical Corporation | The firm was notified by some of their customers that some of the sealed pouches contain tears compromising the sterility of the product. | Packaging | On 12/17/2015 the firm sent Recall Notification Letters to their customers. The letter identified the product, problem, and actions to be taken by the customer. |  | 575 | 575 units | Worldwide Distribution-US (nationwide) including the states of AL, AR, AZ, CA, CT,FL, GA, IA, IN, KS, KY, MA, MD, MI, MN, NC, NH, NJ, OH, SC, TN, TX, UT, VA &amp; WI., and the countries of Austria, Germany, Italy, Japan, Spain, Sweden &amp; Switzerland. | Open, Classified |  |  | 4180
Z-0674-2016 | 72630 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Valeant Pharmaceuticals North America, Small Particle Size Aerosol Generator, SPAG-2 6000 series. Model No. SPAG-2. A nebulizer pneumatic flow system for the administration of Virazole (ribavirin) aerosol. |  |  | September 21, 2015 | 2015 | Valeant Pharmacueticals International | The Serial I.D. Label and some information in the Operation Manual for the SPAG 2 6000 series is incorrect or out of date, including references to accessories that are no longer available. | Finished device change control | Valeant Pharmaceuticals International, Inc. sent Field Correction Notifications /Recall Stock Response Form dated September 22, 2015 through 2nd day air service to their customers. The notification informed the customer about the field correction, how to identify affected product and what actions to take. TThe customers were requested to fill out and return the attached Recall Stock Response Form and return the form via email (bauschandlomb6443@stericycle.com or fax (844.782.5565). A customer service number was provided in case the customer had any related questions or concerns (1.800.321.4576)On October 7, 2015 a second field correction notification was sent to customers. The notification informed the customer about the field correction, how to identify affected product and what actions to take. A customer service number was provided in case the customer had any related questions or concerns (1.800.321.4576). The updated notification clarified that the notification applies to all SPAG-2 6000 model units currently in the market. |  | 1178 | 1178 | USA (nationwide) Distribution and Internationally to Canada and France. | Open, Classified |  |  | 1178
Z-0675-2016 | 72997 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | InVision-Plus and InVision-Plus Junior Needleless IV Connector, Disposable, Single Use |  |  | January 05, 2016 | 2016 | Rymed Technologies, LLC | Potential for the female luer component of the InVision-Plus IV connector to crack during use. | Process design | The firm Ry Med Technology LLC, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 4, 2016 and response form to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use and return the product to: RyMed Technologies LLC, 6000 W. William Cannon Drive, Building B, Suite 300 Austin, TX 78749; the user facilities to change the IV connectors in 24 hour when lipids are being administered; if further distributed, contact other accounts, advise of recall and return outstanding stock; complete and return the RECALL NOTIFICATION ASAP and no later than 10 business day via email to: amccutchen@rymedtech.com, fax to: 512-301-7338, Attention Recall Return Notification or mail to RyMed Technologies LLC, Attention Recall Return Notification, 6000 W. William Cannon, Building B, Suite 300 Austin, TX 78749. If you have any questions concerning the recall, contact Director of Quality Assurance/Regulatory Affairs by email to: amccutchen@rymedtech.com or call 512-301-7338x303. If you have questions about returning/replacement of product contact Logistics Specialist by email to: svara@rymedtech.com or call 512-301-7338x309. |  | 496400 | 496,400 units | US Nationwide Distribution. | Open, Classified |  |  | 543000
Z-0687-2016 | 73038 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The e.cam Dual Signature Gamma Camera system. Used to detect or image the distribution of radionuclides in the body or organ, using the following techniques: planar imaging, whole body imaging, tomographic imaging for isotopes with energies up to 588keV. |  |  | December 31, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | To provide operator instruction manuals for the e.cam Dual Signature; Gamma Camera that were not sent in its entirety when the e.cam systems were refurbished and shipped to the customer. | Packaging process control | An updated Customer Advisory Notice Letter will be provided to each impacted consignee. The receipt of the letter will be documented and additional attempts will be made until such receipt is acknowledged. The Customer Advisory Notice asks that the letter be placed with systems instructions for use along with all provided user documentation. It also advises the consignee to forward the notice to the new customer if the equipment was sold.Customers with questions regarding the Customer Advisory Notice can contact their local Service representative or use the following contact number 1-800-888-7436. |  | 10 | 10 | Distributed in the states of Florida, New York, Indiana, Maryland, South Dakota, Missouri, Tennessee, and Georgia, and the countries of Australia and India. | Open, Classified |  |  | 10
Z-0672-2016 | 72202 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Accusorb MRI |  |  | October 06, 2014 | 2014 | MWT Materials, Inc. | When not used correctly, Accusorb MRI products may overheat and lead to patient injury. | Device Design | MWT issued a letter dated 10/4/2014 via email to all distributors for further distribution to consignees. Tghe letter identified the affected product, problem and actions to be taken. Questions can be direct to info@MWT-Materials.com. |  | 1 | 1 units | US Nationwide Distribution in the states of NY and MA. | Open, Classified |  |  | 187
Z-0683-2016 | 73036 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | IBA Dosimetry GmbH Lift Table Type 4320, Catalog HA03-010#2 and HA03-000#2 |  |  | December 10, 2015 | 2015 | Iba Dosimetry Gmbh | When the electric lift table is connected to the mains power, in a very rare occasion, the table may start to move up in an uncontrolled manner. | Component design/selection | The field safety notice including a confirmation of receipt was sent to the affected hospitals; by email on 12/10/2015 then followed by 12/17/2015 sent by Fed Ex. The firm requested the mains power be disconnected after setting up the table and BP. Replacement part and instructions will follow. |  |  |  | Nationwide and International: Algeria, Angola, Arab Republic of Egypt, Austria, Belarus, Belgium, Bulgaria, Bosnia and Herzegovina, Croatia, Czech Republic, Denmark, Ethiopia, France, Germany, Ghana, Greece, Hungary, Iraq, Ireland, Islamic Republic Iran, Italy, Kazakhstan, Kenya, Kingdom of Saudi Arabia, Kuwait, Lebanon, Libya, Morocco | Open, Classified
Z-0680-2016 | 72943 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aisys, 1011-9000-000. |  |  | December 15, 2015 | 2015 | GE Medical Systems, LLC | The Drive Gas Check Valve of certain Anesthesia products could become stuck in a fixed open position, which could cause pressure to build in the mechanical ventilation cycle. If this issue is left unresolved, it could result in excessive or prolonged pressure in the patient breathing circuit during ventilation potentially resulting in barotrauma. | Nonconforming Material/Component | GE Healthcare sent an Urgent Medical Device Correction letter dated December 15, 2015, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to contact GE Healthcare Service at 1-800-437-1171 or their local Service Representative. |  | 13 | 13 | Worldwide Distribution - US (nationwide including PR except NH, VT, &amp; WY.Internationally to ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BOLIVIA, BOSNIA, BRAZIL, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, CHANA, GIBRALTAR, GREECE, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LITHUANIA,MALAYSIA, MALDIVES, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI lanka, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIETNAM. | Open, Classified |  |  | 4047
Z-0689-2016 | 72965 | Yes | NIO |  |  | Cardiovascular | PMA | ZilverÂż 518RX Vascular Stent with Rapid Exchange Delivery System |  |  | December 22, 2015 | 2015 | Cook Medical Incorporated | Manufacturing of the joint between the inner catheter and stylet of the delivery system does not consistently meet specifications, and the joint may separate during deployment. Failure of this joint during a procedure may result in a partial stent deployment. The clinical effect of a partial deployment may be a need for surgical intervention, or may result in vessel trauma or an embolic event | Device Design | Cook Medical sent an "URGENT: MEDICAL DEVICE RECALL" letter dated December 22, 2015, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at 1-800-457-4500 or 1-812-339-2235.On 2/5/2016 a second recall notification was sent to affected consignees, expanding the recall to All Lots of the affected product. |  | 753 | 753 | Nationwide Worldwide | Terminated | July 25, 2017 | 581 | 753
Z-0651-2016 | 70765 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Vantage Panoramic X-Ray System used to perform dental and orthodontic radiological imaging |  |  | September 04, 2014 | 2014 | Midmark Corp dba Progeny Inc | Failure of the primary collimator may result in unintentional movement of the device. The unintentional movement will not result in the generation of an error message. | Device Design | MidMark sent an Urgent Medical Device Field Correction letter dated September 4, 2014, to all affected customers. The letter included instructions for distributors to: 1) encourage end users to contact Midmark to arrange the field correction after the final solution becomes available; 2) if a partial image is reported to have occurred with a machine, the machine should be removed from use and the end user should initiate a service request per the normal procedure; 3) quarantine any affected machines in the distributor's inventory and contact Midmark Technical Support within 10 business days to provide the Serial numbers to on-site servicing of the affected machines can be scheduled. Distributors with questions about this market action can contact Midmark Technical Support by calling 888-294-3800 x 3 or by e-mail at imagingtechnsupport@midmark.com. |  | 832 | 832 systems | Worldwide Distribution - US including AL, AR, AZ, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI; and Internationally to Australia, Belgium, Canada, Czech Republic, France, Hungary, Lithuania, New Zealand, Poland, Russia, Saudi Arabia, South Africa, Switzerland, Turkey, United Arab Emirates, United Kingdom, and Vietnam. | Open, Classified |  |  | 832
Z-0690-2016 | 72904 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Lipogems SPA Aspiration Cannulae, intended for aspiration of adipose tissue. LGC 13G (13G x 185mm), Sterile. |  |  | December 15, 2015 | 2015 | Globalmed Logistix | The cannulae may bend during procedure. | Under Investigation by firm | Consignees were notified by telephone and letter via FedEx on/about 12/15/2015. |  | 118 | 118 units | CA, CO, CT, FL, GA, NJ and NY. | Terminated | March 09, 2016 | 85 | 118
Z-0694-2016 | 73067 | No | FDF | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fuse 1C Colonoscope |  |  | December 18, 2015 | 2015 | EndoChoice, Inc. | The bending section of the device may partially separate from the insertion tube. Potential for tissue trauma. | Process design | Consignees were notified on12/09/2015 that a sales representative would be contacting them to facilitate the exchange of the colonsocope and the defective device will be returned to EndoChoice. |  | 253 | 253 units | Worldwide Distribution. US nationwide, Colombia, Denmark, France, Germany, Indonesia, Israel, Italy, Japan, Malaysia, Republic of Korea, Spain and Switzerland. | Terminated | March 24, 2016 | 97 | 253
Z-0695-2016 | 72927 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10 Rack Loader; |  |  | December 15, 2015 | 2015 | Bio-Rad Laboratories, Inc. | On a rare occasion, there is a potential to assign a patient result to an incorrect sample ID when running in D10 Rack Loader configuration. | Software design | Urgent Medical Device Correction letters were sent on December 2015 to all affected consignees and subsidiaries. Each subsidiary will be given a Subsidiary Medical Device Correction Notice, Customer Medical Device Correction Notice, and Return Response Form and the number of customers distributed by e-mail. Each Subsidiary much contact customers and translate the Customer Medical Device Correction Notice and Return ResponseForm as needed. A revised letter to subsidiaries and customers was sent on January 20, 2016. |  | 365 | 165 in US, 200 Internationally | Worldwide Distribution-US (nationwide) and the countries of Canada, Australia, Belgium, Brazil, Brunei Darussalam, China, Czech Republic, France, Germany, Greece, India, Israel, Italy, Malaysia, Maldives, New Zealand, Norway, Austria, Poland, Portugal, Russia, Spain, Sweden, Switzerland, Thailand, and UK. | Terminated | May 02, 2016 | 139 | 365
Z-0693-2016 | 72926 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler-Peritoneal Dialysis Cycler |  |  | December 15, 2015 | 2015 | Fresenius Medical Care Renal Therapies Group, LLC | The door latch may not fully close and the door may open unexpectedly. The door may make physical contact with anything near the door of the machine, including the patient. May cause injury, delay in treatment, or Increased Intraperitoneal Volume (IIPV) (overfill). | Device Design | Fresenius Medical Care issued on December 15, 2015, a n Important Product Notification to all affected customers informing them of the problem and how to recognize the proper latch closing. Questions regarding the contents of the Customer Notification or questions regarding the issue, contact Technical Services at 1-800-227-2572 and reference the Field Action Number FA-2015-10-W.A reply form is included as part of the customer communication withinstructions to fill out and return via fax or scanned/emailed copy. |  | 14094 | 14,094 units | Nationwide Distribution | Open, Classified |  |  | 14094
Z-0697-2016 | 72969 | No | LKP | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | EliA SmDP Well, Article Number 14-5624-01 |  |  | December 28, 2015 | 2015 | Phadia Ab | All equivocal and positive results (greater or equal 7 EliA U/ml) on EliA SmDP well lot 0018 may be incorrect and must be considered invalid. Patient samples in complaint investigations caused unspecific signals up to 22 U/ml on EliA SmDP well lot 0018. The nonspecific signals are not caused by anti-Sm antibodies, nor by streptavidin antibodies. | Nonconforming Material/Component | On 12/28/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions call 1-800-3464364, option #2. |  | 76 | 76 | US Nationwide in the states of NC, GA, NJ, NY, PA, TX, IL, MI, MT, NM, and OR. | Open, Classified |  |  | 76
Z-0701-2016 | 72864 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network. |  |  | December 15, 2015 | 2015 | Mindray DS USA, Inc. dba Mindray North America | An issue with the Panorama Central Station may cause the system to spontaneously restart. This may occur after about 49 days of continuous operation. On restart, monitoring will be lost for about 2.5 minutes. Settings and configurations are maintained. | Software design | The firm, Mindray DS USA, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated 12/15/2015 to its customers via certified mail. The letter described the product, problem, and actions to be taken. The customers were instructed to route the recall letter to all potential users of your facility. Mindray will resolve the issue by upgrading the system software on all affected Panorama systems in your facility. Please contact your Mindray Service Representative at 1-800-288-2121 (Monday-Friday 8:30am-5:30pm) to arrange upgrade.If you have any questions, please contact Director, Quality Operations and Regulatory Affair at 201-995-8407 or email to: d.arpino@mindray.com. |  | 63 | 63 products | Nationwide Distribution. | Open, Classified |  |  | 63
Z-0700-2016 | 72908 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Biosure Ratchet Driver, Product No: 72201888 For delivery and placement of orthopedic screws. Arthroscopic surgical instrument |  |  | April 04, 2013 | 2013 | Smith &amp; Nephew, Inc., Endoscopy Div. | Driver does not meet specification, oversized. Driver may not be able to be fully inserted into the screw, or screw may become stuck on the driver. | Employee error | Smith &amp; Nephew sent an "Urgent-Product Recall 1st. Notification" dated April 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Please inspect your inventory and locate all devices from the above listed product and lot numbers and quarantine them immediately. For further questions, please call (978) 749-1000. |  | 9 | 9 | Worldwide Distribution - US Distribution to the states of : AK, AL, CA, CO, FL, IL, IN, MI, NE, NJ, NY, OH, OK, PA, SC, TN and TX., and to the countries of : Austria, Australia, Canada, Dubai, Great Britain, India, Italy, Korea, Malaysia, Netherlands, Portugal, Sweden, Switzerland, Singapore and South Africa. | Open, Classified |  |  | 166
Z-0735-2016 | 72910 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | MediChoice Cohesive Bandages, One (1) roll of cohesive bandage per bag, thirty-six (36) bags per case [all model numbers have equivalent packaging] |  |  | November 17, 2015 | 2015 | Owens &amp; Minor Distribution, Inc. | MediChoice sterile cohesive bandages imported and sold under a private brand by Owens &amp; Minor were tested and found to have microbial contamination. | Under Investigation by firm | The firms recall strategy is as follows : recall was initiated by the Owens &amp; Minor Home Office Recall Officer. The Home Office Recall Officer e-mailed the following documents to the Distribution Center (DC) Recall Officer at each affected DC:1. DC Recall Letter (specific instructions for the DCs),2. Recall Inventory Report form (for the DCs to record the amount of affected product quarantined),3. Customer Letter (specific instructions for the customers),4. Recall Return Response Form (for the customers to report back the details of the quantity of affected product in their inventory),5. Attachment 1 (a document that provides examples of the product labeling),6. Recall Checklist (a checklist for the DCs listing the actions that need to be completed), and7. Recall Combine Queries spreadsheet (listing customers and current on-hand inventory at the DCs).The DC Recall Officer at each DC forwarded the information to the DC Customer Service team. The DC Customer Service team notified each impacted customer via e-mail or letter based upon the customers recorded preference.Level to which the recall was initiated: The recall was initiated to the level of the customer (i.e., hospitals, clinics, surgery centers).Follow-up plans: A minimum of three (3) attempts are made to obtain an acknowledgment from the customer that they received the recall notification. An attempt to determine the status of non-responsive businesses will be made. If the non-responsive business is still viable, an attempt will be made to determine the employee responsible for management of recalls and the recall notice will be sent to that employee. If the business is no longer viable, the steps described below for out-of-business accounts will be followed.Sub-recall: No sub-recall is required. These products are supplied to User Facilities and not distributed to commercial sub-accounts.Plans for effective checks:1. Check records to ensure that the initial recall not |  | 64011 | 64011 | Nationwide Distribution. | Open, Classified |  |  | 202327
Z-0712-2016 | 72599 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD20 Biplane OR Table; Model Numbers: 722025 722020 |  |  | November 03, 2015 | 2015 | Philips Electronics North America Corporation | The use of Allura Xper, Allura Clarity, Allura CV, Allura Centron and UNIQ systems regarding live images and still images can lead to still images being interpreted as live Images. | Device Design | The firm, Philips, issued "URGENT-Field Safety Notices" dated 10/26/2015 and 11/19/2015 to its customers. The notices described the product, problem and actions to be taken. The customers were instructed to ensure the audible signal (X-ray buzzer) is always enabled, to activate only ONE fluoroscopy pedal at time, and to configure the footswitch to the default Philips configuration. Customers were also instructed to review the information with all members of your staff who need to be aware; retain a copy with the equipment IFU and complete and return the Confirmation Respond form to Mail Stop 4302, 3000 Minuteman Road, Andover, MA 01810 or email: Debra.Gibbs@Philips.com. Philips is tracking the delivery of these notices via certified return receipt.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722-9377. |  | 13297 | 13297 in total | Worldwide Distribution: US (Nationwide) including Puerto Rico and countries of: Albania, Algeria, US Virgin Islands, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Brunei Darussalam, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Democratic Congo, Costa Rica, Croatia, Cuba, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Fiji, Finland, France, French Polynesia, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mauritius, Mexico, Moldova, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, RÂżunion, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syria, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 13297
Z-0731-2016 | 72946 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Baxter, RAPIDFILL Syringe Strip, 10 ml. |  |  | December 21, 2015 | 2015 | Baxter Corporation Englewood | Baxter Healthcare has issued an Urgent Product Recall communication for all lots of RAPIDFILL Syringe Strips due to lack of evidence supporting the packaging claim for 12-month shelf life on released product. | Under Investigation by firm | An Urgent Product Recall notification was sent to consignees via USPS, first class mail on 12/21/15. Baxter Healthcare has asked customers to do the following:l. Locate and remove all affected product from your facility. The product code and lot number can be foundon the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return, credit and alternative products.3. Complete the attached customer reply form and return it to Baxter by either fax or scanned e-mail.4. Communicate the contents of the letter to your staff if affected products are distributed to other facilitiesor departments.5. Conduct a consumer-level recall for dealers, wholesalers, distributors, resellers, or Original EquipmentManufacturers that distributed affected product to other facilities. |  | 4695 | 4,695 | Worldwide Distribution-US (nationwide) and the countries of Canada, Sweden, Germany, and the Netherlands. | Open, Classified |  |  | 4695
Z-0732-2016 | 73007 | No | NWV | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | PS-500 Scope Warmer; 10 units per box. |  |  | December 01, 2015 | 2015 | JosNoe Medical, Inc. | Damage to package seals potentially compromising sterility. | Packaging process control | On 12/1/2015, notifications were sent to the affected customers via e-mail. The recall notification included a description of the reason for the recall, the affected product, and instructions for identifying and returning the affected product. |  | 1008 | 1008 boxes | Distributed US (nationwide) and in Canada. | Terminated | March 22, 2016 | 112 | 1008
Z-0737-2016 | 72945 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Denali Filter-Jugular/Subclavian DL950J |  |  | March 13, 2015 | 2015 | Bard Peripheral Vascular Inc | IFU missing contraindications: Patients with uncontrolled sepsis and Patients with known hypersensitivity to nickel-titanium alloys. | Labeling Change Control | The firm sent out Medical Device Recall customer notification letters on 03/13/15.The notification letter stated that there is no required action on the customer's part. Customers were told that if the device did not meet their needs then the firm will provide a replacement unit. Effective immediately the firm is implementing a change in the ordering code for the device to accurately reflect the approved information with the IFU and corresponding retrieval kit.Any questions the firm states to call Bard Canada:1-800-268-2862 Western Canada1-800-387-9473 Ontario1-800-387-7851 Eastern Canada |  | 1183 | 1,183 units | Distributed to CA only. | Terminated | February 01, 2016 | 325 | 1183
Z-0608-2016 | 72960 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ottobock Modular Polycentric EBS Knee Joint 3R60, 3R60=ST, 3R60=KD, 3R60=HD. Used for the prosthetic fitting of amputations of the lower. limb. |  |  | January 04, 2016 | 2016 | Otto Bock Healthcare GmbH | Otto Bock HealthCare GmbH is recalling Modular Polycentric EBS Knee Joint that were serviced between 27 April 2015 and 17 November 2015 due to a service error where the wrong Loctite wasused during repair. | Process control | Starting Jan 4 2016, In coordination with the Manufacturer, the US Agent will notify customers via telephone call of the correction on behalf of the Manufacturer. The customers will be requested to return the affected devices back to the service center for correction. The devices will be returned to the customer following correction. |  | 40 | 40 units | Nationwide Distribution. | Open, Classified |  |  | 40
Z-0736-2016 | 72972 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | HemosIL Plasminogen, Part No. 0020009000, IVD. Instrumentation Laboratory Company. Kit contents: 2 x 2.5 mL Streptokinase reagent and 2 x 2 mL Chromogenic substrate |  |  | December 18, 2015 | 2015 | Instrumentation Laboratory Co. | This Lot is not meeting labeled on-board instrument stability claims of 5 days. | Device Design | Instrumentation Laboratory Co. sent an " Urgent Medical Device Removal" letter dated December 2015, was sent on December 18, 2015. via USPS to US and Canadian customers. The notification requires US and Canadian customers to respond using a Mandatory Response Tracking Form. A Field Safety Notice was e-mailed on December 18, 2015, to the affected IL affiliates, notifying them of the product lot removal due to a stability issue and requesting them to translate the Field Safety Notice, to notify their Competent Authorities and to provide the notification to their affected customer base and secondary distributors. For further questions, please call 1(800) 955-9525. |  | 305 | Domestic: 108 kits, Foreign: 197 kits | Worldwide Distribution - US Distribution and to the countries of : Canada, Argentina, Austria, Belgium, Brazil, China, Czech Republic, France, Germany, Hong Kong, Italy, Lithuania, Panama, Poland, Russia, Saudi Arabia, United Kingdom, and Venezuela. | Open, Classified |  |  | 305
Z-0811-2016 | 71008 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope Video Laryngoscope Reusable Blades (GVL 3, 4, 5 and AVL 2, 3, 4, 5). |  |  | April 13, 2015 | 2015 | Verathon, Inc. | Verathon has revised the Operations and Maintenance Manual (OMM) to specifically state compatibility limitations of the reusable video laryngoscope Blade validated through simulateduse cycles, such that exceeding this number of cycles in the specific chemical agent may affect the potential life of the product. | Nonconforming Material/Component | Verathon sent an Urgent Medical Device Safety Alert letter dated April 13, 2015 and Safety Alert Reply Form to their consignees on April 17, 2015Users are informed that Verathon implemented a Safety Alert regarding updates to the following Operations and Maintenance Manuals (OMM):" GVL Operations and Maintenance Manual: 0900-1204-08-60" AVL Operations and Maintenance Manual: 0900-4200-02-60For a copy of the most recent GVL or AVL Operations and Maintenance Manual, please visit http://verathon.com/support.If you have any questions regarding the Operations and Maintenance Manuals revisions noted above, please contact your Verathon representative or Verathon Customer Care at 800.331.2313 or 425.867.1348. You may also email customerservice@verathon.com.PLEASE NOTE: NO ADDITIONAL ACTIONS NEED TO BE TAKEN. |  | 9088 | total 9088 units ( 5786 units in the US and 3302 units International) | Worldwide Distribution. | Completed |  |  | 9088
Z-0743-2016 | 73033 | No | KNK | Uric acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA Uric Acid CP ref. A11A01670 is an in vitro diagnostic assay for the quantitative determination of uric acid in human serum, plasma and urine based on the enzymatic determination of uric acid using a chromogenic system in the presence of peroxidase and uricase (Trinder method). |  |  | January 11, 2016 | 2016 | Horiba Instruments Inc | N-Acetylcysteine (NAC) present in the blood of patients treated for paracetamol overdose can interfere with the Trinder reaction that uses hydrogen peroxide catalytic on aminoantipyrine and phenol, and can produce falsely low results with the reagents using the Trinder reaction method. | Under Investigation by firm | Customer notification letters were sent on 01/11/16.Customers were provided the following instructions/actions:When performing tests that use the Tinder reaction, the blood draw should be done before treatment is given to the patient.Reagent inserts for the concerned HORIBA Medical Products have been updated and now include the following additional warning: Patients treated with N-Acetylcysteine (NAC) for Paracetamol overdose may generate a false low result. The letter request customers share the information with laboratory staff and retain the notification as part of their laboratory documentation. The letter states it is mandatory for customers to complete and return the enclosed response form within 10 days in order for the firm to verify the understanding of the notice.If any questions contact the local HORIBA Medical representative. |  | 216 | 216 | U.S. distribution to the following; CA, NY, AL, OK, MN, MS, MD, TX, KY, WV, MT, IL, FL, VA, WA, MI, KS, NE, PA, OH, ID, OR, AR, AZ, NM, NC, SC, GA, TNNo foreign distribution. | Open, Classified |  |  | 216
Z-0744-2016 | 72913 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST Medical Systems, BT2000 Automated Manifold Kit (SKU 014613) and CPT2000 (SKU 014623) Automated Manifold Kits, Sterile R, Rx Only, Made in Mexico. |  |  | December 21, 2015 | 2015 | Acist Medical Systems | Reports related to ingress of air into the manifold. | Nonconforming Material/Component | Acist sent an Urgent Medical Device Recall letter dated December 21, 2015, via e-mail to all affected consignees. The letter was addressed to Distribution Partner or Cardiac Cath Lab / Material Manager in the USA. In Europe, the letter was addresses to Purchasing Department. The letter described the problem and the product involved in the recall. Advised consignees to return affected units and to complete and return the Medical Device Recall Form. Distributors were further advised to send the recall letter to their customers. For questions or concerns, contact in US, ACIST Customer Service 1-888-670-7701 or your ACIST Sales Representative. Europe Direct contact ACIST Europe Customer Service +44 (0) 845 548 12 80 or contact your ACIST Sales Representative, Europe Indirect contact ACIST Europe Customer Service +31 45 750 70 25 or contact your Distributor. Distributors contact: ACIST Europe BV, Office: +31 (0)45 750 7015, Mobile: +31 (0)6 233 648 46. |  | 534754 | 48,614 cartons or 486,140 units | Worldwide Distribution - US including all states except DE, MT, RI, UT, &amp; VT and Internationally to Belgium, France, Germany, Iceland, Ireland, Luxembourg, Netherlands, Reunion, United Kingdom, Australia, Austria, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Estonia, Greece, Hong Kong, Hungary, Islamic Republic of Iran, Israel, Italy, Japan, Republic of Korea, Kuwait, Latvia, Lithuania, Mexico, New Caledonia, Norway, Panama, Poland, Portugal, Saudi Arabia, Serbia, Slovenia, Bosnia and Herzegovina, Slovakia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, and Venezuela. | Open, Classified |  |  | 534754
Z-0745-2016 | 73023 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | DS/US Proton Feature with Pinnacle3 Software Version 10.0.0 and 14.0.0 Radiation Therapy Planning System Model #459800200621 |  |  | January 08, 2016 | 2016 | Philips Medical Systems | In DS/US proton planning, there is a correction factor used internally to the dose engine that is calculating improperly in some cases. This can result in an incorrect dose calculation. | Device Design | Philips sent an "Urgent Medical Device Correction" letter dated 16-DEC-2015 to affected customers. The letter described the Affected Products, Problem Description, Hazard Involved, How to Identify Affected Products, Action to be Taken by Customer / User, Actions Planned by Philips and Further Information and Support (contact Philips Customer Care Solutions Center at 1-800-722-9377 ). |  | 3 | 3 | US: Nationwide Distribution in the states of FL, OH, and MO. | Open, Classified |  |  | 3
Z-0748-2016 | 72959 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AQURE System; Model Number: 933-599. The AQURE System manages blood gas and immunoassay analyzers. |  |  | December 29, 2015 | 2015 | Radiometer America Inc | The AQURE System has a design error regarding sample type in which sample type may be specified for some results and left blank for other results. The fact that the sample type field is left blank in some cases could lead to misinterpretation of results. | Device Design | Customers were sent a notification letter on 12/29/15.The notification states that the AQURE System can be used, but to eliminate the potential confusion, blank fields for sample type are to be interpreted as not specified. Customers are told that a Radiometer representative will be in contact to schedule a visit or a remote session. During the visit or remote session the representative will run a database script on the AQURE system which will serve for two purposes:1. Update existing patient results where the sample type field is blank to include Notspecified. 2. Install a so-called database-trigger, which will include the text Not specified in ablank sample type field for future patient results.An upgraded version of the AQURE software will include modified device driversto ensure that Not specified is inserted if the device transmits an empty fieldfor sample type. The new software version will be installed by local engineerwhen available. |  | 375 | 375 | Worldwide Distribution-US (nationwide) including the states of PA, GA, MI, OH, VA, FL, AL, TX, OR, OK, TN, NY, and MA, and the countries of Austria, Czech Republic, Denmark, Estonia, United Emirates, Kuwait, Qatar, Saudi Arabia, Oman, Jordan, Finland, France, Germany, Hungry, Ireland, United Kingdom, Italy, Netherlands, Norway, Poland, Spain, Sweden, Switzerland, Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Russia, Singapore, South Africa, Republic of Korea, and Malaysia. | Terminated | March 29, 2016 | 91 | 375
Z-0692-2016 | 72914 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Emergency Transport Ventilator |  |  | December 22, 2015 | 2015 | Draeger Medical, Inc. | Drager became aware of situations where the error message "Poti unplugged" was generated causing acoustic and visual alarms, the breathing system releases pressure and the ventilation function stops operating. | Device Design | The firm, Draeger Medical, Inc., sent an "Urgent Medical Device Recall" letter dated 12/21/2015 and a copy of the Instructions for Use (IFU) Supplement to consignees in a method that confirms delivery. The letter described the product, problem and actions to be taken. The consignees were instructed to include a copy of the supplement with their Oxylog IFU; inform all affected users in their facility; and contact Draeger Customer Support at 1-800-543-5047 (press 4 at the prompt) to obtain additional copies of the IFU Supplement Free of Charge.If there are any questions regarding this letter, please contact Michael Kelhart at 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your DrÂżger Oxylog ventilator please contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt).On 12/09/2016, Draeger mailed an Urgent Medical Device Recall notification letter to affected customers to inform them of a revised new software for the Oxlog 3000/3000plus (version 1.06 for Oxylog 3000 plus and version 1.23 for Oxylog 3000) that reduces the impact of the error condition. Customers will be contacted by a Drager Service Representative to schedule a time to perform the software update free of charge. The new software changes the device behavior in the event of the error (a controller failure). With the new software, in the event of a controller failure, ventilation will not stop. The device will continue to ventilate with the last setting. An alarm will annunciate and be displayed, alerting the user of the controller failure. |  | 402 | 402 units (314 - Oxylog 3000; 88 - Oxylog 3000 Plus) | Nationwide Distribution -- including states of AR, AZ, CA, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | April 05, 2017 | 470 | 402
Z-0747-2016 | 73000 | No | HTC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Small Electric Drive (SED). For use in general traumatology. |  |  | August 14, 2014 | 2014 | The Anspach Effort, Inc. | The affected Small Electric Drive (SED) may operate solely in reverse mode, not operate in reverse mode when intended, not function, or unintentionally start. | Nonconforming Material/Component | On 11/13/2014 DePuy Synthes send a letter to all their customers informing them of the recall of all the lots of the Small Electric Drive (SED) with instructions to returned and possible exchange. Contact is Complaint Handling Unit Manager at 561-494-3673 |  | 599 | 599 each | Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CO, FL, ID, IL, KS, LA, MI, NC, NY, TX, WA, VA and HI, and the countries of Germany, Hong Kong, Chile, Israel, Poland, Italy, Turkey, Italy, Singapore, Ireland, Korea, Spain, Norway, Denmark, Great Britain, Finland, Belgium, Netherlands, France, Japan, Sweden, Switzerland, Austria, China, India, Thailand, Russia, New Zealand, Zaire, Australia, Taiwan, and Malaysia. | Terminated | February 05, 2016 | 540 | 599
Z-0696-2016 | 73046 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ET-20 Anesthesia Extension Tubes, 2.5 ml approximately, Sterile and Disposable, Wall Medical Inc., Made in China |  |  | January 06, 2016 | 2016 | Global Healthcare Inc | Firm was notified of a potential product defect related to kinking (flattening) of anesthesia extension tubes which may represent a potential health hazard. | Nonconforming Material/Component | Global Healthcare sent an "Urgent: DEVICE RECALL" letter dated January 6, 2015 to its Customer by email. The letter described the product, problem and actions to be taken. The customer was instructed to immediately examine its inventory to determine if you have any of the Model ET-20 on hand; if so, immediately discontinue distributing product; and complete and return the Distributor Device Recall Response Form by email. Additionally, if you have further distributed any of the product, immediately contact your accounts; advise them of recall situation and have them return their outstanding recalled stocks to you. Any returned product will be destroyed by the distributor upon receipt of all affected product.If you require further assistance, please call 770-522-7520. |  | 2000 | 2,000 cases (100 individual units per case) | US Distribution to state of: GA only. | Terminated | March 25, 2016 | 79 | 2000
Z-0746-2016 | 73071 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Astral 100, 150 |  |  | January 13, 2016 | 2016 | Resmed Corporation | Resmed Corporation is recalling Astral 100, 150 External Battery because of a start-up issue. | Software design | ResMed sent an Urgent Safety Alert dated January 13, 2016, to all affected customers. The Alert identified the product, the problem, and the action to be taken by the customer. The notification letter provided the following instructions to distributorsField devices can be upgraded using ResMed Service Software or at an authorized ResMed Service Center. Contact Astral Support for further details.The SR2 software update will be available free of charge and should be implemented for all devices within the two-year service cycle or before at the customers preference.To acknowledge receipt of this Astral Safety Alert, submit this form.For any questions call Astral Support:855-245-4640 |  | 14736 | 14,736 | Worldwide Distribution - US including VA, FL, MO, CA, LA, TX, NJ, NH, NC, WV, MI, ID, OK, MI, MN, UT, CT, IL, AZ, UT, KY, MT, PA, RI, TN, DE, NE, LA, NY, SC, WY, IN, VA, GA, CO and Internationally to CA, CO, and AR. | Terminated | April 22, 2016 | 100 | 14736
Z-0755-2016 | 72391 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry XPT Systems, Software Version 1.0.3, SMN 11127538 |  |  | October 07, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Multiple Software issues. Auto Start-UP Fail, Calibration Interval Resets when a Reagent Blank is run, Control Definition screen assumes range defined is 2 SD, Laboratory Information System (LIS) communication/Laboratory Automation (LAS) issue, Printer Driver Resets, ISE Calibration Ranges are too conservative for Urine Sodium, Archiving and deletion may fail and Workstation services may restart. | Software design | On October 7, 2015 Siemens Healthcare Diagnostics distributed Urgent Field Safety Notices to their custoemrs via Fed Ex in the U.S. and via email to regional countries for implementation outside of the U.S. Customers were provided a work-around if they have experienced any of the following problems: 1) Intermittent auto-startup function fails, 2) calibration interval resets as though a full calibration was performed, 3) Users may be unaware that the QC definition screen with +/- suggests a 2 Standard Deviation range value, 4) a rescanned and re-queried sample may become stuck in a communication loop and all tests may read as "inprocess" if the the initial reaction process is incomplete, 5) The printer stops printing if multiple printouts are requested in quick succession, 6) the ISE calibration for Urine Sodium (Na) may fail due to the ranges being set conservatively during development, 7) The automatic or manually started, archiving and deletion maintenance activity may fail and 8) A workstation services restart screen may appear under multiple scenarios causing the interface to be temporarily inaccessible. If customers have experienced any of the previously mentioned workflow problems, please perform the following actions: 1. Auto Start Failing: Manually perform the Startup Wash, run controls and any calibrations that are due. 2. Calibration interval resets when an RBL is run: Calibration should be run according to the recommendations in the Instructions for Use (IFU) for each assay. Keep a record of required calibrations and ensure that all assays are calibrated before use. 3. Control definition screen assumes that the range is defined as 2 SD values: When creating QC definitions, enter the values associated with a +/- 2 Standard Deviation (SD) range in the +/- ranges field.4. Laboratory Information System (LIS) communication / Laboratory Automation (LAS) issue: Be aware of errors in the event log and take action as needed. Manual intervention and rerunning |  | 87 | 87 units | Worldwide Distribution: US (nationwide) in states of: AZ, CA, MA, NY, TX &amp; WA, and to countries of: Bahrain, Egypt, Slovakia, France, Spain, Italy, Norway, Singapore, New Zealand, Germany, Australia, Republic Korea, Netherlands, Denmark, Finland &amp; United Kingdom. | Open, Classified |  |  | 87
Z-0750-2016 | 73027 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERTÂż FA, bioMÂżrieux, Inc., |  |  | December 24, 2015 | 2015 | bioMerieux, Inc. | Bottles may have been exposed to non-normal shipping conditions which may lead to media degradation and impact growth performance. | Packaging process control | Biomerieux sent a Product Recall Notice letter to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the referenced part and lot number and discard remaining inventory according to their facility's procedure for culture media. Customers were also instructed to contact their local bioMÂżrieux representative for product replacement. |  | 2741 | 2741 units in total | Distributed in the countries of Angola, Austria, Bosnia, Bulgaria, Cameroon, Czech Republic, Democratic Republic of Congo, France, Germany, Greece, Hungary, Italy, Kazakhstan, Kingdom of Saudi Arabia, Kosovo, Lebanon, Netherlands, Niger, Pakistan, Poland, Republic of Belarus, Romania, Russia, Serbia, Sri Lanka, Spain, Sweden, Switzerland, United Arab Emirates, Turkey and Zimbabwe. | Open, Classified |  |  | 2741
Z-0756-2016 | 73060 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | API ZYM Bx2 (Ref 70493) |  |  | January 08, 2016 | 2016 | BioMerieux SA | Discrepant identification results related to API identification strips. These discrepant identification results are due to false negative results on the tests of the strips linked to the ZYM B reagent. | Other | BioMÂżrieux sent an Urgent Field Safety Notice letter DATED jANUARY 8, 2015 to affected customers. The letter identified the affected product, problem and action to be taken. dated . The letter indicated that they could continue to use product in inventory with mandatory requirements that before using a ZYM B ampule and each day of use of API strips that they perform a quality control with ATCC strains as described in each package insert of involved API strips. The letter also requested that they distribute this letter to all appropriate personnel in the laboratory, retain a copy for their files, and forward the information to all parties that may use the product. For additional assistance or have any questions, contact your local bioMÂżrieux Customer Service representative. |  | 13271 | 13,271 | Nationwide Distribution in the states of: AL, CA, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, ME, MD, MA, MI, MN, MS, MO, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TX, VT, VA, WA, WV, WI, WY | Terminated | July 14, 2017 | 553 | 13271
Z-0754-2016 | 72968 | No | MTF |  |  | Immunology | PMA | ADVIA Centaur PSA Assay; Catalog Number: 06574155/SMN 10310292 (100 test kit) and Catalog Number: 02676506/SMN 10310293 (500 test kit) |  |  | December 29, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | PSA assay is not meeting the High Dose Hook Effect expectation in the Instructions for Use (IFU). | Under Investigation by firm | An Urgent Medical Device Correction (UMDC) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on December 30, 2015 to be delivered to customers on December 31, 2015. An Urgent Field Corrective Action (UFCA) was sent out for distribution to all customers outside the United States on December 29, 2015. These notices identified the affected product, problem and actions to be taken. For questions customers were instructed to contact Siemens Customer Care Center or local Siemens technical support representative. |  | 55754 | 55,754 kits | Worldwide - US Nationwide including Puerto Rico, OUS: Albania, Algeria, Angola, Armenia, Austria, Bahrain, Belarus, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, Iceland, Iraq, Italy, Jordan, Kazakhstan, Latvia, Lebanon, Lithuania, Luxembourg, Macedonia, Malta, Morocco, Netherlands, Norway, Oman, Pakistan, Philippines, Poland, Portugal, Qatar, Romania, Russian Fed., Serbia, Slovakia, Slovenia, Spain, Sweden, Switzerland, Tunisia, Turkey, Turkmenistan, Unit.Arab Emir., Vatican, AR, AU, BD, BR, CA, CL, CN, CO, DO, EC, GT, HK, ID, IL, IN, JP, KR, LK, MM, MX, MY, NI, NP, NZ, PE, PH, PY, SA, SG, TH, TW, UY, VN and ZA. | Open, Classified |  |  | 55754
Z-0757-2016 | 73059 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus TJF-Q180V flexible gastrointestinal Duodenoscope |  |  | January 08, 2016 | 2016 | Olympus Corporation of the Americas | Olympus America Inc is conducting a voluntary removal/corrective action of all TJF-Q180V duodenoscopes in order to replace the forceps elevator mechanism Olympus is replacing the forceps elevator mechanism with a new forceps elevator design consistent with the design specification in the recently cleared TJF-Q180V 510k. | No Marketing Application | Olympus America Inc. sent an Urgent Notification letter dated January 15, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory for the affected product, Olympus will contact facilities to make arrangements for return of the TJF-Q180V duodenoscope(s) for the elevator mechanism replacement. For questions call (484) 896-5688. |  | 4436 | 4436 | US Nationwide Distribution except PR | Terminated | January 09, 2017 | 367 | 4436
Z-0763-2016 | 72966 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Mavig Monitor Suspension System, Model GD60, contained in the following Fluoroscopic imaging systems: Advantx LCV+, Advantx LCN+, Advantx LCLP+, Advantx LCA, INNOVA 2000, INNOVA 4100, INNOVA 3100, INNOVA 3131IQ, INNOVA 2121IQ, INNOVA 2100IQ, INNOVA 3100IQ, and INNOVA 4100IQ |  |  | December 18, 2015 | 2015 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a reported incident in which a Mavig suspension arm with monitor assembly fell to the floor. | Nonconforming Material/Component | Consignees were sent on 12/18/2015 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref#12239 dated December 18, 2015. The letter was address to Director of Biomedical Engineering, Director of Radiologyand Chief of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 2111 | 2,111 (1078 US; 1033 OUS) | Worldwide Distribution-US (nationwide) including Guam and Puerto Rico and the countries of Algeria, Angola, Argentina, Australia, Bahrain, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macedonia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sudan, Sweden, Switzerland, Syria, Taiwan, Thailand, Tunisia, Turkey, United Kingdom, Uruguay, Venezuela, and Yemen. | Open, Classified |  |  | 2111
Z-0766-2016 | 71764 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare TraumaDiagnost; Codes: 72001, 720011, 720013, 720020; A radiological examination unit for emergency work and routine skeletal diagnosis. |  |  | July 22, 2015 | 2015 | Philips Electronics North America Corporation | If arm is not fastened securely it could fall down. | Nonconforming Material/Component | Consignees were notified of the recall via letter sent certified return receipt on 7/22/2015. The letter informs customers that the Instruction for Use (IFU) requires the operator not to use the system and call for service, if a mechanical defect or malfunction is suspected. In this case, it is recommended not to position the Ceiling suspension with the Trauma arm directly above a patient as long as this Field Change Order is not yet installed. Philips plans to repair the affected systems. A Philips Service Engineer will contact customers as soon as the Field Action Kit is ready to be implemented. For questions, contact your local representative or use the Technical Support Line at 1-800-722-9377. When communicating with Philips regarding this issue, please reference Field Change Order 72000014. |  | 661 | 399 distributed, 262 are subject to this event | Distributed in the states of NY, WI, MA, AL, FL, GA, RI, CA, KS, and TX, and the countries of Canada, Australia, Austria, Belgium, France, Germany, Ireland, Italy, Netherlands, New Zealand, Portugal, South Africa, Spain, and Switzerland. | Open, Classified |  |  | 661
Z-0759-2016 | 73103 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Patient transport Shuttle with Trendelenburg quick adjustment function (Jupiter, Mars, Titan, TruSystem and Saturn operating table systems). |  |  | January 19, 2016 | 2016 | Trumpf Medical Systems, Inc. | Complaints were received of the table top on the transporter (shuttle) tilting down quickly in a maximum Trendelenburg position. | Component design/selection | TRUMPF Medical sent an Urgent Field Safety Information letter dated January 15, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Customers were instructed to take the following actions:Actions to be taken1. Remove the patient transport Shuttle from operation.2. Trumpf Medical, or a designated service representative, will schedule an appointment with thecustomer to perform service on the impacted device to prevent the issue from occurring.Passing along this information:Please make sure that, in your organization, all users of the patient transport Shuttle as well as any otherpersonnel who must be informed, have been made aware of this safety information.If you have provided the device to third parties, please forward them a copy of this letter, or inform thecontact person listed above.Please store this information at least until this measure has been completed.If you have any questions concerning this request or the procedure as outlined, please contact Trumpf Medical Systems, Inc. at 888-474-9359For questions regarding this recall call 843-534-0606. |  | 1284 | 1284 affected units | Worldwide Distribution - US including MA and TX; Foreign distribution to Austria, Azerbaijan, Belgium, British Virgin Islands, Czech Republic, Denmark, Finland, France, Germany, Greece, Italy, Japan, Luxemburg, Morocco, Netherlands, New Caledonia, Norway, Poland, Portugal, Qatar, Romania, San Marino, Spain, Sweden and Switzerland. | Open, Classified |  |  | 1284
Z-0764-2016 | 72571 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Allura Xper FD20 Ceiling version only; 722028; imaging applications. |  |  | November 03, 2015 | 2015 | Philips Electronics North America Corporation | Normally, movement of the C-arc is initially stopped by the motor (holding torque) then by the electromagnetic brake, leading to a full stop of the C-arc. Because of improper adjustment and tolerances, when the C-arc is initially stopped in an unbalanced position, the air gap of the brake can become too large, resulting in continued C-arc movement (C-arc is looking for balanced position). | Device Design | The letters for the customers in the United States were sent on 11/3/2015, via certified return receipt and were addressed to the Risk Manager.Philips will perform a test on the 15 affected Allura Xper R8.2 FD 20C systems to see if the motion controller amplifier of the brake motor works correctly. If the test will not pass, the motor will be replaced. The Recall Package is sent to all Key Markets worldwide. The Key Market of an affected country will contact the facility where these systems were installed. |  | 15 | 15 subject to correction and removal | Worldwide Distribution. US states of TN and IN and the countries of Australia, Belgium, Egypt, France, Germany, Japan, Republic of Korea, Netherlands, and Saudi Arabia. | Open, Classified |  |  | 17
Z-0762-2016 | 72798 | No | IXY | Wall-mounted radiographic cassette holder. | Radiology | Radiology | 510(K) Exempt | VS300 radiographic Cassette Holder or Wall Stand. For use in radiology department. |  |  | December 03, 2015 | 2015 | Del Medical, Inc. | The wall stand cassette holder can suddenly fall to the bottom of its travel, near the ground and contact patient. | Employee error | The firm sent an email followed by a certified undated letter entitled Safety Related Notification- Field Corrective Action VS-200 and VS-300 Wallstands notifying customers of the recall and actions to be taken.To prevent the possible interruption of service, Del Medical has prepared a service kit which must be installed at the next available opportunity.Del Medical is sending service kits to affected customers. Included within the package will be needed components and, a fixture which will facilitate a one person job. The firm expects the procedure to occupy one hour or, less. Also included with each kit will be an acknowledgement form which is intended to document the fix. The form, along with the original parts are to be returned to Del Medical upon completion of the activity to the following address:Del Medical 241 Covington Drive Bloomingdale, Illinois 60108 Attention: Bill Kostecki 1-847-288-7000 |  | 4000 | 4000 | Worldwide Distribution-US (nationwide) and the countries of Canada, Costa Rica, Libya, Mexico, and Philippines. | Open, Classified |  |  | 4000
Z-0758-2016 | 73098 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | 2.3mm Turbo-Power OTW, Manual, Model No 423-050 |  |  | January 15, 2016 | 2016 | Spectranetics Corp. | The product is labeled with an expiration date that is past its shelf life. | Error in labeling | Spectranectics sent a Medical Device Field Action letter dated January 15, 2016 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and set aside the affected product and contact Customer Service at 1-800-231-0978 to arrange a product exchange. Customers were asked to document the return on the enclosed Field Action Acknowledgement and Receipt Form. Customers with additional questions were instructed to contact their local Spectranetics sales representative.For questions regarding this recall call 719-447-2318. |  | 8 | 8 units | Nationwide distribution to Florida, North Carolina and Pennsylvania. | Open, Classified |  |  | 8
Z-0760-2016 | 72756 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Arctic Sun 5000 and Arctic Sun 2000; Intended for monitoring and controlling patient temperatures. |  |  | November 10, 2015 | 2015 | C.R. Bard, Inc. | Improper targeted temperature therapy was delivered to patients because the Temperature Simulator Key was connected to the patient temperature cable instead of the patient temperature probe. | Use error | Bard Medical Division sent a Safety Alert to each of the customers via FedEx with proof of delivery notice. All consignees have been instructed to carry the notification out to the user-level. |  | 1539 | 1,539 | Nationwide Distribution-DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI. IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | April 21, 2017 | 528 | 1539
Z-0768-2016 | 72974 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Optima XR220amx, Mobile Digital Ready Radiographic System |  |  | December 23, 2015 | 2015 | GE Medical Systems, LLC | Potential safety latch mechanism failure. The safety latch, within the column that supports the horizontal arm, is a backup in the unlikely event the primary safety cable breaks. A fall of the arm assembly could result in bodily harm to a person. | Device Design | Consignees were sent on 12/23/2015 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 10906 dated December 18, 2015. The letter was addressed to Director of Biomedical Engineering, Director of Radiology and Chief of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers may continue to use the system. The firm will correct the affected devices and no charge to the customer; and a GE Healthcare representative will contact customers to arrange for the correction. For questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 2389 | 2389 | Worldwide Distribution -- US, including Washington D.C. Puerto Rico, Virgin Islands, and Guam; and the countries of Afghanistan, Albania, Anguilla, Antigua and Barbuda, Armenia, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, Ethiopia, Finland, France, Germany, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Hungary, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea (Republic of), Kuwait, Lebanon, Malaysia, Martinique, Mexico, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Papua New Guinea, Paraguay, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Taiwan, Thailand, Trinidad , Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Open, Classified |  |  | 3410
Z-0767-2016 | 72210 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 Radiation therapy system |  |  | November 18, 2015 | 2015 | Mevion Medical Systems, Inc. | The treatment table positioning belt can slip gear positions, resulting in deviation of patient setup and treatment positions from the time of slippage until the discovery of slippage and the halting of treatment. | Labeling design | Customer notification letter was sent via certified mail on November 18, 2015. The notification explains the issue and actions to be taken by the customer. |  | 4 | 4 | US nationwide distribution including FL, NJ, OK, and MO. | Terminated | April 03, 2017 | 502 | 4
Z-0786-2016 | 73022 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Bard Distaflo Vascular Bypass Graft; Product Code: DF8006SC; 6 mm x 80 cm. |  |  | January 05, 2016 | 2016 | Bard Peripheral Vascular Inc | Size mislabeling; the graft is a 6 mm diameter graft when the actual packaged product is an 8 mm graft. | Error in labeling | Customer notification letters were sent on 01/05/2016. Customers were asked to not use the affected product, or further distribute it. Customers were asked to check all inventory locations for the affected product code, lot number combination listed. If customers further distributed any of the product they are requested to contact that location, advise them of the recall, and forward the listed instructions. The letter states to remove all identified product from the shelves. If no affected product is going to be returned the firm states to complete and return the attached form stating so. If customers have product it is requested that they also complete the attached form, and state the quantities and lot numbers in stock. The letter requests that customers call the firm's Recall Coordinator, Raye Seisinger, at 1-800-321-4254, option #2 ext. 2501 (M-F 6am to 3pm MST) or email raye.seisinger@crbard.com. Once all information has been verified the customer will be issued a Return Authorization or Consignment Recall Number to facilitate the expedient return of the product. The firm will provide replace product for returned product.Fax the attached form to the firm at 1-800-994-6772. If the customer cannot fax the firm they request that the Recall Coordinator be contact and provided the information verbally.The letter states a mailing label is enclosed to return the affected product. Customers are requested to amek the package as "recalled product" and include the XC or XH number. All products should be returned to the following shipping address:Bard Peripheral Vascular, Inc.1415 W. 3rd St.Tempe, AZ 85281 |  | 4 | 4 units | Distributed in the states of AL, TX, and CA. | Terminated | March 25, 2016 | 80 | 4
Z-0771-2016 | 73193 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic; Catalog/part number 032000 and 033000 Cosmetic cover package. |  |  | January 22, 2016 | 2016 | Accuray Incorporated | The gun box mounting bracket may fail to support the weight of the gun box when in vertical (inverted) position. If this failure occurs the gun box may become loose and could come into contact with a patient. | Device Design | Accuray sent an Urgent Safety Advisory Notification letters dated January 22, 2016 to all affected customers by trackable mail method. The letters identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Acknowledgement form. For questions contact Customer Support at 1-877-668-8667 or visit the Accuray Technical Solution Center at http://www.accuray.com/services-support |  | 226 | 226 | Worldwide Distribution. US Nationwide, Venezuela, United Kingdom, Ukraine, Turkey, Thailand, Taiwan, Switzerland, Spain, Saudi Arabia, Russia, Poland, Pakistan, Myanmar, Mexico, Korea, Japan, Italy, Ireland, India, Greece, Germany, France, Finland, Czech Republic, Colombia, China, Canada, and Belgium. | Terminated | October 20, 2016 | 272 | 226
Z-0792-2016 | 72937 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mizuho OSI Trios, Table Base; |  |  | December 14, 2015 | 2015 | Mizuho OSI | Table Bases may have been assembled incorrectly, missing a retaining ring from the head end assembly. Over time, the parts could go out of alignment, making clutch engagement or disengagement difficult. If the clutch cannot disengage, the user would be unable to rotate the patient from supine to prone. | Employee error | All consignees were notified by telephone and a recall letter was mailed to each facility as a follow up. Customers were asked to remove tables from operation until they were inspected by OSI service personnel. |  | 33 | 33 | US, South Africa, Japan and United Arab Emirates. | Terminated | February 10, 2016 | 58 | 33
Z-0788-2016 | 73141 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Endobon-Xenograft Granules, Zimmer Biomet, Synthetic Bone Grafting Material, used in dental and/or surgical procedures. |  |  | January 21, 2016 | 2016 | Biomet France SARL | Cytotoxicity tests were not passed at 36 months. Use of Endobon-Xenograft Granules past 18 months of its specified shelf life may result in patient illness/injury. | Under Investigation by firm | Zimmer BIOMET France SARL has initiated a voluntary recall via certified mail on January 2, /2016 for the EndobonÂż Xenograft Granules, based on limited (in vitro/bench) results which indicated cytotoxicity tests were not passed at 36 months (real-time) aging. Use of the EndobonÂż Xenograft Granules may result in patient may experiencing soft tissue irritation, up to and including infection. Customers area asked to do the following: 1. Please review this notice.2. Identify, remove from service and segregate all affected Endobon ROX* product in your inventory.3. Complete the attached Business Reply Form (Attachment 1) and:a. Fax it to 877-650-0368 or email it to Biomet5351@stericycle.com.4. If you have product to return, contact BIOMET 3i at the number on the Business Reply Form for an RMA number and return shipping label.a. Return a copy of the Business Reply Form along with any affected product to the address on the Business Reply Form.Replacement product will immediately ship based upon availability.For assistance or other questions that customers may have relative to this notice, please contact BIOMET 3i at 1-800-443-8166 or 1-561-776-6700. Team members are available to assist you 8:00am to 6:00pm (Eastern), Monday through Friday. |  | 145116 | 145,116 units | Worldwide Distribution - US including AL, AK, AR, AK, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MN, MI, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, WV and Internationally to Canada, Japan, Europe. | Terminated | May 31, 2017 | 496 | 145116
Z-0789-2016 | 73091 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Reuseable rotary blade/abrader, product number 899751305. |  |  | November 25, 2015 | 2015 | Richard Wolf Medical Instruments Corp. | Weak weld in the reusable rotary blade/abrader (burr) may result in inner part breaks at the distal or proximal end making the product dysfunctional. | Process control | The firm sent an URGENT Medical Device Recall letter dated 11/25/2015 to consignees informing them of the recall and the actions needed to be taken. These included to not use the affected burr and to destroy the device. The user is also instructed to fill out the Recall Action Response Form to acknowledge receipt of the notification. The form should be scanned to Dsantostefano@richardwolfusa.com or mailed to:Dawn Santostefano, Recall CoordinatorRichard Wolf Medical instruments353 Corporate Woods ParkwayVernon Hill, IL 60061 |  | 10 | Ten units | Distributed in Washington. | Terminated | August 21, 2017 | 635 | 16
Z-0787-2016 | 73076 | No | LRY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SIGMA HP MBT Non-Keel Punch Knee Instrument. |  |  | January 08, 2016 | 2016 | DePuy Orthopaedics, Inc. | Increased punch height, causing a 0.063 in (1.6 mm) gap between the trial and the instrument. This may cause the surgeon to select an insert that is too thin. Surgical delay may result if the punch dislodges from the inserter and needs to be removed from the joint space. Insert Dislocation or Spin Out, or Poor Joint Mechanics may result, if not noted during surgery. May require revision surgery. | Component design/selection | On 1/8/2016, URGENT INFORMATION  RECALL NOTICE for Specific Lots of SIGMAÂż High Performance (HP) MBT Non-Keel Punch Knee Instrument notifications were sent to the affected distributors and Medical Professionals with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For clinical questions from surgeons, please contact DePuy Orthopaedics, Inc.s Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m.  5 p.m. EDT). For questions about the device recall information provided, please contact Kim Earle, Senior Recall Coordinator, at 574-371-4917 (M-F; 8 a.m.  5 p.m. EDT). |  | 351 | 351 | Worldwide Distribution-US Nationwide, Armenia, Australia, Austria, Belgium, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hungary, India, Ireland, Israel, Italy, Japan, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, P R China, Portugal, Poland, Russia, Singapore, Slovenia, South Africa, Spain, Sweden, South Korea, Switzerland, Thailand, Turkey, UAE, UK, and Brazil. | Terminated | March 24, 2017 | 441 | 351
Z-0772-2016 | 73013 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | IMPLANT,TM, MTX |  |  | January 06, 2016 | 2016 | Zimmer Dental Inc | Information specific to the 3.7 mmD and 4.1 mmD Trabecular Metal Zimmer Dental Implants was omitted from the latest revision of the applicable IFU in the Indications, Contraindications, and Precautions sections. | Labeling Change Control | Zimmer sent an Urgent Device Correction letter dated January 6, 2016 to all affected customers.. The letter identified the affected product, problem and actions to be taken. Customers were asked to do the following; review the notice and check inventory for the affected units, ff they wish to keep the product familiarize themselves with the Technical Bulletin information, if they wish to return the affected product call Zimmer Dental Customer Service at 1-800-854-7019 to obtain a Return Authorization Number and to setup a FedEx pick-up to return the products, and complete Attachment 1, the Business Reply Form.It was requested that customers maintain a copy of the noticeand a signed copy of Attachment 1 for your records. For assistance or any other questions contact Zimmer Dental Customer Service at 1-800-854-7019. Customer Service is available 7:00am to 5:00pm (Pacific), Monday through Friday. |  | 1059 | 1059 | Worldwide Distribution. US Nationwide in the states of NY, CA, KS, NJ, CO, MI, MD, WA, CT, KY, OR, OH, IA, SC, AZ, AK, IN, TN, WI, FL, AL, UT, WY, MA, PA, MN, IL, MI, GA, ND, RI, including Puerto Rico; and the countries of Thailand, France, Germany, Colombia, Hungary, Netherlands, Saudi Arabia, Poland, Spain, and Italy | Open, Classified |  |  | 11631
Z-0752-2016 | 73078 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | INFX-8000V Bi-Plane X-Ray Interventional System |  |  | August 14, 2015 | 2015 | Toshiba American Medical Systems Inc | When a fontal DA (Digital Angiography) acquisition was done, scattered x-ray came into the dose meter on the lateral side. As a result, the dose meter sent a minus value to the software. Consequently, the software defined the data as an "abnormal value" and it stopped displaying dose data and the dose data was lost. The following message was displayed, "Dose meter abnormal, Dose info disabled". | Radiation Control for Health and Safety Act | Toshiba's planned action to bring into compliance:1.Affected customers will be notified using the Customer NotificationLetter you provided,2. revised system software will be installed on the systems to prevent this occurrence from happening,3. the corrections will be made free of charge, and4. the corrections will be completed by April 4, 2016.For further questions, please call 1- 800 521-1968. |  | 4 | 4 | US Nationwide Distribution to OH and NY | Open, Classified |  |  | 4
Z-0793-2016 | 73084 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith &amp; Nephew BEAVER Blade 4.0mm |  |  | January 11, 2016 | 2016 | Smith &amp; Nephew, Inc. | Blade exhibiting fracture during hip arthroscopy procedures | Device Design | Smith &amp; Nephew issued letters via Federal Express on January 11, 2016. requested to inspect inventory and locate any unused devices and quarantine them immediately.Complete Inventory Return Certification Form on the following page, indicating the quantities that need to be returned.Contact Smith &amp; Nephews Field Action Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to obtain a return authorization (RA) number. |  | 69043 | 69043 units | Worldwide distribution. US Nationwide and countries of AT, AU, BE, CA, CH, CL, CN, CZ, DB, DE, DK, EE, ES, FI, FR, GB, HK, IE, IL, IN, IT, JP, LT, LV, NL, NO, PL, PR, PT, RO, SE, SG, SI, TH, TR, and ZA. | Open, Classified |  |  | 69043
Z-0796-2016 | 73130 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | T-piece Nebulizer with in-line connectors, Part number 8915-7-50. Single patient use, non-sterile prescription device designed for use in a hospital or home care environment. |  |  | January 05, 2016 | 2016 | Salter Labs | Reports that the T-piece does not fit on the nebulizer. | Nonconforming Material/Component | A customer letter dated 1/13/16 was sent to all customers who purchased the T-piece Nebulizer with in-line connectors because of reports that the T-piece does not fit on the nebulizer. The letter informs the customers that there has been no reports of interruption or delay in therapy. Customers are instructed to contact Salter Labs Customer Service at 1-800-421-0024 to arrange for thereturn of the product and replacement inventory. Customers are instructed to acknowledge receipt of the communication by completing the attached reply form and faxing it to the number indicated on the attached form. Customers are instructed to contact Salter Labs Customer Support at 800-421-0024 or your local Salter Labs representative to document any issues that you have encountered or to request replacement inventory. |  | 70 | 70 units | Distributed in the states of TX, NY, PA, TN, SC, and OH, and in China. | Terminated | April 21, 2016 | 107 | 70
Z-0794-2016 | 72978 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite, a picture archive and communications system. Model Numbers Versions 5.30.0, 5.30.1, 5.30.2, 5.30.2 SB0344, 5.30.3, 5.30.4, 5.30.4 HF0418E, 5.30.5, 5.30.5 HF0404E, 5.30.5 HF0413E, 5.30.5 HF390E, 5.30.5 SB0368, 5.30.6, 5.30.6 HF0405E, 5.30.6 HF0406E, 5.30.7, 5.30.7 HF0429E, 5.30.7 HF0441E, 5.30.7 HF0446E, 5.30.7 HF0470E, 5.30.7 HF0473E, 5.30.8, 5.30.8 HF0477E, 5.30.8 HF0490E, 5.30.8 HF0495 SB0508E, 5.30.8 HF0495E, 5.30.8 HF0495E SB504E, 5.30.8 HF0550, 5.30.8 SB0485E, 5.30.8 SB0486E, 5.30.X SB0381E, 5.35.1, 5.35.1 HF0487E, |  |  | December 17, 2015 | 2015 | Merge Healthcare, Inc. | Potential incorrect Standardized Uptake Values (SUV) measurements in RadSuite. | Software design | Merge Healthcare sent an " Urgent - Medical Device Recall letter dated January 21, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer .If you may have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter. Merge Healthcare is committed to improve efficiencies and enhancing the quality of healthcare worldwide. If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 164 | 164 | Nationwide Distribution to the states of : MN, WI, IL, PA, AL, MO, MT, AZ, CA, OH, FL, SC, MI, TN, CO, NV, IN, NY, MD, AR, KS, NY, IN, IA, TX, OK, NC, WA and DC. | Open, Classified |  |  | 164
Z-0795-2016 | 73053 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | PharmaCaribe PulmoSal 7% Sodium chloride, Inhalation Solution, 7% Sterile, Preservative -free, Non-pyrogenic, pH 7.4 |  |  | December 29, 2015 | 2015 | Pharmacaribe llc | Vials were labeled as USP 7% Hypertonic saline instead of PulmoSal 7% (pH+) | Error in labeling | PharmaCaribe, together with its Contract Manufacturing firm Asept Pak, Inc., have initiated or completed the following actions to issued an Urgent Notice: Field Notification to Medical Device letter dated January t7, 2016. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: Determine what is on hand and what has been distributed.Using the Return Notice provided, report this information as per the instructions on the notice. Please record units as either cases or as cartons (if the cases have been opened and are partial). This document is to be returned via mail or email as per the guidance provided. This will inform the Manufacturer of your inventory status and of your intentions regarding product replacement. Cases of this product may be returned for new cases of the same product (different lot) that have the correct embossing on the vials. Instructions for return are provided in the Field Notification Package Cover Letter. Alternatively, customers may choose to be issued a full refund for all returned product. For cases which have already been distributed to the retail/clinician level and/or are in use by patients, it is acceptable to continue to use the product, with the understanding (and proper communication to all parties) that the tabs on the vials are incorrectly labeled.Contact information:Lisa Early, QA Manager Asept Pak, Inc.Phone: 518-651-2026 Fax: 518-651-2046 E-mail: learly@aseptpak.com |  | 7020 | 7,020 vials | US Nationwide Distribution in the states of FL, LA, TX and VA | Terminated | March 28, 2017 | 455 | 7020
Z-0797-2016 | 73069 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Firebird Spinal Fixation System Torque Limiting Handle (PN 52-1512) |  |  | January 08, 2016 | 2016 | Orthofix, Inc | There is a possibility that the device may deliver a torque value less than the required setting after initial sterilization. In worst case, the unexpected drop in torque value may lead to the spinal construct loosening and requiring corrective surgical intervention. | Component design/selection | Affected consignees were notified via an Urgent Medical Device Recall Notification letter, dated 1/8/16. The letter identified the affected device and stated the reason for the recall. The recalling firm requests customers to remove the affected devices from their inventory and cease further distribution or use. Customers whom the devices were further distributed to should be notified that the affected devices cannot be used and must be returned to Orthofix per the instructions provided. The attached Acknowledgement Form should be completed and returned. Customers can contact their local Orthofix representative or customer service for further information, replacements, or disposal instructions. |  | 36 | 36 units | Worldwide Distribution -- US, Spain, Italy, and Australia. | Terminated | November 16, 2016 | 313 | 36
Z-0688-2016 | 72923 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Moog Curlin Infusion, Ambulatory Infusion Pump |  |  | December 23, 2015 | 2015 | MOOG Medical Devices Group | Moog is recalling Curlin 4000, 6000, and PainSmart Infusion Pumps manufactured or serviced by MMDG between March 18, 2015 and November 6, 2015 because the master pumps used to calibrate them were out of tolerance and may cause an over delivery of medications and solutions at a rate of 1.2 - 6.8% faster than programmed. | Process control | Moog Medical Device Group sent an "Urgent Recall Notice" dated December 22, 2015, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove all of their current Calibration Sets from service before placing these new sets into service. Customers were also instructed to complete and return the Recall Reply Form by fax: 801-264-1051 or email to: MoogMedicalSupport@moog.com and for any questions, please contact: Moog Regulatory Affairs at 801-264-1001. For additional information or clarification call 801-264-1001, ext. 112 or via email at: cdodge@moog.com Monday through Friday, 8:00 AM to 5:00 PM, Mountain Time. |  | 11472 | 11,472 | Worldwide Distribution - US (nationwide), Guam, and countries of: Australia, Canada, Croatia, Turkey, New Zealand, Saudi Arabia, Japan, Korea, and Italy. | Open, Classified |  |  | 11472
Z-0805-2016 | 73092 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System versions 11, 13.0, 13.5 and 13.6; |  |  | January 11, 2016 | 2016 | Varian Medical Systems, Inc. | When using PBC 11.0.31 to calculate the dose for a conventional arc field with more than 100 segments for Eclipse versions 11.0, 13.0, 13.5 or 13.6, the displayed dose does not correspond to the calculated Monitor Units (MU). Potential for unintended radiation exposure. | Software design | Varian sent an Urgent Medical Device Correction letters dated January 11, 2016 to all affected customers. Letters were also sent to Varian Sales and Marketing and Service organizations to notify them of the recall. The letter identified the affected product, problem and actions to be taken. For questions customers were instructed to contact their local Varian Medical Systems Customer Support District or Regional Manager. |  | 9499 | 9499 | Worldwide Distribution - US Nationwide and the countries of Albania, Guadeloupe, Norway,Algeria, Guam, Pakistan, Argentina, Guatemala, Panama,Australia, Hong Kong, Paraguay, Austria, Hungary, Peru,Azerbaijan, Iceland, Philippines, Bangladesh, India, Poland,Belgium, Indonesia ,Portugal, Belorussia, Iran, Puerto Rico,Brazil, Iraq, Qatar, Brunei Darussalam, Ireland, Reunion,Bulgaria, Israel, Romania, Burma, Italy, Russia,'Canada, Jamaica, Saudi Arabia, Chile, Japan, Singapore, China, Kazakhstan, Slovakia, Colombia, Kuwait, South Africa, Costa Rica, Latvia, South Korea,Czech Republic, Lebanon, Spain, Denmark, Libya, Sweden, Dominican Republic, Lithuania, Switzerland, Ecuador, Luxembourg, Taiwan, Egypt, Macau, Thailand, El Salvador, Macedonia, Trinidad and Tobago, Estonia, Madagascar, Tunisia, Finland, Malaysia, Turkey, France, Malta, Ukraine, French Polynesia, Mexico, United Arab Emirates, Georgia, Monaco, United Kingdom, Germany, Morocco, Venezuela, Ghana, Netherlands, Vietnam,Greece, New Zealand, and Zimbabwe. | Open, Classified |  |  | 9499
Z-0806-2016 | 73086 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MRIdian ViewRay Radiation Therapy System, VIEWRAY Inc. |  |  | December 24, 2015 | 2015 | Viewray Incorporated | The system loaded a completion fraction in the incorrect order after a treatment interruption. | Software design | ViewRay sent an Customer Advisory Notice dated December 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Workaround: If you want to move a completion plan to a later date in the Delivery Calendar then do not schedule it until right before you are ready to treat that completion plan.While we will address this issue in the next release of software we recommend that in the event of a completion plan you always review the plan name loaded on the TPDS and verify the remaining time and beams are as expected with the completion plan.If you have any questions or concerns regarding this issue, please do not hesitate to contact ViewRay Customer Support at support@viewray.com or call the ViewRay" Customer Support center at 855-286-8875. |  | 5 | 5 units | US Distribution to the states of : CA, FL, MO &amp; WI., and Internationally to Republic of Korea. | Open, Classified |  |  | 5
Z-0808-2016 | 72802 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Bond Polymer Refine Red Detection DS9390 kits. The common name of the device is detection system for in vitro diagnostic use. The Intended Use states that Bond Polymer Refine Red Detection is a biotin-free, polymeric alkaline phosphatase (AP)-linker antibody conjugate system for the detection of tissue-bound mouse and rabbit IgG and some mouse IgM primary antibodies. It is intended for staining sections of formalin-fixed, paraffin-embedded tissue on the BOND automated system. |  |  | November 27, 2015 | 2015 | Leica Microsystems, Inc. | The Firm informed the importer/distributor, Leica Microsystems, a manufacturing error occurred during the filling of some containers of the Bond Polymer Refine Red Detection DS9390 kits, which may result in inadequate staining. | Mixed-up of materials/components | Leica Biosystems sent an Urgent Medical Device Recall Notice dated November 30, 2015, to all affected customer. The letter identified the product the problem and the action needed to be taken by the customer.Do not use or continue to use the reagent lots listed above, as they have the potential to result in inadequate tissue staining and lead to weak or false negative pathology results. Please also properly destroy any unused or partially used affected lots of the reagent, and indicate by signing and returning the attached Medical Device Recall Acknowledgement Form that this action has been undertaken. As indicated in the Instructions for Use, the clinical interpretation of any staining or its absence should be complemented by morphological studies using proper controls and should be evaluated within the context of the patients clinical history and other diagnostic tests by a qualified pathologist. The consignees are requested to consult with their laboratory director and physician to determine whether a review of past results obtained from the affected lots is clinically warranted.Direct accounts are requested to forward the notice primarily to the end users where the product has been sold and to all those within the organization who need to be aware of this issue.Periodic effectiveness checks will be conducted to address non responders.For further questions please call ( 847) 405-5413 |  | 3745 | 3745 | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Argentina, Australia, Austria, Belgium, Chile, Denmark, France, Germany, Greece, Hong Kong, Italy, Japan , Netherlands, RomaniaRussia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey and UK. | Terminated | January 05, 2017 | 405 | 3745
Z-0809-2016 | 73057 | No | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Exactech Acapella One Interbody Fusion Device Parallel. Orthopedic Spinal Surgery Device. |  |  | September 28, 2015 | 2015 | Exactech, Inc. | Incorrect external and internal packaging labels. Exactech has identified nineteen (19) Acapella One Interbody Fusion Device Lordotic (05-090-01-1610) were packaged as Acapella One Interbody Fusion Device Parallel (05-090-02-1610) | Process control | Exactech sent a recall notification letter dated September 28, 2016 to their only 2 consignees. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately cease distribution or use of these products, extend this information to accounts that have this products in their possession, identify and quarentine any of the product and complete and return the attached recall inventory response form to Exactech. Contact: Kaya Davis at Kaya@exac.com or 1-800-392-2832 |  | 19 | 19 pieces | US distribution in the state of FL and PA. | Open, Classified |  |  | 19
Z-0807-2016 | 73030 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EVOLIS Microplate System, Catalog # 89601. Part number 89788 for the EVOLIS Operator's Manual. in vitro diagnostic |  |  | December 22, 2015 | 2015 | Bio-Rad Laboratories, Inc. | The error "Washer Reagent Clean Fluid" level low alarmed, which prompted the customer to open the system drawer. When the system drawer was opened the customer noted the plate transport was still running and as a consequence the plate was pushed off the plate transporter. | Use error | Bio-Rad sent an Urgent Product Notification letter, dated December 22, 2015, to consignees via Fed-Ex. The firm recognized the typographical error of the PN number on the first letter and sent a revised notification letter, dated 20 January 2016, with the correct PN 89788 to consignees. The letter identified the affected product, problem and actions to be taken. Consignees are advised that in addition to the WARNING in Chapter 4.9.2 in the EVOLIS Operators Manual (PN 89788) mention is made of, "LED lamps on the front panel of the EVOLIS instrument illuminate when a compartment may be accessed. If an error occurs, (for example a bulk reagent becomes depleted) the system displays an error message, pauses, and allows access to the instrumentcompartments."The plate transport may still be moving the microplate to its final destination when access is allowed. To enhance safety features, your Field Safety Engineer (FSE) will apply a (yellow) warning label to the lower front panel. Labels are currently on order and once received will be provided to each FSE. The label will state: WARNING! Before opening instrument, ensure all moving parts have stopped".Consignees can contact Bio-Rad Technical Support directly at 1-800-224-6723, option #2, and then option #3 or by fax at 425-498-1628. |  | 278 | 278 units | US Nationwide Distribution including Puerto Rico. | Terminated | March 13, 2017 | 447 | 278
Z-0810-2016 | 71647 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Actuator for angiographic x-ray monitor ceiling suspension system. Actuator short Assy: THC8AWDS-121 and Actuator long Assy: THC8AWDS-122 used with the following products: 722001 Allura Xper FD 10 C; 722003 Allura Xper FD10; 722005 Allura Xper FD10/10; 722006 Allura Xper FD20; 722008 Allura Xper FD20 Biplane; 722010 Allura Xper FD10; 722011 Allura Xper FD10/10; 722012 Allura Xper FD20; 722013 Allura XPER FD20 BIPLANE; 722015 Allura Xper FD20 OR Table; 722020 Allura Xper FD20 Biplane OR Table; 722022 Allura Xper FD10 OR Table; 722023 Allura Xper FD20 OR Table; 722025 Allura Xper FD20 Biplane OR Table; 722026 Allura Xper FD10; 722027 Allura Xper FD10/10; 722028 Allura Xper FD20; 722029 Allura Xper FD20/10; 722035 Allura Xper FD20 OR Table; 722038 Allura Xper FD20/20; 722039 Allura Xper FD20/20 OR Table; 722058 Allura Xper FD20/15; 722123 Field extensions Xper cardio R7.6; 722124 Field extensions Xper vascular R7.6; 722133 Field ext. Xper cardio systems R7.6; 722134 Field ext. Xper vascular systems R7.6; 72249 INTEGRIS BV5000; 722043 INTEGRIS Allura 15-12 (mono); 722121 Field extensions INTEGRIS cardio systems |  |  | July 06, 2015 | 2015 | Philips Electronics North America Corporation | The actuator assembly became detached and the monitor carriage with a flexvision monitor dropped to its lowest position. | Process control | Philips sent a Field Safety Notice to all affected customers on July 6, 2015. The notice informs consignees of the reason for recall and actions to be taken by the customer/user. On September 15, 2015, Philips plans to issued a mandatory Field Safety Corrective action (FCO72200280) via Certified Return Receipt. This corrective action is the replacement of the affected MCS actuator. For further information, contact your local representative: Technical Support Line: 1-800-722-9377. |  | 3158 | 3158 devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of : Canada, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, Cayman Islands, Chile, China, Colombia, Costa Rica, Cuba, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Republic of Korea, Lebanon, Lithuania, Macedonia, Malaysia, Mexico, Mongolia, Morocco, Myanmar, Nepal, Netherlands, Netherlands Antilles, New Zealand, Norway, Oman, Pakistan, occupied Palestinian territory, Panama, PERU, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom and Viet Nam. | Open, Classified |  |  | 3158
Z-0812-2016 | 72977 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | REF 0038460 Toomey Syringe 70cc Syringe with Catheter Tip and Luer Tip Adapters, Single Use, Rx Only, Latex Free, 50 units. |  |  | January 04, 2016 | 2016 | C.R. Bard, Inc. | Report by a customer of an open package seal. | Packaging process control | BMD sent an Urgent Medical Product Recall letter to each of the customers via FedEx with proof of delivery notice. The leter identified the affected product, problem and actions to be taken. All consignees were instructed to carry out the recall to the user level. The letter included a Recall and Effectiveness Check Form which was to be returned to Bard. |  | 15400 | 15,400 units | US Nationwide Distribution - in the states of: AL, AZ, CA, CO, FL, GA, IL, IN, IA, KS, KY, LA, MD, MI, MS, MO, NE, NV, NJ, NM, NY, NC, OH, OK, PA, SC, TN, TX, VA, WA, WV, and Puerto Rico | Terminated | June 06, 2017 | 519 | 15400
Z-0814-2016 | 73093 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Jadak Barcode Scanner, Part No. 350587-32, 350587-41, 351377-32. |  |  | January 15, 2016 | 2016 | CareFusion 303, Inc. | Customers reported issues when scanning medications with the scanners. When a user scans a medication using the affected barcode scanner, information on a different medication could be communicated to the Pyxis product. | Device Design | BD sent an Urgent Medical Device Recall Notification letter dated January 15, 2015 to all affected customers The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete the enclosed Customer Response Card by selecting the desired deployment option and return the card to CareFusion. Customers with questions are instructed to contact the CareFusion Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact CareFusion Customer Advocacy at (800) 836-7842. Customers with technical questions are instructed to contact CareFusion Technical Support at (800) 727-6102. |  | 9345 | 9,345 units | Worldwide:Distribution - US Nationwide and the countries of Saudi Arabia, Bahrain, Australia, United Arab Emirates, Qatar, Mexico, Guam and the Bahamas. | Terminated | May 08, 2017 | 479 | 9345
Z-0813-2016 | 73224 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Dual Stage Venous Return Catheter, Low-Profile, wire reinforced (36F/46F, Sterile), Model Number: LRD-61046, Expiration 10/31/2018, Product is packed in a sealed pouch, ten (10) pouches per case. Cardiopulmonary bypass catheter. |  |  | January 25, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group initiated field correction for the Dual Stage Venous Return Catheter because the incorrect Instructions For Use (IFU) was packaged in the case. | Employee error | LivaNova sent an Urgent Safety Alert dated January 25, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Advice on action to be taken:"Remove and discard the IFU included within the case packaging for the product."A copy of the correct IFU (Part Number 435085202, Revision B) is attached with this letter"Print copies of the IFU and place with the product.For questions regarding this notice, please contact Customer Service at 1-800-650-2623 (Monday to Friday BAM to 5PM MT) or e-mail USFieldSafetyActions@livanova.com.We sincerely apologize for this inconvenience and thank you in advance for your support |  | 60 | 60 | US Distribution to CA., and Internationally to Canada. | Open, Classified |  |  | 60
Z-0816-2016 | 73164 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LEGION(TM) HEMI STEPPED, TIBIAL SCREW-ON WEDGE, REF 71421147, QTY 1, 5 MM, SIZE 5-6 LT-LAT / RT-MDL; Tibial knee prosthesis, augment. |  |  | January 18, 2016 | 2016 | Smith &amp; Nephew, Inc. | The package for catalog item, a LEGION Hemi Stepped Tibial Wedge, 5 MM Size 5-6 LT-LAT/ RT-MDL, should include both a screw-on hemi stepped tibial wedge and two tibial wedge screws to affix the wedge to the tibial baseplate. A complaint was received indicating that the screws packaged with the wedge were too long. | Packaging change control | The firm notified their direct consignee of the issue on 01/18/2016 by email. The notice requested return of the product. |  | 8 | 8 units | Distributed in the states of TN, WI, OH, and KS, and in the countries of Belgium, Germany, and United Kingdom. | Open, Classified |  |  | 8
Z-0815-2016 | 73120 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | CEA / CD66e Ab-2, Product code RB-368-R7, RB-368-R1; 7 ml each. |  |  | January 18, 2016 | 2016 | Lab Vision Corporation | Certain lots of CEA/CD66e Ab2 were produced with a dilution factor of 1:100 instead of the recommended 1:250. | Nonconforming Material/Component | ThermoFisher sent an Urgent Medical Device recall letter dated January 18, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed if they have any of the affected product to contact Lab Vision with the informationrequested on the attached Acknowledgment Form.Âż Upon identification of the affected lot, please notify Lab Vision of the quantity previously used and the quantity still in inventory.Âż If any of the above listed lots were previously used, please review test results to ensure a valid result was obtained. If a valid result was obtained, no further action is required. If high background staining was seen, standard IHC troubleshooting guidance for high background staining can be used.(See RISK TO HEALTH section of the letter.)Âż Affected product may be returned to Lab Vision. Please indicate on the Acknowledgement Form whether the product will be returned. Upon receipt of the form, RGA information will be sent for product requiring return.Âż If it is preferred to destroy the product at the customer site, please indicate this on the Acknowledgement form.Âż Please use the attached Recall Return Response Acknowledgement &amp; Receipt Form.Customers with questions were instructed to call 269-544-5628. |  | 776 | RB-368-R7 - 30; RB-368-R1 -2 | Worldwide Distribution - US including CA, NY, GA, MS, PA, MA, TX, KY, AZ. and Internationally to Greece, France, Germany, Singapore and Serbia. | Open, Classified |  |  | 776
Z-0817-2016 | 73114 | No | GWF | Evoked response electrical stimulator. | Neurology | Neurology | 510(k) | MEE-1000A Neuromaster, Model No. Accessories: MS-110B (Electric Stimulation), MS-120BK (Current Stimulator Box) JB-116B, JB-132B (Amplifiers) Electromyography. (EMG) |  |  | January 19, 2012 | 2012 | Nihon Kohden America Inc | The MEE-1000A Neuromaster had defective printed circuit boards (PCBs). Gas arrestors on these boards which were intended to provide protection from electrostatic discharge failed to prohibit unintentional current flow. | Device Design | A customer notification letter dated 9/25/12 was sent to customers informing them of a potential malfunction of NIHON KOHDEN amplifier units, models JB-116B andJB-132B and Extension Unit MS-110B, used with the MEE-1000A Neuromaster Neural Function Measuring System. The letter informs the customers of the potential malfunction and that the probability that the malfunction occurs is extremely rare. The letter informs the customers of the preventive measure and the preventive corrective action. Customers are instructed to inform all potential users and operators of the Neural Function Measuring System MEE-1000A Neuromaster about the Corrective Action notice. Customers are instructed that NIHON KOHDEN technical support representative will contact them to arrange an update of their amplifier products. |  | 477 | 477 units total (137 units in US) | Worldwide distribution. US nationwide, Australia, Austria, France, Germany, Greece, Japan, Romania, South Africa, South Korea, Spain, and United Kingdom. | Terminated | February 22, 2016 | 1495 | 477
Z-0818-2016 | 72911 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | High Flow Irrigation Extender; Part Number 72201021. |  |  | July 30, 2013 | 2013 | Smith &amp; Nephew, Inc., Endoscopy Div. | During arthroscopic surgery, when the scope is moved inside the cannula, the body of the irrigation extender separates, potentially obstructing visualization and causing leakage of the irrigation fluid. The device is removed and a back-up device is used to complete the surgery. | Under Investigation by firm | Urgent - Product Recall 1st Notification - Urgent R-2013-13 letters, dated July 30, 2013 were sent overnight via Federal Express to customers. The letters identified the affected product, as well as the potential risk involved. Actions to be taken were provided for both hospital and sales representatives. Hospital representatives were asked to inspect their inventory, locate all devices from the list provided, and quarantine them immediately. The Inventory Return Certification Form should be completed and returned, regardless of whether any affected product is on hand or not. Contact the Smith &amp; Nephew Returns Group at 800-343-5717 option 3) or send an e-mail to ProductRecovery@smith-nephew.com. They will provide instructions for returning affected product and receiving a replacement. |  | 203 | 203 units | Worldwide Distribution -- US, including the states of AL, AR, CA, FL, GA, ID, IL, KS, MA, MI, MO, NC, TN, TX, and AK; and, country of Japan. | Open, Classified |  |  | 203
Z-0819-2016 | 72749 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Endo GIA Ultra Universal 12 mm Short Stapler Handle |  |  | December 02, 2015 | 2015 | Medtronic | Staplers fail to fire or partially fire and reports of the instrument articulating lever disengaging during use | Process control | Medtronic sent an Urgent Medical Device Recall letters dated December 2 2015 and an Updated Urgent Medical Device Recall letter dated January 13, 2016, to affected customers, via Federal Express. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473.For questions or concerns, please do not hesitate to contact your Medtronic representative or Customer Service at (800) 882-5878. |  | 364768 | 364,768 units | Worldwide Distribution - US Nationwide and the countries of Canada, Albania, ALGERIA, Argentina, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Gabon, Georgia, Germany, Greece, Greece, Guatemala, Honduras, Hungary, Iceland, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Mauritius, Mexico, Monaco, Montenegro, Morocco, Namibia, Netherlands, Nicaragua, Nigeria, Norway, Panama, Paraguay, Peru, Poland, Portugal, Romania, Russia, Russian Federation, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, UAE, Ukraine, United Kingdom, Uruguay, Venezuela, and Zimbabwe. | Open, Classified |  |  | 1797442
Z-0828-2016 | 73180 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Aero-AL Impaction handle Catalog Number 48921007. Stryker Spine Manufactured by : Stryker Spine SAS France |  |  | December 18, 2015 | 2015 | Stryker Spine | Reported jamming of the impaction handle when assembled to the Inserter guide at the threads. | Nonconforming Material/Component | Stryker Spine sent Urgent Product Removal notification letters to branches/agencies on 12/18/2015 via FedEx. The letters provided a description of the affected product, the issue, potential hazards and risk mitigations, and actions to be taken. Customers were asked to examine their inventory and hospital locations to identify the product and reconcile with the affected catalog and lot numbers provided. The Customer Response Form and Product Quantity Tracking Spreadsheet for Health Care Facilities should be completed and returned within five days of receipt of the notice to the e-mail or fax number provided. Questions can be directed to Nabil Riaz - Regulatory Compliance Specialist, Regulatory Compliance &amp; Quality Systems at Nabil.Riaz@Stryker.com. |  | 31 | 31 units | Nationwide Distribution | Open, Classified |  |  | 31
Z-0803-2016 | 73251 | No | RBY |  |  |  |  | Olympus Scientific Solutions Americas Corporation ( OSSA ) Delta XRF Analyzer . This is a Analytical X-ray system |  |  | January 13, 2016 | 2016 | Olympus Scientific Solutions Americas | The Firm has discovered a Software bug. | Radiation Control for Health and Safety Act | Olympus Scientific Solutions Americas ( OSSSA) planned action to bring product into compliance: Notification will be given to purchasers per 21 CFR 1003.21. A software modification will be made available to all users contacted. The letter to affected purchasers makes clear that the product will be brought into compliance without charge and offers the purchaser several options for applying the software update.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Tracking of customer/owner self-installation(s) of the software update should be done to the maximum possible extent via serial-number- controlled downloads, follow-up email reporting, return-mail card or some other similarly effective means. For further questions you may call : (866) 446-6689. |  | 5000 | US - 5000 units | Worldwide Distribution - USA (nationwide) Distribution | Open, Classified |  |  | 5000
Z-0826-2016 | 73274 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | software for Syngo Dynamics |  |  | January 29, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is releasing a software update that addresses an issue of mixing data from multiple patients. In rare situations, echo trend graphs may mix data from multiple patients. | Software design | The firm, Siemens Healthcare, sent a "Customer Advisory Notice" dated 1/19/2019 to its customers.The notice described the product, problem and actions to be taken. The customers were instructed to create a new department if immediate resolution is necessary, use workplace for reporting and hand over the customer letter before or during installation of field modification(PSL) distributed with Update Instruction SY013/16P. A software update to resolve reported issues is being released via an Update Instruction SY013/16/P. If you have any questions, contact Regulatory Technical Specialist at 610-448-6478 or email: anastasia.sokolova@siemens.com. |  | 151 | 151 | US Nationwide Distribution including states of: AL AR AZ CA CO CT DE FL GA IL IN KS LA MA MD MI MN MO NC ND NE NJ NM NY OH OR PA SC TN TX UT VA WA and WI. | Open, Classified |  |  | 151
Z-0824-2016 | 73099 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | BCS Automated Blood Coagulation Analyzer System, Multipurpose system for in vitro coagulation studies, Catalog Numbers: 10454728 (BCS RECONDITIONED), 10454729 (BCS RECONDITIONED), 10454742 (BEHRING COAGULATION SYSTEM), 10459303 (BCS RECONDITIONED), 10460659 (BCS INSTRUMENT), 10461881 (BCS INSTRUMENT), IVD --- Device Listing D018966 |  |  | January 14, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | False short clotting times for PT on BCS and BCS XP with Dade Innovin can occur. This has also an effect for false low PT INR values. These false short clotting times were observed only with turbid, hemolytic and/or icteric samples from intensive care patients. The corresponding reaction curves of these samples showed a pre-peak, which can lead to falsely short clotting time. The result was not flagged by the existing check algorithms. | Software design | Urgent Medical Device Correction Letter (dated 1/12/2016) was sent to US consignees via FedEx on 1/14/2016 and to foreign consignees (distribution is determined at country level). The letter re-emphasizes the importance of checking the interference levels given in the respective Application Sheets of the BCS and BCS XP systems. For questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 1680 | 1680 units total | Worldwide Distribution-US (nationwide) and the countries of Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bosnia Herzegovina, Brazil, Canada, Chile, China, Croatia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Guadeloupe, Hungary, Italy, Kuwait, Latvia, Lebanon, Libya, Macedonia, Mexico, Montenegro, Netherlands, Peru, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, and Uruguay. | Terminated | March 16, 2017 | 427 | 1680
Z-0823-2016 | 73095 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | This reagent is for in vitro diagnostic use. Pax-5 (1EW) monoclonal antibody is intended to be used for the qualitative identification by light microscopy of human Pax-5 in formalin-fixed, paraffin-embedded tissue by immunohistochemical staining using the automated Bond system. |  |  | December 30, 2015 | 2015 | Leica Microsystems, Inc. | The Bond" Polymer Refine Detection and Novolink" Polymer Detection System may not provide adequate staining when using a detection protocol where one of its components - Peroxide Block / Peroxidase Block is applied after the primary marker incubation as specified in the Instructions for Use Following the Instructions for Use. | Component change control | The firm sent out a Medical Device Recall letter dated December 29th 2015 to two-hundred (200) consignees in the United States decribing the recall and the neccesary actions to be taken. These included: - Do not use or continue to use the listed antibody with the specified detection kits.- Order replacement detection kits for use with the listed antibody.- There is no need for retrospective review of tissue staining if appropriate positive controls were properly used and interpreted during the immuno / in-situ staining. -pass this notice primarily to the end users where the product has been sold and to all those within your organization who need to be aware of this issue.-Confirm reciept of the Recall Notice letter by signing and dating the attached Acknowledgement Form and faxing it back to us at 1-847-236-3747, or scan and email it to LMGRA@Leica-Microsystems.com Customer are requested to contact their local Leica representative if you have any questions or concerns, or alternatively please contact us at the following email address: Combination.Products@LeicaBiosystems.comPeter LloydLeica Biosystems Newcastle LtdBalliol Business Park WestBenton LaneNewcastle upon Tyne NE 12 8EWUnited KingdomTelephone +44 191 215 0567Facsimile: +44 191 215 1152 |  | 1700 | 1700 units | US Consignees incluude: AR ,AZ ,CA ,CO ,CT ,DC ,DE ,FL ,GA ,HI ,IA ,IL ,KY ,LA ,MA ,MD ,MI ,MN ,MO ,MT ,NC ,NH ,NJ ,NY ,OH ,OK ,PA ,PR ,SC ,TN ,TX ,UT ,VA ,VT ,WA ,WI and WV. Foreign Consignees include:Argentina , Australia , Brazil , Bulgaria , Chile , China , Ecuador ,Egypt , El Salvador , Finland , France , Georgia , Germany , Greece , India , Italy , Japan , Jordan , Kuwait , Lebanon , Malaysia , Philippines , Poland , Portugal , Romania , Russian Fed. , Singapore , South Korea , Spain , Sweden , Switzerland , Thailand , Tunisia , Turkey , United Kingdom and United Arab Emirates. | Terminated | January 05, 2017 | 372 | 1700
Z-0804-2016 | 73137 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Flow-i Anesthesia System |  |  | December 10, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Electrical issues in the control, monitoring or panel sub-systems, caused by the printed circuit boards (PCBs) that control the sub-systems. The problems manifested as multiple re-starts of the sub-systems and, in a few instances, disabling of the system has occurred. Replacement PCBs are now available to correct the issues.Only the Model C20 is affected in the US. | Nonconforming Material/Component | Maquet sent "Urgent Medical Device Field Action Field Correction For Flow-i Anesthesia Systems" letters/Field Correction Response forms dated 12/10/2015 via UPS to affected customers. The notification informed customers of the issue, asked them to distribute the notification to appropriate staff; verify the serial number(s) and informed them they could continue to use the system if they were not experiencing electrical issues. Customers were informed that a Maquet service representative would be contacting them to arrange a time to replace their affected units. Customers were asked to complete and return the enclosed Field Correction Response Form to Maquet at FieldActions@maquet.com or by fax to 973.860.1675.If the customer had any related questions they were instructed to contact Maquet Technical Support at 1.888.627.8383 (option 3; option 1; option 1) Monday through Friday between 8:00 am and 5:00 pm EST. |  | 7 | 7 units within US | US Nationwide in the states of CA, NY, and OR | Terminated | July 06, 2016 | 209 | 7
Z-0830-2016 | 73142 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LPS-FLEX GSF OPT |  |  | January 11, 2016 | 2016 | Zimmer Biomet, Inc. | LDPE bag containing the implant adheres to the highly polished implant surface. | Process control | Zimmer Biomet sent an Urgent Medical Device Recall dated January 11, 2016, to all affected customers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers with questions or concerns were instructed to contact customer service at 1-800-348-2759 or contact their Zimmer Biomet Sales Representative.For questions regarding this recall call 800-447-5633. |  | 84396 | 84396 | Worldwide Distribution - US (nationwide) and Internationally to ARGENTINA AUSTRALIAAUSTRIABARBADOSBELGIUMBERMUDABOLIVIABRAZILCARIBBEANCHILECHINACOLOMBIACOST RICADENMARKDOMINICAN REPUBLICDORAL FLECUADOREL SALVADORENGLANDFRANCEGERMANYGUATEMALAHONDURASHONG KONGINDIAISRAELITALYJAMAICAJAPANKOREAMALAYSIAMEXICONEW ZEALANDNICARAGUAPANAMAPERUSINGAPORESPAINSWEDENSWITZERLANDTAIWANTHAILANDVENEZUELAWEST INDIESCanary IslandsCroatiaCyprusCzech RepublicCanada | Terminated | May 23, 2017 | 498 | 738203
Z-0863-2016 | 73104 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Symbiq infusion pump is a device used in a health care facility to pump fluids into a patient in a controlled manner. It is intended primarily for use in the hospital setting and can be used in other acute and non-acute areas, such as, but not limited to, Nursing Homes, Mobile Intensive Care, Ambulatory Infusion, Centers, Hospice, Sub acute facilities, Outpatient/Surgical Centers, Long-term Care, Urgent Care, Transport, and Physician Offices. |  |  | December 22, 2015 | 2015 | Hospira Inc. | Potential for the Symbiq pole clamp assemblies to be missing therubber stop pad. | Nonconforming Material/Component | On December 22, 2015 Stericycle, Inc. sent an URGENT MEDICAL DEVICE CORRECTION notification to direct consignees of record via traceable mailing. The notification included a description of the recall, risk to health information and necessary actions to be performed by the direct account. These include: an inspection of the affected devices for missing rubber stop pad on the Symbiq pole clamp assemblies at the direct accounts.If missing rubber stop pads were discovered they were advised to contact Hospira's Advanced Knowledge Center at 1-800-241-4002, option 4, to receive a replacement pole clamp assembly or be directed to a field Service Engineer who will replace the pole clamp.Hospira ALso advised informing healthcare professionals at their institution about the issue and complete the attached reply form and return it by e-mail or fax at: 1-800-441-4100/1-800-241-4002 or at ProductComplaintsPP@hospira.comFinally customers where advised to notify FDA of any adverse events through FDA's MedWatch Adverse Event Reporting program. |  | 6634 | 6634 units. | U.S. consignees nationwide.*** Foreign consignees in: Canada. | Open, Classified |  |  | 6634
Z-0858-2016 | 72720 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Healthcare IntelliVue Info Center iX, B.0, 866389 |  |  | November 25, 2015 | 2015 | Philips Electronics North America Corporation | Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads. | Software Design Change | Philips Healthcare issued a Field Safety Notice on 11/25/15 to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will provide a software correction free of charge. On 1/28/16 the Philips issued an updated letter to users.Philips has initiated a correction to address this issue. A software correction will be provided to customers with impacted devices at no charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice: Arrhythmia alarms are not affected. ST alarms, ST values and measurements are also not affected and could be used to assist with the decision to obtain a diagnostic ECG.The ECG waves at the bedside are not affected. Printing ECG waves from the bedside is an alternative to printing from the PIIC iX.If you need any further information or support concerning this issue, contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. |  | 1880 | 1880 units | Worldwide Distribution-US (nationwide) and the countries of Algeria Australia Austria Belgium Bulgaria Canada Denmark Estonia Finland France Germany Hong Kong Hungary Iceland India Indonesia Iraq Ireland Israel Italy Japan Kenya Korea Kuwait Luxembourg Malaysia Mongolia Morocco Netherlands New Zealand Nigeria Norway Poland Qatar Romania Russian Federation Saudi Arabia Singapore Slovenia South Africa Spain Sweden Switzerland Taiwan Turkey United Arab Emirates and United Kingdom. | Open, Classified |  |  | 13120
Z-0852-2016 | 73179 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | AccuLIF PL Tubing Set and AccuLIF TL Tubing Set |  |  | December 16, 2015 | 2015 | Stryker Spine | There were two potential interference conditions identified with the way the tubing set attaches to the inserter. | Device Design | Stryker Spine sent an Urgent Product Removal letter dated December 17, 2015 to affected customers via Fed Ex priority overnight. The letter identified the affected product, problem and actions to be taken. Customers are asked to follow instructions included in the letter. The customer response form should be completed and sent back to email via Spine-RegulatoryActions@Stryker.com or Fax 855-632-9049. For questions call (201)749-8389. |  | 413 | 413 units | US Nationwide Distribution | Open, Classified |  |  | 413
Z-0853-2016 | 73209 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Measurement Module X1 |  |  | January 04, 2016 | 2016 | Philips Electronics North America Corporation | The ST elevation alarm on the Patient Monitor or standalone X2 Measurement Module will not sound when indicated for all chest leads derived using Hexad 12-Lead ECG Monitoring in the Host Monitor | Software Design Change | Philips Healthcare issued the Field Safety Notice on January 4, 2016, and informed customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice: Until your software is upgraded, please make sure that the ST Analysis is switched ON when using STE measurement in the Hexad 12-lead ECG Monitoring. This can be done by entering the ST Analysis Menu and select ST Analysis to On . For more detailed information, please refer to the Instructions for Use (IFU) of your host monitor or to the X2 Measurement Module IFU.Philips will provide a software correction free of charge. For further questions please call (978) 659-3000. |  | 13300 | 13,300 | Worldwide Distribution - Nationwide Distribution and to the countries of : Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iran ( Islamic Republic of ) Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Morocco, Myanmar(Burma) Netherlands Antilles, Norway, Oman, Pakistan, Philippines, Poland, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Province of China, Thailand, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 13300
Z-0855-2016 | 73042 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Stryker Vitallium Wire (6 Pack) Non Sterile. For use in bone procedures. |  |  | December 23, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | The wire packages are correctly marked with a "NON-STERILE" label, however, the enclosed Instructions For Use (IFU) states the device are sterilized via gamma irradiation and should not be resterilized. Since the IFU states the devices are sterile, no instructions for moist heat sterilization are provided. | Under Investigation by firm | Stryker Branches/Agents were notified of the action via email on December 23, 2015. Urgent Medical Device Recall Notification Letters/Urgent Medical Device Recall Notification Acknowledgement Forms dated December 23, 2015 were sent to Branches/Agents via UPS on December 28, 2015. Urgent Medical Device Recall Notification Letters/Urgent Medical Device Recall Notification Acknowledgement Forms dated December 23, 2015 were sent to Hospital Risk Managers and doctors via UPS on 12/23/2015. The notification instructed customers of the related issue with the affected product; how to identify affected product; the potential hazards associated with affected product; risk mitigations; and actions needed to be taken.Customers were asked to complete and return the attached Product Recall Acknowledgement Form within 5 days and either email (strykerOrtho7808@stericycle.com or Fax (1-866-672-0627 ) theresponse form back. Customers were instructed to return all affected products to Stryker C/O Stericycle, 2670 Executive Drive, Suite A, Indianapolis, IN 46241 (Attn RA2015-120-Event 7808). A point of contact was provided in case the customer had any questions 201.831.5272 . |  | 3741 | 3,741 units | Nationwide Distribution. | Open, Classified |  |  | 3741
Z-0862-2016 | 73041 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Plaza Picture archiving and communication system. |  |  | December 22, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Siemens is releasing an updated software version to address several software issues including RGB images will show "?" since calculation of HU is not possible; save as option enabled; changes in access for loading studies; breast region is now properly fitted to segment boundary when clicking fit breast to screen. | Software design | Siemens sent a Voluntary Correction and Removal Action letter dated December 30, 2015. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that a software update VB10A_HF01/06 is being released simulataneously to eliminate the issue. Customers with questions were instructed to call 1-888-826-9702. |  | 68 | 68 | Nationwide Distribution including ALCACTFLINKSLAMAMDMIMNMOMSNENYOHORPATXWAWI | Open, Classified |  |  | 68
Z-0861-2016 | 73168 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Convenience Kits from Centurion that contain an In-Vision- |  |  | January 19, 2016 | 2016 | Centurion Medical Products Corporation | Convenience Kits contain an In-Vision-Plus Needleless IV Connector, code RYM5001 that is being recalled by the manufacturer, RyMed Technologies, LLC. A component molding abnormality may cause female luer threads to crack, leading to potential for the mating tubing set to become detached and for the administration to be interrupted. Other components within these kits are not affected. | Nonconforming Material/Component | The firm, RyMed Technologies, LLC, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated 1/4/2016 to the affected customers via certified mail on 1/19/2016 with instructions for identifying and returning the affected product. The recall letter included a description of the reason for the recall, affected product, actions to be taken, and instructions for responding to the formal recall notification. The customers were instructed to discontinue distribution and/or use and return product to: RyMed Technologies, LLC, 6000 W. William Cannon Drive, Building B, Suite 300, Austin, TX 78749; that user facilities change the IV connector within 24 hours when lipids are being administered; if you further distributed any of the lots, contact your accounts, advise them of recall and return their outstanding stock, complete and return enclosed Acknowledgement and Receipt Form ASAP and no later than 10 business days via E-mail: amccutchen@rymedtech.com, Fax: 512-301-7338 or by mail to address above, Attention Recall Return Notification. Please contact Director of Quality Assurance/Regulatory Affairs at (512) 301-7334 Ext.303 or email: ammccutchen@rymedtech.com should you have questions pertaining to recall notification. Questions pertaining to Returning Product/Replacement Product contact Logistics Specialist at (512) 301-7334 Ext.309 or email svara@rymedtech.com. |  | 1130 | 1130 | US Distribution in states of: TX, FL, VA, TN, PA, IL, GA, and CA. | Terminated | July 19, 2016 | 182 | 1130
Z-0851-2016 | 73155 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | BD FMC7 FITC: Model 430918. |  |  | January 25, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | Becton,Dickinson and Company (BD) has determined that the FMC7 FITC (ASR) has an incorrect expiration date listedon the vial label. | Under Investigation by firm | BD Biosciences sent an Urgent Medical Device Recall letters dated January 2016 via certified mail to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed that the remaining stock of the product can be used until 6-30-16. however, if prefer they can discard all remaining stock of the affected lot and BD will replace.The letter request customers complete the Tracking/verification form and returned to BD Biosciences, per the instructions (mail, fax, or email).For questions contact your local BD representative or BD Customer Support at 877-232-8995 (prompt 3,3). |  | 78 | 78 | Worldwide Distribution - US Nationwide in the staates of TX, NY, GA, NJ, AL, MA, IL, PA, DC, FL, CA, MD, OK. and coutry of Taiwan. | Open, Classified |  |  | 78
Z-0859-2016 | 73139 | No | LGW |  |  | Neurology | PMA | Specify 5-6-5 Lead, Specify 2x8 Lead, Model Numbers: 39565, 39286. Spinal cord stimulation leads. Part of a neurostimulation system for treating chronic, intractable pain of the trunk and/or limbs-including unilateral or bilateral pain. |  |  | January 11, 2016 | 2016 | Medtronic Neuromodulation | Some Specify 5-6-5 and 2x8 surgical lead product labeling does not adequately distinguish between permanent surgical implant and trialing/screening use. | Labeling design | Consignees were sent a Medtronic "Medical Device Correction" letter dated January 2016. The letter was addressed to Healthcare Professional. The letter described the problem and the product involved in the recall. Advised consignees to ensure their staff of the action that is being recommended. Requested consignees to complete and return the "Physician Confirmation Form". For additional information they can contact Medtronic Neuromodulation Technical services at 1-800-707-0933 weekends 7 am - 6 Pm Central time. |  | 51138 | 51,138 (44,006 US; 9,132 OUS) | Worldwide Distribution, including US Nationwide. | Open, Classified |  |  | 51138
Z-0860-2016 | 73138 | chair, geriatric | FRJ | Medical chair and table. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom Bariatric Power Recliner, Product Model # P9096 |  |  | January 20, 2016 | 2016 | Hill-Rom, Inc. | The recliners were shipped without the foot rest ottoman linkage side shields, which could result in consumer injury. | Under Investigation by firm | Hill-Rom sent an Urgent Field Safety Notice by POD (Proof of Delivery) mailing on January 20, 2016, to all affected customers.. The notice identified the product, the problem, and the action to be taken by the customer. Customers willl be contacted for appointments to update the power chairs with missing linkage side shields. Hill-Rom will be installing the missing linkage side shields to the Hill-Rom P9096 Power Bariatric Recliner to address this issue. Customers were instructed to ensure patients and others do not insert an appendage into the foot rest linkage mechanism during opening or closing. Customers may contact Hill-Rom Technical Support at 800-445-3720, if they have any questions. |  | 216 | 216 | Worldwide Distribution - US including AL, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, VT, WA, WI, WV, WY, VA Medical Center and Internationally to Canada and Mexico.**Consumer Recommended Depth - Consumers/User** | Open, Classified |  |  | 216
Z-0865-2016 | 73163 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Flexor Check-Flo Introducer, Catheter introducer, |  |  | January 20, 2016 | 2016 | Cook Inc. | Incorrect packaging. Reports of packages labeled as 6.0Fr actually containing 5.0Fr devices. | Under Investigation by firm | On 1/20/2016, URGENT: MEDICAL DEVICE RECALL notifications were sent to the hospital Risk Management/Recall Administration via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235. |  | 25 | 25 | Distributed in the states of CA, FL, OK, OH, MI, TX, NJ, and NY. | Open, Classified |  |  | 25
Z-0801-2016 | 73056 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Femoral Artery Pressure Monitoring Catheter Sets and Trays |  |  | January 06, 2016 | 2016 | Cook Inc. | A manufacturing process may lead to catheter tip fracture and or separation. Potential adverse events may include loss of device function, medical intervention to retrieve a separated segment, or complications resulting from a separated tip occluding blood flow to end organs. These catheters are used in venous or arterial pressure monitoring, blood sampling, and administration of drugs and fluids. | Process control | On 1/6/2016, URGENT: MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital Risk Management/Recall Administration via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235. |  | 1515 | 1515 | Worldwide Distribution - US Nationwide including DC and PR including the countries of Australia, Germany, China, Chile, Dominican Republic, Panama, Israel, Brazil, Canada | Terminated | August 23, 2017 | 595 | 17827
Z-0866-2016 | 72936 | No | DSF | Paper chart recorder. | Cardiovascular | Cardiovascular | 510(K) Exempt | PRN 50-M+, GE part number 2062759-001, manufacturer part number 600-23310-01 |  |  | December 02, 2015 | 2015 | Gsi Group Inc | The item was recalled due to an increased user/patient risk of exposure to fire and/or electrical shock due to the failure of the PRN 50-M+ Digital Writer | Nonconforming Material/Component | On December 2, 2015 GSI Group Corp. distributed Medical Device Correction Notice letter and confirmation forms dated November 18, 2015 to their customers via FedEx. GSI Group Corp. warned customers about a potential safety issue associated with the flammability and fluid ingress rating of the outer enclosure material. Customers are advised to follow all intended usage, cleaning and maintenance instructions that are found in the PRN 50-M+ Digital Writer Technical Manual. It is recommended that customers complete and return the attached Medical Device Correction Confirmation form via email to customerservice@insidesales-solutions.com. GSI Group Corp. will provide a correction at no charge. Once the confirmation form is received, GSI Group Corp. will make arragements to send replacements. Customers with questions can call 781-266-5700. |  | 2257 | 2257 units | Worldwide Distribution - US Nationwide including Puerto Rico; Internationally: Australia, Bahamas, Canada, Germany, Finland, France, Great Britain, Hong Kong, India, Ireland, Italy, Japan, South Korea, Mauritus, Mexico, Peru, Saudi Arabia, Singapore, Taiwan, United Arab Emirates &amp; United Kingdom | Open, Classified |  |  | 2257
Z-0753-2016 | 72515 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Precision MPi is an all-digital multipurpose tilt-C x-ray system, intended for a multitude of diagnostic procedures, including radiology, fluoroscopy, interventional procedures, vascular and non-vascular procedures, and specialized applications including angiographic studies. |  |  | July 01, 2008 | 2008 | Regulatory Insight, Inc | It was discovered that the Remote Touch Panel (RTP) of the GE Precision MPi X-ray system may not always boot up as intended and needs to be updated to properly accomplish its intended purpose. | Radiation Control for Health and Safety Act | Nordisk Rontgen Teknik ( NRT ) Planned action to bring defect into compliance:1.Nordisk Rontgen Teknik A/S (NRT) will contact customers and initiate a software and hardware update to correct the defect.2.NRT engineers will install the software and perform testing to ensure the software and hardware update was effective.3.The customer notification letter which includes a statement that NRT will without charge, remedy the defect or bring the product into compliance.For further questions please call (720) 962- 5412. |  | 43 | 43 | US Distribution to the states of :TN, CA, WA, OR, PA, WI, FL, MN, VT, NY, NV, MA, CO, IN, MI, MS and OH. | Terminated | July 11, 2017 | 3297 | 43
Z-0864-2016 | 73195 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | EkoSonic Endovascular System, Catalog Number 500-56112. |  |  | January 21, 2016 | 2016 | EKOS Corporation | The EkoSonic kit, including the individually labeled MicroSonic Device (MSD) and Intelligent Drug Delivery Catheter (IDDC) are labeled with a 135 cm working length, and programmed with a 135 cm family code, however the MSD actual working length is 106 cm. | Error in labeling | All affected consignees were contacted by telephone January 20- 22, 2016.The URGENT  VOLUNTARY PRODUCT RECALL letter, dated January 21, 2016, was sent to consignees. Consignees were advised that replacement devices were shipped to their facility on January 21, 2016 and should arrive there on January 22, 2016. Consignees with questions can call: Jocelyn KerstenVice President, Regulatory and Clinical AffairsEKOS Corporation, a BTG International Group Company425-415-3132Jocelyn.kersten@ekoscorp.com |  | 10 | 10 units | Distributed in the states of AZ, CA, MA, MI, OH, and VA. | Terminated | March 15, 2016 | 54 | 10
Z-0798-2016 | 72387 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Shuttle Select Slip-Cath Catheter. Angiographic catheter. |  |  | October 07, 2015 | 2015 | Cook Inc. | Reports of catheter tip splits and or separation. Loss of device function, medical intervention to retrieve separated segment, or complications resulting from separated tip occluding blood flow to end organs may result. Recall expanded 10/07/2015 to include additional catheters. | Component design/selection | On 7/2/2015 Cook Incorporated sent, URGENT: MEDICAL DEVICE RECALL notifications dated July 2, 2015, to affected customers with instructions for returning the affected product. All customers were notified viacourier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Should you have anyquestions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or (812) 339-2235. Firm issued press release 8/3/2015.On 10/09/2015 the firm expanded the recall and reissued press release via www.businesswire.com to include additional catheter products and lots as specified in the release. |  | 9519 | 9519 | Worldwide Distribution - US (nationwide including Puerto Rico and US Virgin Islands), Algeria, Antigua, Armenia, Austria, Australia, Bahamas, Belgium, Bangladesh, Bulgaria, Bahrain, Brazil, Brunei Darussalam, Canada, Canary Islands, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, El Salvador, Finland, France, French Polynesia, Georgia, Germany, Ghana, Great Britain, Greece Guadeloupe, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jersey, Jordan, Kuwait, Kazakhstan, Republic of Korea, Kuwait, Lebanon, Lithuania, Luxemburg, Latvia, Libya, Macau, Malaysia, Malta, Martinique, Mexico, Mongolia, Morocco, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Sri Lanka, Sweden, Switzerland, Spain, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Open, Classified |  |  | 9519
Z-0891-2016 | 73088 | No | GCQ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus, Model No. CF-H180AL, flexible scope, large diameter, video colonoscope |  |  | January 20, 2016 | 2016 | Integrated Medical Systems Inc | If excessive pressure is applied to the distal end of the endoscope, the internal coil pipe assembly at the distal end of the endoscope can become detached and angulation may not operate as intended. | Nonconforming Material/Component | The firm notified their consignees of the issue by Fed Ex on 01/20/2016. The firm plans to coordinate with the medical facilities for inspection of the endoscopes. |  | 232 | 232 instruments | Distributed US (nationwide) and in Australia. | Open, Classified |  |  | 1583
Z-0698-2016 | 72837 | No | KFZ | Therapeutic humidifier for home use. | Anesthesiology | Anesthesiology | 510(K) Exempt | WARM STEAM Personal Steam Inhaler-Adjustable control to regulate flow of steam Soft touch facial &amp; nasal masks" |  |  | November 20, 2015 | 2015 | Bigwall Enterprises | "The public reason for the recall is that the product is being recalled out of an abundance of caution due to one or more consumer complaints, which have been independently investigated and remain unconfirmed. No causal connection between the complaints and the productaÂż"s proper use as instructed have been determined." | Unknown/Undetermined by firm | The firm has only sold the product to Walgreens. Walgreens has notified all of its distribution centers and stores. According to the firm: "Walgreens has issued a stop sale at the store level. Units removed from store shelves and in stock will be sent back to the Walgreens distribution center Valparaiso, Indiana. As units are received from all store locations they will be sent to Bigwall Enterprises." |  | 87503 | 87503 | Nationwide Distribution including Washington, DC, U.S. Virgin Islands, and Puerto Rico. | Open, Classified |  |  | 87503
Z-0886-2016 | 73181 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Auto Endo5 Hem-o-lok Automatic Clip Applier, Medium Large Polymer Clips, Rx Only, Sterile, REF 543965 |  |  | February 02, 2016 | 2016 | Teleflex Medical | The clip in the applier may be missing a boss. A boss is the part of the clip that is used to hold the clip within the applier jaws prior to application. | Nonconforming Material/Component | Teleflex sent an Urgent Medical Device Recall Notification letter dated February 2, 2016 to customers. The letter identified the affected product, problem and actions to be taken. A separate letter was sent to distributors advising them to conduct a sub-recall. Customers were instructed to: 1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed in the letter.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1- 855-419-8507, Attn: Customer Service or email to recalls@teleflex.com.This will document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will document receipt of the letter.For questions contact your local sales representative or Customer Service at 1-866-246-6990 |  | 543 | 543 ea. | Worldwide Distribution in the state of CA, MI, NC, PA, TX, VA, and the countries of Canada, Korea &amp; Belgium | Terminated | August 11, 2017 | 556 | 543
Z-0887-2016 | 73077 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Naturalyte Liquid Acid Concentrate |  |  | January 18, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | On the label calcium (Ca) listed in the right upper corner of the label (red background) incorrectly indicates the product contains 2.5 mEq/L Ca. The concentration indicated in the list of constituents represents the actual calcium concentration of the final dialysate, 3.0 mEq/L. | Error in labeling | The firm, Fresenius Medical Care, issued an "URGENT-MEDICAL DEVICE RECALL" letter dated 1/14/2015 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your stock for the product; if found, discontinue use immediately; place all units in a secure, segregated area; contact your FMCNA Customer Service Team at 1-800-323-5188 for instructions how to return the affected lot (lot# 15STAC072), and promptly complete and return the attached reply form by Fax to: 781-699-9796 Attn: Quality Department or scan and email to: NOTIFYRA@fmc-na.com. Fresenius will pick up and replace your product.If you have any additional medical concerns or questions, please contact Medical Information and Communication at 855-816-2309 or Website: www.fresenius-medinfo.com.Further information or support concerning this issue, contact the Customer Service Care Team at 800-323-5188 and reference the and Support Field Action Number. |  | 848 | 848 (212 cases) | US Distribution to states of: KS, LA, MS, OK, and TX. | Terminated | August 14, 2017 | 574 | 848
Z-1018-2016 | 73126 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | CDI H/S Cuvette, non-sterile, Manufactured by Terumo Cardiovascular Group, Perfusion Pack or Smart Perfusion Pack. |  |  | January 19, 2016 | 2016 | Sorin Group USA, Inc. | The CDI System 500 monitor displays the "H/S DISCONNECT AT CUVETTE" error message when the CDI HIS Cuvette does not make a proper connection with the CDI H/S Probe. When this occurs, blood parameter values for HCT, Hgb and S02 do not display. | Component design/selection | A Safety Alert letter dated January 13, 2015 via registered mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to check the communication of the Cuvette prior to initiating Cardiopulmonary Bypass (CPS).The Terumo Safety Alert did not request removal or return of the product. The instructions that Sarin has communicated to the customers is identical to the information from Terumo. Acknowledgement of the letter and understanding ofthe contents is required for all customer responses. |  | 6859 | 6859 | US Nationwide Distribution and Canada. | Open, Classified |  |  | 6859
Z-0888-2016 | 73307 | No | OTE |  |  | Radiology | PMA | Mammomat Inspiration with Tomosynthesis functionality |  |  | February 12, 2016 | 2016 | Siemens Medical Solutions USA, Inc | A potential system fatal error may occur during tomosynthesis reconstruction with large breasts that cover nearly the entire detector surface and with a breast thickness larger than 90 mm. The data to be processed for the tomosynthesis application can exceed the capacity of the graphical processing unit (GPU). If the fatal error message occurs, reconstruction of the data will be aborted. The syste | Software design | Siemens sent a Field Safety Notice dated February 11, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.How will the issue finally be resolved?Siemens is releasing a field modification to resolve this issue as soon as possible. The upgrade software (VB30P) will be available to the affected customers free of charge. A Siemens service engineer will contact you to schedule implementation of the modification for your system. According to regulations set forth in 21 CFR 7.49(d), we advise that you follow the instructions in this notice and where necessary request that you promptly notify and instruct all the staff at your organization who have to be aware of this problem.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed.We thank you for your cooperation. For further questions, please call (610) 448-6478. |  | 76 | 76 units | Nationwide Distribution including Puerto Rico and to the states of : AR, CA, CO, FL, GA, IA, IL, IN, KS, MA, MD, MI, MO, NC, NE, NH, NJ, NY, OH, PA, SC, TN, TX, VA and WI. | Open, Classified |  |  | 76
Z-0885-2016 | 73258 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | HipLOC CHS Plate/Lag Screw Introducer |  |  | February 16, 2016 | 2016 | Biomet, Inc. | A complaint was received where the Hiploc Plate/Lag Screw Introducer was undersized and would not accept the introducor coupling screw. It appears now that not all parts manufactured to Rev A drawing were reworked. | Process change control | Biomet sent an " URGENT MEDICAL DEVICE RECALL NOTICE" dated February 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The following actions are REQUIRED:Immediately locate and remove the identified device(s) listed below from circulation.Remove the affected product from circulation. Review this notification and ensure that all affected personnel are aware of its contents. Carefully follow the instructions on the enclosed Response Form .Email a copy of the response form to CPWARFieldAction@zimmerbiomet.com prior to return of product.You do not need to complete the Biomet UK response form. Use priority carrier for your shipment.Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA:"Online: www.fda.gov/Safety/MedWatch/HowToReport/default.htm (form available to fax or mail), or"Call (800)FDA-1088Thank you in advance for your assistance and prompt attention. On behalf of Zimmer Biomet, I apologize for any inconvenience this may cause. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 23 | 23 units | US Distribution to the states of : MO, UT, AZ, IL, PA, KY, MI, NC, NJ, MN and GA., and Internationally to The Netherlands. | Open, Classified |  |  | 23
Z-0827-2016 | 73241 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Evita Infinity V500 Ventilator; Babylog Infinity VN500 Ventilator. |  |  | February 04, 2016 | 2016 | Draeger Medical, Inc. | The battery capacity of the optional PS500 power supply unit for the Infinity (ACS) Workstation Critical and Neonatal Care was rapidly and unexpectedly reduced, despite using the most recent power supply firmware version 1.50 released in November 2015 (recall Z-0436-2016). The batteries were prematurely discharged, even though an adequate battery charge status was displayed on the device. | Software design | Draeger Medical, Inc. will send an Urgent Medical Device Recall Letter to consignees in a method that confirms delivery. For consignees that have received new devices with firmware version 1.50 (127 devices) or through device updates (644 devices; reference Recall Z-0436-2016), those devices will be downgraded to firmware version 1.49 and the batteries replaced. Batteries will continue to be replaced for all consignees every six months until a final solution is available. |  | 2553 | 2553 | US Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 2553
Z-0884-2016 | 73156 | No | MAY | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker, T2 K-Wire, 3.2x400mm, CoCr, Recon |  |  | January 21, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | During a review of the packaging, it was determined that the seal integrity of the pouch may be decreased. | Under Investigation by firm | Stryker Orthopaedics sent an Urgent Medical Device Removal letter and Customer Response forms dated January 21, 2016 through UPS. The letter identified the affected product, problem, actions to be taken including forwarding notice if applicable, returning all affected products and completing and returning the Customer Response form to Stericycle.by the instructions provided. |  | 147838 | 147,838 units in total | US Nationwide Distribution | Terminated | June 08, 2017 | 504 | 147838
Z-0966-2016 | 73024 | No | PIF | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EndoVive 3s Low Profile Balloon Kits |  |  | December 23, 2015 | 2015 | Xeridiem Mediem Medical Devices Inc | A feeding (reflux) valve leakage issue. The feeding valve appears to close inconsistently, or close in a delayed manner, when an extension set is disconnected from the Low Profile Balloon device, thus allowing feeding solution or gastric contents to leak out of the device in some instances. | Device Design | Xeridiem sent a customer notification letter dated December 23, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customer was provided the following instructions:1. Segregate all devices from the above table that in Boston Scientific Corporation's inventory and restrict them from being distributed to end users.2. Determine the list of end users who have received devices from the above list, notify them of the quality issue, and have these devices returned to Boston Scientific Corporation. Provide Xeridiem Medical Devices with this list of end users.For any questions the customers is instructed to contact the firm.For questions regarding this recall call 520-882-7794. |  | 38 | 38 | Nationwide distribution to MA only. | Open, Classified |  |  | 7171
Z-1024-2016 | 73333 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SENSE Body Coil 1.5T with identification; used in the ACS NT, Intera 1.5T and Achieva 1.5T MR system |  |  | July 06, 2015 | 2015 | Philips Electronics North America Corporation | Wrong positioning of the coil cables can create unintended resonances via the patient or elevated coil cable temperature due to coupling to the QBC. The combined use of the Sense coil, high SAR level scanning and placement of the cables contrary to instructions for use can lead to patient burns. | Under Investigation by firm | On 6/18/2015, the Field Safety Notice / Medical Device Correction Notice was sent to customers via certified mailing. The letter explains the issue and provides user instructions. A sleeve is being developed that will be applied to the cable on the Sense Body Coil 1.5T and provide consistent spacing between the cable and the patient, thereby reducing the chance of RF interaction and heating. This sleeve will be distributed as part of a Field Change Order with reference FCO781 00437 for Sense Body Coils in the installed base and will be installed by the Philips Field Service Engineers. |  | 5672 | 5672 | US: NationwideOUS: Canada, AfghanistanAlbaniaAlgeriaAndorraArgentinaAustraliaAustriaAzerbaijanBahamasBangladeshBelarusBelgiumBoliviaBosnia and HerzegovinaBotswanaBrazilBurkina FasoChileChinaColombiaColumbiaCongo, DemocraticCosta RicaCroatiaCyprusCzech RepublicDenmarkDominican RepublicEcuadorEgyptEstoniaEthiopiaFinlandFranceGeorgiaGermanyGhanaGreeceGuadeloupeHong KongHungaryIcelandIndiaIndonesiaIranIraqIrelandIsraelItalyJapanJerseyJordanKazakhstanKenyaKorea, Republic ofKuwaitKyrgyzstanLatviaLebanonLibyaLithuaniaLuxembourgMalaysiaMaltaMartiniqueMauritiusMexicoMonacoMongoliaMoroccoMozambiqueNepalNetherlandsNew CaledoniaNew ZealandNorwayOmanPakistanPalestinian territory, occupiedPanamaPapua New GuineaPeruPhilippinesPolandPortugalPuerto RicoRomaniaRussian FederationRwandaSaudi ArabiaSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSri LankaSudanSwedenSwitzerlandSyrian Arab RepublicTaiwanTanzaniaThailandTunisiaTurkeyUgandaUkraineUnited Arab EmiratesUnited KingdomVenezuelaViet NamYemenZimbabwe | Open, Classified |  |  | 5672
Z-1022-2016 | 73010 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT &amp; BrightView XCT Upgrade, gamma cameras; For Single Proton Emission Computed Tomography. |  |  | December 17, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm discovered four issues with the BrightView Family systems.Issue 1: During a cardiac non-Auto Body Contouring (ABC) scan, using Cardiac High Resolution (CHR) collimators, the collimator may come in contact with the bottom edge of the patient pallet. If this occurs, collision sensors would pause the scan.Issue 2:During a scan, the gantry motion may halt prematurely. This cause | Software design | On 12/22/2015, the firm sent Customer Information Letters to their consignees.The letter identified the product, problem and actions to be taken by the customers. |  | 427 | 427 units | Worldwide Distribution-US (nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, VT, WA, WI, WV &amp; WY, and in DC and PR and in the countries of Algeria, Argentina, Armenia, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Cayman Islands, China, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Gabon, Germany, Greece, India, Indonesia, Iran, Ireland, Italy, Japan, Korea, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Palestinian, Panama, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom &amp; Viet Nam. | Open, Classified |  |  | 1221
Z-1019-2016 | 73265 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Virage OCT Spinal Fixation System. Intended to promote fusion of the Occipital-Cervical-Thoracic spine (Occiput-T3). |  |  | February 03, 2016 | 2016 | Zimmer Spine, Inc. | Zimmer Biomet Spine is initiating a Medical Device Correction regarding the Virage OCT Spinal Fixation System by providing additional guidance in the Virage OCT Surgical Technique Manual. | Under Investigation by firm | Distributors, Hospital Risk Managers and Surgeons were notified via an Urgent Medical Device Correction letter beginning 2/3/2016. The letter identified affected product, described the issue and listed their responsibilities. A Certificate of Acknowledgement form was asked to be returned to Zimmer Biomet via email to corporatequality.postmarket@zimmerbiomet.com. Questions or concerns can be directed to Customer Service at 1-866-774-6368 |  | 15753 | 15,753 | Distributed US (nationwide) and in the countries of Australia and Germany. | Open, Classified |  |  | 15753
Z-1023-2016 | 73261 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Custom Perfusion Tubing Pack, Rx, Sterile EO. |  |  | February 04, 2016 | 2016 | Medtronic Perfusion Systems | Medtronic is notifying customers who have or may receive specific lots of Medtronic Perfusion Tubing Packs that include the potentially affected cuvettes manufactured by Terumo Cardiovascular Systems (Terumo CVS). Terumo Cardiovascular Systems (Terumo CVS) has distributed an Urgent Safety Alert regarding the CD IÂż H/S Cuvette. The cuvette is packaged into Medtronic Perfusion Tubing Packs and | Component design/selection | Medtronic sent an Urgent Medical Device Safety Alert dated February 2016 to all consignees on February 4, 2016. The letter described the problem and the product involved in the recall. Advised consignees to review the Terumo CVS Urgent Safety Alert and follow the instructions in the Recommended Action Section of the Urgent Safety Alert letter. Requested consignees to completed the Customer Confirmation Certificate and Email completed form(s) to RS.CFQFCA@Medtronic.com. If email is not available, fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. For questions contact your Medtronic representative or Lifeline Technical Support at 877-526-7890.For questions regarding this recall call 763-391-9000. |  | 9890 | 9890 (7773 US, 2117 OUS) | Worldwide Distribution - US including US: AK, AZ, CA, DE, FL, HI, IL, LA, MD. MA. MN, MS, MO, NE, NJ,. NHY, NC, ND, OH, OR, PAM TN, TX, VA, WA, WV, and WI.Internationally to Australia, Belgium, Canada, Czech Republic, Denmark, Dominican Republic, France, Germany, Israel, Japan, Netherlands, Serbia, South Africa, and Switzerland. | Open, Classified |  |  | 9890
Z-1025-2016 | 73159 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Karma Flexx Wheelchair |  |  | January 27, 2016 | 2016 | Karman Healthcare Inc | Karman Healthcare, Inc. is recalling Karma Flexx Wheelchair because incorrect size screws and corresponding hardware were installed on the front caster fork brackets. | Employee error | On 01/27/16 the firm sent out customer notification letters.In the letter customers were informed that the firm was conducting a product update for original purchasers of KM8522. Customers were provided the following procedures for the product update:1. Completed and signed notification must be returned to Karman Healthcare via fax at (626)-581-2335, or e-mail to edward@karmanhealthcare.com.2. Following the received frame, customers may install their current wheels, and footrest onto the newly updated frame. Kindly package the used frame into the carton in the same manner the new carton was shipped.3. A prepaid return shipping label is included with the letter, which customers may use to return the carton; or contact the firm's customer service line at (800)-805-2762 to have a trained representative schedule a pickup at no charge. |  | 35 | 35 | U.S. distribution to the following states -- WA, CA, NY, and TX. | Open, Classified |  |  | 35
Z-1027-2016 | 72976 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | NeoCoil 3.0T GEM Flex Coil. |  |  | December 18, 2015 | 2015 | NeoCoil, LLC | Potential for higher than specified surface temperatures. Use of 3.0T GEM Flex system may result in patient warming or burns. | Component change control | Distributor was sent a NeoColi "Urgent Medical Device Recall" letter dated 18 December 2015. The letter described the problem and the product involved in the recall. Consignee was advised to quarantine the product and to identify their customers and provide a list to NeoCoil. NeoCoil will contact each customer for correction and removal of the devices. Requested consignee to complete and return the response form. For questions call Michael Leigh at 261-522-6127. |  | 2121 | 2121 | US Nationwide Distribution | Open, Classified |  |  | 2121
Z-1026-2016 | 73308 | No | EBC | Pit and fissure sealant and conditioner. | Dental | Dental | 510(k) | Helioseal Clear Refill 1 x 1.25 g, REF/Product Code/Art. No. 558521AN and Helioseal Clear Assortment 5 x 1.25 g, REF/Product Code/Art. No. 558520AN ---- COMMON/USUAL NAME: Pit &amp; Fissure Sealant |  |  | February 12, 2016 | 2016 | Ivoclar Vivadent, Inc. | Complaints were received claiming the material failed to polymerize under the curing light. After investigation, it was determined that a wrong photo-initiator was used in the production batch U30755 of Helioseal Clear. This failure can lead to incomplete curing of the sealant, if mono-wave LED curing lights are used. | Mixed-up of materials/components | Ivoclar Vivadent sent an Urgent Medical Device Recall letter dated February 11, 2016, to all consignees via Certified Mail, Return Receipt Requested. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to complete, sign and return the attached acknowledgement form and return it to Ivoclar Vivadent via fax (716) 691-2294 or e-mail at recall.us@ivoclarvivadent.com. Contact Ivoclar Vivadent Customer Service at 1-800-533-6825 as soon as possible between the hours of 8:00 am and 6:00 pm EST Monday through Friday and a representative will arrange to pick up your affected materials for return and credit. Please notify your customers of this issue and return the product to us. For questions, please contact Ivoclar Vivadent Customer Service at 800-533-6825. |  | 406 | Domestic: 406 units | Nationwide distribution including NY, PA, FL, IN, IA, TX, and CA. | Terminated | May 08, 2017 | 451 | 406
Z-1028-2016 | 73083 | No | MIJ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Kopans Breast Lesion Localization Needle |  |  | January 08, 2016 | 2016 | Cook Inc. | Product potentially exposed to body fluids during processing. Potential adverse events that may occur as a result of the foreign matter include blood borne pathogen transmission. However, the risk of transmission of blood borne pathogens is thought to be minimal to non-existent since these devices were sterilized after the potential exposure to body fluids. | Process control | Cook Medical sent an URGENT: MEDICAL DEVICE RECALL dated January 8, 2016, to all affected customers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to do the following:Please review the attached list of affected products and lot numbers shipped to your account, and quarantine any affectedproduct that remains unused.2. Immediately collect and return all unused affected products to Cook Medical as soon as possible for credit.3. Please complete the attached Recall Response Form and return to Cook Medical either with the product or separately.4. Please report any Adverse Event to Cook Medical Customer Relations 1-800- 457-4500 or 1-812-339-2235. Mondaythrough Friday between 7:30 a.m. and 5:00 p.m. Eastern Daylight Time or email at FieldActionsNA@Cookmedical.comShould you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235. |  | 134 | 134 | Nationwide distribution including AZ, CA, GA, IL, MO, MS, MT, NC, NH, NY, OH, OK, and PA.. | Terminated | March 23, 2017 | 440 | 134
Z-1030-2016 | 73080 | No | JZG | Beta-2-microglobulin immunological test system. | Immunology | Immunology | 510(K) Exempt | Dimension Vista System |  |  | November 03, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed an increase in the rate of Abnormal Assay errors and calibration failures with the Dimension Vista B2MIC Flex reagent cartridge: Dimension Vista System Flex reagent cartridge and urine stabilizer B2MIC. In the unlikely event that a patient result is generated, a suppressed result would be in conflict with the patients clinical status and other markers of renal injury. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall dated November 2015 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use and discard your remaining inventory of the affected product. Indicate your replacement product needs on the attached Field Correction Effectiveness Check Form and return. For questions call Siemens Customer Care Center- Technical Solutions at 800-441-9250 or your local Siemens technical support representative. |  | 1637 | 1637 units. | Worldwide Distribution and US in the state of CA. | Terminated | January 13, 2017 | 437 | 1637
Z-1031-2016 | 73201 | No | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoris Multicompartmental Knee (MCK) System |  |  | January 15, 2016 | 2016 | Mako Surgical Corporation | Reports where the Mylar/Tyvek pouch for RESTORIS MCK Baseplate in the clamshell packaging configuration exhibited damage, compromising the sterile barrier. | Unknown/Undetermined by firm | The firm, Stryker Orthopaedics, sent an "Urgent Product Recall" letter dated 1/15/2016 to all their consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to inspect the device packaging as the damage to the sterile barrier may be notice prior to the procedure; do not use the device if the seal or package is breached and complete and return the attached Product Recall Acknowledgment Form within 5 days by e-mail StrykerOrtho5317@Stericycle.com or fax at: 888-480-2853. Return all affected products available at your location to: Stryker c/o Stericycle, 2670 Executive Dr., Suite A, Indianapolis, IN 46241, Attn. RA2015-047-Event 5317.If you have any questions, feel free to contact Sr. Director QA at 954-628-0502. |  | 146 | 146 | Worldwide Distribution: US (nationwide) including states of: AL, AZ, CA, CO, CT, FL, GA, IA, ID, IL, LA, MI, MS, MO, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA, WI, and WY ; and countries of: Italy, Scotland, Turkey, Greece, Germany, and Singapore. | Open, Classified |  |  | 146
Z-1029-2016 | 73376 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | ClosureFast Micro Introducer Sheath Set 7Fr/Ch(2.3mm) x 7 cm REF MIS-F707 and ClosureFast Introducer Sheath Set 7Fr/Ch(2.3mm) x 11 cm REF IS-7F11. |  |  | February 17, 2016 | 2016 | Argon Medical Devices, Inc | The incorrect sheath was assembled to the device so that the size 7 products received 11 cm sheaths and the size 11 product received 7 cm sheaths. | Process control | On February 17, a letter was sent to Medtronic notifying them of a voluntary recall. |  | 2000 | 2000 units | Distributed in IL. | Open, Classified |  |  | 2000
Z-0829-2016 | 72824 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | 9131-001 Defibrillation Electrodes, Lot No. 141125-02 |  |  | January 14, 2016 | 2016 | Cardiac Science Corporation | Defibrillation electrodes may increase electrical impedance over time. If impedance becomes too high, the AED will fail the electrode self-test and it will not be Rescue Ready, the AED STATUS INDICATOR will be RED and the AED will beep. This is normal behavior when the impedance limit is exceeded. The high impedance electrodes require replacement. | Nonconforming Material/Component | Cardiac Science sent an Urgent: Voluntary Medical Device Recall letter dated January 2016 starting on January 14, 2014. The letters were either addressed to Distributor or Customer. The letter described the problem and the product involved in the recall. Requested consignees to replace any electrodes from the lot 141125-02 and return them to Cardiac Science. Furthermore, it advised consignees to contact their customers or members of their facilities who need to be made aware by providing a copy of the notification letter and replacement electrodes if they were a distributor. For questions contact Cardiac Science Technical Support at 1-800-426-0337, option 1 in USA or +1-262-953-3500 OUS or via e-mail a9131cardiacscience.com. For questions regarding this recall call 800-426-0337. |  | 10943 | 10,943 | Worldwide Distribution - US all states including PR except ND &amp; RI.International Distribution to Argentina, Australia, Belgium, Bermuda, Canada, Chile, Colombia, Ecuador, France, Germany, Hong Kong, Iceland, Ireland , Israel, Lebanon, Malaysia, Malta, Norway, Oman, Philippines, Poland, Qatar, Romania, Russian Federation, Singapore, Slovenia, Sweden, Turkey, United Kingdom. | Terminated | December 21, 2016 | 342 | 10943
Z-1033-2016 | 73115 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Covidien Kangaroo Polyurethane Feeding Tube Radiopaque Line, Safe Enteral Connections 6.5 Fr/Ch (2.2 mm) x 36" (91 cm), Sterile; Catalog Number: 461438. |  |  | January 21, 2016 | 2016 | Covidien LLC | 8FR 16 PVC tubes were packaged in a 6.5 FR 36 Polyurethane packs and 8FR 20 PVC tubes were packaged in an 8FR 20 Polyurethane packs | Labeling Change Control | Medtronic/Covidien Customers notified via an Urgent Medical Device Recall letter, dated January 22, 2016. The letter identifies the affected product, instructs customers to discontinue use and quarantine any of the product, and return all affected inventory. Customers are also to complete the "Recalled Product Return Form." Questions or concerns can be directed to a Medtronic representative or Customer Service at 800-882-5878. |  | 1010 | 1,010 units | Worldwide Distribution -- US, to the state of CA; and, to Australia and Europe. | Terminated | June 01, 2017 | 497 | 2340
Z-1036-2016 | 73237 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | PICC/CVC Secure Dressing Change kit, Ref: 73775, Lot 213403, Sterile Single Use, Medical Action Industries, Inc., Arden, NC. |  |  | February 03, 2016 | 2016 | Medical Action Industries Inc | Misbranding; Due to the presence of CHG in the Tegaderm PICC/CVC dressing change convenience kit, which is not identified on the exterior labeling | Under Investigation by firm | Medical Action Industries, Inc. sent an Urgent Medical Device Recall letter dated February 3, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to immediately examine their inventory and quarantine all product subject to this recall. Consignees were also instructed if product was further distributed to identify their consignees and notify them at once of this product recall. Recalled product should be returned to Medical Action Industries for credit or replacement by contacting quality@owens-minor.com . Consignees were asked to complete and return the enclosed response form. Consignees with questions were instructed to call 828-338-7630. |  | 9 | 9 cases (180 kits) | Nationwide distribution to MO, NE and TX . | Open, Classified |  |  | 9
Z-1032-2016 | 73225 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | PINPOINT Endoscopic Fluorescence Imaging System |  |  | January 29, 2016 | 2016 | Novadaq Technologies Inc. | It has been determined that the PINPOINT Operator's Manual does not contain validated parameters for gravity-cycle autoclave sterilization of the PINPOINT laparoscopes and light guide cable in the reprocessing tray. | Process design | Novadaq sent an Urgent Device Correction letter dated February 2, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Recipients of this notice have received shipments that include the inaccurate Operator's Manual. Please follow these instructions:1. Discard any old PINPOINT Operator's Manuals Rev J through M.2. Accept and review the new PINPOINT Operator's Manual (Rev N).3. Complete and sign the enclosed "Acknowledgement of Receipt" form, scan it and email to PI NPOINT@novadaq.com. Any questions or concerns may also be sent to this address. For further questions, please call 1-(844) 668-2327, ( Select Option # 1) |  | 110 | 110 units | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the countries of : Bangladesh, Canada, Hong Kong, India, and Israel. | Open, Classified |  |  | 110
Z-1041-2016 | 73254 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Volcano R100 Sterile Equipment Cover; Sterile; |  |  | February 05, 2016 | 2016 | Volcano Corporation | Volcano Corporation has identified that certain lot numbers of Volcano R-100 Sterile Equipment Covers may have the incorrect expiry date on the outer shipping box; individually wrapped inner product bears the correct expiry date. | Labeling False and Misleading | Urgent Medical Device letters were sent to affected consignees the week of February 8, 2016 by Fed Ex or UPS so delivery can be tracked. For OUS consignees, the same letter will be sent and they will execute delivery to customers and track responses. |  | 1355 | 1355 total | Distributed US (nationwide) and the countries of Australia, South Africa, Canada, New Zealand, Belgium, Japan, and Paraguay. | Terminated | March 21, 2017 | 410 | 1355
Z-1045-2016 | 73212 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet Getinge Group 1180.19XO Extension Device. Manufacturer MAQUET GmbH, |  |  | January 07, 2016 | 2016 | Maquet Cardiovascular Us Sales, Llc | MAQUET Medical Systems USA is initiating a field action for correction of the MAQUET Extension Device (Part No 1180.19F0). The Extension Device is manufactured by Maquet GmbH, in Germany, and is a table top module for use with the Maquet MAGNUS operating table system. Maquet GmbH was made aware of a potential issue that can result in the Extension Device becoming unlocked from the table top, causing it to fall down. | Under Investigation by firm | Maquet sent "URGENT PRODUCT RECALL MEDICAL DEVICE CORRECTION" letters/MEDICAL DEVICE CORRECTION RESPONSE FORM dated January 7, 2016, via UPS to the affected customers. The notification informed customers of the issue with the affected product; how to identify affected product; measures to be taken by the users; and that a MAQUET service representative will be contacting them to schedule servicing of the device to install a Support Frame Lock. Customers with questions were asked to contact their local MAQUET Representative or call MAQUET Customer Service at 1-888-627-8383 (Option 2, Option 3) Monday through Friday between 8:00 am and 5:00 pm EDT.Customers were also requested to fill out and return the attached response forms. |  | 249 | 2 within US; 247 OUS | Worldwide Distribution - US including AZ and MO.Foreign distribution to Australia, Austria, Bahrain, Belgium, Brazil, Canada, Denmark, Ecuador, Estonia, Finland, France, Germany, Hong Kong, India, Italy, Japan, Latvia, Malaysia, Netherlands, Norway, Poland, Portugal, Russia, Saudi Arabia, Serbia, South Korea, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | May 25, 2017 | 504 | 249
Z-1043-2016 | 72853 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Columbus Revision Knee System, EnduRo Knee System |  |  | November 30, 2015 | 2015 | Aesculap, Inc. | Aesculap Inc. US has initiated a recall on Tibial and Femur extension sterile pressfit stem implant packaging that are used for Columbus Revision/Enduro Knee Implants because it is labeled as Cementless. US product is indicated in the US for use with Bone Cement only. | Packaging change control | Aesculap sent an Urgent Medical Device Correction letter to Distributors and Sales Rep. The letter identified the affected product, problem and actions to be taken. Customers were notified to add the over label to product packaging. Telephone conferences with Distributors and Sales Rep reiterated need for representative to inform Surgical Handler of the label when providing product for review. Customers were instructed to complete the attached Product Correction Acknowledgement Form and return to Aesculap Quality Assurance department by faxing the form to 610-791-6882 or e-mail to val.strawn@aesculap.com, two (2) weeks of receipt, even if the total inventory in your possession is zero (0). |  | 1232 | 1232 | US Nationwide Distribution | Open, Classified |  |  | 1232
Z-1044-2016 | 72823 | No | LCQ | Gentamicin test system. | Toxicology | Toxicology | 510(k) | Nebulizer Masks Opti-Mist and Opti-Mist Clear |  |  | November 11, 2015 | 2015 | ConvaTec, Inc | The Nebulizer mask malfunctions in that there is reduced or no aerosol output. The malfunction occurs from the reduction in the diameter of a component part inside the nebulizer. | Nonconforming Material/Component | The firm, ConvaTec, sent each consignee affected by the recall an overnight express mail "Urgent Voluntary Recall Notice" and Response Form dated 12/7/2015. The notice described the product, problem and actions to be taken. The recall notice instructed the consignee to immediately stop distributing and quarantine all affected lots; perform a count of all recalled product currently in inventory; complete and return the enclosed Recall Response Form by email at convatec3657@stericyle.com or by fax at 1-888-603-8986, and return all affected product, (ONLY return the lots affected with this recall). If you want replacement product contact customer service by phone at 1-800-634-6003, extension 1103 or by email at respiratoryorders@convatec.com.If you have any questions call 1-877-650-3492. |  | 3605700 | 3,605,700 pieces | Worldwide Distribution: US (nationwide) including states of: AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, ME, MA, MN, MS, MO, NV, NH, NJ, NY, NC, OH, OR, PA, SC, TX VA, WA, WI, and countries of: Australia, Belgium, Burma, Canada, Denmark, Egypt, Estonia, Faroe Islands, Finland, France, Germany, Israel, Italy, Jordan, Kenya, Kuwait, Netherlands, Norway, Russia, Saudi Arabia, Sweden, Switzerland, United Kingdom &amp; Vietnam. | Terminated | July 18, 2016 | 250 | 3605700
Z-1040-2016 | 73172 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 50001 (Inrad) AccuPlace Drace Stereotaxic Needle Guide; Assists the user with inserting a needle at a predetermined angle for computer tomography. |  |  | February 04, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | During an evaluation of the product packaging, it was determined that there is a potential for the product to cause damage to the packaging, creating a possible breach in sterility. | Package design/selection | Stryker sent URGENT MEDICAL DEVICE RECALL NOTIFICATIONS to Customers via Overnight service. Stryker Sales Representatives were notified via email.1. Immediately review this Recall Notification.2. Check all stock areas and/or operating room storage to determine how manyAccuPlace" Needle Guides are at your facility. Discontinue use of theAccuPlace" Needle Guides.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of thisnotification and identify how many, if any, affected items are currently in yourinventory. Please complete and return the BRF even if you dont have anyaffected product on hand.Note: Your signature on the BRF indicates that you received andunderstand this Notification and have followed the instructions in theNotification.4. If you have further distributed this product, please forward this letter and theattached Business Reply Form (BRF) to all affected locations. Pleaseindicate each location on the BRF.5. Fax the completed BRF to Stryker Instruments Regulatory Department at866-521-2762, or scan and email a copy to kara.spath@stryker.com.6. If the BRF for your facility indicates that recalled product is currently onhand, a FedEx label will be emailed to you. This shipping label should beused to return recalled product. Upon receipt of the recalled product, a credit will be issued to your account.In addition, Stryker Instruments sent recall notifications directly to facilities that received recalled product(s) originally shipped by Inrad. Facilities that received products from both Stryker and Inrad will receive one recall notification. The Business Reply Form will be pre-filled to indicate specifically what recalled products were sent to each facility, including the quantity and Product Number(s). |  | 5800 | 5,800 units total | Nationwide Distribution-including the states of AZ, CA, CO, CT, FL, GA, IA, KS, MA, MI, MN, MO, NY, OH, OK, PA, TX, VA, WA, and WI. | Terminated | May 09, 2017 | 460 | 5800
Z-1046-2016 | 73298 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | M/L Taper Hip Prosthesis (Item: 00-7711-012-20, Lot: 62937093 |  |  | February 11, 2016 | 2016 | Zimmer Biomet, Inc. | Zimmer Biomet initiated a lot-specific voluntary recall on 02/11/2016, of the M/L Taper Hip Prosthesis. single manufacturing lot of Size 12.5, Extended Offset, M/L Taper hip stems was incorrectly etched. The device was etched as STD, rather than EXT. The product packaging is labeled correctly as Extended.Offset, and the device is an extended offset stem. | Under Investigation by firm | Zimmer, Inc. initiated a lot-specific voluntary recall on 02/11/2016, of the M/L Taper Hip Prosthesis. Electronic notices were e-mailed to distributors, and letters were mailed via certified mail on 02/11/2016 to distributors and hospitals. A field complaint investigation confirmed that this single manufacturing lot of Size 12.5, Extended Offset, M/L Taper hip stems was incorrectly etched. The device was etched as STD, rather than EXT. The product packaging is labeled correctly as Extended Offset, and the device is an extended offset stem.Distributors will return on-hand product to Zimmer Biomet and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers and surgeons will be provided with a letter identifying the issue and their responsibilities. These responsibilities include: Assisting the Zimmer Biomet sales representative with the quarantine of the product and completion and return of the Acknowledgement Form. Customers are asked to complete the Certificate of Acknowledgement Form and return to corporatequality.postmarket@zimmerbiomet.com.Customers with have further questions or concerns please call the customer call center at 1-800-348-2759 between 8:00 am and 8:00 pm EST. |  | 9 | 9 units | Distributed in the states of SD, MO, NY, MI, PA, KY and TN and in the country of Argentina. | Open, Classified |  |  | 9
Z-1049-2016 | 73248 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery LS, 2286327 |  |  | February 10, 2016 | 2016 | GE Medical Systems, LLC | GE Healthcare has updated the Service Manuals for a number of CT and PET CT Scanners to update an incorrect torque value. | Labeling design | GE Healthcare sent an Urgent Medical Device Correction letter " GEHC Ref# 25467-ABT or 25467" dated February 12, 2016, to all affected consignees . The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology, &amp; Radiologist/Cardiologist. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Consignees were instructed to ensure that the organization that service their equipment uses the latest version of the Service Manual to maintain the safe and proper performance of their medical device. For questions call in the United States: 1-800-437-1171 or your local GE Healthcare Service Representative. For questions regarding this recall call 262-513-4122. |  | 817 | 817 | Worldwide Distribution - US including all states in including PR except ND, VT &amp; WY.Foreign Distribution: ALBANIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BOSNIA &amp; HERZEGOVINA, BRAZIL, BULGARIA, BURKINA FASO, CAMERON, CANADA, CHILE, CHINA, COLOMBIA, DOMINICAN REPUBLIC, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE,HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRA N, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PERU, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA. | Open, Classified |  |  | 1793
Z-1047-2016 | 73151 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Large Volume Paracentesis Kit is packaged in a thermaformed tray which contains slots and snap-in holders for most components. Packed sterile in Tyvek bag. |  |  | January 28, 2016 | 2016 | Gi Supply | Customers have reported that the pinch clamp on the RP Bag was missing. Ascetic fluid from the abdominal cavity may splash User if bag is not clamped prior to removal. | Nonconforming Material/Component | The customers will be made aware of the defect, including pictures provided and are instructed to inspect the RP Bags in the kit prior to use. If the pinch clamp is missing, the customer is instructed to discard the RP Bag and use the sterile RP Bag replacements provided. |  | 390 | 390 Kits | Distributed in DC and the following states CA, FL, IA, KY, MA, MO, NH, NJ, NY, OH, OR, PA, SC, and TN. | Open, Classified |  |  | 390
Z-1054-2016 | 73329 | No | HDW | Contraceptive diaphragm and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Milex Arcing Diaphragm Size 95; Model Number: MXWS95 |  |  | February 16, 2016 | 2016 | CooperSurgical, Inc. | The products have been identified to contain an incorrect size Diaphragm. | Package design/selection | Cooper Surgical sent an "Urgent - Medical Device Recall" letter dated February 16, 2016, to all affected customers. The letter identified the problem, the product and the action needed to be taken by the customer.Customers were instructed to discontinue use of the affected products and complete the attached Acknowledgement and Receipt form for replacement.For further questions, please call (203) 601-9818. |  | 2 | 2 units | US Distribution to the states of : OR and NY | Terminated | March 29, 2016 | 42 | 2
Z-1066-2016 | 72325 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips X-Ray Systems, Allura Xper with R8.2.16 |  |  | October 01, 2015 | 2015 | Philips Electronics North America Corporation | Upon initiating Fluoroscopy the user may encounter a user message Fluoro failed. | Software design | Philips sent a Field Safety Notice to all affected customers on October 1, 2015. The notice informs customers of the issue and the actions planned by Philips. Philips Field Service Engineers are to install Software R8.2.16.1 which addresses the issue. For further information or support concerning the issue, consignees should contact their local Philips Representative or call the Technical Support Line at 800-722-9377. |  | 196 | 196 devices | US Nationwide in the states of : AK, AL, AZ, CA, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, NC, NJ, NY, OH, OR, PA, TN, TX, UT, WI, WV and the country of Canada. | Open, Classified |  |  | 196
Z-1065-2016 | 73177 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Trilogy 100, Trilogy 200, Trilogy O2, Trilogy 202, |  |  | January 18, 2016 | 2016 | Philips Respironics | Software Issue | Software design | Philips Respironics has determined that affected devices may continue to be used in accordance with the device Instructions for Use and the Field Safety Notification that is being issued to customers. Philips Respironics will be providing a software update for affected devices free of charge. Customers will be notified by mail when the software update is available for download.For further questions please call ( 724 ) 387-7500. |  | 104508 | 104,508 units | Nationwide Distribution including Puerto Rico and Internationally to the US Virgin Islands. | Terminated | May 26, 2017 | 494 | 104508
Z-1064-2016 | 73045 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Super Arrow-Flex Percutaneous Sheath Introducer Sets |  |  | January 12, 2016 | 2016 | Arrow International Inc | Arrow has issued a recall for these products due to potential packaging damage. The sterility of the products cannot be guaranteed. If non-sterile product is used there is a potential for infection to occur. | Packaging process control | Arrow International sent an Urgent Medical Device Recall Notification letter dated January 13, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product. Customers were asked to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507 Attn: Customer Service or email it to recall@telefex.com. A customer service representative will contact them with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to the firm. If customers have no affected stock, they should also complete and the enclosed Recall Acknowledgement Form. Customers with questions should call Customer Service at 1-866-246-6990. For questions regarding this recall call 610-378-0131. |  | 56521 | 56,521 units | Nationwide Distribution including ALARAZCACOCTDEFLGAIAIDILINKYLAMAMDMIMNMOMSMTNCNDNENHNJNMNVNYOHOKORPARISCTNTXUTVAVTWAWIWV | Open, Classified |  |  | 56521
Z-0911-2016 | 73243 | No | NKM |  |  | Cardiovascular | PMA | MitraClip Clip Delivery System, product number MSK0101. The MitraClip System |  |  | February 04, 2016 | 2016 | Abbott Vascular | Abbott Vascular has recently received reports of cases on Clip Delivery System devices that contain the One-Way Actuator Knob where a user attempted implanting a MitraClip, but the Clip could not be detached from the delivery system due to a mandrel fracture. | Use error | The firm, Abbott Vascular, sent an "URGENT FIELD SAFETY NOTICE/MITRACLIP/CORRECTION" letter dated 2/4/16 to all the customers who received the MitraClip System. The letter described the product, problem and actions to be taken. The Customers are informed that their inventory is acceptable for safe use following the revised IFU steps described in the customer notification letter and that there is no need for the customers to return any product to Abbott Vascular. Customers with questions are instructed to contact their local Abbott Vascular representative.Customers were also instructed complete and return a Training Form/Effectiveness Check via their Abbott Vascular Representative, fax to 1-951-914-5951 or scan and email to AVRegulatoryCompliance@av.abbott.com. The firm is revising the IFU to incorporate the revised procedural steps.A press release dated 2/26/16 will be posted on the firm's website to inform the public that Abbott has initiated a voluntary safety notice regarding the MitraClip Delivery System to reinforce the proper procedures used to operate and deploy the device. Customers with questions or concerns are instructed to contact Abbott Vascular Representative or call Abbott Vascular at (800) 227-9902. |  | 7600 | 7,600 units total (2,300 units in US) | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Argentina Luxembourg Australia Malaysia Austria Mexico Belgium Netherlands Brazil New Zealand Canada Norway Colombia Philippines Corrientes Poland Czech Republic Portugal Denmark Egypt Romania Finland Saudi Arabia France Slovenia Germany Spain Greece Sweden Hong Kong Switzerland Hungary Thailand Indonesia Turkey Israel United Arab Emirates Italy United Kingdom Kuwait and Vietnam. | Terminated | May 12, 2016 | 98 | 7600
Z-1081-2016 | 72935 | No | FNM | Alternating pressure air flotation mattress. | General Hospital | General Hospital | 510(K) Exempt | Stryker Position Pro Mattress. |  |  | December 03, 2015 | 2015 | Stryker Medical Division of Stryker Corporation | Potential for power cord to melt; hazards include electric shock and fire. | Component design/selection | Stryker sent an URGENT MEDICAL DEVICE CORRECTION NOTIFICATION letter dated November 10, 2015 to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any urgent questions or concerns, please, contact Stryker Customer Service at (1-800-327-0770). Normal business hours are Monday-Friday 8 a.m-5 p.m. (ET). |  | 2694 | 2694 | Worldwide Distribution - US Nationwide and the countries of Brazil, Canada, Chile, France, Hong Kong, Italy, Mexico, South Korea, Switzerland, United Kingdom. | Terminated | May 24, 2017 | 538 | 2694
Z-1076-2016 | 73008 | No | PHC | Infusion pump. | General Hospital | General Hospital | 510(k) | BD Intelliport Medication Management System Sensor |  |  | December 09, 2015 | 2015 | Becton Dickinson &amp; Company | The sterility of the product cannot be assured. This may result in increased risk of infection. | Under Investigation by firm | Becton Dickinson representatives notified their customers in person and a copy of the "Urgent Product Recall" letter and "Recall Response Form" dated 12/9/2015 was provided. The letter identified the reason for the recall; how to identify affected product; and the actions to be taken. The letter instructed customers to immediately review their inventory; complete the enclosed Recall Response Form and fax (1-201-847-4267 it to BD or email it to Becky_Saggau@bd.com even if you do not have any of the affected lot; and return all affected products with the completed Recall Response Form following instructions on the enclosed packing instruction. If customers have any questions or require assistance with the return of the recalled product and/or availability of replacement product, they were instructed to contact 1-201-847-4267 between 8AM and 5 PM ET Monday through Friday. |  | 250 | 250 units | US Distribution to:California and Utah. | Terminated | January 09, 2017 | 397 | 250
Z-1082-2016 | 73140 | No | LXT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Crosslock DVR ePAK -Sterile, Ref 81180050, are complete surgical sets including plates, screws, and instruments for orthopedic fixation |  |  | January 29, 2016 | 2016 | Biomet, Inc. | Components of the DVR ePAKs can oxidize during shipment and storage prior to use. A delay in surgery, less than 30 minutes, may occur if the surgical staff proceeds to re-sterilize the product. If oxidative material would flake off a component contained within the ePAK and enter the surgical field/wound an inflammatory reaction may occur. The inflammatory reaction may result in a foreign body r | Package design/selection | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated 1/29/2016 to the affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The customers were instructed to immediately locate and remove the identified device(s); carefully follow instructions on enclosed Response Form; complete and return Response Form within three (3) business days via email to CPWARFieldAction@zimmerbiomet.com prior to return of product, use priority carrier for your shipment; if you have further distributed the product, you MUST notify hospital personnel of this action. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 21268 | 21,268 | Worldwide distribution: US (Nationwide) and countries of: Panama; Korea; Japan; Netherlands; Ecuador; Chile; Canada; and Australia. | Open, Classified |  |  | 21268
Z-1075-2016 | 72979 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Catalog Number 6801375, Unique Device Identifier Number 10758750001132; and VITROS 5,1 FS Chemistry System (Refurbished), Catalog Number 6801890, Unique Device Identifier Number 10758750001644; IVD. |  |  | December 21, 2015 | 2015 | Ortho-Clinical Diagnostics | Increased U90-382 or 6LU condition codes generated by VITROS 250, 350, 5,1 FS, 4600 and 5600 Chemistry Systems when using Calibrator Kit 9, Lot 954. A trend of complaints regarding customers actions following U90-382 or 6LU condition codes was noted (See RES 72289 - VITROS Calibrator Kit 9, lot 954 recall). | Error in labeling | Ortho Clinical Diagnostics sent an URGENT PRODUCT CORRECTION NOTIFICATION Letter (Ref. CL2015-201, dated December 21, 2015) via FedEx overnight courier and/or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address). The notification letter was sent to all customers who received the affected product. Foreign affiliates were informed by e-mail on 12/20/2015 of the issue and instructed to notify their consignees of the issue and required actions. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Technical Services at 1-800-421-3311. |  | 1405 | US = 994; Foreign = 1405 | Worldwide Distribution - US Nationwide including Puerto Rico, and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela | Open, Classified |  |  | 8387
Z-1083-2016 | 73293 | No | EKM | Gutta percha. | Dental | Dental | 510(K) Exempt | TF Adaptive Gutta Percha Point Size SM3 50pk, Part No. 815-1541 |  |  | February 08, 2016 | 2016 | Ormco/Sybronendo | Ormco Corporation is recalling the TF Adaptive Gutta Percha Point Size SM3-50pk because the inner generic label is identified as ML3. The products outer package is identified as SM3 which matches the product inside. | Error in labeling | An Urgent Medical Device Recall letter dated February 8, 2016 was sent to all customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and fax back the enclosed acknowledgement form to 909-962-5605 to confirm receipt of the recall notification . |  | 5568 | 5,568 units total (2,070 units in US) | Worldwide Distribution - US Nationwide and the countries of Canada, Australia, Mexico, Europe, Hong Kong, MY, IN, JP, ID, SG, TH | Open, Classified |  |  | 5568
Z-1069-2016 | 73365 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Denali Filter-Femoral Delivery System |  |  | February 09, 2016 | 2016 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular, Inc. is recalling Denali Filter-Jugular/Subclavian Delivery System and Denali Filter-Femoral Delivery System because of the stop cocks potentially cracking. | Component design/selection | The firm sent customer notification letters on 02/09/16.The following instructions were provided to customers:1. Do not use or further distribute any affected product.2. Check all inventory locations within the institution for affected product code/lot number combination listed in the recall notice. If further distributed any of the product code/lot numbers, immediately contact that location, advise them of the recall, forward these instructions and have them return the affected product to BPV.3. Remove any identified product from your shelves.4. If have used the affected product, complete and return the attached Recall andEffectiveness Check Form indicating no product will be returned.Once the product affected by this recall has been removed from inventory;5. Fill out the Recall and Effectiveness Check Form. Be sure to state the quantities and lot numbers of each recalled product that you intend to return. It is extremely important the information is received.6. Call Recall Coordinator Raye Seisinger at 1-800-321-4254 Option #2 Ext 2501 (M-F 6am to 3pm MST) or email at raye.seisinger@crbard.com. Once all information has been verified, the Recall Coordinator will issue either a Return Authorization (XC) Number or Consignment Recall Number (XH) to facilitate the expedient return of the product. BPV will provide replacement product for returned product.7. E-mail the completed Recall and Effectiveness Check Form to DenaliRecall@crbard.com or fax it to BPV at 1-800-994-6772. If cannot e-mail or FAX the form, please call the BPV recall coordinator (Raye Seisinger, at 1-800-321-4254 Option #2 Ext 2501) and report the required information verbally.8. A mailing label is enclosed for convenience to return the affected product. Please mark the outside package as "RECALLED PRODUCT" and include the XC or XH number. All products should be returned to the following shipping address:Bard Peripheral Vascular, Inc.1415 W. 3'd |  | 3572 | 3572 | Worldwide Distribution-US (nationwide) including DC and PR, and the states of MN, TX, PA, OH, IL, OR, CA, SC, NC, AK, NY, MD, GA, NJ, AZ, KY, OH, WI, AL, AR, MO, FL, MA, MT, TN, LA, CT, NE, WV, VA, NV, ME, UT, CO, WY, AR, DE, NM, IN, MN, MS, NH, IA, MI, WA, VA, KS, HI, RI, SD, VT, and the countries of Hong Kong, Malaysia, Canada, Chile. | Open, Classified |  |  | 6476
Z-1077-2016 | 73263 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics ADVIA Centaur Calibrator E - 2-Pack |  |  | February 04, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | ADVIA Centaur Systems Calibrator E Lot 42 and 43 Bias with the Testosterone Assay | Under Investigation by firm | Siemens Healthcare issued an Urgent Medical Device Correction (UMDC) to all affected Siemens Healthcare Diagnostics customers in the United States on February 4, 2016 to be delivered to customers on February 5, 2016. An Urgent Field Safety Notice (UFSN) was sent out for distribution to customers outside the United States on February 4, 2016. These notices inform the customers of the lot to lot bias of Calibrator E with the ADVIA Centaur Testosterone assay and advise customers to move to Calibrator E kit lots ending in 44.Fax-back forms will be used to confirm that the customers received theUrgent Medical Device Correction. Questions, contact your Siemens Customer Care Center or your local Siemens technical supportrepresentative |  | 15224 | 15224 (2-pack) | Worldwide Distribution US (nationwide) and Internationally to:CanadaAR AU BD BR CA CH CL CN CO DO EC GT HK ID IE IL IN JP KR LK MM MX MY NI NP NZ PE PH PY SA SG TH TW UY VN ZA | Terminated | August 01, 2017 | 544 | 22198
Z-1067-2016 | 73100 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | iChem VELOCITY Automated Urine Chemistry, Catalog No. 800-3049, 800-3050, 800-3079, 800-3080, 800-3530, 800-7167, 800-7190, 800-7713, 800-7714, 800-7719, 800-7720 |  |  | January 13, 2016 | 2016 | Beckman Coulter Inc. | Iris International noted loose or missing analyte pads found in several locations (strip vial, strip provider module, strip conveyor system or waste container) that are undetectable by the iChemVELOCITY System. | Equipment maintenance | Iris sent an Urgent Medical Device Recall letter dated January 13, 2016 and an updated notification letter dated February 19, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions or concerns are instructed to contact their local support representative:Âż Via our website, at http://www.beckmancoulter.comÂż By phone, call 800 854-3633 in the United States and Canada. |  | 2661 | 2,661 units total (1,424 units in US) | Affected geographies: Argentina, Australia, Austria, Bangladesh, Bermuda, Bosnia and Herzegovina, Brazil, Brunei Darussalam, Canada, Chile, China, Columbia, Croatia, Czech Republic, Ecuador, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Jordan, Republic of Korea, Kuwait, Lebanon, Lithuania, Macao, Malaysia, Mexico, Montenegro, Netherlands, New Zealand, Nigeria, Oman, Panama Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico , Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Republic, United States (including Puerto Rico), and Venezuela | Open, Classified |  |  | 2661
Z-1088-2016 | 73303 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic 0-arm 02 Imaging System |  |  | January 29, 2016 | 2016 | Medtronic Navigation, Inc. | Screws holding the detector panel in as a result the detector panel could make contact with other parts with in the 0-armÂż 02 Imaging System, causing noise and vibration when being positioned or during 3D Imaging could lead to poor image quality | Process control | Medtronic Navigation issued a Field Safety Notice dated January 29, 2016, to notify them of a hardware correction in the unit. A Medtronic service representative will service the system and arrange for an inspection of the system. This representative will properly torque the eight (8) screws if they find them to be loose. In the meantime, if noise or vibration are observed during use of the system customers should contact their Medtronic representative to report the issue. Customers with questions were instructed to call 720-890-3160. For questions regarding this recall call 978-698-6000. |  | 1 | 1 unit | Worldwide Distribution to Austria only. | Terminated | March 10, 2016 | 41 | 1
Z-1084-2016 | 73428 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobile wDR motorized portable diagnostic X-ray systems. |  |  | January 18, 2016 | 2016 | Sedecal S.A. | This is related to previous 806 Notifications about uncommanded movements of a motorized mobile x-ray system. Reason for Recall: Along with the previously noted encoder and gauges problems and corresponding fixes, we have observed that static electricity can also elicit the unwanted behavior. | Device Design | Philips Medical Systems sent an URGENT - Field Safety Notice letter dated August 6, 2014 was sent to all consignees of record informing them of a potential safety issue on the SEDECAL mobile units manufactured from May 2013 on. The firm requests the direct acounts with affected units should take the following steps:1) The unit is equipped with a safety switch in the drive handle. If the drive handle isreleased the unit will stop; 2) Do not try to correct unintended movement using the controls on the drive handle; 3) Instead, stop the system immediately by releasing the drive handle entirely; 4) Alternatively, the emergency button can be pressed; 5) Take special care when the system crosses metal strips or plates on the floor or enters an elevator, as these are situations where electrostatic discharge is likely to occur; and, 6) Should the customer feel uncertain regarding this action, he/she can contact Philips Healthcare.Philips is preparing a Field Action Kit to replace the existing arrestor strap by a new arrestor with higher resistivity. This will reduce the electromagnetic pulse to a level, which is safe to not disturb the electronics. A Philips Service Engineer will contact the direct accounts to distribute Field Action Kit.The firm states that should direct accounts need to communicate with Philips in regards to this recall that they reference FC0-71200138.If they need any further information or support concerning this issue, please contact their local Philips representative. |  | 349 | 349 systems | Nationwide Distribution including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NM, NY, OH, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI. | Open, Classified |  |  | 349
Z-1089-2016 | 71148 | No | EBI | Denture relining, repairing, or rebasing resin. | Dental | Dental | 510(k) | IMPAK ACRYLIC RESIN LIQUID FOR BOTH IMPAK &amp; IMPAK-PF POWDERS 1 qt. size (946 cc), Part Number 3306, and 11 oz. size (325 cc), Part Number 3748. |  |  | November 18, 2013 | 2013 | CMP Industries, Llc | During the current inspection of CMP Industries, the FDA Investigator discovered that the firm had conducted a recall of Impak Acrylic Liquid Batch 103113 in 2013 due to potential rust without notifying the district. | Nonconforming Material/Component | On November 18, 2013, CMP Customer Service contacted customers by telephone requesting return of the affected product and customers were issued a UPS Call Tag. CMP issued a letter to non-responding customers dated December 9, 2013, requesting return of the product because of the possibility of foreign particles. The letter states, the firm found that one of their processes could deliver a material that would pass their QC tests, but may introduce foreign material into the liquid. A customer service representative will be in contact with customers to arrange the return of the product. |  | 246 | 246 units | Worldwide Distribution -- USA, to the states of MI, NJ, PA, and TX; and, the country of Canada. | Completed |  |  | 246
Z-1090-2016 | 73357 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | ROTROL P Control for ROTEM delta Thromboelastometry System, For In Vitro Diagnostic Use Only. Quality control material for monitoring accuracy and precision of test carried out on the ROTEM delta Thromboelastometry System. |  |  | February 18, 2016 | 2016 | TEM Systems Inc | Customer complaints about failed target ranges for alpha angle. Firm has not completed the CAPA identifying the root cause. Projected completion of the CAPA by the firm is July 2016. | Nonconforming Material/Component | An Urgent Important Customer Information notice was sent to accounts on 2/17/2016. The notice requested that they discontinue use of the product and to contact them for a replacement. The letter also included a response form to be returned. |  | 481 | 476 boxes x 5 vials | Distributed in the states of AZ, AR, CA, FL, GA, HI, ID, IL, IN, LA, MD, MA, MN, MO, NE, NJ, NM, NY, PA, OH, SC, SD, TN, TX, UT, WA, WV and the country of Germany. | Open, Classified |  |  | 481
Z-1080-2016 | 72961 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Custom Assembly, Custom FCFFM x/8.2 cm Tubing Extension, Model Number: FCA026-001, Manufactured by: Value Plastics dba Nordson MEDICAL, 805 West 71st Street, Loveland, CO 80538, packaged in Bags consisting of 100, 500 or 1,000 stopcocks. |  |  | December 18, 2015 | 2015 | Value Plastics, Inc. | Potential for the parts to crack. | Process control | Customers were initially notified by telephone call; followed by recall notice via email on 12/18/15. The recall letter instructs customers to identify their customers and notify them of the product recall. Returned product will be quarantined and separated from acceptable materials and scrapped. |  | 33000 | 33,000 | 3 Consignees in NY, PA, MN. No foreign, VA/govt. | Open, Classified |  |  | 33000
Z-1092-2016 | 73393 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare Xhibit Central Station, Model 96102 is used to provide clinicians with central monitoring of patient data for those patients connected to networked Spacelabs Healthcare patient monitors and telemetry transmitters. |  |  | February 26, 2016 | 2016 | Spacelabs Healthcare Inc | The firm has received one report of values for patient height and weight being switched when input at the Xhibit Central Station, Model 96102, causing a bedside monitor Body Surface Area (BSA) calculation to be in error. | Software in the Use Environment | Spacelabs sent the letter Urgent Medical Device Correction -Spacelabs Healthcare Inc. - Xhibit Central Station, Model 96102, dated 26 February 2016 to US consignees.The firm will send a notification letter to international consignees on March 4, 2016.Consignees who have any questions about this corrective action, please contact Spacelabs at 1-800-522-7025 and select 2 for Technical Support. |  | 1075 | total 1075 units (614 units in the US and 461 outside the US) | distributed in AL, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MI, MN, MO, MS, MT, NC, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WY and Guam and in the following countries AFGHANISTAN, ARGENTINA, AUSTRALIA, BAHRAIN, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, EL SALVADOR, FRANCE, GERMANY, INDIA, IRAN (ISLAMIC REPUBLIC OF), ITALY, KUWAIT, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PALESTINE, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, and UNITED KINGDOM. | Open, Classified |  |  | 1075
Z-1091-2016 | 72486 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | BladderScan BVI 9600, Model Numbers: 0270-0451, 0270-0452, 0270-0754, and U270-0451. |  |  | November 30, 2015 | 2015 | Verathon, Inc. | The firm is providing customers with an updated Operations and Maintenance Manual for each of the BladderScan BVI 9600 and AortaScan AMI 9700 devices. The updated Operations and Maintenance Manuals clarify that these devices should not be used for the screening, detection, or diagnosis of abdominal aortic aneurysms (AAAs). | Device Design | Verathon sent an URGENT: MEDICAL DEVICE CORRECTION BLADDERSCAN BVI 9600 WITH AORTASCAN MODE &amp; AORTASCAN AMI 9700 letter dated November 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Verathon will manage other international consignees outside Europe internally and follow a similar approach as the firm requested Stericycle to perform. Verathon will instruct the companys foreign distributors to communicate the Customer Correction Letter and provide the revised OMMs to their international customers.This current Medical Device Correction is intended to reinforce the message in the May 2014 communication regarding the potential variance in measurement and to provide you with updated Operation and Maintenance Manuals that further reinforce the devices intended use and indications for use and that revise warning information, including warnings specified in the May 2014 communication.Customers are instructed to destroy all existing Operations and Maintenance Manuals and replace them with updated versions and complete the information on the Medical Device Correction Reply Form and fax or email the completed form within 10 business days to: Fax: (877) 907-9956 or Email: verathon3108@stericycle.com.Customers with any questions can contact your Verathon representative or Verathon Customer Care at (866) 875-0382. Verathon has live agent support on Monday through Friday from 8am to 5 pm Pacific Time.Customers should report any suspected malfunction or adverse event related to any AortaScan AMI 9700 devices and any BladderScan BVI 9600 devices to Verathon Customer Care at (800) 331-2313 |  | 24230400 | 2524 units total (1441 units of BVI 9600 and 820 units of AMI 9700 in US; 209 units of BVI 9600 and 54 units of AMI 9700 OUTSIDE the US | Worldwide Distribution - US (nationwide) and to the countries of : Guam, United Arab Emirates, Australia, Belgium, Canada , China, Czech Republic, France, Germany, Luxembourg, Netherlands, Norway, Saudi Arabia, Sweden , Taiwan, and United Kingdom | Open, Classified |  |  | 24230400
Z-1093-2016 | 73173 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Verso Shoulder Forked Retractor (Comprehensive Instrumentation 2 Prong Retractor) |  |  | February 16, 2016 | 2016 | Biomet, Inc. | Zimmer Biomet Inc. Verso Shoulder Forked Retractor (Comprehensive Instrumentation 2 Prong Retractor) supplied by Kirkstall Precision was made according to the W1.4034 (420 stainless) which does not conform to 420 S29 indicated on the print.All units manufactured to date were manufactured using 420 stainless. Therefore, all lots of PN: 402852 are considered in scope. | Mixed-up of materials/components | Zimmer Biomet sent an Urgent Medical Device Recall letter dated February 16, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. The following actions are REQUIRED: Immediately locate and remove the identified device(s) listed below from circulation.Carefully follow the instructions on the enclosed Response Form . Email a copy of the response form to CPWARFieldAction@zimmerbiomet.com prior to return of product.Use priority carrier for your shipment.If you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed Dear Risk/Recall Manager notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. However, you are charged with the location and return of these products.Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA:Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 1261 | 1261 units | Worldwide Distribution - US including CA, NY, VA, IL, IN, SC, NC, AR, MO, ND, AL, PA, MI, WA, OR, LA, FL, GA, OH, KS, KY, TX, MA, AZ, NJ, UT, IO, NM, WV, MS, NE, and MD OUS: Australia, Chile, Columbia, Costa Rica, Argentina, Korea, Japan, Panama, Malaysia, Puerto Rico, Canada, and The Netherlands. | Terminated | May 30, 2017 | 469 | 1261
Z-1035-2016 | 73014 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope Titanium Single-Use Video Laryngoscope. It is a single-use video laryngoscope system (GlideScope Titanium System) consists of a sterile-packaged, single-use video laryngoscope, multi-use Smart Cable and portable GlideScope video monitor. |  |  | January 29, 2016 | 2016 | Verathon, Inc. | The firm has become aware of the potential for disruption (flickering) in the video laryngoscopy image when GlideScope Titanium SU laryngoscope blades are in use. Video "flickering" appears as the intermittent break-up of the on-screen video image, appearing as either distorted horizontal or vertical bars of displaced video signal. | Process control | The firm, Verathon, began mailing the "URGENT: MEDICAL DEVICE RECALL"GLIDESCOPE VIDEO LARYNGOSCOPE TITANIUM SINGLE-USE BLADES letter to their consignees on January 29, 2016. The letter describes the product, problem and actions to be taken.Verathon will instruct the company's foreign distributors to communicate the Customer Recall Letter to their international customers.Verathon has contracted with a business services company, Stericycle, to coordinate the initial consignee communications in the United States, Australia and Canada. Stericycle will send the Customer Recall Letters, and conduct consignee follow-up communications for return of the Recall Response Forms.Verathon will replace all affected GlideScope Titanium Single Use (SU) blades that remain in your facility's inventory with new product that has been passed the enhanced screening test for video flickering.Customers were advised to take the following actions:-Fill out the attached Recall Response Form and return it to Verathon by fax (888) 943-5170, or email, verathon5165@stericycle.com. Please return the form even if you do not have any blades subject to the recall.If consignees need to expedite a product replacement contact Verathon representative or Verathon Customer Care at (800) 331-2313 (U.S. and Canada). The firm has live agent support Monday through Friday from 8:00am to 5:00pm Eastern Time. Consignees can also send email to CCTechSupport@verathon.com.Consignees who have any questions about this Product Recall, please contact Verathon representative or call (888) 943-4207 (U.S. and Canada) or send email to verathon5165@stericycle.com.Please report any suspected malfunction or adverse event related to any GlideScope Video Laryngoscope devices to Verathon Customer Care at the telephone numbers above and to FDA's MedWatch Program (1-800-FDA-1088 or www.fda.gov/medwatch). |  | 7733 | 7733 units (6377 units in the US and 1356 units International) | Worldwide Distribution - US (nationwide) and to countries of: Australia, Belgium, Canada, Denmark, France, Germany, Hong Kong, Italy, Luxembourg, Saudi Arabia, Singapore, Spain, Sweden, The Netherlands, U.A.E., and United Kingdom. | Terminated | March 01, 2017 | 397 | 7733
Z-1094-2016 | 73034 | No | MOQ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Battery Pack |  |  | December 23, 2015 | 2015 | Pro-Dex Inc | The sterile battery may contain particulates within the sterile package. | Process change control | A customer notification letter was sent 12/23/15.The letter instructs the customer to cease distribution and use of the recalled lots. It also recommends that the inventory and distribution records be evaluated immediately to ensure that affected product is removed from any potential point of use.Return product to:Returns Department (Recall complaint ClS-001)Pro-Dex, Inc2361 McGaw AvenueIrvine, California 92614 USAReference the Recall Complaint # on the outside of the shipping box.If any questions or concerns, please contact theQARA group at Pro-Dex at 949 769 3200. |  | 1995 | 1995 total | Distributed in FL. | Terminated | July 21, 2017 | 576 | 1995
Z-1057-2016 | 73266 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | FiberOptix Ultra 8 IAB: 8Fr 40cc; Product Code: IAB-05840-LWS |  |  | February 10, 2016 | 2016 | Arrow International, Inc., Division of Teleflex Medical Inc. | The sheath body may become separated from the sheath hub. If the separation occurs, there is a potential for bleeding from the device. | Under Investigation by firm | The firm, Teleflex Arrow International, sent an "Urgent Medical Device Recall Notification" letter dated 2/11/2016 via FedEx to its Customers. The letter described the product, problem and actions to be taken. The customers were informed of the reason for recall and were instructed to take the following actions: 1. If you have any products referenced in the attached enclosure, immediately discontinue use and quarantine those products appropriately. 2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507 Attn: Customer Service or email it to recalls@teleflex.com. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Arrow International. 3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email it to recalls@teleflex.com. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 13405 | 13,405 US and 33,735 OUS in total | Worldwide Distribution: US (Nationwide) and countries of: Argentina, Australia, Belgium, Brunei Darussalam, Brazil, Bahamas, Canada, Chile, China, Colombia, Ecuador, Guatemala, Hong Kong, Indonesia, India, Japan, Korea (Republic of [South] Korea), Myanmar, Mexico, Malaysia, Nepal, New Zealand, Panama, Peru, Philippines, Pakistan, Singapore, Thailand, Trinidad and Tobago, Taiwan, Uruguay, Viet Nam, and South Africa (Zuid Afrika). | Open, Classified |  |  | 13405
Z-1098-2016 | 73421 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Intra operative MRI Systems: Neuro II-SE , Neuro III-SV, iMRI 1.5T A, and iMRI 3T S. |  |  | February 10, 2016 | 2016 | Deerfield Imaging | Emergency helium venting lines (i.e. quench lines) may not have been properly installed. | Device Design | Consignees were sent a IMRIS "Urgent Field Safety Notice" letter dated February 10, 2016. The letter described the problem and the product involved in the recall. The firm also stated that it is important that customers review the portion of the quench line vent described and ensure that it is installed to the specifications supplied as part of the Planning Guide. Customers were reminded that it is important to have an emergency plan in place in the event Helium gas escapes into the magnet room or any other area. For questions customers can contact their local customer support hotline or +1-866-475-0525. |  | 49 | 35 USA, 14 OUS | Worldwide Distribution -- US, including the states of CA, CT, FL, GA, IL, IN, MD, MA, MN, MO, NH, NC, OH, OR, PA, TX, UT, VA, and WI; and ,countries of CHINA, CANADA, AUSTRALIA, FRANCE, JAPAN, and GERMANY. | Terminated | April 05, 2017 | 420 | 49
Z-1096-2016 | 73222 | No | NPM | Bone grafting material. | Dental | Dental | 510(k) | NuOSS Cancellous |  |  | August 25, 2015 | 2015 | Collagen Matrix Inc | On 8/25/2015, Collagen Matrix, Inc discovered that one of the 171 units of NuOss Cancellous product shipped to a distributor may have included an incorrect size. | Unknown/Undetermined by firm | Collagen Matrix Inc. notified their sole customer on November 17, 2015.On August 25, 2015, during a review of inventory, tt was discovered that of the 171 units product shipped, 19 units may have been the incorrect size.For further questions please call (201) 405-1477 |  | 171 | 171 units | US Distribution to the state of : MA | Open, Classified |  |  | 171
Z-1097-2016 | 73206 | No | NPM | Bone grafting material. | Dental | Dental | 510(k) | NuOss XC Sinus |  |  | May 13, 2013 | 2013 | Collagen Matrix Inc | During the 36 month real time shelf life testing of NuOss XC Sinus (BMCU-Umbrella) to verify the 3 year shelf life designated for the device, 3 of 13 products did not pass the recovery test, which requires that the compressed product fully expand within 1 minute of hydration time. | Process control | The firm, Collagen Matrix Inc., sent an email dated May 14, 2013 to its sole customer. The email described the product, problem and actions to be taken. The customer was instructed to return all product with expiration dates 12/2014 or earlier under Returned Goods, the firm will replace these products with new inventory; and return to Collagen Matrix, Inc., 15 Thornton Road, Oakland, NJ 07436 or quarantine products with expiration date 1/2016 until further notice.If you have any questions,call 201-405-1477 x 304 or email: www.collagenmatrix.com. |  | 347 | 347 units | US Distribution to state of: MA. | Terminated | June 21, 2017 | 1500 | 347
Z-1163-2016 | 72900 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Product 65 consists of all product under product code: KTT and same usage: |  |  | January 11, 2016 | 2016 | Zimmer Manufacturing B.V. | As a result of the insufficient sealer calibration data for product packaged in firm's Building II between August 2010 and April 2013. The affected products are sterile. | Packaging | Zimmer Biomet sent an Urgent Medical Device Recall Letter, dated January 11, 2016, to Distributors, Sales Representatives, Operation Managers, and Risk Managers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to 1.Review the notification and ensure that relevant personnel are aware of the contents.2. Assist your Zimmer Biomet sales representative with the quarantine of any affected product.3. Your Zimmer Biomet sales representative will remove the recalled product from your facility.4. Complete and return the attached Certificate of Acknowledgment form to corporatequality.postmarket@zimmerbiomet.com.5. If after reviewing this notification you have further questions or concerns please call Customer Service at 1-800-348-2759, or contact your Zimmer Biomet Sales Representative. |  | 5538 | 5538 | US Nationwide Distribution | Terminated | April 13, 2017 | 458 | 665035
Z-1172-2016 | 73253 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Accu-Chek Inform II Base Unit as a part of the Accu-Chek Inform II Blood Glucose Monitoring System |  |  | February 11, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Accu-Chek Inform II Base Unit might produce physical transmission errors in the form of data loss in the communication between the meter and the Data Management Systems (DMS). The issue can lead to the data loss or in the worst case to an erroneous assignment of the patient data (patient mismatch). The issue will only occur at sites using POTCT1-A communication via USB. | Software Design Change | Roche sent an URGENT MEDICAL DEVICE RECALL CORRECTION letter dated February 11, 2016, via UPS Ground (receipt required) to the consignee. The Consignee instructions were as follows:Âż If currently using USB connectivity, follow the instructions outlined in this Urgent Medical Device Correction (UMDC) to determine if your base unit is affected by the issue.Âż Until the new software is available, do not use USB connectivity if the base unit is affected by this issue; switch the connectivity method to Ethernet (TCP/IP) or Wi-Fi.Roche is not requesting return of affected product. Any returned product will be managed according to current procedures.Consignees with questions were instructed to contact the ACCU-CHEK Customer Care Service Center at 1-800-440-3638.For questions regarding this recall call 317-521-3911. |  | 97529 | US 5,604 devices, OUS 91,925 | Nationwide Distribution including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI, and WV. | Terminated | May 30, 2017 | 474 | 97529
Z-1171-2016 | 73129 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Sekure Chemistry Sekisui Diagnostics Acetaminophen L3K Assay, In-vitro Diagnostic |  |  | January 22, 2016 | 2016 | Sekisui Diagnostics P.E.I. Inc. | Crystal formation in the Color Reagent (R2) when left on board automated analyzers causing samples to under recover | Device Design | Sekisui sent an Important Field Correction - Customer Notification letter dated January 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Discard all reagent in house with a current expiration date of 2016-09.2. Update the expiration on all remaining kits on hand with the new expiration dates provided 3. When adding additional reagent to the analyzer, a new reagent wedge should be used.4. If crystals are observed within reagent stored on-board, reagent should be discarded and replaced with fresh reagent.5. Complete, sign and return the attached CONFIRMATION OF NOTIFICATION form by FAX to the number provided. This will indicate receipt of this Field Correction, act as confirm action of compliance with the actions,and allow for issuance of replacement material for expired product on hand (lots 47677, 47694, 47695).Questions contact Technical Services at 1-800-565-0265 or viaemail at PEIDiagnosticTechnical@sekisui-dx.com |  | 14407 | 14,407 kits. | Worldwide Distribution - US Distribution to the states of : AZ, CA, FL, IA, ID, IL, MA, , MN, MT and ND., and to the countries of Canada, France and Hong Kong. | Open, Classified |  |  | 14407
Z-1174-2016 | 73183 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | CORDIS AVANTI + CATHETER SHEATH INTRODUCER SYSTEM |  |  | December 31, 2015 | 2015 | Cordis Corporation | Incorrect cannula of the sheath introducer (smaller than intended). | Mixed-up of materials/components | The firm, Cordis, sent an 'Urgent MEDICAL DEVICE CORRECTION (Field Safety Notice) dated 12/31/2015 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to do the following:1) Read the Description of the problem section carefully to fully understand theissue involved.2) Check your inventory to determine if you have any remaining affected productin your possession. Check all storage and usage locations. The purpose is toidentify the product, not to remove the product.3) Keep a copy of this notice with any affected product.4) Review, complete, sign and return the enclosed Acknowledgement Formdirectly to Cordis by fax to (908) 429-8287 Alternate fax (904) 928-5077. number or scan and email to MiamiFRA@its.jnj.com.5) Share this letter with anyone who needs to be informed in your facility and in any facility where potentially affected devices may have been transferred.6) Maintain awareness of this notice until all affected product has beenconsumed. If vessel dilator to sheath incompatibility is experienced, you canreport the incident through the standard complaint process (1-800-327-7714Option 3).For any health care professional with medical concerns, please contactCordis to speak to a clinician at 1-800-327-7714 Option 3, Mondaythrough Friday from 9:00 AM to 5:00 PM EST.For questions related to the Correction that are not adequately addressedin this letter, please contact the Cordis Field Action Coordinator at (786) 313-2365.For questions related to the Acknowledgement Form and its return, pleasecontact Cordis QA at (786) 313-8730." |  | 1985 | 1,985 units | US Nationwide Distribution including states of:AL, AZ, CA, CO, CT, FL, GA, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NH, NJ,NY, OH, OR, PA, RI, TN, TX, VA, WA, WV, and Hawaii. | Open, Classified |  |  | 1985
Z-1173-2016 | 73208 | No | NPM | Bone grafting material. | Dental | Dental | 510(k) | ZCORE Porcine Xenograft Particulate in Syringe |  |  | November 17, 2015 | 2015 | Collagen Matrix Inc | The distributor, Osteogenics Biomedical, notified the firm that the syringes were not working properly and the syringes were being held in quarantine. | Component design/selection | Collagen Matrix Inc. notified their sole customer on 11/17/2015. |  | 200 | 200 units | US in the state of TX | Open, Classified |  |  | 200
Z-1176-2016 | 73422 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | The Spirit TM Select bed is a Med-Surg bed intended to support and transport |  |  | February 18, 2015 | 2015 | CHG Hospital Beds Inc | It was identified that in some situations the bed exit alarm may not function as intended. The software code for the bed exit system has the potential to auto-reset erroneously. In some situations, the software code does not allow enough time for the weight value to fall to zero once the patient egresses from the bed, in which case, the bed has the potential to reset the bed exit alarm even without a patient on the bed.The bed exit on the Spirit Select bed may incorrectly reset if the Auto Reset option is selected. A patient fall may result in soft tissue or bone injury which may require medical or surgical intervention to prevent permanent disability. If the bed exit alarm incorrectly resets then medical assistance after a fall or during entanglement may be delayed. Patients with dementia or disorientation are at greatest risk. There have been two reported injuries associated with these issues to date. A Medical Risk Assessment determined that there is a remote probability of occurrence of harm requiring medical intervention. | Software design | On 2/18/2016, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any questions or concerns, please contact Stryker Customer Service (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 1892 | 1892 | Distributed US (nationwide) and in Canada. | Terminated | May 23, 2017 | 825 | 1892
Z-1177-2016 | 73213 | No | HJK | Polymethylmethacrylate (PMMA) diagnostic contact lens. | Ophthalmic | Ophthalmic | 510(k) | Volk G-4 Small Ring, No Flange, Gonio Lens with case, UPC: 811608013602 |  |  | February 02, 2016 | 2016 | Volk Optical Inc | The firm discovered that the incorrect lot number was engraved on the product. | Employee error | Volk Optical sent an Urgent Medical Device Recall letter dated February 4, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine any affected product. Also, if product was further distributed customers should identify their customers and notify them at once of the product recall and include a copy of the letter. The affected devices should be returned to Volk with an assigned Return Merchandise Authorization (RMA) number. Customers should contact Volk at (440) 942-6161 or toll free at (800) 345-8655 as soon as they receive the notification to setup a RMA for their device(s). Customers were instructed to complete and return the enclosed response form as soon as possible. Customers with questions should call (440) 510-0745. |  | 9 | 9 lens | Worldwide Distribution - US (Virginia only) Internationally to Canada, France &amp; Indonesia. | Terminated | April 26, 2017 | 449 | 9
Z-1178-2016 | 73267 | No | GDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Pilling Modular Verres Body/Obturator, Product Code 140880 |  |  | January 08, 2015 | 2015 | Teleflex Medical | Packaging error in which customers who ordered product code 140880 (Verres cannula for use with a separately available needle tip) were incorrectly provided with product code 140878 (Verres cannula with integral needle tip). | Packaging process control | Teleflex sent an Medical Device Advisory Notification letter dated January 8, 2016, to all affected customers notifying them of the packaging error.Our records indicate that you have received products that are subject to this notification.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at recalls@teleflex.com or 1-866-246-6990. |  | 8 | 8 | US Distribution to the state of : UT | Terminated | April 14, 2017 | 827 | 8
Z-1180-2016 | 73009 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Quick Connects. Steris Corporation. Model numbers: QKC1736E, QKC1736INT, QKC1737E, and QKC1737INT. Each Quick Connect contains a flow unit consisting of tubing, tethers, and adapters. For sterilant and rinse water used in endoscope washers. |  |  | December 04, 2015 | 2015 | Steris Corporation | The firm discovered that one of the hysteroscope models listed for processing in four existing Quick Connects has one inlet port rather than two inlet ports as indicated in the Quick Connect processing instructions. | Nonconforming Material/Component | On 1/18/2016 the firm sent Voluntary Field Correction NOtices to their customers. |  | 809 | 809 units | Worldwide distribution. US nationwide, Australia, Bahrain, Bulgaria, Canada, Greece, Italy, Malaysia, Mexico, Qatar, Russia, Spain, and United Kingdom. | Open, Classified |  |  | 809
Z-1181-2016 | 73096 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 980 Ventilator System, PB980 Ventilator (980xxxxxxxx). |  |  | January 15, 2016 | 2016 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Graphical user interface (GUI) unresponsive to touch and Loss of primary ventilation under certain circumstances. Covidien Respiratory and Monitoring Solutions, now a part of Medtronic, issued a field corrective action notice for two issues on all models of Puritan Bennett 980 (PB980) ventilator. | Under Investigation by firm | Customers were notified via an Urgent Medical Device Safety Notification, dated February 8, 2016, on 1/14/16. The letter identified the affected device and the issues involved. Consignees are requested to forward the recall letter to anyone they may have transferred or assigned for use one of the potentially affected ventilators. Medtronic is investigating the root cause of these conditions and will provide a service update to resolve the issues as soon as corrections can be implemented. Customers are to complete the attached form and return it per the instructions to confirm receipt and understanding of the information. Questions or issues should be directed to the Technical Support Department at 800-255-6774. |  | 1864 | 1,864 units | Worldwide Distribution -- US, including but not limited to the territory of Puerto Rico; and Ireland, Jordan, Saudi Arabia, Bahrain, Australia, Japan, Israel, Turkey, Chile, South Africa, Costa Rica, United Arab Emirates, Qatar, Peru, Panama, Singapore, Kuwait, Lebanon, Brazil, Mexico, Colombia, Russia, Uruguay, Slovenia, Morocco, Argentina, Egypt, and Romania. | Terminated | July 11, 2017 | 543 | 1864
Z-1179-2016 | 73238 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Alaris Pump VersaSafe Infusion Set, Model No. 2122-0007 |  |  | February 15, 2016 | 2016 | CareFusion 303, Inc. | CareFusion is recalling the Alaris VersaSafe Infusion Set because of separation and/or leakages at the VersaSafe split septum port and tubing may lead to leakage of the administration set. | Other | A customer notification letter dated 2/15/16 will be sent to customers to inform them that CareFusion is recalling the Alaris VersaSafe Infusion set number 2122-0007, Lot number 13085603. The letter informs the customers that Separation or leakage at the VersaSafe split septum port and tubing may lead to leakage of the administration set. The separation and leakage may occur during infusion. The letter informs customers of the potential risk and actions to be taken. Customers with recall related equations are instructed to contact CareFusion Support Center at (888) 562-6018, 7am-4pm PST. Customers with adverse event reports are instructed to contact Customer Advocacy at (888)812-3266, customerfeedback@carefusion.com, 24 hrs/day, 7 days/wk. Customers are instructed to complete and return the enclosed mandatory customer response card to acknowledge the receipt of the customer communication. |  | 5760 | 5,760 units | Michigan | Open, Classified |  |  | 5760
Z-1187-2016 | 73438 | No | NEY | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Short Percutaneous Antenna with Thermosphere" Technology |  |  | February 23, 2016 | 2016 | Covidien Medical Products (Shanghai) Manufacturing L.L.C. | Medtronic is voluntarily recalling specific item codes and production lots of Covidien Emprint" Percutaneous Antenna with Thermosphere" Technology due to the potential for the generator to shut down in the precence of high temperature measurements in the Emprint" Percutaneous Antenna. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter dated February 23, 2016 to all affected customers via Federal Express or Certified Mail. The letter identified the affected product, problem and actions. Customers were advised to quarantine and discontinue use of the affected product and to Return product. For questions or concerns, contact your Medtronic representative or Customer Service at (800) 882-5878. |  | 4920 | 4,920 units | Worldwide Distribution - US Nationwide, VA/govt/military consignees. | Terminated | June 21, 2017 | 484 | 4920
Z-1182-2016 | 73452 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Horizon Medical Imaging (HMI) products versions 4.6.1 to including 11.9 and McKesson Radiology (MR) products 12.0 and 12.1.1. |  |  | February 18, 2016 | 2016 | Mckesson Medical Immaging | McKesson has identified a design deficiency where under rare circumstances, imported images/studies may re-use a non-unique image directory. This issue may cause incorrect images to be displayed for a patient. | Software design | McKesson sent an Urgent Field Safety Notice dated February 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Immediate actions to take to reduce the risk to patientAll customers receiving this safety notice must contact McKesson Support as soon as possible at 1-800-663-2533 US/Canada and International Toll Free Radiology - 00 800 626 20009:1) To acknowledge that they have read and understood this Field Safety NoticeAnd2) Arrange installation of software update/upgrade.This notice should be distributed to all personnel within your organization who need to be aware of this issue and the required actions, and alert other affiliated parties that are affected by this Field Safety Notice.For further questions, please call (404) 338-3556 |  | 631 | 631 devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and to the states of : AL, AK, AZ, AR, CA, CO, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY., and the countries of : Guam, Great Britain, France, Canada, Ireland and Australia | Open, Classified |  |  | 631
Z-1184-2016 | 73403 | No | INK | Powered wheeled stretcher. | Physical Medicine | Physical Medicine | 510(k) | Prime Series Stretchers (Fifth Wheel, Big Wheel and Zoom) |  |  | February 05, 2016 | 2016 | Stryker Medical Division of Stryker Corporation | The hydraulic jack assemblies (PN: 0753-002-101, 0753-002-070, 0753-002-001, and 0753-002-001) were manufactured using the incorrect oil. The oil used does not contain the same wear resistance additives as the specified oils and therefore its use increases the potential for wear which may cause a leak event or jack drift. Testing was done to evaluate the jacks with incorrect oil and leak events we | Mixed-up of materials/components | On 2/5/2016, URGENT MEDICAL DEVICE NOTIFICATION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any urgent questions or concerns, please, contact Stryker Customer Service at (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m.-5 p.m. (ET). |  | 36 | 36 | Distributed in DC and the states of AL, CA, FL, MN, MT, NC, NE, NY, PA, RI, TN, TX, UT, and WV. | Open, Classified |  |  | 36
Z-1183-2016 | 73479 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens Syngo Dynamics-a Picture Archiving and Communication System (PACS) Model Numbers: 10091805, 10091807, 10091637, 10091673 |  |  | January 25, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens' conducting a recall due to a potential issue when using the measurement package of the VA10 version of syngo Dynamics. | Software design | The firm, Siemens Healthcare, sent a "Customer Advisory Notice" dated 1/19/2016 to customers. The notice described the product, problem and actions to be taken. The customers were instructed to follow the instructions provided on what steps the user can take to avoid the issue and to forward the notice to the new owner, if you sold this device/equipment and its no longer in your possession. Siemens is preparing a modification of syngo Dynamics that will resolve this potential malfunction.If you have any question, contact Regulatory Technician Specialist at 610-448-6478 or email: Anastasia.sokolova@siemens.com. |  | 228 | 228 units | US Nationwide Distribution including states of: MN, NC, NE, FL, TX, CA, AR, MI, MA, IL, DE, VA, SC, OH, CO, WI, MO, NY, CT, UT, PA, WA, NJ, GA, IN, MD, TN, LA, OR, AZ, ND, KS, AL, NM. | Open, Classified |  |  | 228
Z-1186-2016 | 73352 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Stockert Heater-Cooler 3T; Item Number 16-02-8, H3T/240V/60Hz ; 16-02-82, H3T208V/60Hz ; 16-02-85, H3T/120V/60Hz. Device Listing E572259 |  |  | August 15, 2012 | 2012 | Sorin Group USA, Inc. | On August 15, 2012, Sorin informed consignees via letter of changes made to create Version 12 of the Heater-Cooler 3T IFU. The revised IFU included enhancements to achieve water quality that meets the standard for German drinking water. | Device Design | Sorin Germany, via Sorin Colorado, distributed an important information letter to its US customers on August 15, 2012 informing them about an update of the cleaning and infection procedure for the U.S. market that were included in Revision 12 of the IFU. |  | 1125 | 1125 units | Worldwide Distribution - US Nationwide, VA/govt/military | Terminated | May 26, 2017 | 1745 | 1125
Z-1192-2016 | 73150 | No | LGW |  |  | Neurology | PMA | Algostim LLC, Algovita Implanted Pulse Generator stimulator, Model 2408 ( 3 x 8 channel) |  |  | January 26, 2016 | 2016 | QiG Group LLC | The implantable pulse generator (IPG) model 2408 had suboptimal performance establishing a connection to the external devices unless they were brought into close proximity (as close as 5 cm). The following functions may not be available: Reprogramming the stimulator, Adjusting stimulation parameters (on/off, intensity modifications, stimulation program selection), and Recharging the stimulator. | Nonconforming Material/Component | Consignee was sent on 1/26/2016 a Algostim, LLC "Urgent Field safety Notice" dated January 21, 2016. The letter described the problem and the product involved in the recall. The letter provided Background information as well as described the Clinical Implications, Rate of Occurrence, Recommendations and Mitigations. Advised consignee to deliver the respective letter to each physician who have implanted affected product and collect signed Response forms. |  | 40 | 40 | Distributed in Germany and Luxembourg. | Terminated | April 05, 2017 | 435 | 105
Z-1190-2016 | 73363 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | InstinctÂż JavaÂż System (IFU reference: 046WAN0000T) . |  |  | February 12, 2016 | 2016 | Zimmer Spine, Inc. | Zimmer Biomet is initiating a Medical Device Correction Notice to provide updated cleaning and disinfecting instructions for use of the InstinctÂż JavaÂż System (IFU reference: 046WAN0000T) and Universal ClampÂż System (IFU reference: SNA027-N-90001). | Process control | Consignees were sent on 2-12-2016 a Zimmer Biomet "Urgent Medical Device Correction Notice" letter dated February 12, 2016. The letter was either addressed to Hospital Representative or to Distributors, Sales Representatives, and Distribution Operations Managers. The letters described the problem and the product involved in the recall. Advised consignees to immediately utilize the updated and attached Instructions For Use. Requested consignees to complete the "Acknowledgement of Field Notification" form. For questions contact Customer Service at 1-866-774-6368 or e-mail corporatequality.postmarket@zimmerbiomet.com. |  | 2506 | 2506 | US (nationwide)-including the states of AL, AZ, AR, CA, CO, FL, GA, ID, IL, IN, IA, LA, ,MD, MA, MI, MN,MO, NV, NH, NY, NC, OH, OK,. PA, TN,TX, WA, and WI. | Open, Classified |  |  | 5990
Z-1191-2016 | 73314 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si (IS3000) and da Vinci Xi (IS4000) Surgeon Side Console (SSC)Touchpad; Model number 187505-04; Intended to assist in the accurate control of Intuitive Surgical Endoscopic Instruments. The system can also be employed with adjunctive mediastinotomy to perform coronary |  |  | February 16, 2016 | 2016 | Intuitive Surgical, Inc. |  | Other | Urgent Medical Device Recall letters were sent on March 2, 2016. Letters advised users of the issue and that Intuitive Representative will visit each site to replace the affected compact flash card. |  | 542 | 542 (351 SCC Touchpads and 191 PSC touchpads) | US and Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Finland, France,Germany, Hong Kong, India, Israel, Italy, Japan, Malaysia, Monaco, Netherlands, Philippines, Portugal,Qatar, Romania, Russian Federation, Saudi Arabia, South Korea, Sweden, Switzerland, Taiwan, Turkey, United Kingdom and Venezuela. | Open, Classified |  |  | 542
Z-1195-2016 | 73204 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Arthrex IBalance TKA Tibial Tray Size 2 - Catalog # AR-503-TTTC, Size 3 - Catalog # AR-503-TTTD, Size 4 -Catalog # AR-503-TTTE, Size 5 - Catalog # AR-503-TTTF, Size 6 - Catalog # AR-503-TTTG , Size 7 -Catalog # AR-503-TTTH, and Size 8 - Catalog # AR-503-TTTJ. |  |  | December 17, 2015 | 2015 | Arthrex, Inc. | Devices reported to have a smooth texture to the outer surface of the metal, which differed from a previous that had a textured outer surface. | Nonconforming Material/Component | On February 5, 2016 Arthrex sent Urgent Medical Device Voluntary Recall letters to their distributors and Surgeons/Medical Facilities reporting the removal. The letter identified the reason for the recall, the potential patient impact, and actions to be taken. Users were to discontinue use of the identified affected devices and follow the instructions provided to return the affected devices that have not been implanted. To return the devices, contact Arthrex at 1-866-267-9138 or by fax at 1-239-591-6943 or via e-mail at Complaints@Arthrex.com. Customers must complete and return the Acknowledgement and Receipt. |  | 2378 | 2,378 units. | Nationwide Distribution -- AR, AZ, CA, CO, DC, DE, FL, GA, IL, IN, KS, MA, MD, MI, MN, MO, NC, NJ, NM, NV, NY, OH, OK, PA, RI, TN, TX, WA, WI and District of Columbia. | Open, Classified |  |  | 2378
Z-1175-2016 | 73064 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Toshiba DRAD-3000E FPD Wireless System |  |  | January 14, 2016 | 2016 | Toshiba American Medical Systems Inc | When a user performed radiography using the wireless FPD, a message window appeared on themonitor stating that imaging transmission was not completed with no radiographic image. It also indicated to select the "OK" button to re-acquire the image data or to select theÂż'Cancel" button to cancel the re-acquisition. As instructed the user selected "OK" and the same messagewindow appeared. This sam | Radiation Control for Health and Safety Act | Toshiba's planned action: 1."Toshiba America Medical Systems, Inc. (TAMS) will advise contact customers to stop using the system if any abnormalities are found using the system.2. TAMS will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that TAMSwill without charge, remedy the defect or bring the product intocompliance.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15working days of receipt of this letter in the manner specified in 21 CFR1003.21 and 1003.22. This office and the Food and DrugAdministration (FDA) district office coordinator noted below are to beincluded in the notification.2. TAMS will implement this CAP by June 30, 2016.Customers with questions were advised to call 301-796-5910.For questions regarding this recall call 714-730-5000. |  | 28 | 28 | Nationwide distribution to PA, UT, FL, NY, LA, WV, NJ, OH, WI and TX. | Open, Classified |  |  | 28
Z-1229-2016 | 73121 | No | EZL | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Covidien Dover PREC PRM 18FR SILVER FLY TY |  |  | January 22, 2016 | 2016 | Covidien LLC | Aplicare Castile Soap towelettes included in the Foley Trays and Mid stream specimen kits are contaminated with bacteria | Nonconforming Material/Component | Medtronic/Covidien notfiied customersl via Federal Express on January 22, 2016,, The letter instructs customers to not use the Castile Soap Towelettes contained in the product kits listed on Attachment 1 and discard the affected towelettes. The Castile Soap packet included in these kits does not impact the integrity of the remaining products in the kit. Questions or concerns, contact your Medtronic representative or Customer Service at (800) 882-5878.For questions regarding this recall call 203-492-5000. |  | 3280 | 3280 | US (nationwide) and to Canada. | Terminated | June 20, 2017 | 515 | 909552
Z-1230-2016 | 73384 | No | NCD |  |  | Microbiology | PMA | QFT TB Antigen Tube. |  |  | February 02, 2016 | 2016 | Qiagen Sciences LLC | Potential for obtaining a false positive result due to the possibility of endotoxin presence. | Nonconforming Material/Component | QIAGEN sent an "Urgent - Notification of Market Withdrawal" dated February 2, 2016, to all affected customers. The letter identify the product the problem and the action needed to be taken by the customer.Immediately discontinue use of this lot and discard any remaining inventory. Please acknowledge receipt of this notification and actions taken by completing the information on the next page, and return to: TechService-NA@qiagen.com or fax +1-661-775-7479.QIAGEN apologizes for any inconvenience this market withdrawal has caused, and continues to take every effort possible to ensure the highest quality of product is always available to our valued customers around the globe. Please do not hesitate to contact your local QuantiFERON sales representative or customer service representative should you have any questions or concerns.For further questions, please call ( 240) 686-7500. |  | 152900 | The total amount of product distributed 152,900 Tubes | Worldwide Distribution - US Distribution to the state of : Georgia., and to the countries of : Austria, Belgium, Switzerland, Czech Republic, Germany, Denmark, Estonia, Spain, France, Great Britain, Ireland, Korea (South) Luxembourg, Netherlands, Poland, Portugal, Qatar, Russian (Federation), Singapore, Turkey and South Africa. | Open, Classified |  |  | 152900
Z-1233-2016 | 73292 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas p 512 pre-analytical system |  |  | February 23, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Due to a false triggering or detection of the lifting gripper READY signal, sample tubes are not correctly placed back in the Rack Tube Transport (RTT) after the decapping process. Therefore, open sample tubes can be dropped in the cobas p 512, spilling the sample material. | Software design | Roche Diagnostics sent an "Urgent Medical Device Correction" letter via UPS Ground (receipt required) on February 23, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.To prevent this malfunction, a software patch has been developed and validated. The installation of this software patch is mandatory and will be performed by a Roche Field Engineering Specialist (FES). Your Roche FES will contact you to schedule the installation of the software patch. Until the software patch is implemented, Roche recommends that instrument operators monitor their systems for the occurrence of Error 140 Gripper has lost tube. If this error occurs, all components in contact with a spill must have extra cleaning. Refer to the cobas p 512 pre-analytical instrument Operators Manual Version 1.5.Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 for further questions. |  | 35 | 35 Units | US Distribution including Puerto Rico and to the states of :TX, OH, TN, AZ, WA, MI and GA | Open, Classified |  |  | 35
Z-1232-2016 | 73481 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo X Workplace is a medical workstation for real-time viewing, image manipulation, 3D-visualization, communication, and storage of medical images and data on exchange media. It is used for diagnostic image viewing and post processing during interventional procedures. |  |  | February 24, 2016 | 2016 | Siemens Medical Solutions USA, Inc | After importing, the segmentation results appear mirrored at the CARTO system and can't be used for the ablation procedure. | Software design | Siemens sent an Important Customer Safety Notice dated February 24, 2016, to all affected customers notifying them that the issue will be remedied with a software update via AX004/16/S. Following the installation of this update, the segmentation results can be used at the CARTO system as described in the operator manual. Customers with questions were instructed to call 800-888-7436. For questions regarding this recall call 610-448-6461. |  | 6 | 6 | Nationwide Distribution to IL, NY, MT, and MN. | Open, Classified |  |  | 6
Z-1244-2016 | 72991 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System (Software Version 3.0.0 and after) |  |  | December 15, 2015 | 2015 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Due to Operator Console software anomaly, a change in treatment completion status of the last patient of the day may be altered from "Completed (or Discontinued)" to "Untreated" under specific conditions. | Software design | Due to Operator Console software anomaly, a change in treatment completion status of the last patient of the day may be altered from "Completed (or Discontinued)" to "Untreated" if all the following conditions are met: (1) At least one patient has been treated on the day; (2) A major error occurs when VeroTM/MHI-TM2000 is operated in a mode, other than Clinical Mode, which means Daily Check Mode or Maintenance Mode. (3) The major error is NOT associated with a function that uses therapeutic/ MV radiation parameters, such as an error in Isocenter Check function, etc. Note 1: If the major error is associated with a function that uses therapeutic/MV radiation parameters, the problem that treatment completion status is changed will never occur. The Operator Console changes the treatment completion status in the course of processing the major error. Note 2: Although treatment completion status is changed to "Untreated", its delivered dose result has been correctly saved into R&V; System. It has not been lost. Advice on Actions by Customer/Users: 1. Please make sure to perform the operations as described below in case a major error occurred in a mode other than Clinical Mode. (1) Cancel the error.(2) Click [Clinical] button on the left of Operator Console screen. (3) Verify, on R&V; System screen, that treatment completion status of the patients who were treated on the day are all correct. The treatment completion status is indicated in different colors as follows depending on the status: Untreated: BLUE, Completed: GRAY, Discontinued: RED. Customers requiring support concerning this issue, please contact the Customer Service Representative: 1-800-597-5911 (Mon - Fri) or via email us.support@brainlab.com. If any other problem is detected, please stop operation and contact the Customer Service representative. |  | 25 | 25 devices (4 domestically in U.S.A. and 21 internationally) | Distributed in the states of Florida, New York, Ohio &amp; Texas, and the countries of France, Germany, Japan, Italy, Korea, &amp; Belgium. | Terminated | April 20, 2017 | 492 | 25
Z-1246-2016 | 73197 | No | GDO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Irrigation Clip 30; Intended to position and hold irrigation tubing during surgical procedures. |  |  | November 02, 2015 | 2015 | The Anspach Effort, Inc. | Assembled with the incorrect (smaller) front clips (component 44-0495 instead of 44.0498). | Nonconforming Material/Component | On Nov 2, 2015 Anspach sent a letter to their customers informing them of the recall. Customer Immediate Actions: Please complete the attached reply form confirming your receipt of this letter. Return the completed form by fax to the number or email provided on the form. Returning the form promptly will prevent you from receiving repeat notices. If you distribute any of the products to otherservices or facilities, please forward this information as appropriate. We understand this product recall removal) may have an impact on your facility. The FDA has recommended that healthcare facilities follow their risk management procedures as it relates to this notification. We apologize for any inconvenience that this Product Removal may create and appreciate your cooperation with our request. If you have any questions regarding this correction, please contactthe Complaint Handling Unit Manager at 561-494-3673 or contact your Anspach Sales Representative. |  | 50 | 50 clips | Distributed in the states of PA and MN, and the country of Switzerland. | Open, Classified |  |  | 50
Z-1245-2016 | 73316 | refractor, manual, non-powered, including phoropter | HKN | Manual refractor. | Ophthalmic | Ophthalmic | 510(K) Exempt | RT-5100 (Epic 5100) Refractor and RT 3100 Refractor: |  |  | February 29, 2016 | 2016 | Nidek Inc | There has been a reoccurrence of an issue on the RT 5100 and RT 3100 Refractor of the Near Point Chart Arm lowering spontaneously with potential for slight injury to face or nose of patient. | Component design/selection | Nidek sent a Recall letters dated February 29, 2016 to all affected site. The letter identified the affected product, problem and actions to be taken. Nidek Customer Service will be replacing the affected product with new parts to address this recall. For questions call Customer Service at 1-800-722-0219. |  | 2703 | 2703 separate serial numbers | US Nationwide Distribution | Open, Classified |  |  | 2703
Z-1247-2016 | 73407 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | Neurotherm Curved Radiofrequency (RF) cannnula 10 cm |  |  | February 26, 2013 | 2013 | Neurotherm, Inc. | Straight needle labeled as a curved needle | Process control | NeuroTherm notified consignees by telephone between February 2- March 15, 2013 about the affected product, problem and actions to be taken. Customers were instructed to examine inventory and return devices. |  | 135 | 135 cannulas (27 boxes/5) | US Nationwide Distribution in the states of CA ,CO, OH, OR | Terminated | March 25, 2016 | 1123 | 135
Z-1239-2016 | 73315 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Alaris PC Unit, Infusion Pump Model 8000, Part No. TC10005092 |  |  | February 12, 2016 | 2016 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC unit because a component on the PC unit power supply may cause a "System Error" or "Missing Battery" error code (120.4630). | Under Investigation by firm | A customer notification letter dated 2/25/16 will be sent to all their customers who purchased the Alaris PC unit (Model 8000) to inform them that CareFusion is recalling it because a component on the PC unit power supply board may cause a "System Error" or "Missing Battery" error code (120.4630). The letter informs the customers of the potential risk involved and the actions to be taken. Customers with recall related questions are instructed to call CareFusion Support Center at (888) 562-6018, 7am-4pm PST, Mon-Fri or email at supportcenter@carefusion.com. Customers with clinical inquires, product complaints or clinical troubleshooting are instructed to contact customer advocacy at (888) 812-3266, 24 hrs/day, 7 days/wk or email at customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris system are instructed to contact technical support at (888) 812-3229, 6am-5pm PST, Mon-Fri or email at DL-US-INF-TechSupport@carefusion.com. |  | 170 | 170 units | Nationwide in US: AR, AZ, CA, FL, GA, IN, KY, MD, MI, MN, MO, NJ, NM, NY, OH, OR, PA, SC, TN, TX, VA, WI. | Terminated | June 23, 2017 | 497 | 170
Z-1237-2016 | 73398 | No | LJE |  |  | Gastroenterology/Urology | 510(k) | UltraxxTM Nephrostomy Balloon Set |  |  | January 27, 2016 | 2016 | Cook Inc. | Cook Medical is voluntarily recalling specific lots of the Ultraxx" Nephrostomy Balloon Set and the AscendÂż AQÂż Ureteral Dilation Balloon Catheter Set because of a labeling mix-up. Cook Medical has received reports of packages labeled as Ultraxx" Nephrostomy Balloon Set actually containing theAscendÂż AQÂż Ureteral Dilation Balloon Catheter Set and vice versa. | Labeling mix-ups | Cook Medical sent an Urgent Medical Device Recall Notification letter dated January 27, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to quarantine any unused affected product. Immediately collect and return all unused affected products to Cook Medical If you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235.For questions regarding this recall call 800-457-4500. |  | 20 | 20 | Worldwide Distribution - US including GA, ID, LA, OR, SC, TX and Internationally to CHINA. | Terminated | May 17, 2017 | 476 | 20
Z-1236-2016 | 73366 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Ambu, Aura Gain - Single Use Laryngeal Mask- Sterile |  |  | February 10, 2016 | 2016 | Ambu Inc. | Ambu, Aura Gain laryngeal Mask, was found to have lack of Sterility Assurance. The firm discovered the problem during final quality control at manufacturing site. Damage to the sterile barrier (the pouch) may compromise a medical device's sterility,leading to contamination of the product and,ultimately,patient complications. | Process control | Filed Safety Notice was sent via USPS Certified Mail to all customers on Dec 10,2015. |  | 45990 | Total Numbers of unit distributed: 45990 pes (USA) | Please see attached consignee list | Open, Classified |  |  | 45990
Z-1234-2016 | 73337 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | BioTex Reusable Adapter Kit, Part Number 401-021-1010, packaged non-sterile in a dedicated case for steam sterilization |  |  | February 22, 2016 | 2016 | Medtronic Navigation, Inc. | Medtronic Navigation is recalling the Biotex Adapter Kit because it was commercially distributed by BioTex without a cleared premarket [510(k)] from FDA. | No Marketing Application | The firm, Medtronic, sent an "URGENT Medical Device Recall" letter dated 2/2016 to its consignees. The letter described the product, problem and actions to be taken. The identified consignees will be visited by the assigned Medtronic field representative to explain the need for the field action. The consignees are instructed to locate and remove/secure the kit from use. The Medtronic field representative will assist the consignee with identifying and removing the kit, and complete a Customer Visit record for each consignee they are assigned/visit to document the results of their visit. The consignee is instructed to sign the customer visit form presented by the representative. A copy of this record and a letter explaining the FCA will be left with the consignee.Please contact Medtronic Technical Services at 1-800-595-9709 if you have any questions regarding the subject action or the content of the letter. |  | 4 | 4 | US Distribution to states of: FL, NJ, AZ, and NC. | Open, Classified |  |  | 4
Z-1235-2016 | 73343 | No | LQE |  |  | Ophthalmic | PMA | KAMRA Inlay, ACI 7000 Corneal Inlay, Part No. 76195-01 and 76195-10 |  |  | February 11, 2016 | 2016 | AcuFocus, Inc. | AcuFocus is recalling KAMRA ACI 7000 Corneal Inlays due to concerns with the labeled shelf life. | Under Investigation by firm | The firm, AcuFocus, sent an "URGENT FIELD SAFETY NOTICE" letter dated 2/11/16 to all their customers to inform them that AcuFocus is recalling 2 lots of KAMRA ACI 7000 Corneal Inlays due to concerns with the labeled shelf life. The described the product, problem and the actions to be taken. Customers are instructed to report any abnormal appearance of the implanted KAMRA inlay to AcuFocus, inc; stop using and remove from your inventory all affected lots; complete and return the attached Customer Recall Response Form via email: customerservice@acufocus.com, Fax: +1 (949) 585-9545 Attn: Recall Coordinator, or mail to: AcuFocus Inc., Attn: Recall Coordinator, 32 Discovery, Suite 200, Irvine CA 92618, within 3 business days even if you have no inventory affected by this recall; and provide the attached Customer Field Safety Notice and Customer Recall Response Form to your customers who have or may have received product from the affected two lots. Customers with product complaint events are instructed to report regarding KAMRA inlays involved in this recall event, and to inform AcuFocus, Inc. by calling +1 (949) 585-9511 ext 608. Customers with any questions related to the return process, are instructed to contact an AcuFocus Customer Service Representative at +1 (949) 585-9511 ext 608. |  | 182 | 182 units | International distribution in countries of: Canada, Germany, Saudi Arabia, Turkey, and United Kingdom. | Open, Classified |  |  | 182
Z-1243-2016 | 73052 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE 2000 EstradiolTest Code E2 |  |  | January 13, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | The drug fulvestrant (FaslodexÂż) may cause falsely elevated estradiol results. | Device Design | Siemens issued an Urgent Medical Device Correction (UMDC) to affected Siemens Healthcare Diagnostics customers in the United States on January 13, 2016 to be delivered to customers on January 14, 2016 via Federal Express. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on January 13, 2016. The letters identified affected product, problem and the actions to be taken. Customers may continue to use the estradiol assays to report patient results however, the assays should not be used for patients being treated with fulvestrant. Customers where instructed to complete and return the Correction Effectiveness Check attached. For questions contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 9334 | 9334 units | Worldwide Distribution  US Nationwide , AR, AU, AU, BD, BO, BR, CA, CH, CL, CN, CO, DO, EC, EG, HK, ID, IE, IL, IN, JP, KR, LK, LY, MX, MY, NI, NP, NZ, PE, PH, PY, SA, SG, SV, TH, TW, UY, VE, VN, and ZA. | Open, Classified |  |  | 54134
Z-1266-2016 | 73157 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | The S.C.O.R.E.S. Unit (self contained operating room equipment |  |  | January 13, 2016 | 2016 | PMBS, LLC | Water retention issues. Units with filters that remain wet after the sterilization may compromised sterility. | Device Design | On January 13, 2016 PMBS, LLC sent letters to their customers informing them of the correction. Per product's instructions for use, the SCORES cabinet is to be run for a minimun of a 4 minute sterilization cycle, followed by a 30 munites dry time followed by a 30 minutes door crack time. not all facilities generate completely dry SCORES filters following these instructions. It is imperative that the filters be checked in the Operation Room prior to useing the surgical instruments sterilised within the cabinet. If the filters are wet, do not use the contents. In the letter is a IFU addendum including instructions for running the equiptment with extended cycles. For more information please contact Ms. Michele Mauzerall, CEO at 908-300-7093. |  | 58 | 58 units | FL, GA, MN, NY, OH, OR SC, and TN | Open, Classified |  |  | 58
Z-1248-2016 | 72890 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Cardiosave Hybrid and rescue IABP |  |  | December 16, 2015 | 2015 | Maquet Datascope Corp - Cardiac Assist Division | Maquet has recieved information that in some Cardiosave IABPS, the scroll compressor did not meet the specifications for output pressure or vacuum at specific flow rates. | Other | Maquet sent an "Urgent - Medical Device Field Correction Notice and Response form dated December 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers are informed of the issues and provided corrective actions. Customers are asked to completed the attached response form in order to return to 1-973-629-1518, or via email to cardiosavescroll@maquet.com. Questions can be directed 1-800-777-4222, option 3, Monday through Friday 8 am-6 pm. EST. |  | 2350 | 1243 units US; 1107 units OUS | Worldwide Distribution- Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 2350
Z-1265-2016 | 73400 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard XP Knee System Porous Femoral Component |  |  | February 05, 2016 | 2016 | Biomet, Inc. | The anterior/posterior (AP) dimension may be oversized by 0.016 . If the femoral component does not fit as tightly as intended, bony ingrowth may not occur causing loosening of the femoral component. A revision surgery may be necessary. | Device Design | On 2/5/2016, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors, hospital risk managers, and implanting surgeons via courier. Distributors were instructed to locate and return the affected product, hospitals to locate and discontinue use, and implanting surgeons to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 890 | 890 | UTNetherlands | Terminated | March 28, 2017 | 417 | 890
Z-1249-2016 | 73494 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Reveal LINQ insertable Cardiac Monitor, Model LNQ11. |  |  | February 23, 2016 | 2016 | Medtronic Inc. | Medtronic has identified an issue with the sensitivity of an algorithm used in the Reveal LINQ ICM that may prematurely trigger the Recommended Replacement Time (RRT) alert in some devices. Battery capacity is not affected and the device will continue to support data collection and manual data transmissions. As part of the normal behavior of the device, 30 days after RRT status is reached, Reveal | Software design | Consignees in US were hand delivered a Medtronic "Medical Device Correction" letter dated February 2016. The letter was addressed to Physician or Healthcare Professional. The letter described the Issue &amp; Patient Management Guidance. For questions contact your local Medtronic Representative or Medtronic Diagnostic Technical Services at 800-929-4043.Outside the U.S., the letter will be mailed or hand delivered to impacted consignees. |  | 189009 | 189,009 | Worldwide Distribution-US (nationwide) including DC and the countries of Andorra, Australia, Austria, Belgium, Botswana, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic,Denmark, Faroe Islands , Finland, France, Germany, Greece, Hong Kong,Hungary, Iceland, Indonesia, Iran, Ireland, Israel, Italy, Kazakhstan, Kenya,Kuwait, Latvia, Lithuania, Luxembourg, Malaysia, Malta, Martinique, Namibia,Netherlands, New Caledonia, New Zealand, Norway, Oman, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore,Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tanzania,Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 189009
Z-1253-2016 | 73402 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | 10mm X 12o Solitaire Ti Medium Spacer  Solitaire Anterior Spinal System, Model Number 1400-1230. |  |  | February 12, 2016 | 2016 | Biomet Spine, LLC | Zimmer Biomet Spine initiated a recall of certain Solitaire Ti Spacers because they may have a mislabeled lordotic angle of 12degrees when the angle is actually 6 degrees. | Mixed-up of materials/components | Zimmer Biomet sent an Urgent Medical Device Recall Notice letter dated February 26, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the enclosed Response Form. For questions call 303-465-8929. |  | 10 | 10 | Worldwide Distribution - US Nationwide - one foreign consignee (not Canada). | Open, Classified |  |  | 10
Z-1264-2016 | 73543 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clear Vue Prone Breast Treatment Table and Access ClearVision Prone Breast Device |  |  | January 07, 2016 | 2016 | Wfr/aquaplast Corp | Some of the Access ClearVue Prone Breast devices may contain scales that are misaligned from the zero reference with respect to other Access ClearVue Prone Breast devices. This can result in differences in scale orientation. | Process control | Qfix sent a Field Safety Notification letter dated January 7, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached Inventory Survey form and mail it back. For questions call (610) 268-0585. |  | 216 | 18 units of RT-4545CV-17 and 12 units of RT-4545CV-20 | Worldwide Distribution - US Nationwide in the states of NY, PA, CA, IL, and TX. Foreign Countries: Singapore, Australia, Saudi Arabia, and Spain. | Terminated | January 06, 2017 | 365 | 216
Z-1250-2016 | 73533 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 |  |  | March 14, 2016 | 2016 | Terumo Cardiovascular Systems Corporation | Specific CDIÂż Blood Parameter Monitoring System 500 devices are being voluntarily recalled because the BPM Sensor Head Assemblys Thermistor, which provides the blood temperature value that results in accurate display values on the monitor, does not meet specification. This may cause inaccurate temperature measurement and inaccurate analyte display values on the CDI System 500 monitor.Terumo Cardiovascular Systems (Terumo CVS) received complaints of inaccurate temperature measurements for specific devices distributed since November 5, 2015. These include both new production devices and recently serviced devices.POTENTIAL HAZARDThere have been no reported illnesses or injuries as a result of this issue.A user who is not aware that the CDI System 500 is displaying inaccurate temperature values may not manage patient temperature appropriately. Using inaccurate information to manage warming and cooling strategies for a procedure could result in prolonged time on bypass. It could also lead to unnecessarily aggressive temperature management, resulting in excessive hyper or hypothermia, with potential neurologic and organ dysfunction, or increased blood component damage.Inaccurate temperature measurement could also cause inaccurate measurements of other BPM values including potassium (K+), pO2, pCO2, and pH. The greater the temperature inaccuracy, the greater the degree of inaccuracy of these other BPM values due to the dependence of their algorithms on temperature for the calculations. Inaccurate measurement of these values could result in inappropriate patient management strategies being employed to address them with the potential to result in moderate patient injury. | Nonconforming Material/Component | Terumo sent an " URGENT MEDICAL DEVICE RECALL" notification dated March 14, 2016, to all affected customers via mail. The letter identified the product the problem and the action needed to be taken by the customer. The Recall Notification has instructions for responding to the formal recall notification. If you have questions, contact Terumo CVS Customer Service: 800.521.2818 Customer Service Hours: Monday  Friday 8 a.m.  6 p.m. ET |  | 222 | 222 | Worldwide Distribution - Nationwide Distribution and to the countries of : Australia, Belgium, Canada, Japan , Singapore and United Arab Emirates. | Terminated | March 28, 2017 | 379 | 222
Z-1251-2016 | 73324 | No | MHW | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | IMCO brand Conforming Stretch Gauze Bandages, 4 x 4.1 Yds/10 cm x 3,7 cm, Reorder No. 704-IMC, Bar code 6 86864 04351 5, 96 per case (8 boxes x 12 sterile bandages), Single Use |  |  | February 24, 2016 | 2016 | Amd Ritmed | Product does not meet sterility requirements based on FDA analysis and tested positive for Bacterium Bacillus. | Process control | AMD Ritmed sent an Urgent - Medical Device Recall letter dated February 25, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are asked to examine all stocks immediately to determine if any of the affected product is on hand. If so, please discontinue distributing the product immediately and promptly return.If customers further distributed any of the product, please immediately contact your accounts advise them of the recall situation, and have them return their outstanding recalled stocks to you. Return these stocks as indicated above. Please return the enclosed response form immediately providing the requested information by fax 1-866-871-3455 or by email: AMDCustomerservice@amdritmed.com. Customers with questions can call 1-800-445-0340 or email AMD Customer Service: AMDCustomerservice@amdritmed.com. |  | 126 | 126 cases total (87 cases domestically and 39 cases internationally) | US Distribution including Puerto Rico and to the states of : DE, FL, KY, MI, NE and NJ., and Internationally to the country of Canada. | Open, Classified |  |  | 126
Z-1262-2016 | 73504 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The E.CAM gamma camera system |  |  | February 26, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | We received a customer complaint that Siemens reported via the MedWatch process (1423253-2015-00002). The result of a Hazard Analysis and CAPA investigation, we have determined that for e.cam or Symbia E systems that although the existing user manual is correct, additional supplemental instructions, with the inclusion of visual aids, may further clarify and emphasize the proper collimator change | Labeling design | Siemens sent a Customer Advisory Notice CAN 001 -2016 to all affected customers describing the recall and additional actions to take.Siemens updated user instructions, drafting customer advisory letter, and obtaining internal and external approvals, and translation of this content. Customer Advisory Notice Letter provided to each impacted consignee will document receipt of the letter and additional attempts will be made until such receipt is acknowledged. Instructions provided in letter explain that Hospital staff should inform end users of the enhancements that have been made to the user manual and continue to use your system. It also instruct end user that the advisory notice and addendum instructions should be placed with the system's instructions for use. The system use information should be available to all operators of the system.Siemens advises that if the end user has sold their equipment and it is no longer in their possession, They kindly ask that the end user forward the notice to the new owner of this equipment. They should inform Siemens about the new owner of the equipment.Adverse events or quality problems experienced with the use of this product should be reported to Siemens through the contact information provided below and may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. If the end user has any questions regarding the important notice, they should contact their local Service representative at the contact numbers provided below.-America: 1-800-888-7436-Europe, Middle East, and Africa: +49 9131 940 4000-Asia and Australia: +86 (21) 3811 2121 |  | 3090 | 3090 units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Puerto Rico.***** Foreign Consignees: Albania, Algeria, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Botswana, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Georgia, Germany, Ghana, Greece, Hong Kong, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kosovo, Kuwait, Libya, Malaysia, Mauritius, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Nigeria, Norway, Oman, P.R. China, Pakistan, Palestinian Ter, Peru, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad,Tobago, Tunisia, Turkey, U.A.E., Ukraine, United Kingdom, Uruguay and Vietnam. | Open, Classified |  |  | 3816
Z-1252-2016 | 73386 | No | KYF | Aqueous shunt. | Ophthalmic | Ophthalmic | 510(k) | Molteno3 Glaucoma Implants G-Series: |  |  | February 22, 2016 | 2016 | Innovative Ophthalmic Products, Inc | Innovative Ophthalmic Products (IOP) is recalling the Molteno 3 G-Series Glaucoma Implant the package insert/Instructions for Use (IFU) for the device has been updated to add a caution statement to minimize off-label use. | Labeling False and Misleading | Innovative Ophthalmic Products, Inc ( IOP ) sent an Customer Notification letter dated February 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken. The letter informs the customers that IOP is recalling because the package insert/Instructions for Use (IFU) for the device has been updated to add a caution statement. Specifically, the new IFU caution statement is "Caution should be taken with supra-Tenon implantation of this device as erosion of the conjunctiva is known to result from the raised edge of the oval pressure ridge."For further questions please call (202) 857-6464 |  | 278 | 278 units | Nationwide Distribution | Open, Classified |  |  | 278
Z-1268-2016 | 73460 | No | EAZ | Dental operating light. | Dental | Dental | 510(K) Exempt | Helios HL3T and HLT Track Mounted Dental Lights |  |  | March 08, 2016 | 2016 | Pelton &amp; Crane / Marus / DCI Equipment / KaVo | The firm received two customer complaints local Pelton &amp; Crane distributors that the Helios 3000 track mounted dental lights were missing the two 8/32x3/16 set screws during the initial installation of the dental light. | Nonconforming Material/Component | Pelton &amp; Crane sent an Urgent Medical Device Recall letters to affected customers on March 8, 2016 and will be completed by 3/11/2016. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Field Correction Acknowledgement Form in the enclosed self-address stamped envelope to Pelton &amp; Crane. For questions or concerns regarding this recall please call 1-800-659-6560, ext. 7227. |  | 156 | 156 units | US Nationwide Distribution in the states of AZ, AR, CA, CO, CT, FL, IA, ID, IL, IN, MA, MI, MN, NJ, NY, NC, OH, OR, PA, TN, TX, UT, VA, WI and Canada | Open, Classified |  |  | 156
Z-1269-2016 | 73417 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MicroScan LabPro Information Manager System, Software Versions 1.0 to 4.41 |  |  | February 11, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter is recalling the MicroScan LabPro Information Manager System because the software incorrectly allows the operator to manually edit the carbohydrate substrates when manually reading dried overnight gram negative panels with an ID Hold status. | Unknown/Undetermined by firm | An Urgent Medical Device Recall Letter dated 2/11/2016 was sent to all customers who purchased the MicroScan LabPro Information Manager System, Version 1.0 to Version 4.41. Beckman Coulter is expanding the recall that was initiated in July 2015 regarding FA-25776 for the MicroScan LabPro Information Manager System. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Customer Support Representative: http:www.beckmancoulter.com, (800) 677-7226 in US and Canada. Customers outside of US and Canada, are instructed to contact their local Beckman Coulter representative. |  | 2702 | 2,702 units total (1,032 units in US) | Affected geographies: Bahamas, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, French Guiana, Germany, Ghana, Greece, Guadeloupe, Guatemala, India, Israel, Italy, Japan, Kazakhstan, Lithuania, Mexico, Morocco, Netherlands, Northern Mariana Islands, Panama, Paraguay, Peru, Poland, Portugal, Puerto Rico, Romania, Russia, Singapore, South Africa, Spain, Taiwan, Tunisia, Turkey, United States, Uruguay, Virgin Islands, Zimbabwe. | Open, Classified |  |  | 2702
Z-1267-2016 | 73448 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | The Spectra-System Dental Implant 2008 system is comprised of dental implant fixtures and prosthetic devices that compose a two-piece implant system. |  |  | March 04, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | Implant Direct Sybron Manufacturing, LLC is recalling GoDirect Implants because some consignees did not receive the correct tool to drive the implant to bone level. | Nonconforming Material/Component | Implant Direct sent an Urgent: Medical Device Field Corrective Action letter dated March 4, 2016, to all affected customers. The letter states that Implant Direct will correct the issue and send the proper tool. The proper tool is included with the notification. The letter requests that customers complete and return the Acknowledgement and Recall Return Form within 48 hours.It customers are an authorized Implant Direct Sybron Manufacturing distributor it is requested that customers are identified and contact them to inform them of the issue within 48 hours of receiving the notification. Distributors should also provide customers with the correct tool.If you have any of the affected product listed above and have any questions contact Implant Direct Sybron Manufacturing LLC Customer Care at 1-888-649-6425.For further questions regarding this please call (818) 444-3300 Ext. 3323 |  | 2033 | 2033 | Worldwide Distribution - Nationwide Distribution to the states of : FL, AZ, IL, MS, TX, CA, ID, AL, MO, MN, NY, CO, GA, ME, NJ, NC, NV, LA, KS, MN, KY, PA, UT, OH, MI, SC, MD, WA, WI., and to the countries of : Canada, Honduras and Dominican Republic. | Open, Classified |  |  | 2033
Z-1271-2016 | 73426 | No | LKK |  |  | General Hospital | PMA | Medtronic, Model 8781 Ascenda" Intrathecal Catheter kit, Rx Only, Sterile Ethylene Oxide. |  |  | February 09, 2016 | 2016 | Medtronic Neuromodulation | Medtronic Neuromodulation is initiating a voluntary removal of a single Model 8781 Ascenda" Intrathecal Catheter kit. We have identified that this kit may have been produced with incorrect outer package labels. As a result, it may not contain all components necessary to complete the implant procedure. | Under Investigation by firm | Medtronic sent an "Medical Device Removal" letter dated February 2016. The letter described the problem and the product involved in the recall. The Action Required was to segregate the unused affected product. Their Medtronic Representative will be contacting consignee. For questions contact your Medtronic representative, or call for further questions. (800) 707-0933 |  | 9 | 9 (1 still unused) | Internationally to Japan. | Terminated | August 11, 2016 | 184 | 9
Z-1279-2016 | 69339 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Foot Switches used with Philips MultiDiagnost Eleva with Flat Detector; Product Codes: 708037, 708038 |  |  | September 24, 2014 | 2014 | Philips Medical Systems, Inc. | Loss of key image functionality due to a bent pedal of the Footswitch. | Device Design | Field Safety Noticed was mailed on September 24, 2014 via Certified Return Receipt. This FCC consists of a Mechanical Hardware update of the Footswitch in order to prevent bent pedals and therefore prevents an intermittent or continuous inability of making live fluoro images or exposures.Customers with questions were instructed to contact their local Philips representative. |  | 168 | 168 | Worldwide Distribution - US (nationwide) and Internationally to:CZECH REPUBLICALBANIAALGERIAARGENTINAAUSTRALIAAUSTRIAAZARBAIJANBAHAMASBANGLADESHBELGIUMBRAZILBULGARIACAYMAN ISLANDSCHILECHINACOLUMBIACOSTA RICACROATIACUBACZECHOSLOVAKIADENMARKDOMINICAN REPUBECUADOREGYPTESTONIAFINLANDFRANCEGABONGERMANYGREECEHONG KONGHUNGARYICELANDINDIAINDONESIAIRANIRAQIRELANDISRAELITALYJAPANJORDANKAZAKHSTANKENYAKOREA SOUTHKUWAITLATVIALEBANONLIBYALITHUANIAMACEDONIAMALAYSIAMAURITIUSMEXICOMOLDAVIAMONGOLIAMOROCCONEPALNETHERLANDSNETHERLANDS ANTNEW ZEALANDNORWAYOMANPAKISTANPALESTINEPANAMAPapua New GuineaPARAGUAYPERUPHILIPPINESPOLANDPORTUGALPUORTO RICOQATARREUNIONROMANIARUSSIAN FEDERATSAUDI ARABIASERBIASINGAPORESLOVENIASOUTH AFRICASPAINSUDANSWEDENSWITZERLANDSYRIATAIWANTANZANIATHAILANDTUNISIATURKEYUKRAINEUNITED ARAB EMIUnited Arab EmiratesUNITED KINGDOMURUGUAYUSAUZBEKISTANVENEZUELAVIETNAMYEMENGUADELOUPESYRIAN ARAB REPUBLIC | Open, Classified |  |  | 4348
Z-1276-2016 | 73411 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Estradiol III; Elecsys and cobas e analyzers; REF 06656021 190; M 6.5 mL, R1 9 mL, R2 9 mL, |  |  | March 07, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Due to the risk of a recently identified cross reactivity, the Elecsys Estradiol assay should not be used when monitoring patient being treated with fulvestrant. Customers should be informed that fulvestrant will increase the apparent concentration of estradiol in women being treated with this drug. If treatment with fulvestrant has been altered or discontinued as a result of falsely elevated e | Device Design | Roche sent an Urgent Medical Device Correction letter dated March 7, 2016 via UPS Ground (receipt required) to the consignee list. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to complete the attached fax form and fax it to 1-866-891-1984. For questions contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336. |  | 23630 | 23,630 | US Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 125942
Z-1308-2016 | 73436 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | VERO MHI-TM2000 Linear Accelerator System --- Common Name: Medical Linear Accelerator |  |  | March 07, 2016 | 2016 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | In rare cases the VERO/MHI-TM2000 Operator Console could set an incorrect Gantry angle for the first beam of an IMRT treatment if certain specific conditions are met. If Gantry angle for the first beam of IMRT is incorrectly set and if nobody is aware of it, an erroneous treatment would be given to the patient. | Software design | A Field Safety Notice dated March 2, 2016, was sent to all affected customers via e-mail on March 7, 2016. The Field Safety Notice informed users of the problem and provided a safe workaround until corrected software is available. The correction will be free of charge and a Customer Service Representative will contact all affected customers to schedule its installation. Customers with questions were advised to contact their Customer Service Representative. |  | 29 | 25 Worldwide, US: 4 units | Nationwide Distribution including New York, Texas, Florida, and Ohio | Terminated | April 20, 2017 | 409 | 29
Z-1303-2016 | 73477 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens branded SOMATOM Emotion and SOMATOM Sensation families of CT systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles. |  |  | February 15, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is initiating this recall due to a potential malfunction and hence potential hazard to patients and personnel when using the SOMATOM Emotion or SOMATOM Sensation systems in conjunction with the use of the cardio licenses "Cardio Pro" and "Cardio Retro". | Device Design | A Safety Advisory Notice was sent to customers informing them of a potential malfunction and hence potential hazard to patients and personnel when using the SOMATOM Emotion or SOMATOM Sensation systems. Siemens informed them that they will deliver an Update Instruction for their Service Engineers to modify the installed Firmware. This will correct the rare synchronization errors of Scanner Firmware and Scanner Software to start the application of the x-ray in compliance with the optimum triggering point for best image quality. |  | 124 | 124 units | US (nationwide) including Guam and the states of AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, KY, MA, MD, ME, MI, MN, MO, NC, NJ, NM, NY, OH, PA, SC, TN, TX, VA, and WA. | Open, Classified |  |  | 124
Z-1278-2016 | 73605 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Cios Alpha, mobile X-ray system |  |  | February 19, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software issues on Cios Alpha mobile C-Arm system | Software in the Use Environment | Siemens sent an Customer Safety Advisory Notice on February 26, 2016, to inform them of some issues that may arise on the Cios Alpha unit. The letter identified the product the problem and the action needed to be taken by the customer. Users were instructed to change the date format to DDMMYYYY to prevent issues, not to release radiation as long as the DVD icon is blinking and to disconnect and re-connect the trolley if the system is needed immediately. In the meantime they are preparing a modification of Cios Alpha that will resolve the above-mentioned issues. As soon as the modification is available, a service engineer will get in touch with each customer and implement the modification to the Cios Alpha system free of charge. If the device was sold, customers were instructed to forward the recall notification to the new owner.For further questions please call (610) 448-6478. |  | 89 | 89 systems | Nationwide Distribution to the states of : AL, AR, CA, CO, CT, DE, FL, GA, IA, IL, KY, MD, MI, MN, MO, NC, NE, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, WA and WI | Terminated | March 20, 2017 | 395 | 89
Z-1299-2016 | 73548 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | system, x-ray, angiographic, model # 10094135, 10094137, 10094139, 10094141, 10280959 |  |  | February 26, 2016 | 2016 | Siemens Medical Solutions USA, Inc | In Artis zeego systems, angulations in the vicinity of the C-arm collision area can cause the cable inlet to become clamped and can result in mechanical damage. In Artis systems with A100Plus generators, an attempt to resume operation following detection of a fault (such as a short circuit in the X-ray tube) can result in the failure of a module in the high-voltage generator. | Device Design | The firm, Siemens, sent an "Important Customer Safety Notice" dated 2/16/2016 to its customers. The notice informed them of a corrective action that will be performed to prevent a possible hazard to patients, operators, other persons or equipment. The notice described the product, problem and actions to be taken. Siemens plans to contact each customer to arrange a date to add additional protection mechanisms by means of a system software update.If you have any questions, contact Engineer, Regulatory Technical Specialist at 610-448-6461 or email: Meredith.adams@siemens.com. |  | 552 | 552 units | US Nationwide Distribution including states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, GU, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and PUERTO RICO. | Terminated | June 21, 2017 | 481 | 552
Z-1309-2016 | 73550 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MRIdian ViewRay Radiation Therapy System, ViewRay Treatment Planning and Delivery System (also known as the MRIdianÂż System) is indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body when radiation treatment is indicated |  |  | March 07, 2016 | 2016 | Viewray Incorporated | When editing the isocenter or the couch position of the plan while in the treatment workflow (in the Points screen) and re-optimizing, the software will not prompt the user to shift the couch to the new isocenter. Asa result there is the potential to deliver dose to the initial isocenter rather than the new location. | Software design | ViewRay sent an Urgent Device Correction letter dated March 7, 2016. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to contact ViewRay Customer Support at support@viewray.com or call the ViewRay Customer Support center at 855-286-8875. |  | 5 | 5 units | Worldwide Distribution - US including CA, FL, MO &amp; WI., and International distribution to the Republic of Korea and The Netherlands. | Open, Classified |  |  | 5
Z-1311-2016 | 73589 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/zeego (system, x-ray, angiographic) for single and biplane diagnostic imaging and interventional procedures such as cardiac and neuro angiography, general and rotational angiography, operating room angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | February 19, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Possible leak in the cooling system, fluids could potentially seep into the equipment cabinet. This leak can occur sporadically in affected systems. | Other | Siemens mailed out an Important Customer Safety Notice to customers informing them of a corrective action that will be performed to prevent a possible leak in the detector cooling system of Axiom Artis and Artis zee systems. The notice stated the affected plastic couplings will be replaced with metal couplings to prevent fluid leaks. Siemen's service organization will contact each customer to arrange a date to perform the corrective action. Customers asked to call 1-800-888-7436 to schedule an appointment for the corrective action to be implemented. |  | 1757 | 1757 systems | Nationwide distribution: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and Puerto Rico | Open, Classified |  |  | 1757
Z-1282-2016 | 73606 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis One, Interventional, Fluoroscopic, x-ray system |  |  | March 11, 2016 | 2016 | Siemens Medical Solutions USA, Inc | After importing segmentation results of the left atrium created on the Artis One system, the possibility exists for the results to be mirrored to an electro-anatomical 3D mapping system, CARTO from Biosense Webster Inc. | Software in the Use Environment | Siemens mailed an Important Customer Safety Notice to customers to inform them of the issue and the corrective actions they plan to take to address the problem. The letter stated that Siemens service organization will contact customers to arrange a date for the installation of the software update. Customers were instructed to call Siemens service organization for an appointment at 1-800-888-7436. |  | 1 | 1 angiography system | Nationwide Distribution to Michigan only | Open, Classified |  |  | 1
Z-1281-2016 | 73546 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | CIOS ALPHA (Image-intensified fluoroscope x-ray system) |  |  | February 16, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens found a defective part in their production line which could increase risk of an electrical shock for service personnel when working inside the device during system maintenance. The risk of an electrical shock is only present when a service technician ignores several steps listed in the service manual as precautionary actions during device maintenance. | Nonconforming Material/Component | Siemens sent a Customer Safety Advisory Notice dated February 18, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that Siemens is planning to replace the affected parts as a precautionary measure. A Siemens service engineer will contact each customer to schedule implementation of this precautionary measure. |  | 4 | 4 units | US Nationwide Distributed in the states of AR, CA, NE and UT | Open, Classified |  |  | 4
Z-1283-2016 | 73672 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis zee/zeego, system, x-ray, angiographic |  |  | February 18, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Due to a leak in the cooling system, fluids could potentially seep into the equipment cabinet. Loss of coolant can result in a cooling system failure and subsequent system failure. This leak can occur sporadically in affected systems. | Unknown/Undetermined by firm | Siemens sent an Important Customer Safety Notice to customers on February 18, 2016 informing them that their service organization will contact them to arrange a date to perform the corrective action (replacement of affected plastic couplings with metal couplings to prevent fluid leaks). Customers were advised to contact Siemens service organization for an earlier appointment at 1-800-888-7436. |  | 986 | 986 systems in total | US Nationwide Distribution in the states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, GU, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY including Puerto Rico. | Terminated | August 08, 2017 | 537 | 986
Z-1289-2016 | 73380 | orthosis, limb brace | IQI | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | Rebound Cartilage |  |  | February 24, 2016 | 2016 | Ossur H / F | Ossur Americas is recalling , Knee Shield / Patella Cup, Knee Undersleeve Gray, NecLoc Extrication Collar, Rebound Cartilage, Rebound Cartilage Custom, Rebound Air Walker, Rebound Diabetic Walker, Back Support Suspenders Clip on Black, Back Support Suspenders Sew On Black, Extreme Custom, Unloader Custom XT, Unloader Custom Lite, Unloader Custom, and Papoose Infant Spine Immobilizers because the p | Labeling False and Misleading | The firm sent out customer notification letters on 02/26/16.The notification letter instructed customers to include the latex caution notice in any existing inventory of affected products. The latex caution notice can be retrieved at the conclusion of the response survey or by going directly tosafetyalerts.ossur.com. Should customers require Âżssur to provide pre-printed notices for existing inventory, contact Âżssur customer service. For the Âżssur Miami JÂż Cervical Collars with lot numbers before MX150727 customers are instructed to remove and discard the assist strap, which contains natural rubber latex, located in the patient pack.Recipients of the notice with should take the following actions:1. Pass the notice to those who need to be aware within the organization or to any organization where the affected devices have been transferred.2. Maintain awareness on this notification for an appropriate period.3. If have further distributed the products, identify customers and notify them at once of the product alert. It is recommended to include a copy of the alert notification letter.4. If any of your customers are currently wearing a product, check for any skin irritations or other reactions consistent with natural rubber latex allergies.5. To verify have received the notification, visit safetyalerts.ossur.com to complete the Response Survey.If any questions, contact Âżssur Customer Service at (800) 233-6263.The notification letter for the Papoose Infant Spine Immobilizer contained the following additional actions/instructions:Examine inventory, quarantine products subject to the recall and contact customer service for a return authorization to return the product to Âżssur.If any of customers are currently wearing a product, it is recommended replacing the unit with one shipped after the transition lot number MX150727. Check for any skin irritations or other reactions consistent with natural rubber latex allergies. |  | 1030 | 1030 | U.S. distribution to the following; AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, RI, TN, TX, UT, VA, VT, WA, WI, WV | Open, Classified |  |  | 27899
Z-1312-2016 | 73600 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Interlocking Bolt, 151 mm (General and Plastics Surgery Devices) Guide, Surgical Instrument |  |  | February 16, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is initiating a voluntary medical device recall of certain lots of the Interlocking Bolt which are a part of the Synthes 4.5 mm LCP Proximal Tibia Plate System. The Synthes 4.5 mm LCP Proximal Tibia Plate are indicated for the treatment of nonunions, malunions and fractures of the proximal tibia including: simple comminuted, lateral wedge, depression, medial wedge, bicondylar, combi | Device Design | DePuy Synthes sent an Urgent Notice to inform customers that they were initiating a voluntary recall of lot numbers: 1776197, 5706588, 1782340 and 1868678 of the Interlocking Bolt which are part of the Synthes 4.5mm LCP Proximal Tibia Plate System. Customers were asked to review their inventory for the affected lots, call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, complete the Verifcation Section (page 3 of the letter) identifying that the affected product has been located and return it to Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. A copy of the Verification Section of the letter was to be faxed to 866-737-4701 or emailed to: Synthes2856@stericycle.com. |  | 42 | 42 plate systems | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Open, Classified |  |  | 42
Z-1313-2016 | 73325 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | NEL GBS Medium 20 ml in 100x15mm plate (Standard Petri Dish) |  |  | December 21, 2015 | 2015 | Northeast Laboratory Services, Inc. | Product marketed without a 510 (k) | No Marketing Application | Northeast Laboratory services issued recall letter dated 2/29/16and advised users to: Immediately examine your inventory for all lots of these products you may have in stock. Please destroy all remaining product in stock. Users requested to fill out the questionnaire immediately and return the questionnaire via email or Fax.Previously Northeast Laboratory Services notified users by letter dated 12/21/15 that the firm has discontinued the products effective January 2016. Alternative products were referenced in the letter for users to transition to replacement product. |  | 9300 | 9300 | Nationwide Distribution. | Terminated | April 07, 2016 | 108 | 42040
Z-1315-2016 | 73381 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | TeleSentry Battery, 3.7 Li-Ion Polymer Battery, CAUTION Do not expose to high temperature. Do not disassemble. Dispose of properly. Use specified charger only. 60pC/140pF The ScottCare Corporation. |  |  | February 25, 2016 | 2016 | Scottcare Corporation | The firm received a complaint of the battery exploding while recharging. | Device Design | On 2/26/2016 the firm sent "Urgent: TeleSentry Battery Recall" notification letters to their customers. The letter identified the affected device and the reason for the recall. Customers were asked to immediately examine their inventory and quarantine the affected device. If the product had been further distributed, then a copy of the letter should be provided to those customers. Customers are to follow the steps provided for returning the affected device, including contacting ScottCare Service at 1-800-243-9412. In addition, the attached response form should be completed and returned. Questions should be directed to ScottCare Customer Service at 1-800-243-9412. |  | 9067 | 9,067 batteries | Worldwide Distribution -- USA, including the states of CA, CT, DE, FL, IL, KS, KY, MA, MD, MI, MN, MO, NC, NH, NJ, NM, NV, NY, OH, PA, SC, TN, TX, &amp; WV; and, the country of Canada. | Open, Classified |  |  | 9067
Z-1231-2016 | 73364 | No | RCE |  |  |  |  | GE Inspection Technologies, LP, Cabinet X-ray system used for industrial quality control. |  |  | December 16, 2015 | 2015 | GE Inspection Technologies, LP | GE Inspection Technologies, LP ("GE") has discovered that certain of its cabinet x-ray systems may not fully comply with the Food and Drug Administration's (FDA) Cabinet XÂż ray System performance standards, 21 C.F.R. Âż 1020.40(c)(1)(i) | Vendor change control | GE Inspection Technologies Planned Action to bring into Compliance:1. GE Inspection Technologies will use their updated adhesive procedure to install new brackets to the bottom left and right lead strips in the control door. The new procedure and brackets were developed in November 2013.2. A second notification of customers with affected systems will be sent within5 days of approval of this CAP. This notification will describe the repair above and propose that a GE technician visit each site and install the described repair.3. Repairs will be completed within 120 days after approval of this CAP and will be made free of charge. For further questions, please call (717) 447-1278.; |  | 55 | 55 US Distribution | US Distribution to the states of : TX, MD, MI, CA, TN, NH, NY, AL, NC, ID, MA, AZ, IN, FL, PA, NJ, WI, VA, MN, MD CO, GA, SC, OH and DC. | Open, Classified |  |  | 55
Z-1274-2016 | 73283 | No | PGX |  | Microbiology | Microbiology | 510(k) | Simplexa Group A Strep Direct (MOL2850) with Direct Amplification Disc (MOL1451, MOL1452, MOL1455) |  |  | February 10, 2016 | 2016 | Focus Diagnostics Inc | Focus Diagnostics is recalling the Direct Amplification Discs (DAD) supplied with the Simplexa HSV 1 &amp; 2 Direct (MOL2150), Simplexa Flu A/B &amp; RSV Direct (MOL2650) and Simplexa Group A Strep Direct kits (MOL2850) because it may result in a false negative result, a false positive result or Error code (s) 500 or 505, and it may result in higher invalid rates. | Under Investigation by firm | A customer notification letter dated 2/10/16 was sent to all their customers who received the Direct Amplification Discs (DAD) supplied with the Simplexa HSV 1 &amp; 2 Direct (MOL2150), Simplexa Flu A/B &amp; RSV Direct (MOL2650) and Simplexa Group A Strep Direct kits (MOL2850). The letter informs the customers that Focus Diagnostics is recalling MOL2150 and MOL2650 kits because it may result in a false negative result, a false positive result or Error Code(s) 500 or 505. Additionally, some customers are reporting insufficient specimen volume errors which may result in higher invalid rates. In those cases, the error results in a Report Note Statement of Insufficient Specimen Volume . The letter informs the customers of the recommendations to be taken for the identified kits and the risks involved for positive results, negative results, and delayed results. The letter informs the customers of the actions to be taken. Customers are instructed to sign the achnowledgement form and email the form to Technicalinfo@focusdx.com or fax back to Focus Diagnostics Technical Services at (562) 240-6526 within 10 business days.On 2/11/16, Focus Diagnostics sent to their customers an amendment to the customer notification letter to clarify what is considered an "early Ct", Focus would consider Internal Control (IC) Ct's less than 25 as "early Ct". Customers are informed that if an internal control Ct is less than 25, additional testing may be necessary.On 2/15/16, Focus Diagnostics sent to their customers an amendment to the customer notification letter to inform them to determine if they had instances of early Cts less than 25, then to please refer to Section III  Analyzing Runs, pages 6-10 and 6-15 in the Focus Diagnostics, Operator Manual for Integrated Cycler Studio 6.0  For Use with IVD Assays in the USA. Also, for a more detailed analysis, customers are instructed to view excerpts from the ICS Operator Manual covering Section III, pages 6-10 and 6-16.On 3/17/16, |  | 41 | 41 units | Worldwide Distribution - US Nationwide in the states of: TX, CA, MA, NY, WA, MI, ME, IN, NJ, WV, OH, PA, HI, NH, CO, MN, FL, SC, IN, IA, TN, OR, AZ, WI, TN, IL. and the countries of: Australia, Brazil, Canada, Dominican Republic, Denmark, Ecuador, France, Germany, Israel, Kuwait, Spain, Sweden, Slovenia, Thailand, and United Kingdom. | Terminated | June 16, 2016 | 127 | 18614
Z-1316-2016 | 73281 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | evaluemed Infant Heel Warmer 301-1223; Distributed by Tri-anim |  |  | February 10, 2016 | 2016 | Philips Electronics North America Corporation | eValueMed Infant Heel Warmers and eValueMed Infant Transport Mattresses bear no expiration date, nor any warning that product used beyond its expected shelf life may warm to temperature above its specification. | Labeling design | Philips sent an Urgent Medical Device Recall leter dated January 2016. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following: Actions for Tri-Anim Customers to Take1.) Discontinue use of eValueMed Infant Heel Warmers and eValueMed Infant Transport Mattresses.2.) Check for any inventory of eValueMed Infant Heel Warmers or eValueMed Infant Transport Mattresses.3.) Destroy all products on hand (e.g., by activating or puncturing each unit). Make a note of the total number of cases of products that were destroyed as this will be needed for step 4 below.4.) Complete and sign the Customer Reply Form provided in this letter or Tri-anims version thereof, to notify Tri-anim of how many eValueMed Infant Heel Warmers and eValueMed Infant Transport Mattresses were destroyed. Submit your signed reply form only after you have completed all steps. Please do not return any product to Philips.Customers with questions should call 336-449-5230 or 262-798-9802.For questions regarding this recall call 978-659-3000. |  | 539.156 | 539.156 | Nationwide distribution to Ohio only. | Terminated | August 01, 2017 | 538 | 539.156
Z-1318-2016 | 73489 | No | KRR | Knee joint patellofemoral polymer/metal semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | All instrumentation associated with the Prelude PF Resurfacing Knee System. |  |  | March 09, 2016 | 2016 | Biomet, Inc. | Inadequate design control. | Device Design | Zimmer Biomet sent an "Urgent Medical Device Recall Notice" dated March 9, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The following actions are REQUIRED:1. Immediately locate and remove the identified device(s) listed below from circulation.2. Carefully follow the instructions on the enclosed Response Form .3.Email a copy of the response form to CPW ARFieldAction@zimmerbiomet.com prior to return of product.4. Use priority carrier for your shipment.Please confirm receipt of this notice by sending back the response form within three (3) business days. Thank you in advance for your assistance and prompt attention. On behalf of Zimmer Biomet, I apologize for any inconvenience this may cause. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 576 | 576 | US Distribution to the states of : CA, CO, GA, IN, KS, MI, MO, NC, NJ, OH, TN and TX. | Terminated | May 23, 2017 | 440 | 576
Z-1319-2016 | 73334 | apparatus, melting point, paraffin | IDT | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Richard-Allan Scientific Paraffin Type 3 |  |  | February 16, 2016 | 2016 | Richard-Allan Scientific Company | Multiple complaints were received from customers indicating that they ordered product 8335(Type 3 Paraffin) but when they opened the box, they found product 8336 (Type 6 Paraffin) inside. | Labeling mix-ups | ThermoFisher sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated February 15, 2016, to all affected customers. The letter contained the following instructions: If you have the above lot in your possession, please determine whether the bags are labeled as Type 3 (PN 8335) or Type 6 (PN 8336).Âż If the product is labeled as Type 3 (PN 8335) please complete the attached Response form with the quantity in inventory and return the form to Richard-Allan Scientific. The product may be used and no further action is required.Âż If the product is labeled as Type 6 (PN 8336) please complete the attached Response form with the quantity in inventory. The product may be used as is or returned to Richard-Allan Scientific. Please indicate on the Response form if the product will be used or a return is desired. Return the completed form to Richard-Allan Scientific.Âż If affected product has been used previously, please complete the attached Response form with the quantity that was used. No further action is required.By returning the attached Acknowledgment Form will facilitate reporting of this matter to the FDA. Customers with questions were instructed to call 269-544-5628. |  | 29 | 29 cases (10 bags per case) | Worldwide distribution to AZ , CA , FL , GA , IL , MI , MN , NM , NY , OH , TX , UT and Internationally to Germany and the United Kingdom.6 distributed to OUS accounts: Germany, and the United Kingdom | Open, Classified |  |  | 29
Z-1328-2016 | 73373 | No | GEA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MicroAire PALÂż Cannula |  |  | February 08, 2016 | 2016 | MicroAire Surgical Instruments, LLC | MicroAire has initiated a recall on product PAL-R4011XL (40cm x 4mm Del Vecchio French Rasp cannula), due to our receipt of three (3) complaints that it bent or broke during surgery.Accordingly, the PAL-R4011XL has also been put on production hold until further notice. | Nonconforming Material/Component | The firm is sending letters to all customers who have received PAL-R4011XL fromAugust 2013- October 2015. The firm sent these letters by UPS to expedite the delivery and for traceability. |  | 145 | 145 manufactured and distributed | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 145
Z-1325-2016 | 73487 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008 Series Hemodialysis Delivery Systems: 2008T, 2008K2, 2008K and 2008K@home: |  |  | March 10, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | Added Warning Statement: Fresenius 2008 Series Hemodialysis Machines False Blood Leak Alarm When Dialyzing Patients Treated with Hydroxocobalamin (or any form of Vitamin B-12) | Device Design | Fresenius Medical Care issued an Urgent Medical Device Alert dated March 10, 2016, to all affected consignees. Consignees were instructed to place notification and Warning Addendum with their 2008Âż Series Hemodialysis Machine Operators Manual. A reply form was included as part of the customer communication with instructions to fill out and return via fax or scanned/emailed copy.Further information or support concerning this issue, customers were instructed to contact Fresenius Technical Services at 800-227-2572. |  | 8384 | 8384 units (8369 US) (15 OUS) | Nationwide Distribution | Open, Classified |  |  | 8384
Z-1323-2016 | 73259 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MicroAire PAL-R4011XL - 40cm x 4 mm Del Vecchio French Rasp aspiration cannula |  |  | February 08, 2016 | 2016 | MicroAire Surgical Instruments, LLC | Complaints that the cannula bent/broke during surgery. | Component design/selection | The firm is sending letters to all customers who have received PAL-R4011XL from August 2013- October 2015. The firm sent these letters by UPS to expedite the delivery and for traceability. |  | 37 | 37 | Worldwide Distribution. US nationwide including Puerto Rico, Australia, Bangladesh, Czech Republic, India, Japan, Lebanon, New Zealand, Norway, Singapore, Switzerland, Taiwan, and Turkey. | Open, Classified |  |  | 37
Z-1324-2016 | 73387 | Yes | FJI | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Hemoflow F3 Dialyzer |  |  | February 26, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | Dialyzer header leak due to possible improper torque | Process control | The firm, Fresenius Medical Care Renal Therapies Group, LLC issued an "URGENT MEDICAL DEVICE RECALL" letter dated 2/24/16 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of your Hemoflow F3 Dialyzer; place all affected units in a secure, segregated area; contact your FMCRTG, LLC Customer Service Team at 1-800-323-5188 for instructions on how to return the recalled product, and fill out and return the attached Reply Form via fax to: 781-699-9796 ATTN: Quality Department or scan and email to: NOTIFYRA@fmc-na.com at your earliest convenience.If you have any additional medical concerns or questions, please contact Medical Information and Communication at 855-616-2309 or Website: www.fresenius-medinfo.com. |  | 1650 | 1,650 units | US Nationwide Distribution including states of: AZ, CA, CT, DC, GA, IL, IN, MI, MN, MO, MS, NY, OH, PA, SC, SD, TX, and WA. | Terminated | June 26, 2017 | 486 | 1650
Z-1320-2016 | 73497 | No | OHG | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | Enforcement Discretion | Gomco Style Circumcision Clamp Trays is an instrument used to compress the foreskin of the penis during circumcision of a male infant. |  |  | February 29, 2016 | 2016 | Cardinal Health | Two lots of Gomco Style Circumcision Clamp Trays, SSI-0034 and 10-1100 contained a 1.3cm circumcision clamp. They both were labelled as containing a 1.1cm clamp. | Labeling Change Control | Cardinal Health sent an URGENT: PRODUCT RECALL letter dated February 29, 2016, to affected consignees explaining the reason for the product recall and actions that should be taken by the end user. These included the following actions:1. STOP using affected product.2. RETURN the enclosed acknowledgment form via fax to 847-689-9101 or 614-652- 9648, whether or not you have affected product, as Cardinal Health is required to confirm receipt of this notification from you, and to prevent further notices.3. NOTIFY any customers to whom you may have distributed product affected by this recall.4. CONTACT the appropriate Customer Service group to arrange for return of any affected product you may have:" Hospital800.964.5227" Distributor800.635.6021 The firm asks that consignees who have question or requests of special assistance relating to the recall should contact Cardinal Health Customer Advocacy at 800.292.9332.For questions regarding this recall call 847-473-1500. |  | 440 | SSI-0034: 380 units and 10-1100: 60 units | Nationwide Distribution including AR, KS, WV, , NC, TX, AL, MS and VA. | Open, Classified |  |  | 440
Z-1332-2016 | 73456 | No | MCH | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | STD KIT GK260, Standard Set 360 Z |  |  | December 22, 2015 | 2015 | Applied Medical Resources Corp | Applied Medical Resources Corporation is recalling CA500 Epix Universal Clip Applier, kit models STD GK258 AME Standard Kit 44, STD KIT GK213, AME Standard Kit 15, and STD KIT GK260, Standard Set 360 Z, which contain the affected CA500 lot, because of a potential loading mechanism nonconformance. | Nonconforming Material/Component | Customer notification letters were sent on December 22, 2015 for the U.S. and on January 6, 2016 for those in Europe.The customer notification letter requested that the following actions be taken:- Check inventory for recalled product.- Complete the attached Recall Notification Confirmation Form to acknowledge therecall, indicate if facility is returning or has already used this lot, and provide a no-charge PO number if replacements requested.- Return the Recall Notification Confirmation Form to Applied Medical by emailing torecall60684384@appliedmedical.com or faxing to (949) 713-8936.- Return affected product and a copy of the Recall Notification Confirmation Form to Applied Medical For product return questions contact Karen Mitchell, Sales Operations Supervisor, RGA Dept. at (949) 713-8622 or by email at kmitchell@appliedmedical.com.For regulatory questions contact Lauren Contursi, at (949) 713-8767 or by email at lcontursi@appliedmedical.com. |  |  |  | U.S .Nationwide Distribution in the states of GA, AL, LA, TX, MN, AK, CA, ID, OR, UT, WY, NM, and AZ.Foreign distirbution to the following; Austria, Belgium, Denmarl, Canada, Japan, France, Germany, Ireland, Italy, Greece, Netherlands, Russia, Spain, Switzerland, Turkey, Great Britain | Terminated | October 03, 2016 | 286
Z-1322-2016 | 73012 | No | LNR |  |  | Gastroenterology/Urology | PMA | The THERAKOS CELLEX Photopheresis System is a photoimmune therapy. Catalog CLXUSA. |  |  | December 21, 2015 | 2015 | Therakos Inc | The Mallinckrodt (Therakos) post-market monitoring process detected an increased trend in tubing leaks within the pump tubing organizer, associated with a portion of Kits from Lot D 342. | Nonconforming Material/Component | Therakos sent an Urgent Medical Device Recall letter dated December 23, 2015, to all affected customers by UPS Overnight Mail. The letter identified the product, the problem, and the action to be taken by the customer. The consignees were instructed to immediately identify, discontinue use and return affected product. Customers with questions were instructed to call Therakos Customer Service at 1-855-422-9115, option 4.For questions regarding this recall call 908-238-6813. |  | 729 | 729 units | Nationwide Distribution to AL, AZ, CA, CO, DC, FL, ID, IL, IN, MA, ME, MN, MO, NC, NJ, NY, OH, PA, SC, TN, TX, UT, VT, AND WI. | Open, Classified |  |  | 729
Z-1326-2016 | 73561 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | Titanium Low Profile Neuro Screw, Self Drilling, 3mm, part numbers: 400.833, 400.833E |  |  | February 23, 2016 | 2016 | Synthes (USA) Products LLC | Out of specification at the thread. The thread height of these screws is under-sized. The cross section of the affected area is less than that of conforming screws such that the thread height of the screws is undersized. | Nonconforming Material/Component | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated February 23, 2016 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were notified by mail of a potential patient impact. They were instructed to review existing inventories and call DePuy Synthes at 1800-479-6329 to obtain a Return Authorization (RA) Number and replacement order. After completing the verification section on page 4 of the letter, they were advised to return the verification section (page 3) with the product to Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. A copy of the completed verification section is to sent to DePuy Synthes by: Fax 877-931-0643 or scan/email Synthes8097@stericycle.com. |  | 535 | 500 units distributed in USA, and 35 units distributed OUS | Worldwide Distribution - US Nationwide in the states of MA, VA, TN, GA, WA, IA, NJ, TN, PA, NY, CO, CT, VT, and the countries of Canada and Switzerland. | Open, Classified |  |  | 685
Z-1321-2016 | 73525 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TOPAZ MicroDebrider 0.8 mm REF AC4040-01; TOPAZ MicroDebrider IFS 0.8 mm REF ACH4040-01; TOPAZ MicroDebrider 0.8 mm REF Q6000-01; TOPAZ EPF MicroDebrider 0.8 mm 45 REF Q6002-01 |  |  | March 03, 2016 | 2016 | ArthroCare Corporation | During functional testing of devices after real time aging, it was determined that product failed to meet shelf-life at the expiration date indicated on the product labels. | Package design/selection | Consignees were notified via letter on 3/3/2016. Consignees were instructed to quarantine the recalled products and complete an Inventory Return Certification Form. |  | 26903 | 26,903 devices | Worldwide Distribution-US (nationwide) and the following countries-Austria, Belgium, Bulgaria, Canada, Denmark, Estonia, Finland, France, Germany, Gibraltar, Great Britain, Greece, Ireland, Israel, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Spain, Sweden, Switzerland, China, Hong Kong, India, Lebanon, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arabic Emirates, Vietnam, Australia, Japan, New Zealand, Brazil, and Chile. | Terminated | May 17, 2017 | 440 | 26903
Z-1341-2016 | 73498 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona The Personalized Knee System Articular Surface Posterior Stabilized (PS) Left Height 18mm Sterile For use in total knee arthroplasty. REF 42-5114-008-18 |  |  | March 17, 2016 | 2016 | Zimmer Manufacturing B.V. | The firm received a customer complaint establishing that the product (Persona PS Articular Surfaces)information shown on the box differed by sides. One side showed P/N: 42-5114-008-18/ Lot: 62632101 and the other side showed P/N: 42-5114-005-14/ Lot: 62646580. | Employee error | Distributors, Sales Representatives, Operation Managers, and Hospital Risk Managers will be provided with an "Urgent Medical Device Recall-Lot Specific" letter. |  |  | TBD | Distributed in US (AL, AZ, CA, GA, HI, ID, KS, LA, MA, MN, MO, NY, OH, OK, PA, TX, UT, VA, and WI) and foreign countries (Korea, Japan, Germany, Singapore, Austria, Italy, South Africa, and Spain). | Terminated | March 14, 2017 | 362
Z-1343-2016 | 72957 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | NAMIC CONVENIENCE KIT SELF REGIONAL HEALTHCARE PV LEG PACK, REF Catalog No. 651919311, UPN Product No. H7496519193111 |  |  | November 03, 2015 | 2015 | AngioDynamics Inc. | Procedure Products initiated the recall due to the inner lumen of the needles not meeting the intended specification. Specifically, at least a portion of the affected lot is unable to accept a 0.038" guidewire as intended. In some cases the inner lumen was found to be 0.036". | Process control | On November 03, 2015, Navilyst Medical distributed Urgent Voluntary Medical Device Recall notices to their customers via courier service. Customers were advised to immediately remove the recall product from inventory and to segregate the products in a secure location for return to AngioDynamics, Inc. Customers should immediately forward a copy of the recall notification letter to all sites if these products were further distributed. If affected product is located in your institution, please call AngioDynamics, Inc. Customer Service at 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday - Friday EST) to obtain a replacement or credit for your returned product. Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return). The Reply Verification Tracking form can be submitted via fax to: Attn: Convenience Kits - Procedure Products Needles 1-800-782-1357 or via email to rdenino@angiodynamics.com or sbaxter@angiodynamics.com. |  | 136 | 136 kits | US Nationwide in the states of CA, IL, MA, SC, TX | Terminated | March 16, 2017 | 499 | 289
Z-1348-2016 | 73429 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System, Unique Device Identifier No. 10758750002740 &amp; 10758750007110. |  |  | February 16, 2016 | 2016 | Ortho-Clinical Diagnostics | The two sets of wires (4 total wires) exiting from the EMI filters were switched, creating a situation where a module that is intended to be de-energized will remain energized only if the associated AC power cord is unplugged. | Process control | On February 16, 2016, Ortho Clinical Diagnostics distributed Urgent Product Correction Notification notices to their customers via FedEx overnight courier, e-communications and via USPS. Customers are advised that there is no impact to the results generated with an affected system. As a precaution, your Ortho-trained service representative will contact you to schedule an inspection of the wiring within the system at your facility. Following the inspection, reconfiguration of the wiring on your system will be performed if appropriate. Please store a copy of the notification with your user documentation until the wires within your system(s) are inspected. In addition, complete the receipt form to indicate that you have been informed of this anomaly. Please return the form to Joe Falvo via email at confirmationadmin@its.jnj.com or via fax to 1-888-557-3759 or 1-585-453-4110. Customers with questions can contact the Technical Solutions Center at 1-800-421-3311 anytime. |  | 2101 | 2,101 Total units (1,043 domestically &amp; 1,058 internationally) | Worldwide Distribution-US (nationwide) including Puerto Ricoand the countries of Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy &amp; Spain. | Open, Classified |  |  | 2101
Z-1275-2016 | 73553 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Guardian II Hemostasis Valve, models 8210 and 8211. The following models are only available Outside the US: FH101, FH101-T, FH101-25, and FH101-50. |  |  | March 04, 2016 | 2016 | Vascular Solutions, Inc. | Vascular Solutions became aware of a potential problem with the click version of the Guardian II hemostasis valves. The low pressure seal may not close properly, which may allow air to be introduced into the device and may lead to risk of an air embolism. | Employee error | Consignees were sent on 3/4/2016 via first class mail a Vascular Solutions "Urgent Medical Device Recall" letter dated March 3, 2016. The letter described the problem and the product involved in the recall. It included the "Immediate Action Required of Customer / User" which instructed the customer to identify the product, remove it from inventory, place it in a secure area, and complete and return the VSI Account Inventory Form to the Customer Service Department. |  | 5283 | 5283 | Nationwide Distribution in the states of AL, AR, CA, CO, DE, FL, GA, ID, IL IN, IA,KY, LA, MD, MA, MI, MO, NE,NV, NH, NY, NC, OH, PA, SD, TN, TX, UT, VA, WA, and WV.OUS: Ireland | Terminated | October 19, 2016 | 229 | 5283
Z-1349-2016 | 73475 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | A Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images, including digital mammography images. |  |  | January 18, 2016 | 2016 | Siemens Medical Solutions USA, Inc | To inform users about the possible incorrect values for Distance Measurements when using certain modalities in combination with syngo Imaging. | Software design | Siemens mailed a Customer Advisory Notice to their customers instructing them to check their modalities for the usage of the mentioned DICOM attributes with different values. They recommended upgrading the workplace(s) to syngo Imaging Studio version VB36D if they have such modalities and it cannot be ensured by the modality provider that only one attribute is filled or both attributes are filled with the same value. Customers were asked to contact their local service organization to clarify the compatibility of their syngo Imaging server version with workplace version VB36D. |  | 49 | 49 imaging systems | US (nationwide) including the states of AL, CA, CO, CT, FL, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OK, OR, PA, TN, TX, and WI. | Terminated | August 03, 2017 | 563 | 49
Z-1270-2016 | 73468 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Ellex Medical Fixed eye Safety filters Model Integre S LP5532. |  |  | February 01, 2016 | 2016 | Laserex Systems Inc. | The Firm discovered that there is a design error which is traceable back to 2012 when a design change on the safety filter optic adaptor was implemented. | Device Design | Ellex Medical Pty Ltd. Planned Action to bring into Compliance:1.Field Corrective Action replacing all affected safety filters of all units manufactured between July 2012 and February 2016, 2.All repairs will be made at no cost to the purchaser; and3.The corrections will be completed by March 30, 2017.CDRH approves the CAP subject to the following conditions:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and1003.22. For further questions please call (800) 824-7444. |  | 74 | 72 units plus 2 spare units | Worldwide Distribution - US Distribution | Open, Classified |  |  | 74
Z-1347-2016 | 73555 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SIGNA Pioneer MR Systems |  |  | March 01, 2016 | 2016 | GE Medical Systems, LLC | Ferrous material in the SIGNA Pioneer Table Control Box (TCB) could cause unexpected attraction to the MR system's high strengthmagnetic field during servicing activities. This could create a potential risk of serious injury to service personnel who attempt to inappropriately remove the TCB fromthe magnetic field. There have been no events reported as a result of this issue. Clinicalscanni | Labeling Change Control | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 68002 dated March 1, 2016. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions or concerns regarding this notification,contact GE Healthcare Service at 18004371171 or your local Service Representative. |  | 16 | 16 (2 US; 14 OUS) | Worldwide Distribution - US Nationwide in the states of SC, NC and the countries of Australia, Canada, France. Germany, Japan and Turkey. | Open, Classified |  |  | 16
Z-1357-2016 | 72922 | No | MJC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | BARDEX I.C. Anti-Infective 2-Way 3cc Foley Catheter, Rx only. Intended for urinary bladder drainage in female patients requiring catheterization for management of incontinence, voiding dysfunction and surgical procedure. |  |  | December 18, 2015 | 2015 | C.R. Bard, Inc. | Labeling error: Catalog contains Instructions for Use (IFU) error; incorrect inflation volume. | Error in labeling | C. R. Bard, Inc. sent a "NOTIFICATION OF MISLABELING" letter, dated 12/18/2015 to all affected consignees. The letter identifies the product, problem, and actions to be taken by the customers. |  | 1428000 | 1,428,000 units (total) | Distributed US (nationwide) and the countries of Canada and Australia. | Open, Classified |  |  | 1428000
Z-1360-2016 | 73579 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Arthrex Deltoid Ligament Reconstruction Implant System. |  |  | January 11, 2016 | 2016 | Arthrex, Inc. | Potential for a component contained in the Implant System to be non-sterile. | Other | On 01/22/2016, all distributors and end user facilities were sent an Urgent Medical Device Voluntary Recall letter by Arthrex, Inc. The firm is voluntarily recalling and replacing the referenced device. Immediately discontinue distributing and use of the indicated product and return to Arthrex. Additionally the Recall Acknowledgement of Receipt card should be completed and returned. Contact Customer Service at 800-934-4404 for replacement and return information. |  | 117 | 117 devices | Distributed to the states of AR, AZ, CA, CO, FL, IL, IN, KS, KY, LA, MD, MO, NC, NJ, NV, NY, OH, PA, SD, TN, TX, and WI. | Open, Classified |  |  | 117
Z-1351-2016 | 73338 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64:Computed Tomography X-ray Systems, Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes |  |  | February 05, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The firm was notified of a software error in which the system may not automatically send all image/data series to remote devices. | Software design | Philips Medical Systems sent a Customer Information dated February 5, 2016, to all affected customers via certified mail. The letter identified the product, the problem, and the action to be taken by the customer. To correct the issue Philips will release field change orders (FCO) 72800652 and 72800644 to install a software update on the affected systems.Customers with questions were instructed to contact their local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer CareSolutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts).For questions regarding this recall call 440-483-2997. |  | 285 | 285 units | Worldwide Distribution - US to the District of Columbia: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA &amp; WV.Product was also shipped to the following countries: Afghanistan, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Cuba, Denmark, Egypt, Ethiopia, Finland, France, Germany, Ghana, Gibraltar, Greece, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kyrgyzstan, Lebanon, Lithuania, Macedonia, Malawi, Mexico, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestinian, Philippines, Poland, Republic of Korea, RÂżunion, Romania, Russian Federation, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 1002
Z-1350-2016 | 73639 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition AS, a family of CT systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission dta from either the same axial plane taken at different angles or spiral planes taken at different angles. |  |  | March 22, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens became aware that the front cover on some Siemens CT system was not properly welded. It is possible for the welded-in threaded bolt to break out of the ground plate of the front cover if the cover is opened during service. This could not happen during clinical operation. Therefore, there is no potential risk for the user or patient and the CT System has no risk for a malfunction during clinical operation | Other | Siemens sent an "Informational Only" notice to customers on 03/22/2016 informing them to be aware that the front cover on some Siemens CT systems was not properly welded. The letter was only to notify customers that the issue is being addressed with the completion of a service visit. In addition, Siemen plans to replace each cover and hand delivered each notification to customers at the time the service is performed. |  | 1 | 1 CT system | Distributed to: Cincinnati, OH. | Terminated | August 07, 2017 | 503 | 1
Z-1361-2016 | 73410 | No | GXD | Radiofrequency lesion generator. | Neurology | Neurology | 510(k) | Neurotherm NT2000iX RF Generator |  |  | October 09, 2013 | 2013 | Neurotherm, Inc. | NT2000iX software shipped with the international setting turned on. The international settings include access to Corodotomy, Bi-Polar, and No Temperature modes were not cleared in the United States for use. | Software Manufacturing/Software Deployment | Neurotherm notified consignees via telephone and guided through the process to update the software manually. Thecalls were initiated by Neurotherm representatives on October 09,2013. |  | 25 | 25 | CA, CO, IA, ME, NJ, NY, SC, TX, WI | Terminated | April 14, 2016 | 918 | 25
Z-1362-2016 | 73171 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | USE (Ebony) PTA .014" rapid exchange (RX) Peripheral Dilatation Catheter. Distributed by US Endovascular, St. Louis, MO. |  |  | February 11, 2016 | 2016 | Us Endovascular | Product may kink during use rendering the product unusable. | Device Design | NATEC Medical ltd.sent an Urgent Medical Device Advisory Notice dated March 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Warnings:This device should be used only by physicians who are experienced and have a thorough understanding of the clinical and technical aspects of angioplasty.What action do you need to take:1. Always use a guiding catheter when contralateral access is needed.2. Always insert cautiously the PTA catheter inside the guiding catheter, and if any resistance is felt, the cause should be determined via fluoroscopy.3. For contralateral procedure, the USE014OTW has to be preferred to Ebony 014 Rx.NATEC Medical Ltd and our distributor US Endovascular are committed to keeping you and your healthcare professional informed of issues and solutions concerning our products and services. Should you have any questions regarding this or any other matter regarding your USE 014Rx PTA catheter, please do not hesitate to contact our distributor: US Endovascular at (314) 341-4930. |  | 843 | 843 catheters | US Distribution to the states of : AR, CA, FL, LA, MO, MI, MS, NY, TN, and TX. | Open, Classified |  |  | 843
Z-1363-2016 | 72894 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Medix NatalCare ST-LX Intensive Care Incubator. Neonatal incubator. |  |  | January 08, 2016 | 2016 | Natus Medical Incorporated | NatalCare LX Incubators were manufactured using a defective oval port hinge assembly with the result that the screws holding the port in place can become loose and possibly result in the port detaching from the incubator. | Process control | Natus of Seattle, WA began sending the Urgent Field Safety Notice NatalCare LX Incubator letter (DOC-017939A) to consignees on 1/8/16.A repair kit to correct this issue is now available for repair of affected units. Consignees should complete the NatalCare LX Incubator Oval Port Hinge Repair Kit Order Form and FAX it 847-949-6031 or scan and email this form to Natus_Quality_Programs@natus.com.Alternately you can telephone Natus Technical Service at 888-496-2887 and order the kit from a specialist. Natus will send the kit(s) free-of-charge to the address you specify. Each upgrade will require up to Âż - 1 hour of time and require simple hand tools.If you have any questions, please contact Technical Service at 888-496-2887 or send email to: customer.service@natus.com |  | 75 | 75 units (65 units in the US and 10 units in Pakistan) | Distribution US nationwide and Pakistan. | Open, Classified |  |  | 75
Z-1336-2016 | 73412 | No | MDS |  |  | Clinical Chemistry | PMA | Dexcom G4 PLATINUM (Professional) Receiver |  |  | February 23, 2016 | 2016 | Dexcom Inc | Customers may not receive an intended audible alarm or alert if relying on hearing the alarm or alert. | Device Design | Dexcom sent out a customer notification letter by certified mail to all direct customers on 02/23/16 notifying customers that they may not receive an intended audible alarm or alert if relying on hearing the alarm or alert. As a result customers may not detect a severe hypoglycemic (low glucose) or hyperglycemic (high glucose) event.The letter instructs the customer to peridically test the alarms and alerts on their receiver to make sure that the alarms and alerts are functioning properly (especially if the receiver gets wet of is dropped). Receiver part numbers and instructions for testing the alarms and alerts can be found on the back of the notification as well as in the User's Guide.Should customers find that alarms and alerts are not functioning properly or if any questions regarding the notification contact the deciated hotline tool free at 844-607-8398 or dial 858-291-1700.The healthcare professional notification that was sent out provided similar instructions as the direct customer notificaiton in which it is recommended that healthcare professionals periodically check the speaker function by following the testing instructions attached in the notification.The distributor notification letter instructed them to take the following actions:Hold shipments of Dexcom receivers and sticker all existing Dexcom receiver boxes on the outside front of the box with the enclosed notification stickers (see the attached visual guide for sticker placement). Once completed can begin shipping the products.In order to ensure that the notification is timely received by patients select one of the following two (2) options by no later than February 25, 2016:Send, via certified mail, the notification directly to those patients that have purchased a Dexcom receiver from you. If elect this option, Dexcom will provide the required quantity of notifications enclosed in certified mail envelopes. In order to facilitate this effort email the number of notifi |  | 5722 | 5,722 | Worldwide Distribution-US (nationwide), Canada, Italy, Netherlands, Belgium, Norway, New Zealand, Poland, Sweden, Denmark, Slovenia, Turkey, South Africa, India, Israel, Hungary, Hong Kong, UK, Ireland, France, Finland, Spain, Portugal, Colombia, Chile, Germany, Austria, Switzerland, Kuwait, United Arab Emirates, Australia, Czech Republic, Bahrain, Lebanon, and Qatar. | Terminated | August 11, 2017 | 535 | 270111
Z-1364-2016 | 73528 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Diana Automated Compounding System, Diana Onco Plus, Item No. ASN220 |  |  | March 08, 2016 | 2016 | ICU Medical, Inc. | ICU Medical Inc. has received a report of compounded volumes not meeting the accuracy specification by the Diana Compounding System when mixing volumes under 5ml. | Use error | ICU Medical, Inc. sent customer notification letter dated March 8, 2016, to customers to inform them that ICU Medical Inc. has received a report of compounded volumes not meeting the accuracy specification by the Diana Compounding System when mixing volumes under 5ml. The letter informs the customers that ICU Medical will provide an updated Operator's Manual to provide clarity on the difference between Operating Range and Accuracy Range as soon as they are available. Customers are informed that the updated specifications are provided with the customer communication and to be used until the Operators Manual is provided. Customers are instructed to return the completed Medical Device Safety Communication Response form to ICU Medical via Fax (801) 264-1755 or by email at recall@icumed.com. Customers with any questions or require assistance relating to the customer notification letter are instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time:(866)8299025 and select option 8, or email the following address:productreturns@icumed.com. |  | 36 | 36 units | Nationwide Distribution to Nationwide in AL, AZ, CA, CO, CT, DC, GA, IL, MD, MI, MN, MO, NC, NH, NJ, NY, OH, OK, PA, TN, TX, VA..Worldwide: AU, BE, BR, CA, CH, DE, DZ, EE, EG, ES, FI, FR, HK, HR, IL, IT, JO, KR, MX, NL, PL, SA, SE, SI, TR. | Open, Classified |  |  | 196
Z-1371-2016 | 73390 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | QUADROX-i oxygenators |  |  | February 23, 2016 | 2016 | Maquet Cardiovascular Us Sales, Llc | Maquet Cardiopulmonary (MCP) is conducting a voluntary recall based on the potential risk of exposure to endotoxin for various products. | Unknown/Undetermined by firm | The firm, Maquet Inc., sent an "URGENT MEDICAL DEVICE RECALL" Letter/Response Form dated 2/23/2016 via Fed Ex to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to review the list to determine if you have the affected products; monitor patients for signs; monitor and treat signs according to your protocols and clinical judgments; support shock symptoms and maintain circulatory and hemodynamics per your facility's protocol and care guidelines, and complete and and return the enclosed Response Form by email to: MCPrecall2016@maquet.com or fax to: 1-973-629-1518 within three (3) business days.For any questions, please contact your Maquet sales and representative or Maquet Customer Service at 1-888-627-8383 (press option 2, followed by option 2) Monday through Friday, between the hours of 8:00am and 6:00pm EST. |  | 72000 | 72,000 units in total | US Nationwide Distribution. | Open, Classified |  |  | 72000
Z-1377-2016 | 73649 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System (IS4000) user manual English and OUS. |  |  | March 16, 2016 | 2016 | Intuitive Surgical, Inc. | Intuitive Surgical has found that the use of instruments other than the Cardiac Probe Grasper (e.g. LargeNeedle Driver Instrument) to manipulate the ablation probes during cardiac surgery can result in an increase of microscopic metallic particulates generated from the contact with the probes. | Device Design | Intuitive sent a Field Safety Notice/Urgent Medical Device Correction letters dated March 22, 2016 by trackable mail method to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to 1. In addition to existing instructions for use, observe the following warning when manipulating the cardiac ablation probe, Use only the Cardiac Probe Grasper to manipulate a cardiac ablation probe.2. Distribute a copy of this letter to all da Vinci S, Si, and Xi users at your facility.3. Place a copy of this letter with your user manual.4. Complete the attached Acknowledgement Form and return it to Intuitive Surgical as instructed.5. Retain a copy of this letter and the Acknowledgement Form for your files.Intuitive Surgical will provide a user manual addendum to incorporate the warning related to this risk.For questions contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service. |  | 248 | 248 - total all models | Worldwide Distribution - US Nationwide in the countries Australia, Belgium, Brazil, Canada, Chile, Finland, France, Germany, Greece, Iceland, India, Israel, Italy, Netherlands, Qatar, Russia, Singapore, South Korea, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 248
Z-1375-2016 | 73630 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | Boston Scientific TSX Transseptal Needle, Model Number M004TSX60o; |  |  | March 15, 2016 | 2016 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is voluntarily conducting a recall of one lot of the TSX" Transseptal Needle due to a discrepancy between the labeling and the actual configuration of the device in the packaging. The curve dimensional values on the labels of this lot incorrectly indicate a 50Âż curve, rather than the correct 86Âż curve | Software design | The consignee was provided verbal and written notification of the field action on March 15, 2016. There was a total of one(1) consignee, an OEM customer in the USA, who received 25 units. Written notification included product identification information, instruction to immediately quarantine any devices and discontinue distribution, and instructions to initiate a product recall in accordance with their company's established policies for customers to whom they shipped the affected Merit product. An email notification was sent to the Merit sales representative of the affected account. The sales representative was provided a Customer Response Form applicable to their account. The sales representative was instructed to assist the account with the completion of the Customer Response Form and the return of the affected units back to the Merit Field Assurance Department. |  | 25 | 25 | Distributed to one customer in the US. | Open, Classified |  |  | 25
Z-1370-2016 | 73719 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, PT Graphix Magnet J-TIP Guidewire |  |  | March 28, 2016 | 2016 | Boston Scientific Corporation | Boston Scientific is voluntarily implementing a Medical Device Recall of one batch of PT Graphix Magnet J-TIP Guidewires. Boston Scientific has found that, while the batch is labeled as a J-tip model, the guidewires are straight and do not have the pre-formed J-tip. The most severe injury expected to occur is a limited procedure delay caused by the need to replace the guidewire with one that has | Process control | Boston Scientific sent an "Urgent Medical Device Recall" letter dated March 28, 2016. The letter was addressed to Cath Lab Materials Manager / Field Action Contact. The letter described the problem and the product affected by the recall. Advised consignees to cease immediately further distribution or use of the product and to notify their customers. They also requested consignees to complete and return "The Reply Verification Tracking Form" For questions contact your local Boston scientific sales Representative of the Boston Scientific Filed Action Center.. |  | 47 | 47 units | US nationwide in the states of ME, LA, and TN | Open, Classified |  |  | 47
Z-1381-2016 | 73383 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aisys CS2 anesthesia.. |  |  | February 17, 2016 | 2016 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential safety issue involving a momentary, self-correcting anesthetic agent bolus when using 21% oxygen on all Aisys CS2 and upgraded Aisys anesthesia devices. | Nonconforming Material/Component | Consignees were sent on 2/17/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 34075 dated February 17, 2016.The letter was addressed to Chief of Anesthesia, Director of Biomedical / Clinical Engineering, and Health Care Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GEHealthcare Service at 1-800-437-1171 or your local Service Representative. |  | 5496 | 5496 | US: AL, AZ, AR, CA, CO, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, NE, NV, NJ, NY, NC, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI.OUS: ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BERMUDA, BOLIVIA, BRAZIL, CANDA, CHILE, CHINA, COLOMBIA, CROATIA, CRÂżCHE REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY,GREECE, HELSINKI, HONG KONG, HUNGARIA, ICELAND, INDIA, IRAQ, IRELAND, ITA LY, JAPAN, JORDAN, KAZAKHSTAN, KOREA, KUWAIT, MALAYSIA, MEXICO, MOLDOVA, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, PERU, POLAND, PORTUGAL QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM,URUGUAY, VIETNAM. | Open, Classified |  |  | 6790
Z-1380-2016 | 73665 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ DIAGNOSTIC ULTRASOUND SYSTEM, Model EPIQ 5C, EPIC 5G, EPIQ 5W, EPIQ 7C, EPIC 7GC, and EPIQ 7W. |  |  | March 28, 2016 | 2016 | Philips Ultrasound, Inc. | The fasteners securing the control panel assembly to the base of the Philips EPIQ Ultrasound System may loosen over time, which could subsequently lead to the detachment of the entire assembly from the ultrasound system. | Device Design | The firm, Philips, sent an "Urgent-Medical Device Correction" Philips EPIQ Ultrasound System (MDC 79500381/2), letter dated 2016 MAR 23 to consignees on 3/28/16. The letter describes the product, problem and actions to be taken. The customers were instructed to review the information with all members of your staff who need to be aware of the contents of this communication.If, at any time, the control panel assembly on your ultrasound system wobbles or feels loose, stop using your system immediately and contact your local Philips representative or Philips Customer Service at 1-800-722-9377. Otherwise, you may continue to use your system.Philips will contact all EPIQ customers to arrange for service to replace the fasteners connecting the control panel assembly to the system with fasteners less likely to loosen with repeated handling. This service will be performed free of charge. If you need any further information or support concerning this issue, please contact your local Philips representative or Philips Customer Service at 1-800-722-9377. |  | 11085 | 11,085 units total (4909 units in the US and 6176 units outside the US) | Worldwide Distribution: US (nationwide) including Washington, D.C., and countries of: Algeria, Argentina, Armenia, Australia, Austria, Bahrain, Bangladesh, Belgium, Bermuda, Bolivia, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, French Guiana, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jersey, Jordan, Kazakhstan, Kenya, Korea, Republic of, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mayotte, Mexico, Monaco, Mongolia, Morocco, Myanmar, Netherlands, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Palestine, State of, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, RÂżÂżunion, Romania, Russia , Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, and Viet Nam. | Open, Classified |  |  | 11085
Z-1378-2016 | 73610 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Spetzler Claw Tip, Universal - 5450-800-311 |  |  | March 22, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | The Spetzler Claw Tip was packaged with the incorrect Tip Sleeve Assembly. The Tip Sleeve Assembly included in the packaging is the incorrect size.Risk to Health:If the Tip Sleeve Assembly does not fit the Spetzler Claw" Tip, a surgical delay may occur while an alternate Tip Sleeve Assembly is retrieved. | Mixed-up of materials/components | On 3/22/2016, URGENT MEDICAL DEVICE RECALL NOTIFICATION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments:Kara Spath269-389-4518kara.spath@stryker.com |  | 140 | 140 units | Distributed in the states of AZ, CA, CO, CT, IL, IN, MD, MI, MN, NY, PA, SC, TN, WA, WI, and VA, and in the NETHERLANDS. | Terminated | July 19, 2016 | 119 | 140
Z-1379-2016 | 73595 | No | LWH | Surgical sponge scale. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | SurgiCounter scanner, part of the SurgiCount Safety-Sponge System. |  |  | March 21, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | The affected scanner does not correctly interact with the SC360 software. When docked pre/postoperatively, the affected SurgiCounter scanner may display an error message (code=2). Also, the SurgiCounter scanner software version number is not appropriately displayed in the SC360 application during installation. | Software design | Stryker sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letters dated March 21, 2016 to affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments:Kelly Jo Davis269-389-2921kellyjo.davis@stryker.com |  | 286 | 286 | US Nationwide Distribution in the states of AR, AZ, CA, CT, FL, GA, IN, LA, MD, MN, MO, MS, NJ, NM, NY, OH, SC, SD, TN, TX and UT. | Terminated | September 06, 2016 | 169 | 286
Z-1397-2016 | 72992 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Catalog Number: 9001771 MAYO STAND COVER STERILE |  |  | January 07, 2016 | 2016 | Customed, Inc | Recall initiated due to inadequate validation/qualification of sterilization processes, and package integrity for products released for distribution as sterile. | Packaging | Recall Notification Letter sent to direct accounts via certified mail with return receipt on January 7, 2016 and January 8, 2016. |  | 540 | 540 | Distributed only in Puerto Rico. | Terminated | May 03, 2017 | 482 | 62224
Z-1388-2016 | 73054 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | The Concerto Basic Shower Trolley is intended for assisted hygiene care, especially showering and bathing of residents in care environments such as senior/assisted living, group home, special care, nursing homes, hospitals and home care. The Concerto Basic Shower Trolley is intended for indoor use. |  |  | December 21, 2015 | 2015 | Arjo, Inc. dba ArjoHuntleigh | Reports of the bolts connecting the stretcher to the hydraulic piston and the stretcher tilting mechanism becoming loose. | Process change control | An Urgent Field Safety Notice and Customer Response form was mailed to affected customers describing the reason fro the recall and the actions to be taken. The instruction to customers is to review the Field Safety Notice and complete the Customer Response Form to acknowledge receipt and understanding of the Field Safety Notice and the required actions.These actions taken are as follows:1. Ensure that all caregivers and users of the ArjoHuntleigh Concerto or Basic Shower Trolley are made aware of this Field Safety Notice (FSN) and all affected Concerto Shower Trolleys at your facility will be available for inspection by an ArjoHuntleigh Service Technician. 2. Ensure that a copy of this field safety notice (FSN) is placed in the Concerto/Basic Shower Trolley Instructions for Use manual for future reference and training. 3. Complete and sign the enclosed Customer Response Form and return this form to the ArjoHuntleigh Regulatory Department. Note: ArjoHuntleigh will contact the designated facility representative listed on the Customer Response Form to schedule service to inspect and upgrade your device, free of charge. ArjoHuntleigh advises customers to distribute the notice to those individuals who need to be aware within their organization or to any organization where the potentially affected devices have been transferred. Additionally the they should maintain awareness of the notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. The Notice includes a Note for the customer to inform FDA of adverse reactions or quality problems experienced with the use of this product may be reported to the FDAs MedWatch Adverse Event Reporting program either online, by regular mail or by fax. The letter states that if the cusstomer has additional questions or they require assistance completing the Customer Response Form, please contact ArjoHuntleigh at 1-800-323-1245, ext. 57985 or via email at AHNA-RA |  | 1596 | 1,596 units | US Consignees; 222 units distributed in the following states: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, MI, MO, MS, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI and WV. ***Foreign Consignees; 1374 units distributed in the following countries: Australia, Austria, Belgium, Bosnia &amp; Hercegovina, Brazil, CANADA, China, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, Iceland, Ireland, Italy, Japan, Lithuania, Luxemburg, Netherlands, New Zealand, Nigeria, Norway, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan and UK. | Open, Classified |  |  | 1596
Z-1394-2016 | 73430 | No | NFG | Medical image communications device. | Radiology | Ophthalmic | 510(K) Exempt | VERION Reference Unit (Vision Planner) |  |  | March 01, 2016 | 2016 | Alcon Research, Ltd. | Alcon is conducting a voluntary medical device correction of all VERION Reference Units (Vision Planner) that are shared with the Alcon LenSx Laser System after receiving reports concerning the inclusion of unplanned arcuates on printed, saved (.pdf) or exported surgical plans. | Software change control | Alcon sent an Urgent Medical Device Correction letter dated February 29, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that an Alcon Technical Services Engineers will contact customers, schedule Service Calls, and perform onsite corrections. For questions call 1-800-862-5266. |  | 97 | 97 units | Worldwide Distribution - US Nationwide throughout the US and the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Czech Republic, Chile, China, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, France, Georgia, Greece, Guatlemala, Hungary, India, Indoesia, Isreal, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Nicaragua, Norway, Panama, Philippines, Portugal, Romania, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, UK, Ukrain, United Arab Emirates | Terminated | April 13, 2017 | 408 | 97
Z-1395-2016 | 73454 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | BARD ArcticGel Pads |  |  | March 04, 2016 | 2016 | C.R. Bard, Inc. | Complaints were received of restricted/inaccurate flow rate. | Process control | Bard Medical sent an Urgent Medical Product Recall letter dated March 4, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to not use any of the affected products and to return any unused product Bard Medical Division via the provided Customer Recall and Effectiveness Form. BMD will be providing a replacement for the returned product. If the product was further distributed, the consignee was instructed to identify their customers and notify them of the product recall. |  | 21412 | 21,412 units | Worldwide Distribution - Domestic: US Nationwide; Foreign: Argentina, Australia, Austria, Belarus, Belgium, Bosnia, Brazil Canada, Chile, China, Colombia, Croatia, Cyprus, Denmark, Egypt, Finland, France, Germany, Greece, Oman Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, , South Korea, Kuwait, Macedonia, Malaysia, Mexico, New Zealand, Norway, Oman, Pakistan, Paraguay, Peru, Philippines, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, The Netherlands, Turkey, United Arab Emirates, Ukraine, United Kingdom, Uruguay and Vietnam. | Open, Classified |  |  | 21412
Z-1392-2016 | 73552 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Oblique Posterior Atraumatic Lumbar (OPAL) System, Polymeric spinal fusion cage, non-sterile, OPAL Spacer 10 mm x 24 mm, 11 mm Height - Revolve |  |  | February 18, 2016 | 2016 | Synthes (USA) Products LLC | The label on the outer packaging does not match the product inside and intended use of the device. | Labeling mix-ups | DePuy Synthes sent an Urgent Notice letter dated February 18, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review inventory, immediately remove the affected lots from their stock and complete the attached Verification Section. For questions call 610-719-5450 or contact your DePuy Synthes Sales Consultant |  | 6 | 6 units | Distributed in the US: OH and ND | Open, Classified |  |  | 6
Z-1389-2016 | 73568 | No | MNL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Cygnus Medical First StepBedside ADDWATER Kit,non-sterile. |  |  | March 17, 2016 | 2016 | Cygnus Medical | Foreign material in container | Packaging | Cygnus Medical LLC sent an Recall Notice letter certified mail to direct customers and receivers on March 17, 2016. The letter identified the product the problem and the action needed to be taken by the customer. The letter was preceded with an email or fax to expedite notification. Customers to respond via fax or email regarding the quantities in their possession, including a response where the quantity is zero. Returns to made to firm. Contact Quality Manager Cygnus Medical LLC, phone: (203) 488-4554 x 105 fax: (203) 764-1858.Email: qualitycontrol@madpoly.com.For further questions, please call (203) 488-4554. |  | 4680 | 39 cases of 120 units | US Distribution to the states of : CA , CO, FL, IL. MN, MT, NY, OH, PA, TX, VA, VT, WA and WI. | Open, Classified |  |  | 4680
Z-1390-2016 | 73597 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | BD-Retic-Count Reticulocyte Reagent System; |  |  | March 18, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | The carton label for one lot of BD Retic-Count Reticulocyte Reagent System had the incorrect storage conditions. The labels on the reagent as well as the Instructions for Use are correct. | Under Investigation by firm | BD Biosciences sent an Urgent Product Recall letter on March 18, 2016, to all affected customers by certified mail. The letter identified the product the problem and the action needed to be taken by the customer. Customer Actions:1. Examine your data files obtained using reagents from the affected lots for low performance.2. Discard all remaining stock of BD Retic-Counter reticulocyte reagent system, Catalog Number 349204, lot number 5175559, and request replacement product. BD will replace product at no cost.3. As a requirement of our Regulatory Tracking Process, we request that you complete the enclosed Tracking/Verification form , even if you no longer have the affected reagents listed in your inventory.Please return the Tracking/Verification form to BD Biosciences, per the instructions (mail, fax, or email).If you require further assistance, please contact BD Customer Support at 877.232.8995 (prompt 3, 3) in the United States. For customers outside the US, contact your local BD Biosciences representative or distributor. |  | 205 | 205 | Worldwide Distribution - US Distribution to the states of : CT, MN, FL, NY, CA, MA, DC, UT, VA and IN., and to the countries of : Belgium, Brazil, Canada, China, India, Japan, Korea, Singapore and Taiwan. | Open, Classified |  |  | 205
Z-1497-2016 | 73339 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AMT G-JetÂż Low-Profile Transgastric-Jejunal Feeding Device 14F 1.5cm 45cm SINGLE USE ONLY Rx ONLY STERILE. |  |  | January 26, 2016 | 2016 | Applied Medical Technology Inc | The firm received complaints that the product is shorter in length than indicated on the label. | Employee error | On 1/26/2016 the firm telephoned their customers notifying them of the recall. On 2/19/2016 the firm sent recall notification letters to their customers. The letter identified the product, problem, and actions to be taken by the customer. |  | 20 | 20 units | Distributed in the states of AZ, CA, DE, IN, MI, NM, PA &amp; VA. | Open, Classified |  |  | 20
Z-1391-2016 | 73544 | No | PLV | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | X-Guide Surgical Navigation System, X-Clip Bulk Kit |  |  | February 23, 2016 | 2016 | X-NAV Technologies, LLC | The X-Guide system utilizes an X-Clip (PN P007665) which contains fiducials to assist with registration of the patient anatomy. The X-Clip is a device worn intraorally by a patient during the CT scan and guided surgery procedures. The X-Clip contains three fiducial markers made from a Stainless Steel material. The X-Nav specification requires these fiducials be fabricated from Stainless Steel 3 | Nonconforming Material/Component | The firm, X-Nav Technologies, LLC, sent a "Medical Device Recall" letter date 2/23/2016 via mail and email to customers indicating the issue and risks involved. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory of X-Clips and remove any product containing the affected lots; complete the enclosed Acknowledgement and Receipt form to record the lot numbers located on the labeling for parts and discard affected product and return form to firm by Email to: Vigilance@X-Navtech.com or by Mail to: X-NAV TECHNOLOGIES, LLC, 1555 BUSTARD ROAD, SUITE 75,LANSDALE, PA. 19446, ATTN: Director, Quality Assurance and Regulatory Affairs, even if you have already used or no longer have possession of the X-Clips identified. Should you have any questions regarding this notification, please contact our Customer Service representatives at 267.436.0420 from 8:30a.m. to 5:30p.m. (EST) Monday-Friday. |  | 845 | 845 units (P007665) | US Distribution to: FL, OR, NE, PA, NM, FL, VA, OH, WA, CA, MD, LA, ID, CN, TN, and DC. | Open, Classified |  |  | 845
Z-1503-2016 | 73685 | No | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ad-Tech Electrode Connection System/Cables. Lightweight TECH-ATTACH Cable Models: |  |  | March 29, 2016 | 2016 | Ad-Tech Medical Instrument Corporation | Ad-Tech is recalling unopened/unused sterile cables intended to be used only with all compatible Ad-Tech electrodes for the purpose of connecting these electrodes to third party monitors/stimulators. This recall has been initiated due to package failures (i.e. compromised sterile barriers) that were identified during the execution of a shipping simulation investigation. | Package design/selection | Consignees were sent a "Medical Device Recall" letter dated March 29, 2016. The letter described the problem and the product being recalled. Advised consignees to immediately examine their inventory, quarantine and return the product to AD-Tech. If the product has been further distributed identify customers and notify them of the recall. Requested consignees to complete and return the "Medical Device Recall Return Response" form to their Ad-Tech Clinical Specialist at customersupport@adtechmedical.com; FAX: 262-634-5668 or Telephone 262-634-1555. For questions contact Ad-Tech Clinical Specialist.. |  | 197 | 197 | Distributed in the states of GA and MI, and the countries of FRANCE, UNITED KINGDOM, COLOMBIA, and BRAZIL. | Terminated | August 24, 2016 | 148 | 197
Z-1499-2016 | 73557 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Leica Microsystems (Schweiz) Ag. The common name of the device is surgical microscope. It has been assigned the product code FSO by the FDA and is classified as a Class I medical device. |  |  | January 11, 2016 | 2016 | Leica Microsystems, Inc. | An electronic component of the Leica M525 OH4 and Leica M720 OH5 surgical microscopes with 400W illumination could potentially overheat, resulting in smoke and smell emanating from the ventilation openings of the illumination units located opposite and away from the surgeon / patient side of the surgical microscope. | Nonconforming Material/Component | The URGENT MEDICAL DEVICE RECALL letter was sent via USPS mail and email on January 25th 2016 to consignees describing the reason for the recall and the necessary steps to take. These included the following:Continue to use the system according to the User Manual. Both illumination units should be checked prior to beginning any operation case. The Leica M525 OH4 and Leica M720 OH5 provide two redundant illumination units. Independent from the issue explained, the user should be reminded that if they notice a failure in the illumination unit that they are currently using, they should switch to the back-up illumination at the next opportunity during surgery by using the "Change Lamp" page 37 for M525 OH4 and page 40 for M720 OH5 button on the control unit "Main" screen (see Picture 2 below) as described in user manual. Changing lamps will interrupt the illumination for less than 5 seconds.They should contact their Leica Microsystems representative for repair of the defective illumination unit before using the instrument for the next surgery.Note: The user manuals of the Leica M525 OH4 and Leica M720 OH5 state:Âż Never begin an operation with only one functioning xenon lamp! Âż Like any other instrument in the operating theater this system may fail. Leica Microsystems therefore recommends that a backup system be kept ready during the operation. The recall further describes transmission of the field notice:The notice must be passed on to all those within their organization who need to be aware of it and to any organization where the affected devices has been sold, rented, leased or otherwise transferred.The end users are requested to confirm receipt of the notice by signing, dating and returning the attached Acknowledgement Form' with the serial number of their instrument back to Leica Microsystems by email to:LMGRA@leica-microsystems.comwithin 10 working days after receipt of field safety notice. |  | 381 | 381 units | Worldwide Distribution-US (nationwide) including Puerto Rico and the states of MA, CA, SC, NJ, TX, CO, FL, PA, NY, MO, WI, NC, UT, ID, AZ, MI, KY, CT, OR, , IL, AL, MN, MT, OH, NE, GA, IN, SD, TN, OK, AK, KS, NV, NM, , ND, VA, WA, WY and MD, and the countries of China, Belgium, France, Germany, Great Britain, Italy, Norway, Portugal, Spain, Sweden, Switzerland, India, Japan, Argentina, Chile, Colombia, Mexico, Peru, Belarus, Egypt, Greece, Morocco, Poland, Qatar, Romania, Russian Fed., Saudi Arabia, South Africa, Turkey, Canada, , Australia, Malaysia and Taiwan. | Terminated | April 24, 2017 | 469 | 381
Z-1504-2016 | 73776 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Extended Brilliance Workspace NM Nuclear medicine image display and processing application suite. AutoSPECT Pro Software Application. Philips Medical Systems, Cleveland, OH. |  |  | March 04, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The AutoSPECT Pro application was only designed to reconstruct cardiac SPECT data obtained with detectors positioned at 90Âż or 180Âż relative to one another. However, certain gamma cameras allow for other relative detector angles. Data acquired at these other angles will not be correctly reconstructed by AutoSPECT Pro and the results will likely be erroneous. | Software design | On March 4, 2016, the firm sent an URGENT - Field Safety Notice Medical Device Correction letter to their customers. The letter identified the affected device, the reason for the recall, and under what circumstances the problem occurs. Customers are to contact Philips as described if they are unsure whether their gamma cameras supports the modes. The firm will implement the AutoSPECT Pro software upgrade and contact customers to schedule the appointment. If further assistance or support is needed, customers are to contact their local Philips representative. |  | 126 | 126 Units | Worldwide Distribution -- US, including the states of CO, DE, FL, GA, IL, IN, KS, MO, NC, NV, NY, OH, OR, PA, TN &amp; TX; and, the countries of Australia, Austria, Belgium, Canada, China, Denmark, Egypt, France, Germany, Greece, Ireland, Israel, Italy, Lebanon, Malaysia, Saudi Arabia, Slovenia, South Africa, Switzerland, Turkey, &amp; United Kingdom. | Open, Classified |  |  | 196
Z-1508-2016 | 73202 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Robotic Arm Interactive Orthopedic System (RIO) |  |  | April 30, 2014 | 2014 | Mako Surgical Corporation | Revision of the field manual | Process control | As per our conversation with firm representative on March 23, 2016, they stated "no notification was sent to the customers", and referred to Ref: Part 806 File number: RA 2015-017 where it states "Stryker Orthopaedics has completed a retrospective review of other revisions to the field service manual, and we have confirmed there are no other similar events for which Part 806 reports should have been submitted". For further questions, please call (866) 647-6256.Internal actions taken only. |  | 249 | 249 | Worldwide Distribution - Nationwide Distribution and to the countries of : Australia, Italy, Japan, Hong Kong, Scotland, Germany, Turkey , Greece, Singapore, Taiwan and South Korea. | Open, Classified |  |  | 249
Z-1509-2016 | 73643 | No | MZO |  |  | Microbiology | PMA | OraQuick HCV Visual Reference Panel |  |  | January 15, 2016 | 2016 | OraSure Technologies, Inc. | OraSure Technologies, Inc. discovered the package insert included with the OraQuick HCV Rapid Antibody test Visual Reference Panel may be incorrect due to a complaint received from one of their customers. The customer reported they received the Ora Quick Ebola Visual Reference Panel instead of an OraQuick HCV Visual Reference Panel. | Unknown/Undetermined by firm | Letters, dated January 15, 2016, and correct HCV VRP insert were sent to all customers who received HCV VRP lot 6648737. The letter stated the reason for the recall; and, asked users to discard the Package Insert that was included with the product and replace it with the corrected one provided with the letter. Questions regarding the recall should be directed to OraSure Sales Representative or our Customer Care Department at 1-800-ORASURE. |  | 244 | 244 Visual Reference Panel Kits | Nationwide Distribution -- AK, AL, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KY, LA, MA, MI, MO, MT, NC, ND, NE, NH, NM, NY, OH, OR, PA, RI, SC, SD, TN, TX, VA, WA, and WI. | Open, Classified |  |  | 244
Z-1507-2016 | 73505 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | VariAx Compression Plating System |  |  | October 10, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received two reports indicating that during screw insertion under a high lateral angle a bone screw went through a oblong hole. | Device Design | Stryker Orthopaedics sent notification letters/product accountability forms on October 10, 2013, via email. Customers are provided the product issue, hazards and risk mitigation along with actions needed. Customers should receive the affected products and complete the BRF and return to Stryker. Questions can be directed to Stryker. For further questions, please call (201) 831-5000. |  | 118 | 118 units | Nationwide Distribution | Terminated | April 21, 2016 | 924 | 118
Z-1342-2016 | 73633 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Fetch 2 Aspiration Catheter, Catalog No. 109400-001. |  |  | March 22, 2016 | 2016 | Boston Scientific Corporation | Boston Scientific is recalling all models of Fetch 2 Aspiration Catheter as a result of receiving reports from the field of catheter shaft breaks; including cases requiring the use of a snare or other intervention to retrieve the http://webapps.cdrh.fda.gov/division-tracking/images/trash.pngbroken shaft from the patient. | Process control | The firm,Boston Scientific, "Urgent Medical Device Recall Notice - Immediate Action Required" letter dated 3/22/2016. The letter was addressed to Material Manger / Field Action Contact. The letter described the problem and the product involved in the recall and actions to be taken. The customers were instructed to immediately cease further distribution or use of the product and complete and return the Reply Verification Tracking Form via email: BSCFieldActionCenter@bsci.com or fax to: BSC Field Action Center at 1-866-213-1806. Distributors should notify their customers of the recall and inform them that this recall is to the hospital level. For questions regarding this communication, please contact your local Boston Scientific Sales Representative or the Boston Scientific Field Action Center at 763-494-1133 or email: BSCFieldActionCenter@bsci.com. |  | 21155 | 21,155 units | Worldwide Distribution - US (nationwide) including DC &amp; VI except DE, ID, SD, and WY and the countries of Australia, Austria, Belgium, Bulgaria, Canada, Chile, Croatia, Cyprus, Ecuador, France, Germany, Great Britain, Hungary, Italy, Japan, Jordan, Korea, Malaysia, Mauritius, Mexico, Netherlands, Philippines, Portugal, South Africa, Spain, Sweden, Thailand, Trinidad, Tobago, and Uruguay. | Terminated | August 11, 2017 | 507 | 21155
Z-1510-2016 | 73382 | component, traction, non-invasive | KQZ | Noninvasive traction component. | Orthopedic | Orthopedic | 510(K) Exempt | Ossur Lil Angel Vest P2 Kodel. Pediatric sized vest is a support structure used with the halo system for traumatic or chronic disorders requiring traction along the cervical spine. |  |  | March 14, 2016 | 2016 | Ossur H / F | There is a potential for partial or complete fracture of the plastic stabilizer strap on the vest. | Component design/selection | Customer notification letter were sent on 03/14/16 via traceable method.The notification asked customers to examine their inventory, quarantine product subject to the recall and contact customer service for a return authorization to return the product to Ossur.Customers were asked to take the following actions:1. Pass the notice to those who need to be aware within the organization or to any organization where the potentially affected devices have been transferred.2. Maintain awareness on the notice for an appropriate period.3. If have further distributed this product identify the customers and notify them at once of this product alert. It is recommended that a copy of the notice is included.4. If any customers are currently wearing a product it is recommended a detailed inspection of the product vest structurebe conducted to ensure there is no visible cracks on the vest. |  | 435 | 435 total all models | Distribution US Nationwide and Hong Kong | Open, Classified |  |  | 435
Z-1516-2016 | 73705 | No | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | GDC-10 360 SOFT 10MM X 20CM SR Detachable Coil |  |  | March 21, 2016 | 2016 | Stryker Neurovascular | The incorrect DFU was packaged with two lots of GDC 360 degree Detachable Coils. | Under Investigation by firm | The firm, Stryker Neurovascular, sent an "URGENT: Field Safety Notification" dated 3/21/2016 via courier to its consignees/customers. The notification describes the product, problem and actions to be taken. The customers were instructed to immediately check your inventory; follow the instructions listed in the Directions For Use (DFU); circulate and maintain the notice internally to all interested/affected parties; inform Stryker if any of the subject devices have been distributed; complete and return the FIELD SAFETY CORRECTIVE ACTION AKNOWLEDGMENT FORM to Stryker via fax at 1 (866) 876-4355 within 7 calendar days, as well as notify any other organizations that the device may have been distributed to.Should you have any queries concerning this matter contact Quality Manager at 510-413-2593 or email: Geraldine.ahern@stryker.com |  | 316 | 316 total, both models | Worldwide Distribution: US (nationwide) including states of: OH, CO, WA, IL, MI, CA, TX and countries of: Hong Kong, Italy, Germany, France, Spain, Australia, Canada, Sweden, India, Turkey, Cyprus, Belgium, Switzerland, Czech Republic, Uruguay, Mexico, Saudi Arabia, Russia, Brazil, Netherlands, and Ukraine. | Open, Classified |  |  | 316
Z-1522-2016 | 73655 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Emerald Cleaner, product number 09H46-02. |  |  | March 24, 2016 | 2016 | Abbott Laboratories | Three lots of CELL-DYN Emerald Cleaner may have occurrences where the analyzer generates Quality Control Out of Range Low for parameters RBC and PLT. | Under Investigation by firm | All customers who were shipped the affected lots were sent a Product Recall letter on March 24, 2016 via Federal Express.Letters were titled Product Recall - Immediate Action Required.An additional four lots were discovered to be affected by the recall and the firm send letters to customers affected by these lots on June 27, 2016 |  | 19774 | 10,364 plus 9410 additional units for recall expansion. | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Austria, Bosnia &amp; Herze, Cameroon, Egypt, Germany, Indonesia, Iraq, Lebanon, Mali, Montenegro, Netherlands, Philippines, Tanzania, Thailand, UAE, New Zealand, Malaysia, Mexico, South Korea, Jamaica, Turks &amp; Caicos, Italy, Canada, Vietnam, Singapore, Argentina, Colombia, France, Paraguay, Australia, Uruguay, Bahrain, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Ethiopia, Finland, Greece, Switzerland, United Kingdom, Brazil, Hong, Kong, Hungary, India, Ireland, Israel, Kazakhstan, Kenya, Kuwait, Latvia, Lithuania, Luxembourg, Turkey, Jamaica, St. Martin, Nigeria, Oman, Pakistan, Philippines, Poland, Qatar, Romania, Russia, Saudi Arabia, Serbia, Slovakia, Slovenia, Spain, Guatemala, Peru, and Uruguay. | Terminated | January 04, 2017 | 286 | 19774
Z-1519-2016 | 73593 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens SOMATOM Definition AS; Intended to produce cross-sectional images of the body. Computed tomography (CT) x-ray system. |  |  | March 02, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software bugs in these SOMATOM systems could possibly cause scan aborts, rescans of patients with additional dose, delayed diagnosis and as worst case scenarios, could possibly cause the need of additional contrast medium, accidental x-radiation to the user or impede customer workflow. | Software design | Siemens initiated a Customer Advisory Notice on 03/07/2016 to inform customers about actions for bug-fixing the following systems: SOMATOM Definition AS, SOMATOM Definition Flash and SOMATOM Definition Edge, which could possibly cause scan aborts, res cans of patients with additional dose, delayed diagnosis and as worst case scenarios, could possibly cause the need of additional contrast medium. Siemens mailed certified as well as via email distribution a Customer Advisory Notice to their customers to inform them about corrective actions for bug-fixing the following systems running with software version VA44A_SP3a, VA44A_SP3b, VA44A_SP3c, correction of potential safety issues. The issues described will be corrected with software VA44_SP4, handled as update CT051/15/P. Following the installation of this update, the mentioned malfunctions will be corrected. |  | 1096 | 1096 | Nationwide Distribution-including DC, and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Open, Classified |  |  | 1096
Z-1523-2016 | 73752 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Mach 4 Mouse Probe vial within the Mach 4 Universal HRP-Polymer Kit; |  |  | March 31, 2016 | 2016 | Biocare Medical, LLC | Biocare Medical has identified a single lot (082715) of its Mach 4 HRP Polymer kit in which some few Mach 4 Mouse Probe vials may have been mislabeled as HRP-Polymer. | Labeling mix-ups | Biocare Medical sent a Medical Device Recall Notification letter dated March 31, 2016 to affected US customers by USPS. Foreign consignees were also notified per their various government regulations. The letter identified the affected product, problem and actions to be taken. Customers were requested to complete and return by April 15, 2016 the enclosed Return Response Form. |  | 93 | 93 | US Nationwide Distribution in the states of: KS, CA, AZ, OR, FL, KS, CO, MA, KY, NY, IL, OR, MO, NC, TX, PA, OK. SC, TN, PA, TN, GA, and OR.Hungary, Canada, Boliva | Open, Classified |  |  | 93
Z-1530-2016 | 73509 | No | PCD | Powered patient transport. | Physical Medicine | Physical Medicine | 510(K) Exempt | Acorn 180 Curved Stairlift |  |  | February 17, 2016 | 2016 | Acorn Stairlifts Inc | Aluminum rivets holding the base squab to the framer under prolonged customer use were seen to elongate and subsequently break. | Component change control | Customers were notified by phone calls, documents mailed with a Rationalization Action plan to follow. Contact information 1-888-563-0410 x1500 or 1400 |  | 10897 | 1,290 USA and 9,607 Foreign | Worldwide Distribution - USA (nationwide) Distribution to the states of : AL, AZ, CA, CO, DE, FL, GA, IL, IN, KS, KY, LA, MA, ME, MD, MI, MO, NH, NJ, NY, NC, OH, PA, RI, SC, TN, TX, UT, VT, VI, WA, WV, WI, and District of Columbia and to the countries of UK and Germany. | Open, Classified |  |  | 10897
Z-1500-2016 | 73588 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | ADVIA 560 Hematology Systems, Siemens Material Number (SMN) 11170842, IVD |  |  | March 10, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Software anomaly; Siemens identified that software version 1.4.2133 on the ADVIA 560 Hematology System does not trigger the following flags: The G or L morphology flags for immature granulocytes (IG) and atypical lymphocytes (ATYP), respectively. The results from patient samples which have immature granulocytes or atypical lymphocytes will not generate the flags when they should. As defined in t | Process design | Siemens sent an Urgent Medical Device Correction Letter dated March 11, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. .Actions to be Taken by the Customer:To ensure that the patient sample amount for processing is met, follow the requirements for the amount of sample required in the ADVIA 560 Hematology System Operators Guide. Review all patient results for abnormal white blood cell results and process as defined by your lab criteria for abnormal samples and confirm using manual differentials as needed.Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days.Please review this letter with your Medical Director.Your Siemens Customer Service Engineer will contact you to schedule time for installation of the 1.4.2333 software when it is available.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions, please call (312) 275-7795. |  | 72 | US: 8 systems; Foreign: 64 systems | Worldwide Distribution - US Distribution including Puerto Rico and to the states of : California, Illinois, Washington and Wisconsin., and to the countries of : Angola, Austria, Canada, Chile, Colombia, Czech Republic, France, Germany, India, Italy, Kenya, Netherlands, Norway, Portugal, Singapore, Slovenia, Spain, Turkey, Uganda and United Kingdom. | Open, Classified |  |  | 72
Z-1525-2016 | 73604 | No | EZL | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Self-CathÂż Pediatric Catheter Size CH 06, Product # 5044101400, Catakig # 306. |  |  | March 22, 2016 | 2016 | Coloplast Manufacturing US, LLC | Coloplast Corp. is voluntarily recalling a single lot of Self-Cath Pediatric French Size CH 06 cathetersbecause a small portion of the lot was identified to have missing eyelets. This recall is based upon acomplaint Coloplast received regarding a catheter malfunction. | Nonconforming Material/Component | Coloplast sent an "Urgent Medical Device Recall" letter dated March 22, 2016 to customers. The letter described the problem and the product involved in the recall. Advised consignees to immediately examine, quarantine and return the affected product. If the product has been further distributed, consignees should identify their customers and notify them of the recall. It also requested consignees to complete and return the Acknowledgement Form. For questions contact Coloplast Customer Service directly at 800-533-0464. |  | 31320 | 31,320 | US Nationwide Distribution in the states of: AL, AR, C, FL, IL KY, LA, MD, MA, MI, MN, MO, NJ,NY, NC, OK,OR, PA, TN, TX, VA and WA. | Open, Classified |  |  | 31320
Z-1529-2016 | 73602 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 for Proton Radiation Therapy |  |  | March 15, 2016 | 2016 | Mevion Medical Systems, Inc. | Possibility of Gantry Motion with Aperture or Compensator only partially inserted. An Aperture may fall out during Gantry motion. It is possible that it can hit a patient and cause traumatic injury or death. | Device Design | Customer notification letters were sent to customers on March 15, 2016, by certified mail, return receipt requested, as well as by email. The letter explains the problem, description of potential harm, and the cause containment, and mitigation. The cause of the malfunctioning sensors was discovered to be a non-conforming limit switch assembly. A Mevion Field Service Engineer will immediately inspect all such parts in the field for similar non-conformance, and replace those parts if necessary. In addition, a new periodic preventive maintenance check will be performed to individually test each of the redundant sensors at all sites. In the interim period, on a daily basis, until the inspection, repair, and the preventive maintenance check is completed at any site, Mevion Field Service will check that the Applicator Door interlock sensors are working properly. Mevion will also update the inspection procedure for incoming parts to ensure that any non-conformance is identified early in the manufacturing process. |  | 5 | 5 | US nationwide distribution to MO, OK, FL, and NJ. | Open, Classified |  |  | 5
Z-1524-2016 | 73549 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Cervical Spine Locking Plate (CSLP) System, 4mm Cancellous Self Tapping Screw, Length 16mm Lot # 9666776 and 4mm Cortex Standard Screw, Length 16mm Lot # 9671879 |  |  | January 06, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is initiating a Voluntary Medical Device Recall of lot # 9666776 and 9671879 of Cancellous and Cortex 4.0 mm Screws, which are a part of the Cervical Spine Locking Plate (CSLP) System due to discovery of a single lot of 450.136 screws (4mm Cancellous Self Tapping Screw, Length 16mm, lot 9666776) being swapped with a single lot of 450.016 (4mm Cortex Standard Screw, length 16 mm, lot | Packaging process control | DePuy Synthesis sent an Urgent Notice Medical Device Recall letter dated January 6, 2016 to customers. The letter identified the affected product, problem and actions to be taken. The letter requested customers to complete and return the attached Verification Section (page 3 of the letter). For questions call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 290003280 | 30 units of Lot # 9666776 and 11 units of Lot # 9671879 | Worldwide Distribution - US Nationwide in the states of NC, CA, NY, TX, NJ, CO and the countries of: Ontario and Canada. | Open, Classified |  |  | 290003280
Z-1540-2016 | 73609 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TC-PLUS(TM), Primary Fixed Tibial Component VKS Right, SIZE 2 RIGHT, STERILE R, REF 12 000 009, S&N; 7500264, Smith &amp; Nephew Orthopaedics AG |  |  | March 11, 2016 | 2016 | Smith &amp; Nephew, Inc. | The Affected Devices may contain a dimensional deviation that could potentially result in thesurgeon being unable to seat the polyethylene (PE) Insert on the affected VKS tibial baseplate during surgery. | Process control | The firm, Smith &amp; Nephew, issued an "URGENT FIELD SAFETY NOTICE: Medical Device Field Safety Corrective Action/Recall" letter dated 3/10/2016 to its local QA/RA personnel/customers in the affected countries by email on 03/11/2016. The letter described the product, problem and actions to be taken. The customers were instructed to locate and quarantine affected unused devices immediately; return quarantined product to your national Smith &amp; Nephew agency/distributor; complete and return the Return Slip via fax to your national Smith &amp; Nephew agency/distributor; and pass on this safety information to all those who need to be aware in your organization. The local QA/RA staff will then notify end customers according to local procedures. In Germany as well as in Switzerland, sales representatives will inform the affected customers in-person and retrieve affected devices from their stock.If you have any questions, call 978-749-1330 or email: joe.metzger@smith-nephew.com |  | 1 | 1 unit | International Distribution only in countries of: Germany and Switzerland. | Terminated | April 10, 2017 | 395 | 191
Z-1553-2016 | 73349 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX Pentra 400 (version 5.0.8 or lower) |  |  | February 19, 2016 | 2016 | Horiba Instruments Inc | Horiba Instruments, Inc. is recalling ABX Pentra 400( version 5.0.8 or lower) and Pentra C400 (version 1.1.2 or lower) because clinical chemistry analyzer malfunctions when the following certain alarms appear on the system. | Software design | HORIBA Medical sent an "Urgent Safety Notice" dated February 19, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A new software version to solve the issue is currently under development; however the firm requests customers use the provided procedure until the new software version is installed on the system, or be extremely vigilant with results of one of the alarms mentioned appears on the system. The letter instructs customers to share the information with laboratory staff and retain the notification as part of laboratory documentation. The letter states it is mandatory for customers to complete and return the enclosed response form within 10 days, in order to verify understanding of it.Any questions or concerns contact the local Horiba Medical representative.For further questions, please call (949) 250-4811. |  | 201 | 201 | USA (nationwide Distribution) to the states of : CA, NY, AL,OK, MS, MD, NV, MN, TX, KY, WV, MI, NY, MT, IL, FL,WA, ID, MO, WY, NJ, VA, KY, GA, PA, TN, MS, KS, AZ, NE, SC, OR, OH, IA, RI, MI and NC. | Open, Classified |  |  | 201
Z-1555-2016 | 71667 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Spirit One A-C Powered Hospital Bed |  |  | July 02, 2015 | 2015 | Stryker Medical Division of Stryker Corporation | Customer complaints associated with faulty brake system | Device Design | Stryker Corp.sent an Urgent Medical Device Recall letter on July 21, 2015, via FedEx to all affected customers. The letter will identified the product the problem and the action needed to be taken by the customer. Customers are instructed the following actions: 1.Locate the units listed on the attached business reply form.2.Remove these units from service. (Note: The original mattress purchased with the bed is not within scope of this action. The mattress will be compatible with your replacement option.)3.Your Stryker Sales Representative will contact your facility to coordinate removal of the recalled beds and present replacement options.4.Return the enclosed business reply form to your Stryker Sales Representative to confirm receipt of this notification or fax (269)488-8691 or email productfieldaction@stryker.com 5.If you have loaned or sold any of the beds listed in this letter, please forward a copy of this notice to the new users and advise us of their new location in the space provided on the business reply form.6.If you have disposed of any of the beds and they are no longer in use, please advise us of their obsolescence by providing us with their serial number in the space provided on the business reply form.*For additional information, please refer to the products operations and/or maintenance manuals. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax.If there are any questions or concerns, please contact Stryker Customer Service (1-800-327-0770) during normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 143 | 143 total | US Distribution to the states of : AL, AZ, FL, GA, HI, IA, IL, KS, LA, MI, MN, MO, NE, NJ, OH, PA, SC, TN and TX., and Internationally to the countries of : Canada and UK. | Terminated | March 24, 2017 | 631 | 143
Z-1556-2016 | 73131 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Spirit One A-C Powered Hospital Bed |  |  | July 02, 2015 | 2015 | Stryker Medical Division of Stryker Corporation | Inaccurate scale systems | Device Design | Stryker Corp. will initiate the recall by distributing letters on 07/21/2015 via FEDEX. Customers are asked to do the following: 1. Locate the units listed on the attached business reply form. 2. Remove these units from service. (Note: The original mattress purchased with the bed is not within scope of this action. The mattress will be compatible with your replacement option.) 3. Your Stryker Sales Representative will contact your facility to coordinate removal of the recalled beds and present replacement options. 4. Return the enclosedbusiness reply form to your Stryker Sales Representative to confirm receipt of this notification or fax (269)488 -8691 or email productfieldaction@stryker.com 5. If you have loaned or sold any of the beds listed in this letter, please forward a copy of this notice to the new users and advise us of their new location in the space provided on the business reply form. 6. If you have disposed of any of the beds and they are no longer in use, please advise us of their obsolescence by providing us with their serial number in the space provided on the business reply form. *For additional information, please refer to the products operations and/or maintenance manuals. Adverse reactions or quality problems experienced with the use of this product may be reported tothe FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. If there are any questions or concerns, please contact Stryker Customer Service (1-800-327-0770) during normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 143 | 143 | Distributed in the states of AL, AZ, FL, GA, HI, IA, IL, KS, LA, MI, MN, MO, NE, NJ, OH, PA, SC, TN, and TX, and the countries of Canada and UK. | Terminated | March 24, 2017 | 631 | 143
Z-1528-2016 | 73205 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trinity Acetabular Shell |  |  | January 29, 2016 | 2016 | Corin USA Limited | Potential for abrasion of the sterile barrier pouches of the new packaging system. | Packaging process control | Corin sent an Urgent Medical Device Recall letter dated February 3, 2016. The letter identified the affected product, problem and actions to be taken. On January 29, 2016, Eight (8) of 11 distributors were spoken to during a Corin National Sales Meeting informing them of the recall. The 3 remaining distributors were notified by phone calls same date. Also, on February 2, 2016 Corin sent a Recall Letter to their Foreign consignee informating them of the recall of the Trinity Acetabular Shell and the actions to be take by the costumers including the return of the Customer Response Form and Contact information as Corin LTD Compliance Director at +44 (0) 1285 649190 Monday thru Friday 8:30am to 5:00pm UK Time |  | 629 | 629 devices in total | Worldwide Distribution - US Nationwide in the states of CA, FL, ID, TX, OK, OH, MA, and NC and the countries of Australia Belgium, Egypt, Germany, Ghana, Italy, Japan, Poland, South Africa and UK. | Open, Classified |  |  | 629
Z-1562-2016 | 73617 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aisys and Aisys CS2. |  |  | March 15, 2016 | 2016 | GE Medical Systems, LLC | The vacuum suction tubing could be kinked within the machine resulting in suction less than the 20 lpm required by ISO 10079-3. This could result in inadequate suction performance to clear the airway and inadequate ventilation of the patient. | Employee error | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" GEHC Ref# 34076 letter, dated March 15, 2016. The letter was addressed to the Chief of Anesthesia / Health Care Administrator / Risk Manager and Director of Biomedical / Clinical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The letter states that users can continue to use all functions of the device. However, the vacuum suction option can only be used if it passes the evaluation. A GE representative will contact customers to arrange for the correction at no cost to the customer. For questions or concerns regarding this notification, contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 952 | 952 (488 US; 464 OUS) | Worldwide Distribution -- US, to the states of AZ, CA, DC, FL, GA, ID, IL IN, KY, MD, MI, MN, MO, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI; and, to the countries of AUSTRALIA, BELGIUM. BOLIVIA, CANADA, CHILE, DENMARK, FINLAND, FRANCE, GERMANY, HUNGARIA, INDIA, IRELAND, JAPAN, KAZAKHSTAN, KOREA, KUWAIT, NETHERLANDS, NEW ZEALAND, SAUDI ARABIA, SPAIN, SWITZERLAND, THAILAND, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 952
Z-1561-2016 | 73761 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Radionuclide Radiation Therapy System |  |  | April 19, 2016 | 2016 | Elekta, Inc. | The latches of the frame adapter can be locked even if the locating pins of the frame adapter is not inserted into the corresponding holes in the coordinate frame. | Use error | Elekta sent an Important User Notice (100-03-202-27) letter to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the Acknowledgement Form attached. For questions contact your local Elekta office. |  | 217 | 217 | Worldwide Distribution - US Nationwide in the states of AK, AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, KY, LA, MA, MI, MN, MS, MO, MT, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SD, TN, TX, UT, VA, WA, WV, WI, and countries of Australia, Austria, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Egypt, France, Germany, Greece, Hong Kong, India, Indonesia, Iraq, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Netherlands, Norway, Philippines, Poland, Portugal, Russia, Saudi Arabia, Serbia, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom | Open, Classified |  |  | 217
Z-1559-2016 | 73783 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | Trypticase TM Soy Agar with 5% Sheep Blood (TSA IITM) // MacConkey II Agar |  |  | March 28, 2016 | 2016 | Becton Dickinson &amp; Co. | Trypticase TM Soy Agar with 5% Sheep Blood (TSA IITM) // MacConkey II Agar is contaminated with Listeria Monocytogenes. The device is a non-sterile prepared plated media. | Material/Component Contamination | BD sent an Urgent Product Recall letter dated March 2016 to all US customers via UPS or email. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the affected lots and discard any remaining product. BD will replace the discard inventory. Customers were asked to complete and fax the attached form to : Attention Regulatory Compliance, Fax: 410-316-4258 or email to BDRC2@bd.com. |  | 88300 | 88,300 plates (883 cartons) | US Nationwide Distribution in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, KS, LA, MA, MD, ME, MI, MN, MS, MT, NC, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TX, UT, VA, WA. | Terminated | November 14, 2016 | 231 | 88300
Z-1558-2016 | 73838 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force, Computed Tomography x-ray system |  |  | April 01, 2016 | 2016 | Siemens Medical Solutions USA, Inc | The Neonate Head protocol with the pediatric kernel Hp38 could result in artefacts and possibly lead to a misdiagnosis (i.e. either non-existing blood or liquid is mimicked in the images or actual existing blood or liquid is not depicted as expected). There is also a risk of a potential misdiagnosis in using this protocol for surgery planning. Investigations revealed an incorrect parameterization of the reconstruction algorithm applied for the pediatric head kernel as cause of the problem. | Software design | Siemens distributed a Safety Advisory Notice dated April 1, 2016, to affected customers The letter identified the affected product, problem, how to avoid potential risks, how the issue will be resolved. and actions to be taken. Customers were asked to observe the safety notice and comply with the corresponding measures until further notice. |  | 37 | 37 CTsystems | US Nationwide Distribution : NC, NH, IL, MA, WA, PA, MD, MN, TX, NY, SC, OH, CA, KY, MT, IA | Terminated | April 17, 2017 | 381 | 37
Z-1552-2016 | 73706 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3, PowerCross" 0.018" OTW PTA Dilatation Catheter, model: AB18W030200150, Sterile EO, Rx Only. |  |  | March 29, 2016 | 2016 | ev3, Inc. | Medtronic is conducting a voluntary Urgent Medical Device Recall of two lots of the PowerCross 0.018" Âż Over the Wire (OTW) Percutaneous Transluminal Angioplasty (PTA) Dilatation Catheter because the product is mislabeled. | Labeling False and Misleading | Consignees were sent on 3/29/2016 a Medtronic "Urgent Medical Device Recall" letter dated March 2016. The letter was addressed to Risk Manger or Health Care Professional.The letter described the problem and the product involved in the recall. |  | 27 | 27 (25 US, 2 OUS) | US: CA, GA, LA, MI, MS, MO,NV,NH. NY, OH, PA, SD, TX.OUS: SINGAPORE | Terminated | May 19, 2017 | 416 | 27
Z-1557-2016 | 73673 | No | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | eCareCoordinator allows the customer to schedule patient tasks (i.e., take weight measurement, take blood pressure). |  |  | March 10, 2016 | 2016 | Philips Visicu | The adherence check generates a software error.Two issues occur as a result of the software error: 1) No adherence flag is generated for that day; 2) No task is generated for the date 14 days later. An error message is generated and forwarded to operations. Subsequently, for days in which no task was generated, patient will not get a reminder on the tablet to take a measurement. If the pa | Under Investigation by firm | 1. Field Safety Notice(FSN) will be sent to all customers using eCareCoordinator v.1.3 platform. The FSN describes the problem and the circumstances in which the design defects occur and the action planned by Philips to correct the problem. The FSN will also detail the functionality change in detail.2. A software correction will be made and released. Customers will be upgraded to eCC v.1.4. The v.1.4 is scheduled to be released on or about 8 March 2016. The defect will have been corrected and tested. |  | 9 | 9 | US Distribution -- to the stated of FL, MA, MI, MS, PA, and KS. | Open, Classified |  |  | 9
Z-1560-2016 | 73857 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion Lightning CT System iStation cover, SID: 70000246, Item Code: TSX-0354A/2D.03 |  |  | April 12, 2016 | 2016 | Toshiba American Medical Systems Inc | Toshiba American Medical Systems, Inc. is recalling the TSX-035A because of a potential problem with the Aquilion Lightning CT System iStation cover falling off of the gantry. | Device Design | An Urgent Medical Device Correction letter dated 4/12/16 was sent to a customer to bring to the customer's attention with a potential problem with the Aquilion Lightning CT System iStation cover. The letter informs the customer that the iStation cover may come off due to deficient dimensional geomtry. The letter informs the customer of the problem identified and corrective action to be taken. The customer is instructed to complete and return the attached form and fax it to (877) 349-3054. If the customer has any questions regarding the recall letter, then the customer is instructed to contact Paul Biggins at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 1 | 1 unit | US in the state Nevada | Open, Classified |  |  | 1
Z-1563-2016 | 73117 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | 3M Steri-Vac Gas Sterilizer/Aerator, models 5XL and 8XL. |  |  | January 22, 2016 | 2016 | 3M Company - Health Care Business | 3M discovered a valve with incorrectly labeled "input" and "output" orientation. This defect can result in the valve being installed backwards, creating a remote possibility of valve failure. A failure may result in EO being vented outside the sealed chamber during a limited period in the EO injection phase of the cycle. | Nonconforming Material/Component | 3M sent an Urgent Medical Device Field Correction letter to customers beginning 01/22/2016. The letter identified affected product, and stated that a 3M Service representative will contact the consignee to inspect the sterilizer. Questions can be directed to 3M Health Care Service Line at 1-800-228-3957. |  | 74 | 74 valves | AL, CO, CT, FL, GA, IN, NC, NJ, NY, OK, OR, PA, RI, VA, WI, and OUS to include: Argentina, Canada, Chile, China, Columbia, El Salvador, Germany, India, Indonesia, Japan, Korea, Mexico, Peru, Philippines, South Africa, Switzerland, Thailand, Trinidad, Turkey, United Arab Emirates, Vietnam. | Open, Classified |  |  | 74
Z-1567-2016 | 73660 | No | JCT | Tracheal prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific Ultraflex Tracheobronchial Distal Release Stent System 12mm X 30mm |  |  | March 18, 2016 | 2016 | Boston Scientific Corporation | Mislabeled: 10mm diameter x 30mm length Ultraflex Tracheobronchial stents were labeled as 12mm diameter x 30mm | Process control | Boston Scientific issued notification letters dated March 17, 2016. The notification instructs the customer to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form. Customers with questions should call 866-868-4004. For questions regarding this recall call 508-382-9555. |  | 4 | 4 units | Nationwide Distribution to AL, CA, MA, and TX. | Open, Classified |  |  | 4
Z-1566-2016 | 73674 | No | HIF | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | PneumoSure XL High Flow Insufflator and PneumoSure High Flow Insufflator; |  |  | March 28, 2016 | 2016 | Stryker Endoscopy | When operating the device in high flow mode (40L/min, 15 mmHg) and connecting to a house gas inlet source, within seconds of insufflating the flow rate would drop to 0L/min and the red Check Gas Supply symbol would appear on the display, immediately suspending insufflation. | Nonconforming Material/Component | Urgent Medical Device Recall letters were sent to domestic customers on March 28, 2016 by certified mail. Stryker International distribution sites were notified and are responsible for providing information to affected customers. |  | 489 | 489 | Worldwide distribution. US nationwide including Puerto Rico, China, Hong Kong, India, Japan, Korea, Romania, Vietnam, Singapore, Taiwan, Netherlands, Peru, Australia, Canada, Mexico, Argentina, Chile, Poland, South Africa, and Guatemala. | Open, Classified |  |  | 489
Z-1564-2016 | 73677 | No | DXX | Steerable catheter control system. | Cardiovascular | Cardiovascular | 510(k) | Magellan Robotic Catheter Accessory Kits, |  |  | March 25, 2016 | 2016 | Hansen Medical Inc | Firm's quarterly dose audit testing showed that one batch was radiated below the augmented dose and therefore did not meet the augmented dose of 28.4 kGy. | Under Investigation by firm | Hansen Medical sent an Urgent Notice Product Recall notification dated March 28, 2016, to all affected customers. The notification identified the product, the problem, and the action to be taken by the customer. Customers were instructed to provide the product to the Hansen Medical field representative to assure credit appropriately applied to their account. Customers with questions were instructed to contact their local Hansen Medical representative or call 650-404-5581. |  | 8 | 8 kits | Nationwide Distribution to OH, TX. FL, CA and GA. | Open, Classified |  |  | 8
Z-1565-2016 | 73435 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Automated peritoneal dialysis (APD) cycler |  |  | February 16, 2016 | 2016 | Baxter Healthcare Corp. | Baxter Healthcare Corporation is sending this communication to inform you of incomplete instructions in the AMIA Automated Peritoneal Dialysis (PO) System Clinician Guide for calculating the total recommended solution therapy volume. Specifically, the instructions do not specify the need for an extra 200 mL of PD solution in order to prime the patient line and for air purge operations. The tot | Under Investigation by firm | An Urgent Device Correction communication was sent to the one affected customer via U.S.P.S., first class mail on 02/11/2016.A Clinician Guide Addendum was created to instruct clinicians of the extra 200 mL of PD solution required to prime the patient line and for air purge operations. This Addendum was provided as an attachment to the Urgent Device Correction communication letter. |  | 10 | 10 units | US Consignee: NY ***Foreign Consignee: Ecuador | Open, Classified |  |  | 10
Z-1570-2016 | 73748 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Trucount Absolute Counting Tubes: |  |  | March 31, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | One lot of lot of BD Trucount tubes may contain tubes from another lot. There is a 3.3% difference in bead counts between the two lots found within some of the pouches and presents a low risk to patient safety or impact to data if this product is used. | Packaging | BD Biosciences sent an Urgent Product Recall letter dated March 2016 to all affected customers on March 31, 2016 by Certified Mail, e-mail and direct delivery by Sales Associates.Letters instructed that any remaining inventory of the affected lot be discarded and that date files obtained using reagents from the affected lot be reviewed for the possibility of incorrect data.Customers with questions were instructed to contact BD Customer Support at 877-232-8995 (prompt3, 3) in the united State. For customers outside the US, customers should contact their local BD Biosciences representative or distributor.For questions regarding this recall call 408-954-6080. |  | 1743 | 1743; 467 US and 1276 International | Worldwide Distribution - US including MD, CA GA. WA. PA. CO. TN. FL. NC. DE. VA, CT, LA. NC, NY, OH, MN, NH and Internationally to China, Japan, and India. | Open, Classified |  |  | 1743
Z-1568-2016 | 73736 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist Stapler 45 and Stapler 30 instruments used on the da Vinci Xi systems (IS4000) with p5 software. |  |  | March 25, 2016 | 2016 | Intuitive Surgical, Inc. | Potential for unexpected motion of the Xi Stapler jaws on the da Vinci Xi System with p5 software relating to a combination of the p5 software and the surgeon quickly transitioning from the clamp to the fire pedal during use. | Software design | Intuitive Surgical sent an Urgent Product Safety notification to customers via e-mail on March 25, 2016. Customers were advised to STOP USE on the EndoWrist Stapler 45 and 30 instruments on da Vinci XI systems with p5 software. Product letters followed and were mailed on March 31, 2016 by Fed Ex.Urgent Medical Device letters were sent via Fed Ex on March 31, 2016 (dated April 1, 2016). These letters reiterated what was explained in the previous e-mail and provide more information of the unexpected motion issue. In addition to the stop use request, letters requested that all personnel be fully informed of the issue.Further information or support concerning this issue, please contact your Clinical Sales Representative or contact ISI Customer Service at the number:North America: 800-876-1310 Option 3 (6 am to 5pm PST) |  | 43 | 43 | US Nationwide Distribution in the states of: WI, FL, OH, TX, CA, AZ, IL, MD, MN, NY, OK, HI, MO, IA, KS, PA, and AL. | Open, Classified |  |  | 43
Z-1571-2016 | 73688 | No | JSJ | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | Hardy Diagnostics GBS Detect |  |  | March 21, 2016 | 2016 | Hardy Diagnostics | Performance failure; poor hemolytic reaction of non or weakly hemolytic group B streptococcus (GBS) with the target QC organism, Streptococcus agalactiae (ATCC 13813) due to product deterioration. | Under Investigation by firm | On 03/21/2016 customers were notified by phone, and were sent a notification letter.Customers were instructed via phone call to discard remaining plates of the implicated lot, and to notify the firm of how many plates are left. The firm stated that they will be issuing replacements. The firm stated that they will be sending written notification as well. The firm requested full name and e-mail or fax number in order to follow-up with the written notification.The customer notification letter instructs customers to complete, sign, scan/e-mail or fax back enclosed form stating compliance with the action regarding the specific lot. The firm requests the form be returned to the attention of the Quality Assurance Department at techservice@hardydiagnostics.com, fax number 805-614-9274. If the lab is unable to be reached via telephone and would like replacements or credit contact Customer Service Department at 800-266-2222, option 1. Replacements will be sent out at no charge.Any questions contact Technical Services Department at 800-266-2222, option 2 or via e-mail at techservice@hardydiagnostics.com. |  | 3530 | 3,530 (353 pk/10) | U.S. Distribution to the following states: AL, SC, FL, CT, VA, WI, MA, MI, WV, OH, IN, IL, NC, and TX. | Open, Classified |  |  | 3530
Z-1569-2016 | 73616 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen (TM) Laboratory Automation System configured with Thermo Scientific Recapper Module (230V or 110V -- COMMON/USUAL NAME: enGen (TM) Track System -- NOTE: Outside the US the system is also known as VITROS Automation Solution (VAS) TC Automation System (configured with Thermo Scientific Recapper Module from VITROS Automation Solutions), Product Code 6844300 --- The enGen (TM) Lab Automation System is a modification to VITROS Analyzers that includes the addition of an automated track sub-system (TC Automation) for pre-analytical and post-analytical sample and data management. In addition to an Automated Track Sub-system, the enGen (TM) System includes one or more VITROS Chemistry, Immunodiagnostic and/or Integrated Systems (clinical analyzers) as well as a Data Management Subsystem (Middleware), and a Software Configuration (gsb) File. The enGen System is typically interfaced with a Laboratory Information System (LIS) via connectivity facilitated at the Data Management Subsystem level. The enGen (TM) Laboratory Automation System contains a module (i.e. enGen (TM)Laboratory Automation Recapper Module) that applies caps to the sample collection container after the sample has been aspirated. In this module, an electrical cable supplies power to the cap vibratory feeder and is supposed to be located above the vibratory feeder support plate. The cap vibratory feeder sits upon a feeder support plate. Un-capped sample containers pass underneath this support plate prior to being re-capped for storage. The Recapper Module is supplied by Thermo Fisher Scientific. |  |  | March 17, 2016 | 2016 | Ortho-Clinical Diagnostics | The electrical cable that supplies power to the cap vibratory feeder of the Recapper Module may loosen from its correct position above the vibration feeder support plate and migrate to under the support plate. If this occurs, the uncapped sample tubes may hit the cable, and sample fluid (e.g., serum) may splash from the tube and into another tube or onto the under carriage of the support plate. | Component design/selection | On 3/17/2016 URGENT PRODUCT CORRECTION NOTIFICATION letter (Ref. CL2016-074, dated 3/17/2016)) was sent notifying the users of the issue and advise them of the required actions. The notification was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected product and advised them of the issue and the required actions. Foreign affiliates were informed by email on 3/17/2016 of the issue and instructed to notify their consignees of the issue and required actions. |  | 49 | US: 19, Foreign: 30 | Distributed in the states of MD, NY, ID, NC, CO, OH, PA, TN, MS, MO, CA, DE, FL, GA, AZ, NJ, and WV and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, France, India, Italy, Norway, Portugal, Saudi Arabia, Spain, Sweden, Thailand, and United Kingdom. | Open, Classified |  |  | 49
Z-1578-2016 | 73680 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hand-Held Scanner USB IT3800 |  |  | March 03, 2016 | 2016 | Roche Molecular Systems, Inc. | The hand-held barcode scanner model IT3800 used with the COBAS AmpliPrep instrument mis-identified a sample barcode ID. | Labeling design | Roche Molecular Systems sent an "Urgent Medical Device Correction (UMDC) notification letter dated March 3, 2016, to their affected customers. Customers were instructed to provide the UMDC to other affected parties who may need to be aware of this issue. Customers were asked to complete the attached fax form and fax it to 1-866-737-5396. Customers with questions should contact Roche Support Network Customer Support Center at 1-800-526-1247. For questions regarding this recall call 908-253-7200. |  | 6939 | 6,939 pieces | Nationwide Distribution | Terminated | May 25, 2017 | 448 | 6939
Z-1575-2016 | 73693 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network |  |  | March 09, 2016 | 2016 | Mindray DS USA, Inc. dba Mindray North America | Panorama Central Station including the work station View Station, View Station and the eGateway will revert to the year 2000 when Daylight Savings Time (DST) occurs on March 13, 2016, or under various conditions. | Unknown/Undetermined by firm | Mindray DS USA Inc. issued a Medical Device Service Notice dated March 9, 2016. The letter identified the product, the problem, and the action to be taken by the customer. Customers requiring technical assistance were instructed to contact Mindray's Technical Support team at 1-800-288-2121. Mindray will address the systems response to DST in a future software update which will be provided to the customer at no cost. |  | 1492 | 1447 units US, 45 units OUS | Nationwide Distribution | Open, Classified |  |  | 1492
Z-1573-2016 | 73741 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Slice CT Scanner System PN: 989605858501 |  |  | March 24, 2016 | 2016 | Philips and Neusoft Medical Systems Co., Ltd. | The following issues are found in NeuViz 16 systems with software version 1.1.4.21425 and version 1.1.4.21426: 1) During the filming operation on MX 16-slice console software, the clipboard used for copying and pasting images is not cleared between patients. If the operator fails to copy the current patient's images before pasting, a previous patient's image may be present in the clipboard and be | Under Investigation by firm | Affected consignees will be notified of the recall via letter on March 24, 2016. Mandatory Field Change Order and FCO kits will be released to all affected systems in June, 2016. |  | 37 | 37 units | NC, OH, NE, SC, TX, LA, PR, MO, FL, CT | Open, Classified |  |  | 37
Z-1579-2016 | 73747 | No | NUA | Cystic fibrosis transmembrane conductance regulator (CFTR) gene mutation detection system. | Immunology | Immunology | 510(k) | CF InPlex ASR Card; Catalog Number: 95-0501. |  |  | March 31, 2016 | 2016 | Hologic, Inc | Complaints of false positive HET mutations and leaking cards; could lead to unnecessary invasive testing during pregnancy and the associated risks of such invasive testing. | Nonconforming Material/Component | Customer notification letters were sent 03/31/16. Customers were instructed to immediately discontinue the use of the affected lot numbers.If a HET or MUT result was reported with the affected lots, Hologic recommends that customers notify the recipient of the result and request the patient be retested with another test method.Any HET or MUT results that were previously retested and reported out as negative do not require further testing.The firm requests that customers segregate any remaining inventory of all affected lots listed on the Customer Response Form. After segregation is complete reconcile the number of units (ASR Card or IVD Molecular Test Kit) which remain in inventory upon receipt of notification against the number originally shipped to site, and destroy all remaining inventory of the affected lots.Document the following in the appropriate fields on the Customer Response Form:1. If have zero of the affected products on hand check the first box.2. If have some inventory of the affected products on hand check the second box and then document the quantity on hand below, in the Qty On-Hand column.Note: When counting the number of units remaining in inventory count each unit (ASR Card box or IVD Molecular Test Kit), whether partial or full, as a unit.3. Print name, phone and email address, in case the firm needs to follow up.4. Sign and date the form. Signing acknowledges compliance with this recall notification, including destroying all remaining inventory of the affected products.When this documentation is complete fax the CRF to Hologic Technical Support using fax number: 858-410-8250. Alternatively, you can scan the document(s) and email to Hologic Technical Support at molecularsupport@hologic.com.If laboratory has shared any of the affected cards noted in the CRF with another laboratory contact Hologic Technical Support immediately using the phone number below, so that Hologic may contact the other laboratory t |  | 2580 | 2580 total | Distributed in the states of OH, AL, IL, AZ, AR, NJ, TX, GA, CO, DE, NC, MI, FL, ME, CA, IN, OR, KS, KY, MO, WA, NY, OK, WV, CT, IA, LA, RI, MA, and PA, and in the country of Canada. | Open, Classified |  |  | 2580
Z-1576-2016 | 73650 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Drew Scientific Inc. Drew3-PAC. This D3 Reagent Pack includes: diluent reagent 5L, enzymatic cleaner 750 mL, and a lysing reagent 125 mL. |  |  | March 10, 2016 | 2016 | Jas Diagnostics/Drew Scientific | The reagent pack may result in elevated platelet background during Drew3 Hematology Analyzer routine check. | Process control | JAS Diagnostics/ Drew Scientifics sent an Important Product Advisory letter dated March 10, 2016, to all affected customers. The domestic customers were notified via letter and international customers were notified via email. Customers were instructed to contact Drew Scientific Inc. Customer Service or Technical Service at 1-800-433-0945 for product replacement. The firm suggested a temporary workaround of performing five background runs three times per day in order to ensure background counts are running within specification. |  | 7691 | 7,691 | Worldwide Distribution - US Distribution including the states of : AL, AR, CA, FL, GA, IA, IN, KS, KY, LA, MA, MN, MO, MS, NC, NE, NY, SC, TN, and TX. and to the countries of : Bolivia, Cairo, Columbia, Ecuador, Italy, Moscow, Netherlands, Pakistan, Peru, Slovenia, Spain, and Venezuela. | Open, Classified |  |  | 7691
Z-1574-2016 | 73770 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 5.0 mm Periprosthetic Locking Screw Self-tapping 12 mm Sterile; |  |  | April 04, 2016 | 2016 | Synthes (USA) Products LLC | This lot of the 5.0 mm Periprosthetic Locking Screw, Self-tapping, 12mm Sterile, was shipped after the expiration date of May 31, 2012. The label on the product has the correct expiration date. | Process control | DePuy Synthes mailed to customers an Urgent Notice informing them they were voluntarily recalling 5.0 mm Periprosthetic Locking Screw Self-tapping 12 mm-Sterile, part # 02.221.462S, Lot # 5539094, expiration date May 31, 2012, due to the units being shipped after the expiration date of May 31, 2012. Customers were instructed to call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, complete the Verification Section on page 3 by checking the appropriate box and return the form and product to Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. |  | 15 | 15 units | Distributed in IL and IN. | Open, Classified |  |  | 15
Z-1572-2016 | 73886 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON X700 Ultrasound System with software version 1.0.04. |  |  | September 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Due to a communication error between the software and V5Ms transducers rotation function, ACUSON X700 ultrasound systems at software versions 1.1.04, display an IMG_15 error message causes the system to lock up, which requires a reboot to recover system operation. | Software design | September 5, 2015 we initiated a correction via update program US020/15/S where we visited all four units and updated their software to version 1.0.05 |  | 4 | 4 | Distributed in Germany, U.A.E., Hungary and Brazil. | Terminated | April 27, 2016 | 235 | 4
Z-1581-2016 | 73779 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | reVive Light Therapy Pain Relief Device (Model Number: RVPNSYS) |  |  | March 30, 2016 | 2016 | LED Technologies, LLC | Incorrect Power Supply in the package, rendering the device inoperable. | Nonconforming Material/Component | Customers were notified on March 30, 2016 via telephone call. LED Technologies, Inc. provided the consignee with a letter to distribute to customers who purchased the reVive Light Therapy Pain Relief device, and included instructions to receive a replacement of the correct power supply. |  | 32 | 32 units | Distributed to 1 consignee located in the United States. | Open, Classified |  |  | 32
Z-1583-2016 | 73512 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | RESTORIS PST RIO Offset Shell Impactor |  |  | February 23, 2016 | 2016 | Mako Surgical Corporation | Potential disassociation of the orientation pin. | Nonconforming Material/Component | On March 4, 2016 Stryker Orthopaedics initiated a voluntary product recall of the RESTORIS PST RIO Offset Shell Impactor informing the customer in an Urgent Product Recall to do the following: 1. Inspect all metal instruments for any damage or difficulty disassembling/assembling instruments2. Return all affected products at their location to Stryker C/O Stericycle, 2670 Executive Dr. Suite A, Indianapolis, IN, 46241, Attn of Event 8763, Clayton Odor.3. Complete and sign the Business Reply Form and send it back via e-mail or fax within five business days.Contact 866-241-0741 or via e-mail strykerortho8763@stericycle.com |  | 96 | 96 | Nationwide Distribution -- AL, CA, CO, CT, FL, GA, IA, ID, IL, IN, LA, MA, MD, MI, MN, MS, NC, NH, NJ, NY, OH, OK, OR, PA, RI, TN, TX, VA, WA, and WI. | Open, Classified |  |  | 96
Z-1582-2016 | 73676 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Pleur-evac Chest Drainage System, Dry Suction/Dry Seal Control, REF A-6000-08LF, Rx Only, Teleflex Medical. |  |  | March 30, 2016 | 2016 | Teleflex Medical | The label on the Tyvek bag is missing. This is used by customers to identify material code, lot number, and expiration date once the product is removed from the shipper box. | Error in labeling | Teleflex sent an Urgent Medical Device Recall Notification dated March 30, 2016 to customers and distributors. The letter identified the affected product, problem and actions to be taken. The letter requested that they discontinue use and quarantine any products. The letter instructed customers to complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com.. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 12124 | 12,124 ea. | Worldwide Distribution - US Nationwide in the states of: AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, LA, MD, MA, MI, MS, MO. MT, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WV, Puerto Rico, Belgium, Canada, Chile, Dominican Republic &amp; Germany | Open, Classified |  |  | 12124
Z-1585-2016 | 73745 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | X-Ray Diagnostic Table, Model No. YSF-300 |  |  | March 31, 2016 | 2016 | Shimadzu Medical Systems | Report of unintentional movement of table to reverse tilting. | Device Design | Shimadzu sent an Urgent Voluntary Medical Device Recall Notice to all affected customers on April 8, 2016. Shimadzu Corporation has identified a potential issue in the Shimadzu X-ray Diagnostic Table. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns are instructed to contacted Shimadzu National Technical Support at (800) 228-1429.Shimadzu will send out an updated customer notification letter to inform customers that the liquid penetration may cause an electrical short circuit, causing the table to unintentionally tilt in reverse. If this unintentional reverse tilting occurs, there is a potential risk of injury to the patient and/or the user.For questions regarding this recall call 310-217-8855, ext 139. |  | 185 | 185 units | Nationwide Distribution | Terminated | January 02, 2017 | 277 | 185
Z-1584-2016 | 73581 | No | MQJ | Assisted reproduction micromanipulators and microinjectors. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Integra 3; Model Numbers: 6-54-100, 6-54-110, 6-54-150, 6-54-160 |  |  | March 11, 2016 | 2016 | RESEARCH INSTRUMENTS LTD - SITE B | The units have been fitted with a potentially faulty power supply cable. | Device Design | Research Instruments sent a recall notification dated 03/11/2016, via UPS with confirmed delivery receipt, to all affected customers. The letter identified the affected product, problem and the action needed to be taken by the customer. The letter asked customers to contact Research Instruments First Line Support if they had any questions. |  | 380 | 380 units | Worldwide Distribution - US Nationwide in the states of: MA, NY and countries of : Algeria, Australia, Austria, Bahrain, Belgium, British Virgin Islands, Bulgaria, Canada, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Hong Kong, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Nigeria, Norway, Pakistan, Peru, Poland, Romania, Serbia, Singapore, Slovakia, South Africa, Spain, Sri Lanka, Sweden, Taiwan, Thailand, Tunisia, Turkey, UAE, Ukraine, United Kingdom, and Vietnam, | Open, Classified |  |  | 380
Z-1587-2016 | 73486 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The Livongo Health In TouchÂż Blood Glucose Monitoring System (BGMS), is an Over-The-Counter (OTC) system designed for the self- monitoring of blood glucose by persons with diabetes in home settings to aid in their diabetes management. |  |  | February 18, 2016 | 2016 | Livongo Health | The instructional insert for the In Touch Control Solution for use with the In Touch meter and In Touch strips as a quality control check did not contain the limitation of sensor accuracy: (1) with the use of acetaminophen (Tylenol)(2) with the use of Ascorbic Acid (vitamin C) supplementation(3) in the presence of uric acid | Error in labeling | The recall was conducted in three distinct events. Members were first sent an e-mail of the Recall Notification - Livongo Test Strip Insert (Event 1) containing information regarding the test strip insert discrepancy. Following the e-mail distribution, a hard copy of the same information was distributed (Event 2) to members via United States Postal\ Service (USPS) mail. Lastly, a second e-mail (Event 3) containing the same information was sent to members. The distribution list for the second (hardcopy distribution) and third (e-mail distribution) events was a subset of the initial member distribution. |  | 63000 | 63,000 units | US consignees: Arizona, California, Delaware, Florida, Georgia, Illinois, Kentucky, Maryland, Michigan, Minnesota, Mississippi, Missouri, Montana, New Jersey, New York, Ohio, Pennsylvania, South Carolina, Tennessee, Texas, Virginia and Washington. | Open, Classified |  |  | 63000
Z-1586-2016 | 73433 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AnalystÂż MD Version 1.6.1 and 1.6.2 Software used with the following instruments: API 3200MD" LC/MS/MS System, Instrument Part Number: 5024501; 3200MD QTRAPÂż LC/MS/MS System, Instrument Part Number: 5024500; Triple Quad" 4500MD LC/MS/MS System, Instrument Part Number: 5031257; QTRAPÂż 4500MD LC/MS/MS System, Instrument Part Number: 5031231 |  |  | February 24, 2016 | 2016 | Ab Sciex | Wrong quantitative results may be displayed in a report from the device, which may potentially lead to an incorrect patient diagnosis. | Software design | Sciex sent an Urgent Medical Device Correction letter dated February 29, 2016, to all affected US customers via Fedex on March 2, 2016. International customers were notified beginning on March 3, 2016. The letter informed customers of the issue and requests that they take temporary steps to eliminate the potential for erroneous events. In addition, a software update is in development and is expected to be available in approximately 4 weeks. Customers with questions were instructed to contact SCIEX at 1-289-982-2531. |  | 279 | 279 | Worldwide Distribution - US (Nationwide) and Internationally to the following countries:AUSTRIACANADACHINAFRANCEGERMANYINDIAIRELANDITALYKOREAMALAYSIANETHERLANDSNORWAYRUSSIASINGAPORESPAINSWITZERLANDUNITED KINGDOM | Open, Classified |  |  | 279
Z-1589-2016 | 73974 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE, Revolution CT. |  |  | April 20, 2016 | 2016 | GE Medical Systems, LLC | An out of tolerance torque tool was used on three joints during the manufacturing process. If the X-ray Tube becomes loose on the CT gantry during operation, this could result in beam tracking errors, potential scan aborts and/or a possible image artifact. There is a secondary safety mechanism to prevent components from being expelled from the gantry if the bolts fail. | Process control | GE Healthcare Systems sent an Urgent Medical Device Correction letter dated April 20, 2016, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 26 | 26 (US = 4;OUS = 22) | Worldwide Distribution - US including FL, GA, NC, and TX and Internationally to AUSTRALIA, CANADA, CHINA, FRANCE, GERMANY, ITALY, JAPAN, REPUBLIC OF KOREA, NORWAY, PERU, SWEDEN UNITED ARAB EMIRATES, UNITED KINGDOM. | Open, Classified |  |  | 26
Z-1588-2016 | 73199 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | MAKOplasty Hip Array Kit Laminate 150230 |  |  | November 25, 2015 | 2015 | Mako Surgical Corporation | Stryker Orthpaedics became aware that there is the potential for the 210540 MAKOplasty Hip Array Kit Laminate 150230 to list two part numbers for the RIO Base Array, which is item #5 on the product table. | Component change control | On 12/09/2015, Product Field Action Letters and Product Accountability forms were sent to Branch/Agencies via overnight courier with a return receipt; Hospital Risk Management were sent Notification letters and a Product Recall Acknowledgement form with return receipt on 12/9;/2015. |  | 10 | 10 units. | CO, KY, MN, WA and WV | Open, Classified |  |  | 10
Z-1591-2016 | 73862 | No | MTD | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Pentaray NAV High-Density Mapping ECO Catheter, Catalog No. D128207, D128208, D128209, D128210, D128211, D128212 |  |  | April 14, 2016 | 2016 | Biosense Webster, Inc. | Biosense Webster is recalling the Pentaray Catheters to clarify the contraindication language in the Instructions For Use (IFU) and product labeling for this catheter relative to patients with prosthetic valves. | Use error | The firm, Biosense Webster, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated 4/14/2016 to its customers. The letter described the product, problem and actions to be taken. The letter informs customers that the recall action is NOT a product removal, and that they may continue to use PENTARAY Catheters in accordance with the updated contraindication as stated in the recall letter. Customers are instructed to read and follow the Labeling Correction Notice carefully; pass on this letter to anyone in your facility that needs to be informed; maintain a copy of this letter with the PENTARAY Catheters, and sign and return the attached ACKNOWLEDGEMENT FORM immediately by email: BWI_FA_Coordinator@its.jnj.com, fax: 909-839-8514 or mail: Biosense Webster, Inc., Attn: FCA Coordinator,15715 Arrow Highway, Irwindale, CA 91706. The updated the IFU to clarify the contraindication statement as follows: "Do not use PENTARAY Catheters in patients with prosthetic valves". For questions related to this medical device correction, please contact your BWI sales representative or call 909-839-8500. |  | 59672 | 59,672 units | Worldwide Distribution: US (Nationwide) and countries of: Canada, China, Hong Kong, Indonesia, India, Japan, Korea, Malaysia, New Zealand, Singapore, Thailand, Taiwan, United Arab Emirates, Austria, Belgium, Bulgaria, Switzerland, Czech Republic, Germany, Denmark, Algeria, Spain, Finland, France, United Kingdom, Greece, Hungary, Ireland, Israel, Iraq, Iceland, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Sweden, Turkey, South Africa. | Open, Classified |  |  | 73584
Z-1502-2016 | 73634 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The Spacelabs Healthcare Xhibit Telemetry Receiver (XTR), Model 96280 is used to transmit patient(s) physiological measurements and data to a monitor. |  |  | March 28, 2016 | 2016 | Spacelabs Healthcare Inc | Reports of an abnormally high frequency of squelch events on Xhibit Telemetry Receivers (XTR), which is indicative of the loss of the telemetry data signal. In some cases, the Xhibit Central Monitor was also reported to display an Xhibit Offline error message for all patients being monitored. | Other | On March 28, 2016, a customer letter will be sent via priority service, return receipt requested, to all U.S. customers. On April 4, 2016, a customer letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record. Spacelabs recommends that you brief your staff regarding the failure modes listed above. If you continue to use affected Telemetry Receiver(s), please note:**When there is a loss of signal, a signal loss alarm is generated and a triangular-shaped squelch waveform is displayed for any channel exhibiting signal loss. If unacknowledged, this alarm will escalate to medium priority after five minutes. After an additional five minutes, this alarm will escalate to high priority. This failure is usually transient.**When the XTR goes offline, all telemetry waveform(s) associated with the XTR will disappear from the Central Station and an offline message will appear in each telemetry waveform zone. This failure is not recoverable. Please circulate this notification to all persons affected by the information it contains and add a copy to the appropriate monitor user manual until such time as corrections have been implemented by Spacelabs. Spacelabs will contact you to schedule a convenient time for Spacelabs to update your Xhibit Telemetry Receivers at no cost. If you have any questions about this corrective action, please contact Spacelabs at 1-800-522-7025 and select 2 for Technical Support. |  | 115227 | 217 units of XTRs and 531 units of PCBAs (164 XTRs and 425 PCBAs in US and 53 XTRs and 106 PCBAs outside the US) | distributed in AL, AZ, CA, CO, CT, FL, ID, IN, KY, MI, MN, MO, MS, MT, SC, TX, WY and in the following countries: AUSTRALIA, CANADA, CHINA, FRANCE, NETHERLANDS, PHILIPPINES, ROMANIA, and UNITED KINGDOM. | Open, Classified |  |  | 115227
Z-1498-2016 | 73612 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | iVue and iVue 500 (now called iScan) OCT systems with software version 2016.0.0.127 (USA only) and 2016.1.0.127 (out of US) |  |  | March 16, 2016 | 2016 | Optovue, Inc. | A software anomaly was detected that affects Optovue iVue and iVue 500 with Normative Database that may display images of the same eye rather than right and left if user scrolls on the map. | Under Investigation by firm | Recall Letters were sent to customers on March 22, 2016 to inform them of the recall. |  | 40 | 40 | PA, CA, IL, HI. MO, CO, GA, ME, IN, AZ, WV, VA, FL, AR, UT, NE, IN, NV, LA, TN, KS, OK. Foreign distribution forthcoming | Terminated | April 04, 2017 | 384 | 40
Z-1595-2016 | 73601 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | TCM CombiM |  |  | November 18, 2015 | 2015 | Radiometer America Inc | Radiometer America, Inc. is recalling TCM CombiM on rare occasions they may have an error that can cause the alarm to sound higher than the pre-set maximum volume. | Software design | RadioMeter sent an Product Notification letter out on November 18, 2015, January 10, 2016 and March 29, 2016.The notification sent on November 18, 2015 and January 10, 2016, instructed customers a temporary solution of switching off the analyzer and then on again in order to re-establish the pre-set alarm volume. If alarm sound is higher than the pre-set volume customers are instructed to contact their local Radiometer service representative. If any questions contact Radiometer America Technical Support at 1-800-736--600 opt. 4.The notification sent on March 29, 2016, provided the same instructions and temporary solution. The notification also informed customers that the firm will be in contact with them to discuss the upgrade process. |  | 1488 | 1488 | Worldwide Distribution - USA (nationwide) Distribution to the states of : MO, NE, AL, PA, IL, MN, TX, ID, CO, UT, NV, CA, NY, RI, OR, NJ, AZ, GA, IN, VA, LA, NM, MD and AR., and to the countries of : Canada, Austria, Czech Republic, Denmark, Finland, France, Germany, Hungary, Ireland, United Kingdom, Italy, Netherlands, Norway, Spain, Sweden, United Arab Emirates, Kuwait, Qatar, Saudi Arabia, Oman, Jordan, Croatia, Malta, Iceland, Brazil, Thailand, Philippines, Peru, Honduras, Mexico, Kenya, Israel, Romania, Russia, Poland, Turkey, Australia, New Zealand, China, Hong Kong, South Africa, India, Korea, Singapore and Taiwan. | Open, Classified |  |  | 1488
Z-1592-2016 | 73854 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound imaging system with software version VB10C and using transesophageal (TEE) transducer Z6Ms, V5Ms or V7M. |  |  | April 08, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | While imaging with a transesophageal (TEE) transducer (Z6Ms, V5Ms, or V7M), user may lose the ability to control the color region of interest, the pulsed wave or continuous wave Doppler gate or cursor, the M-mode cursor, the 2D field of view, or the RES region of interest with the trackball. | Software design | Under Siemens update program US013/16/S a Customer Advisory Notification will be mailed to all affected customers by Certified mail with return receipt beginning March 30, 2016. Letters provide customers information on how to avoid the issue until the software patch is installed. |  | 87 | 87 | Worldwide distribution. US nationwide, Austria, Canada, China, Germany, India, Italy, Spain, Sweden, and Switzerland. | Open, Classified |  |  | 87
Z-1594-2016 | 73742 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CARESTREAM Image Suite V4; Image Suite V4: DICOM STORE SCP: REF/Catalog # 1056191, DICOM STORE SCP/FOR IMAGE SUITE V4: REF/Catalog # 6566988; Image Suite V4 Bundles: IMAGESUITE STANDALONE PACS: REF/Catalog # 1741289, IMAGESUITE STANDALONE PACS FOR INDIA: REF/Catalog # 1741297; Carestream PRO (Image Suite V4): PRO Medical Wireless GOS System-Desktop: REF/Catalog # 1741891, PRO Wireless System Laptop: REF/Catalog # 1741925, PRO Wireless System - w/o Computer: REF/Catalog # 1741933, PRO Tethered System Desktop: REF/Catalog # 1741941, PRO Tethered System Laptop: REF/Catalog # 1741958, PRO Tether System - w/o Computer: REF/Catalog # 1741966, PRO Fixed System - w/o Computer: REF/Catalog # 1741974, PRO Medical Wireless CsI System-Desktop: REF/Catalog # 1741982, PRO Wireless System Laptop: REF/Catalog # 1742006, PRO Wireless System - w/o Computer: REF/Catalog # 1742014, PRO Tethered System Desktop: REF/Catalog # 1742022, PRO Tethered System Laptop: REF/Catalog # 1742055, PRO Tether System - w/o Computer: Catalog # 1742063; PRO Fixed System - w/o Computer: REF/Catalog # 1742089 -- Made in U.S.A. by: Carestream Health, Inc., 150 Verona Street, Rochester, NY 14608 |  |  | April 15, 2016 | 2016 | Carestream Health Inc | Carestream Health received a complaint related to CARESTREAM Image Suite 4 from a foreign hospital stating that the annotation on the overlay is displayed as "L (Left)", when it should be "R(Right)". | Software design | Carestream Health sent an URGENT MEDICAL DEVICE CORRECTION letter dated April 6, 2016, and Consignee Notification Acknowledgement form to the US consignees on April 15, 2016, via Certified Mail, Return Receipt Requested. Foreign consignees are being notified based on the individual country regulations. Customers who have this Image Suite configuration but do not use it for CT or MR Imaging do not need to take any action. Customers who use the product for CT or MR Images are asked to make users aware of the problem and take additional steps to check correct orientation of any markers, until the software is updated. Customers with questions or concerns were instructed to contact the Carestream Customer Care Center in the US at 1-800-328-2910. Customers outside the US were instructed to contact their local Service Support number.For questions regarding this recall call 1-800-328-2917. |  | 306 | US: 37 units, Foreign: 269 units | NY and OH for further distribution nationwide; Foreign distribution to the following countries: Afghanistan, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia &amp; Herzgovena, Brazil, Bulgaria, Cambodia, Canada, Chad, Chile, China, Colombia, Croatia, Czech Rep., Djibouti, France, Germany, Greece, Hungary, India, Indonesia, Iran, Iraq, Ireland, Italy, Jamaica, Japan, Jordan, Kuwait, Kyrgyzstan, Laos, Latvia, Lebanon, Liechtenstein, Lithuania, Macao, Malaysia, Mexico, Montenegro, Myanmar (Burma), Nepal, Netherlands, New Zealand, Niger, Norway, Oman, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Scotland, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Taiwan, thailand, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Vietnam. | Terminated | March 16, 2017 | 335 | 2855
Z-1600-2016 | 73884 | No | LGW |  |  | Neurology | PMA | Nuvectra, Algovita Spinal Cord Stimulation System Clinical Programmer, Model 4500. For use with Algovita Spinal Cord Stimulation Systems. The version of the software in the Clinician Programmer is v1.1.5. |  |  | April 06, 2016 | 2016 | Nuvectra | If used, the optional Swap feature copies programming parameters from the external pulse generator (EPG) as well as unintended calibration data to the implantable pulse generator (lPG). This may result in the user being unable to recharge the lPG.This safety notice only applies to CPs Model 4500. | Software design | Nuvectra sent an "Urgent Field Safety Notice" letter dated April 6, 2016 to customers. The letter identified the affected product, problem, recommendations and Mitigations, and actions to be taken. |  | 17 | 17 OUS | Germany | Open, Classified |  |  | 17
Z-1597-2016 | 73767 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | INFINITI Ultrasound Pack. Catalog Numbers: 8065751716, 8065751717, 8065751718, 8065751720, 8065751721, 8065751723. |  |  | October 06, 2014 | 2014 | Alcon Research, Ltd. | The recalling firm received complaints related to leakage from the tubing near the white luer of the irrigation line for the INFINITI Ultrasound Packs. | Process change control | The recalling firm notified affected customers on 10/6/2014 via a customer letter with an attached Medical Device Removal Response Form. On 10/23/2014, a second notification was sent to 76 customers who did not respond to the initial notification. |  | 6597 | 6597 packs | Nationwide Distribution. | Terminated | March 31, 2017 | 907 | 6597
Z-1596-2016 | 73585 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK 2 Gram Positive Susceptibility Test Cards (AST-P605), REF 22 325. 20 Cards/Kit. bioMerieux, Inc, Durham, North Carolina 27712. |  |  | February 24, 2016 | 2016 | Biomerieux Inc | Product may lead to false susceptible results of Staphylococcus spp., Enterococcus spp., and S. agalactiae to Minocycline. | Other | Biomerieux, Inc. sent an Urgent Notice Product Correction on February 24, 2016, to all affected customers. The recalling firm notified their foreign subsidiary of the Field Safety Corrective Action (FSCA) by email. The local subsidiary in the country impacted has informed users impacted and provided instructions to resolve the issue.Customers were instructed to confirm the Field Safety Corrective Action letter has been distributed and reviewed by all appropriate personnel within theirorganization.Âż For the referenced test kit (VITEKÂż 2 ASTP605, Lot 485368220), refrain from reporting Minocycline (MNO) results for any organism until a new lot is implemented in your laboratory.Âżo Please note that a bioART rule can be created to suppress the analysis/reporting of Minocycline (reference VITEKÂż 2 Online Software User Manual). The bioART rule can be disabled upon implementation of a new ASTP605 lot number.Customers were instructed tostore the letter with their bioMÂżrieux instrument documentation.Âż Complete the attached Acknowledgement Form and return it to their local bioMÂżrieux representative.For questions regarding this recall call 314-731-8526. |  | 6740 | 337 Kits of 20 Cards/Kit | Worldwide Distribution to China  Taiwan region; No US/military/government distribution. | Terminated | May 26, 2017 | 457 | 6740
Z-1599-2016 | 73668 | No | IXW | Automatic radiographic film processor. | Radiology | Radiology | 510(K) Exempt | Smart CR Digital Computed Radiography System; Model Number: CR-IR346RU |  |  | April 01, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | Reports of failure of the device's power supply. | Device Design | Fujifilm sent an Medical Device Correction letter dated March 22, 2016, via FedEx to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Instructions for turning the unit off when not in use as an interim action. FMSU will conduct a field replacement of the power supply in all installed units.For further questions, please call (888) 385-4633. |  | 2062 | 2,062 devices | Nationwide Distribution. | Open, Classified |  |  | 2062
Z-1598-2016 | 73947 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Aera/Skyra Systems, whole body magnetic resonance imaging systems, Model numbers 10432914, 10432915. Docking station top cover (part no# 10591430) of the removable patient table. |  |  | April 13, 2016 | 2016 | Siemens Medical Solutions USA, Inc | The docking station top cover of the removable patient table may contain a magnetic metal plate. When docking/undocking the table, the magnetic top cover may become attracted to the magnet. | Vendor change control | Siemens mailed to customers on April 13, 2016, a Customer Safety Advisory Notice, MR013/16/S, to inform them of the actions to take to prevent a possible hazard to persons and equipment. Customers were informed that they will be contacted by one of Siemens' service organization who will schedule an appointment to perform a correction, which is to exchange the top cover. |  | 8 | 8 systems | US nationwide distribution to OH, FL, TN, MN, OH, NY, VA, and MN. | Terminated | April 17, 2017 | 369 | 8
Z-1536-2016 | 73744 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Dialog+ Hemodialysis machines: Conductivity Sensor (BIC) Item 3456103A |  |  | April 01, 2016 | 2016 | B Braun Medical Inc | Potential leakage of conductivity sensors of the Dialog Dialysis machines. | Process control | B. Braun mailed an Urgent Medical Device Correction notice to affected customers informing them of the Dialog+Hemodialysis System due to the potential for Bicarbonate and End conductivity cells (3456102A and 3456103A respectively) to exhibit cracks. Notice provided customers with actions for Qualified Service Technicians, Clinical Personnel Administering Dialysis treatment and adverse event reporting. All Dialog+Hemodialysis machines will be inspected at the customer's site by the customer Biomed, who has been trained and qualified by B. Braun Avitum. The inspection of each machine will be performed in accordance to a Field Service Information (FSI). Customers with questions were instructed to contact the Technical Support Team at 800-621-0445 or the Customers Service Team at 800-848-2066.For questions regarding this recall call 610-997-4717. |  | 147 | 147 | Nationwide Distribution to ND, OH, CA, VA, FL, LA, SC, IL, UT, AZ, KY, IA, WI, OK, TN, SD, AL, NJ, TX, NM, NY, KS, WA, MO, CO, PA, CT, DE, GA, MD, VA, Puerto Rico, and St. Thomas USVI | Terminated | July 11, 2017 | 466 | 1033
Z-1605-2016 | 73780 | Yes | NVZ |  |  | Cardiovascular | PMA | Medtronic CareLink" Monitor (2490C) and Medtronic CareLink Express" Monitor (2020B) |  |  | March 31, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | A recent firmware update developed by Medtronic for the 2490C CareLink Monitors and 2020B CareLink Express Monitors included incorrect data on the country analog modem dial-up configuration table used to dial into the CareLink Network. This incorrect data resulted in the removal of the 0 prefix necessary for select countries as well as an incorrect alignment of phone numbers for other select cou | Software design (manufacturing process) | Medtronic sent an "Urgent Field Safety Notice" dated April 2016to their customers. The letter was addressed to Healthcare Professional. The letter identified the affected product, problem and customer actions to be taken. For questions contact your local Medtronic Representative. |  | 5038 | 5038 | AUSTRIA, AUSTRALIA, BELGIUM, SWITZERLAND, GERMANY, DENMARK, FRANCE, GREAT BRITAIN, HUNGARY, JAPAN, NETHERLANDS, NEW ZEALAND, POLAND, PORTUGAL, SWEDEN. | Open, Classified |  |  | 5038
Z-1601-2016 | 73907 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Dynamics; Kinetdx Picture Archiving and Communications System |  |  | April 14, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens has identified changes in default carry forward workflow using VA10B_HF03 version of Syngo Dynamics. If the template design has measurements inside a carry forward-configured report section, this will result in carrying forward the entire section of the report including data from previous studies. This may result in the physician making analysis based on the old measurements. | Software design | Siemens mailed out a "Customer Advisory Notice" #SY009/16/P to customers informing them of the issue and what to do if it was to occur. Siemens informed customers that they are working on the "carry forward" feature enhancement and is planning to release a software update VA10B_HF06 and they should follow the instructions in the advisory notice until the update is performed. |  | 84 | 84 units | Nationwide: AZ, CA, IL, IN, MA, MD, MI, MN, NJ, NY, OH, OR, PA, TN, TX, UT, WA, WI | Terminated | March 20, 2017 | 340 | 84
Z-1607-2016 | 74005 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Toshiba American Medical Systems Celesteion, Item Code: PCA-9000A/2, 9000A/3B, 9000A/2B |  |  | April 24, 2016 | 2016 | Toshiba American Medical Systems Inc | Potential degradation of detector performance. | Device Design | Toshiba sent an Urgent Medical Device Correction letter dated April 24, 2016 to all affected customers. The letter informed customers that Toshiba America Medical Systems is bringing attention to a potential problem with the Celestion PET/CT System. The letter informed the customers of a potential degradation of detector performance and informed the customers of the problems and the corrective actions to be taken. Customers were instructed to complete and return the attached form and fax it to (877) 349-3054 or be sent via email to regulatoryaffairs@toshiba.com. Customers with questions were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 3 | 3 units | US Distribution to WI and NV. | Open, Classified |  |  | 3
Z-1602-2016 | 73869 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Medtronic Midas Rex Dissecting Tool Legend 10 cm 4mm BA DIAM. |  |  | April 08, 2016 | 2016 | Medtronic Sofamor Danek Usa, Inc - Dallas Distribution | Incorrect color coding and attachment reference were included on the labeling of 10BA40D dissecting tools with lot number H5237248. | Labeling mix-ups | A notification letter was provided to each consignee, translated as necessary, on April 8, 2016. |  | 81 | 81 tools | Distributed in France, Ireland, Netherlands, Portugal, Switzerland, and United Kingdom. | Open, Classified |  |  | 81
Z-1604-2016 | 73641 | stretcher, patient restraint | NZD | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Poly-bags containing the strap labeled with "Adult" sticker. |  |  | March 10, 2016 | 2016 | Ferno-Washington Inc | On March 9, 2016, the firm was contacted regarding the possibility of incorrect labeling on the outside poly- bag of the traction splint straps. The firm investigated and determined its supplier had labeled the outside packaging for the pediatric and adult traction splints straps incorrectly. The packaging labeled Adult contained a pediatric splint strap and the packaging labeled Pediatric conta | Process control | Ferno-Washington, Inc , has contacted all customers via telephone on March 10, 2016. All products will be replaced with new products.For further questions, please call: (937) 283-2885. |  | 270 | 270 | Worldwide Distribution - US Distribution to the states of : AK ,AZ ,CA, CO, FL ,GA ,HI, IL ,IN ,KY ,MA ,MI ,MS ,NC ,NE ,NJ ,NY ,OH,ON,TX ,VA ,WI , and WV. There are no government accounts for this recall. There is one Canadian account for this recall. There are no Mexican accounts for this recall. and to the countries of : UAE, MALAYSIA, UNITED KINGDOM, SLOVAKIA, SOUTH AFRICA, QATAR and AUSTRALIA.. | Open, Classified |  |  | 423
Z-1606-2016 | 73631 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit, Model 8015 |  |  | March 24, 2016 | 2016 | CareFusion 303, Inc. | The Alaris PC units model 8015 may display a system error code 133.6080 due to failure with the super capacitor (C245) at power up on the Alaris PC unit logic boards. | Nonconforming Material/Component | The firm, Carefusion, sent a "Medical Device Recall Notification" letter dated March 30, 2016 to its customers. The letter described the product, problem, the potential risk involved and the required action for users. The customers were instructed to do the following: If you observe System Error code 133.6080, remove the PC unit from use and contact the CareFusion Support Center, and promptly complete and return the enclosed Customer Response Card. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018 from 7am-4pm PST, M-F, or by email at supportcenter@carefusion.com. Customers with clinical inquiries, product complaints, or clinical troubleshooting are instructed to contact Customer Advocacy at (888) 812-3266, 24 hrs/day, 7 days/week or by email at customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3229, 6am-5pm PST, M-F or by email at dl-us-inf-techsupport@carefusion.com. |  | 23397 | 23,397 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Great Britain and Canada. | Open, Classified |  |  | 23397
Z-1608-2016 | 73318 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TLCÂż Self Retaining Retractor System Contents 1-RETRACTOR FRAME 6-ELASTIC STAYS STERILE EO Rx ONLY |  |  | February 17, 2016 | 2016 | Applied Medical Technology Inc | The firm received a complaint of an open pouch which breaches the sterilization validation. | Packaging | AMT sent an Urgent Medical Device Recall letter dated February 17, 2016 to their customers. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Service at 440-717-4000. |  | 5334 | 5,334 units | Product was shipped to the following states: AL, AR, AZ, CA, CT, FL, GA, IA, IN, KS, KY, MA, MD, MI, MN, MO, NC, NE, NH, NJ, OH, PA, SC, TN,TX, UT, VA &amp; WI.Product was also shipped to the following countries: Austria, Denmark, Germany, Italy, Japan, Norway, Spain, Sweden, Switzerland &amp; United Kingdom. | Open, Classified |  |  | 12451
Z-1620-2016 | 73882 | No | FFT | Stomach pH electrode. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Given Imaging BravoÂżÂż pH capsule delivery device, 1-pack. |  |  | April 15, 2016 | 2016 | Covidien LLC | Potential of an allergic reaction in a patient with nickel sensitivity. | Device Design | The firm, Medtronic, sent an "URGENT: Medical Device Recall" letter dated 4/15/2016 to its Customers via Federal Express. The letter describes the product, problem and actions to be taken. The customer was instructed to promptly forward the information from the letter to recipients, if you have distributed any of the affected devices; and complete and return the Acknowledgement and Receipt Form by email to: Quality Compliance MITGFCA@Covidien.com or Fax to: (203) 492-7719. Medtronic is adding the following statement to the BravoÂż pH monitoring system User Manual and product package insert so that physicians are aware of the nickel content and the potential of an allergic reaction in a patient with nickel sensitivity. "The Bravo capsule contains a trocar needle that is made of stainless steel. Use caution in patients with known sensitivities or allergies to the metals that are contained including chromium, nickel, copper, cobalt and iron. The BravoÂż pH test lasts from 48-96 hours."If you have any questions, please contact Technical Service at 1-800-448-3644, option 3. |  | 1305 | 1,305 units | Worldwide Distribution: US (nationwide) and countries of: Argentina, Australia, Belgium, Brazil, Canada, Cyprus, El Salvador, Greece, Hong Kong, Israel, Mexico, Panama, Singapore, South Korea, and Thailand. | Terminated | July 13, 2017 | 454 | 14200
Z-1618-2016 | 73594 | No | LAO |  |  | Toxicology | 510(k) | Syva Emit Methotrexate Assay, is a homogeneous enzyme immunoassay intended for use in the quantitative analysis of methotrexate in human serum or plasma. |  |  | March 08, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics confirmed that Syva Emit Methotrexate Application (DVMTX.1) for Dimension Vista System does not meet On-Board Stability claim of 20 hours. Testing using reagent at the end of the 20 hours On-Board Stability has shown that values of methotrexate for patient and quality control samples increased. | Unknown/Undetermined by firm | Siemens mailed an Urgent Medical Device Correction letter to customers informing them that they were withdrawing the Emit Methotrexate Application on the Dimension Vista System. Customers were asked to contact their local Siemens sales representative for alternative testing solutions for Methotrexate. Siemens advised customers to change the On-Board Stability for the Methotrexate application on the Dimension Vista System to four hours: alternatively, customers can run the Methotrexate application in batch mode, along with QC. It was also requested in the notification that customers complete and return the Field Correction Effectiveness Check Form that was attached with the notification within 30 days. |  | 18364 | 18364 units | Units were shipped to MI, FL, NJ, NY, DC, CA, GA, MI, MT, VA, NC, TX, PA, Australia, Belgium, Czech Republic, Finland, Japan, Republic of Korea, Italy, Saudi Arabia, Slovakia, Slovenia, Spain and Switzerland. | Terminated | August 07, 2017 | 517 | 18364
Z-1616-2016 | 73657 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Atlas Spine Torque Limiting Handle 6N-m, P/N 70089 Model # ITN-001-000 |  |  | July 26, 2013 | 2013 | Atlas Spine, Inc. | The device was not appropriately set to mitigate the risk of implant fracture upon rotation inside the disc space. | Device Design | Atlas Spine, Inc. informed all consignees of the recall and sub-recall on July 26, 2013, in person, electronically, and/or via email. Firm takes the responsibility of tracking of set return and inventory control of the product.For questions regarding this recall call 561-354-1108, ext 102. |  | 16 | 16 | Nationwide Distribution to FL, MO, and WV. | Completed |  |  | 16
Z-1621-2016 | 73737 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Direct Drive Clip Applier CA090, indicated for ligation of tubular structures or vessels in laparoscopic and general surgical procedures. |  |  | April 01, 2016 | 2016 | Applied Medical Resources Corp | Increased customer complaints indicating inconsistent clip application; may lead to unoccluded vessels. | Nonconforming Material/Component | Customer notification letters were sent on 04/01/16 via UPS.The notification requests that customers check their inventory for recalled products, complete the attached Recall Notification Confirmation Form to acknowledge the recall and indicate if the facility is returing or has already used any of the products. If customers are a distributor it is requested that they notify any facilities to which have been distributed the affected product. Customers are asked to return the Recall Notification Confirmation Form to Applied Medical by emailing to recall60698158@appliedmedical.com or faxing to (949) 713-8921. Customers are to return affected product and a copy of the Recall Notification Confirmation Form to Applied Medical.For product return questions contact Karen Mitchell, Sales Operations Supervisor, RGA Dept. at (949) 713-8622 or by email at kmitchell@appliedmedical.com.For regulatory questions contact Lauren Contursi, at (949) 713-8767 or by email at lcontursi@appliedmedical.com. |  | 559274 | 559,274 | U.S. distribution to the following: IL, IN, WY, PA, CA, LA, MN, AZ, OH, OR, CT, WI, NC, AK, NE, MI, KS, FL, MN, TN, SC, MA, MD, KY, TX, GA, TX, NJ, MO, CO, MS, MT, ID, ME, VA, NY, UT, MI, AL, PR, NM, DE, OK, WA, DC, NH, ARForeign distribution to the following; Australia, Spain, New Zealand, Italy, Israel, Estonia, United Arab Emirates, Germany, Netherlands, France, Kuwait, Korea, Qatar, Mexico, Switzerland, Austria, Bahrain, Lebanon, Brazil, Belgium, United Kingdom, Finland, Japan, Argentina, Ireland, Iran, Denmark, Norway, Colombia, Serbia, Uruguay, Sweden, Canada, Costa Rica, South Africa, Czech Republic, Chile, Finland, India, Thailand, Malaysia | Open, Classified |  |  | 559274
Z-1613-2016 | 73977 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NovaPACS  Diagnostic Viewer versions 8.3.7, 8.4.2, 8.4.3, and 8.4.4. Novarad Corporation |  |  | November 16, 2015 | 2015 | Novarad Corporation | The SUV values that are being calculated in the PET/CT fusion tool are incorrect. | Software design | Consignees were notified on 11/16/15 via letter to contact Novarad Technical Support to arrange for an upgrade to a corrected version of NovaPACS software. |  | 2386 | 2,386 | Nationwide. 3 Canadian and 33 foreign consignees. No VA/gov/military. | Open, Classified |  |  | 2386
Z-1622-2016 | 73592 | No | CDQ | Urea nitrogen test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista Blood Urea Nitrogen (BUN) Flex reagent cartridge |  |  | February 16, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has determined that Dimension Vista Blood Urea Nitrogen (BUN) lots 15215AE, 15243BB, 15264BA, 15299BB, 15300BA and 15341AC may exhibit inaccurate patient and/or Quality Control results. The issue has been isolated to Flex reagen cartridges manufactured from specific mold cavities. | Nonconforming Material/Component | Siemens Healthcare sent an Urgent Medical Device Correction letter dated February 2016, to all affected customers.It was determined that Dimension Vista Blood Urea Nitrogen (BUN) lots 15215AE, 15243BB, 15299BB, 15300BA, 15320BB and 15341AC may exhibit inaccurate patient and/or Quality Control results. The issue has been isolated to Flex reagent cartridges manufacture from specific mold/cavity identifies as described in Table 2 of the Urgent Medical Device Correction notification sent to customers.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions, please call (302) 631-0528. |  | 18668 | 18668 units | Worldwide distribution - US (nationwide) distribution including Puerto Rico and to the states of : AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. , and to the countries of : Canada, United Arab Emir, Australia, Austria, Belgium, Switzerland, Bahamas, Czech Republic, Germany, Denmark, Spain, Finland, France, Italy, Netherland, Norway, Portugal, Slovakia, Slovenia, Japan, Republic of Korea, New Zealand. | Open, Classified |  |  | 18668
Z-1617-2016 | 73827 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Arctic Sun Temperature Management System, 5000, 100-120V and 220-240V |  |  | November 09, 2015 | 2015 | C.R. Bard, Inc. | The product may contain electronic components that lead to a premature drainage of the internal control panel coin cell battery responsible for maintaining the system close and Static Random Access Memory (SRAM). This premature coin cell battery drainage could render the device unresponsive upon System startup. | Unknown/Undetermined by firm | BARD sent an Urgent Medical Product Recall to all affected consignees on November 9, 2015. The letter indicates that a technician will contract the firm for the evaluation and potential replacement of the affected electronic components. The technician will fill out the Field Action Effectiveness Check Form after evaluation.Customers with questions were instructed to call 1-770-784-6471 or email bmd@crbard.com.For questions regarding this recall call 770-784-6220. |  | 2613 | 2613 | Worldwide Distribution to AL, AK, AZ, AR, CA, CO, CT, DEL, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Finland, France, Germany, Gibraltar, Greece, Guam, Hong Kong, India, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Luxembourg, Malaysia, Mexico, Netherlands, Norway, Oman, Peru, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom | Terminated | June 06, 2017 | 575 | 2613
Z-1615-2016 | 73801 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur CP Immunoassay System, Siemens Material Number (SMN) 10309710, IVD. --- CLASSIFICATION NAME: Immunoassay Analyzers. |  |  | March 30, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | The firm identified an issue in sample processing when using EZee-Nest Insert Cups due to two inaccuracies that exist in the ADVIA Centaur CP Operators Guide. | Use error | Siemens sent an Urgent Medical Device Correction letter dated March 2016, to all affected customers. The letter was sent to the US customers via Fed Ex beginning on March 30, 2016., and were e-mailed dated March 20, 2016, to regional countries for implementation outside the US. Field service personnel were sent a support bulletin describing the issue and instructing them how to manage customer questions. " Complete and return the Field Correction Effectiveness Check form attached to this letter within 30 days." Please review this letter with your Medical Director.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions, please call (914) 524-2868. |  | 3679 | US: 442 units; Foreign: 3237 units | Worldwide Distribution - USA (nationwide) and to the countries of : Afghanistan, Albania, Algeria, Angola, Argentina, Australia, Austria, Bangladesh, Bosnia Herzeg., Brazil, Bulgaria, Canada, Chile, China, Colombia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hungary, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Kenya, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Mexico, Netherlands, New Caledonia, Norway, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Uganda, United Kingdom, Uruguay, Vietnam, Yemen. | Terminated | March 29, 2017 | 364 | 3679
Z-1623-2016 | 73856 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Distal Femoral Replacement (patient specific, custom made orthopedic implant device). PIN 19960 |  |  | April 12, 2016 | 2016 | Stanmore Implants Worldwide Ltd. | Incorrect component used to manufacture distal femoral replacement. | Mixed-up of materials/components | A notification was sent to the prescribing physician on 04/12/2016. |  | 1 | 1 | US distribution to OH. | Open, Classified |  |  | 1
Z-1626-2016 | 73835 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Launcher Guiding Catheter |  |  | April 04, 2016 | 2016 | Stryker Sustainability Solutions | Device packaging may not be sealed. | Process control | Customer notification letter was sent 04/04/2016.The customer was provided the following instructions:Use of the affected lots of the recalled product should be discontinued. A RecallEffectiveness Check Form has been enclosed. Note that the recall is limited to the specific lot number shipped to facility.Once inventory has been checked complete the Recall EffectivenessCheck Form. The form needs to be completed even if no affected product isfound.Return the completed and signed Recall Effectiveness Check Form utilizingone of the following methods:Return to your local Stryker Sustainability Sales RepresentativeEmail to ssspfa@stryker.com, orMail to:Stryker Sustainability Solutions181 O West Drake DriveTempe, AZ. 85283Attn: Jodie RueckertIf the form indicates that any affected devices remain in inventory, a prepaidshipping label will be issued for the return of the product. Customers will receive acredit for all affected devices returned.Adverse reactions or quality problems experienced with the use of this productmay be reported to:Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433x5555http://www.stryker.com/productexperience/The FDA's MedWatch Adverse Event Reporting program either online, byregular mail, or by fax. |  | 2 | 2 | US distribution to KY. | Open, Classified |  |  | 6
Z-1624-2016 | 73816 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton G5 with software version between v2.41, v2.42, v2.50. |  |  | March 09, 2016 | 2016 | Hamilton Medical, Inc. | After performing the suctioning maneuver, including disconnecting the patient, suctioning , and reconnecting the patient, the preset pattern of ventilation many not continue as expected. | Software design | Hamilton Medical Inc. sent an" Medical Device Field Safety Corrective Action" letter dated March 7, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Hamilton Medical then had to compile a list of to whom they had distributed the affected ventilators and sent letters on April 6, 2016. Letters included the Medical Device Field Safety Corrective Action from Hamilton AG. Letters included immediate actions to be take by use and that reps will be visiting each site to upgrade ventilators to v 2.60 to alleviate the issue.Distributors will update software and will provide an updated Operator's manual as soon as it is available.Questions may be directed to 800-426-6331, Ext, 215. |  | 1115 | 1115 | US distribution only. | Terminated | June 29, 2017 | 477 | 1115
Z-1633-2016 | 73442 | No | NRP | Sirolimus test system. | Toxicology | Toxicology | 510(k) | Dimension clinical chemistry system, Sirolimus (SIRO) Flex reagent cartridge (DF306 SMN 10464331) Lot EB6064, Open Well Instability |  |  | February 09, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received customer complaints regarding QC shifts and imprecision after one day when using DimensionÂż SIRO lot EB6064. They have confirmed SIRO lot # EB6064 does not meet the open well stability claim of 2 days. | Under Investigation by firm | Siemens sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be Taken by the Customer:Discontinue use of and discard the kit lot listed in Table 1.Please review this letter with your Medical Director.Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days.If you have received any complaints of illness or adverse events associated with the product listed in Table 1, immediately contact your local Siemens Customer Care Center or your local Siemens technical support representative.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions, please call (302) 631-0528. |  | 394 | 394 | Worldwide distribution and US including Puerto Rico and to the states of : TX, FL, CA, LA, DC, GA, IL, AR, WI, MI and MD. | Open, Classified |  |  | 394
Z-1648-2016 | 73669 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Covidien Kangaroo Connect Enteral Feeding Pump |  |  | April 01, 2016 | 2016 | Medtronic | Kangaroo Connect Feeding Pump Occlusion alarms fail to alarm | Device Design | Medtronic(formerly) Covidien notified accounts via Urgent Medical Device Recall letter, dated 4/1/16, that specific serial numbers with software versions 1.12 and 1.13 have a software anomaly. The firm shipped replacement units with software V 1.14. Affected devices must be returned to the firm once the replacement units have been received. A Recalled Product Return Form required to be completed and returned. Questions or concerns, contact Medtronic representative or Customer Service at (800) 882-5878. |  | 223 | 223 units | Worldwide Distribution -- USA, Australia, Canada, France, and Singapore. | Open, Classified |  |  | 368
Z-1632-2016 | 73837 | Yes | LWS |  |  | Cardiovascular | PMA | Ellipse ICD (Implantable Cardioverter Defibrillator) Model CD2411-36Q, CD2411-36C, CD1411-36Q, CD1411-36C |  |  | March 31, 2016 | 2016 | St Jude Medical Inc. | St. Jude Medical is recalling the Ellipse ICD (Implantable Cardioverter Defibrillator) due to the potential inability to deliver high voltage therapy. | Other | A customer letter dated 4/7/16 was sent to all customers to inform them that St. Jude Medical is recalling the Ellipse ICD (Implantable Cardioverter Defibrillator) due to the potential inability to deliver high voltage therapy. The letter informs the customers that the recall involves information related to 127 units of our Ellipse ICD device distributed worldwide but not implanted in patients. Customers with any questions about the advisory are instructed to contact their local Sales Representative or St. Jude Medical Technical Services at 800-722-3774, 24 hours a day, seven days a week. |  | 127 | 127 units | Worldwide Distribution - US Nationwide in the states of CT, NJ, SC, FL, GA and countries of: Finland, France, Germany, Great Britain, Greece, Israel, Italy, Japan, Kuwait, Lebanon, Netherlands, Northern Ireland, Poland, Romania, Slovakia, Spain, Tunis, United Arab Emirates, and Vietnam. | Open, Classified |  |  | 127
Z-1634-2016 | 73846 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | CinchLock SS Knotless Anchor with Inserter; |  |  | April 12, 2016 | 2016 | Stryker Corporation | Stryker Endoscopy is recalling the CinchLock SS Knotless Anchor with Inserter and the CinchLock Flex Anchor with Inserter after 13 reports of s of an internal wire breakingupon deployment of the anchor. | Device Design | Urgent Medical Device Recalls letters were sent on April 12, 2016 to US customers by certified mail. Stryker International distribution sites will be notified by e-mail. Letters describe the reason for recall with the affected part and lot numbers. Any recalled product should be segregated and returned to Stryker. |  | 3100 | 3100 total device, both model numbers. | Distributed US (nationwide) and in Australia, Switzerland, Netherlands, Spain and Sweden. | Open, Classified |  |  | 3100
Z-1638-2016 | 73578 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Ascom Mobile Monitoring Gateway (versions 4.1.1 and 4.2.0) and Ascom Unite Connect for Nurse Call (version 2.2.0). |  |  | December 10, 2015 | 2015 | Ascom US, INC. | Due to a malfunction of software, the secondary module will not automatically take over messaging if the primary module fails after a software upgrade to one of the other affected versions. Messages from GE Carescape/Nurse Call will not forward to handset until primary module is restored. | Software design | On January 4, 2016 , the firm is initiating a field corrective action to update the software in the installed base and resynchronize the primary and secondary servers. All customers were sent a letter January 4, 2016, informing them of the voluntary recall of Ascom Mobile Monitoring Gateway and Unite Connect for Nurse Call. The reason stated for why the recall was initiated was that the secondary module would not take over the messaging function when the primary module failed, following a software upgrade. Ascom instructed customers to check if their software was one of the affected versions they should contact their local Ascom representative to review the MMG configuration and determine if the system is subject to malfunction. The same instructions were directed towards those who were not able to determine their software version. If it is determined that the program is subject to malfunction, the primary and secondary modules should be re-synchronized. A test of module redundancy should be conducted on re-synchronized modules. The second module should take over for the primary module within thirty seconds. Model redundancy should be tested frequently and following system upgrades. For more information call (919) 535-6143 , 8am till 4:30pm. |  | 319 | 319 | CA, CT, FL, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MO, NJ, NM, OH, OK, PA, TN, TX, VA, WA, WI and Hawaii.Australia, Austria, Belgium, Canada, France, Germany, Netherlands, Poland, Qatar, Sweden, Switzerland, and United Arab Emirates. | Open, Classified |  |  | 319
Z-1637-2016 | 73873 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Axial Torque Limiting Driver (1st generation) |  |  | February 02, 2015 | 2015 | Bradshaw Medical Inc. | A component failure could cause a situation where the applied torque was significantly higher than the calibrated torque setting. | Device Design | Consignees were sent a Bradshaw Medical "Product Withdrawal Notice" letter, date marked starting February 2, 2015. The letter specified the Part Number, Lot Number (s) and Shipped date for each consignee. The letter described the product and advised consignees to discontinue use of the product. Requested consignees to complete and return the Customer /Distribution verification Form to BMI Regulatory Affairs FAX (262) 925-1380 and to further send the "Withdrawal Notice and Customer Verification Form" to all customers to whom they have distributed the product. |  | 3760 | 3,760 units | US Nationwide Distribution in the states of: CA, FL, IN, MA, MI, OH, PA, TN UT, VA and the country of UNITED KINGDOM | Open, Classified |  |  | 3760
Z-1627-2016 | 73908 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | GEM(TM), FLOWCOUPLER(R), 2.0 mm, STERILE EO, Rx only, 1 FLOWCOUPLER, REF GEM2752-FC |  |  | April 21, 2016 | 2016 | Baxter Healthcare Corp. | Instructions for use booklet may puncture the outer Tyvek lid. | Package design/selection | Baxter sent an Urgent Product Recall letter dated April 21, 2016 to customer. The letter identified the affected product, problem and actions to be taken. The letters requested that direct consignees contact Synovis MCA to arrange for return and credit. Companies that further distributed the product were asked to notify other facilities or departments. For questions contact Synovis MCA at 800-510-3318 or 205-941-0111. |  | 315 | 315 units | Nationwide and Germany, Sweden, Netherlands, Israel, Denmark, United Kingdom, Canada | Open, Classified |  |  | 2058
Z-1654-2016 | 73667 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA CHEMISTRY ENZYMATIC CREATININE_2 |  |  | March 17, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has become aware of N-Acetylcysteine (NAC) and Metamizole (Dipyrone) interference with Trinder and Trinder-like reaction assays. Siemens has confirmed that falsely depressed results may occur on samples drawn from patients receiving N-Acetylcysteine (NAC) or Metamizole. NAC is the accepted antidote for acetaminophen toxicity and is justified in patients at signific | Device Design | Siemens mailed an Urgent Medical Device Correction letter to customers making them aware of the problem. Special instructions were also given for the customer to review the letter with their Medical Director and ensure Venipuncture occurs before drug administration of NAC or Metamizole as indicated in the letter. Baseline assay values before administration of NAC or Metamizole therapy would not be affected. Customers were asked to complete and return the Field Correction Effectiveness Check Form attached with the letter within 30 days. Furthermore, if customers received any complaints of illness or adverse events associated with the specified lots, they were to immediately contact their local Siemens Customer Care Center or local Siemens technical support representative. |  | 137422 | 137422 units total | Nationwide: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, Puerto Rico, St. Thomas, VI, Guam, Pacific Mariana Islands | Open, Classified |  |  | 137422
Z-1650-2016 | 73832 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 64 Multi-Slice CT Scanner System (consist if two variants: NeuViz 64e, NeuViz 64i) |  |  | May 02, 2015 | 2015 | Neusoft Medical Systems Co., Ltd. | It was found by R&D; in April, 2014 that there was a defect in NeuViz 64 system software version 1.0.5+P09. The icons of side decubitus didn't meet the patient position description, when the scanner position was set to "Right" in the system setting - scanner options. | Device Design | An upgraded software version (1.0.5.2619+P11) was released on May 7, 2015. The FCO was released to the service engineers of Neusoft Medical Systems USA, which is subsidiary of the reporting firm, and the engineers performed FCO for each affected device together with the distributors' engineers. The corrections (FCOs to upgrade software) on all the affected systems in the USA have been completed. At present (4/27/16) there is no longer this issue on the systems, so no further actions are needed for customers/distributors. However, the recalling firm shall inform the direct accounts of this fix before May 5, 2016. The recalling firm will submit the communication evidence ASAP once they have evidence of the communication. For further questions, please call (281) 453-1205. |  | 7 | 7 units | US Distribution - Including Puerto Rico and the states of IL, CT, SC, NE | Open, Classified |  |  | 7
Z-1647-2016 | 73413 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Vitek 2 Gram Positive Susceptibility Test Cards (AST-P646 REF 420144), 20 cards/kit. |  |  | February 24, 2016 | 2016 | Biomerieux Inc | Product insert fails to identify performance limitation related to EUCAST breakpoints for Teicolplanin. | Labeling Change Control | Biomerieux sent an Urgent Product Correction Notice letter via email on February 24, 2016 to affected customers . The letter identified the affected product, problem and actions to be taken. Customers were asked to complete the attached Acknowledgement Form (Attachment A) and return it to Biomerieux, Inc. For questions contact your local bioMÂżrieux representative. |  | 1035 | 1,035 kits | Foreign distribution in the countries of Czech Republic, Hong Kong (China), Hungary, Japan, Poland, South Africa, Switzerland, and Turkey. | Open, Classified |  |  | 37476
Z-1577-2016 | 73701 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Battery Pack Used with the Capnostream"20 and Capnostream"20p Patient Monitor, Battery Packs Model Number - 016400, 010520. |  |  | April 13, 2016 | 2016 | Medtronic | Medtronic Respiratory &amp; Monitoring Solutions Announces a Voluntary Recall of a Battery Pack Used with the Capnostream"20 and Capnostream"20p Patient Monitor Due to the Battery Pack Might Experience Thermal Runaway with the Potential of Fire Risk to the Monitor. | Process change control | Medtronic sent an Urgent Medical Device Recall letter dated April 2016, by postal mail or courier to all affected customers. The letter instructed users to remove and properly dispose of current batteries. If product was further distributed customers were instructed to forward the recall letter. Customers were asked to completed and return the Acknowledgement and Receipt Form by fax or email. Customers with questions should contact Technical Services at 800-635-5267, option 1, ten 1 and 1 or by email at HQTSWEB@COVIDIEN.com.For questions regarding this recall call 303-876-8859. |  | 9817 | 9,817 batteries (8,170 with monitors and 1,647 as spare parts) | Nationwide Distribution to AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. Product was also distributed worldwide. There is 1 VA customer. | Open, Classified |  |  | 9817
Z-1669-2016 | 73765 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Lumenis M22 System Model Number: GA-0005400 (M22 IPL + YAG Module) with Acne Filter (KT-1014971). |  |  | November 17, 2015 | 2015 | Lumenis Ltd | Lumenis Ltd Announces a Field Action of the M22 IPL Acne Filters for the Lumenis M22 IPL Hand Piece due to the Risk of Superficial Burns When Using the Device. | Device Design | Customers were notified on November 17, 2015 by Customer Notification Letter and customers were asked to return the device. |  | 33 | 33 filters | Distributed in the states of GA, CA, CO, CT, KY, LA, MA, MN, NJ, NY, RI, and SC, and the countries of Italy, Germany, France, and China. | Terminated | November 30, 2016 | 379 | 33
Z-1670-2016 | 73635 | No | CDT | Triglyceride test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry Triglyceride_2, concentrated; TRIG_c; Catalog # SMN # 10697575. |  |  | March 07, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed the ADVIA Chemistry Triglyceride concentrated reagent (TRIG_c) does not meet Instructions For Use (IFU) linearity claims at the upper limit of the assay range as it approaches end of shelf life for lots # 348297 and 359932. | Under Investigation by firm | Siemens mailed an Urgent Field Safety Notice, dated March 2016, to customers confirming the ADVIA Triglyceride concentrated reagent (TRIG_c) does not meet Instructions for Use (IFU) linearity claim at the upper limit of the assay range as it approaches end of shelf life for lots # 348297 and 359932. Siemens asked customers to discontinue use of and discard kits from the two referenced lots, complete and return the Field Correction Effectiveness Check attached to the letter within 30 days. Customers were to review their inventory to determine their laboratory 's replacement needs. Customers are asked to retain the letter for their records and forward a copy to anyone that received the product. Customers may contact Siemens Customers Service or their local technical support. |  | 7433 | 7,433 units (4,228 for lot 348297 and 3,199 for lot # 359932) | Worldwide Distribution -- USA, including the states of CA, TN, UT, OK, FL, MI, MN, NY, CT, NC, KS, TX, CO, WA, NM, and OR; and, the countries of Austria, Australia, Belgium, Brazil, Bulgaria, Denmark, Estonia, France, Germany. Great Britain, Italy, Latvia, Netherland, Norway, Portugal, Romania, Slovakia, Slovenia, Spain, Sweden, Czech Republic, India, New Zealand, Thailand, Turkey, and Russian Fed. | Open, Classified |  |  | 7433
Z-1671-2016 | 73786 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AB Sciex Triple Quad 4500MD LC/MS/MS System, Part No. 5031257; and QTRAP 4500MD LC/MS/MS System, Part No. 5031231. |  |  | March 25, 2016 | 2016 | Ab Sciex | Under certain conditions, mass spectrometers may report incorrect quantitative results. | Under Investigation by firm | AB SCIEX issued Urgent: Medical Device Correction letter, dated 3/25/16, advising user of the problem and advising not to use the system until corrected. The letter also discussed the risk to health, as well as how to recognize if the issue has occurred. Sciex Field Service Employee (FSE) will be sent to customer site to perform the field corrective action on the affected system. Questions regarding the letter should be directed to SCIEX at 1-289-982-2531. |  | 2 | 2 units | Distributed to the state of NC. | Completed |  |  | 2
Z-1686-2016 | 73694 | No | LSE | Epstein-Barr virus serological reagents. | Microbiology | Microbiology | 510(k) | Diamedix Is-EBV-VCA IgG Test Kit |  |  | October 20, 2014 | 2014 | Diamedix Corporation | A gel-like substance was found in kit component, sample diluent , that could potentially compromise product performance. | Unknown/Undetermined by firm | The firm, Diamedix, sent letters dated October 14, October 20 and December 5, 2014 to its customers directly informing them of the correction action. Notification of all domestic customers via Certified Mail. Notification fo the international customers via Federal Express and/or e-mail. With the letter, a replacement bottle of Sample B Diluent was sent and the customer was advised to discard any unused bottles of affected product. New lot number labels were provided to identify kits with the replacement Sample B Diluent. Customers with questions were told they could contact Diamedix Technical Services (1-800-327-4565) if they had any questions or required any additional guidance regarding the correction. |  | 588 | 588 kits | Worldwide distribution: US nationwide including states of: AZ, CA, CO, FL, GA, IL, IN, KY, MA, MI, MO, NC, NE, NH, NJ, NY, OH, OK, OR, PA, RI, TN, TX, VA, WI and Puerto Rico. and Country of: ITALY. | Open, Classified |  |  | 1379
Z-1672-2016 | 74096 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Visualase Cooled Laser Applicator System (VCLAS). The VCLAS system consists of the Visualase Laser Diffusing Fiber (LDF) and the Visualase Cooled Catheter System (CCS). |  |  | May 09, 2016 | 2016 | Medtronic Navigation, Inc. | Potential for the Visualase Cooled Laser Applicator System (VCLAS) catheters to melt during procedures. | Labeling design | Medtronic will be sending a Medical Device Safety Notification letter FedEx on May 18, 2016. Each customer will be asked to complete and return a response form to Medtronic. This is a safety alert notification and does not require return of devices.We are committed to patient safety and appreciate your prompt attention to this matter. If you have any questions regarding this action, please contact Medtronic Technical Services at 1-800-595-9709. |  | 120 | 120 | Distributed within the US, Netherlands, Canada and Australia and to govt/VA/military consignees. | Open, Classified |  |  | 120
Z-1676-2016 | 73981 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Periarticular plating system, cancellous bone screw, partially threaded, hex head, 4.0 mm diameter, 16 mm length |  |  | May 04, 2016 | 2016 | Zimmer Manufacturing B.V. | A small hole may be present in the inner tray of specific lots of sterile Hip and Trauma Implant Screws. | Packaging change control | All distributors will be notified via electronic mail. Distributors with affected lots will also be notified via courier. Hospitals risk managers and surgeons will be notified via recall notification letter, dated May 4, 2016. |  | 99 | 99 | Worldwide distribution. US nationwide, Korea, Japan, Argentina, Brazil, Singapore, Taiwan, China, El Salvador, Australia, India, Hong Kong, Chile, Germany, Italy, France, Spain, UK, Sweden, Canada. | Terminated | July 05, 2017 | 427 | 11400
Z-1614-2016 | 73849 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | NeuroBlate System SideFire Select 2.2 mm Directional Laser Probe, Sterile EO. Model #s SFS000-01,SFS002-01, SFS122-01, SFS222-01, SFS322-01, SFS422-01, and SFS522-01. |  |  | March 25, 2016 | 2016 | Monteris Medical Corp | Monteris received a report that the sapphire lens on a 2.2SF probe fractured during a LITT procedure, with resulting introduction of C02 into the patient's brain. | Device Design | Monteris Medical Corporation field personnel visited each consignee starting March 25, 2016 and April 1, 2016. The devices were physically removed from each site. Customers from whom the devices were removed will either receive a credit or receive replacements of other models of laser delivery probes. |  | 121 | 121 | Nationwide Distribution in the states of KS, MO, NC, OH, and VA. | Terminated | March 10, 2017 | 350 | 121
Z-1689-2016 | 73194 | No | LFZ | Cytomegalovirus serological reagents. | Microbiology | Microbiology | 510(k) | Diamedix Is-CMV IgG Test Kit |  |  | March 26, 2015 | 2015 | Diamedix Corporation | Product contained an incorrect substrate. | Other | On March 26, 2015 Erba Diagnostics sent a communication to their clients informing them of the recall and sending t hem a replacement sustrate HRP to replaced the wrong one sent on the kit. |  | 150 | 150 kits | AZ, CA, FL, GA, KY, MA, MI, MN, MO, NC, NE, NY, OH, OK, OR, TN, TX, VA, WI and Puerto Rico. | Open, Classified |  |  | 150
Z-1688-2016 | 73883 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | LenSx Laser System. Indicated for use in patients undergoing cataract surgery and refractive surgery. |  |  | April 14, 2016 | 2016 | Alcon Research, Ltd. | Alcon is conducting a voluntary recall of all LenSx Lasers with Corneal Flap Capability after receiving reports concerning incomplete corneal flap creation, which manifests as an un-dissected area(s) within the side cut. | Software design (manufacturing process) | Recall notification letters were sent to consignees on April 14, 2016. The letter details precautions to be taken by the surgeon and requests the consignee sign and return an acknowledgement form. |  | 81 | 81 units | United States, Albania, Australia, Brazil, Chile, Costa Rica, Czech Republic, France, Germany, Greece, Hungary, India, Italy, Romania, Slovakia, Spain, Taiwan, Thailand, Turkey, UAE, Ukraine | Open, Classified |  |  | 81
Z-1690-2016 | 73691 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 1) 24 G x 0.75 in. BD Angiocath" Autoguard" shielded IV |  |  | March 18, 2016 | 2016 | Becton Dickinson &amp; Company | The device may have a defect in the catheter. In some instances this defect could result in catheterseparation or breakage. | Device Design | BD isssued an urgent voluntary product recall lettter/recall response card on 3/18/2016. |  | 1305150 | 1,305,150 units | Nationwide | Terminated | July 19, 2017 | 488 | 1305150
Z-1693-2016 | 73777 | No | NGY | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | 23G and 25G Constellation Totalplus Paks provides access into the posterior segment of the eye for surgical instruments within the frame of 23G and 25G Pars-Plana-Vitrectomy (PPV) surgeries. |  |  | August 11, 2015 | 2015 | Alcon Research, Ltd. | The product has potential to leak beyond their design specification. | Process control | The affected consignees were contacted via letter on 8/11/15 to communicate the potential for leakage. |  | 225000 | 225,000 units | United States, Argentina, Armenia, Australia, Azerbaijan, Bangladesh, Brbados, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Curacao, Cypress, Denmark, Ecuador, Finland, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Lsreal, Italy, Japan, Korea, Lithuania, Malta, Mexico, Monaco, Netherlands, New Zealand, Pakistam, Philippines, Poland, Prtougal, Puerto Rico, Romania, Russia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay | Terminated | April 20, 2017 | 618 | 339149
Z-1668-2016 | 73738 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | ANSPACH, Rx only, Otologic Curved Micro (OCCM) Diamond Ball Burr, 2.0 MM, Part # OCM7-2SD |  |  | March 28, 2016 | 2016 | The Anspach Effort, Inc. | Certain lots of the Otologic Curved Micro (OCM) Diamond Ball Burrs may be missing the diamond coating and the ball may be the incorrect size. | Process control | Consignees were sent an Urgent Notice of this Medical Device Recall on 3/28/2016 for their product Otologic Curved Micro (OCM) Diamond Ball Burr, with the following actions required: We have record that your facility has the product(s) subject to this recall (removal). DePuy Synthes asks that you review your inventory and immediately remove the affected lots from stock.Please take the following actions: If you DO HAVE any of the identified devices, please take the following steps: Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided. Please contact Product Support at (800) 327-6887, option 1, to arrange for return of any unused product. A credit will be issued for returned product. Send a copy of the completed Verification Section (page 3) to Product Support by fax to (561) 627-2682. Send a copy of the completed Verification Section (page 3) to the Complaint Handling Unit byscan/email to RA-DPYUS-chu.pb@ITS.JNJ.com. If you DO NOT HAVE the identified product, please take the following steps: Complete the Verification Section by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledgesyour receipt of medical device recall (removal) information.Âż Return the documents to The Anspach Effort by:Âż Fax: (561) 627-2682Âż Scan/email: RA-DPYUS-chu.pb@ITS.JNJ.comÂż Giving them to your DePuy Synthes Sales ConsultantÂż Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual of the notification. |  | 27 | 27 devices | US Distribution in the states of KY and TX. | Open, Classified |  |  | 27
Z-1636-2016 | 72350 | No | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | ABG-HM-1 Hummi Micro Draw Blood Transfer Device |  |  | October 01, 2015 | 2015 | Hummingbird Med | Hummingbird Med Devices, Inc. is recalling ABG-HM-1 Hummi Micro Draw Blood Transfer because the connection between the "Hummi" Y connector and the yellow cannula hub may not be sufficiently secure enough and could separate while in use. | Component design/selection | On 10/01/15 the firm sent out customer notification letters.The letter states for the following actions to be taken:1. Customers using the Hummi Micro Draw device are recommended not to use the identified lot.2. A local distributor/sales representative will be contacting customers to arrange a return of the identified lot for credit and arrangement of replacement.3. Replacement product should be available within 10 days of 10/30/15.4. During the period of product unavailability customers might want to choose to temporarily to return their previous method for blood drawing.The firm requests a Reply Form be completed and returned via fax at 949-583-2775 or email to JorgeHaider@hummingbirdmed.com.On 11/19/2015 the firm sent out a second notification letter involving additional recall lots. The firm provided the same instructions to customers as their previous notification; however stated replacement product should be available by mid December 2015. |  | 11500 | 11500 units | Distributed in the states of CA, KY, MD, GA, and IL. | Terminated | January 26, 2017 | 483 | 11500
Z-1700-2016 | 73441 | No | DHA | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | 1) VITROSÂż Immunodiagnostic Products Total Beta-hCG II Reagent Pack |  |  | February 18, 2016 | 2016 | Ortho-Clinical Diagnostics | When testing was conducted using certain lots, VITROSÂż Systems generated results within the measuring range for samples known to not contain measurable hCG. The system reported results up to approximately 7.40 mIU/mL (IU/L) for patient samples that should have been less than the measuring range of the assay (&lt;2.39 mIU/mL (IU/L)). | Under Investigation by firm | On February 18, 2016, Ortho Clinical Diagnostics distributed Urgent Product Correction Notification letters and Confirmation of Receipt Forms to their customers (Ref. CL2016-043) via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail. Customers are advised to immediately discontinue using and discard all remaining inventory of the listed lots. Ortho Clinical Diagnostics will replace all remaining inventory or provide account credits as indicated on the confirmation of receipt form. Customers are advised to review previously reported results using the affected lots and to discuss these findings with your Laboratory Medical Director to determine the appropriate course of action. Post the notification by your VITROS System or with your user documentation. Complete and return the Confirmation of Receipt form to Joe Falvo via email at confirmationadmin@its.jnj.com or by fax to 1-888-557-3759 or to 1-585-453-4110. Clients with questions can contact Customer Technical Services at 1-800-421-3311 at any time. |  | 14780 | 14,780 pieces | Nationwide Distribution. | Open, Classified |  |  | 14780
Z-1702-2016 | 73815 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | AngioVac Circuit Pack (Tandem Packs), REF/Catalog No. 25-187, UPN H965251870, STERILE, Rx Only, Single Use Only -- The AngioVac Circuit and AngioVac Circuit Pack are designed to be used with the AngioVac Cannula. |  |  | March 21, 2016 | 2016 | Angiodynamics Inc. (Navilyst Medical Inc.) | Maquet Cardiopulmonary, the manufacturer of the RotaFlow Centrifugal Pump, determined that the products may possess levels of endotoxin that exceed regulatory limits. AngioDynamics confirmed that the affected RotaFlow Centrifugal Pumps were included in packaged AngioVac Circuit Packs. | Environmental control | On 3/21/2016, AngioDynamics sent recall notification packages (dated 3/21/2016) to the consignees via Federal Express. Consignees are instructed to follow the instructions provided on the Maquet Cardiopulmonary recall letter (dated 2/23/2016) and return the Reply Verification Tracking Form, provided in the recall notification. |  | 377 | Domestic: 377 kits | Nationwide Distribution. | Open, Classified |  |  | 716
Z-1703-2016 | 74171 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | WHIN Infusion Set - 90 Huber Needle and Wing (22 ga/1in). Intravenous delivery through intravascular ports, 20 units per case, Catalog Number 471734. |  |  | April 26, 2016 | 2016 | B. Braun Medical, Inc. | Potential for the tip of the Huber needle on the WHIN Infusion Set to protrude from the needle guard. | Component design/selection | The firm, B. Braun Medical Inc., sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 4/26/2016 to customers. The letter described the product, problem and actions to be taken. The customers were instructed to evaluate their current inventory for the affected lots; fill out the Product Removal Acknowledgement form and record the total number of units found, and return the form to B. Braun Medical Inc. Quality Assurance department by faxing the form to 610-849-1197 or email to PA_Quality Assurance.BBMUS_Service@bbraun.com within two weeks of receipt. A BBMI Customer Support Representative will contact the customer to provide instructions for handling the affected product and arrange for return to BBMI.Should you have any questions or concerns regarding the attached information, please contact our Customer Support Department at (800) 227-2862. |  | 5980 | 5,980 units [Lot # 0061482060 (3,340 units); Lot # 0061486821 (2,640 units)] | Worldwide distribution: US Distribution to states of: IN, NV, MA, IL, AR, TX, OH, CA, GA, FL, MS, NJ, AL, PA, TN, KY, VA, SC, and country of: Canada. | Open, Classified |  |  | 5980
Z-1699-2016 | 73678 | [6 empty] | Orapharma (Onpharma) Onset Mixing Pen Rx only, Product/Part Number 1054-02. |  |  | January 29, 2016 | 2016 | Valeant Pharmacueticals International | It has been determined that the Onset Mixing Pen has an optimal use period of 18 months from start of use. However, the current labeling and IFU does not stipulate this time period. | Labeling False and Misleading | Valeant sent a letter entitled "Field Correction to Orapharma Onset Mixing Pen," dated 1/29/2016, to customers via UPS. The letter identified the affected device and the issue involved. The firm recommends that users stop using the device if it had been in use for more than 18 months. Customers were to refer to the list of the affected lots to determine if their device needed to be replaced. If the device was one of the ones identified, users were to contact Valeant Pharmaceuticals to arrange for a replacement at no charge. Questions should be directed to customer service 800-321-4576. |  | 7716 | 7,716 | Nationwide Distribution | Open, Classified |  |  | 7716
Z-1704-2016 | 73823 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Liquid Mouse Monoclonal Antibody Calretinin, NCL-L-CALRET-566. |  |  | March 22, 2016 | 2016 | Leica Microsystems, Inc. | Novocastra Liquid Mouse Monoclonal Antibody Calretinin when used in combination with specific Lots of antibody diluent will give reduced staining. | Nonconforming Material/Component | A Medical Device Recall Notification letter was sent via e-mail on April 5, 2015 to the consignees describing the nature of the recall and action that needed to be taken. The firm requested that the consignees take the following actions: Only use Novocastra Liquid Mouse Monoclonal Antibody Calretinin with the lot numbers of diluent specified.There is no need for retrospective review of tissue staining if appropriate positive controls were properly used and interpreted during the immuno / in situ staining.Complete and return the attached acknowledgment.The firm requests that consignees contact their local Leica representative if they have any questions or concerns, or alternatively please contact the firm at the following e-mail address; LMGRA@leica-microsystems.com |  | 290 | 290 units | Worldwide Distribution -- U.S., to the following states: AK, CA, FL, GA, MN, MT, NJ, OH, TX and WA; and, the country of Canada. | Open, Classified |  |  | 290
Z-1709-2016 | 73945 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 3.2 mm Three-fluted drill bit, sterile, Catalog ID # 03.010.060S |  |  | April 18, 2016 | 2016 | Synthes (USA) Products LLC | The sterile packaging may be compromised. A protective cap is to be secured on the tip of the drill bit. If the protective cap is not secured, the drill bit may puncture the sterile barrier. | Device Design | DePuy Synthes sent an Urgent Notice Medical Device Letter dated April 19, 2016 to customers. The identified the affected product, problem and actions to be taken. The letter instructed customers to examine their inventory to determine if they currently have the affected product. Customers were asked to complete and return the Verification Section of the recall notification letter. For questions call 610-719-5450. |  | 7363 | 7363 units | US Nationwide distribution in the states of: AE, AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX,UT, VA, VT, WA, WI | Open, Classified |  |  | 8637
Z-1708-2016 | 73983 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Reprocessed Daig Supreme Fixed Curve Diagnostic Electrophysiology Catheters; Model #: 401450, 401877. Indicated for temporary intracardiac sensing, recording, stimulation, and electrophysiology mapping of cardiac structures. |  |  | April 28, 2016 | 2016 | Stryker Sustainability Solutions | Mislabeled for size | Under Investigation by firm | The firm sent out customer notification letters on 04/28/16.Customers are instructed to discontinue the use of affected lots of recalled product. Other EP Catheters reprocessed by Stryker Sustainability Solutions should be considered as alternatives. A Recall Effectiveness Check Form including the quantity of affected devices shipped to facility is enclosed.Please complete the Recall Effectiveness Check Form and indicate if any devices from the affected lots remain in your inventory. This form needs to be completed even if no affected product is found.Please return the completed and signed Recall Effectiveness Check Form to your local Stryker Sustainability Sales Representative or email to SSSPFA@stryker.comor mail to:Stryker Sustainability Solutions1810 West Drake DriveTempe, AZ, 85283Attn: Jodie RueckertIf the form indicates that affected devices remain in inventory, a prepaid shipping label will be issued for the return of the product. Customers will receive a credit for all affected devices returned.Adverse reactions or quality problems experienced with the use of this productmay be reported to:- Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433X5555- http://www.stryker.com/productexperience/- The FDAs MedWatch Adverse Event Reporting program either online,by regular mail, or by fax.On 05/12/16 the firm sent out a clarification letter to emphasize that use of affected lots of product should be discontinued. In addition, SSS would like to note that, use of devices larger than expected may result in excess force applied to the blood vessel. The risk to the patient can vary from a prolongation in procedure time, vessel rupture, organ damage or other unforeseen consequences. Some of these injuries may require surgery or other procedures to remedy. Any questions contact the Stryker Sustainability Solutions Sales Representative or Regulatory Affairs. |  | 18 | 18 devices | Distributed in the state of NY and the country of Canada. | Open, Classified |  |  | 18
Z-1707-2016 | 73968 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery IGS 740. Indicated for use in generating fluoroscopic and rotational images of human anatomy. |  |  | April 15, 2016 | 2016 | GE Medical Systems, LLC | Potential non-recoverable loss of image acquisition. The affected Discovery systems may experience multiple X-Ray abort errors before or during a real-time fluoroscopic Interventional procedure. | Device Design | Consignees were sent on 4/15/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref#12246 dated April 14, 2016. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology and Radiologists/Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 25 | 25 | Worldwide Distribution-US (nationwide) including the states of AL, CA, CO, CT, IL, IN, MD, MN, MS, MO, NE,NV,NH, NM, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI, and the countries of ALGERIA, BELGIUM CANADA, CHINA, FRANCE, GERMANY, HUNGARY, IRELAND, ITALY, JAPAN, MEXICO, NETHERLANDS, PERU, SAUDI ARABIA, SINGAPORE, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 79
Z-1705-2016 | 74087 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iGUIDE System, for patient positioning, with assistance of a 30 Tracking System in a radiotherapy environment. |  |  | May 12, 2016 | 2016 | Elekta, Inc. | If by mistake the initial pre-treatment imaging was performed BEFORE the HexaPOD was moved to the DRIVE (*START) position, iGUIDE offers the possibility to branch into a specific catch-up workflow, which was implemented to use the initial scan data. | Software design | Elekta sent an Important Field Safety Notice that was distributed to all affected customers on May 12, 2016. The notice informs end users about the problem and the potential clinical impact. It also gives a recommendation how the risk can be mitigated. The notice also advises the customer that a patch will be released that will introduce a patch that avoids an inaccurate position and thus eliminates an unsafe situation.Corrective Action #2: Permanent Solution  Software UpgradeA software patch will be released that will correct the behavior of the iGUIDE software. The target release date for this patch is January 2017.Service teams will have 6 months from the date of release to correct all affected devices in the field. For further questions, call (770) 300-9725. |  | 20 | 20 | IL, LA, Austria, Australia, Botswana, Germany, Denmark, France, India, Japan | Open, Classified |  |  | 20
Z-1720-2016 | 73949 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Accriva- APTT Cuvette for use on the Hemochron Jr. test system, catalog #J103, lot B6JCA012. |  |  | April 22, 2016 | 2016 | Accriva Diagnostics Inc., dba ITC, dba Accumetrics | Clot activating reagent mis-positioned or visibly absent in the cuvette's test channel causing APTT test to either have an error code or an erroneously high result. | Process design | Accriva sent an " Urgent Field Safety Notice" letter dated April 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were requested to take the following actions:1. Forward communication to all those within organization who need to be aware of this matter. In addition, forward notification to all other organizations where affected devices may have been transferred to.2. Check inventory to determine if you have any APTT Test Cuvettes, LotNumber: B6JCA0123. If have inventory of APTT Test Cuvettes, with the subject Lot Number destroy the product and submit the attached Destruction Form along with your request for credit.4. If have inventory DO NOT SHIP TO CUSTOMERS any kits labeled with APTT Test Cuvettes and the subject Lot Number.5. If have shipped this product to customers please inform them immediately to STOP using the product. Provide the customers with a copy of this notification, request for the destruction of any unused product and the return of the attached forms.6. Follow the instructions on the attached Field Corrective Action Form, complete the form and return it as instructed on the form.7. If have questions, please contact Accriva Diagnostics directly contact information is provided below.Accriva Diagnostic, Attn: Regulatory AffairsUSA Tel: 858.263.2347Email: Fieldaction16-007@accriva.com |  | 135 | 135 boxes | U.S. distribution to the following; state: VA. Foreign distribution to the following; China, Italy, Germany. | Open, Classified |  |  | 135
Z-1813-2016 | 73984 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Volcano s5, s5i, CORE and CORE Mobile systems with software version 3.4 and v3.4 software kits. |  |  | April 25, 2016 | 2016 | Volcano Corporation | Volcano has become aware of an incompatibility issue between Impacted Systems and hospital network scans. In specific circumstances, an impacted System will encounter unexpected data from the hospital network and be forced to reboot | Software in the Use Environment | The firm, Philips, sent an "URGENT MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter dated 4/25/2016 to its customers beginning April 29, 2016 which include a technical service bulletin to explain how to assure the incompatibility issue does not affect system performance. The customers were instructed to do the following:-If you have an Impacted System but it is not connected to a network, this issue does not affect your system. Additionally, if network scans are not performed, this issue does not affect your system.-If you have an Impacted System that is connected to a network and staff performs network scans while the Impacted System is in operation, this issue does affect you and the circumstances exist for a manual reboot to occur during an active procedure. To avoid this, the following corrections must promptly be taken: 1. Disconnect your Impacted Systems from the facilities network. If you need to reconnect it while not being used in a procedure, make sure it is disconnected again before starting any procedure; OR 2. Do not permit your staff to perform network scans while the Impacted Systems are in operation.-Please complete, sign, and return the attached form indicating that you received this Field Corrective Action notification via Fax to Volcano Customer Service at (916) 638-8812 or email to CS@volcanocorp.com.If you have any questions, please call (800) 228-4728, option 5. |  | 5875 | 5875 (2685 in US; 3190 ROW) | Worldwide Distribution: US (nationwide) and countries of: Japan, South Africa, Germany, Sweden, Belgium, France, Italy, Spain, India, Russian Federation, Hungary, Turkey, Poland, Portugal, Costa Rica, Finland, Great Britain, Croatia, India, Chile, Thailand, Mauritius, Iraq, Denmark, Norway, Ireland, Netherlands, Bulgaria, Czech Republic, Luxembourg, UAE, Austria, Poland, Pakistan, Sri Lanka, Kazakhstan, Slovakia, Saudi Arabia, Unite Arab Emirates, Egypt., Jordan, Kuwait, Lebanon, Bulgaria, Romania, Latvia, Lithuania, Bosnia and Herzegovina, Greece, Malta, Cyprus, Romania, Tunisia, Kenya, Algeria, Israel, Jordan, Hong Kong, Republic of Korea, Malaysia, New Zealand, Venezuela, Brazil, Philippines, Australia, Taiwan, Peru, Singapore, Panama, Mexico, Argentina, Costa Rica, Dominican Republic, China, and Paraguay, | Open, Classified |  |  | 5875
Z-1711-2016 | 73870 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SMR glenosphere impactor/extractor; Used to implant the glenosphere in the SMR Reverse Shoulder replacement. |  |  | April 08, 2016 | 2016 | Limacorporate S.p.A | Complaints of intra-operative breakage of the glenosphere impactors/extractors. | Component design/selection | The affected firms received written notification to all distributors. The letter instructed the consignee to identify and remove all affected product in stock and return parts to Lima USA. |  | 314 | 314 units | Distributed US (nationwide) and in Australia. | Open, Classified |  |  | 314
Z-1814-2016 | 74025 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | PerFuse Decompression Instrument PN: 800-0541, Lot:100650. Designed to access the femoral head for decompression. |  |  | April 11, 2016 | 2016 | Biomet, Inc. | The Trocar and Plunger Assemblies are missing from the instrument, which could result in a delay in surgery greater than 30 minutes. | Process control | Zimmer Biomet has initiated a voluntary recall of the PerFuse Decompression Instrument, following an investigation which identified that the Trocar and Plunger Assemblies are missing from the instrument. The firm contacted customers via telephone and e-mail on March 28, 2016; and official recall notices were mailed via FedEx on 4/11/2016. Customers were instructed to quarantine recalled items, Complete the Certificate of Acknowledgment and email a copy toCPWARFieldAction@zimmerbiomet.com; Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA: www.fda.gov/Safety/MedWatch/HowToReport/default.htm (form available to fax or mail), or Call (800)FDA-1088 .Questions related to this recall should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 22 | 22 | Distributed in the US states of CA, GA, MN, and OH, and in the Netherlands. | Terminated | May 30, 2017 | 414 | 22
Z-1812-2016 | 73959 | No | CAD | Anesthetic vaporizer. | Anesthesiology | Anesthesiology | 510(k) | Draeger D-Vapor/ D-Vapor 3000 Vaporizers with Baxter Desflurane Agent and Bottles. Catalog # M35500 (D-Vapor) and M36700 (D-Vapor 3000). Anesthesia Vaporizer |  |  | April 21, 2016 | 2016 | Draeger Medical, Inc. | Operators noticed unwanted release of evaporating anesthetic agent at the end of the filling procedure. Eye contact with volatile anesthetics can result in serious irritation. Immediate eye flushing and consultation of a physician is recommended. Updated Material Safety Data Sheets (MSDS) for Desflurane recommend wearing safety glasses when handling these substances. | Process change control | Drager mailed a "Medical Device Safety Alert" dated April 21, 2016, to customers informing them of the problem and provided instructions for the "correct filling procedure" and recommendations on how to prevent evaporation and spraying of Desflurane after bottle release to resolve the "High Friction of gas bottle valve" problem. |  | 16078 | 16078 units distributed within the United States | US nationwide distribution, including Puerto Rico and Guam. | Open, Classified |  |  | 16078
Z-1712-2016 | 73575 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Therapy Treatment Planning System Stand-alone Software 3.0, 3.5, 4.0, 4.5 and 4.7., designed for treatment planning and analysis of radiation therapy. |  |  | February 10, 2016 | 2016 | RAYSEARCH LABORATORIES AB | A software issue with editing tools that use the left mouse button held down, for drawing in and interacting with the patient views and beams eye views. The views can become unsynchronized with the stored data if simultaneously right clicking, pressing Ctrl-S or Ctrl-Z while the left mouse button is held down. This bug does not affect dose computations, which are based on the stored system state. | Other | On February 10, 2016, RaySearch Laboratories distributed Field Safety, Medical Device Correction notices and Reply Forms to their customers via email. To determine if the software version you are using is affected, open the About RayStation dialog in the RayStation application and check if the build number reported. This bug does not affect dose computations. Dose is always computed based on the stored system state. Actions to be taken by the user include the following: Always release the left mouse button before using other keys or mouse buttons, If the warning message "Temporary items should not be visible on approval" is displayed or if there is some suspicion that the bug may have triggered in structure definition, either review the structures in the Plan approval dialog, or close and reload the patient case and review the structures again. All customers should educate planning staff and all users about this workaround. All customers are advised to inspect product and identify all installed units with the referenced software version number, then confirm that you have read and understood the notice by completing and returning the reply form. The completed reply form should be returned to Freddie Cardel, freddie.cardel@raysearchlabs.com, fax 888-501-7195. Customers with questions can call 1-877-778-3849, available east coast business hours all US working days. |  | 1264 | 1,264 Total units (552 units domestically &amp; 711 units internationally) | **Domestic** AR, AZ, CA, CT, CO, DE, FL, GA, IL, LA, MI, MD, MS, MA, NY, NV, NJ, OH, OR, PA, NC, RI, TN, WA &amp; Washington D.C; **Internationally** Algeria, Australia, Austria, Belgium, Canada, China, France, French Polynesia, Germany, Iran, Italy, Japan, Netherlands, New Zealand, Norway, Republic of Korea, Spain, Sweden, Switzerland, Thailand &amp; United Kingdom | Terminated | April 24, 2017 | 439 | 1264
Z-1829-2016 | 74031 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Streamline OCT Occipito-Cervico-Thoracic System, Crosslink Counter Torque Sleeve. Spinal fixation device instrument. |  |  | April 22, 2016 | 2016 | RTI Surgical, Inc. (dba Pioneer Surgical Technology, Inc.) | Potential for a product defect where the cannula may be obstructed. This obstruction could prevent a driver from going through, which increases the potential for the driver to slip, resulting in possible patient injury. | Process design | RTI sent an Urgent Voluntary Product Notification dated April 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.RTI Surgical is conducting a voluntary recall of the Streamline OCT Occipito-Cervico-Thoracic System, Crosslink Counter Torque Sleeve due to the potential for a product defect where the cannula may be obstructed. This obstruction could prevent a driver from going through, which increases the potential for the driver to slip resulting in possible patient injury.For further questions, please call (906) 226-4489. |  | 17 | 17 devices | US nationwide distribution in CA, FL, IN, MI, MO, NC, NY, and TX. | Open, Classified |  |  | 17
Z-1830-2016 | 73703 | No | OWD |  |  | Molecular Genetics | PMA | cobasÂż EGFR Mutation Test, v2 and cobasÂż cfDNA Sample |  |  | March 15, 2016 | 2016 | Roche Molecular Systems, Inc. | An error was found within the Hungarian translations of the cobasÂż EGFR Mutation Test v2 Instructions for Use (M/N07340761001-01HU, Doc Rev. 1.0, Dated 08/2015) and the cobasÂż cfDNA Sample Preparation Kit Instructions for Use(M/N 07573758001-01HU, Doc. Rev. 1.0, Dated 05/2015). | Packaging | Roche Molecular Diagnostics Inc. issued an Urgent Field Safety Notice dated March 15, 2016 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Actions to be taken by the customer:" If you have utilized 2 ÂżL of DNA PBB during testing of plasma specimens, the samples must be retested using the correct volume (2mL DNA PBB).Communication of this Field Safety Corrective ActionThis notice must be passed on to all those who need to be aware within your organization or to any organization/individual where the potentially affected devices have been distributed/supplied.Please maintain awareness of this notice and resulting action for an appropriate period to ensure the effectiveness of the corrective action.The undersigned confirms that this notice has been notified to the appropriate Regulatory Agency.For questions regarding this recall call 908-253-7200. |  | 8 | 8 kits | Hungary | Open, Classified |  |  | 8
Z-1816-2016 | 74038 | No | GEA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Various trauma and sports medicine instruments and implants. |  |  | April 01, 2016 | 2016 | Biomet, Inc. | Insufficient seal strength on sterile barrier packaging.The seal could be jeopardized causing the sterile barrier to be compromised and the product could become nonsterile.If the product becomes non-sterile and is used or implanted, it may lead to an infection. | Equipment maintenance | Zimmer Biomet sent an Urgent Medical Device Recall Notice dated April 1, 2016, to all affected customers. Customers were instructed to locate and return the affected product, and implanting surgeons were instructed to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  |  |  | Worldwide Distribution - US (Nationwide) and globally to Argentina, Australia, Brazil, Canada, Chile, Colombia, Costa Rica, Ecaudor, The Netherlands, Japan, Korea, Malaysia, Mexico, Peru. | Terminated | May 23, 2017 | 417
Z-1784-2016 | 73734 | No | HBA | Neurosurgical paddie. | Neurology | Neurology | 510(k) | American Surgical Delicot 4 mm x 50 mm |  |  | March 31, 2016 | 2016 | American Surgical Company | Product packaging defective compromising sterility | Packaging process control | American Surgical issued an Urgent Recall and Field Safety Notice letterdated March 31, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were provided instruction to address affected devices in their inventory. They should address this by either selecting Action A or B: Action A: Users should perform a check of the device before use and, as stated in the IFU, patties should not be used if in an open or damaged pouch. For the affected lots, special attention should be given to lower portion of the pouch where the Tyvek and clear film are sealed.Action B: Place the affected lots in quarantine and contact American Surgical Company for an RMA and replacement. All affected devices are eligible for replacement by ASC.Complete and return confirmation of actions described in Field Safety Notice.Contact:Tel: 1-800-343-0060 Option 2, Fax: 781-595-5460Email: CustomerService@AmericanSurgical.com |  | 99 | 99 units | Worldwide Distribution - US Nationwide and the countries of Australia, Belgium, Canada, France, Germany, Greece, Hong Kong, Hong Kong, Israel, Italy, Italy, Japan, Netherlands, New Zealand, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, and United Kingdom. | Open, Classified |  |  | 7701
Z-1831-2016 | 73923 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CADD Administration Sets with Flow Stop, Product Reorder No: 21-7321-01, 21-7321-24, 21-7322-01, 21-7322-24, 21-7323-24, 21-7324-01, 21-7324-24, 21-7333-24, 21-7336-01, 21-7336-24, 21-7339-01, 21-7339-24, 21-7359-01, 21-7359-24, 21-7383-01, 21-7383-24, 21-7390-01, 21-7390-24, 21-7391-01, 21-7391-24, 21-7394-01, 21-7394-24, 21-7395-24. |  |  | April 22, 2016 | 2016 | Smiths Medical ASD, Inc. | Smiths Medical became aware that under delivery of medication may occur on CADD Administration Sets with Flow Stop. | Nonconforming Material/Component | Smiths Medical sent an "Urgent Medical Device Field Correction Action" letter dated April 20, 2016. The letter was addressed to "Clinicians who oversee the use of the CADDÂż Administration Sets with Flow Stop and Distributors thereof". The letter described the Affected devices, Reason for Field Correction Action, Risk to Health, and Instructions to Customers (Distributors and Clinicians). Requested consignees to complete and return the "Field Corrective Action Confirmation Form". For questions regarding this notification, contact Smiths Medical Customer Service Department at 1-866-831-8399. |  | 13696830 | 13,696,830 units | US: ALL STATES IN CONTINENTAL USA INCLUDING DC AND PR EXCEPT DE.OUS:UNITED ARAB EMIRATES, AUSTRIA BELGIUM, CANADA, CHINA, GERMANY, DENMARK, SPAIN, FINLAND, FRANCE, GREAT BRITAIN, HONG KONG, IRELAND, INDIA, ITALY, JAPAN, LUXEMBOURG, MAURITIUS, NEW CALEDONIA, NEW ZEALAND, PANAMA, PORTUGAL, SAUDI ARABIA, SWEDEN, SINGAPORE, TAIWAN, VIET NAAM, SOUTH AFRICA. | Open, Classified |  |  | 13696830
Z-1835-2016 | 73377 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Video Duodenoscope |  |  | February 22, 2016 | 2016 | Pentax Medical Company | Pentax Medical learned of reports of carbapenem-resistant Enterobacteriaceae (CRE) infection at a medical facility. | Other | Pentax Inc. issued an "Urgent Field Correction" Letter dated February 19, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter provides information regarding the updated Reprocessing instruction for use along with some main changes in the Operation IFU. Instructions are also provided regarding prior versions of the Reprocessing and Operation Instructions for Use. A web posting went live on 2/19/2016.For further questions, please call:Tel:800-431-5880 ext. 2064 (8:30AM- 5:00 PM, Monday-Friday, EST) |  | 423 | 423 units | Nationwide | Open, Classified |  |  | 423
Z-1834-2016 | 73661 | No | GET | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Pneumatic Impactor, General Surgery, With Microaire Hose Connector |  |  | March 11, 2016 | 2016 | MicroAire Surgical Instruments, LLC | MicroAire discovered that 20 suspect Impactor instruments ( MicroAire Pneumatic Impactor) were shipped to end users and were missing the lock and unlock symbol markings on them. | Other | MicroAire sent an Urgent Medical Device Recall letter dated March 28, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.We request that you please respond to this letter promptly to determine how many of these units you have and will be returning to MicroAire. Upon return of the instrument to MicroAire, MicroAire will provide you with replacements.Please complete and send back to MicroAire page two of this letter, indicating that you do have this instrument at your facility and whether you will be returning the instrument for replacement.You can email the form to MicroAire Customer Service at notification@microaire.com or fax the form to 1-800-648-4309.You can contact MicroAire Customer Service at 1-800-722-0822 in the United States or 1-434-975-8000 outside the United States, between the hours of 8:30 AM and 5:30 PM (ET) Monday through Friday if you have any questions. |  | 20 | 20 units | United States | Open, Classified |  |  | 20
Z-1836-2016 | 74188 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SEAL,IS4000 PORTS,5MM-8MM,BOX of 10; Model numbers: 470361-03, 470361-04, 470361-05, 470361-06; Part number: 470361. The da Vinci Xi Surgical System cannula sterile single use disposable cap for each cannula . When attached to the cannula, the seal maintains insufflation when the cannula is empty and when the instrument is inserted in the patient during a surgical procedure. |  |  | May 09, 2016 | 2016 | Intuitive Surgical, Inc. | Potential for particulates in insufflation stopcocks on the da Vinci Xi Surgical System cannula seals. The particulates could be introduced into the cannulalumen when insufflation is connected during a surgical procedure. | Manufacturing material removal | Upon discovery of the issue, Intuitive sent Urgent Product Safety Notices to all customers by e-mail to initially alert them of the situation. This will be followed by a formal letter that is scheduled to arrive on May 23, 2016. Both notification request that users Do NOT use the da Vinci Xi 5 mm8 mm Universal Cannula Seal or the da Vinci Xi 12 mm &amp; Stapler Universal Cannula Seal for insufflation, or open the stopcock during intracardiac procedures. |  | 46 | 46, 191 boxes - total both sizes | Worldwide distribution. US, Australia, Austria, Belgium, Chile, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Monaco, Netherlands, Norway, Puerto Rico, Portugal, Qatar, Romania, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Open, Classified |  |  | 46
Z-1842-2016 | 73123 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hip Preservation System Signature Series PreBent Burs under the following labels: 1) PreBent Spherical Bur, 5.5mm x 19cm Catalog Number HPS-HB11, 2) PreBent Polishing Bur, 5.5mm x 19cm Catalog Number HPS-HB12 and 3) PreBent Oval Bur, 6.0mm x 19cm Catalog Number HPS-HB13 |  |  | December 21, 2015 | 2015 | ConMed Corporation | Metal shavings released from burs during use are due to contact between the bur and the hood of the PreBent Bur assembly. During the manufacturing process an assembly error occurred which caused inadvertent mixing of two sizes of bearings. | Mixed-up of materials/components | ConMed sent an Urgent Medical Device Recall Notification letter dated December 17, 2015, to all affected customers. Customers are urged to stop use and return these devices immediately. Customers who have further distributed these items are urged to contact all of those organizations immediately and provide a copy of the recall notification letter. It is imperative that all end users of these devices receive the recall notice and respond immediately. Customers with questions can call 315-624-3237 or email largohpshb@conmed.com. The devices should be returned to ConMed Corporation NCR-47057 11155 Concept Blvd. Largo, FL 33773. For questions regarding this recall call 315- 624-3533. |  | 991 | 991 pieces | Worldwide Distribution - Nationwide to CA, CO, FL, IN, MD, MN, NC, NY, OH, TX, UT &amp; VA; International: Canada, Wales | Open, Classified |  |  | 991
Z-1848-2016 | 73427 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Kerlix" X-Ray Detectable Laparotomy Sponges, Tray Pack, 16" x 17" (40.6 cm x 43.2 cm) |  |  | March 03, 2016 | 2016 | Covidien LLC | Compromised sterility due to breach of barrier | Packaging process control | Covidien initiated recall by letter on March 4, 2016. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Any questions or concerns, contact your Medtronic representative or Customer Service at (800) 882-5878 |  | 114720 | 114720 | Distributed US (nationwide) and the countries of Canada, Austria, Belgium, Columbia, Croatia, France, Germany, Guadeloupe, Ireland, Italy, La Reunion, New Zealand, Norway, Poland, Portugal, Saudi Arabia, South Africa, Spain,Sweden and United Kingdom. | Terminated | June 15, 2017 | 469 | 8748710
Z-1860-2016 | 73469 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Endoscope; ED-530XT Operation Manual |  |  | December 23, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | This correction is in response to publicized reports of multi-drug resistant bacteria on endoscopes used for Endoscopic Retrograde Cholangiopancreatogram (ERCP) procedures. | Other | Fujifilm issued a urgent field correction/Tracking/Verification form on 12/23/2015. |  | 1473 | 1473 units | Nationwide Distribution. | Open, Classified |  |  | 3031
Z-1844-2016 | 74128 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | EZ Slide Battery/line powered hydraulic table with removable leg section and 21 top slide |  |  | April 06, 2016 | 2016 | Skytron, Div. The KMW Group, Inc | The potential exists within the identified tables that incorrectly manufactured parts have been used in assembly. If a table is found to have these incorrectly manufactured parts, it is possible for the tables elevation column slide cap to become detached from the elevation column and could result in the loss of control of the table top and injury to the patient or staff. | Reprocessing Controls | Skytron sent an Urgent Medical Device Recall letter dated A,pril 4, 2016, to all affected customers. On April 6, 2016, URGENT: Medical Device Recall notifications were sent to the affected hospitals and distributors via certified mail. The notification instructed the hospital to identify and remove the affected product from use. The notification instructs the distributor to contact the hospital to schedule replacement. The recall notification included a description of the reason for the recall, affected product, and consignee responsibilities. The notification did not include instructions for responding to the formal recall notification, however the company did request acknowledgement via follow-up email.Please call Technical Services, at 616-656-1187 or 616-656-6403 with any questions. For after-hours support, please call 1-800-SKYTRON. |  | 458 | 458 | Nationwide Distribution to AL, AZ, CO, IA, WI, PA, TN, MA, TX, MI, FL, MS, LA, and CA. | Terminated | July 20, 2017 | 470 | 458
Z-1843-2016 | 73664 | No | LDH |  |  | General Hospital | 510(k) | AllergiEnd ST-9 Multiple Skin Test Applicator (Item 1000) and AllergiEnd ST-9 Multiple Well Test Tray (Item (1001). Contents: 8 applicators to perform up to 72 skin tests. |  |  | March 22, 2016 | 2016 | Medscience Inc | During an FDA inspection it was found that the products are marketed without a cleared 510k. | No Marketing Application | MedScience sent an Urgent Recall letter dated March 28, 2016, to their customers to inform them of the product recall. The letter instructed customers to discontinue and refrain from further use of recalled products and to contact MedScience to arrange the return of products still in possession by the consignee. MedScience set up a phone line and email address for customers to ask questions and obtain further direction: (800) 393-8817 and recall@medscienceinc.com. |  | 11400 | 11,400 (Item 1000) &amp; 380 (Item 1001) | Nationwide Distribution to AZ, CA, CO, CT, FL, GA, LA, MD, MI, NJ, NY, PA, TX, VA, and WA. | Open, Classified |  |  | 11400
Z-1865-2016 | 73826 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The ARCHITECT Estradiol Kit |  |  | March 24, 2016 | 2016 | Abbott Laboratories | Interaction of Fulvestrant with the ARCHITECT Estradiol assay, leading to falsely elevated estradiol results in patients treated with Fulvestrant. | Device Design | The firm sent a Product Correction letter dated 3/24/2016 to all affected consignees describing the nature of the recall as well as instructions to follow. The necessary actions included the following:Consignees are asked to review the letter with their Medical Director. Patients undergoing Fulvestrant therapy should not be tested with the ARCHITECT Estradiol assay. Complete the Customer Reply Form. If the consignee has forwarded the product listed to other laboratories, a copy of Product Correction letter should be provided to them.The customer is asked to retain the letter for their laboratory records.Customers with questions should contact either Customer Service at 1-877-4ABBOTT for the US; or, the local area customer service for outside of the US. |  | 39346 | 39,346 units | Worldwide Distribution -- US, to the following states: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NV, NY, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY; and, the countries of ALBANIA, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOVI, BRAZIL, BRUNEI, BULGARIA, CAMEROON, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GAZA &amp; JERICHO, GEORGIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HUNGARY, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, IVORY COAST, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KUWAIT, KYRGYZSTAN, LATVIA, LEBANON, LIBYA, LITHUANIA, LUXEMBOURG, MACEDONIA, MALAYSIA, MALDIVES, MALI, MAURITIUS, MEXICO, MONTENEGRO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, OMAN, PAKISTAN, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REPUBLIC OF YEMEN, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TOGO, TRINIDAD &amp; TOBAGO, TUNISIA, TURKEY, UAE, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM and ZIMBABWE. | Open, Classified |  |  | 39346
Z-1867-2016 | 73684 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Integra MicroFrance Monopolar and Bipolar Electrosurgical Instruments |  |  | March 07, 2016 | 2016 | Integra LifeSciences Corp. | Reports relating to unintended patient or operator burns or operator electric shock for some of Integra's MicroFrance electrosurgical instruments. | Other | Integra sent an "Urgent Medical Device Correction" Field Safety Notice and an addendum dated March 7, 2016 to their customers via traceable courier service. |  | 2487 | 2,487 Instruments | AR, CA, CO, FL, IN, KY, MA, MI, NC, NJ, NY, NM, OH, TN, TX, WA and WI | Open, Classified |  |  | 32633
Z-1846-2016 | 73871 | No | NUJ | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Reprocessed Microline Scissor Tips intended to be used with a reusable hand piece and designed for minimally invasive and open surgical procedures to facilitate cutting, preparation, mobilization, and coagulation of tissue. Model #'s MIC3112, MIC3122, MIC3132, MIC3142, MIC3152, MIC3162. |  |  | April 26, 2016 | 2016 | STERILMED, INC. | Sterlilmed is recalling all unexpired Scissor Tips because of a potential manufacturing defect that can possibly result in a patient injury. | Process design | Consignees were sent on /26/2016 a Sterilmed "Urgent Medical Device Recall" letter dated April 26, 2016. The letter was addressed to either Materials Manager, Risk Manager and Operating Room Director or Distributor. The letter described the problem, the product involved in the recall and the actions required. Advised consignees to immediately quarantine and return the affected product and to complete and return the "Business Reply Form" to Sterilcycle.Advised Distributors to forward the letter to their customers if product was sent to them.For questions contact Sterilmed's Customer Care Team at 1-888-541-0078 . |  | 11326 | 11,326 | US: AL, AR, CA, CO, FL, GA, IN,IA, KY, MD, MA, MI, MN, MO, NV, NJ, NC, OH, OK, OR, PA,, SC, TN, TX, VA, WA, WV, WI.Canada | Terminated | April 05, 2017 | 344 | 11326
Z-1872-2016 | 73961 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography X-ray Systems |  |  | April 08, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The firm became aware of a problem where the system may not map Varian drive after CT user logout/login. | Software design | The firm sent a Customer Information letters dated March 8, 2016 to their customers. The letter identified the affected product, problem and actions to be taken. For information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 158 | 158 units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV &amp; WY. Product was also shipped to the following countries: Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, China, Colombia, Cuba, Estonia, Finland, France, Germany, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Myanmar, Netherlands, New Zealand, Norway, Palestine, Panama, Philippines, Poland, Portugal, RÂżunion, Russian Federation, Senegal, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania,Thailand, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Vietnam. | Open, Classified |  |  | 1097
Z-1845-2016 | 73640 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVIS EXERA Duodenovideoscope Olympus TJF Type 160F/VF |  |  | March 15, 2016 | 2016 | Olympus Corporation of the Americas | New reprocessing instructions for the Olympus TJF-160/VF duodenoscopes, consisting of revised manual cleaning and high level disinfection procedures they plan to implement as well as a new cleaning brush (MAJ-1534), which is enclosed with the notification. | Under Investigation by firm | Olympus notified customers on 03/14/2016 via an Urgent Safety Notification letter that referenced discontinuation of previously distributed copies of the TJF-160F/VF Reprocessing Manual. Customers were asked to inspect their inventory and discard any existing inventory of TJF-160F/VF Reprocessing Manuals. They were asked to update their information with the enclosed TJF-160F/VF Reprocessing Manual, which contains new brushing and flushing steps for the TJF-160F/VF elevator mechanisms and the forceps elevator recesses. Instuctions for cleaning and a new cleaning brush (used for cleaning the device) was included with the notification. Customers were asked to ensure all reprocessing personnel are completely knowledgeable and throughly training on the new reprocessing instructions in the new Reprocessing Manual. |  |  |  | Nationwide: | Open, Classified
Z-1839-2016 | 74139 | No | NBE |  |  | General &amp; Plastic Surgery | PMA | COSEAL Surgical Sealant Kit, 4 mL, Product Code: 934071; For use in vascular reconstructions to achieve adjunctive hemostasis by mechanically sealing areas of leakage. |  |  | May 12, 2016 | 2016 | Baxter Healthcare Corp | Potential for incomplete dissolution of the polyethylene glycol (PEG) component during the reconstitution of the product, which may affect the consistency of the hydrogel formation during use. | Nonconforming Material/Component | An Urgent Product Recall letter dated 5/13/16 was sent to customers to inform them that Baxter Healthcare Corporation is issuing a voluntary product recall for the product codes and lots listed below due to the potential for incomplete dissolution of the polyethylene glycol (PEG) component during the reconstitution of the product, which may affect the consistency of the hydrogel formation during use. The letter provides the customers with the list of affected products, hazards involved, and actions to be taken. Customers with questions regarding the recall communication, are instructed to contact Baxter Product Surveillance at (800) 437-5176, 8-5pm, Monday-Friday. |  | 6804 | 6,804 units | Distributed US (nationwide) including Puerto Rico and in the Bahamas, Canada, and Singapore. | Terminated | March 23, 2017 | 315 | 13206
Z-1841-2016 | 74079 | No | GDZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Miltex Premium Grade Self-Lock Chisel Blade Handle, Catalogue No. 4-405, manual surgical instrument. |  |  | March 11, 2016 | 2016 | Integra York PA, Inc | One lot of self-lock chisel blade handles have a wider slot that holds the chisel blade that may result in the inability to firmly secure the blade in the handle. | Nonconforming Material/Component | Integra mailed out an Urgent: Voluntary Medical Device Recall letter to customers informing them to evaluate their inventory for the affected product. If product was shipped to customers, they were asked to identify their customers and to stop shipping immediately. Integra requested that customers complete the attached form that was sent with the notification indicating whether or not they had the affected product on hand and customer service will contact them to provide a Return Authorization number for the product's return. Customers could reach Customer Service at 866-854-8300. |  | 678 | 678 Chisel Blade Holders | Worldwide distribution. US nationwide, Mexico, and Korea. | Terminated | April 17, 2017 | 402 | 678
Z-1871-2016 | 73704 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ergon III Series Skyboom Monitor Bracket. Part # 3FCM1 Flatscreen Bracket, single monitor mount, 27inch max. screen size |  |  | March 28, 2016 | 2016 | Skytron, Div. The KMW Group, Inc | Potential for the weld to fail at the pivot point of the monitor bracket. This weld failure could cause the display interface and bracket to fall from its mounted location, resulting in injury. | Process design | Skytron LLC initiated a voluntary recall of twenty-four monitor brackets used as a mounting point for display interfaces utilized in clinical applications for viewing purposes and to provide an articulating point of movement on the ceiling mounted radial arm suspension system of the Ergon 3 equipment pendant /booms via telephone calls to customers, and written notices sent via certified USPS mail on 04/01/2016.All potentially affected users will receive a field correction notification directly from Skytron, LLC along with the contact information of their local Skytron distributor. Skytron distributors must contact identified users within the universe of affected products to schedule a time replace the affected monitor bracket(s).Skytron will provide a replacement of the monitor bracket. The replacement monitor bracket(s) and installation must be conducted by you, the Skytron distributor. The affected monitor bracket(s) and Service Report indicating that the monitor bracket(s) was replaced must be returned to Skytron.Please contact Larry Perez, Vice President of Technical Services, at 616-656-1187 or lperez@skytron.us, or Thao Selleck, Complaint Resolution Analyst, at 616-656-6403 or tselleck@skytron.us with any questions. For after-hours support, please call 1-800-SKYTRON. |  | 24 | 24 | Worldwide distribution. US states: VA, VT, UT, GA, NJ, and IL; Japan. | Terminated | July 25, 2017 | 484 | 24
Z-1883-2016 | 73948 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Torque Limiting Screwdriver: |  |  | April 26, 2016 | 2016 | Greatbatch Medical | Greatbatch Medical has initiated a global Field Recallfor Torque Limiting Devices. Greatbatch identified through laboratory testing that the required sterility assurance level (SAL) of 10-6 is notachieved on torque limiting devices when subjected to steam sterilization cycles identified in the InstructionFor Use (IFU) provided with the device. | Device Design | Consignees were sent on 4/26/2016 a Greatbatch "Urgent Medical Device Recall" letter dated April 26, 2016. The letter described the product involved in the recall, the reason, and the risk to health. The letter advised consignees to identity and quarantine the product and to complete and return the Filed Recall Response Form. For questions contact your local Greatbatch Medical representative, at 1-763-951-8235, or e-mail FieldActionCenter@Greatbatch.com. |  | 1110 | 1110 | US: IN, MI, TX.OUS: GERMANY, FRANCE, ITALY, SWITZERLAND. | Open, Classified |  |  | 14128
Z-1881-2016 | 73978 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Firebird NXG Spinal Fixation System Set Screw Driver used to insert a set screw into a spinal fixation construct. Rx Only Provided non Sterile. |  |  | April 27, 2016 | 2016 | Orthofix, Inc | There is a possibility that the subject set screw driver could disengage from the torque limiting handle when the counter torque wrench is allowed to contact the counter torque handle. | Device Design | Affected distributors will be provided a written notification of field correction via USPS Certified Mail or FedEx on 4/27/16. Distributors will be additionally instructed to provide field correction notifcations to hospitals, surgeons, or other distributors that they may have further ditributed the product to. |  | 36 | 36 units | KS, NV, TX, CA, ID, FL, MA, MI, NJ, AZ, NH, SC, UT, WA | Terminated | November 23, 2016 | 210 | 36
Z-1880-2016 | 73766 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | STAT-Check Resuscitation Bag |  |  | March 15, 2016 | 2016 | Ventlab LLC | Stat-check CO2 indicators are non-functional. The indicator is yellow when not in use and should be purple when not in use. The yellow indicates a 5% CO2 level, cardiac massage and ventiliation being performed correctly. The color will not change during use and will not be an indicator of CO2 level or improper performance of cardiac message and ventilation. | Process control | Ventlab sent an URGENT: MEDICAL DEVICE PRODUCT REMOVAL notification dated March 14, 2016, on March 15, 2016, electronically to all affected customers. Customers were instructed to return the affected product and notify affected consignees of the recall. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any questions regarding this action, please call Ventlab at 1-844-635-5326, or email at PFA@ventlab.com. |  | 2430 | 2430 | Nationwide distribution to AZ, MS, NY, MO, FL, WI, TX, FL, and CA. | Open, Classified |  |  | 2430
Z-1876-2016 | 73802 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastr Liquid Mouse Monoclonal Antibody Renal Cell Carcinoma Marker (RCC) is recommended for use as part of a panel of antibodies in the definition of tumors of proximal nephrogenic differentiation. |  |  | April 05, 2016 | 2016 | Leica Microsystems, Inc. | Testing of retained units indicated that Lot numbers of CDX2 and RCC indicated are not stable up to the expiry date on the product labeling. | Nonconforming Material/Component | An Urgent: Medical Device Recall Notification dated 4/5/2016 was sent to consignees with the affected lots of CDX2 and RCC marker. The customers were notified of the problems as well as the following actions to be taken by the consignees: Do not use or continue to use the reagent lots listed above, as they have the potential to result in inadequate tissue staining and lead to weak or false negative pathology results.Appropriately destroy any unused or partially used affected lots of the reagentÂż Complete the attached Recall Notification Acknowledgement Form that this action has been undertaken.The firm also request that the notification is passed to all those within their organization who need to be aware of this issue.The firm requests consignees to return the completed acknowledgement form to LMGRA@Leica-Microsystems.com to receive credit for affected lots that have been destroyed by the consignees laboratory. |  | 270 | 270 units | Distributed in the following states AR, CA, CT, FL, ID, IL, IN, MT, OH, PA, TN and TX, and in Canada. | Open, Classified |  |  | 720
Z-1885-2016 | 73885 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista 500 Intelligent Lab System running on software versions V.3.6.1 SP1 or V.3.6.2, Device listing # D011374, All serial numbers and lots are affected. |  |  | March 24, 2016 | 2016 | Siemens Healthcare Diagnostics Inc. | Discrepant patient results on Dimension Vista Intelligent Lab Systems. Siemens Healthcare confirmed a software defect which, in a very specific set of circumstances, results in the Dimension Vista System omitting an aliquot probe rinse between sample aspirations when processing tubes in Sample Racks that are front loaded on the Dimension Vista System. | Software design | Siemens sent an Urgent Medical Device Correction letter dated March 24, 2016, to all affected customers. Customers were advised to follow the instructions contained in the "Urgent Medical Device Correction" notification letter to help minimize or eliminate the impact that this specific software defect could cause until a software correction could be made (i.e instructions to minimize the impact of the software defect due to "Bad Sample Detect " False Transition" errors when processing front loaded samples) and (how to eliminate the impact of the software defect when processing front loaded samples). Customers were also asked to review the recall notification letter with their Medical Director and complete and return the Field Correction Effectiveness Check Form within 30 days. Customers with questions were advised to contact Siemens Customer Care Center - Technical Solutions at 800-441-9250 or their local Siemens technical support representative.For questions regarding this recall call 302-631-8248. |  | 1301 | 1301 units | Nationwide Distribution to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, , ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Puerto Rico. | Terminated | March 13, 2017 | 354 | 5202
Z-1878-2016 | 74004 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | Replacement HEPA filter for Natus Sterile-Drier Model 43 and Model 44. Filter Catalogue Part Number: 54343, 54344, and 100105-01. |  |  | April 22, 2016 | 2016 | Natus Medical Incorporated | The replacement HEPA Filters have been manufactured with sealant that may not be properly cured. Improperly cured sealant can emit a toxic gas when heated, and can drip down during use, contaminating the interior chamber of the Sterile Drier and its contents. | Nonconforming Material/Component | Natus sent an " URGENT RECALL NOTIFICATION/ FIELD SAFETY NOTICE" Olympic Sterile Driers Model 43 &amp; 44, Replacement HEPA Filter Contamination letter (dated April 22, 2016) was sent to all affected customers on May 2, 2016. for the first set of customers and then sent on May 12, 2016, for the second set of customers, who purchased 57 replacement filters. Customers are advised "do not install any replacement HEPA filter " if the serial number is one of the 21 shown as affected serial numbers on the letter. The URGENT RECALL NOTIFICATION/ FIELD SAFETY NOTICE Olympic Sterile Driers Model 43 &amp; 44, Potential HEPA Filter Contamination (dated May, 2016 and released on May 12, 2016,) was sent to customers on May 16, 2016, who purchased 11 sterile driers with the installed filters. Customers are advised "do not use the Sterile drier " if it contains a HEPA filter that is identified with one of 21 affected serial numbers. Customers are advised not to install any defective replacement HEPA filters or do not use sterile drier if it currently contains defective HEPA filter. Please discontinue use of the Sterile Drier immediately if it contains defective HEPA filter as heat may cause the uncured adhesive to drip into the chamber.Customers contact Natus to obtain the RMA number and return unused filters and Natus will provide a replacement filter.Please contact Natus at Seattle_Technical_Service@natus.com or call 866-940-7143 (option 1, Olympic products) for next steps regarding a contaminated Sterile Drier. |  | 21 | 21 filters | distributed in the US including AZ, CA, FL, IL, KS, MA, MD, MN, NC, NH, NY, OH, OK, PA, VA, WA, WA DC,and WI and in Canada and Singapore. | Open, Classified |  |  | 21
Z-1870-2016 | 73743 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | General Surgery Pack, part number 006880-3. This is a custom convenience surgical instrument disposable kit. |  |  | March 21, 2016 | 2016 | Windstone Medical Packaging, Inc. | These packs contain item #4037 Kerlix X-ray Detectable Laparotomy Sponges which were recalled by Medtronic due to the potential for sterile barrier to be comprised at the packaging seal. | Packaging | Aligned Medical Solutions (AMS) sent the Urgent Medical Device Recall letter, dated March 21, 2016, and the Recall Reply Form to their consignees via email.AMS informed consignees that their AMS packs (surgical kits) contained Kerlix X-Ray Detectable Laparotomy Sponge which was recalled by Medtronic.Consignees were advised to follow the over-labeling instructions and dispose the Kerlix X-Ray Detectable Laparotomy Sponge at the time the surgical kit is opened for use. Consigneeswere instructed to follow the "Instruction for placing Recall Notice Label to Packs" and label any kits remaining inventory.Consignees should return the Reply Form via fax at 407-865-7456 or call 321-527-7714 for any questions about this recall. |  | 170 | 170 units | Nationwide distribution to Louisiana and Montana. | Open, Classified |  |  | 1192
Z-1612-2016 | 73825 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator |  |  | April 14, 2016 | 2016 | Philips Electronics North America Corporation | MRx monitor/defibrillators could fail to charge because the therapy printed circuit board may have been loaded with an incorrect electronic component. | Employee error | Philips Healthcare sent an Urgent Medical Device Correction letter dated April 2016, on April 14, 2016, to all affected customers. The letter informed customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of providing a replacement MRx unit, free of charge, for all units affected by this issue. When the replacement unit arrives, please use the return shipping label and packaging provided with the replacement MRx unit to return the affected unit to Philips. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: You can continue to use your MRx prior to receiving a replacement unit. But, if you observe a red X in the Ready for Use (RFU) indicator, you should follow the troubleshooting instructions in the MRx IFU.Customers with questions were instructed to contact their local Philips representative or call 1-800-722-9377. |  | 17 | 17 | Worldwide Distribution - US including OH, CA, NC and Internationally to China, Iraq, Japan, and Korea. | Open, Classified |  |  | 17
Z-1906-2016 | 73986 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Non-Vented Dispensing Pin with One-Way Valve and Luer Lock Ultrasite Valve |  |  | April 14, 2016 | 2016 | B. Braun Medical, Inc. | BBMI recently identified a potential for a tear in the outer blister packaging on a limited number of units. The tears have been identified in the blister package paper, where the unit label writing is located. | Process control | B. Braun notified customers via letter to inform them of the reason for the recall, the risks and affected lots involved were identified and customers were asked to discontinue use of the product and complete the Product Removal Acknowledgement form and return to B. Braun Medical Inc. Quality Assurance department by faxing it to (610)849-1197 or e-mail to: PA_QualityAssurance.BBMUS_Service@bbraun.com and a BBMI Customer Support Representative will contact the consignee to provide instructions for handling the affected product and arrange for its return to BBMI. |  | 11000 | Approximately 11,000 units available in commercial channels | US Nationwide Distribution in the states of: AL, AR, AZ, BC, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN. TX, UT, VA, VT, WA, WI, WV and Puerto Rico | Terminated | March 13, 2017 | 333 | 11000
Z-1900-2016 | 73200 | No | MRA |  |  | Orthopedic | PMA | Novation Ceramic AHS Cluster-Hole Shell, Plasma Coated, 54mm, O.D. (Acetabular Cup), Catalog Number 140-01-54; For use in primary total hip arthroplasty. |  |  | April 12, 2016 | 2016 | Exactech, Inc. | Devices were inadvertently sent to packaging and finished goods prior to the PMA site-change supplement approval. | Packaging process control | Customers were notified 4/12/2016 with a Product Recall Advisory Notice of affected products.In order to comply with applicable regulations and Exactech policies: " Immediately cease distribution or use of these products. " Extend this information to your accounts that may have this product in their possession. " Identify and quarantine any of the subject devices in your inventory " Complete and Return the attached Recall Inventory Response Form to Exactech, per the instructions provided on the notice itself. According to policy, please contact your Exactech inventory representative within 5 business days to confirm quantities at your location. If you have any inventory restocking questions related to this issue, please contact Kaya Davis at kaya@exac.com or 1-800-392-2832. Please return the products to the following address:Exactech Distribution Center 2411 NW 66th Court Gainesville, Florida 32653Attention: Kaya Davis kaya@exac.com 352-337-3915 (Fax) |  | 3 | 3 devices | Distributed in FL, NJ, NY, OH, OK, RI, TN, VA and HI. | Open, Classified |  |  | 20
Z-1879-2016 | 74164 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | May 18, 2016 | 2016 | Elekta, Inc. | When DICOM exporting a 3D Monaco plan and the "Composite Field Sequencing" (CFS) checkbox is selected, the Y jaws will snap to the port shape when they should remain where they were defined. | Software design | Field Safety Notice 382-01-MON-004 was sent to customers on May 18, 2016. The notice informs users of the specific product and version numbers affected by both issues, and any work around that can be used to avoid the issue. The customers were instructed to complete and return the Field Safety Notice Acknowledgement form to their local Elekta as soon as possible. |  | 1401 | 1401 units | Worldwide Distribution - US Nationwide in the states of AK, AZ, CA, CO, FL, GA, IL, IN, KY, MI, MN, MO, MT, NV, NH, NJ, NY, NC, OK, OR, PA, SD, TX, VA, WA, WI, and the countries of: Alger, Antiqua &amp; Barbuda, Argentina, Australia, Austria, Belarus, Botswana, Brazil, Bulgaria, Canada, China, Columbia, Croatia, Cyprus, Czech Republic, Ecuador Egypt, Estonia, France, Germany, Greece, Hong Kong, India, Indonesia, Iran, Israel, Italy, Japan, Lithuania, Malaysia, Malta, Mexico, Morocco, Nabibia, Netherlands, New Zealand, Norway, Panama, Peru, Poland, Portugal, Russia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka,Sweden, Switzerland, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, Zimbabwe | Open, Classified |  |  | 1401
Z-1691-2016 | 74106 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Sedecal SA Mobile Diagnost w DR x-ray system |  |  | March 23, 2015 | 2015 | Sedecal USA, Inc. | Due to a software defect, the system may sporadically apply the default x ray exposure parameters for an adult ( patient type : Normal ) even though the patient type " Newborn" was selected and is displayed in the generator control area of the Eleva User Interface. | Radiation Control for Health and Safety Act | Sedecal USA has developed a new software release to address this issue. The distributor of this product, Philip, will modify the systems in the field to correct the defect at no cost to the customer.1. You will contact customers and initiate a software update to correct the two defects.2. You will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.For further questions, please call (847) 394-6960. |  | 145 | US - 145 | US Distribution - TX, OH, NY, MD, IN, PA, DE, WA, NH, AZ, GA, NC, UT, VT, CA, LA, NJ, MN, FL, WI, HI, IA, MA, ME, TN, CT, ID, MI, AR and KS. | Open, Classified |  |  | 145
Z-1833-2016 | 73799 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Plum A+ is a dual-line volumetric infusion system designed to meet the growing demand for hospital-wide, as well as alternate site and home healthcare, standardization. With its primary line, secondary line, and piggyback fluid delivery capability, the Plum A+ is suited for a wide range of medical/surgical and critical care applications. Full compatibility with LifeCare Plum Series administration sets and accessories. |  |  | March 25, 2016 | 2016 | Hospira Inc. | Firm's investigation determined that the Proximal Occlusion Alarms observed in the field were attributed to the mechanism calibration process being performed with an out of specification metal cassette, resulting in an offset in proximal pressure. | Process control | Hospira Inc, sent an "URGENT MEDICAL DEVICE CORRECTION" letter on March 31, 2016, to the direct consignees of record via traceable mailing which included a data capture log containing a list of devices located with each customer. Delivery will be confirmed via traceable mailing. Consignees who further distributed the product were asked to notify their customers. Please take the following actions related to Urgent Medical Device Correction letter notification: 1. Inform potential users in your organization of this notification.2. Complete the attached reply form and return it to the fax number or e-mailaddress on the form, even if you do not have the affected product.3. If you have distributed the product further, notify your accounts that receivedthe product identified above of this notification and ask them to contactStericycle at 1-877-479-8088 (M-F, 8am-5pm ET) to receive a reply form.Hospira will contact you to make arrangements for your impacted Plum 360 I Plum A+ devices to be re-calibrated to correct this issue.: |  | 2567 | 2567 Plum A+ infusion pumps. | U.S. distribution : Natrionwide. Foreign :Brazil, Canada, Chile, Colombia, Ecuador, Guatemala, Mexico, Oman, Peru, Taiwan, Turkey, United Kingdom and Uruguay. | Open, Classified |  |  | 4057
Z-1909-2016 | 73979 | No | PFX |  | Immunology | Immunology | 510(k) | CytoScanÂż Dx Assay. The CytoScan Dx Assay consists of five reagent modules, wash buffers (WS), a microarray kit, and analysis software. |  |  | April 14, 2016 | 2016 | Affymetrix Inc | During a recent Ohio site inspection by the FDA it was discovered that warning statements needed to be included with the product shipment. | Labeling design | A letter dated April 14, 2016 was sent to customers. The letter stated the reason for the recall. It also informed customers that there is no impact on the performance of the assay. Customers were asked to either visit the IVD web portal (www.affymetrix.com/ivd) and review the warning statements; or, request a hard copy free of charge. |  | 507 | 507 kits | Distributed in the US to the states of AR, CA, DC, MI, MA, MN, NC, NM, NY, PA, SC, TN, TX, WA and WV. | Open, Classified |  |  | 507
Z-1911-2016 | 74140 | No | LHJ | Staphylococcus aureus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Sure Vue Color Staph 150T, Sure Vue Color Staph 300T, Sure Vue SELECT Staph, biokit SELECT Staph 75T |  |  | February 19, 2016 | 2016 | Integrated Science Systems Inc | Customer complaints of very weak reactions with the kit, or false negative results with different Staph aureus cultures. | Process control | Integrated Science Systems sent a Recall Notification letter dated March 29, 2016. Integrated Science Systems requested that unopened kits in inventory be returned. Once the root cause has been determined, the kits will be reworked and returned to INOVA for distribution. For questions call 706-922-8686. |  | 2625 | 2625 kits | Nationwide distribution to CA only. | Open, Classified |  |  | 2625
Z-1695-2016 | 74083 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare DigitalDiagnost stationary X-ray system |  |  | November 10, 2015 | 2015 | Philips Healthcare | The detector may signal that it is ready for acquisition when it actually is not, resulting in failure to properly acquire the X-ray image. | Radiation Control for Health and Safety Act | Philips Healthcare plans to correct the defect at no cost to customers by implementing a field correction involving the installation of new software.1. You will contact customers and initiate a software update to correct the defect.2. You will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.For further questions, please call 1-800- 722-9377. |  | 160 | 160 | USA (nationwide) Distribution to the states of : ( DuraDiagnost) to the states of : OH, CA, WA and NE.(DigitalDiagnost) to the states of : AK, AZ, AL, CA, CT, CO, DE, FL, GA, HI, ID, KY, LA, MA, MN, MD, MS, ME, MO, NE, NY, ND, OH and PA. | Terminated | May 03, 2017 | 540 | 165
Z-1910-2016 | 74070 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker FlowPort II Adapter: Model number: 00CAT00778 |  |  | May 04, 2016 | 2016 | Stryker Corporation | Complaints were received for the Stryker FlowPort II Adapter, and investigation found that the scope sat in the adapter 180 degrees in the wrong direction | Device Design | A recall letter was not sent as all affected devices have been returned. A ship hold was placed on the device on October 16, 2015. |  | 21 | 21 units | GA, UT, MD, NJ, MT | Terminated | June 03, 2016 | 30 | 21
Z-1715-2016 | 74149 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Brilliance iCT SP Computed Tomography X-Ray System |  |  | September 06, 2015 | 2015 | Philips Healthcare | Software Defects resulting in: (1) sagittal result shortened for axial scans; (2) single series displayed on console viewer when Creating multi-planner reconstruction; (3) scan length changes with a change in field of view; (4) DoseRight algorithm miscalculations leading to incorrect CTDI values; (5) scan length changes during subsequent axial results, (6) surview scan lengths near 135mm or 184mm | Radiation Control for Health and Safety Act | Philips plans to initiate a field corrective action that will install a new software release to address the issue. Philips will modify the systems in the field to correct the defect at no cost to the customer.1. You will contact customers and initiate a software update to correct the defect.2. You will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.For further questions, please call (1-800) 722-9377, Option 5 |  | 423 | US-423 | US Distribution - AZ, PA, FL, DC, W VA, NC, SC, AR, OK, OH, NY, NV, TX, PA, IN, CA, MN, MA, DC, MS, MD, OR, IL, CO, MO, VT, IA, AL, LA, GA and WA. | Open, Classified |  |  | 423
Z-1912-2016 | 74080 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Synthes 8.0 mm Flexible Shaft Reamer, 360 mm, Orthopedic manual surgical instrument. Part number 351.43 |  |  | May 02, 2016 | 2016 | Synthes (USA) Products LLC | Devices did not pass the biological safety evaluation for cytotoxicity following exposure to test conditions. The high growth inhibition levels observed during testing could be attributed to corrosion of the device at solder points. This could potentially be reproduced during use and reprocessing. | Process design | Depuy Synthes sent an Urgent Notice: Medical Device Recall letter on 05/02/2016 to all affected consignees. Consignees were instructed to do the following: review your inventory immediately to see if any of the affected products were being stored, and call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement order, to complete the Verification Section (page 3) and return with the product to Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. |  | 244 | 244 units | US nationwide distribution. | Open, Classified |  |  | 244
Z-1915-2016 | 74196 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Revolution CT |  |  | May 09, 2016 | 2016 | GE Healthcare, LLC | GEHC has become aware of a potential safety issue in the Revolution CT Patient Table that can result in unintended motion in cases involving large patients. The Revolution CT cradle can descend on its own after activation by the user. This can put the patient at risk for harm. The harm can involve the impinging or crushing of a hand if caught between the cradle and CT inner gantry bore. No injurie | Process control | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter ,GEHC Ref# 25468, dated May 9, 2016 to its Consignees. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. The consignees were instructed to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions; continue to use your Revolution CT product; follow the Patient Positioning warning outlined in the Revolution CT User Manual; to prevent pinching or crushing of the patients extremities, keep the patients hands and feet away from the edge of the moving table top/cradle and its surrounding equipment, or between table base and side panels of the table; and take special care when positioning physically large patients. If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 197 | 197 (US = 68; OUS = 129) | Worldwide Distribution: US (nationwide) including states of: AL, CA, CT, FL, GA, HI, IL, IN, KS, NE, NY,NC, OH, OR, PA, TN, TX, UT, A, WA, WI; and countries of: Australia, Belgium, Brazil, Canada, China, Denmark, Estonia, France, Germany, Hong Kong, Indonesia, Italy, Japan, Korea (Republic of), Kuwait, Mexico, New Zealand, Norway, Peru, Poland, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 197
Z-1914-2016 | 73887 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System, Catalog Number 6802413, Unique Device Identifier No. 10758750002740; and VITROS 5600 Integrated System (refurbished), Catalog Number 6802915, Unique Device Identifier No. 10758750007110; IVD. |  |  | April 06, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics confirmed a software timing anomaly in which the VITROS 3600 and 5600 Systems may unexpectedly aspirate from or dispense into an unintended sample that may lead to erroneous or contaminated results. | Software design | On 4/6/2016,URGENT PRODUCT CORRECTION NOTIFICATION letters (Ref. CL2016-076, dated 4/06/2016) were sent via FedEx overnight and/or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected product and advised them of the issue and the required actions. Follow enclosed instructions to help decrease the probability of the software anomaly until the software update is installed. Discuss any concerns with previously report results with their Laboratory Medical Director to determine the appropriate course of action. Install Software Version 3.2.3 at their earliest convenience. Foreign affiliates were informed of the issue by e-mail on 4/6/2016, and instructed to notify their consignees of the issue and required actions. |  | 7881 | VITROS 5600: US: 1161 units, Foreign: 1063 units; VITROS 5600 Refurbished: US: 19 units, Foreign: 38 units | Nationwide, Puerto Rico and Foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 8686
Z-1921-2016 | 73658 | No | JTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Oxoid Nitrocefin (+ Reconstitution Fluid Batch) REF: SR0112C, Oxoid Ltd. For the rapid chromogenic detection of beta-lactamase activity in bacterial cells. in vitro diagnostic |  |  | March 24, 2016 | 2016 | Remel Inc | Individual bottles of product may produce weak beta-lactamase reactions. Use may result in delayed or false negative reporting. | Under Investigation by firm | Customers were notified by letter beginning on March 24, 2016. |  | 2 | 2 units were imported and distributed domestically | US distribution in MO. | Open, Classified |  |  | 2
Z-1920-2016 | 74053 | No | JEY | Bone plate. | Dental | Dental | 510(k) | 1.5 mm Rapid Resorb Orbt1 F1 P1 w/Bending Template / 35 mm - Sterile, Rapid Resorbable Fixation System. 1.5 mm Rapid Resorbable Orbital Floor Plate with Bending Template. |  |  | May 02, 2016 | 2016 | Synthes (USA) Products LLC | Labeling errors. The outer packge label expiration date may exceed the expiration date of the subcomponents. | Under Investigation by firm | DePuy Synthes sent an Urgent Medical Device Recall Notice dated May 2, 2016, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory immediately and remove the affected lots from their stock. Customers were instructed to call 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement product., customers were also asked to complete the Verification Section and scan/email to Synthes4096@stericycle.com. Customers with questions were instructed to call 610-719-5450. |  | 14 | 14 units | Nationwide Distribution to DE, NC, CA, OK, PA, SC, WI, CO, NJ, TX, UT, IL, MO, MI, AL, KS. | Open, Classified |  |  | 55
Z-1917-2016 | 74148 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona Trabecular Metal Tibia; Persona Trabecular Metal Tibial Plate Instruments and Modular Brackets; Used to facilitate the implantation of joint replacement implants. |  |  | May 11, 2016 | 2016 | Zimmer Biomet, Inc. | Zimmer Inc. is initiating a recall of specific Persona" Trabecular Metal" Tibial Plate Instruments and Modular Brackets, due to incidents of radiolucent lines and loosening in implants. | Component design/selection | Zimmer Inc. is initiating a voluntary recall of Persona" Trabecular Metal" Tibia / Tibial Plate Instruments and Modular Brackets following an increase in complaints of radiolucent lines and loosening, and earlier recall of the Persona" Trabecular Metal" Tibial Plates . All sizes and lots of the affecteddevices are being removed from distribution. Recall notices were issued to affected distributors, hospitals, and surgeons on 05/11/2015 via mail. Customers may contact the following with any questions: 1-877-946-2761 between 8:00 am and 5:00pm EST. Customers are asked to do the following: 1. Review the notification and ensure affected personnel are aware of the contents.2. Assist Zimmer Biomet sales representative with the quarantine of any affected product.3. Zimmer Biomet sales representative will remove the recalled product from customer facilities.4. Complete the Certificate of Acknowledgement Form and return to: corporatequality.postmarket@zimmerbiomet.com. |  | 10256 | 10,256 | Worldwide Distribution -- US, including the states of AK, AL, AZ, CA, FL, GA, HI, ID, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, and WI; Foreign: Australia, New Zealand, Korea,Austria, Belgium, Germany, Italy, France, Luxembourg, Netherlands, South Africa, Spain, Sweden, Switzerland, and United Arab Emirates. | Terminated | June 02, 2017 | 387 | 10256
Z-1922-2016 | 74125 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Actifuse ABX/Actifuse MIS System |  |  | August 18, 2015 | 2015 | Apatech Limited | Baxter Healthcare is issuing a voluntary recall for all lots of Actifuse ABX and Actifuse MIS System products with expiry before July 29, 2017 due to the possibility that the products may have endotoxin levels above specification criteria. | Process control | The firm, Baxter, sent an "Urgent Product Recall Notification" letter dated 8/18/2015 and a follow-up letter date 4/29/2016 to consignees via FedEx, overnight mail. Baxter also notified affected customers by phone. The letters described the product, problem and actions to be take. Baxter Healthcare has instructed customers to do the following: 1. Locate and remove all affected product from their facilities.2. Contact Baxter Healthcare Center for Service to arrange for return and credit at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time, Monday through Friday.3. Complete the attached customer reply form and return it to Baxter by either fax it to 224-270-5457 or scanned e-mail to fca@baxter.com.4. If you distribute this product to other facilities or departments with your institution, notify all affected parties of this communication.For general questions regarding this communication, please call Baxter's Therapeutic Medical Affairs at 805-657-0007 or 862-432-7941. |  | 11696 | 11696 units | Worldwide Distribution: United States: Nationwide (including Puerto Rico) and countries of: Canada, Austria, Belgium, Switzerland, Germany, Denmark, Spain, Finland, France, Great Britain, Ireland, Italy, The Netherlands, Norway, Sweden, Slovenia, Czech Republic, Kuwait, Mexico, Brazil, Australia, Taiwan, and New Zealand. | Open, Classified |  |  | 11696
Z-1924-2016 | 74176 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Tearaway Introducer, Model # VS203, VS303, 510 K # 130687, packaged individually in a pouch, 5 pouches per carton, lot # MBZ140, MBZL450, MBZV930 MBZZ490 |  |  | May 05, 2016 | 2016 | Medical Components, Inc dba MedComp | This recall has been initiated due to the product labeled with the incorrect expiration date. | Labeling mix-ups | MedComp sent a Product Recall letter dated May 5, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to evaluate their inventory and quarantine for the return of all un-used affected product. They were asked to contact Customer Service for a Return Goods Authorization (RGA) number at 215-256-9191. |  | 3122 | VS203 Lot# MBZX140 (100 units); VS303 Lot # MBZL450 (60 units); VS303 Lot # MBZV930 (65 units); VS303 Lot # MBZZ490 (33 units). | US Distributed to: FL, TX, NC | Terminated | June 16, 2017 | 407 | 3122
Z-1926-2016 | 74059 | No | LGW |  |  | Neurology | PMA | Medtronic RestoreSensor, Model 37714; and, RestoreSensor SureScan, (MRI), Model 97714, Implantable Neurostimulators. |  |  | May 03, 2016 | 2016 | Medtronic Neuromodulation | Loss of therapy during recharging of device. Reportedly, a charging session was terminated prior to obtaining recharge threshold voltage of 3.615, which triggered rapid battery depletion state. As a result, the device depleted to 1.925 volts in one to two days rather than the typical 30 days. Insufficient coupling between recharger and device during recharge session was found to be key factor. | Under Investigation by firm | Consignees were sent on 5/3/2016 a Medtronic "Urgent Medical Device Safety notification" letter dated May 2016. The letter was addressed to "Healthcare Professional". The letter described the problem and the product involved in the recall. The letter listed the recommendation to patients to follow current recharge instructions, paying particular attention to Charging Efficiency and Battery Charge Level indicators on the recharger. Consignees were also requested to complete and return the "Physician Reply Form:.For questions contact Medtronic Neuromodulation Technical services at 1-800-707-0933 weekdays 7 am - 6 pm CST. |  | 60941 | 60,941 | Worldwide Distribution | Terminated | August 04, 2017 | 458 | 60941
Z-1916-2016 | 74107 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Universal SlingBars |  |  | March 08, 2016 | 2016 | Hill-Rom, Inc. | The center bolt of the sling bar, which connects the bar to the patient lift, has failed during use. If the sling bar is not used as intended the bolt may be weakened. If the bolt is weakened there is a potential risk for breakage with the result of a free fall of the patient. This hazard could cause potentially minor to catastrophic injuries to a patient. | Labeling design | On 3/8/2016, MEDICAL DEVICE CORRECTION notifications were sent to the affected US customers via email and courier. On 3/24/2016, MEDICAL DEVICE CORRECTION notifications were sent to the affected Canadian customers via email and courier. On 3/29/2016, MEDICAL DEVICE CORRECTION notifications were sent to the affected Global customers via email and Docapost. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The notification provided new instructions for users with a lift manufactured prior to 11/19/2013, instructed the user on how to inspect for damage, and offers replacement at a 50% reduction in cost. If you have any questions concerning this Field Safety Notice, please contact Hill-Rom Technical Support, your distributor, or your Hill-Rom representative. |  | 163561 | 163,561 | Worldwide and US Nationwide Distribution AE, AO, AR , AT , AU , AZ , BD , BE , BG , BH , BM , BN , BR , CA , CH , CL , CN , CO , CS , CZ , DE , DK , EG , ES , FI , FR , GB , GR , HK , HR , HU , ID , IE , IL , IN , IR , IS , IT , JO , JP , KR , KW , LB , LI , LK , LT , LU , LV , LY , MA , MT , MX , MY , NL , NO , NZ , OM , PH , PK , PL , PT , QA , RO , RS , RU, SA , SE , SG , SI , SK , SY , TH , TN , TR , TW , UA , US , UY and ZA. | Terminated | May 15, 2017 | 433 | 163561
Z-1923-2016 | 74002 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | 4 ml FX Sodium Fluoride/Potassium Oxalate tube, Ref 454297 |  |  | April 26, 2016 | 2016 | Greiner Bio-One North America, Inc. | The 4 ml FX Sodium Fluoride/Potassium Oxalate tube may have no or low additive (sodium fluoride and potassium oxalate), which can effect analytical results for glucose and lactate. | Process control | Greiner Bio-One North America, Inc. sent an Urgent Product Recall letter dated April 26, 2016, to all affected consignees. The letter requested that they stop distribution of the product and to provide Greiner with a list of their sub-accounts. The letter included a Product Disposition/Site Confirmation form which is to be returned to Greiner Bio-one. Consignees with questions were instructed to call 800-515-8112. For questions regarding this recall call 704-261-7823. |  | 250800 | 250,800 pieces (209 cases) | Nationwide Distribution to CA, FL, IL, KY, NY, PA, VA and Puerto Rico | Open, Classified |  |  | 250800
Z-1925-2016 | 73999 | No | LWG | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical Table Cushion (HybriGel Pads) |  |  | April 19, 2016 | 2016 | Berchtold Corp. | The Surgical Table Cushion (HybriGel Pads) could potentially contain staples that's secure two underlying fabrics, which would be captured on certain scans on patients, such as x-ray. | Nonconforming Material/Component | Berchtold Corporation sent an Urgent Medical Device Recall letter dated April 19, 2016, to all affected customers. The consignees were instructed to return the enclosed acknowledgment form by email to ComRecallCoordinator@stryker.com and then discard the HybiGel Pads. |  | 227 | 227 units | Nationwide Distribution | Terminated | March 31, 2017 | 346 | 227
Z-1935-2016 | 74225 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | MEGAÂż 7.5Fr. 30cc Intra-Aortic Balloon Catheter with Insertion Kit &amp; STATLOCKÂż |  |  | May 13, 2016 | 2016 | Datascope Corporation | It was discovered during post-quarantine packaging on 2/16/2016 that the IFU for the STATLOCK Sheath Stabilization device for Percutaneous Sheath Introducer (Vendor PN 2403097-1405R) was erroneously packaged with the Mega and Sensation Plus IAB Catheter Kits instead of the IFU for the STATLOCK Catheter Stabilization device for IAB catheters (Vendor PN 24030905-1406R). | Unknown/Undetermined by firm | Maquet Inc sent an Urgent Medical Device Field Correction Letter/Response Form dated 5/13/2016 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 31030 | 31,030 units (US); 15,604 units (OUS) | Worldwide Distribution. | Open, Classified |  |  | 31030
Z-1928-2016 | 74215 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Optima 1.5T MR430s MR Scanner, a diagnostic imaging device |  |  | May 09, 2016 | 2016 | GE Healthcare, LLC | The gas venting may not occur properly. A site reported a magnet quenching with subsequent venting of the cryogen gas into the magnet room. Cryogen gas that vents into the magnet room could cause critical deprivation of oxygen to users and patients. | Nonconforming Material/Component | GE Healthcare sent an " Urgent Medical Device Correction" letter dated May 9, 2016, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Detail, Product Correction &amp; Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 101 | 101 | US: AL, AZ, CA, CO, FL, ID, IL, IA,ME, NE, NV,NH, NJ, NY, NC, OH, PA, SC, SD, TN, TX, UT,VA, WA, WI.OUS:ArgentinaAustraliaAustriaBrazilCanadaChileColombiaDenmarkEgyptFinlandFranceGermanyGreeceHong KongIcelandItalyJapanKorea (Republic of)KuwaitMalaysiaNetherlandsNorwayPeruPolandPortugalRussiaSingaporeSouth AfricaSpainSwedenSwitzerlandUkraineUnited Arab EmiratesUnited Kingdom | Open, Classified |  |  | 198
Z-1929-2016 | 74265 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac Vero is a Patient Positioning System for Radiation therapy. |  |  | May 04, 2016 | 2016 | Brainlab AG | Potentially incorrect positioning when using Implanted Marker Detection with Brainlab ExacTrac Vero 3.5 | Device Design | A FIELD SAFETY NOTICE / PRODUCT NOTIFICATION was sent to customers on May 4, 2016 and distributed via email. An active reply by the consignees will be followed, if necessary via further different communication media.As of now for new installations- the Product Notification letter is handed over to the customer during installation (before clinical use) by the according Brainlab Service representative,- a signature is required by the customer that the contained User Corrective Action will be implemented before clinical use by the facility (to be adhered to by the users).Brainlab Instructions and actions required to be performed by the customer are described below.1. For Positioning and Verification, carefully review plausibility of implant marker fusions resulting in rotation angles of more than 10Âż. If in doubt consider the following options before proceeding:- Deselect one or several markers defined in the CT data set, if you suspect these markers to have migrated- Apply bony fusion instead of marker based fusion- Apply ConeBeam CT correction and verification instead of marker based fusion- Acquire a new CT scan and repeat treatment planning and definition of implanted marker positions based on the new CT data. For further details regarding the suggested proceedings, please consult the according Clinical User Guide. 2. For Dynamic Tracking continue to follow the instructions of the Clinical User Guide regarding the review of the correlation model based on implanted marker detections. Specifically: - Examine the correctness of localization of markers and targets in the X-ray sequence.- Verify the consistency of the Target Detection Curve and/or Correlation Model Curve. Irregularities of the curves are an indication for incorrect implanted marker detection. In this case please consider the options described in 1. before proceeding.If the customers require further clarification, they should contact their local Brainlab Customer |  | 27 | 27 Systems | Distributed in the states of Florida, New York, Texas and Ohio, and in the countries of Belgium, France, Germany, Italy, Japan and South Korea. | Open, Classified |  |  | 27
Z-1933-2016 | 73921 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cholesterol, Catalog No. OSR6x16 |  |  | April 13, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter has identified that Dipyrone (Metamizole) may cause erroneous, false low results in the following assays: Cholesterol OSR6X16, Triglyceride OSR6X118, Uric Acid OSR6X98, Lactate OSR6X93, and Lipase OSR6X30. | Device Design | Beckman Coulter sent an Urgent Medical Device Recall letter dated April 13, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions regarding the customer notification letter are instructed to contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support; (800) 223-0130 in the US and Canada. Outside of the US and Canada, customers are instructed to contact their local Beckman Coulter Representative. |  | 43877 | 43,877 units total (10,496 units in US) | Worldwide Distribution - US Nationwide including Puerto Rico and the countries of: Albania, Algeria, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia-Herzegovina, Botswana, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Eritrea, Ethiopia, Finland, France, Germany, Ghana, Gibraltar, Greece, Guadeloupe, Guam, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Lebanon, Lesotho, Liberia, Libyan Arab Jamahiriya, Lithuania, Malawi, Malaysia, Mexico, Morocco, Mozambique, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, Zambia, and Zimbabwe. | Open, Classified |  |  | 43877
Z-1942-2016 | 74172 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Norian Drillable Inject 3 CC-Sterile; catalog # 07.704.003S |  |  | May 11, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is initiating a voluntary medical device recall of unexpired and unopened part and lot numbers for the Norian Drillable Injects-Sterile due to the rotary pouch within the referenced lots and associated part numbers may potentially be labeled with the incorrect powder volume/size; however, the rotary pouch contains the correct powder volume/size according to the outer label. | Under Investigation by firm | DePuy Synthes sent an Urgent Medical Device Recall letter dated May 9, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with the affected products were asked to call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement product(s), Complete the Verification Section (page 3 of this letter) and Return the Verification Section (page 3 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Send a copy of the completed Verification Section by:Fax: 855-695-8597 or Scan/email: Synthes6797@stericycle.com. If they had any questions they were asked to call 610-719-5450. For questions regarding this recall call 610-719-5443. |  | 1246 | 1246 units | Worldwide Distribution - Nationwide to AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD., ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, UT, VA, VT, WA, WI, WV, WY and Internationally to Canada... | Open, Classified |  |  | 6156
Z-1698-2016 | 74189 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ellipse ND: YAG Laser applicator for Ellipse Nordlys system |  |  | May 05, 2016 | 2016 | Ellipse A/S | Failure to provide the required instructions specifying a procedure and schedule for calibration of the measurement system. | Radiation Control for Health and Safety Act | Ellipse A/S has evaluated our instructions in the Service Manual and brought those parts into the Operator's Manual as well which we believe is relevant and usable for the end-user of the product. The details in efficiently implementing the distribution of the manual has been evaluated and coordinated with our service provider in the US.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Any future customer notification letters to be sent must comply with the requirements of 21 CFR 1003.21. This includes, but is not limited to, the prescribed packaging and the statement that you will remedy the defect or bring the product into compliance at no charge.For further questions, please call (760) 798-9550. |  | 25 | US - 25 | Nationwide Distribution to the states of : NJ, PA, MI, CO, CA, FL, OK, NY, NH, CT, AZ and MO. | Open, Classified |  |  | 39
Z-1940-2016 | 73998 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | CENTURION FMS (Fluidics Management System) Pack for the CENTURION Vision System and identified Custom Paks |  |  | April 28, 2016 | 2016 | Alcon Research, Ltd. | Alcon is conducting this Voluntary Medical Device Removal for specific lots of its CENTURION FMS packs due to the possible presence of a molding irregularity found within the aspiration luer which may result in a reduction of vacuum within the aspiration line. | Process control | The affected consignees were sent a recall notification letter on 4/28/16. |  | 156712 | 156,712 units | US, Argentina, Australia, Austria, Belarus, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colunbia, Croatia, Cyprus, Czech Republic, Denmark, Finland, Germany, Greece, Hong Kong, Hungary, India, Ireland, Italy, Japan, Korea, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Peru, Philippines, Poland, Portugal, Russia, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uruguay | Terminated | March 31, 2017 | 337 | 156712
Z-1939-2016 | 74170 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Maestro Air Motor |  |  | April 28, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Due to the buildup of grease within the hose and motor, the product may require multiple attempts to start. The grease may also migrate into the attachment end of the motor. There is a negligible possibility that the orange grease may leak from the attachment end of the motor and contaminate the surgical wound resulting in soft tissue inflammation. | Component design/selection | Stryker sent an "Urgent Medical Device Recall Notification" dated April 28, 2016, to all affected customers. The recall notification included a description of the reason for the recall, affected product, customer responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments at (269) 389-4518. |  | 189 | 189 | Domestic only: ALARCAFLGAILINKSKYMDMIMONCNMOHOKSCTNTXVAWAWI | Terminated | March 24, 2017 | 330 | 189
Z-1938-2016 | 74034 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Sterile Sheridan/HVT Cuffed, Murphy Eye, Tracheal Tube, Ref: V5-10314, ID: 7.0 mm, Rx Only, Manufactured by: Teleflex Incorporated. |  |  | May 05, 2016 | 2016 | Teleflex Medical | The labeling of some units indicates that it contains a size 7mm tracheal tube, but may contain a size 7.5mm tracheal tube. | Error in labeling | The firm, Teleflex Medical, sent an "Urgent Medical Device Recall Notification" letter dated May 5, 2016 to its accounts/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue distribution and quarantine product; communicate this recall to any of your customers who have received product; have your customers return to you any affected product with Recall Acknowledgement Form; contact Teleflex Medical Customer Service for more information at 1-866-246-6990; and complete and return Recall Acknowledgment Form via fax to 1-855-419-850, Attn: Customer Service or email to recalls@teleflex.com, even if you and your customers have no affected stock/product. Teleflex will provide you with instructions for the return of product. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 6390 | 6,390 ea. | US Distribution to states of: AL, AZ, CO, FL, GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, NY, NC, OK, PA, RI, TX and UT. | Open, Classified |  |  | 6390
Z-1941-2016 | 74213 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Signature TKA GDE/MDL Set 04-05 |  |  | May 03, 2016 | 2016 | Materialise USA LLC | A patient specific case of Signature Personalized Patient Care System Knee Guides included a femur guide which does not match the patient associated with the case.If the issue is not identified prior to surgery, the personalized patient specific guide would provide an improper fit and would not be able to be used. | Process control | Materialise sent an URGENT: Signature Personalized Patient Care System Knee Guides Notification dated May 3, 2016. It was sent to the affected customer via email with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Contact Materialise customer service for additional information or to find a status of the second patient specific surgical guide Monday through Friday, 8:00AM to 4:30 PM, Eastern Time.734-259-7016 |  | 1 | 1 | One consignee, Indiana. | Completed |  |  | 1
Z-1937-2016 | 73828 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Henry Schein Single Use Syringe Luer-Lock, 10 mL, Intended for use by health care professionals for general purpose fluid aspiration/injection |  |  | March 25, 2016 | 2016 | Becton Dickinson &amp; Company | BD 10 mL Syringe Luer-Lok Tip with BD PrecisionGlide Needle 21G x 1 (0.8 mm x 25 mm) and Henry Schein 10 mL Single Use Syringe Luer-Lock are being recalled due to possible adulteration and non-conforming manufacturing practices. | Error in labeling | Becton Dickinson sent via email and UPS Ground an "Urgent Voluntary Product Recall" Letter dated March 25, 2016 along with a Recall Response Form and packaging instructions to their sole customer. Customers were instructed to immediately review their inventory , quarantine any affected product, and immediately discontinue shipment of the affected product. Customers should complete the Recall Response Card from and fax it to BD at 1-201-847-6990. Return all affected product with the completed Recall Response Card. If product was further distributed please identify your customers and notify them immediately of the product recall. Customers were also instructed to complete the Recall Response Form even if they do not have any affected lots in their inventory. Customers with questions were instructed to call 1-201-847-4267. For questions regarding this recall call 201-847-5612. |  | 811200 | 811,200 units | Nationwide Distribution to NY only | Open, Classified |  |  | 987200
Z-1949-2016 | 74082 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | LCP Dia-Meta Volar Distal Radius Plate 15 H Shaft/LT-Sterile; Indicated for fractures, osteotomies, and non-unions of the radius and other small bones. |  |  | May 02, 2016 | 2016 | Synthes (USA) Products LLC | Product incorrectly packaged. The labels on the outside of the box and on the LCP Dia-Meta Volar Distal Radius Plate inside the box are correct; however, the preprinted box which contains the LCP-Dia-Meta Volar Distal Radius Plate is incorrect. The box is preprinted for a 4.5 mm LCP Condylar Plate - Distal Femur. | Under Investigation by firm | DePuy Synthes sent out an Urgent Notice to customers informing them of a package mixup and the impact it could have on users. Customers were asked to evaluate their inventory and immediately remove the affected lot (9890811) from stock. Customers were asked to call DePuy Synthes at 1800-479-6329 to obtain a Return Authorization Number for return of the affected product, fill out the verification section of the letter and mail with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Any questions, they were said to call 610-719-5450 or to contact their DePuy Synthes Sales Consultant. |  | 6 | 6 units | Distributed to one customer: Colorado | Open, Classified |  |  | 6
Z-1945-2016 | 70054 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | The XN-Series modules (XN-10, XN-20) are quantitative multi-parameter automated hematology analyzers intended for in vitro diagnostic use in screening patient populations found in clinical laboratories. |  |  | December 18, 2014 | 2014 | Sysmex Corporation | There is a potential for elevated hemoglobin (Hgb) and impedance platelet count (PLT-I) values. | Device Design | Sysmex Corporation sent a Product Notification Field Correction letter were to all affected customers of record. The letters included instructions for customers to: 1) Review results with an unexpected increase in Hgb, MCH, MCHC, and/or PLT-I. The increase may result in a low or normal value being artificially elevated to either a normal result or a result above the reference ranges, or a high value being artificially increased. There is negligble impact to the RBC when the issue occurs; 2) The relationship with an elevated Hgb would cause the indices (MCH and MCHC) values to appear to be abnormally elevated and may result in the results being flagged for verification or indicated to be outside of reference ranges; 3) Review any "Turbidity/Hgb Interf?" flag which will automatically trigger at an MCHC of 37.5 g/dL; 4) If Hgb is affected, the PLT-I may be also. Review any PLT-I result when an unexpected increase in MCH/MCHC or "Turbidity/Hgb Interf?" flag occurs; 5) Gross interference is likely to trigger a flag, however there is a potential any interference could increase the Hgb or the PLT-I count from a lower level to an "normal" level. Follow laboratory protocol for confirming unexpected results. Customers were further advised that they will be contacted by a Sysmex Field Representative within 180 days to schedule an appointment to replace the rinse cup on the analyzer(s) and that version 00-16 software will be installed to change the rinsing sequence for the aspiration probe. Customers with questions are requested to e-mail the Sysmex Technical Assistance Center. Urgent questions can be answered by calling the Technical Assistance Center at 888-879-7639 (24 hours a day / 7 days a week). |  | 1137 | 1,137 units (US) 54 units (Canada) 9 inventory | Worldwide Distribution - US to AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WY + District of Columbia + Guam; Internationally to Australia, Bahrain, Bangladesh, Cambodia, Canada, China, Germany, Hong Kong, India, Indonesia, Malaysia, Maldives, Mongolia, Mynamar, Nepal, New Zealand, Pakistan, Philippines, Singapore, South Korea, Sri Lanka, Taiwan, Thailand, and Vietnam.THE CORRECTION IS LIMITED TO XN-10/XN-20 UNITS DISTRIBUTED BY SYSMEX AMERICA, INC., INSTALLED AT END-USER SITES IN US, PUERTO RICO, AND CANADA. | Terminated | July 25, 2016 | 585 | 1137
Z-1948-2016 | 74178 | No | NQN | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | Ariol is an automated scanning microscope and image analysis system. It is intended for in vitro diagnostic use as an aid to the pathologist in the detection, classification, and counting of cells of interest based on particular color, intensity, size, pattern, and shape. |  |  | May 06, 2016 | 2016 | Leica Biosystems Richmond Inc. | An issue during the manufacturing process caused systems to have an improperly activated Windows 7 OS, even though a valid Windows 7 license was associated with each manufactured system. | Nonconforming Material/Component | According to the firm a decision to conduct a Field Corrective Actions of the Ariol Image Analysis and Scanning Systems was done out of an abundance of caution. Leica Biosystems Richmond sent Medical device Field Correction notification letters to the customers of the potentially affected product. The initial written communication was sent to the customers via first class USPS on May 6, 2016. The Medical Device Field Correction letter describes the issues identified with the Ariol Image Analysis and Scanning Systems and the Windows 7 operating system. The letter provided the customer with actions they are required to take which include the following: 1) whether they have a properly activated Windows 7 operating system. 2) If they do not have a properly activated operating system they should to contact Leica Technical Support for additional information and to request a service appointment for correction to their system. For USA customers, they ask that they call 18005374669. International customers are asked to go to http://www.leicabiosystems.com/contactus/contactusonline/ and then choose your country.3) Customers are asked to confirm receipt of the notice by signing and dating the attached Acknowledgement Form and emailing a copy to RichRA@leicamicrosystems.com or faxing it back to us at 18472363800.Service personnel will be sent to the affected customer locations to replace the hard drives of the affected systems. |  | 33 | 33 systems | U.S. consignees: Nationwide; Foreign Consignees: China, Japan, Russia, Italy and England. | Open, Classified |  |  | 33
Z-1947-2016 | 74097 | Yes | NIM |  |  | Cardiovascular | PMA | Cordis PRECISE (R) RX Nitinol Stent System (Biliary) |  |  | May 04, 2016 | 2016 | Cordis Corporation | Inability to deploy the stent or partial stent deployment. | Component design/selection | Cordis sent an Urgent Medical Device Recall (Removal) letter dated May 4, 2016, to all affected customers.Action to be taken by the customer: 1) Immediately check your inventory to confirm whether you have any units from affected lots in your possession. Identify and set aside any units from the affected lot in a manner that ensures the affected product will not be used. Check all storage and usage locations.2) Review, complete, sign and return the enclosed Acknowledgement Form directly to Cordis at the fax number on the form. Note that you will receive multiple letters and Acknowledgement Forms if your facility received affected product at both a legacy Ship-To account number and a Cardinal Health ShipÂż To account number at the same address. In that scenario, please sign and return both forms.3) Return all affected products to the address listed on the form, with reference to Return Authorization Number ( i.e. RA #)#which is listed on the form. Replacement devices will be shipped after the affected product is returned. Consigned Inventory will also be replaced after the product is returned and an analysis of SKU mix needed to properly support the customer is conducted.4) Share this letter with others in your facility who need to be made aware of this recall, including the areas in some facilities where these products are frequently used (see list above) Contact any other facilities that have beenprovided with units of affected lots. Maintain awareness of this notice until all affected product has been returned to Cordis.5) Keep a copy of this notice with any affected product until returned.For questions related to assistance returning product or billing concerns or questions please contact your sales representative or Cordis Customer Service at 1-800-327-7714 Option 1, Monday through Friday from 7:00AMto 8:00 PM EDT. |  | 7500 | 4,400 units in US; 3,100 units to Foreign Countries | NATIONWIDE and FOREIGN CONSIGNEES: Japan, Taiwan, Thailand and South Korea. | Open, Classified |  |  | 7500
Z-1955-2016 | 73536 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00720504022 ACETAB.CUPW/SH,0DEG,22MMX40MM |  |  | April 15, 2016 | 2016 | Zimmer Trabecular Metal Technology, Inc. | Zimmer Biomet is initiating a recall of specific lots of foil pouched, sterile Trabecular Metal implants that were manufactured between April 2011 and October 2012. The foil pouches are part of a dual barrier system to provide sterile integrity. | Unknown/Undetermined by firm | Zimmer issued an Urgent Medical Device Recall Letter dated 4/15/2016. |  | 19884 | 19,884 total | Nationwide and Foreign. | Open, Classified |  |  | 19884
Z-1888-2016 | 74111 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | AVEA Comprehensive ventilator, 17310-xx (xx = 0 - 14) |  |  | May 17, 2016 | 2016 | Carefusion 211 Inc dba Carefusion | CareFusion has identified a potential risk associated with AVEA Ventilator caused by an incorrect F1 fuse on the AVEA Ventilator Transducer Communication Alarm (TCA) Board which may fail/blow, causing loss of power to the User Interface Module (UIM). | Under Investigation by firm | CareFusion sent a Field Safety Notice dated May 17, 2016, to all affected customers. The notice informed customers that CareFusion has identified a potential risk associated with AVEA Ventilator caused by an incorrect F1 fuse on the AVEA Ventilator Transducer Communication Alarm (TCA) Board which may fail/blow, causing loss of power to the User Interface Module (UIM). This malfunction creates a blank AVEA Ventilator UIM and INOP condition during pre-use check or during use with cessation of ventilation to the patient. By design an audible clinical alarm is activated. Customers were informed of the problems identified and the actions to be taken. Customers were not required to return the affected devices or suspend use, customers would receive an copy of the identified affected serial numbers, as well as a response card, customers were requested to promptly return the response card to expedite the correction process and acknowledge receipt of the notification, Customers would be contacted by CareFusion to arrange onsite remediation of the affected devices, in the interim if any AVEA ventilator unit in their facility exhibits a blank AVEA Ventilator UIM and INOP condition during pre-use check or during use with cessation of ventilation, immediately remove the ventilator from service, provide alternate ventilation and contact CareFusion Technical Support per the contact information listed below to report the issue. Customers with questions should call 888-562-6016, for technical support 800-231-2466 or support.vent.us@carefusion.com. |  | 922 | 922 units total (501 units in US) | Worldwide Distribution - US (Nationwide) and the countries of Australia, Hungary, Paraguay, Belgium, Indonesia, Philippines, Bolivia, Ireland, Poland, Bermuda, India, Saudi Arabia, Canada, Japan, Singapore, Chile, Jordon, South Africa, China, Kuwait, Taiwan, Ecuador, Malaysia, Thailand, Egypt, Mexico, Turkey, France, Netherlands, United Kingdom, Guatemala, and Nicaragua. | Open, Classified |  |  | 922
Z-1957-2016 | 74156 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The Red Head 2 Reusable Fat Harvesting Canister, Product # RH-2, Non-Sterile. |  |  | April 21, 2016 | 2016 | Miami Fat Supply, Inc | During an FDA inspection it was found that the products are marketed without a cleared 510k, nor is approved for multiple use. | No Marketing Application | Customers were sent an Urgent: Medical Device Recall/Correction letter, dated 04/19/2016, on 04/21/2016 that served as a notification of the Red Head 2 voluntary recall. Customers were informed that there were previous complaints of compromised product integrity and performance following sterilization. Lids are subject to fracture and fail after steam sterilization. Customers were advised to destroy used product and to ship unused product back to the manufacturer. Customers were to return an acknowledgement form to confirm their notification of the voluntary recall. |  | 443 | 443 units | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, FL, GA, IL, MD, MI, MO, NM, NY, OH, TX, UT, and VA; and, the country of Dominican Republic. | Open, Classified |  |  | 443
Z-1989-2016 | 74045 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Excelsior Disposable Syringe Heparin Lock Flush, USP syringe (5 mL Fill in 10 mL syringe). For flushing of IV catheters and IV tubing only. |  |  | April 28, 2016 | 2016 | Excelsior Medical Corp | Certain lots of Heparin Lock Flush, USP Syringes (5 ml Fill in 10 ml Syringe) failed Anti-Factor IIa assay for potency during testing of retention samples for heparin potency. | Nonconforming Material/Component | Excelsior Medical Inc. sent a recall letter on 4/29/2016 via certified mail. Customers are instructed to discontinue using the product subject to this recall. Customers are requested to complete the customer reply form and fax it (866) 807-3714, so arrangements can be made to return the product to Excelsior. |  | 184500 | 184,500 | US Nationwide distribution. | Terminated | August 08, 2017 | 467 | 184500
Z-1990-2016 | 74207 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker PowerPRO (XT, IT, or TL) Cots |  |  | May 05, 2016 | 2016 | Stryker Medical Division of Stryker Corporation | Complaints of smoke coming from the foot end of the cot. This reportedly occurred when either the battery was placed into the cot or shortly after placing the battery into the cot. In addition to the alleged smoking, one of the reports had also noted that when the user attempted to lower the cot (prior to the alleged smoking) the device went into high speed retract and then a few moments later i | Vendor change control | On 5/5/2016, IMPORTANT MEDICAL DEVICE NOTIFICATION notifications dated 5/3/2016 were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any questions or concerns, please contact Stryker Customer Service (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 770 | 770 | Distributed US (nationwide) and the countries of Australia, Canada, Hong Kong, Spain, Switzerland, and United Kingdom. | Open, Classified |  |  | 770
Z-1958-2016 | 74084 | No | PLF | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | BioSync Anatomic Cotton Wedge, 16 x 4.5 mm.; Used for internal bone fixation for bone fractures, fusions or osteotomies in the ankle and foot. |  |  | April 22, 2016 | 2016 | Arthrex, Inc. | Potential for polystyrene particulate presence. | Material/Component Contamination | On 04/22/2016 Arthrex contacted all distributors that received potentially affected devices and instructed them to place the product in quarantine. Distributors were then instructed to return affected devices to Arthrex. Surgeons who have implanted the device will also be receiving a recall notification. Surgeons, risk managers, and distributors received mailed recall notification identifying the recalled products and batches, the reason for recall, and instructions for returning affected product. Customers are to call Arthrex's Customer Service Line (800-934-4404) to obtain return authorization and further instructions. Customers were also required to complete a "Recall Acknowledgement of Receipt" postal card to return to Arthrex, Inc. in an enclosed envelope. |  | 138 | 138 total products | Distributed in the states of CA, CO, FL, GA, IL, IN, KS, MD, NC, NY, PA, SC, VA, WA, and WI. | Open, Classified |  |  | 138
Z-1964-2016 | 74067 | No | KYB | Intraocular lens guide. | Ophthalmic | Ophthalmic | 510(K) Exempt | STAAR AQ Cartridge-FP, guide and delivery system for Intraocular lenses. |  |  | May 02, 2016 | 2016 | Staar Surgical Co. | The STAAR AQ Cartridge-FP may malfunction and may crack during loading or delivery of the IOL (Intraocular Lenses). | Process control | STAAR sent a Medical Device Recall Letter dated May 11, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. STAAR Surgical is recalling the STAAR AQ Cartridge-FP because it may malfunction and may crack during loading or delivery of the IOL. Customers with questions are instructed to contact their STAAR customer service representative at (800) 352-7842. |  | 2612 | 2,612 units | US | Open, Classified |  |  | 2612
Z-1966-2016 | 73598 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Vaxcel PICC with PASV Catheter Kit under the following labels: 1) 3F, UPN H965454350, Rx ONLY, 2) 4F, UPN M001454550, Rx ONLY &amp; 3) 5F, UPN M001454700, Rx ONLY. |  |  | March 14, 2016 | 2016 | Navilyst Medical, Inc., an AngioDyamics Company | PICC catheters that contain valve(s) manufactured prior to July 15, 2014 lack a manufacturing specification change that reduces the incidence of hemolysis during blood sampling through the PICC catheter. | Device Design | Navilyst Medical distributed Urgent Voluntary Medical Device Recall notification letters on March 14, 2016, and Reply Tracking forms to their customers via courier service. Customers were asked to immediately remove any affected (recalled) product from their inventory (whether in labs, central supply, shipping and receiving or any other location). Segregate this product in a secure location for return to Navilyst Medical. Immediately forward a copy of therecall notice to all sites to which you have distributed affected product. If affected product is located in your institution, please call Navilyst Medical Customer Service at 1-800-772-6446 Monday to Friday, 8 am - 7 pm, EST to obtain a replacement or credit for your returned product. Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return) by fax to 1-800-782-1357.For questions regarding this recall call 518-792-4112. |  | 2716 | 2,716 units | Worldwide Distribution US (Nationwide) Internationally to AU, GR, SE, SA, IN, CA, GB, BR, BE, NL, ES, KR, HK, AE, IT and ES. | Open, Classified |  |  | 17283
Z-1991-2016 | 74229 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Piccolo Composite Tibia and Femoral Nailing Systems - Ball Tip Guide Wire for fixation of fractures that occur in and between the proximal and distal third of the long bones being treated. |  |  | May 10, 2016 | 2016 | CarboFix Orthopedics, Ltd. | Firm received complaints where the Piccolo Composite Ball Tip Guide Wire penetrated the internal package, and thus compromised the product sterility. | Under Investigation by firm | On May 10, 2016, an advisory notice was issued to all distributors that currently may possess (or supplied to hospital) the product. The notice was sent to a representative of the distributor companies via an e-mail (using confirm receipts for delivery and notification when the email is read), together with a letter to hospital(s) for the representative to use in case the distributor supplied the product to the hospital. The letter requested that the product be returned and also requested a sub-recall. |  | 638 | 638 | AL, AZ, CA, CO, DC, FL, GA, IN, KS, KY, MD, MI, NE, NV, NY, OH, OK, TX, WA | Open, Classified |  |  | 638
Z-1965-2016 | 74114 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TS Series Depth Gauge for Integra Ti6 Internal Fixation System 2.0/2.5mm Screws |  |  | May 10, 2016 | 2016 | Integra LifeSciences Corp. | As a result of an internal investigation of a non-conformance it has been determined: 1) Depth Gauges from Lot UP15H2 exceed the overall length specification requirements. 2) Depth Gauges from prior lot numbers manufactured before Integra acquired the product line reflect an incorrect number for the first mm depth marking on the gauge. The number should be 4, it is 6. 3) The space / holder | Process design | The firm, Integra Lifesciences, sent an "Urgent: Voluntary Medical Device Recall" letter dated 5/10/16 to affected consignees/customers via traceable courier service on 5/10/16. The letter describes the product, problem and actions to be taken. The customers were instructed to stop using and distributing product immediately; complete and return the Recall Acknowledgement and Return Form via fax to 1-609-275-9445 or email to FCA2@integralife.com, even if you do not have any of the product.Should you have any questions, please contact Customer Service at 1-866-854-8300. |  | 71 | 71 units | US Distribution to states of: CO, FL, IL, MI, NJ, NY, SC, TX, and VA. | Open, Classified |  |  | 71
Z-2001-2016 | 73169 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE Phototherapy system, Catalogue Part Number: 010066. |  |  | April 11, 2016 | 2016 | Natus Medical Incorporated | The firm is adding the warning statement to the Instruction for Use for the neoBLUE LED Phototherapy System and provide a laminated Quick Guide to affix to each device. | Labeling False and Misleading | The URGENT: FIELD SAFETY NOTICE PRODUCT LABELING CHANGEneoBLUE LED Phototherapy System, dated April 2016, was went to US customers between April 11, 2016 and April 27, 2016. Natus distribution partners outside the US will receive the Letter to Distributors and they are requested to contact their downstream customers to distribute the information. The estimated date of notifying international consignees is May 31, 2016.Enclosed with this letter, customers would receive:1) Quick Guides (2 each) which provide operating steps, information about measuring intensity, and warnings (PN 051693)2) User manual addendum3) Customer reply formNatus requests that this information be distributed to all associated parties within your organization that use or service this product. Please note: this communication is in addition to any prior labeling changes that you may have received from Natus.Customers should submit the Reply Form via fax to 847-949-6031 or email to mailto:Natus_Quality_Programs@natus.com.If there are any questions regarding this letter and addendum, please call Natus Technical Service at 888-496-2887. |  | 13541 | 13541 units total (7840 units in the US and 5701 units outside the US) | Worldwide Distributed - US Nationwide in the US including American Samoa, Guam Puerto Rico and in the following countries:Argentina, Australia, Austria, Bangladesh, Belgium, Bermuda, Bolivia, Brunei Darussalam, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, , Kenya, , Kuwait, Latvia, Lithuania, Malaysia, Moldova Republic of, Morocco, Namibia, Netherlands, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia (Slovak Republic), Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Viet Nam, Virgin Islands (British), and Yemen. | Open, Classified |  |  | 13541
Z-1996-2016 | 71372 | No | GEL |  |  | General &amp; Plastic Surgery | 510(k) | EVARREST Fibrin Sealant Patch, Manufactured by: Omrix Biopharmaceuticals, Ltd, Israel. |  |  | April 11, 2016 | 2016 | Ethicon, Inc. | Ethicon is recalling EVARREST Fibrin Sealant Patch because the characters "0-bad" were printed in the expiration date field on the foil pouch labels of the Fibrin Patch units instead of the valid expiration date (2016-12-28). | Error in labeling | Ethicon sent an "Urgent: Product Recall (Removal)," dated April 14, 2016 via UPS next day mail to their affected customers. The letter identified the affected product and the reason for the recall. The letter also states that product from the affected lot is still within the expiration date, and health care practitioners, who have treated patients with the product should continue to follow patients in the usual manner. Users are to examine their inventory and quarantine the affected product. The affected product should be returned per the instructions outlined in the letter. Ethicon will provide replacement product upon receipt. Customers are to complete and return the Business Reply Form. If customers need assistance returning the product, they can contact Stericycle at 1-866-761-9472. If there are questions regarding the replacement product, customers can contact Ethicon Customer Support Center at 1-877-ETHICON. |  | 228 | 228 units | Nationwide Distribution | Terminated | August 02, 2017 | 478 | 228
Z-1994-2016 | 73203 | No | MOQ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Universal Battery Charger II (UBC II) |  |  | January 15, 2016 | 2016 | The Anspach Effort, Inc. | The device is reported to enter safe mode and is no longer able to charge batteries. | Process design | DePuy Synthes sent an "Urgent Notice- Medical Device Recall (Removal) letter dated January 15, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter informing them of the recall (removal) and asking them to return the UBC II when it display the blue flashing lights that indicate service and to contact DePuy Synthes Customer Support at 800-327-6887 option 2 for return device and arrange to repair or replacement. Any other question please contact the Complaint Handling Unit Manager at 561-494-3673. |  | 1992 | 1,992 | USA Distribution to the states of : AR, AZ, CA, CO, FL, IL, IN, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NY, OK, OR, PA, SC, TN, TX, VA, WA, WI and Hawaii | Open, Classified |  |  | 1992
Z-2002-2016 | 74341 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Small Patient ECC Pack, REF BEQ-TOP 24100, Product Usage: |  |  | May 09, 2016 | 2016 | Datascope Corporation | Custom Tubing Set BEQ-TOP-24100 (Small Patient ECC Pack) was identified as possessing a level of endotoxin that exceeded the regulatory limits. | Under Investigation by firm | Maquet sent via express mail and hand delivered a recall letter and response form dated May 9, 2016 to their sole customer. |  | 17 | 17 units | US in the state of GA | Open, Classified |  |  | 17
Z-1993-2016 | 73473 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK15 monitor/defibrillator with End-Tidal C02 (EtC02) feature installed. |  |  | January 27, 2016 | 2016 | Physio-Control, Inc. | The firm became aware that when using EtC02 in the kPa or % setting and in a situation where the reading is above 9.9 kPa, the display of the LIFEPAK 15 respiratory rate may partially obscure a portion of the leading digit of the EtC02 value. This affects all LIFEPAK 15 with an EtC02 feature installed and configured to the kPa or % setting. | Software design | Physio-Control sent the letter with the title Information Regarding LIFEPAK 15 MONITOR/DEFIBRILLATOR WITH End-tidal CO2 OPTION, dated January 2016, to customers in the United Kingdom (UK) as REQUESTED by the Medicines and Healthcare product Regulatory Agency (MHRA) on January 27, 2016.Physio initiated a voluntary field action on February 23, 2016 and started notifying US consignees on March 22, 2016 and started notifying international consignees on April 18, 2016. The Information Regarding LIFEPAK 15 MONITOR/ DEFIBRILLATOR WITH End-tidal CO2 OPTION letter, dated March 2016, was sent.Customer Letters began mailing in EMEA region, beginning on 06/20/2016 via designated Courier.Customer Letters began mailing in Canada on 09/09/2016 via FedEx.Customer Letters began mailing in Latin America on 09/16/2016 via email.Customer Letters began mailing in APAC region on 09/22/2016 via email.Second Customer letters with Confirmation Sheets began mailing in the US October 2016 via FedEx and USPS Certified Mail.Actions should be taken by users:1. Please forward this information to all of your sites, trainers and users that may have a LIFEPAK 15 monitor/defibrillator.2. If you use your LIFEPAK 15 monitor/defibrillator to display EtC02 values you should also display the EtC02 waveform which includes the range along the right side of the display.3. If you use either kPa or % settings, activate EtC02 alarms to identify EtC02 values outside the normal range.4. If your protocols allow this option, consider changing the EtC02 to measure in the mmHg setting.The firm reported that the revised software for LifePak15 version 4 (v4) and the customer install instructions would be available on LIFENET by April 25, 2016.The revised software for LifePak 15 versions 1 and 2 (v1 and v2) and the customer install instructions are estimated to be available on LIFENET in mid May 2016.Customers who have any questions about this subject, please contac |  | 44798 | 44798 units total (23189 units in US; 21609 units international) | Worldwide Distribution - US (nationwide) including Guam and Puerto Rico and Internationally to Âżland Islands, Algeria, Argentina, Aruba, Australia, Austria, Azerbaijan, Bahrain, Barbados, Belgium, Bosnia and Herzegovina, Botswana, Brazil, Brunei Darussalam, Cambodia, Canada, Cayman Islands, Chile, China, Costa Rica, Croatia, CuraÂżao, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Eritrea, Faroe Islands, Finland, France, French Guiana, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guinea, Guyana, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Islamic Republic of, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of, Kosovo, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, the former Yugoslav Republic of, Malaysia, Maldives, Malta, Martinique, Mauritius, Mayotte, Mexico, Morocco, Mozambique, Namibia, Netherlands, New Caledonia, New Zealand, Norway, Oman, Pakistan, Panama, Papua New Guinea, Poland, Portugal, Qatar, RÂżunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Suriname, Sweden, Switzerland, Taiwan, Province of China, Thailand, Trinidad and Tobago, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Yemen. | Terminated | March 31, 2017 | 429 | 44798
Z-1992-2016 | 73692 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Protura Robotic Couch System Software, versions 1.6.1 (Part Number MT6XSM1.6.1), 1.7.0 (Part Number MT6XSM1.7.0), and 1.7.1 (Part Number MT6XSM1.7.1), CIVCO, Orange City, IA 51041 |  |  | March 31, 2016 | 2016 | Med Tec Inc | When using the recalled product in split shift mode, the non-sub-millimeter positioning has been found to be inconsistent with the setting parameters. | Software design | Civco Medical Solutions sent a "Urgent Field Correction Notice" dated March 31, 2016, to all affected customers. The letter identified the product the problem and the action need to be taken by the customer.The software's other configuration, "All 6 Degrees Protura" does not have the same translational issues. The firm send notification letters on February 23, 2016. In this letter, the Product Manager verified each site's software is configured to this operational mode which operates without the identified problem. In a later and separate notification on March 31, 2016, the firm informed customers that a new software version will be installed to address the problem. No timeframe was given in the letter.For further questions, please call ( 712) 737-8688. |  | 9 | 9 | Worldwide Distribution - US Distribution including the countries of Switzerland, Germany.. | Open, Classified |  |  | 9
Z-2000-2016 | 74239 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BIOGRAPH mCT S(20) - 4R, MATERIAL NUMBER 10528958 |  |  | April 25, 2016 | 2016 | Siemens Medical Solutions USA Inc. | The bracket that is used to lift the cover has bolts that were not welded in potentially causing the cover to fall down during service. | Process design | Siemens sent a Customer Safety Advisory Letter to all affected customers. The letter identified the affected product, problem and actions to be taken. Siemens service organization will be contacting customers to schedule the update as per the attached Update Instruction and will be provided the replacement bracket to resolve the issue. For questions contact your Siemens representative. |  | 1 | 1 unit | Worldwide Distribution - US Nationwide and the countries of China and Malaysia. | Open, Classified |  |  | 4
Z-1995-2016 | 74197 | No | NVP | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Flexichamber Anti-Static Valved Collapsible Holding Chamber, used by patients to administer aerosolized medication from most pressurized Metered Dose Inhalers(pMDIs). |  |  | May 16, 2016 | 2016 | FSC Laboratories, Inc | Product had torn/loose seal around the mouthpiece. | Process control | Letters were sent to consignees on 5/16/2016 informing them of the recall and requesting that they immediately discontinue use of the product. The letter included a reply form to be returned to FSC Laboratories. |  | 1435 | 1435 units | US Nationwide distribution. | Open, Classified |  |  | 1435
Z-2003-2016 | 74206 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Dome Centric component (Reference 01.04227.005) from the Anatomical Shoulder (AS) System 2.0. |  |  | May 17, 2016 | 2016 | Zimmer Gmbh | In some cases it has been difficult or not possible to disassemble the adjusted Dome Centric from the AS Humeral Rasp after initial positioning of the AS Humeral Trial Head which resulted in the whole construct of the Dome Centric with the still assembled AS Humeral Rasp taken out of the humeral canal. This could cause a delay in surgery time or the surgeon could decide to close the wound without finishing the surgery, increasing the infection risk or second use of anesthesia necessary. | Device Design | Zimmer Biomet sent an URGENT MEDICAL DEVICE RECALL  LOT SPECIFIC notification dated May 17, 2016, to all affected consignees via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Consignees were instructed to review the notification and ensure affected personnel were aware of the contents. Assist their Zimmer Bioment sales representative with the quarantine of any affected product from the table above. Complete the Certificate of Acknowledgement Form and return to corporatequality.postmarket@zimmerbiomet.com.If after reviewing this notification consignees have further questions or concerns please call the customer call center at 1- 800-348-2759 between 8:00 am and 8:00 pm EST. |  | 177 | 177 | Nationwide Distribution to AZ, CA, CO, GA, IL, KS, KY, MA, MD, MO, OH, PA, TN, and TX. | Open, Classified |  |  | 177
Z-2004-2016 | 74263 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Musculoskeletal Transplant Foundation, AFT Diverted Tube (3/4/Filled), indicated for treatment of surgically-created osseous defects or osseous defects created from traumatic injury. It can be used in the: extremities, pelvis, and spine. Product Code: 227005. |  |  | April 21, 2016 | 2016 | Musculoskeletal Transplant Foundation, Inc. | AFT Diverted Tubes are being recalled due to a customer complaint involving one unit which had an unsealed Tyvek pouch. The non-sterile foil pouch may introduce contamination into the sterile field. | Packaging process control | Musculoskeletal Transplant Foundation (MTF) sent an Voluntary Field Action Notification Letter dated April 21, 2016, and via telephone to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were notified of the situation with the product and provided instructions on how to handle the product. Questions can be directed to Customer Service Customer at 1-800-433-6576. |  | 114 | 114 units | nationwide | Open, Classified |  |  | 114
Z-2038-2016 | 74179 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | SIOOA I (EP I84) Rabbit Monoclonal Antibody; |  |  | May 10, 2016 | 2016 | Cell Marque Corporation | Cell Marque became aware of some lots of SIOOA I (EP I84) Rabbit Monoclonal Antibody, catalogue numbers 408R-14, 408R-15, and 408R-16, that exhibited an unacceptable IHC test result. | Unknown/Undetermined by firm | Product recall notifications were sent to customers on May 12, 2016. Letters advised that all unused or partially-used vials be destroyed. Customers should contact customer service for replacement and credit. |  | 23 | 23 units | Distributed in the states of MA, CA, and WA and the countries of Australia, Denmark, Turkey, Germany, Russia, Slovenia, Netherlands and Zambia. | Open, Classified |  |  | 23
Z-2039-2016 | 74245 | No | JOY | Automated cell-locating device. | Hematology | Hematology | 510(k) | CytoVision Image Analysis and Capture System, an automated cell-locating device. The Cytovision is a rapid metaphase finder, image acquisition and computer aided chromosome analysis system which assists the operator in viewing chromosomes and looking for cellular anomalies. CytoVision enables a qualified Cytogeneticist to rapidly and accurately analyze the chromosome banding pattern. |  |  | May 02, 2016 | 2016 | Leica Biosystems Richmond Inc. | Systems have an improperly activated Windows 7 OS, even though a valid Windows 7 license was associated with each manufactured system. The system appears to be unlicensed. These systems cannot be re-activated using standard Windows licensing procedures and are requiring checks of each potentially affected system to confirm Windows 7 OS is properly activated. | Software Manufacturing/Software Deployment | Leica Biosystems Richmond sent Medical device Field Correction notification letters to the customers of the potentially affected product. The initial written communication was sent to the customers via first class USPS on May 2, 2016. According to the firm a decision to conduct a Field Corrective Actions of the Ariol Image Analysis and Scanning Systems was done out of an abundance of caution. The Medical Device Field Correction letter describes the issues identified with the CytoVision Image Analysis and Capture System and the Windows 7 operating system. The letter provided the customer with actions they are required to take which include the following: 1) whether they have a properly activated Windows 7 operating system. 2) If they do not have a properly activated operating system they should to contact Leica Technical Support for additional information and to request a service appointment for correction to their system. For USA customers, they ask that they call 18005374669. International customers are asked to go tohttp://www.leicabiosystems.com/contactus/contactusonline/ and then choose your country. 3) Customers are asked to confirm receipt of the notice by signing and dating the attached Acknowledgement Form and emailing a copy to RichRA@leicamicrosystems.com or faxing it back to us at 18472363800. Service personnel will be sent to the affected customer locations to replace the hard drives of the affected systems. |  | 283 | 283 systems | Worldwide distribution. US nationwide; Portugal, Spain, Thailand, Italy, China, Romania, Egypt, India, France, Hungary, Mexico, Germany, Norway, Turkey, Pakistan, Malaysia, UK, Luxembourg, Israel, Canada, Switzerland, Taiwan, South Africa, Saudi Arabia, Korea, Singapore and Japan. | Open, Classified |  |  | 283
Z-2040-2016 | 74113 | No | GMJ | Histoplasma capsulatum serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Histoplasma Yeast Complement Fixation Antigen Dilute Ref# H30150 |  |  | May 09, 2016 | 2016 | Immuno-Mycologics, Inc | The device was found to have bacterial contamination. | Process control | Immuno-Mycologics, Inc sent an Urgent Medical Device Recall letter dated May 9, 2016, to all affected customers. The letter identified the affected product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their stock and destroy affected stock on hand. Immediately complete the attached Acknowledgement and Receipt Form even if they do not have affected stock. Customers who require replacement indicate the number of vials needed in the attached Acknowledgement Form. Customers with questions were instructed to call 1-405-360-4669. |  | 467 | 467 units | Nationwide Distribution to UT, MA, OH, IN, NC, NY, VA, CA, IA and KY. | Terminated | January 10, 2017 | 246 | 467
Z-2041-2016 | 74036 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Healthcare ADVIA Centaur Tnl-Ultra Assay . In-Vitro Diagnostic for quantitative determination of cardiac troponin For in vitro diagnostic use in the quantitative determination of cardiac troponin I (cTnl) in serum, heparinized plasma, or EDTA plasma using the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT/ADVIA Centaur CP systems |  |  | May 02, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | Exhibits a greater than 10% change in results in samples with biotin levels up to 10 ng/mL (41 nmol/L). Instructions for Use (IFU) states that specimens that have up to 10 ng/mL (41 nmol/L) of biotin | Component design/selection | Siemens Healthcare sent an Urgent Medical Device Correction letter dated May 2, 2016, The letter identified the product, the problem and the action to be taken by the customer. Customers were instructed to continue use of the affected product, review the letter with their doctor, complete and return the Field Correction Effectiveness Check attached to the letter. If customers received any complaints of illness or adverse events associated wit the affected product they should immediately contact their local Siemens Customer Care Center or their local Siemens technical support representative. Customers with questions should contact their Siemens Customer Care Center or their local Siemens technical support representative.For questions regarding this recall call 508-668-5000. |  | 16033 | SMN 10317708 = 16434 kits and SMN 10317709 = 16033 kits | Worldwide Distribution - US (Nationwide) and Internationally to Albania AustriaBelgium Bulgaria Burkina Faso Canada Croatia Cyprus Czech Republic Denmark Egypt Estonia Finland France French Polynesia Georgia Germany Great Britain Greece Guadeloupe Hungary Iceland Italy Kazakhstan Latvia Lithuania Luxembourg Macedonia Malta Martinique Netherlands Norway Poland Portugal Romania Russian Federation Serbia Slovakia Slovenia Spain Sweden Switzerland Tunisia Turkey United Arab Emirates and Uzbekistan. | Open, Classified |  |  | 16033
Z-2044-2016 | 74241 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Crit Line in a Clip (CLiC). CliC P/N: CL 10041001 |  |  | May 24, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | Use of the Blood Pressure Graph located on the Crit-Line screen of the 2008T Hemodialysis Machine may result in a frozen screen, possibly causing the blood pump to stop. | Software design | Fresenius Medical Care Renal Therapies Group, LLC issued Recall Letter issued 5/24/16 via certified mail. Clinics actively using the CLiC option have been verbally communicated to stop using the CLiC option at . The other clinics that did not activate the CLiC option have been verbally communicated to not activate the CLiC option at this time. A Software update to all 2008T Hemodialysis machines will be required.A reply form is included as part of the customer communication with instructions to fill out and return via fax or scanned/emailed copy.For technical Inquiries, please contact Technical Services at 1-800-227-2572. Additional medical concerns or questions, please contact the Medical Information and Communication line at 855-616-2309 or website www.fresenius-medinfo.com. |  | 875 | 875 units | Distributed US (nationwide) | Open, Classified |  |  | 875
Z-2042-2016 | 74230 | No | GAW | Nonabsorbable polypropylene surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sharpoint Polypropylene Blue Monofilament 6"/15cm 9-0 USP, Product No. J2558N |  |  | May 20, 2016 | 2016 | SSC- Surgical Specialties Corporation | Surgical Specialties is recalling Sharpoint Polyproplyene Blue Monofilament 6"/15cm 9-0 USP because it may contain Black Nylon 9-0 suture. | Process control | A customer notification letter dated 5/20/16 was sent to all customers to inform them that Surgical Specialties is recalling Sharpoint Polyproplyene Blue Monofilament 6"/15cm 9-0 USP because it may contain Black Nylon 9-0 suture. Customers are instructed to complete the attached customer acknowledgement form and return via fax to Surgical Specialties' Quality Assurance department at (610) 404-4010 or return via email to kknappenberger@surgicalspecialties.com. Customers with any questions are instructed to contact Kelly Knappenberger at (484) 220-2402 or Kris Liberacki at (262) 994-8083. |  | 36 | 36 boxes (432 units) | TX, ME, MA, MI, OH, KS, MN, KS, TN, GA | Open, Classified |  |  | 36
Z-2005-2016 | 74192 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph mCT-X 4R, MATERIAL NUMBER 10248670 |  |  | April 22, 2016 | 2016 | Siemens Medical Solutions USA Inc. | Siemens Medical Solutions, Molecular Imaging has become aware of a potential for unexpected SUV values if a non-Siemens phantom is used for calibration. | Nonconforming Material/Component | Siemens sent a Customer Safety Advisory Letter on April 22, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your local Siemens representative at the contact numbers provided in the letter. |  | 83 | 83 units | Worldwide Distribution - US Nationwide | Terminated | May 08, 2017 | 381 | 1075
Z-2045-2016 | 74191 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Myriad handpiece |  |  | February 24, 2014 | 2014 | Nico Corp. | This recall has been initiated due to a defective component within the handpiece which willcause the handpiece to fail. The failure will result in the inability to cut tissue with thehandpiece, rendering it non-functional. Should this failure occur, a replacement Myriadhandpiece will be required. If a replacement handpiece is not available, alternate instrumentation for tissue removal will be necessary. | Device Design | NICO sent an Urgent: NICO Myriad Handpiece Recall Notification on February 24, 2014, to all affected customers requesting return of the affected product. The recall notification included a description of the reason for the recall, affected product, customer responsibilities, and instructions for responding to the formal recall notification. If you have any questions regarding this matter, please call (317) 660-7118, ext. 104. |  | 63 | 63 units | US Distribution to the states of : OR, IL, IA, NC, TX, CA, GA, OH, WI, AR, MI, MD, IN. | Completed |  |  | 63
Z-2043-2016 | 74372 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Fujifilm Synapse PACS software version 4.4.000, |  |  | May 10, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | Synapse cannot display image files, DICOM SR files, and/or Annotation files. The "Image Not Loaded" message displays instead | Software design | The firm, FujiFilm, sent an "MEDICAL DEVICE CORRECTION" letter dated May 10, 2016 via Federal Express to its customers. The letter describes the product, problem and actions to be taken.The customers were instructed to: Extend the Ad hoc deletion retention period or discontinue use of Ad Hoc Deletion, and complete and submit the attached Customer Response Form to Fujifilm via fax to (203)-251-7863. Affected systems will be corrected via on-site visits by FMSU Customer Service Engineers upgrading to Version 4.4.100.Questions or concerns, contact the FUJI FILM Customer Support Center at 1-888-FUJIMED Option #1 . |  | 27 | 27 units | US Distribution to states of: CA, FL, KS, NE, NY, and PA. | Open, Classified |  |  | 27
Z-2047-2016 | 74015 | No | MCF | Manual toothbrush. | Dental | Dental | 510(K) Exempt | Germ Terminator, Toothbrush Steam Sterilization for Oral Health. Model GT100. The Germ Terminator uses steam heat to sanitize two manual toothbrushes or two toothbrush heads. Water is poured into a reservoir in the device's housing. The steam cycle is followed by a drying cycle. |  |  | April 05, 2016 | 2016 | Ceg Enterprises Llc | CEG Enterprises, LLC is recalling Germ Terminator Toothbrush Sanitizer GT100 due to the following issues: UL label believed to be not genuine; discoloration of power cords (could indicate some degradation); reports of excessive heat. | Unknown/Undetermined by firm | CEG Enterprises sent an Urgent Germ Terminator Toothbrush Sanitizer GT100 Recall letter dated April 19, 2016, via UPS and email to all affected customers. The letter identified the affected product, the reason for recall, stated to quarantine product, and to return all remaining product. The letter also instructed that Tuesday Morning retail stores post the recall notification letter on the retail shelves where the product was offered for sale, as well as on their website. Consumers are encouraged to return the device to a Tuesday Morning retail store. Questions can be directed to 612-337-1821 |  | 100548 | 100,548 units | Nationwide Distribution through Tuesday Morning retail stores. | Open, Classified |  |  | 100548
Z-2046-2016 | 74208 | tray, impression, preformed | EHY | Preformed impression tray. | Dental | Dental | 510(K) Exempt | 6534-08N InterActiveÂż Implant Open-Tray Transfer Narrow |  |  | May 19, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | The incorrect sized product, InterActive Implant Open-Tray Transfer Wide, 6534-08W, was packaged instead of the InterActive Implant Open-tray Transfer Narrow, 6534-08N, causing a possible issue with the impression. This discrepancy may lead to the possibility of an oversized crown fabrication. | Packaging | The firm sent out customer notificaiton letters on 05/19/16 via FedEx.Customers were asked to reivew the table to determine if they have any of the affected product in inventory. Customers were being informed that the proper product was being sent with the notificaiton. Customers were asked to complete and return the the Acknowledgement and Recall Return Form within 48 hours. If customers are an authorized Implant Direct Sybron Manufacturing distributor, the firm requests that they identify those customers that may have been shipped the affected product lot and contact these customers to inform them of this issue within fortyeight (48) hours of receipt of this notification in order to provide the customers with the correct tool. |  | 48 | 48 | Worldwide Distribution - U.S. Nationwide in the states of OR, NE, VA, FL, MO, WA, TX and the countries of CZ, GB, NL, MF, DK, IT, IE, DE | Open, Classified |  |  | 48
Z-2048-2016 | 74276 | Yes | NIP |  |  | Cardiovascular | PMA | EV3 EverFiex" Self-expanding Peripheral Stent System, Sterilized using Ethylene Oxide, Rx Only, PRB35-08-040-080. |  |  | May 23, 2016 | 2016 | Ev3, Inc. | Medtronic is issuing a medical device recall to return product potentially associated with an outer boxmislabel issue for two Protege EverFiex" Self-expanding Peripheral Stent Systems, model PRB35-08-040-080, from lot A164580. The size on the outer box is labeled as 8x40x80mm while potentially containing a6x40x120mm device as indicated by both the inner pouch and device labels. This potential issue is limited totwo EverFiex stent system devices from lot A164580 that were repackaged by Medtronic and distributed.This does not affect any other Medtronic product or implantable devices. | Labeling mix-ups | Medtronic sent an "Urgent Medical Device Recall" letter dated May 2016, to all affected customers. The letter described the problem and the product involved in the recall. Consignees were instructed to identify, quarantine and return unused product to Medtronic. Consignees were requested to complete the Customer Confirmation Certificate and email it to RS.CFQFCAÂżMedtronic.com or fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. Customers with questions were instructed to call 651-367-0612.For question regarding this recall call 763-398-7000. |  | 2 | 2 | Nationwide Distribution to LA and WI. | Open, Classified |  |  | 2
Z-2049-2016 | 74226 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PrismaSate BGK2/0 Dialysis Solution for Continuous Renal Replacement Therapy |  |  | May 23, 2016 | 2016 | Baxter Healthcare Corp. | Presence of leaks near top of the PrismaSate bags | Other | Baxter Healthcare Corporation sent an Urgent Product Recall communication to affected customers via U.S.P.S. first class mail on 05/23/2016. |  | 10054 | 10,054 units | Distributed to Columbia and the following US states: AZ, CA, CO, DE, FL, GA, IA, ID, IL, IN, KY, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NV, NY, OH, OR, PA, TN, TX, UT, VA, WA, WI | Terminated | March 28, 2017 | 309 | 10054
Z-2056-2016 | 63990 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq Two Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids |  |  | February 04, 2013 | 2013 | Hospira Inc. | Frequent Proximal Occlusion Alarms (POAs) occurring on Symbiq devices with list numbers 16026 and 16027. When there are false alarms, the pump will continue to alarm even though there is no occlusion present. False POAs could result in a delay/interruption in therapy, and require the clinician to reset the alarm multiple times or to replace the pump. | Software design | Hospira sent letters on May 2, 2013, to customers informing them of the discontinuation of sales, service, and support for the Symbiq pump were sent via traceable mailing to all affected customers. For questions regarding this recall call 224-212-4892. |  | 15285 | 15,285 | Worldwide Distribution - USA (nationwide) | Completed |  |  | 35596
Z-2051-2016 | 74445 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | ViperWire Advance Peripheral Guide Wire, Sterilized with Ethylene Oxide, Rx Only. Part number 72023-02, Model Number VPR-GW-17. Five (5) guide wire pouches in one (1) guide wire box. For use with CSI Peripheral Orbital Atherectomy Devices (OAD), including Stealth 360 and Diamondback 360. Cardiovascular. |  |  | June 13, 2016 | 2016 | Cardiovascular Systems Inc | The pouch label was missing the use by date (UBD) of 2018-04. | Process control | Consignees were sent on 6/13/16 a Cardiovascular Systems, Inc. "Urgent Correction" letter date June 10, 2016. The letter described the problem and the affected product. Advised consignees to set affected product aside for their sales representative which will be applying the correct pouch label to the affected product. Requested consignees to fill and sign the Customer Acknowledgement Form that the sales representative will be presenting to them. For questions contact Quality at Cardiovascular Systems, Inc. 1-651-259-2068 1225 Old Highway 8 NW Saint Paul, MN 55112. |  | 190 | 190 units (38 - 5 pack shelf cartons) | US Nationwide distribution. | Open, Classified |  |  | 190
Z-2050-2016 | 73608 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Bigger-Better-Bladder With 3/8" ID tubing, ITEM BBB38 |  |  | March 03, 2016 | 2016 | Circulatory Technology Inc | The seal between the balloon and the housing may leak. | Device Design | Circulatory Technology Inc. sent an Urgent Product Recall notices dated March 3, 2016, to their customers via email. Circulatory Technology Inc. has identified this action as a voluntary recall, however they did not withdraw the product from the market. Instead, Circulatory Technology Inc. has advised customers to check the integrity of the seal between both the inlet and outlet tube and the housing as per the Instructions For Use. In addition, for this lot, Circulatory Technology Inc. recommends that you bend the inlet as well as the outlet tube away from the housing before conducting the leak test. There should be no visible separation between the tubing and the housing. Customers with questions can call 516-624-2424. |  | 72 | 72 units | US Distribution to the states of: CA, IN, MI, NY, PA, TX &amp; VA | Open, Classified |  |  | 72
Z-2054-2016 | 73395 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Sodium Chloride Inhalation Solution, USP 7%, 60 x 4 mL Sterile Unit-Dose Vials, Rx Only, NDC 0487-9007-60. |  |  | November 16, 2015 | 2015 | Nephron Pharmaceuticals Corp. | Lack of Assurance of Sterility: Internal dye ingress test failure to assure the integrity of the vial. | Process control | Nephron Pharmaceuticals Corporation sent an Urgent Drug Recall letter to all affected customers on November 16, 2015. The letter identified the product, the problem, and the action to be taken by the customer. NPC requested customers remove the affected lots from their inventory, including store shelves. Customers with questions should call 1-844-374-6841 |  | 13692 | 13,692 cartons | Worldwide Distribution - US including AL, AR, AZ, CA, CO, CT, FL, GA, IL, KY, LA, MA, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, VT, WA, WI, WV and Internationally to Canada. | Terminated | January 09, 2017 | 420 | 24900
Z-2052-2016 | 74378 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | June 17, 2016 | 2016 | Elekta, Inc. | It is possible that that the selected Plan-of-the Day could be accidently changed by scrolling the mouse wheel up and down and inadvertently selecting another plan. | Software design | Elekta sent an Important Field Safety Notice to all customers on June 17, 2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. A software patch will be issued for Mosaiq version 2.62 in order to the correct the defect. For further questions please call (404) 993-5886. |  | 69 | 69 | US Distribution including Puerto Rico and to the states of : AK, AZ, CA, FL, GA, IL, IN, LA, MD, MI, MO, NJ, NY, OH, OR, PA, TN, TX, UT, VA, WA, and to the countries of : Canada, Denmark, Finland, Germany, Israel, Japan, Netherlands, Norway, Sweden, United Kingdom | Open, Classified |  |  | 69
Z-2072-2016 | 73894 | No | CAL | Flexible laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Intubation Scope |  |  | April 06, 2016 | 2016 | Pentax Medical Company | Pentax has become aware that operational/cleaning accessories and therapeutic devices (e.g., clips, stents, balloons, etc.) can become lodged in an endoscope's instrument channel. | Device Design | Pentax sent an Urgent Field Correction letter dated April 6, 2016 to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were reminded to carefully check that all accessories are intact, that no parts have fallen off and become lodged within the endoscope's instrument/suction channel and to ensure that any therapeutic devices passed through the instrument channel are accounted for after use. If the instrument/suction channel does become blocked or clogged due to the accumulation of debris, an accessory that cannot be removed, or other cause, do NOT attempt to correct the blockage or continue to use the endoscope. Customers were instructed to contact their local PENTAX Medical Service facility to have the endoscope repaired. Customers with questions were instructed to call 800-431-5880, ext 2064. |  | 1220 | 1,220 units | Worldwide Distribution - US (Nation Wide) including Puerto Rico and Internationally to Denmark, Guam | Open, Classified |  |  | 31293
Z-2057-2016 | 73584 | container, specimen, non-sterile | NNI | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | dynarex Specimen Containers, Screw-On, Leak-Resistant Lid, Tamper-Evident Seal, Sterile, Single Use Only |  |  | March 03, 2016 | 2016 | Dynarex Corporation | Exterior carton is mis-labeled as Sterile. Non-sterile, bulk specimen containers were packed in incorrect exterior carton. | Labeling False and Misleading | On March 2, 2016, dynarex Corporate Headquarters distributed Voluntary Device Recall notices and Recall Verification Forms to their customers via email. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to examine their stock immediately to determine if they have any of the lots listed on hand. If so, discontinue distributing and quarantine the product immediately. Promptly contact the Dynarex Credit &amp; Return representative Avi Celnick (845-365-8200 Ext 6644) to arrange for product return. Customers account will be credited for all returned products. All customers are advised to complete the Voluntary Recall Verification Form and returm immediately via fax or e-mail to Vijay Sachdev (Fax 845-365-8238; v.sachdev@dynarex.com). |  | 9975 | 399 cases/25 sleeves/20 cups &amp; lids per sleeve | US Nationwide in the states of AL, AR, FL, GA, MA, NC, NY, NJ, NH, PA, OH, TX &amp; VA; International: Canada | Open, Classified |  |  | 9975
Z-2084-2016 | 74224 | No | HXI | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | K-wire, Smooth Double-pointed 0.8 x 90mm, a component of the METASURG Capture Screw System |  |  | May 20, 2016 | 2016 | Integra LifeSciences Corp. | The affected K-wires do not meet specification requirements. Specifically, a) K-wires were made of 316LVM stainless steel instead of 304 stainless steel. b) K-wires have passivated surface finish instead of electro-polish surface finish. c) K-wires have incorrect length and diameter. | Nonconforming Material/Component | Integra sent an Urgent Voluntary Medical Device Recall letter dated May 20, 2016, to all affected consignees regarding the recall via traceable courier service. The letter identified the product, the problem, and the action to be taken by the customer. Customers who have supplied affected product should forward a copy of the Recall Notice to them. Customers should review their inventory for affected product, complete the form as to whether they have any affected product or not, return the form by email or fax as indicated on the form. Customers with questions should contact Customer Service at 1-800-654-2873. For questions regarding this recall call 609-936-6822. |  | 182 | 182 units | Nationwide Distribution to OK, CA, TX, WA, GA, NY, MS, NY, PA | Open, Classified |  |  | 182
Z-2064-2016 | 74142 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris System PC Unit Model 8015 with software versions 9.17 and 9.19; |  |  | May 12, 2016 | 2016 | CareFusion 303, Inc. | A patient weight can be populated incorrectly under certain conditions when using the RESTORE feature to restore infusions running on the Alaris LVP module model 8100 and the Alaris Syringe module model 8110. | Software design | An Urgent Medical Device Recall Letter was sent to customers on 5/31/16 to inform them that BD had identified an issue with the Alaris System PC unit model 8015 with software versions 9.17 and 9.19. Firm has received reports that a patient weight can be populated incorrectly under certain conditions when using the RESTORE feature to restore infusions running on the Alaris LVP module model 8100 and the Alaris Syringe module model 8110. The letter informs the customers of the potential risk and the required action for users. Customers are informed that BD will contact them within 60 days to initiate the scheduling process for the software upgrade installation. |  | 69693 | 69,693 units | Distributed US (nationwide) and in Canada. | Open, Classified |  |  | 69693
Z-2081-2016 | 74282 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | CENTURION Vision System |  |  | May 24, 2016 | 2016 | Alcon Research, Ltd. | The grounding pin within the Power Entry Switch on certain systems was susceptible to detachment when the AC power cable is removed. | Nonconforming Material/Component | Alcon has already notified the affected customer verbally on May 24, 2016, and via written correspondence on June 1, 2016. The device was corrected on May 26, 2016, and the Power Entry Switch that was removed from the customers CENTURION Vision System did not exhibit a grounding pin susceptible to detachment. For further questions, please call (817) 293-0450. |  | 1 | 1 unit | US Distribution to Illinois | Terminated | April 13, 2017 | 324 | 1
Z-2063-2016 | 74234 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Liquid stable 2 part Homocysteine Reagent, for in vitro quantitative determination of total homocysteine in human serum and plasma - H7575-57 kit includes one Rl and one R2 bottle and one bottle of calibrator level 1 and one bottle of calibrator level 2. |  |  | April 03, 2015 | 2015 | Medtest Holdings, Inc. | Deterioration in performance of the Liquid Stable (LS) 2-Part HOMOCYSTEINE REAGENT resulting in under recovery.The manufacturer has indicated that the HOMOCYSTEINE REAGENT lot listed above shows deterioration in assay performance manifesting itself by under recovery of Homocysteine in Controls and Patient Samples.The cause of this deterioration has been attributed to a loss of reductant activity in Reagent 1.Testing performed indicated that the affected lot is at the limit of acceptable clinical performance. This observation was consistent across the reportable range of the assay. For this reason, do not continue to use affected reagent lot. Test data generated to date will be acceptable where validity control specifiCations have been met. | Unknown/Undetermined by firm | On 4/3/2015, Urgent Field Safety Notice notifications were sent to the affected customers via email with instructions for identifying and destroying the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. |  | 13 | 13 | MI, MN, NJ, IL, NY.US Only. | Completed |  |  | 13
Z-2085-2016 | 74275 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Saw Sagittal Catalog Number: ES 300; Saw Sagittal Refurbished Catalog Number: ES300-R; Saw Saggital, Finger Catalog Number: ES310; Saw Saggital, Finger Refurbished Catalog Number: ES310-R; Used to dissect bone in surgical applications. |  |  | May 13, 2016 | 2016 | Medtronic Sofamor Danek Usa, Inc - Dallas Distribution | Medtronic Powered Surgical Solutions testing of the Sagittal Saw Rocker Arm Seal identified a potential for leakage of fluid into the motor during use. | Device Design | Recall notification letters were sent to consignees on 05/13/2016. |  | 368 | 368 devices | Worldwide Distribution-Including the states of AL, AR, AZ, CA, CT, FL, GA, IA, ID, IL, IN, MA, MD, NC, NH, NY, SC, TN, TX, and UT. | Terminated | June 13, 2017 | 396 | 368
Z-2062-2016 | 66372 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | *** 1) Product labeled in part: 19683-28; LIFESHIELD; MICRODRIP SOLUSET WITH 1.2 MICRON FILTER 150ML BURETTE PLUMSET, CONVERTIBLE PIN, CLAVE SECONDARY PORT, OPTION-LOK, NON-DEHP; NOMINAL LENGTH: 118 IN; PRIMING VOLUME: 14 ML; |  |  | September 09, 2013 | 2013 | Hospira Inc. | Intravenous fluid may leak at the CLAVE secondary port of Hospira PlumSet IV administration sets. | Device Design | Hospira sent an URGENT MEDICAL DEVICE CORECTION letters dated September 9, 2013 to customers. The letter identified the affected product, problem and actions to be taken. the letter included additional instructions on the proper handling/use of the CLAVE devices. The letters request that all incidents of separation or breakage of CLAVE from the secondary port of the cassette be reported to Hospira by calling 1-800-441-410 (Monday - Friday, 8:00 AM - 5:00 PM CST) or by e-mail at ProductComplaintPP@hospira.com. Customers are further requested to complete and return a response form included wih the URGENT MEDICAL DEVICE CORRECTION letter and to notify any customers of this advisory. Hospira will also be supplementing the instructions for CLAVE use as an interim measure until a new method for bonding the CLAVE to the secondary port of the Plum cassettes can be devised and implemented. |  | 3920932 | 3,920,932 units in total | US (Nationwide) Distribution | Completed |  |  | 10212296
Z-2099-2016 | 74408 | No | NYQ |  |  | Immunology | PMA | Dako HER2 CISH pharmDx Kit, product code: SK 109 |  |  | June 10, 2016 | 2016 | Dako North America Inc. | An incorrect HER2 probe concentration was used when producing the probe component (Vial 3) in this kit lot. This has resulted in the HER2 probe concentration being too low. | Process control | Dako sent a recall notification dated June 10, 2016 to all customers. The letter identified the affected product, problem, the investigation findings, the affected results, and the actions to be taken. Customers are instructed to complete and return the enclosed device recall form to Dako QA Vigilance by Dako.dkvigilance@agilent.com. Customers with questions regarding the notification, are instructed to contact their sales representative. |  | 31 | 31 kits | Worldwide Distribution - US Nationwide in the states of: PA, CA, AL, NC and countries of Austria, Bosnia and Herzegovina, Belgium, Bulgaria, Brazil, Canada, Switzerland, Germany, Spain, France, Greece, Italy, Macedonia,Qatar, Romania, Serbia, Sweden, and Turkey. | Open, Classified |  |  | 31
Z-2100-2016 | 74198 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK CR Plus and/or LIFEPAK EXPRESS are Automatic External Defibrillator(s) and are non-wearable. |  |  | May 25, 2016 | 2016 | Physio-Control, Inc. | LIFEPAK CR Plus Automated External Defibrillators (AED) or LIFEPAK EXPRESS AED may fail to initiate voice prompts when the ON/OFF button is pressed and the lid is opened due to an internal component (reed switch) that can intermittently become fixed in the closed position. A defibrillator in this condition will fail to deliver a shock, with the potential result that therapy is not delivered and a patient is not resuscitated. | Component change control | The firm, Physio Control, sent an "URGENT MEDICAL DEVICE CORRECTION-ACTION REQUIRED"- LifePak CR Plus AE and LifePak Express AED letter and the Confirmation sheet, dated May 2016, to US customers on 5/25/16. Physio will notify international consignees during the week of 6/13/16.The letter described the product, problem and the actions to be taken. The customers were instructed to-URGENTLY bring this letter to the immediate attention of the person(s) responsible for maintaining/monitoring your Automated External Defibrillators (AEDs); follow the instructions on the Confirmation Sheet that outline specific actions to take for the serial numbers listed, and if you have a routine check process, please continue this process. If you have not established a routine check process, please refer to section 5 of the Operating Instructions for recommended actions. In addition, customers should follow the instructions on the Confirmation Sheet and submit this form to Physio by: -Fax to: 1-866-448-9567-Email to: rsrecall@physio-control.com -Mail to: Physio-Control, Inc. P.O. Box 970006, Dept. 15N Redmond, WA, 98073-9706Physio-Control will contact customers with LIFEPAK CR Plus and LIFEPAK EXPRESS AEDs that contain the affected reed switch assembly. A device correction including provision of loaner devices and replacement of the reed switch component will be arranged for all affected devices. Customers who have any questions regarding this notification, please call Physio-Control at 1-800-442-1142, 6:00 a.m. to 4:00 p.m. (Pacific), Monday  Friday. |  | 25178 | 25178 units (10,418 in US and 14,760 outside US) | Worldwide Distribution-US (nationwide) including Guam and Puerto Rico and countries of: Argentina, Austria, Australia, Bahamas, Bangladesh, Belgium, Brazil, Brunei Darussalami, Canada, Chile, China, Croatia, Cyprus. Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Lithuania, Luxembourg, Malta, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan (Province of China), Turkey, United Arab Emirates, United Kingdom, New Caledonia, and New Zealand,. | Open, Classified |  |  | 25178
Z-2094-2016 | 74271 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LEGION(TM) HK FEMORAL ASSEMBLY, SIZE 5 LEFT, A/P 62 MM, M/L 70 MM, REF 71421375 |  |  | May 26, 2016 | 2016 | Smith &amp; Nephew, Inc. | The surface inside of the femoral component, where cement is applied for adhesion, is below specification. | Process control | The firm sent an Urgent Product Recall letter dated May 26, 2016 to customers, via email and FedEx Overnight delivery . The letter identified the affected product, problem and actions to be taken. Customers were instructed to 1. Inspect inventory and locate any unused devices from the listed product and batch numbers, and quarantine them immediately.2. Complete the Inventory Return Certification Form on the following page, indicating the quantities that need to be returned. If you do not have product to return, place an x in the column No Product to Return. Include your contact information in the spaces provided at the bottom of the form.3. Contact Smith &amp; Nephews Field Actions Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to return the completed Inventory Return Certification Form and obtain a return authorization (RA) number if you have product to return.4. Return any affected product to the address listed on the Inventory Return Certification Form. Indicate the RA number on your return shipment. |  | 10 | 10 units | Worldwide Distribution and US Nationwide including Puerto Rico and the countries of Australia, Austria, Belgium, Canada, France, Germany, Italy, Netherlands, South Africa, Sweden, and UK. | Open, Classified |  |  | 171
Z-2089-2016 | 73629 | No | EBI | Denture relining, repairing, or rebasing resin. | Dental | Dental | 510(k) | IMPAK Liquid Acrylic Resin sold under the following labels: |  |  | March 02, 2016 | 2016 | CMP Industries, Llc | Possible presence of foreign material (rust). | Material/Component Contamination | CMP Industries/Nobilium sent an Urgent IMPAK Recall letter dated March 2, 2016, to all affected customers via certified mail with return receipt in an attempt to verify if customers were in possession of the affected product. Customers were advised to immediately examine inventory and to quarantine product subject to the recall. In addition, if customers further distributed the product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of the notice. If customers have any of the product in their posession subject to recall, please contact the firm by calling 518-434-3147, ext 144, Monday - Friday 8-4:30 pm, EST. |  | 2596 | 236 pieces (25 pieces/11 oz, 211 pieces/1 Quart) | Worldwide Distribution - US including CA, MI, LA, OH, NY, IN, WA; Internationally: to CA | Terminated | March 16, 2017 | 379 | 2596
Z-2088-2016 | 74390 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza, Picture archiving and communication system (PACS), Model Numbers - 10863171, 10863172, 10863173, |  |  | May 09, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software error in previous software versions in which two references for the same image may exist in the database. | Software design | Siemens sent an Customer Information letter dated May 10, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter with software instruction to customers informing them about an upcoming software VB10B under(SY003/16/P) update change and issue that the new software version will address. For further questions, please call (610) 448-6478. |  | 74 | 74 systems | US Distribution to the states of : AL, CA, CT, FL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NE, NY, OH, OR, PA, TX, WI | Terminated | March 20, 2017 | 315 | 74
Z-2087-2016 | 74280 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hettich Centrifuge Modules: |  |  | May 11, 2016 | 2016 | Siemens Healthcare Diagnostics Inc. | Premature bucket failures | Under Investigation by firm | Siemens mailed to customers an Urgent Medical Device Correction (UMDC#LAI16-02.A.US) Urgent Field Safety Notification letter informing customers of the problem with recommendations to follow the existing maintenance instructions in the product operator manuals as well as a new requirement to spin at cycle speeds not exceeding 3500 rpm. |  | 487 | 487 | Nationwide: AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, PA, SC, SD, TN, TX, VA, WA, WI, WV. | Open, Classified |  |  | 487
Z-2097-2016 | 74061 | No | EBI | Denture relining, repairing, or rebasing resin. | Dental | Dental | 510(k) | IMPAK Liquid Acrylic Resin sold under the following labels: 1) IMPAK ELASTIC ACRYLIC RESIN LIQUID FOR MOUTHGUARDS &amp; BRUXISM SPLINTS, Contents: 1 Quart (946 ML), REF 3306, Rx ONLY; 2) IMPAK ELASTIC ACRYLIC RESIN LIQUID FOR MOUTHGUARDS &amp; BRUXISM SPLINTS, Contents: 11 oz (325 ML), REF 3748, Rx ONLY, 3) BRUX-EZE AN ELASTIC ACRYLIC RESIN FOR THE CONSTRUCTION OF DENTAL TREATMENT APPLIANCES, 1 Quart Liquid, PN 7450; Indicated for relining a denture surface. |  |  | April 28, 2016 | 2016 | CMP Industries, Llc | Possible presence of foreign material (rust). | Material/Component Contamination | On April 28, 2016, Nobilium, Division of CMP Industries LLC distributed Urgent Recall notices and Response Forms to their customers via courier service. Customers were advised to immediately examine inventory and to quarantine product subject to the recall. In addition, if customers further distributed the product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of the notice. If customers have any of the product in their possession subject to recall, please contact the firm by calling John Nicpon at 518-434-3147, ext. 144. Hours of operation are Monday - Friday 8-4:30 pm, EST. Return completed recall response forms to nicpon@nobilium.com. |  | 77 | 77 units Total | Domestic - MI, MN, NJ, NM, NY &amp; PA; International - Canada | Open, Classified |  |  | 77
Z-2086-2016 | 74173 | No | HSN | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Cadence Size 5 Left Tibial Tray Implant, designed to treat ankle arthritis through replacement of the ankle joint with a prosthesis, thereby reducing pain, restoring alignment, and allowing for movement at the replaced joint. Catalog Number: 10207205; Lot Number #QJ0248 |  |  | April 29, 2016 | 2016 | Integra LifeSciences Corp. | As a result of an internal review of all labeling for the Cadence System, it was determined that the outer package label for the Size 5 Left Tibial Tray Implant reflects the material composition incorrectly. The labeling incorrectly reflects the material as Cobalt Chromium instead of Titanium. | Labeling mix-ups | On 4/29/16, US consignees were notified by electronic mail delivered by traceable method. |  | 12 | 12 units | NC and PA | Open, Classified |  |  | 12
Z-2105-2016 | 73918 | tips and pads, cane, crutch and walker | INP | Cane, crutch, and walker tips and pads. | Physical Medicine | Physical Medicine | 510(K) Exempt | Universal Quick Adjust under the following labels: 1) Aluminum crutch, adult, 1 pair per case, 2) Aluminum crutch, tall adult, 1 pair per case |  |  | February 17, 2016 | 2016 | Medical Depot Inc. | Tip crutch failure involving the bottom of the metal portion which comes in contact with ground | Nonconforming Material/Component | On February 17, 2016, drive DeVilbiss Healthcare issued a preliminary notice their customers informing them of the situation via courier service. Secondly, On February 25, 2016, drive DeVilbiss Healthcare distributed another notice informing customers that the potential issue was confirmed to a limited number of lots. Customers were advised to quarantine all products in inventory as replacement products will be provided. The company notified customers that they will commence shipping orders for the product on Monday, February 29, 2016. Customers were also advised that a substantial inventory of replacement crutch tips for distribution will begin on or about March 7, 2016. On March 11, 2016, drive DeVilbiss distributed a third notice to their customers via courier service advising them of customer reimbursement for time and labor costs associated with this action. Drive Devilbiss Healthcare will reimburse based on the total number of tips requested to be shipped. Lastly, on May 12, 2016, drive DeVilbiss distributed a fourth Urgent Medical Device notice and response form to their customers via email. drive DeVilbiss advised customers to: 1) quarantine any stock that has not been inspected already and compare serial numbers of stocks to list provided, 2) any product received after 3/7/16 has been fully inspected and has an approved crutch tip that is free of defect, 3) if products were further distributed, contact those customers who received product between 7/2015 to 12/2015 and determine if items are still being used, 4) If the customer is still using this item, contact your sales representative so the return and replacement of this product can be processed and 5) complete and return the response form vial email. Customers with questions can contact by telephone 516-998-4600 ext. 4422. |  | 123125 | 123,125 pairs | Nationwide Distribution | Open, Classified |  |  | 123125
Z-2114-2016 | 74511 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core 128, model #728323; Intended to produce cross- sectional images of the body. |  |  | June 21, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Supplier design change contributed to a thinner plastic which was used as part of the cover on the devices. The thinner plastic cover may crack causing injury to patients, operators, by-standers and service personnel. | Nonconforming Material/Component | Philips is informing the consignee of this issue through:Customer Information Letter (CIL) 72800656 .International consignees will be contacted by Philips Healthcare Global Market representatives, who will distribute the CIL. A Field Service Engineer (FSE) will contact the customers to schedule a time to conduct the appropriate correction through Field Change Order (FCO) 72800656 once it is available. This FCO will address the above identified issue.A response form sent to the customer via certified mail. This package will also contain a return receipt. |  | 14 | 14 units | Distributed in China, Denmark, Germany, Hungary, Italy, Philippines, Spain, Sweden, Switzerland, Thailand, and United Kingdom. | Open, Classified |  |  | 29
Z-2110-2016 | 74507 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Big Bore Radiology tomography X-ray system scanner, model number 728244. |  |  | June 14, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Socket Head Cap Screws (SHCS) used in the Z-Axis Adjustment Plate were found to be made of stainless steel rather than the specified alloy steel. | Nonconforming Material/Component | Philips Medical Systems (Cleveland), Inc. sent a Customer Information letter dated June 14, 2016, to all affected customers via certified mail. The letter identified the product, the problem, and the action to be taken by the customer. No action is required of the customer for this issue. Systems that may have been manufactured with these stainless steel fasteners still meet the safety specifications required for the adjustment cradle and may continue to be used clinically. For further information or support customers were instructed to call 1-800-722-9377.For questions regarding this recall call 404-483-2015. |  | 89 | 89 | Worldwide distribution to Belgium, China, Egypt, Germany, India, Korea, Republic of, Myanmar, Nigeria, Peru, Russian Federation, Rwanda, Singapore, Spain, Sweden, Taiwan, Thailand, and United Kingdom | Open, Classified |  |  | 136
Z-2107-2016 | 74422 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | SYNGO Breast Care, visualization and image enhancement tools to aid radiologist in the review of digital Mammography images and tomosynthesis datasets. |  |  | June 06, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software issues. Siemens is voluntarily initiating a recall after they became aware of the following system behavior: 1) At times, the view and laterality marker is overlaid by patient demographic information, thus becoming difficult to read.2) The facility name and address are not shown per default (in the current software versions this information is displayed only when the reader chooses an appropriate image text display mode). | Other | Siemens mailed customers a letter to inform them about improvements that were made to the SYNGO Breast Care reading software to correct certain behaviors that were being seen during use. Customers were asked to upgrade their systems to this new version. |  | 29 | 29 units | Distributed to: CA,NY,TX,OH,CO,NY,TX,CA,NE,NJ,TX,FL,IL,TX,MO,CA,PA,FL,NJ,MO,ND | Open, Classified |  |  | 29
Z-2108-2016 | 74393 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Free T3 Assay, Ref No. A13422 |  |  | June 08, 2016 | 2016 | Beckman Coulter Inc. | Demonstrate an upward shift in patient results. | Device Design | Beckman Coulter sent an " Urgent Medical Device Recall" letter dated June 8, 2016, to all affected customers. Beckman Coulter has determined through customer feedback and internal testing that the Access Free T3 Reagent lots listed above demonstrate an upward shift in patient results. The preliminary results indicate this upward shift may be related to a June 2015 formulation design change that was introduced to improve the Access Free T3 open reagent pack stability. The upward shift in patient test results, therefore, is expected to be maintained for all future lots. The letter informs the customers of the impact and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with any questions are instructed to contact Customer Technical Support Center at http:www.beckmancoulter.com, call (800) 854-3633 in US and Canada. Customers outside US and Canada, are instructed to contact local Beckman Coulter representative. |  | 205744 | 205,744 kits total (13,358 kits in US) | Worldwide Distribution - US Distribution and to the countries of : Algeria, Australia, Austria, Bangladesh, Belgium, Bolivia, Bosnia and Herzegovina, Botswana, Brazil, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Djibouti, Ecuador, Egypt, El Salvador, Estonia, France, Germany, Ghana, Gibraltar, Greece, Guam, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Lithuania, Macao, Malaysia, Mayotte, Mexico, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Rwanda, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, Zambia, Zimbabwe | Terminated | April 18, 2017 | 314 | 205744
Z-2120-2016 | 74273 | No | GZO | Galvanic skin response measurement device. | Neurology | Neurology | 510(K) Exempt | The ZYTO Tower is the input device to program the software with various virtual items. |  |  | November 23, 2015 | 2015 | ZYTO Technologies, Inc. | ZYTO Technologies Inc. Announces a Voluntary Recall of the ZYTO Tower and ZYTO Select and Elite Software Due to Claims Exceeding the 510(k) Clearance. | Labeling False and Misleading | Consignees were notified on 11/23/2015 via letter and email of an update to the ZYTO Select and Elite Software that would disable all tower functionality.. |  | 1252 | 1252 total | Nationwide. Canadian and other foreign consignees. No VA/govt/military. | Open, Classified |  |  | 1252
Z-2118-2016 | 74427 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core 128-Computed Tomography X-ray system |  |  | April 01, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field that certain Ingenuity Core128 systems running software version 3.5.4 exhibited intermittent swirl-like ring artifacts that may appear on reconstructed images. A patient rescan may be required if the images cannot be used for interpretation due to the swirl-like artifact. | Software design (manufacturing process) | Philips Medical Systems (Cleveland) Inc. sent an Urgent Field Safety Notice letter to affected customers , via certified letter with receipts returned for tracking purposes. The letter identified the affected product, problem and actions to be taken. Philips Healthcare is implementing software update version 3.5.5 to correct the problem. For questions contact your local Philips representative or local Philips Healthcare office. |  | 19 | 19 | US Nationwide Distribution in the states of CA, FL, ID, KS, MI, NY,OH,TX | Open, Classified |  |  | 262
Z-2119-2016 | 74363 | No | JMO | Iron-binding capacity test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products Calibrator Kit 29, REF/Product Code 680 2344, used in conjunction with VITROS Chemistry Products dTIBC Reagent Generation 30 (GEN 30), REF 680 2001, and Assay Data Disk (ADD), REF/Product Code 6801876 (Unique Device No. 10758750001576), Data Release Versions (DRVs) 5870 through 5883 inclusive; IVD --- Ortho-Clinical Diagnostics |  |  | May 24, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics started receiving customer complaints for biased results when using VITROS dTIBC Reagent GEN 30 product. Ortho's investigation confirmed that incorrect calibration mathematics were assigned to the Calibrator Kit 29 Lots 2915 and 2995 supporting dTIBC Reagent GEN 30 product. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated May 24, 2016, via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (Ortho Direct, Drop-Ship and US Federal Government) to inform them of the issue and advised them to discontinue using the affected product until Assay Data Diskette (ADD) Data Realease Version (DRV) 5884 or above was installed on their VITROS System. Foreign affiliates were informed of the issue via e-mail on 5/24/2016 and instructed to notify their consignees of the issue and required actions.Customers with questions were advised to contact Ortho's Customer Technical Services at 1-800-421-3311.For questions regarding this recall call 908-218-8776. |  | 842 | US: 1517; Foreign: 842 | Worldwide Distribution - US (Nationwide, Puerto Rico) and Internationally to Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, Poland, and Venezuela. | Open, Classified |  |  | 842
Z-2124-2016 | 74391 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Magnetic Resonance System G-scan Brio (Part Number /Model No.101002000) |  |  | June 06, 2016 | 2016 | Esaote S.p.A. | The device is being recalled because dynamic loading forces from larger patients can fracture the extremity portion of the patient table. Additionally, a weight limit is added to the patient seat for weight-bearing examinations. | Other | Esaote initiated a voluntary recall of the G-scan Brio Magnetic Resonance System on 06/06/2016 with letters sent via certified, due to a potential weakness on the extremity areas of the patient table that could lead to a stress fracture of the table. The firm has instructed customers to refer to the Medical Device Safety Notification provided with the recall notice, date and sign this letter in the space provided on the last page and send it back via email to quality@esaoteusa.com. If customers have questions regarding the recall, please call 1-317-813-6030. For additional information, please contact us by phone during normal business hours, Monday-Friday from 8:00am  5:00pm Eastern Standard Time, via email at quality@esaoteusa.com, or at any time by visiting our website. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDAs MedWatch Adverse Event Reporting program either online at www.fda.gov/medwatch/report.htm, by regular mail using the postage-paid, pre-addressed Form FDA3500 available at www.fda.gov/MedWatchgetforms.htm or by fax to 1-800-FDA-0178. |  | 40 | 11 USA; 29 OUS | Domestic: FL, TX, OH, MN, NV; Foreign: No customers in Canada;No VA/DOD. | Terminated | May 30, 2017 | 358 | 40
Z-1908-2016 | 73461 | No | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | ABG-HM-1 Hummi Micro Draw Blood Transfer Device; Used for blood transfer and collection from peripheral arterial line catheters. |  |  | March 11, 2015 | 2015 | Hummingbird Med | Hummingbird Med Devices, Inc. is recalling the Hummi Micro Draw because of insertion difficulties. | Process design | The firm, Hummingbird Med Devices, Inc., sent a "Safety Alert" on March 3, 2015 and Addendum on March 11, 2015, to all affected customers. The notification sent on March 3, 2015, instructed customers that: 1) the Hummi Micro device is not recommended to be used with the Baxter Interlink device with the Hummi Micro draw blood transfer device as the Interlink T Connector septum may be damaged.2) the use of Nexus Medical Micro-T Split Septum T Connector NMT8046 with all uses of the Hummi Micro Draw Blood Transfer device. The connector has been tested and validated for use in combination with the Ameritus Medical's Hummi Micro-Draw Blood Transfer Device.3) Distributors of the Hummi Micro Draw Device contact and inform all Hummi Micro-Draw Blood Transfer Device users to recommend change to the Nexus Split Septum Micro-T Connector NMT8046 for all uses with the Hummi Micro Draw device. Inservice materials and training on proper Hummi Micro Draw device blunt cannula insertion into the connector's septum will be provided.4) Distributors complete the attached form indicating hospitals and date contacted, name, and contact information of responsible person and resulting outcome. and5) Scan and email the attached form to: jorgehaider@humigbirdmed.com or fax to: 1-949-583-2775. On March 11, 2015, the firm sent out a second email to their distributor to inform customers again that only the split septum T-Connector recommended for use with the Hummi Micro Draw is the Nexus Micro T split septum T-Connector. The Nexus Micro T-Connector is designed to provide improved insertion force and eliminate potential leaking while providing a microbial barrier when used with the Hummi Micro Draw device.The Hummi Micro device manufactured in 2015 now has improved lubrication on the blunt tube which makes insertion into smaller septum opening of the Mexud Micro T-Connector smooth and easier to accomplish. The firm states in the letter that the distributor will offer |  | 62400 | 62,400 | US Distribution in the state of CA. | Terminated | February 07, 2017 | 699 | 62400
Z-2126-2016 | 73198 | No | MOQ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Battery Power Line II Battery Oscillator, DePuy Synthes Power Tools, The Anspach Effort Inc. |  |  | November 25, 2015 | 2015 | The Anspach Effort, Inc. | Has the incorrect operating directions on the mode switch ring, i.e. its is labeled "FWD, LOCK, REV" instead of "ON, LOCK, ON". | Labeling design | DePuy Synthes sent an Urgent Notice Recall Notification dated November 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to review their inventory and immediately remove any affected devices, complete the verification section of the letter and return it to the Anspach Effort Inc., and return any affected product to their DePuy Synthes Sales Consultant. Any questions regarding this recall / removal contact the Complain Handling Unit Manager at (561) 494-3673 or their DePuy Synthes Sales Consultant.For further questions regarding this, please call (561) 627-1080. |  | 20 | 20 hand pieces | US Distribution to the state of : VA and Internationally to:Thailand and UK | Open, Classified |  |  | 20
Z-2125-2016 | 74344 | Yes | NVZ |  |  | Cardiovascular | PMA | Medtronic, MyCareLink" Patient Monitor, Model 24950, Rx Only. |  |  | May 26, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | Recently, a new software version was automatically sent to a subset of Model 24950 MyCareLink Monitors. After release, Medtronic identified an issue with the software that prevents implanted device data from being available to clinicians on the CareLink" Network. While the transmission appears successful to the patient, the transmitted data, including CareAlerts, are not visible to the clinic. | Software design | Medtronic sent an Urgent Field Safety Notice dated May 2016 to all affected customers. The letter described the issue and the product involved in the recall. The letter also included the "Customer Actions" for consignees to follow. Requested consignees to complete and return the Customer Confirmation Certificate to |  | 262 | 262 | Worldwide Distribution to Austria, Belgium, China , Czech Republic , Denmark , Finland , France, Germany, Hong Kong, India , Ireland , Israel, Italy , Malaysia, Netherlands, Norway, Poland, Portugal , Saudi Arabia, Slovakia, South Africa, Spain, Sweden, Switzerland, UAE, and United Kingdom. | Open, Classified |  |  | 262
Z-2123-2016 | 74300 | No | EOX | Esophagoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Covidien Cytosponge Cell Collection Device. |  |  | June 06, 2016 | 2016 | Covidien, LLC | Medtronic is recalling all lots of the Covidien Cytosponge Cell Collection device after two reports of the device detaching from the removal string during withdrawal from the patient's esophagus. | Process control | The firm, Medtronic, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 6, 2016 to affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine and discontinue use; return affected product to Medtronic/Covidien, Quality Department, 540 Oakmead Parkway, Sunnyvale CA 94085; complete and return the RECALLED PRODUCT RETURN FORM by email to Customer Service feedback.customerservice@Covidien or fax to (800)-895-6140 if product purchased directly from Medtronic; or by email to Quality Compliance MITGFCA@Covidien.com or fax to (203) 492-7719 if Product purchased through distributor or you have "no inventory"; and to promptly alert any other person or facilities if they have further distributed the product. All affected lots must be returned.If you have any questions or concerns, please do not hesitate to contact yourMedtronic representative or Customer Service at (800) 882-5878. |  | 620 | 620 total units (260 US; 360 OUS) | Worldwide Distribution-US Distribution to TN, NC, IL, CA, MN, OR and NY; and countries of: United Kingdom and Iran. | Open, Classified |  |  | 620
Z-2128-2016 | 74317 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Novation Press-Fit Splined RDD, Extended Offset, HA Coated, 12/14 Taper Femoral Stems, Size 18. |  |  | May 26, 2016 | 2016 | Exactech, Inc. | Labeled with the incorrect Global Trade Item Number (GTIN). The GTIN on the label identifies the devices as another. | Under Investigation by firm | Exatech sent an Important Notification letter dated May 16, 2016, to all affected customers. The letter notified customers of the product's incorrect Global Trade Item Number. Customers were instructed to cease distribution/use of the product, extend the information to accounts possessing the recalled product, identify and quarantine any devices in inventory, and to complete and return the Recall Inventory Response Form acknowledging that they received the recall notification. Customers were instructed to call 800-392-2832 with any questions. For questions regarding this recall call 352-377-1140. |  | 14 | 14 | Nationwide Distribuiton to OK, ME, NY, &amp; Hawaii. | Open, Classified |  |  | 22
Z-2133-2016 | 74329 | Yes | CBP | Nonrebreathing valve. | Anesthesiology | Anesthesiology | 510(k) | 5195GM, Cadillac First Aid Kit, #23459698, brand General Motors. No UPC. |  |  | May 16, 2016 | 2016 | Lifeline First Aid LLC | Cadillac First Aid Kit is recalled because it contains the CPR Shield with One Way Valve and Barrier Filter (mask) and this mask is cleared only for prescription use and is not cleared for over-the-counter (OTC) use. | Other | The firm sent the PRODUCT RECALL NOTICE- LIFELINE FIRST AID KITS WHICH INCLUDE CPR MASK letter, dated 5/11/2016, to their consignees.CPR shields will be replaced without charge and acceptable Lifeline First Aid kits will be re-worked and returned using our standard RGA process. In addition, the firm sent the PRODUCT RECALL NOTICE- CADILLAC FIRST AID KIT letter, dated May 31, 2016, to General Motors. General Motors will be sending out a dealership bulletin to inform their customers of this issue. Once confirmed, replacement masks will be sent to the requested dealerships and made available to be ordered (at no cost) as needed.Please complete and fax the attached Customer Response Form to Lifeline at (503)783-3721.Consignees with questions should call Lifeline First Aid, LLC at (503)783-3660. |  | 2975 | 2975 kits | Distributed in AZ, CA, CO, FL, GA, IL, KY, MD, MN, MO, MT, NE, NJ, NY, OK, OR, PA, SC, TN, UT, VA, WA, WA, and WI. Also distributed in New Zealand. | Open, Classified |  |  | 103905311
Z-2135-2016 | 74454 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MEXL-3010/A5- Vantage Titan 3T Magnetic Resonance Imaging System |  |  | April 14, 2016 | 2016 | Toshiba American Medical Systems Inc | A potential problem associated to the gradientcoil (ASGC) in the Titan 3T MRI System. | Device Design | An Urgent Medical Device Notification letter was sent to customers to alert them to a potential problem associated to the gradient coil (ASGC) in the Titan 3T MRI System. The letter informs the customers of the problem identify and the corrective action. Customers are informed that their Toshiba service representative will contact them for an appointment to install the modifiedsoftware on their system when it becomes available. Customers are instructed to complete the attached customer response form and fax it to the toll-free number shown at the top of the form; sent via email to TAMS-RegulatoryAffairs@toshiba.com. Customers with any questions are instructed to contact Paul Biggins at (800) 421-1968 or to contact their local Toshibarepresentative at (800) 521-1968. |  | 33 | 33 units | US Nationwide Distribution | Terminated | March 30, 2017 | 350 | 33
Z-2131-2016 | 74442 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR750w |  |  | June 08, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with the patient bore heating on the Discovery MR750w.A small area on the bore surface, which is normally accessible to the patient, can get warm enough to cause a serious burn in the event the patient touches the bore and proper padding is not used.No injury has been reported to GE as a result of this issue. | Nonconforming Material/Component | Consignees were sent on 6/8/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 60901 dated June 8, 2016. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 471 | 471 (184 US; 287 OUS) | US: AK, AZ, CA, COM, CT , DC, FL, GA, GU, HI, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, NV, NH, NJ, NY, NC, OH,OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WY.OUS:AlgeriaArgentinaAustraliaAustriaBahrainBelarusBelgiumBermudaBrazilBulgariaCanadaChinaEgyptFranceGermanyHungaryIndiaIndonesiaIrelandItalyJapanKazakhstanKorea (Republic of)KuwaitLebanonMexicoNetherlandsNew ZealandPolandQatarRomaniaRussiaRussian FederationSaudi ArabiaSingaporeSpainSwedenSwitzerlandTaiwanThailandTurkeyUnited Arab EmiratesUnited KingdomUruguay | Open, Classified |  |  | 471
Z-2134-2016 | 74515 | No | FET | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Image1 SPIES X-LINK Camera Control Unit (CCU), Product Model No. TC301, when used with Image1 SPIES D1 camera head, Product Model No. TH116 &amp; TH115, and DC1 Video Mediastinoscope |  |  | June 20, 2016 | 2016 | Karl Storz Endoscopy | Potential electrical interference issue when the Image1 SPIES X-LINK and Image1 SPIES D1 camera head are used in conjunction with the DCI Video Mediastinocope and an electro-surgery generator (ESG). | Device Design | KARL STORZ sent an Customer Notification letter dated June 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. KARL STORZ has become aware become aware of a potential electrical interference issue when the lmage1 SPIES X-LINK and lmage1 SPIES D1 camera head are used in conjunction with the DCI Video Mediastinoscope and an electro-surgery generator (ESG). Customers are instructed to complete the attached acknowledgement Form and to return it to KARL STORZ. Customers with questions are instructed to contact Technical Support Dept. at (800) 421-0837, press 2, and then press 1. |  | 619 | 619 units total (90 units in US) | Worldwide Distribution - US Distribution to the states of : IL, NM, TX, MD, TN, PA, FL, KY, MO, NY, MN, DC, MA, WI, CT, MI, ID, CO, GA, NJ and AR., and to the countries of : Argentina, Australia, Belgium, Bulgaria, China, Germany, Finland, Great Britain, India, Italy and Japan. | Terminated | January 05, 2017 | 199 | 619
Z-2136-2016 | 74362 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROW, ARROWG+ARD AND ARROWG+ARD BLUE PLUS PRESSURE INJECTABLE CENTRAL VENOUS CATHETER. |  |  | June 02, 2016 | 2016 | Arrow International Inc | Arrow is recalling these products due to the possibility that although the kit is labeled as containing a 20 cm catheter, a 16 cm catheter may be packaged in the kit. | Process control | Arrow mailed an Urgent Medical Device Recall Notification letter to customers requesting they immediately discontinue distribution and quarantine any products with the affected lot numbers. A Recall Acknowledgement Form was included with the notification letter for customers to complete and email to recalls @teleflex.com indicating the amount of product on hand. |  | 1170 | 1, 170 units | US Nationwide Distribution: in the states of MA, CA, IL, TN, IL | Open, Classified |  |  | 1170
Z-2083-2016 | 74158 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic Navigation O-arm 02 surgical x-ray imaging system |  |  | March 17, 2016 | 2016 | Medtronic Navigation, Inc.-Littleton | Review of on-system electronic product labeling and Instr uctions for Use in response to an internal management concern uncovered gaps in compliance with the requirements for information to be provided to users in 21CFR part 1 020.30(h). | Radiation Control for Health and Safety Act | Medtronic planned action to bring device into Compliance:1. You will contact customers and provide an Errata sheet to them updating the missing and incorrect information.2. Medtronic field service personnel will perform radiation output measurements via validated procedures to attain compliance with 21CRF part 1020.32(k)(6).3. You will update the systems software-based instructions for use by a software update which includes the information supplied by the Errata sheet.For questions call (800)-595-9709. |  | 63 | US - 63 | Nationwide Distribution | Open, Classified |  |  | 63
Z-2130-2016 | 74311 | No | NVE | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The System 83 Plus, Endoscope Washer/Disinfector |  |  | May 06, 2016 | 2016 | Custom Ultrasonics, Inc. | To warn customers that Custom Ultrasonics, Inc. System 83 Plus AERs should not be used for cleaning and/or high-level disinfection of duodenoscopes until further notice. | Process control | Custom Ultrasonics, Inc. sent an Urgent Medical Device Recall letter dated May 6, 2016, to all affected customers. The letter advised customers not to use the System 83 Plus AERs for cleaning and/or high-level disinfection of duodenoscopes until further notice. In addition, they sent a warning label for customers to affix in a prominent location on the System 83 Plus AER's on or before June 3, 2016 stating the following: "This device is not indicated for reprocessing of duodenoscopes. Do not reprocess any duodenoscopes in this device until further notice. For alternative reprocessing options, please contact the duodenoscope manufacturer."For questions regarding this recall call 215-364-1477. |  | 2345 | 2,345 units | Nationwide Distribution to IN, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD., ME, MI, MN, MO,MS, MT, NC, NE, NH, NJ, NM, NV, NY OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT,WA, WI, WV, and WY. | Terminated | April 18, 2017 | 347 | 2345
Z-2129-2016 | 74464 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | PoleStar system (N20 / N30) Warning Lamps Control Box, Warning Lamps Control Box Cables, Thermoflex Chiller |  |  | June 20, 2016 | 2016 | Medtronic Navigation | Medtronic Navigation is recalling certain components of the PoleStar system (N20 / N30) because of misplaced ground pins which may cause electrical shock to the user. | Process change control | Medtronic sent a Field Safety Notice letter dated August 1, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that Medtronic will perform an inspection on site to verify and correct the grounding connection of the PoleStar System. For questions contact Medtronic Navigation Technical Services at 1-800-595-9709. |  | 52 | 52 | Worldwide Distribution - US Nationwide. | Terminated | June 21, 2017 | 366 | 52
Z-2082-2016 | 74153 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic 0-arm Imaging mobile x-ray system. |  |  | March 17, 2016 | 2016 | Medtronic Navigation, Inc.-Littleton | It has been discovered that replacement x-ray generator and system motion batteries could become damaged during shipment, resulting in unacceptable risk of failure. | Radiation Control for Health and Safety Act | Medtronic's planned action to bring device into Compliance. 1. You will provide an Errata sheet to customers updating the missing and incorrect information.2. Field Service Engineers (FSEs) will perform radiation output measurements via validated procedures to attain compliance with21CFR part 1020.32(k)(6).3. You will inspect and test X-ray generator and system motion batteries for damage, in accordance with the plan submitted to CDRH. Damaged batteries will be replaced.4. You will update the systems software-based instructions for use by a software update which includes the information supplied by the Errata sheet.For further questions, please call (800) 595-9709. |  | 581 | US - 581 | Nationwide Distribution | Open, Classified |  |  | 581
Z-2138-2016 | 74418 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Ti Button Plate 7 Holes, Fastener, Fixation, nondegradable, Soft Tissue |  |  | May 31, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is voluntarily initiating a recall due to affected devices being distributed and or labeled with MR Safety information that does not meet the current ASTM standard. | Under Investigation by firm | Synthes mailed to customers (Hospital Personnel and Sales Consultants) an URGENT FIELD SAFETY NOTIFICATION letter informing them the insert they received, which contained incorrect MR Safety Information, does not meet the current ASTM standards. Customers were asked to review the package insert provided with the notification letter, discard the outdated package insert, complete the attached verification Section (page 3 of this notification) and send a copy of the completed Verification Section by Fax: 844-782-5568 or scan/email: Synthes7923@stericycle.com. For further questions call (610) 719-6500. |  | 6255 | 6,255 units | USA Distribution to the states of : AK,AL,AR,AZ,CA,CH,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NC,ND,NE,NH,NJ,NM,NV,NY,OH,OK,ON,OR,PA,RI,SC,SD,TN,TX,UT,VA,VT,WA,WI, &amp; WV | Terminated | April 25, 2017 | 329 | 6255
Z-2137-2016 | 74078 | No | IRT | Powered heating pad. | Physical Medicine | Physical Medicine | 510(K) Exempt | HeaterBaum Foot &amp; Ankle Heating Pad, Heating Soft Compression Therapeutic Boot with Precise Temperature Control Thermostat, 110 V with AC Adapter. |  |  | January 19, 2016 | 2016 | Ergoactives, Llc | Products are marketed without a cleared 510k. | No Marketing Application | Customers were notified by phone on 01/19/2016. Two follow-up notifications were sent to customers via e-mail. Customers were asked to immediately examine their inventory for the affected product; and, to notify and identify any customers they may have further distributed to. Ergoactives LLC requests that all Heaterbaum products be sent back to their headquarters: 20725 NE 16th Avenue, Suite A-26, Miami, FL 33179, along with the order number and note of where the product was purchased. The enclosed response form should be completed and returned as well. If customers have any questions, they can call Ergoactives LLC at (305)359-5031. Customers will receive a full refund for the return of their Heaterbaum products. |  | 278 | 278 units | Worldwide Distribution -- US, including the states of AL, AR, CA, FL, GA, KS, MA, MD, ME, MI, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, PA, TN, TX, UT, VA, and WA; and, the countries of Ecuador &amp; Venezuela. | Open, Classified |  |  | 278
Z-2143-2016 | 74354 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation System, Product Code engen, Software Versions enGen Select 3.2 and Above and enGen Custom 2.0 and Above; IVD |  |  | May 16, 2016 | 2016 | Ortho-Clinical Diagnostics | If the user selects the Comment or Interpretation option (in the enGEN .gsb file), and the numerical results fall between Negative and Reactive, the IM incorrectly displays the interpretive text as Borderline and sends the incorrect text to the LIS. The IFU states the interpretation term for these numerical result values is Retest?.Âż The difference in terminology exists for US Markets only. | Software design | Ortho Clinical Diagnostics sent an "IMPORTANT PRODUCT CORRECTION NOTIFICATION" Customer Letter dated May 16, 2016, to all affected customers in the US with enGen Systems via FedEx or US Priority Mail (for Post Office Box addresses) or by e-mail to subscribers of ORTHO PLUS e-Communications to advise them of the issue and the required actions . Required Actions: 1. Use the numerical test results to assess actions to be taken for these assays, per the IFU.2. Post this notification by your enGen System or with your user documentation.3. Complete and return the Confirmation of Receipt form by 25 May 2016. with the Instructions for Use (IFU) directions).If you have any questions regarding this notification, please call the Ortho Technical Solutions Center at any time at 1-800-421-3311. |  | 30 | US: 30 units | Nationwide | Open, Classified |  |  | 30
Z-2145-2016 | 74346 | No | LOJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | VITROS Immunodiagnostic Products AFP Reagent Pack, REF/Product Code 680 0784, Unique Device Identifier No. 10758750000876, IVD; --- Ortho-Clinical Diagnostics Percoed, Bridgend, UK --- Please note: the VITROS Immunodiagnostic Products AFP Calibrators, Unique Device Identifier No. 10758750000913, REF 680 0863, were included in the communication because they are lot-linked to the Reagent Packs and cannot be used with any other lot of reagent. There is no known failure with the Calibrators. |  |  | May 09, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | Ortho Clinical Diagnostics (OCD) confirmed that specific lots of VITROS Immunodiagnostic Products AFP Reagent Pack product exhibit an increase in the reaction signal over time potentially resulting in the inability to calibrate and use the product. | Nonconforming Material/Component | The firm, Ortho Clinical Diagnostics, sent an "IMPORTANT PRODUCT CORRECTION NOTIFICATION" Customer Letter (Ref. CL2016-098) dated 5/9/2016, via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received the affected product. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue distributing and discard the affected lots; complete and return the Confirmation of Receipt form by May 18, 2016 via Fax to 1.888.557.3759 or 1.585.453.4110 or scan to PDF and email to: ConfirmationAdmin@its.jnj.com, and forward this notification if you have distributed these systems outside of your facility.If you have any questions, contact Customer Technical Services at 1-800-421-3311. |  | 13577 | US: Lot 6471 - 282, Lot 6480 - 344 | US Nationwide Distribution | Open, Classified |  |  | 13577
Z-1887-2016 | 73636 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System: HeartWare Battery, Lithium Ion, 14.8 V, Rx only. Catalog 1650 (US) and 1650, 1650-DE and A00035 (OUS) |  |  | January 07, 2016 | 2016 | HeartWare Inc | Heartware Ventricular Assist System Battery Cell experiencing premature depletion. | Under Investigation by firm | The firm, HeartWare, sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 7, 2016, to all their consignees informing them that of the HeartWareÂż HVAD Battery recall (device exchange). The letter described the product, problem and actions to be taken. The customers were instructed to:1. Promptly complete the attached Acknowledgment Form and return it to HeartWare no later than 30 days from the date of this letter via email to HeartWare7786@stericycle.com or provide verbal Acknowledgement by calling Customer Service line , 1(844)491-7871; or fax to +1 (844)721-3044..2. Immediately identify and quarantine affected batteries in hospital inventory, and apply a RECALLED PRODUCT  DO NOT USE sticker (provided with this notice) as illustrated below. Replace quarantined inventory AC Adapters and batteries by requesting a Return Goods Authorization (RGA) through a qualified representative or directly through HeartWare Customer Service (via telephone 1-877-367-4823 or email cs@heartware.com). A qualified representative will update the Controller software for affected Controller units in inventory.3. Identify, collect and quarantine affected product in patient possession. Arrange for current patients to bring their Batteries to a clinic appointment as soon as possible, (at least within the next three months), with urgency dependent upon individual patient circumstances. During their regularly scheduled appointment, a qualified representative will assist with quarantining and replacing affected product. Apply a RECALLED PRODUCT  DO NOT USE sticker as illustrated on affected batteries. Replace affected patient product using hospital inventory, and replace consumed hospital inventory by requesting an RGA as described above.4. Return affected product to HeartWare. Your HeartWare representative will assist you with returning affected product to HeartWare.5. Complete and return the attached Completion Form no later than 6 months from the date of this let |  | 18631 | 18, 631 | US Nationwide distribution in the states of AL, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, and Washington D.C. | Open, Classified |  |  | 18631
Z-2146-2016 | 74356 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | QuickSlide MGS-80, Automated Mutli-Slide Gram Stainer/GramPRO-80. |  |  | May 24, 2016 | 2016 | Hardy Diagnostics | Discrepant culture and gram stain results, which could potentially lead to misdiagnosis or improper treatment of a patient. | Under Investigation by firm | The firm began notifying consignees via phone on 05/23/16. The firm also sent out Urgent Notification letters, dated June 15, 2016.The firm's phone script instructs customers to decrease the number of rinses on the slide and increasing the number of rinses in the chamber for better outcomes and performance with the instrument. Reducing the rinses on the slide may prevent bacteria from coming off the slide, and adding rinses to the empty chamber will help ensure that the chamber is clean before the next slide is introduced.The firm's customer notification letter recommended an immediate update to the automated rinse cycle to decrease the number of slide rinses and increase the number of empty chamber rinses to improve instrument performance. A step-by-step procedure for updating the rinse cycle settings was discussed via phone and should now be implemented.In addition to the changes, the firm is updating the GramPro 80 User Manual to reflect new recommendations, and it is strongly recommended that all users perform methanol fixation of slides for automated staining and discontinue the use of heat fixation to prevent cross contamination of the unit. Once updates to the User Manual have been completed, the firm will provide an updated revision for customer records.For further information or technical assistance, contact QuickSlide Technical Team at gibbsd@HardyDiagnostics.com. |  | 25 | 25 | U.S. Distribution to the following states: CA, TX, NE, OK, LA, NC, PA, AZ, NY, PA, VA, NM, MI, and WV. | Open, Classified |  |  | 25
Z-2147-2016 | 74490 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | AORTIC ARCH CANNULA, Catheter, Cannula And Tubing, Vascular, Cardiopulmonary Bypass (part numbers NA-55X7 and NA-55X8) |  |  | June 22, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group USA is recalling Aortic Arch Cannulae (part numbers NA-55X7 and NA-55X8) because the distal end may be less than specified. The defective cannula are showing an angle that is about 15-20 degrees instead of approximately 45 degrees. | Process control | The firm, LivaNova, sent an "Urgent Safety Alert" field communication letter dated June 22, 2016 via certified mail to consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to: check your inventory to determine if you own Cannulae, if no inventory, complete the response form and return it per instructions indicated on form.; check the tip angle to determine if it is conforming to specifications, if specifications not met, file a complaint through your normal process whether you decide to keep it, use it or return it to Sorin Group; If you are not willing to use it, return affected product parts to Sorin Group USA,; contact Customer Service at 1-800-650-2623 or email CustomerService@livanova,com to have a RMA issued, and complete and return the Customer Response Form no later than July 20, 2016 via Fax at 303-467-6502 or by e-mail to USFSN@livanova.com. If you need support for the execution of these actions, please contact your LivaNova sales representative or call Customer Service at 1-800-650-2623 (Monday to Friday 8AM to 5PM MT).For questions regarding this notice, please contact Customer Service at 1-800-650-2623 (Monday to Friday 8AM to 5PM MT) or e-mail USFSN@livanova.com. |  | 5740 | 5740 | Worldwide Distribution: US (nationwide) to states of: LA, Ml, MO, OH, TN, and TX; and countries of: Italy and Canada. | Terminated | March 27, 2017 | 278 | 5740
Z-2144-2016 | 74522 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge OfficePACS software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Potential data loss occurs as a result of the software archiving not working properly. | Software design | Merge Healthcare sent an Urgent Medical Device Recall letter dated January 26, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to provide a copy of the notification to all users. For additional questions customers were advised to call Merge Customer Service at (87) 741-5369. For questions regarding this recall call 262-367-0700. |  | 216 | 216 sites have the affected software | The software was distributed nationwide to medical facilities. There was no government/military/foreign distribution. | Open, Classified |  |  | 216
Z-2142-2016 | 68318 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | The LifeCare Pump Model 4200 PCA Infuser makes possible the practical implementation of patient-controlled analgesia. The health care professional sets the prescribed dose, rates, and limits for the drug to be infused. The infuser can be set to operate in one of three modes: PCA Mode, Continuous Mode, and PCA + Continuous Mode. The LifeCare Pump Model 4200 PCA Infuser is designed with a number of mechanical, electrical, and software safeguards that prevent patient injury or drug abuse. |  |  | March 08, 2013 | 2013 | Hospira Inc. | Hospira has received reports of PCA pumps not detecting distal occlusions, one of which resulted in a serious injury. The issue is caused by normal wear and tear on the Half Nut (the component/nut that travels up and down the lead screw) which prevents it from properly detecting the pressure build-up associated with a distal occlusion. | Component design/selection | Hospira sent an IMPORTANT DEVICE INFORMATION letter dated March 8, 2014 to all affected customers.. The letter advised customers to: 1) Perform the Performance Verification Test (PVT) Occlusion Test as defined in the PCA Technical Service Manual (TSM); 2) If the device does not pass the test, remove it from clinical service and contact Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) to report the issue; and 3) Perform the appropriate troubleshooting and repair activities defined by your facility, which may include returning the device to Hospira for further diagnosing and servicing. The letter advised the Hospira is in the process of establishing a useful life for the half nut and to determine when it will require replacement. A new Technical Service Manual will be issued in late 2014 that adds a requirement for annual PVT Occlusion testing. The System Operating Manual will include new instructions to reduce the excessive wear on the half nut resulting from incorrectly using the vial information. Customers who further distributed the devices are requested to notify their customers of the recall and ask them to call Stericycle at 888-965-5801 (Monday - Friday, 8:00 AM - 5:00 PM EST) to receive a reply form. Questions about this recall should be direct to Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 6268 | 6,268 infusion pump systems | Nationwide Distribution to AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY) + District of Columbia + Puerto Rico + Virgin Islands; *** FOREIGN: | Completed |  |  | 106513
Z-2149-2016 | 74585 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Surface Monitoring System{OSMS, Varian Private Label} This is used on Varian's TrueBeam EDGE; |  |  | June 23, 2016 | 2016 | Varian Medical Systems, Inc. | Varian Medical Systems has received reports that unintended changes were made to planned couch parameters, specifically the couch rotation parameter, during patient set-up. This can occur when a user sets a different couch rotation from the plan couch rotation while performing patient alignment with Optical Surface Monitoring System [OSMS]. | Software design | The firm, Visionrt, sent an "URGENT MEDICAL DEVICE CORRECTION NOTIFICATION-URGENT FIELD SAFETY NOTIFICATION" letter dated June 24, 2016 to all affected customers. Letters describe the product, problem and actions to be taken. The customers were instructed as follows: DO NOT USE the "Move Couch" function if the couch rotation value set in the AlignRT device DOES NOT match the planned couch rotation for the treatment field; retain a copy of this document along with your most current product labeling; pass the letter to all those who need to be aware of it within their organization; always confirm that the patient is correctly aligned prior to commencing treatment; is not required to return their devices; promptly inform Vision RT if they believe that patient harm occurred, and complete Appendix 1- Customer Acknowledgment Response and return it via email to corrections@visionrt.com.Should you have any queries on this letter, please do not hesitate to contact Vision RT by telephone on +44 20 83464300 (866-778-2379 from the US) or as per http://www.visionrt.com/contact/details. |  | 189 | 166 OSMS in US, 23 OSMO - International, | Worldwide Distribution: US (nationwide) and countries of: Austria, Azerbaijan, Brunei Darussalam, Bulgaria, Denmark, Georgia, Germany, India, Italy, Malaysia, Morocco, Norway, Portugal, Switzerland, and United Kingdom. | Open, Classified |  |  | 211
Z-2158-2016 | 74283 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5 mm and 4.0 mm Cannulated Screw, Product code HWC, Device Listing Number D029094, Screw, Fixation Bone |  |  | February 16, 2016 | 2016 | Synthes (USA) Products LLC | Labeling does not match the cleared indications for use in the United States and Canada. | Labeling False and Misleading | DePuy Synthes initiated a Field Safety Notice to remove the affected Technique Guides and marketing literature from the field. Sales Consultants and Synthes Trauma Marketing will receive training on affected technique guides and marketing literature which promoted off label use. All technique guides and literature have been put on hold and removed from DePuy Synthes website. Effectiveness of this field action will be tracked using their training system. They will confirm training when completed as a means to measure effectiveness. |  | 19497844 | 19,497,844 units in total | US Nationwide Distribution in the states of :AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, N/A, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | June 01, 2017 | 471 | 19497844
Z-2148-2016 | 74030 | [6 empty] | Breathe Right Clear, 30 ct. carton., UPC 7-57145-00246-7 (US), UPC 9-300673-871749 (AU/NZ) |  |  | April 25, 2016 | 2016 | GSK Consumer Healthcare | Some cartons used in the secondary packaging of Breathe Right Clear, Large, 30 ct. Nasal Strips were incorrectly labeled for cartons intended for the AU/NZ market. | Error in labeling | GSK Consumer Healthcare sent an Urgent - Product Recall Letter dated April 4, 2016, to all affected customers and an updated recall letter dated April 25, 2016, via FedEx.The recall letters identified the product and problem. The April 25, 2016, was sent to further clarify, the lot number on the Breathe Right carton. The customers were asked to immediately examine their inventory, quarantine and cease distribution of the specific lot of the affected product. Customers were asked to carry out a physical count and record the data on the response form. The response form should be returned to GSK within five (5) business days, even if you do not have the recalled lot. The response form should be faxed to (336) 499-8795. If customer needs shipping assistance or questions about the recall process, they can call 1-844-722-2104, and refer to Event Number 535STRIPS16. If any customer has questions, they can direct them to GSK Consumer Healthcare Consumer Relations team at 800-858-6673. |  | 1440 | 1,440 cases ( 34, 560 retail - units ) | US Nationwide | Terminated | July 19, 2017 | 450 | 1440
Z-2159-2016 | 74510 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge OfficePACS software. The firm name on the label is Merge Healthcare, Heartland, WI. |  |  | January 31, 2016 | 2016 | Merge Healthcare, Inc. | Potential data loss occurs as a result of product archiving not working properly. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/31/2016 explaining the reason for recall and instructing the customer to provide a copy of the notification to all users. |  | 49 | 49 sites have the affected software | Distribution was made to medical facilities located in AL, AZ, CA, CO, CT, FL GA, ID, IL, KY, MA, MO, NE, NH, NJ, NY, NC, OK, PA, SC, TN, VA, and WA. There was no government/military/foreign distribution. | Open, Classified |  |  | 49
Z-2169-2016 | 74425 | No | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Target Detachable Coils: |  |  | June 06, 2016 | 2016 | Stryker Neurovascular | Stryker Neurovascular has become aware that some units of Target Nano product do not have the intendedstretch resistance. The sutures inside the coil that provide stretch resistance may have been damaged during manufacturing which can lead to an increased occurrence of coil stretching. | Process change control | Stryker Neurovascular sent an Urgent Medical Device Voluntary Recall letter dated June 3, 2016, to all affected customers via FED Ex. Letter includes recall letter, list of impacted lot numbers, customer acknowledgement form and frequently asked questions. Customers were instructed to immediately check their inventory. segregate the effected inventory in a secure location for return to Stryker, circulate the Field Safety Notice internally to all interested/affected parties. Inform Stryker if any of the subject devices have been distributed to other organizations. Customers were also instructed to complete the attached customer response form and return the completed form to their nominated Stryker Representative. |  | 45302 | 45,302 (8,188 in the U.S.) - total, all model numbers | Worldwide Distribution - US (nationwide) and Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Canary Islands, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Ireland, Italia, Japan,Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Mexico,Netherlands, Norway, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, United Kingdom, United States, Uruguay, and Vietnam. | Open, Classified |  |  | 45302
Z-2151-2016 | 74556 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NovaPACS  versions 7.4, 7.5, 7.6, and 8.0, Diagnostic Viewer. Distributed by Novarad Corporation. Picture archiving and communication system (PACS) |  |  | May 09, 2013 | 2013 | Novarad Corporation | Potential for images to be flipped while streaming, which could incorrectly display image orientation markers. | Software design | Consignees were notified on 05/09/13 via letter and email to upgrade software and to turn off streaming function. |  | 2695 | 2695 | Worldwide distribution including US nationwide. No Canadian distribution. | Open, Classified |  |  | 2695
Z-2202-2016 | 74489 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Femoral Arterial Line |  |  | June 02, 2016 | 2016 | Arrow International Inc | Labeling error. Some shipping cartons labeled as ASK-04018-VUH contained ASK-04020-VUH kits; and some shipping cartons labeled as ASK-04020-VUH contained ASK-04018-VUH kits. | Labeling mix-ups | Arrow International mailed an Urgent Medical Device Recall Notification letter and acknowledgement form dated June 20, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to 1. Immediately discontinue distribution and quarantine any products with the product codes and lot numbers listed above.2. Inspect affected product within your control to identify if the shipping corrugate label correctly identifies the products inside the shipping carton. Products with an incorrect label on the shipping carton can be returned to Arrow. 3. Once you have completed inspection of all of the products from your own inventory, check the box on the enclosed Acknowledgment Form that indicates that you have completed the inspection and fax it to 1-855-419-8507, Attn: Customer Service, or email it to recalls@teleflex.com.For questions contact your local sales representative or Customer Service at1-866-246-6990. |  | 50 | 50 units in total | US distribution in the states of TN and KY. | Terminated | August 21, 2017 | 445 | 50
Z-2191-2016 | 74416 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Part # SD313.011, Matrix Neuro Screwdriver Blade MQC/MR Conditional/Short, Lot # UR75220, mfg 7/18/2007 |  |  | May 31, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is voluntarily initiating this recall due to affected devices that were distributed, etched and or labeled with MR Safety information not meeting the current ASTM standard. | Labeling False and Misleading | DePuy Synthes mailed an URGENT MEDICAL DEVICE FIELD SAFETY NOTIFICATION letter to their customers (Hospital Personnel, including Imaging Department Personnel and DePuy Synthes Colleague) to inform them of the recall and actions to take. The letter identified the affected, problem and actions to be taken. The letter instructed customers not to use the affected Instruments which are labeled or etched MR Safe and MR Conditional during MRI scans. Customers were asked to review the information contained in the Field Safety Notification letter and complete the Verification Section located on page 3 of this notification. For questions call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 18 | 18 units | US Nationwide Distribution in the states of: AR,CA,CO,CT,DC,ID,IL,KY,LA,MA,MD,MN,MO,MS,MT,NY,PA and TX. | Open, Classified |  |  | 749
Z-2203-2016 | 74154 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | cobasÂż Liat Influenza A/B Quality Control Kit; For use with cobasÂż Liat Influenza A/B Assay. |  |  | May 17, 2016 | 2016 | Roche Molecular Systems, Inc. | Kit labeled with the incorrect expiration date. | Incorrect or no expiration date | An Urgent Medical Device Recall (UMDR) notice (Customer Letter) was sent to the 35 individual affected customer sites by UPS on May 17 2016.Customers were asked to take the following actions: " Discontinue use of cobasÂż Liat Influenza A/B Quality Control kit lot numbers 51222E, 60121H, 60201D, and 60229F, and discard them per your local regulations. " Complete the enclosed UMDC faxback form (6941-00-0516) and fax it to 1-844-449-8506. " File this UMDC for future reference. |  | 194 | 194 distributed kits | Distributed in the states of: TX, NY, OR, HI, MD, PA, NC, MA, IA, WA, OR, AK, ME, MN, VA, MT, PA, and LA. | Open, Classified |  |  | 194
Z-2207-2016 | 74389 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | HX2" Temperature Management Systems; |  |  | June 08, 2016 | 2016 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) is conducting a voluntary recall of its heater-cooler devices due to the distributor no longer selling the Interchlor sanitizer (Product Code 205) referenced in its heater-cooler cleaning instructions. | Nonconforming Material/Component | Terumo Cardiovascular Systems (Terumo CVS) initiated a voluntary recall on 06/08/2016, by mailing certifed letters to customers. Terumo CVS is notifying all users of its heater-cooler devices that the distributor, Intercon Chemical Company, no longer sells the Interchlor sanitizer (Product Code 205) referenced in its heater-cooler cleaning instructions. However, the identical sanitizer is available from other sources.This notice applies to the following Terumo CVS heater-cooler devices:Âż Sarns" TCM and TCM II Cooling and Heating SystemsÂż Sarns" Dual Cooler-HeatersÂż HX2" Temperature Management SystemsTerumo CVS will provide users with updated user instructions regarding the 10% Sodium Hypochlorite sanitization solution. Instructions will be provided under separate cover, when available; updated manual and cleaning guide will be completed in December 2016.Customers are asked to do the following: Review this notice and ensure that all users receive notice of this issue.2. Return the attached Customer Response Form acknowledging receipt of this notice.3. Source a replacement for the Interchlor sanitizer. See information above or call Terumo CVS Technical Support for assistance.If you have questions about this notification, please contact your Terumo representative or Terumo CVS Technical Support:800.441.3220, extension 6932Technical Support Hours:Monday  Friday8 a.m.  6 p.m. ET(24/7 Emergency Hours) |  | 6412 | 6412 units | Domestic: AK, AL, AR, AZ, CA , CO, CT, DC , DE , FL, GA, HI, IA, ID, IL, IN, KS, KY,LA , MA, MD, ME, MI , MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV , NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY Foreign: Argentina, Australia, Belgium,Brazil, Canada, Chile, China, Colombia, Costa Rica, Dominican Republic, Ecuador, El Salvador, England, Finland, France, GermanyGreece, Hong Kong, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Korea, Lebanon, MalaysiaMalta, Mexico, Netherlands, New Zealand, Norway, Philippines, Portugal, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates , Uruguay , Venezuela, Vietnam | Terminated | August 04, 2017 | 422 | 6412
Z-2204-2016 | 74562 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | 5.0mm Round Fluted Bur, Super Long |  |  | June 16, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments initiated a voluntary recall of specific lots of Round Fluted Burs, due to tarnishing or corrosion which may be present on the recalled burs and could result in a foreign body reaction (inflammation) necessitating surgical intervention. | Under Investigation by firm | Stryker initiated a voluntary recall of the Stryker Round Fluted Burs (used with Stryker CORE" System) on June 16, 2016, via certified mail due to the potential for the devices to have tarnishing or corrosion present. Non-responding accounts will be contacted by phone, email, and/or fax as necessary in an effort to obtain signed Business Reply Forms. Affected product will be removed from the field. Further corrective and preventative actions will be determined through the associated CAPA. The existing label will be defaced, packaging will be opened and products will be destroyed. Actions to be taken by the Customer/User:1. Immediately review this Recall Notification. 2. Check all stock areas and/or operating room storage to determine how many affected Round Fluted Burs from the affected lots are at your facility. Quarantine and discontinue use of the recalled Round Fluted Burs.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if you dont have any affected product on hand. Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification.4 If you have further distributed this product, please forward this letter and the attached BRF to all affected locations. Please indicate each location on the BRF. 5. Fax the completed BRF to Stryker Instruments Regulatory Department at 866-521-2762. 6. If the BRF for your facility indicates that recalled product is currently on hand, a FedEx label will be emailed to you. This shipping label should be used to return recalled product. Upon receipt of the recalled product, a credit will be issued to your account.Report any serious adverse events or product quality problems to Stryker Instruments: 1-800-253-3210 Health care professionals and consumers |  | 25 | 25 | US Distribution to the states of : CA, IL, SDForeign: Canada, Netherlands | Terminated | June 06, 2017 | 355 | 25
Z-2211-2016 | 74297 | No | OIJ | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(K) Exempt | Mammotome Revolve Dual Vacuum Assist Biopsy System Model number: MST1009 &amp; MST0815. Each disposable probe is packaged within a thermoformed tray and covered with a pre-printed (product code and graphics) Tyvek lid. Each tray is labeled on-demand after sealing with a lot number and expiration date. Trays are kitted into sets of five probes and packaged within a sales carton. Two sales cartons are then packed into a corrugated case and shipped to the global distribution center. Device is labeled as sterile. |  |  | May 19, 2016 | 2016 | Devicor Medical Products Inc | A 10 Gauge needle, 9cm long was shipped in an 8 Gauge needle, 15 cm package. | Packaging | An Urgent: Medical Device Recall of Mammotome Revolve Dual Vacuum Assist Biopsy Probe due to Potential Product Code Mislabel communication was sent to the customers via 1st class USPS on June 6, 2016. The labeling and the Customer response form was also attached. Customers were asked to examine their inventory and quarantine the affected product. If the product was further distributed, those customers are to be identified and notified. The letter stated that the assigned sales rep would be in contact to support the identification, inspection and replacement of the recalled products. Customers are to complete and return the enclosed response form. Questions are to be directed to Mammatome Customer Service 1-877-926-2666. |  | 360 | F11537107D1 = 360; F11536423D1 = 80; F11537109D1 = 360 | Worldwide Distribution -- USA, including the states of AL, AR, CA, CT, DC, FL, GA, IL, KS, KY, LA, MA, MD, ME, MI, MN, MS, NC, NJ, NY, OH, PA, SC, TN, WA and WI; and, the countries of Canada and Japan, Korea, and The Netherlands. | Terminated | August 03, 2017 | 441 | 360
Z-2213-2016 | 73696 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | AMD-RITMED brand Conforming Stretch Gauze Bandages, Sterile, 6 in x 4/1 yds./15 cm x 3.7 cm, Reorder Number/Code D606, Bar code 6 86864 04230 3, Made in China -- Manufactured for &amp; distributed by: AMD-RITMED INC. Montreal Quebec H8T 3J8 -- US Distribution: AMD-RITMED INC. Tonawanda, NY 14150 -- Type of packaging: 1 roll/pouch, 12 pouches/inner box, 4 inner box/case, 48 rolls per case. |  |  | March 30, 2016 | 2016 | Amd-Ritmed, Inc. | The product lot did not meet sterility requirements based on FDA sampling and analysis. | Material/Component Contamination | The firm, AMD-Ritmed, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated 3/30/2016 and Medical Device Recall Response Form to the US distributors/customers via e-mail. The letter described the product, problem and actions to be taken. The customers were instructed to examine your stock to determine if you have any; immediately discontinue distributing product; promptly return product via parcel post to Tonawanda/Buffalo warehouse, ATTENTION: RETURNED GOODS, AMD-Ritmed Inc., 295 Firetower Drive, Tonawanda, NY 14150; if you have further distributed the product, immediately contact your accounts, advise them of the recall situation and have them return their outstanding stock, and complete and return the enclosed MEDICAL DEVICE RECALL RETURN RESPONSE-Acknowledgment &amp; Receipt Form via email to : AMDCUSTOMERSERVICE@AMDRITMED.COM or fax to: 1-866-871-3455, ATTN: AMD-Ritmed Customer Service.If you have any questions, please contact us at (1-800-445-0340) or at AMD Customer Service AMDCustomerService@amdritmed.com. |  | 107 | US: 107 cases (48 rolls per case) | US Distribution in states of: NY, CA, and TN. | Terminated | July 15, 2016 | 107 | 107
Z-2206-2016 | 74618 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation 3.0, RayStation 3.5, RayStation 4.0, RayStation 4.5, RayStation 4.7 and RayStation 5 --- Radiation Therapy Treatment Planning System --- designed for treatment planning and analysis of radiation therapy, provides treatment unit set up parameters and estimates dose distributions. |  |  | July 01, 2016 | 2016 | RAYSEARCH LABORATORIES AB | An issue was found with the evaluation of biological clinical goals in RayStation 3.0, RayStation 3.5, RayStation 4.0, RayStation 4.5, RayStation 4.7 and RayStation 5. Biological clinical goals for an adapted plan based on another planning CT than the original planning CT will show incorrect values. -- To the best of the firm's knowledge, the issue has not caused any patient mistreatment or other incidents. However, the user must be aware of the following information to avoid incorrect plan evaluation during treatment planning. | Under Investigation by firm | RaySearch Laboratories AB sent Field Safety Notice/Medical Device Correction #11905 Letter (RSL-D-61-298, dated 7/1/2016) to end users via e-mail on 7/1/2016. The letter informs end users of the issue and provides actions to be taken by the users (work around). |  | 170 | US: 170 units | California, Connecticut, Delaware, Florida, Hawaii, Maine, Missouri, New York, Ohio, Texas and Washington | Open, Classified |  |  | 170
Z-2216-2016 | 74094 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | cobasÂż LiatÂż Analyzer |  |  | May 09, 2016 | 2016 | Roche Molecular Systems, Inc. | Device was released with default pcal/gain settings forphotometers instead of the pcal/gain settings that are required for release. The default settings can potentially lead to the generation of erroneous or invalid results. | Under Investigation by firm | Customers were notified of the recall via phone and email sent on 5/9/2016. They were instructed to take the following actions: Immediately discontinue use of the affected instruments until they have been reworked by a Roche Field Engineering Specialist. A Roche representative will contact you to schedule this service. Complete the attached fax form and fax it to 1-317-521-4815. File this Urgent Medical Device Correction (UMDC) for future reference. Please contact the Roche Support Network Customer Support Center 1-800-800-5973 if you have questions. |  | 2 | 2 | Distributed in NC and WA. | Completed |  |  | 2
Z-2214-2016 | 74415 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Depuy Synthes Trauma- Trochanteric Fixation Nail Advanced (TFNA) Blade/Screw Guide Sleeve, Surgical Instrument |  |  | June 07, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is initiating a voluntary medical device recall due to Blade/Screw Guide Sleeves (Part # 03.037.017) may have been manufactured with a slightly larger than specified outer diameter which could prevent the sleeve from going through the TFNA aiming arm properly, which may cause surgical delay. | Nonconforming Material/Component | DePuy Synthes sent an Urgent Notice Medical Device Recall letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to conduct functional testing to inspect the instrument to determine if the Guide Sleeve is able to fit through the aiming arm properly. Customers were asked to only return the device if the product does not pass.Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales Consultant. |  | 2064 | 2,064 units | Nationwide Distribution toAK,AL,AR,AZ,CA,CO,CT,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NC,ND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,PR,RI,SC,SD,TN,TX,UT,VA,VT,WA,WI,WV, and WY | Terminated | April 25, 2017 | 322 | 2064
Z-2222-2016 | 74437 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | T030-000 -30" 20 Drop IV Admin Secondary Set. Packaged in a Tyvek pouch, with 50 pouches per case. |  |  | June 14, 2016 | 2016 | WalkMed Infusion, LLC | WalkMed Infusion, LLC Announces a Voluntary Field Action of Multiple Triton Administration Sets Due to the Potential Presence of Particulate In or On the Administration Sets. | Nonconforming Material/Component | Terumo sent a notification letter dated June 8, 2016 to customers. The letter identified the affected product, problem and action to be taken. Customers were instructed to 1. Review this notice and ensure that all users receive notice of this issue. 2. Return the attached Customer Response Form acknowledging receipt of this notice. 3. Source a replacement for the Interchlor sanitizer. See information above or call Terumo CVS Technical Support for assistance. 4. Terumo CVS will provide users with updated user instructions regarding the 10% Sodium Hypochlorite sanitization solution. Instructions will be provided under separate cover, when available.For questions contact your Terumo representative or contact Terumo CVS Technical Support at 800.441.3220, extension 6932. |  | 2453 | 2453 cases | US Nationwide Distribution in the states of AL, AZ, CA, FL, IL, MA, MS, NJ, NV, NY, OH, PA, TN, UT, and WA. | Open, Classified |  |  | 24249
Z-2217-2016 | 74652 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Carestream Touch Prime, Catalog Number 1738830, and Carestream Touch Prime XE, Catalog Number 1738822 --- Common Name: Touch Ultrasound |  |  | July 08, 2016 | 2016 | Carestream Health Inc | A software issue related to the generic volume measurement functionality could result in the volume of the subject anatomy being overestimated. For example, measurement results of the anatomy may appear enlarged or distended when it actually is not. | Software design | The firm, Carestream Health Inc., sent an "URGENT: MEDICAL DEVICE CORRECTION" Letter dated 7/7/2016 and response form the end users/customers on 7/8/2016 via Certified Mail, Return Receipt Requested. The letter described the product, problem and actions to be taken. The Customers were instructed not use the generic volume measurement tool until Carestream has updated the system software. In the interim, calculate volumes by using the specific volume tool from the desired calculation package. If you have distributed the device outside your facility, please alert your customer(s) of this field correction and contact Carestream Customer Care Center as listed above.Carestream has created a software update for Touch Ultrasound to resolve this issue and will update your software. Carestreams representative will contact you and schedule a convenient time to upgrade the software.For questions or concerns, please contact the Carestream Customer Care Center in the US at 1-800-328-2910 (available on a 24/7 basis). Outside of the US, please call your local Service support number. |  | 7 | 7 Units (US: 5 units, OUS: 2 units) | Worldwide Distribution: US Distribution to states of: GA, IA, and TX; and country of: Italy. | Terminated | May 02, 2017 | 298 | 7
Z-2219-2016 | 74436 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The pump (with cushioning foam inserts) is packaged in a single pump box. Four pump boxes are placed in an over-shipper for distribution. |  |  | June 14, 2016 | 2016 | WalkMed Infusion, LLC | WalkMed Infusion, LLC Announces a Voluntary Field Action of the Triton Infusion Pump (model 300000) and Triton FP Infusion Pump (model 400000) Due to Unapproved Changes to the Software and Specifications of the Triton Infusion Pump (model 300000) and Unapproved Changes to the Intended Use of the Triton FP Infusion Pump (model 400000). | Device Design | WalkMed sent an Urgent : Medical Device Recall letter dated June 14, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Action to be taken by Users: WalkMed recommends the following actions be taken with respect to the recalled product:Our records indicate that you have received Triton pumps. Immediately review your inventory and determine whether you have any of the Triton or Triton fp infusion pumps (model numbers 300000 and 400000).Please complete and promptly return the enclosed recall response form, even if you do not have any affected stock subject to this recall notice.WalkMed recommends that users stop use of and replace the Triton and Triton fp Infusion pumps with alternative devices as soon as possible.To ensure that patients who use WalkMed Triton infusion pumps can continue to receive infusion therapies without interruption, WalkMed will support customers in their transition to alternative devices. To ensure continuity of patient care, WalkMed will continue to provide IV administration sets to existing customers, provide preventive maintenance and service for distributed pumps and make available any parts necessary for the repair or service of units for up to 3 months from the date of this recall notice.Contact WalkMed at (303) 420-9569 to arrange for shipment of the product being returned.If an owner of a Triton or Triton fp infusion pump is interested in information about options for a replacement product, please contact WalkMed. For further questions please call (303) 420-9569. |  | 2482 | 2482 units | Nationwide to AL, AZ, CA, FL, IL, KS, MA, MD, MS,NJ, OH, PA, TN, TX, UT, and WA. No foreign/VA/govt/military. | Terminated | April 28, 2017 | 318 | 5448
Z-2188-2016 | 74473 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems Angio WorkStation: XIDF-AWS801v6.00, v6.01, and V6.10 System: INFX-8000V Bi-Plane System |  |  | February 24, 2016 | 2016 | Toshiba American Medical Systems Inc | During a procedure the Peak Skin Dose (PSD) value was displayed on the DTS larger than the DTS expected value. It was found that the software incorrectly calculated the PSD value when the x-ray condition of the frontal plane and the lateral plane are the same. It incorrectly used in the calculation a dose value from a previous exposure. | Radiation Control for Health and Safety Act | Toshiba Medical Systems (TAMS) planned action to bring product into Compliance:1.TAMS will install the software and perform testing to ensure the software update was effective.2. The customer notification letter which includes a statement that TAMS will, without charge, remedy the defect or bring the product into compliance.. If you have any questions regarding this, please feel free to contact your local Toshiba representative at (800) 521-1968. |  | 23 | 23 | US Distribution to the states of : NC, NY, TX, GA, CA, AZ, IL, FL,MA, MO and DE. | Terminated | April 25, 2017 | 426 | 23
Z-2218-2016 | 74449 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith and Nephew TWINFIX QUICK-T Knot Pusher Suture Cutter; Part Number: 7209417, for arthroscopic use. |  |  | June 15, 2016 | 2016 | Smith &amp; Nephew, Inc. | The device could be subject to breaches of its sterile packaging. | Package design/selection | Urgent Recall Notification Letters were sent via Federal Express on June 15, 2016, to the customers listed in the firm's distribution report. Customers were asked to inventory their devices and contact the firm for return authorizations. |  | 563 | 563 | Worldwide distribution. US nationwide, AT, AU, DE, FR, GB, JP, PT, and RU. | Open, Classified |  |  | 563
Z-2238-2016 | 74639 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | 2.3MM Tapered Router, |  |  | June 24, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | A variation in flute depth on the routers was observed. The variation in flute depth may cause the routers to bemore susceptible to breakages.Risk to Health:Injury to critical soft tissue requiring surgical intervention may occur. | Process design | Stryker sent an Urgent Medical Device Recall Notification letter dated June 24, 2016, to the affected customers via courier, and to Sales Representatives and international affiliates via e-mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments at 269-389-2921. |  | 3163 | 3,163 | Worldwide Distribution - US (Nationwide) Netherlands, Canada, Japan, Switzerland, and United Kingdom | Terminated | March 23, 2017 | 272 | 3163
Z-2239-2016 | 74295 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | ATEC CANISTER. Suction Canister with Lid 490 cc., RX Only; Suction canister used to collect fluid during a breast biopsy. |  |  | May 31, 2016 | 2016 | Hologic, Inc | Cracks ATEC Canister Lids. | Process control | Customers were notified of the recall via certified letter beginning on May 31, 2016. The letter requests that customers immediately inventory your supplies of the Recalled product in all appropriate areas and cease use of all affected stock. In addition, they are asked to complete the enclosed Recall Response Card indicating whether you have any of the Recalled product on hand, as well as the quantity of product that you have in stock. |  | 16870 | 16,870 | Worldwide distribution. US: AK AL AR AZ CA CO CT DC DE FL GA HI IA ID IL IN KS KY LA MA MD ME MI MN MO MS MT NC ND NE NH NJ NM NV NY OH OK OR PA PR RI SC SD TN TX UT VA VT WA WI WV WY OUS: United Arab Emirates, Australia, Argentina, Belgium, Canada, Switzerland, Chile, Colombia, Germany, Denmark, France, Great Britain, Hong Kong, Ireland, Israel, Italy, Korea, Kuwait, Lebanon, Malaysia, Netherlands, New Zealand, Qatar, Singapore, Thailand, Taiwan, and South Africa. | Terminated | May 24, 2017 | 358 | 16870
Z-2243-2016 | 74547 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Parker Trach-Vac Endotracheal Tubes. Tyvek pouches, 10 per box. |  |  | June 02, 2016 | 2016 | Parker Medical | Parker Medical announces a voluntary field action for the Parker Trach-Vac Endotracheal Tubes - (Endotracheal tube,with Parker Flex-Tip,Sub-Glottic Suction, High Volume low Pressure) due to tube kinking during use. | Nonconforming Material/Component | Customers were notified on June 2, 2016 that a recall was initiated via telephone. The information communicated identified the affected product, problem and actions to be taken. Customers were instructed to segregate the affected product for immediate pick-up. A Sales Representatives went to each hospital and retrieved all unused product. |  | 5500 | 5,500 units | US Nationwide Distribution in the states of KY, MO, FL, TN, and TX. | Terminated | January 31, 2017 | 243 | 5500
Z-2241-2016 | 74468 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension VistaÂż Homocysteine Flex Reagent Cartridge (HCYS), Device Listing No.: D017878 is an in vitro diagnostic test for the quantitative measurement of total homocysteine in human serum, heparinized plasma and EDTA plasma on the Dimension Vista System. |  |  | May 17, 2016 | 2016 | CSL Behring GmbH | Reagent lots # 15243MA and 15243MB shows a reduced stability once opened that does not meet the Instructions For Use claim of seven (7) days. A maximum bias of -37% was observed. | Under Investigation by firm | Siemens mailed an Urgent Medical Device Recall to all customers who received the affected lots notifying them of the following: 1) Discontinue use and discard any remaining inventory of the affected lots, and 2) Indicate their replacement product needs within 30 days. |  | 898 | 898 units | Distributed to: AL, AR, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, LA, MA, MD, MI, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, WI, WV | Open, Classified |  |  | 898
Z-2242-2016 | 74484 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | MODIFIED LINEAR STANDARD/LATERAL NECK TRIALS WITH SPRING SIDES REF S-200635 and S-200636 |  |  | June 21, 2016 | 2016 | Encore Medical, Lp | Reported failure of the instrument. The retaining ring component of a neck trial became disassembled from the body of the neck. | Process design | The recalling firm notified affected consignees via letter on 6/21/16. The consignees were instructed to return the affected devices. |  | 14 | 14 (7 sets) | CO, ID, TX | Terminated | January 10, 2017 | 203 | 14
Z-2237-2016 | 74486 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Replace Select Tapered TiU NP 3.5x11.5mm, Article No. 36105 |  |  | June 22, 2016 | 2016 | Nobel Biocare Usa Llc | Incorrect top label on an implant package. | Labeling False and Misleading | Nobel Biocare sent a customer notification letter dated June 23, 2016 to all affected customers to inform them that Nobel Biocare has detected that the lot#13030677 of Replace Select Tapered TiU NP 3.5 x 11.5 mm (product#36105) carries an incorrect top label. The letter informs the customers that the incorrect top label states "Replace NP 3.5 x 11.5 CC" (ie. Conical Connection). The product, Replace Select Tapered TiU NP 3.5 x 11.5 mm, has a Tri-lobe connection. Customers were instructed to complete the attached Quality Message Acknowledgement form. Customer with questions were instructed to contact customer service at 714-282-4808. |  | 53 | 53 units total (33 units in US) | Worldwide Distribution - US including OR, OH, NY and Internationally to Taiwan and BrazilNon US: Taiwan, SC | Completed |  |  | 53
Z-2244-2016 | 74513 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiographic Injector and Syringe, MEDRAD MRXperion MR Injection System, MEDRAD MRXperion Sterile Disposable MRI Kit |  |  | June 09, 2016 | 2016 | Bayer Healthcare | Bayer Healthcare is initiating this recall due to complaints that were received from customer sites describing a 4205 error message when the injector is used with a 3T scanner. | Device Design | Bayer sent an Urgent Medical Device Field Corrective Action letter dated June 8, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to take the following steps should a 4205 error message is encountered. For questions contact Bayer Customer Support Team at 1-800-633-7237. |  | 42 | 42 units | Worldwide and the US state of NEW YORK | Open, Classified |  |  | 42
Z-2240-2016 | 74576 | No | JDO | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TSP Hip Fracture Plating System |  |  | June 10, 2016 | 2016 | Biomet, Inc. | The screw may fracture at the junction of the shaft and the head during implantation. The patient may retain a foreign body if the fractured screw cannot be removed. Soft tissue irritation will not occur as the retained portion of the screw is below the surface of the plate. A delay in surgery less than 30 minutes may occur. | Device Design | Zimmer Biomet sent an Urgent Medical Device Recall letter dated June 10, 2016, to all affected customers via courier. Distributors and hospitals were instructed to locate and return the affected product, and implanting surgeons were instructed to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-6725, Monday through Friday, 8 a.m. to 5 p.m. |  | 5458 | 5458 units | Worldwide Distribution - US (Nationwide) Internationally to Australia, The Netherlands, and Japan | Terminated | July 20, 2017 | 405 | 5458
Z-2247-2016 | 74559 | [6 empty] | The Synchron Systems No Foam DxC (Reference No. B64130) affects the UniCel DxC System (chemistry analyzer). All lots of No Foam DxC reagent are affected by this recall. |  |  | June 27, 2016 | 2016 | Beckman Coulter Inc. |  | Under Investigation by firm | A Customer notification letter dated 6/27/16 informs their customers that Beckman Coulter has identified an incompatibility between No Foam DxC reagent (REF B64130) and the DxC hydropneumatic parts associated with No Foam delivery. The letter informs their customers that the DxC No Foam reagent is intended to prevent formation of foam and/or bubbles in the DxC waste system of the hydropneumatic subsystems. Over time, this incompatibility may lead to: Breaks in No Foam Cap Assembly (PN A84827) and Cracks in No Foam Waste Collector (PN A60007/A64665). The letter informs the customers of the impact and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding the recall notice are instructed to contact our Customer Support Center Via http://www.beckmancoulter.com/customersupport/supportÂż Via phone, call 1-800-854-3633 in the United States and Canada. Customers outside the United States and Canada, are instructed to contact their local Beckman Coulter Representative. |  | 4221 | 4,221 units total (2,033 units in US) | Affected geographies: United Arab Emirates, Curacao, Angola, Argentina, Australia, Bangladesh, Belgium, Bulgaria, Brunei, Canada, Switzerland, Colombia, Costa Rica, Czech Republic, Germany, Ecuador, Egypt, Eritrea, Spain, Finland, France, United Kingdom, Greece, Guatemala, Hong Kong, Ireland, India, Iraq, Italy, Jordan, Japan, Korea, Republic of, Kuwait, Lebanon, Lithuania, Morocco, Macao, Mexico, Malaysia, Netherlands, New Zealand, Oman, Panama, Philippines, Pakistan, Poland, Portugal, Paraguay, Qatar, Romania, Serbia, Russia, Rwanda, Saudi Arabia, Singapore, Slovakia, Swaziland, Thailand, Tunisia, Turkey, Taiwan, China, United States of America (including Puerto Rico), Viet Nam, Mayotte, South Africa. | Open, Classified |  |  | 4221
Z-2245-2016 | 74620 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via, picture archiving and communications system software controlled. |  |  | June 23, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Incorrect values for the volume calculation. Software update VB30B via Update Instructions SY018/16/P to resolve software errors. | Software design | Siemens released and mailed to customers a notice with software update VB30B instructions via update instruction SY018/16/P for customers to follow until the update can be made. The letter also described the software issues involved; and asked customers to include the letter in the operator's manual until the update is performed. Customers are to ensure that all necessary personnel are aware of the situation. |  | 50 | 50 units | Nationwide Distribution -- AR, CA, CO, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, ND, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TX, and UT. | Open, Classified |  |  | 50
Z-2248-2016 | 74264 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OT1000 Series Orthopedic Surgical Tables |  |  | May 23, 2016 | 2016 | Steris Corporation | STERIS has identified that the pivot pin assembly located on the table top may become loose preventing the table from articulating properly. | Other | Recall Notification Letters to the affected Consignees was sent on May 23, 2016 via FedEx with tracking numbers for delivery confirmation. STERIS Field Service Technicians will visit each domestic Customer location to perform the inspection and reinstallation of the pivot pin assemblies. For international direct accounts who have further distributed the surgical table, Steris asked to identify the Customers, and to notify them of this product field correction, and that not to operate the tables until the inspection has been done of the pivot pin assembly on all affected tables. |  | 148 | 148 | AL;AK, AZ; AR; CA; CO; CT; DC; FL; GA; IL; IN; IA; NY; NJ; NC; OH; OK; OR; PA; PR; SC; TX; UT and VA. Foreign accounts: Canada; France; Korea; Panama; Qatar and Thailand. | Completed |  |  | 148
Z-2249-2016 | 74249 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Amerigel Hydrogel Wound Dressing Advanced Formula for wound management 1 oz. |  |  | May 20, 2016 | 2016 | Amerx Health Care Corp. | Product did not meet specifications for Zinc Acetate (OOS). | Other | Initial recall notification was sent to distributors, informing them of the voluntary recall and providing instructions on expanding the recall to the wholesale level. Distributors were instructed to identify recalled product in their inventory, quarantine any recalled product, and to complete and fax back a Recall Response Form. Contact information was provided for those with questions: 1-800-448-9599 and recall@amerxhc.com. |  | 6993 | 6,993 Units | AR, AZ, CA, CO, CT, FL, GA, IL, IN, MA, MD, MI, MO, NC, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, VA, &amp; Alaska. Bahrain &amp; Colombia. | Open, Classified |  |  | 6993
Z-2246-2016 | 74308 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DeVilbiss Model 7305P-D Powered Suction Unit |  |  | May 23, 2016 | 2016 | DeVilbiss Healthcare LLC | A small percentage of units of 7305 Series Vacu-Aide Suction Units are experiencing premature failure during operation. | Component design/selection | DeVilbiss Healthcare sent an Urgent - Medical Device Recall letter on May 23, 2016, to all affected customers notifying them of the issue. The letter identified the product the problem and the action needed to be taken by the customers. Customers were instructed to isolate all of the affected units included on the serial number list and return for replacement. For further questions, please call (814) 443-7602. |  | 48 | 48 units | US Distribution to the states of : NC, OH, TX, MO, TN, WI, KY, FL, LA, MN, IN, and SC | Terminated | May 30, 2017 | 372 | 48
Z-2250-2016 | 74604 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Convenience Kits from Centurion that |  |  | April 29, 2015 | 2015 | Centurion Medical Products Corporation | Kits from Centurion contain a Medtronic/Covidien Light Glove that is being recalled by the manufacturer, Medtronic.According to the notice received from Medtronic, the recall was initiatedbecause some light covers may have splits or holes, which may result in the transfer ofmicroorganisms from the light handle into the patient wound. | Nonconforming Material/Component | On 4/28/2015, URGENT PRODUCT RECALL NOTICE notifications were sent to the affected customers via certified mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. |  | 483 | 483 | KS, IN, IA, KY, MT, SD | Completed |  |  | 483
Z-2252-2016 | 74537 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Trucount Tubes: |  |  | June 24, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | One lot of BD Trucount Absolute Counting Tubes (IVD) may not be sealed properly and exposure to ambient environment can rapidly degrade the performance of the product. | Packaging process control | Urgent Product Correction notifications letters were sent on June 24, 2016 to all affected customers. Letters identify the issue and product and provide illustrations of the Kit Box and pouch labels.Correction letters identified the issue and the product involved with photos of product labels. BD letters states that they want to reinforce the need to visually inspect the desiccant contained in each pouch . If the desiccant has turned from blue to lavender, discard the remaining tubes.The acknowledgement form should be completed and returned regardless if any inventory remains. If you require further assistance, please contact BD Customer Support at 855.236.2772 (prompt 3) in the Unites States. For customers outside the US, contact your local BD Biosciences representative or distributor |  | 844 | 328 US, 516 International | USA Distribution to the states of : NC, WA, WI, NJ, PR, MD, NY, MA, CA, NY , MD, KY, IN, VA, SD, PA, OH, AR, AZ and Foreign to : Canada, Belgium and Chile | Open, Classified |  |  | 844
Z-2256-2016 | 74309 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIAÂż Centaur XPT system software versions V1.0.1 (Bundle 1.0.912 SMN 10819704), V1.0.2 (Bundle 1.0.1086 SMN 11219806), V1.0.3 (Bundle 1.0.1108 SMN 11220781) and V1.1 (Bundle 1.1.243 SMN 11221979). |  |  | May 26, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Eight (8) issues were identified which may affect the results generated by the system software version. | Under Investigation by firm | On May 26, 2016, Siemens Healthcare Diagnostics Inc. distributed an Urgent Medical Device Correction notices to their customers via courier service. The letter identified the affected product, problem and actions to be taken. Customers are advised to complete and return the Field Correction Effectiveness Check form within 30 days via fax to (312) 275-7795. For questions contact your local Siemens technical support representative. |  | 446 | 446 units Total (3 domestically &amp; 443 internationally) | Worldwide Distribution - US Nationwide in the states of MD, MN, MS and the countries of Australia, Austria, Bangladesh, Belgium, Canada, Chile, Czech Republic, Denmark, Egypt, Finland, France, Georgia, Germany, Guadeloupe, Hong Kong, Hungary, India, Ireland, Italy, Japan, Jordan, Latvia, Malaysia, Netherlands, New Zealand, Norway, P.R. China, Philippines, Poland, Portugal, Qatar, Republic Korea, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Turkey, U.A.E, United Kingdom &amp; Vietnam. | Open, Classified |  |  | 446
Z-2253-2016 | 74447 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Anspach G1 High Speed Electric Handpiece Air Cooled (eG1A) |  |  | June 15, 2016 | 2016 | The Anspach Effort, Inc. | The graphics at the attachment interface indicating locked and unlocked position are reversed. | Employee error | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE-RECALL (Removal)" letter dated June 15, 2016 to its customers. The letter described the product, problem and actions to taken. The customers were instructed to review your inventory and immediately remove the product subject to this recall from stock; If you DO HAVE any of the identified devices, please take the following steps:" Ensure anyone in your facility impacted by this notification reads this letter carefully.Customer Letter Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating product subject to the recall has been located. Indicate the number of devices found." The RMA information has been filled in for your convenience." Please include your name, title, address, telephone number, signature and date in the spaces provided." Send a copy of the completed Verification Section to DePuy Synthes, Customer Quality Department by:" Fax: (561) 627-2682" Scan/email: DPYUS-PowerToolsFieldActions@its.jnj.com" Giving it to your DePuy Synthes Sales Consultant." This return documentation acknowledges your receipt of medical device removal information." Return the affected product to your DePuy Synthes Sales Consultant" Keep this notice visibly posted for awareness until all products subject to this recall has been returned. While processing your returns, please maintain a copy of this notice with the product subject to this action and keep a copy for your records.If you DO NOT HAVE the identified product, please take the following steps:" Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicating that no affected product has been located. Include your name, title, address, telephone number, signature and date in the spaces provided." Send a copy of the completed Verification Section to DePuy Synthes, Customer Quality Department by Fax: (561) 627-2682" Scan/email: DPYUS-PowerToolsFiel |  | 18 | 18 | Worldwide Distribution: US Distribution to states of: GA and FL; and country of: Switzerland. | Open, Classified |  |  | 18
Z-2254-2016 | 74221 | No | OWD |  |  | Molecular Genetics | PMA | cobas EGFR Mutation Test V2 |  |  | April 05, 2016 | 2016 | Roche Molecular Systems, Inc. | During internal studies using contrived plasma specimens, several mutations (L8568R, Exon 19 deletion, T790M) inconsistently generated "No Mutation Detected" (ie False Negative ) result with the cobas EGFR mutation Text V2.0 when utilizing the cobas cfDNA Sample Preparation kit. | Nonconforming Material/Component | Roche Molecular Diagnostics sent an Urgent Field Safety Notice letter dated April 5, 2016 to affected customers. Then letter identified the affected product, problem and actions to be taken. The Instructions for use are provided for the user until new ones are available.. |  | 1244 | 735 devices EGFR; 509 devices cfDNA | Foriegn Distribution Austria, Belgium, Croatia, CzechRepublic, Denmark, France, Germany,Greece, Hungary , Italy, Lithuania,Luxembourg, Netherlands, Norway, Poland,Romania, Slovakia, Slovenia, Spain,Switzerland, Turkey, United KingdomcfDNA: Austria, Belgium, Croatia, CzechRepublic, Denmark, France, Germany,Greece, Hungary , Italy, Lithuania,Luxembourg, Netherlands, Poland,Romania, Slovakia, Slovenia, Spain,Switzerland, Turkey, United Kingdom | Open, Classified |  |  | 1244
Z-2265-2016 | 74505 | No | KTN | Infectious mononucleosis immunological test system. | Immunology | Immunology | 510(k) | Acceava Mono II Rapid Test Strip (Whole Blood, Serum, Plasma), Catalog No. 92404. Mononucleosis Rapid Test Device a simple test that utilizes an extract of bovine erythrocytes to qualitatively and selectively detect IM heterophile antibodies in whole blood, serum or plasma. |  |  | June 22, 2016 | 2016 | Alere San Diego, Inc. | The Positive Control information referencing human plasma as the antibody material described in the product labeling and packaging is incorrect. | Unknown/Undetermined by firm | A Medical Device Correction letter dated 6/22/16 will be sent to all customers to inform them that Alere is recalling the MONO Rapid Test products because the Positive Control Information referencing human plasma as the antibody material described in the product labeling and packaging is incorrect. Customers are informed that the labeling states the Positive Control contains the following: Diluted human plasma containing IM heterophile antibodies, 0.09% sodium azide. The Positive Control contains the following: Goat anti-mono antibody, 0.09% sodium azide. Customers are instructed to complete and fax or email the enclosed verification form within 10 days. Customers with questions are instructed to contact Alere San Diego at (844) 804-4945 or email at verfications.ts@alere.com. |  |  |  | US | Open, Classified
Z-2096-2016 | 74301 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics Soft Vu Omni Flush Angiographic Catheter |  |  | June 01, 2016 | 2016 | Stryker Sustainability Solutions | Product defect; tip of two (2) Angiographic Catheters model # 10732203 broke off, prior to being used on the patient. | Incorrect or no expiration date | URGENT MEDICAL DEVICE RECALL RESTERILIZED ANGIODYNAMICS SOFT VU OMNI FLUSH ANGIOGRAPHIC CATHETER customer notification letters were sent 06/01/16.Customers were instructed to discontinue use of the recalled product immediately. Customers were told to complete the Recall Effectiveness Check Form and list the affected lots shipped to the facility. The form should be completed, signed, and returned to the local Stryker Sustainability Sales Representative, emailed to ssspfa@stryker.com, or mail to:Stryker Sustainability Solutions1810 West Drake DriveTempe, AZ 85283Attn: Jodie Rueckert Customers will receive a credit for all affected devices returned. |  | 167 | 167 | Worldwide Distribution -- USA, to the states of TX, UT, PA, CA, VT, MA, WA, VA, NY, MD, and CT; and, the country of CANADA. | Terminated | October 19, 2016 | 140 | 167
Z-2212-2016 | 74465 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | DIAMOND-FLEX CIRCULAR RETRACTOR, ANGLED, 40MM 5MM, Product Code 89-6114, Lot Code(s): 878971; 879366; 879404, and Date Code: C16. |  |  | July 01, 2016 | 2016 | Carefusion 2200 Inc | BD, formerly CareFusion, has identified a potential risk associated with a weld failure which could result in the wire protruding thru the tip of the instrument when articulated. If this failure were to occur while in use in a procedure it has the potential to damages tissue or organs | Material/Component Contamination | BD, sent an "URGENT: Medical Device/ Safety Alert/Recall Notification" letter dated 6/30/2016 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to return the affected units, along with the enclosed CUSTOMER RESPONSE FORM to: BD formerly CareFusion, 75 North Fairway Drive, Vernon Hills, Illinois, 60061 Attn: Customer Advocacy; to expedite the correction process and acknowledge receipt of the notification. The firm will issue a credit upon receipt of the returned affected product.For questions and support 1-800-323-9088 Prompt 3 or email: GMB-US-Complaint-Intake@carefusion.com. |  | 10 | 10 | Worldwide Distribution - US, including the states of WA and MD; and, the country of Japan. | Terminated | March 29, 2017 | 271 | 10
Z-2272-2016 | 74569 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC Synchron Clinical Systems and UniCel DxC Synchron Access Integrated Clinical Systems |  |  | April 22, 2014 | 2014 | Beckman Coulter Inc. | Possible leaking of No Foam solution from the No Foam Bottle Assembly on the UniCel DxC Synchron Systems. | Nonconforming Material/Component | An Important Product Notice dated 4/22/14 was sent to all their customers to inform them that Beckman Coulter has identified a manufacturing issue with a plastic fitting on the No Foam Bottle Assembly which may crack and cause No Foam solution to leak. The letter informs the customers of the impact, action to be taken, and resolution. Customers are asked to inspect the No Foam Bottle Assembly for leaks prior to refilling the bottle. If leaks occur, customers should contact Customer Support Center following the instructions provided. Customers are instructed to complete the enclosed response form within 10 days. Customers with questions are instructed to contact the firm's Customer Support Center- Via our website, http://www.beckmancoulter.com/customersupport/support- Via phone, call 1-800-854-3633 in the United States and Canada- Outside the United States and Canada, contact your local Beckman Coulter Representative. |  | 3540 | 3,540 units total (1,067 units in US) | Worldwide Distribution -- United Arab Emirates, Angola, Australia, Belgium, Bulgaria, Canada, Switzerland, China, Colombia, Costa Rica , Czech Republic, Germany, Denmark, Egypt, Eritrea, Spain, Finland, France, United Kingdom, Greece, Guatemala, Hong Kong, Indonesia, Ireland, India, Italy, Jordan, Japan, Korea, Republic of, Kuwait, Lebanon, Lithuania, Morocco, Mexico, Malaysia, Namibia, Netherlands, Norway, New Zealand, Oman, Panama, Philippines, Pakistan, Poland, Portugal, Paraguay, Qatar, Romania, Serbia, Russian Federation, Rwanda, Saudi Arabia, Sweden, Singapore, Slovakia, Swaziland, Thailand, Tunisia, Turkey, Taiwan, US (including Puerto Rico), Vietnam, South Africa, and Zimbabwe. | Open, Classified |  |  | 3540
Z-2263-2016 | 74043 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Dual Chamber Container fill with gravity/attached transfer sets, 250 mL / 1250 mL, 300 mL / 1700 mL, 500 mL / 2500 mL, 500 mL / 3500 mL |  |  | June 23, 2016 | 2016 | The Metrix Company | An issue was identified after a redesign of the auto compounding device. The compounder cannot meet the designed accuracy specification when dispensing into bags. After reviewing the initial design of the device, it was determined the dispensing volume validation was not properly done and none of the compounders on the market can meet the stated specifications. | Device Design | On 6/23/2016, the recalling firm sent written recall notices to their distributors and end users notifying them of the recall. Customers and consignees were instructed to not use the devices, locate and quarantine all affected devices, and to complete the response form. |  | 18474 | 18,474 case | Product distributed to medical facilities in the following US states: AL, AZ, CA, CO, CT, FL, IA, IL, KS, LA, MA, MD, MN, NC, OH, PA, PR, TN, TX, UT, and WA. Product was also distributed internationally to the following countries: Australia, Canada, Dominican Republic, Malaysia, Turkey, and United Kingdom | Open, Classified |  |  | 180049
Z-2251-2016 | 74637 | No | MOT |  |  | Radiology | 510(k) | Best Theratronics cabinet x-ray systems intended for use in the irradiation of blood and blood products (packaged in transfusion bags) to inactivate T-lymphocytes for the prevention of Graft Versus Host Disease. |  |  | May 06, 2016 | 2016 | BEST THERATRONICS LTD. | It was discovered that both the new device and the predicate device failed to comply with performance standard. | Radiation Control for Health and Safety Act | Best Theratronics ( BTL) planned action to bring into compliance:BTL proposes that the repairs are to be completed during annual routinepreventive maintenance visits at the user's site to minimize downtime, which disrupts the user's ability to irradiate blood. A standard repair kit has been developed. For further information, please contact Best Theratronics Customer Service at 1-866-792-8598 |  | 128 | 128 | Nationwide | Terminated | August 01, 2017 | 452 | 128
Z-2273-2016 | 74355 | No | CJY | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products Ca Slides, GEN 29, REF/Product Code 145 0261, Unique Device Identifier No. 10758750009114 or 20758750009111, IVD -- Sales Unit: 300 slides/pack |  |  | May 18, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) confirmed a potential for biased calcium (Ca) results to be generated when using VITROS Ca Slides, GEN29, and the VITROS 250/350/950/5,1 FS, 4600 Chemistry System or the VITROS 5600 Integrated System. An increase in negatively biased calcium results complaints prompted the investigation that identified this product issue. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an URGENT PRODUCT CORRECTION NOTIFICATION letter dated May 18, 2016 to affected customers. . The letter identified the affected product, problem and actions to be taken. The letter instructed to discontinue use of the affected lot(s) and discard all remaining inventory. Customers were asked to complete and return the Confirmation of Receipt form. For questions contact Customer Technical Services at 1-800-421-3311. |  | 32382 | US: 27,625 sales units; Foreign: 4757 sales units | Worldwide Distribution - US Nationwide in the countries of Puerto Rico and Foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, Poland and Venezuela. | Open, Classified |  |  | 32382
Z-2274-2016 | 74567 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Allegra X-5 rotor yoke, Catalog No. (Instruments) B30589 and B30590 (Rotor) B30593. Product Labeling: IFU B29071AA. |  |  | February 05, 2016 | 2016 | Beckman Coulter Inc. | A defective rotor yoke on Allegra X-5 allows improper rotor seating on the hub resulting in possible rotor failure if the buckets are loaded in an imbalanced configuration. | Nonconforming Material/Component | Beckman Coulter sent an Important Product Notice Recall letter dated February 5, 2016 to customer. The letter identified the affected product problem and actions to be taken. Customers are instructed to complete and return the enclosed Response Form. Customers with questions are instructed to contact our Customer Support Center" Via website, http://www.beckmancoulter.com/customersupport/support" Via phone, call 800-369-0333 the United States and Canada" Outside the United States and Canada, contact your local Beckman Coulter Representative. |  | 62 | 62 units total (33 units in US) | Worldwide Distribution - US Nationwide (including Puerto Rico) and the countries of: Australia, Canada, Colombia, Germany, Guatemala, Netherlands, Singapore, Tunisia, , Spain | Open, Classified |  |  | 62
Z-2236-2016 | 74554 | No | RCE |  |  |  |  | GE Inspection Technologies, LP - Phoenix v¦tome¦x m240 cabinet x-ray systems. |  |  | March 02, 2016 | 2016 | GE Inspection Technologies, LP | Radiation emissions which exceeded the limit in the performance standard were detected during maintenance of one system. The identified cause was insufficient attachment of shielding in the door. Jostling during delivery was identified as the probable cause of this failure to comply with the performance standard. | Radiation Control for Health and Safety Act | GE plans to notify its customers with the systems through a TIL. The TIL will inform customers that GE will add lead strips to the inside window frame under the lead glass window of the main cabinet door. GE recommends that it visit each customer site and repair the system by adding lead strips along the window frame of the cabinet door, beneath the lead glass window, within 90 days following FDA's approval of the repair plan. |  | 4 | 4 | US Distribution to the states of : FL, TN, and TX. | Terminated | August 07, 2017 | 523 | 4
Z-2279-2016 | 73985 | No | ODG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ultrasound Gastroscope |  |  | April 06, 2016 | 2016 | Pentax Medical Company | Pentax Medical Video Processor EPK-i5010 has been shipped with a reusable Pentax Medical water bottle assembly (OS-H4) which has not been cleared for use with this devices in the United States. | Device Design | Pentax Medical sent an "Urgent Field Correction Letter and Response Form" dated April 6, 2016 via Certified Mail to their affected customers. The recall letter identified the product, problem and action to be taken by customer. The customer is asked to complete the response form even if the customers does not have any OS-H2/OS-H4 water bottles to discard. If customer needs assistance in completing the response form, they can contact their local sales representative. For questions regarding the recall, customers can contact Mr. Paul Silva, Pentax Americas Recall Coordinator at: (800) 431-5880 ext. 2064, from (8:30 am - 5:00 pm, Monday - Friday, EST), via fax: (201) 799-4063 or via email to: recall.coordinator@pentaxmedical.com. |  | 819 | 819 units | US Nationwide Distribution | Terminated | July 17, 2017 | 467 | 23755
Z-2288-2016 | 74563 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AnastoCiip Universal Clip Remover (6 Pack} |  |  | June 22, 2016 | 2016 | LeMaitre Vascular, Inc. | US marketing brochure for the AnastoClip AC ClosureSystem states that the Anastoclip catalog numberscan be used for dura closure, but the catalog numbers are only intended for cardiovascular use. | Unknown/Undetermined by firm | The firm, LeMaitre Vascular, Inc., sent an "URGENT: Medical Device Correction" letter dated June 20, 2016 to its distributors and hospitals via Priority mail. The letter describes the product, problem and actions to be taken. The customers were informed of the intended use of the devices and the correct catalog numbers to order for the dura application. The customers were instructed to return unused devices for an exchange or refund; destroy any Anastoclip related marketing materials that you have received from Lemaitre between October 25- 2015 and June 15, 2016, and complete and return the reply form for reconciliation of this recall via fax at 781-221-2223, email: CSUS@LEMAITRE.COM or by mail. LeMaitre Vascular, Inc. will replace any affected device with a new device labeled for dura closure or issue a credit per request.Questions concerning this safety notice, please contact me at 781-221-2266 ext. 183. |  | 55 | 55 units (6 pack) | US Nationwide Distribution | Open, Classified |  |  | 1051
Z-2281-2016 | 74500 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Trident Constrained Liner Inserter/Impactor Tip, 22 mm, 28 mm, and 32 mm. Non Sterile |  |  | June 03, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Impactors do not meet the required sterility assurance level (SAL) of 10-6 when using the recommended steam sterilization method. | Process design | Stryker notified its branches/agencies of this recall by e-mail on May 26, 2016. They were asked to quarantine the affected devices. Stryker sent a recall letter and business reply form dated June 3, 2016 to their affected customers via UPS. Stericycle will be handling the returns. |  | 1864 | 1,864 units | Worldwide distribution. US nationwide (AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MS, NJ, NY, OH, PA, TX, WA, WI, WV), Sweden, South Africa, Chile, Australia, New Zealand, Canada, United Kingdom, Switzerland, Colombia, and Poland. | Open, Classified |  |  | 1864
Z-2290-2016 | 74596 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Quickie Q7 wheelchair. Model EIR4 |  |  | July 01, 2016 | 2016 | Sunrise Medical (US) LLC | There is an observed failure of the backrest bracket which can fatigue over the life time of the product. When this occurs there is the possibility of the bracket fracturingcausing the backrest to detach from the chair frame. | Device Design | The firm began calling dealers to alert them to the recall and obtain primary contacts information. Dealers were then sent their list of Quickie Q7 wheelchairs that records show being sold by them and dealers are asked to return a list of chairs they are able to locate for correction. The first dealer letter was sent via e-mail on July 7, 2016. |  | 13852 | 12,142 devices in US, 1,710 devices internationally. | US and Mexico, Japan, Chile, Kuwait, Columbia, Hong Kong, China, Uruguay, South Africa, Costa Rica, South Korea, Singapore, New Zealand, El Salvador, Bahrain, Uruguay, Lebanon, United Arab Emirates, Brazil, Argentina, Netherlands Antilles, British Virgin Islands, | Open, Classified |  |  | 13852
Z-2282-2016 | 74452 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instructions for Use (IFU) for Progesterone 3 Calibrator Set, repackaged as one set/package. |  |  | June 07, 2016 | 2016 | Tosoh Bioscience, Inc. | Instructions For Use had incorrect lot number and calibrator values | Error in labeling | Tosoh Bioscience sent an Urgent Correction Notification letter dated June 7, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions call Tosoh Technical Support at 1-800-248-6764. |  | 123 | 41 single sets of Progesterone 3 calibrators | US Nationwide Distribution in the states of AL; CA; CO; FL; KY; MA; NY; OH; SC; TX; UT; VA and WI. | Open, Classified |  |  | 123
Z-2291-2016 | 74529 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Molift Air 205, Molift Air 205 Propulsion, Molift Air IRC, Molift Air 205 IRC Propulsion, Molift Air 300, Molift Air 300 Propulsion, Molift Air 300 IRC, Molift Air 300 IRC Propulsion |  |  | June 24, 2016 | 2016 | Moller Vital | Potential for the bolt on the suspension coupling on the Molift Air to come loose from the housing. | Device Design | Moller Vital sent a Field Safety Corrective Action on June 22, 2016. Field correction instructions were provided to affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to read the instruction, and return the recipient confirmation, report all serial numbers distributed in their region/country to Etac Supply Gjovik, order the upgrade kit from Etac Supply Gjovik, Customers were also instructed to follow up with the customers and users and confirm that all hoists in their market are upgraded. Customers with questions were instructed to contact Etac Supply Gjovik.at molift.qa@etac.com. |  | 516 | 516 units (374 US) | Worldwide Distribution - US (nationwide) and Internationally to SZ, TR, BE, CH, DE, DK, ES, FI, FR, GB, IS LV, NL, NO, SE, SI, AU, CN, IL, JP, NZ, IN, QA, PE, AE | Terminated | June 23, 2017 | 364 | 516
Z-2292-2016 | 74693 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Avenir MÂżller Stem 4 Standard |  |  | June 30, 2016 | 2016 | Zimmer Gmbh | Product mix-up. The Avenir MÂżller Stem 6 lateral uncemented might be placed in the packaging of the Avenir MÂżller Stem 4 standard uncemented and vice versa.Potential for delay in surgery, and/or potential for use of incorrect implant which may in turn lead dismetry in patient and possibly the need for revision. | Labeling mix-ups | Zimmer Biomet sent an Urgent Medical Device Recall letter dated June 30, 2016 to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to review the notification immediately and ensure affected personnel are aware of the contents and assist Zimmer Biomet sales representative with the quarantine of any affected device. Your Zimmer Biomet sales representative will remove any affected device. Customers were also instructed to complete the Certification of Acknowledgement form and return to corporatequality.postmarket@zimmerbiomet.com. Customers with further questions or concerns were instructed to contact the customer call center at 1-800-348-2759 between 8:00 am and 8:00 pm EST.For questions regarding this recall call 800-348-2759. |  | 2 | 2 | Nationwide Distribution to OH and NM | Terminated | March 27, 2017 | 270 | 2
Z-2294-2016 | 74222 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Endotine Transbleph 3.5 (CFD-080-0167) |  |  | May 06, 2016 | 2016 | MicroAire Surgical Instruments, LLC | MicroAire has initiated a limited recall (29 Units)on the Endotine Transbleph 3.5 (CFD-080-0167), Lot # 285373, due to minor mislabeling ( the front label is correct, but the side label lists an incorrect part number, i.e. CFD-080-0177). | Labeling Change Control | The firm is sending recall letters to the customers who have received the Endotine Transbleph 3.5 (CFD-080-0167), lot #285373. Letters sent by UPS to expedite the delivery. |  | 29 | 29 | The product was distributed in the United States and Sweden | Completed |  |  | 29
Z-2293-2016 | 74668 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | July 15, 2016 | 2016 | Elekta, Inc. | It is possible that a change to an Order Set will not be saved in the current open Care Plan. This can result in a wrong chemo dose in the order. | Software Design Change | Elekta sent Important Field Safety Notice 371-03-MSQ-001 to all customers on 7/15/2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. The letter also included a response form which is to be returned to Elekta. |  | 383 | 383 | Worldwide distribution. US nationwide, Antigua and Barbuda, Australia, Austria, Bahamas, Bahrain, Canada, Czech Republic, Germany, Hungary, South Africa, Spain, Sweden, Turkey, and United Kingdom | Open, Classified |  |  | 383
Z-2296-2016 | 74583 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Integra External Fixation System Rocker Bottom Model # 12225400. |  |  | July 01, 2016 | 2016 | Integra LifeSciences Corp. | The silicone insert embedded in the rocker bottom has the potential to become dislodged or loose from the rocker bottom mold. | Device Design | All affected consignees were notified via an "Urgent: Voluntary Medical Device Recall" letter sent on 7/1/16. The letter identified the affected product as well as the reason for the recall. Customers were asked to check their traceability records for shipments of the affected product. If the affected parts are found, customers are to stop use and distribution. A copy of the letter should be provided to any customers that the product may have been distributed to. The attached "Acknowledgement and Return Form" should be completed and returned regardless if customers have affected product or not. Questions regarding the instructions should be directed to Customer Service at 1-800-654-2873. |  | 27 | 27 units | Distributed in the US to the states of FL, VA, OK, CO, DC, and IN. | Open, Classified |  |  | 27
Z-2298-2016 | 74524 | No | KLS | Theophylline test system. | Toxicology | Toxicology | 510(k) | SyvaÂż EMITÂż 2000 Theophylline - Catalog # 4P019UL /SMN#10445324 is a homogeneous enzyme immunoassay intended for use in the quantitative analysis of theophylline in human serum or plasma. |  |  | June 17, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed the SyvaÂż EMITÂż 2000 Theophylline reagent lot G3, when run ona Beckman AU Clinical Chemistry System, and the ADVIAÂż Chemistry Theophylline_2 (THEO_2) reagent lot334018, may exhibit increased imprecision for commercially available Bio-Rad Liquichek TDM Quality ControlLevels and patient samples. | Process control | Siemens Healthcare Diagnostics mailed a letter to customers on June 17, 2016, asking that they discontinue use of and discard the affected kit lots identified in the recall notification letter. Customers were also asked to complete and return the Field Correction Effectiveness Check/Product Replacement Form that was mailed with the Recall Notification Letter. |  | 1004 | 1,004 units | Distributed to: CA, CT, DE, FL, GA, IA, IN, KS, KY, LA, MA, MD, MI, MO, NC, NJ, NM, NY, OH, OK, PA, TN, VA, WA | Terminated | March 15, 2017 | 271 | 1662
Z-2300-2016 | 64758 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Imagecast PACS with Centricity RIS-IC versions prior to 10.6 Update Package 18 |  |  | February 18, 2013 | 2013 | Ge Healthcare It | A software defect was discovered that causes images to be out of context with clinical information. | Software design | GE sent an "Urgent Medical Device Correction" letter to all affected customers on April 2, 2013.The letter identified the product the problem and the action customers need to take.The software has been modified to correct these defects. The modified software will be released and installed in user sites through GE Healthcare FMI #85204.For further questions, please call ( 847) 277-4468. |  | 13 | 13 | US Nationwide Distribution. | Open, Classified |  |  | 13
Z-2295-2016 | 74523 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Babytherm 8004/8010, Infant Warmer, Class II device, 510 K K971198, Catalog # FR00105 (Babytherm 8010), 2M30404 (Babytherm 8004), non sterile, software controlled. |  |  | June 22, 2016 | 2016 | Draeger Medical, Inc. | Risk associated with operating the device in skin temperature mode with temperature sensors attached but no patient connected. | Nonconforming Material/Component | Drager mailed to customers an Urgent Medical Device Recall Notification Letter dated June 22, 2016 to affected customers. The letter identified the affected product, problem information on how to operate the device and replacement of parts information. For questions regarding this letter please call 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your DrÂżger Babytherm 8004/8010 please contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 1599 | 1,599 units | US Nationwide Distribution including Puerto Rico and the states of AK,AL,AR,CA,CO,CT,DE,FL,GA,HI,IA,IL,IN,KS,KY,LA,MA,MD,MI,MN,MO,MS,NC,ND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,RI,SC,TN,TS,TX,UT,VA,WA,WI,WV, and WY. | Open, Classified |  |  | 1599
Z-2297-2016 | 74617 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SPHERICAL HEAD REFLECTION(R) 6.5 MM CANCELLOUS SCREW, REF 71332525, QTY: (1), STERILE R, smith&nephew; |  |  | July 06, 2016 | 2016 | Smith &amp; Nephew, Inc. | The labels correctly state that the product should be a 25mm screw, but the box contained a longer, 35mm screw. | Packaging process control | Smith &amp; Nephew sent an Urgent Product Recall Notification letter dated July 6, 2016, to all affected consignees via email and FedEx Overnight delivery on July 6, 2016. The letter identified the product, the problem, and the action to be taken by the consignee. Customers were asked to inspect their inventory and located any unused affected devices and quarantine them immediately, complete the Inventory Return Certification Form and fax to 1-901-566-7975 indicating the quantities that need to be returned. Consignees should contact Smith &amp; Nephew's Field Actions Department via email at FieldActions@smith-nephew.com to obtain a return authorization (RA) number.For questions regarding this recall call 901-399-5520. |  | 185 | 185 units | Worldwide Distribution - US (nationwide) and Internationally to AU, CA, GB, IN, ZA, and TR | Terminated | July 01, 2017 | 360 | 185
Z-2302-2016 | 74366 | No | KLI | Total thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products Total T4 Reagent Pack, REF/Product Code 874 4468, Unique Identifier No. 10758750005147, IVD --- Ortho-Clinical Diagnostics, |  |  | June 02, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | Ortho-Clinical Diagnostics confirmed that sample results generated from VITROS Immunodiagnostic Products Total T4 Reagent Packs may show a negative bias prior to the open pack storage limit of &lt;= 8 weeks that is located in the Instructions for Use (IFU) for VITROS Immunodiagnostic Products Total T4 Reagent Pack. | Under Investigation by firm | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter (Ref. CL2016-112) dated June 2, 2016, to all affected customers via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received the affected product to inform them that sample results generated using VITROS Total T4 Reagent Packs that are opened for more than 4 weeks may produce a negative bias and advised them to only use the Total T4 Reagent Pack up to 4 weeks after the pack has been opened. Foreign affiliates were informed of the issue by e-mail on 6June 2, 2016, and instructed to notify their consignees of the issue and required actions. Customers with questions were instructed to contact Global Technical Solutions Center at 1-800-421-3311. |  | 278 | US: 12,015; Foreign: 55 Sales Units to Canada, 223 Sales Units shipped to Bermuda | Worldwide Distribution - US (Nationwide including Puerto Rico) and foreign distribution to Canada and Bermuda | Open, Classified |  |  | 278
Z-2303-2016 | 74651 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Hartland, Wisconsin. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | A migrated study that has annotations will display the annotations, but when a prior study is loaded that should have annotations, the second study will not display annotations. | Software design | The recalling firm, Merge, issued an "URGENT: MEDICAL DEVICE RECALL" letter dated 1/27/2016 via e-mail on 1/30/2016 or via certified mail if they did not have an e-mail address for the customers. The letter described the product, problem and actions to be taken. The customers were instructed to not discontinue use of Merge PACS; to open prior comparison studies in the secondary viewer to check if there were annotations and read the report or visit note pertaining to that study to see if there are noted findings that were not otherwise apparent; to receive an upgrade call Merge Customer Service at (877) 741-5369 or send an email to support@merge.com to schedule the upgrade (Reference this call #2016-010); complete the online form: http://www.merge.com/Landing-Pages/Customer-Upgrade-Decline-Form.aspx no later than 2/22/2016; If you decide to decline the maintenance release, complete the attached Customer Decline Response form and return to Merge Healthcare via recall@merge.com, and ensure all users of the product are provided with this notification. If you have further distributed this product, identify your customers and notify them at once of this product recall.If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 208 | 208 sites have the affected versions | Worldwide Distribution: US (Nationwide) and countries of: Australia, Belgium, Canada, Jordan, New Zealand, and the United Kingdom. | Open, Classified |  |  | 208
Z-2255-2016 | 74593 | No | RBY |  |  |  |  | Olympus Scientific Solutions Americas Corporation ( OSSA ) Model: Analytical X-Ray System X -Ray Fluorescence. |  |  | June 22, 2016 | 2016 | Olympus Scientific Solutions Americas | Trigger lock option does not always prevent x-ray generation. Specifically if the unit is in test mode or repeat test mode then the trigger lock may not prevent x-ray initiation when the trigger is pulled. The correction is via repair by a software update. | Radiation Control for Health and Safety Act | The firm notified customers via an "Important - Electronic Product Radiation Warning" letter. It described the issue; and, told customers to not rely on the Trigger Lock feature until the software upgrade had been completed. The letter also provided instructions for customers to not only determine their software version; but, also instructions to bring their affected device into compliance. The firm will remedy the issue without charge. Customers can contact their local representative or OSSA at ani.technicalsupport@olympus-ossa.com, phone: +1 866 446 6689, fax: +1 781 419 3980, with questions or if additional information is needed. |  | 212 | 2, 210 units distributed in the USA | Nationwide Distribution. | Open, Classified |  |  | 212
Z-2301-2016 | 74648 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Studies coming over via telmed were missing patient's DOB, procedure, and referring physician. | Software design | The recalling firm issued Urgent: Medical Device Recall letters dated 1/27/2016 and 5/26/2016. The letter informed the customer of the affected product, the nature of the issue, the potential harm involved, and actions to be taken. The firm released a fix for the issue; and, customers were asked to complete and return the enclosed form by June 17, 2016. Questions can be directed to recall@merge.com. |  | 109 | 109 sites have the affected versions | Worldwide Distribution -- USA, New Zealand, and United Kingdom. | Open, Classified |  |  | 109
Z-2304-2016 | 74312 | No | LEH | Vancomycin test system. | Toxicology | Toxicology | 510(k) | ONLINE TDM Vancomycin 100 tests; ONLINE TDM Vancomycin 200 tests; Hitachi Vancomycin (Modular P); Hitachi Vancomycin (917, MOD P) |  |  | May 26, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Roche Diagnostics Operations, Inc. has issued a voluntary recall for the Online TDM Vancomycin assay on the cobas c 311/501/502 analyzers and the Modular Analytics P module due method sheets that state an incorrect method comparison against the COBAS INTEGRA 800 analyzer. | Under Investigation by firm | Roche sent an Urgent Medical Device Correction letter dated May 26, 2016, to all affected customers via UPS Ground (receipt required) . The letter identified the product the problem and the action needed to be taken by the customer.Customers will be informed to disregard the method comparison data in the package insert and to follow the instructions outlined in UMDC 16-083. Non-responding accounts will be defined as users for which confirmation of receipt of the Urgent Medical Device Correction cannot be obtained. These accounts will be monitored on an ongoing basis and follow-up attempts will bemade accordingly. Affected product under Roche control was placed on hold. Affected product will be reworked to add a sticker referring to customers to UMDC 16-083. Roche is not requesting return of the affected product. Any returned product will be managed according to current procedures. The incorrect method comparison will be removed from the method sheets. Updated method sheets will be released at a later time. Customers may contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336. |  | 53897 | 53,897 Units distributed | Nationwide Distribution to the states of : AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN,TX, UT, VA, VT, WA, WI, WV, WYForeign:NoneVA/DOD: See list below | Terminated | March 23, 2017 | 301 | 53897
Z-2312-2016 | 74477 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | 3T 16ch Flex SPEEDER Coil, Large, NeoCoil Part Number NC046200, Toshiba Medical Systems Model MJAJ-222A. The NeoCoil 3T 16ch Flex SPEEDER Coils are tuned to receive RF frequency corresponding to the proton precession in a 3 tesla magnetic field, which is governed by the Larmor equation. |  |  | June 17, 2016 | 2016 | NeoCoil, LLC | There is a potential for higher than specified surface temperatures involving the 3T 16ch Flex SPEEDER Large Coil. | Nonconforming Material/Component | NeoCoil sent an "Urgent Medical Device Recall" letter dated 17 June 2016. The letter described the problem and the product involved in the recall. Advised consignee to quarantine the product and if the product was further distributed to identify the customers and provide the list to NeoCoil. The recall will be carried to the user level. Requested consignee to complete and return the response form. For questions call 262-522-6127. |  | 39 | 39 (11 US, 28 OUS) | Worldwide Distribution - US and Internationally to Georgia, Germany, Korea, Netherlands, Pakistan, Spain, Japan, and the United Kingdom. | Open, Classified |  |  | 39
Z-2311-2016 | 74600 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | AmediCheck Instant Test Cup, Part No. C-C0501A-3A |  |  | July 05, 2016 | 2016 | Branan Medical Corporation | Branan Medical is recalling the AmediCheck Instant Test Cup, a 5 panel drug cup, because it may erroneously contain 12 panel cups due to a packaging error. | Labeling mix-ups | BMC sent an Urgent Medical Device Recall Notice letter on July 6, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the form even if they do not have affected product. The letter informs the customers that replacement product will be provided free of charge for any unused portion of the affected lot based on completion and return of the attached Customer Verification Form. Customers are instructed to return the completed Customer Verification within 10 business days of receiptby FAX at 949-598-7166, or by email at branan.complaints@brananmedical.com. Customers with anyquestions about the information contained in the notification are instructed to contact Branan Medical Corporation at 949-598-7166 or by e-mail atbranan.complaints@brananmedical.com. |  | 96 | 96 units | US Nationwide distribution in the states of: TN, NC, FL, and PA | Open, Classified |  |  | 96
Z-2307-2016 | 74731 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | MultiDiagnost Eleva; x-ray system. Product Codes: 708032, 708036 |  |  | July 19, 2016 | 2016 | Philips Electronics North America Corporation | The two bolts that attach the clamping plate to the lateral shaft may break, which could cause the c-arm to move laterally forward, slide out of its bearings, and sag | Device Design | On July 19, 2016, Philips sent a Field Safety Notice to all affected customers. Philips Field Service Engineers will add an independent safety catch free of charge. |  | 993 | 993 total | Worldwide distribution. US nationwide, Canada, Algeria, ARGENTINA, Australia, Austria, Belgium, China, Colombia, Czech Republic, Denmark, Egypt, Estonia, Faroe Islands, Finland, France, French Guiana, Germany, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Jersey, Jordan, Republic of Korea, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Philippines, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, TURKEY, Ukraine, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 993
Z-2308-2016 | 74568 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU480 Chemistry Analyzers, Reference No. B28624 |  |  | July 30, 2015 | 2015 | Beckman Coulter Inc. | Beckman coulter has received complaints regarding leaking and cracked C cups. When the cup cracks, coolant can leak onto the connectors of the Printed Circuit Boards (PCB) in the analyzer, resulting in a short-circuit causing the instrument to go into a stop mode. | Nonconforming Material/Component | In March 2011, the material and molding process of the Coolant cup (C cup) was changed. BEC did not issue a customer notification letter, as we do not believe that a notificationletter was either warranted or of any potential use or value to our customers, because the coolant cup ( C cup ) is an internal component that only a BEC Field Service Engineer (FSE) can access and inspect  i.e., the C cup is not in a customer-accessible part of theanalyzer. Thus, there are no specific actions the customer could take to inspect, identify, or mitigate the potential for a crack or leakage of the C cup. For questions regarding this recall call 714-961-4933. |  | 3355 | 3,355 units total (617 units in US) | Worldwide Distribution - US (nationwide (including Guam and Puerto Rico), and Internationally to Albania, Algeria, Australia, Austria, Bangladesh, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Botswana, Brazil, Bulgaria, Canada, China, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Eritrea, Ethiopia, Finland, France, Germany, Ghana, Gibraltar, Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Lebanon, Liberia, Lithuania, Malawi, Malaysia, Mexico, Morocco, Mozambique, Myanmar, Namibia, Nepal, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, United States Uruguay, Viet Nam, Zambia, Zimbabwe. | Open, Classified |  |  | 3355
Z-2200-2016 | 74424 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, Temporary Pacing Lead System, Model 6416, Sterile EO. |  |  | June 16, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | This Pacing Lead System does not comply with Section 8.5.2.3 of IEC 60601-1, and corresponding provisions of FDA 21 CFR 898, which relate to design standards to prevent connecting a patients lead to a possible hazardous voltage. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter dated June 2016 to affected customers. The letter was addressed to Health Care Professional. The letter described the problem and the product involved in the recall. The letter described the "Customer Actions" which advised consignees to identify, quarantine and return affected product to Medtronic. Customers were instructed to complete the enclosed Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality or scan and e-mail to RS.CFQFCA@medtronic.com. For questions, please contact your Medtronic CRHF Sales Representative. |  | 11697 | 11697 (5523 OUS, 6174 US) | Worldwide Distribution. US nationwide, Albania, Algeria, Australia, Botswana, Bulgaria, Canada, Colombia, France, Gabon, Germany, Guam, Hong Kong, Hungary, Ireland, Israel, Italy, Jordan, Kazakhstan, Macedonia, Mauritius, Netherlands, Netherlands Antilles, Norway, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Serbia, Slovenia, South Africa, Spain, St Maarten, Sweden, Switzerland, Tajikistan, Turkey, Uganda, United Arab Emirates, United Kingdom, and US Virgin Islands. | Open, Classified |  |  | 11697
Z-2323-2016 | 74548 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | V-TEK TWIST D 3.5X90MM 30MM Can 12MM Rnd; The twist drill is used to drill holes into bones/tissues for foot, ankle and hand implant systems. |  |  | June 27, 2016 | 2016 | Zimmer Gmbh | Zimmer Biomet initiated a voluntary recall of the Twist Drill &amp; V-Tek Twist Drill due to to a potential tolerance stack up identified between the twist drill and the mating Kirschner Wire (K-Wire) component. | Under Investigation by firm | Zimmer GmbH initiated a voluntary recall of the Normed - TWIST DRILL Cannulated &amp; Normed- V-TEK TWIST DRILL Cannulated, with written notices on June 27, 2016 via certified mail due to a potential tolerance stack up identified between the twist drill and the mating Kirschner Wire (K-Wire) component. Customers are asked to do the following: Review the notification and ensure affected personnel are aware of the contents.2. Locate all affected product identified above and quarantine them immediately.3. Return any affected product within your possession. Clearly mark the outside of all return packages, Recall, and include a copy of the Inventory Return Certification form (Attachment 1) with your returnshipment(s).4. If after reviewing this notification you have further questions or concerns please call the customercall center at 1-800-348-2759 between 8:00 am and 8:00 pm EST.Adverse reactions or quality problems experienced with the use of this product may be reported to theFDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax." Online:www.fda.gov/medwatch/report.htm" Regular Mail: use postage-paid, pre-addressed Form FDA 3500 available at: www.fda.gov/MedWatch/getforms.htm.Mail to address on the pre-addressed form." Fax: 1-800-FDA-0178Affected units returned as part of this recall activity will be scrapped |  | 169 | 169 | Domestic: CO, ILForeign: Argentina, Costa Rica, India, Malta, Japan, Philippines, Vietnam, Portugal, Dominican Republic, Argentina, Australia, Russia, Saudi Arabia, Ireland, Ecuador, China, Singapore, Indonesia, Canada, Mexico, South Korea , Spain, Netherlands, Sweden, Malaysia, New Zealand, Switzerland, El Salvador, England, Norway, Finland, Venezuela, Taiwan, Italy, Colombia, Greece, Thailand, France, Israel, Jamaica, Japan, Jordan, Germany, Korea, Lebanon, Belgium, Hong Kong, Philippines, Iran, Brazil, Turkey, United Arab Emirates (UAE), Uruguay | Terminated | May 30, 2017 | 337 | 169
Z-2316-2016 | 74715 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Synergy Spine. The software application is sent in CD format with an IFU, wrapped in plastic with a label for shipping purposes. |  |  | July 18, 2016 | 2016 | Medtronic Navigation, Inc. | Medtronic Navigation, Inc. announces a voluntary field action for the Medtronic Navigation StealthStation Software applications affected by Neurologica BodyTom/CereTom floor-based scanners. | Software design | Medtronic sent an Urgent Medical Device Safety Alert to all affected customers on August 20, 2016. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached confirmation form and return to RS.NavFCA@medtronic.com or fax it to Medtronic Technical Service at 651-367-7075. For questions contact Neurologica Technical Support at 1-888-564-8561 (US/Canada) or +1-978-564-8561 (International) or email support@neurologica.com |  | 27 | 27 in total | Worldwide Distribution - US Nationwide in the state of: NY, AZ, AR, NC, CA, OH, MI, IN, MA, TX, OH, OR, CO, MO, UT, IL. and the countries of: India, Australia, and Russian Federation. | Open, Classified |  |  | 27
Z-2320-2016 | 74624 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | RePlantÂż Angled Abutment |  |  | July 05, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | Issue with design specification; oversized condition. | Process control | Implant Direct sent an Urgent: Medical Device Field Corrective Action letter dated July 6, 2016, to all affected customers. The letter requests that it be reviewed and determined if customer has any affected product in inventory. Customers are asked to complete and return the Acknowledgement and Recall Return Form within 48 hours. If customers are an authorized Implant Direct Sybron Manufacturing distributor, they request that customers identify those customers that may have been shipped the affected product lot and contact these customers to inform them of this issue within fortyeight (48) hours of receipt of the notification in order to provide the customers with the correct tool.If customers have any of the affected product listed the firm will send a replacement part. If any questions contact Implant Direct Sybron Manufacturing LLC Customer Care at 1-888-649-6425. |  | 205 | 205 | U.S. distribution to the following; CA, WA, UT, GA, AZ, NE, OR, MI, TX, WI, FL, AL, AR, MO, VA, MA, TN, LA, NC, KS, NY, SC, OK.Foreign distribution to the following; CA, DE, IT, GB, PL, DE, GB, NL, GR, ES, HR, BG. | Open, Classified |  |  | 205
Z-2325-2016 | 74691 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | BK Medical ApS Ultrasound System Scanner bk2300 - |  |  | June 28, 2016 | 2016 | B-K Medical A/S | In the system (bk3500) software generic volume measurement (HxWxLxl) function the user may beunaware of the calculation factor (1) and assume the generic volume calculation factor is similar to that ofthe other (prostate, adenoma, testis, kidney, and bladder) volume calculation factors provided in thesoftware. | Software design | BK Medical sent an Important Corrective Action Notification letter dated June 27, 2016, to all affected customers. The letter explains the issue and asks customers to discontinue the use of the generic volume measurement and contact BK medical ApS Service for assistance. Customers with questions were instructed to call 800-645-4442. |  | 6 | 6 | Worldwide Distribution - US to TX, FL, and MA. Internationally to Australia | Open, Classified |  |  | 6
Z-2329-2016 | 74599 | No | MGB |  |  | Cardiovascular | PMA | 6F/7F MynxGrip Vascular Closure Device; |  |  | July 13, 2016 | 2016 | AccessClosure, Inc., A Cardinal Health Company | A total of 70 6F/7F Mynx Grip devices were shipped from a distributor using improper transport/packaging configurations. | Process control | Cardinal Health sent an Urgent Medical Device letters dated July 15, 2016 to affected customers. The letters provided the reason for recall and the exact lot number and quantity shipped to each location. Customers were instructed to quarantined affected products and complete and return the Field Action response form indicating the number of devices that will be returned. Instructions for returning the devices are included with an RMA number and Fed Ex airbill. Questions should be directed to Customer Advocacy at 408-610-6570 between 9 am and 5 pm PST. |  | 55 | 55 units | US in the states of CA, and MS. | Terminated | August 15, 2016 | 33 | 55
Z-2333-2016 | 74610 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | SureStep Tray, with Add-a-Foley, COMPLETE CARE, Drainage Bag, and STATLOCK Stabilization |  |  | June 29, 2016 | 2016 | C.R. Bard, Inc. | Labeling problem | Labeling mix-ups | Bard Medical sent an Urgent Medical Device Product Recall letter dated June 29, 2016. The letter identified the product, the problem, and the action to be taken by the customer. The letter requested that customers examine their inventory and quarantine the product. The letter also requested a sub-recall. The letter included a reply form which is to be returned to Bard. Customers were also instructed if product was further distributed they identify customers and notify them at once. Customers with questions should contact the BMD Recall Coordinator at 1-800-793-8110. |  | 530 | 530 units | Nationwide Distribution to CA, FL, IL, IA, KY, NC, MN, NC, OR, RI, SD, TX, UT, and WA. | Open, Classified |  |  | 530
Z-2327-2016 | 74669 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Horizon 24 Large Titanium Clips with Adhesive Backing, Ref 004204, Rx Only, Sterile. |  |  | August 08, 2013 | 2013 | Teleflex Medical | The manufacturing and expiration dates are not printed on the blister pack label. | Error in labeling | Teleflex sent an Urgent Medical Device Recall Notification dated August 8, 2013, to all affected customers. The letter requested that they quarantine the product and to return it to Teleflex. An acknowledgment form was included with the letter which also was to be returned to Teleflex. A separate letter was sent to distributors requesting they conduct a sub-recall.If you have any further questions, please call (610) 378-0131. |  | 8220 | 8,220 ea. | USA Distribution to the states of :AL, AZ, CA, CT, FL, IL, KY, MD, MI, MN, NM, NY, NC, OR, PA, SC, TX, VA, WA, WI, and Internationally to Canada | Completed |  |  | 8220
Z-2328-2016 | 74751 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IntelliVue Patient Monitor |  |  | July 26, 2016 | 2016 | Philips Electronics North America Corporation | If an affected Patient Monitor has been powered on continuously for several months, any displayed waveforms will contain outdated data and therefore fail to reflect the patients current condition. | Nonconforming Material/Component | Philips sent an Urgent Medical Device Correction letter dated July 20, 2016 to all affected customers. The Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Customers were informed that Philips will provide a software correction free of charge. For any further information or support concerning this issue, contact your local Philips representative at 1-800-722-9377. |  | 50495 | 50,495 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 50495
Z-2332-2016 | 74560 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Spirit Select Bed, A-C Powered Hospital Bed |  |  | June 28, 2016 | 2016 | CHG Hospital Beds Inc | Stryker Medical is initiating a voluntary recall of the Spirit Select and Spirit Plus A-C powered hospital beds due to reports of hi-lo actuators broken at the mount ends, which could cause the bed to unexpectedly lower resulting in patient injury. | Component design/selection | The firm, Stryker Medical, issued an "URGENT MEDICAL DEVICE RECALL" letter dated 7/8/2016 by FedEx to its customers and included a revised preventive maintenance checklist. The letter described the product, problem and actions to be taken. The customers were instructed to: locate the units listed on the attached business reply form; remove these units from service, if not possible to remove units from service, place units in the height position according to instructions; file the revised preventative maintenance checklist; return the Business Reply Form to confirm receipt of the notification by fax (269)488-8691 or email productfieldaction@stryker.com. If customers have loaned or sold any of the beds listed in this letter, please forward a copy of the recall notice to the new users and advise us of their new location in the space provided on the business reply form. Your Stryker Field service Representative will contact your facility to add support brackets to your beds. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax."Online: www.fda.gov/medwatch/report.htm "Regular Mail: use postage-paid FDA form 3500 available at: www.fda.gov/MedWatch/getforms.htm. Mail to MedWatch , P.O. Box 3002, Rockville, MD 20847-3002 "Fax: 1-800-FDA-0178 If you have any questions or concerns, please contact Stryker Customer Service (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 4308 | 4,308 units total | Worldwide Distribution: US (nationwide) including states of: AK , AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, and WI; and country of: Canada | Terminated | June 08, 2017 | 345 | 4308
Z-2330-2016 | 74608 | [6 empty] | EnVision FLEX+, Mouse, High pH (Link), Product No. K8002, K8004. |  |  | July 06, 2016 | 2016 | Dako Denmark A/S | The pH of Target Retrieval Solution in the kit measures 4-5 instead of the correct pH of 9.1. | Process control | A recall notification letter dated 7/6/16 was sent to customers to inform them that Dako is recalling EnVision FLEX+ kit because the pH of Target Retrieval Solution in the kit measures 4-5 instead of the correct pH of 9.1. The letter informs the customers that with a pH of 4-5 instead of pH 9 there could be weak staining, false negative or false positive non-specific staining depending on the antibody used. There is the possibility that this staining could lead to a false result. Customers are to discard the affected product, determine the impact as instructed, and confirmed that they have received replacement kits. Customers are instructed to complete the Device Recall Form and contact local sales representative for any questions regarding the recall notification. |  | 147 | 147 kits | Worldwide Distribution -- US, Canada, and Colombia. | Open, Classified |  |  | 147
Z-2326-2016 | 74565 | No | FTR | Silicone gel-filled breast prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | Mentor MemoryGel Breast Implant; Indicated for females for breast augmentation and breast reconstruction. |  |  | July 01, 2016 | 2016 | Mentor Texas, LP. | The box of 300cc MemoryGel Breast Implant was labeled with null manufacturing and expiration dates. | Labeling mix-ups | The affected consignees were notified via customer letter on 7/1/16, and a follow-up with respective sales representative via visit. |  | 35 | 35 units | Distributed in the states of NY, WI, CA, TX, CO, NC, NJ, OH, TN, RI, and IL. | Open, Classified |  |  | 35
Z-2341-2016 | 74577 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge HEMO software. |  |  | July 02, 2012 | 2012 | Merge Healthcare, Inc. | The International Normalized Ratio (INR) value displayed in the study report does not match the value that is imported on the pre-procedure labs screen from HL7 | Software design | The firm issued an "Important Product Information" bulletin to customers dated 7/2/2012 via e-mail letting them know there was an upgrade available. On 1/30/2016, the recalling firm initiated phone calls to their customers letting them know there was an upgrade available and if they declined it, they would like a written response. |  | 37 | 37 | The medical facilities who currently have the affected versions are located nationwide. There was government distribution but no military or foreign distribution. | Terminated | June 26, 2017 | 1820 | 37
Z-2343-2016 | 74549 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | 91496 Ultraview SL Command Module, Options A, B, C, and I. |  |  | July 01, 2016 | 2016 | Spacelabs Healthcare Inc | The firm has received multiple reports of the Non-Invasive Blood Pressure (NIBP) parameter becoming non-functional with an associated message (No Reading). In addition, there is one of the following three error messages: Inflate Error, HW Error, or No Data. | Vendor change control | Spacelabs expanded the recall and sent an Urgent - Medical Device Correction letter dated August 30, 2016, to additional affected customers via priority mail (return receipt requested). On September 6, 2016, a customer letter (translated as necessary) was sent via email to all international subsidiaries and distributors of record.**********************************************************************************************Spacelabs Healthcare Inc, sent an Urgent - Medical Device Correction letter dated July 1, 2016, to all affected customers (return receipt requested). On July 8, 2016, a customer letter (translated as necessary) was sent via email to all international subsidiaries and distributors of record.Spacelabs Healthcare will contact customers to schedule a convenient time for a Spacelabs service technician to update their affected Command Module(s) at no cost.For additional information or technical assistance, please contact:Technical Support 1-800-522-7025 and select (2) for Technical Support |  | 5580 | 3801 modules and 116 PCBAs &amp; valves distributed in the US; 1297 modules and 366 PCBAs &amp; valves distributed outside the US | Worldwide Distribution - USA Distribution and to the countries of :AFGHANISTAN, ARGENTINA, AUSTRALIA, BAHRAIN, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM; CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DOMINICAN REPUBLIC; ECUADOR, EGYPT, ESTONIA, FRANCE, GERMANY, GUATEMALA, HONG KONG, INDIA, IRAN (ISLAMIC REPUBLIC OF), ISRAEL, ITALY, JORDAN, KUWAIT, MEXICO, MOROCCO, NETHERLANDS, NICARAGUA, PAKISTAN, PALESTINE, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, SAUDI ARABIA, Slovakia, SPAIN, SRI LANKA, SWITZERLAND, TAIWAN, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VIET NAM. | Terminated | May 22, 2017 | 325 | 5580
Z-2215-2016 | 74469 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Oxylog 3000, Disposable Pediatric Patient Circuit, Catalog Number: 5704964 |  |  | May 31, 2016 | 2016 | Draeger Medical, Inc. | Isolated episodes of leakage at the check valve were found | Device Design | Drager sent an Urgent Medical Device Recall letter dated May 31, 2016 and Customer Reply and Order Card to all customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to inspect stock and dispose of affected products. For questions regarding this letter call 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your DrÂżger device/accessory contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 42922 | 42,922 units worldwide (US distribution 1,530) | US Nationwide Distribution including states of: AK, CA, FL, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NY,OH,OK,OR,PA,TN,TX,UT,VA,WA, and WI. | Open, Classified |  |  | 42922
Z-2337-2016 | 74636 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ZimmerÂż Air Dermatome II Width Plate, 2 inch. Part # 885120200. Width plate is packaged inside of a padded-foam corrugate box. A product label is applied to the lid of the box. The width plates, when sold individually, are placed in a labeled poly zipper bag, which is then placed into a corrugated shipper carton or padded envelope. |  |  | July 04, 2016 | 2016 | Zimmer Surgical Inc | The surface coating applied to the device is blistering, peeling and discoloring after usage over time | Process design | The firm, Zimmer Biomet Surgical, sent an "URGENT MEDICAL DEVICE RECALL CORRECTION NOTICE" dated 7/6/2016 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to do the following:1. Carefully review this letter and ensure all users of the Zimmer Air Dermatome II at your facility have been informed of this notice.2. Inspect the Zimmer Air Dermatome II devices within your facility and indicate on the attached Certificate of Acknowledgment the condition of the device. Complete and sign the attached Certificate of Acknowledgment and return it to corporatequality.postmarket@zimmerbiomet.com.3. Before each use, continue to examine the surface coating for blistering/peeling. If the blistering/peeling characteristics are notpresent, you may continue to use. Immediately after each use, re-examine the surface coating for blistering/peeling. If the blistering/peeling characteristics are present, discontinue use. In the event that alternate devices or therapies are not available, surgeons may, at their discretion, determine the use of the device as a medical necessity and continue to use. After use, follow your facilitys wound irrigation protocol or Zimmer Biomet recommends copious irrigation with a normal saline solution or immersion of the graft in a bowl of normal saline. Zimmer Biomet recommends irrigating the donor site with saline as well. Carefully examine the skin graft and donor site for gray coating flakes and continue to irrigate as necessary. Report this as an adverse event to Zimmer Biomet Surgical at SurgicalRegulatoryReporting@zimmerbiomet.com.If after reviewing this notification you have further questions or concerns please call 1-330-364-0989 between 8:00 am and 5:00 pm EST, Monday through Friday. |  | 90 | 90 units | US Nationwide Distribution including states of: AL, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, and WI. | Terminated | June 21, 2017 | 352 | 1270
Z-2346-2016 | 74402 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice whole body computed tomography X-ray system. Imaging diagnostic tool. |  |  | April 20, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The following issues have been found in MX 16-slice systems with software version 1.1.4.21426:1. During the filming operation on MX16-slice console software, the clipboard used for copying and pasting images is not cleared between patients. If the operator fails to copy the current patients images before pasting, a previous patients images may be present in the clipboard and be copied onto the film of the current patient. 2. During the Bolus Tracking scan, if the Auto Voice in Tracker scan is enabled, the Tracker scan will be aborted unexpectedly and the diagnostic scan after the tracker scan needs to be manually started.3. If the scan protocol with SAS (Spiral Auto Start) function is selected to plan the scan, the SAS option may not be displayed on the Contrast tab of the scan protocol parameter area or the SAS option is displayed but not enabled as pre-configured in the protocol. This issue only occurs on the first helical scan after system startup that applies SAS function | Software design | On April 20, 2016 the firm sent Field Safety Notices to their customers. Philips will release a Field Change Order to upgrade the software free of charge. A Philips Field Engineer will contact their customers to schedule this correction. |  | 899 | 899 units | Products were sold in the following states: AL, AR, CA, FL, GA, IL, IN, MS, NC, ND, NJ, NV, NY, OH, PA, SC, TN &amp; VA.Products were also shipped to the following countries: Albania, ALGERIA, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Burundi, Cape Verde, Chile, China, Colombia, Congo, Democratic, Costa Rica, Cyprus, Dominican Republic, Ecuador, Egypt, France, Germany, Greece, Guatemala, Hungary, India, INDONESIA, Ireland, Italy, Japan, Jordan, Latvia, Lebanon, Libyan Arab Jamahiri, Lithuania, Malawi, Mali, Mexico, Morocco, Myanmar, Namibia, Nepal, Netherlands, Nigeria, Oman, Palestine, Paraguay, Philippines, Poland, Romania, Russia, Senegal, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, United Kingdom, Venezuela, Viet Nam &amp; Zambia. | Open, Classified |  |  | 979
Z-2344-2016 | 74655 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/ zeego, Artis Q/ Q.zen, stand alone system, software controlled |  |  | July 07, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Due to an error in the 19 Live Display, image reproduction may fail in the examination room and the potential exists for the loss of images immediately after system startup. | Nonconforming Material/Component | Siemens sent an Important Customer Safety Notice dated July 7, 2016, by certified mail to customers to inform them of the corrective action they plan to take to prevent potential loss of the image display for live images in the examination room after system startup. Siemens service organization will contact each customer to arrange a date to resolve the issue with a software update of the monitor firmware via AX012/16/S. For questions regarding this recall call 610-448-6461. |  | 51 | 51 units | Nationwide Distribution to AL,AZ,CA,CO,IA,IL,MA,MD,MI,MN,MO,NE,NJ,NY,OH,OR,PA,RI,TX,VA,WA,WI,WV | Open, Classified |  |  | 51
Z-2352-2016 | 74764 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products Estradiol Reagent Pack, Unique Device Identifier No. 10758750005017, REF/Catalog Number/Product Code 855 2630; IVD -- Ortho-Clinical Diagnostics, Pencoed, Bridgend CF35 5PZ, UK --- The VITROS Estradiol test is performed using the VITROS Estradiol Reagent Pack and Calibrators on VITROS ECi/ECiQ Immunodiagnostic Systems, VITROS 3600 Immunodiagnostic Systems and VITROS 5600 Integrated Systems. |  |  | July 14, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | Ortho Clinical Diagnostics has issued a recall of their VITROS immunodiagnostics Estradiol Reagent assay. The firm was made aware of a potential cross-reactivity between the drug Fulvestrant and theVITROS Immunodiagnostic Estradiol Reagent assay. The firm conducted an internal investigation and was able to confirm biased positive (ranging from 1303-4046%) on Estradiol results, obtained from postmenopausal females taking Fulvestrant. | Device Design | On 7/14/2016, URGENT PRODUCT CORRECTION NOTIFICATION Customer Letter (Ref. CL2016-141, dated 7/14/2016) was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received product since 6/1/2015 to inform them that Fulvestrant is an interferent of the Estradiol Reagent Pack and causes positively biased results. Discuss any concerns you may have regarding previously reported estradiol results with your Laboratory Medical Director to determine the appropriate course of action. The results from this or any other diagnostic test should be used and interpreted only within the context of the overall clinical picture. For questions, contact Customer Technical Services at 1-800-421-3311. Foreign affiliates were informed of the issue via e-mail on 7/14/2016 and instructed to notify their consignees of the issue. |  | 7597 | Domestic: 1620; Foreign: 7597 | Nationwide, Puerto Rico, and foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Poland, Singapore, Spain, Venezuela. | Open, Classified |  |  | 7597
Z-2350-2016 | 74705 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | CURVED-TIP STAPLER 30,IS4000, model number 470530; |  |  | July 14, 2016 | 2016 | Intuitive Surgical, Inc. | An increase in the number of da Vinci EndoWrist Stapler complaints where the customer experienced aStapler Firing Failed message as well as review of da Vinci Xi Surgical System logs found an increased trend | Software design | Field Safety Notifications - Urgent Medical Device Corrections and Removal letters were sent on July 15, 2016 (dated July 14, 2016 advised users of the Stapler Firing Failed message and actions to take should the message be received. |  | 2065 | 2065 total | US, including Puerto Rico and Australia, Austria, Belgium, Chile, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Italy, Japan, Monaco, Netherlands, Norway, Portugal, , Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Venezuela. | Terminated | January 03, 2017 | 173 | 2065
Z-2345-2016 | 74609 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Medivance ArcticGel Neonatal Pad |  |  | June 30, 2016 | 2016 | C.R. Bard, Inc. | An internal review of an available product found that the Information For Use only contained English and Japanese languages when a CE-marked 23 language version was expected. | Other | A Customer Communication Correction Letter was sent to customers on June 30, 2016. The letter requested that they examine inventory and remove the original IFU and substitute a copy of the IFU that was included with the Correction letter. The letter also requested a sub-recall, and included a reply form which was to be completed and returned. If you have any further questions please call (401) 825-8690. |  | 64 | 64 | US Distribution to the states of : CA, FL, GA, IL, IN, MI, MN, MO, NY, OR, PA, TX, WA, and Internationally to Germany, Italy and Switzerland | Terminated | June 06, 2017 | 341 | 64
Z-2351-2016 | 74557 | formalin, neutral buffered | IFP | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Richard Allan Scientific 10% Neutral Buffered Formalin, |  |  | June 13, 2016 | 2016 | Richard-Allan Scientific Company | potential contamination with potassium hydroxide may have an impact when performing immunohistochemistry testing. It is unknownwhether the contamination will affect tissue antigenicity. If it were to affect tissue antigenicity, theoutcome would be risk of false positive or false negative test results which may lead to seriousinjury. There is the possibility the patient may need to be re-biopsied. | Nonconforming Material/Component | On 6/13/2016, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected customers via certified mail with instructions for identifying and disposing of the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. |  | 18 | 18 cases | Distributed in the states of FL, PA, NY, MD, CA, and AZ. | Open, Classified |  |  | 18
Z-2353-2016 | 73255 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | S3 MedSurg Bed, Model 3005.Stryker Medical |  |  | February 02, 2016 | 2016 | Stryker Medical Division of Stryker Corporation | CPU board failures cause fowler (backrest) electronic controls to stop working. | Component change control | Stryker Medical initiated a voluntary recall of the S3 MedSurg Bed, Model 3005, with notices sent on February 2, 2016 via certified mail. Stryker medical identified that Fowler does not raise up to desired angle; and the Fowler does not raise electronically, but manual function remains operational through the manual CPR release. Actions Needed for Customers:1. Locate the bed(s) listed by serial number on the attached business reply form.2. A Stryker ProCare Representative will contact your facility to schedule a time for repair.3. Fax (269)488-8691 or email productfieldaction@stryker.com the enclosed business reply form to confirm receipt of this notification.4. If you have further distributed the beds listed in this letter, please forward a copy of this notice to the new users and advise us of their new location. If you have any urgent questions or concerns, please, contact Stryker Customer Service at (1-800-327-0770). Normal business hours are Monday-Friday 8 a.m-5 p.m. (ET). |  | 8490 | 8490 | Distribution US nationwide, Canada, and Mexico. | Terminated | June 12, 2017 | 496 | 8490
Z-2365-2016 | 74581 | No | NFB | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Invacare Precise Rx Pediatric Flowmeter Model Number: IRCPF16 AW. Packed one per carton. |  |  | June 21, 2016 | 2016 | Invacare Corporation | The rear gradation scale was screen printed upside down on the flowmeter gauge. | Process control | lnvacare Corporation sent an Urgent Medical Device Recall letter dated June 27, 2016, to all affected consignees with return receipt to notify them of the recall. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees with questions were instructed to call 877-413-6008. |  | 395 | 395 units | Worldwide Distribution - US to CA, CO, CT, FL, GA, IA, ID, IL,IN,KS,KY,LA,MA,MD,MI,MS,MT,NC,NE,NJ,NM,NY,OH,PA,SC,TN,TX,UT,VA,WA and Internationally to Canada, Germany and New Zealand. | Open, Classified |  |  | 395
Z-2363-2016 | 74776 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Inc., Hartland, WI. |  |  | February 05, 2016 | 2016 | Merge Healthcare, Inc. | When measuring a lesion on an unmagnified mammography image then performing the same measurement on an image magnified by the Hologic imager (different image), the measurements are not the same. | Software design | The recalling firm's customer support notified customers via phone call beginning 2/5/2016 to inform them of the issue. Customers were informed that a software fix that corrects the issue was available. Customers were asked not to use the measurements on mammo images until upgraded. If customers want an upgrade, then someone would contact them to schedule the upgrade. |  | 49 | 49 sites potentially have the affected versions | Worldwide Distribution -- USA and Australia. | Open, Classified |  |  | 49
Z-2375-2016 | 74782 | No | DZI | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | Instructions for Use (IFU) and Manual NobPrl CC Guided Surgery Kit |  |  | July 20, 2016 | 2016 | Nobel Biocare Usa Llc | Incorrect drill protocol in the Instructions for Use (IFU) | Labeling design | The firm, Nobel Biocare, sent "IMPORTANT-QUALITY MESSAGE" Field Correction Letters and response forms dated July 22, 2016. The letter describes the product, problem and actions to be taken and informs the customers that the titles of the IFU and the Manual are "NobelGuide for NobelParallel Conical Connection" (ID number 2002) and "Guided Surgery with Nobel Parallel Conical Connection" (ID number 79423). The Customers are instructed to review if you have a printed version of the IFU or the Manual related to the article and discard them; and to complete the attached "Quality Message Acknowledgement" form and fax to 714-998-9348 or email to: joann.radovich@nobelbiocare.com. Customers are informed that an English version of the corrected IFU is attached to the letter and translated versions will be available on the Nobel Biocare website for electronic IFU's http://www.nobelbiocare.com/IFU within 30 days. If you have further questions, contact our customer service department at 714-322-5001. |  | 231 | 231 kits | Worldwide Distribution: US (nationwide) and Countries of: Austria, Belgium, Canada, Germany, Denmark, Spain, France, Great Britain, Italy, Netherlands, Norway, Poland, Portugal, Russia, and Turkey. | Open, Classified |  |  | 231
Z-2342-2016 | 74761 | [6 empty] | NucliSENS easyMAG Magnetic Silica |  |  | July 08, 2016 | 2016 | BioMerieux SA | Several customer complaints about amplification performance issues when using different batches of MagSil with the NucliSENS easyMAG extraction platform. | Under Investigation by firm | Consignees were notified of the recall by letter on July 18, 2016. The letter informed them of what actions need to be taken if they experience problems. The letter included a reply form to be returned. Subsidiaries and distributors were notified electronically on 7/8/2016. |  | 7845 | 7,845 units | Worldwide Distribution - US Nationwide in the states of: AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MO, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WA, WV, WI, and the countries of Angola, Argentina, Austria, Australia, Austria, Botswana, Brazil, Burkina Faso, Cambodia, China, Columbia, Djibouti Antilles, Estonia, Georgia, Germany, Greece, India, Israel, Italy, Kuwait, Latvia, Malaysia, Mexico, Netherlands, Poland, Qatar, Russia, Senegal, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, Zambia, and Zimbabwe | Open, Classified |  |  | 7845
Z-2377-2016 | 74634 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Modular Handle Triathlon Instruments, Non-Sterile, REF 6541-4-808, Lot 123, NDC M546654148080S |  |  | June 23, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedic is voluntarily recalling the Triathlon Modular Handle because analysis revealed that the press fit specifications between the dowel pin and the mating hole in the shaft were not being achieved potentially resulting in disassociation from the instrument. | Device Design | Stryker sent an Urgent Product Recall letter dated June 24, 2016 to Branches/Agencies by e-mail and Business Reply Forms was sent to Branches/Agents/Hospital Risk Management via UPS (with return receipt) on June 24, 2016. The letter identified the affected product, problem and actions to be taken. For questions call (201) 831-6693. |  | 12469 | 12,469 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 12469
Z-2373-2016 | 74622 | disk, abrasive | EHJ | Abrasive device and accessories. | Dental | Dental | 510(K) Exempt | OptraPol Polishing Discs, Rx ONLY: OptraPol Refill Cup /10: REF/Product ID 634459, OptraPol Refill Small Flame/10: REF/Product ID 634458, OptraPol Assortment: REF/Product ID 634464 --Device Listing # B106952. One step polishing system for all common dental composites and amalgam. |  |  | July 11, 2016 | 2016 | Ivoclar Vivadent, Inc. | After a few seconds of polishing with moderate force, parts of the OptraPol cup are breaking away. OptraPol Small Flame is also affected. Two complaints were filed referring to OptraPol, claiming that the abrasive polishers wear faster than expected. | Under Investigation by firm | URGENT MEDICAL DEVICE RECALL Letter (dated 7/11/2016) and Response Form were sent to customers on 7/11/2016 via Certified Mail, Return Receipt, Requested. The letter instructs customers to notify their customers of the recall if the product had been distributed and to return the product to Ivoclar Vivadent. For questions, please contact Ivoclar Vivadent Customer Service at 1-800-533-6825. |  | 86 | US: 86 units total | Distributed in the states of CA, FL, GA, IA, IL, IN, KS, KY, MA, MI, NV, NY, OH, PA, TX, and WA. | Open, Classified |  |  | 86
Z-2378-2016 | 74788 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Tritest CD3 FITC/CD4 PE/CD45 PerCP (IVD); |  |  | July 18, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | Internal review found that some vials of CD3/CD4/CD45 state that the product contains gelatin. The product does not contain gelatin and should state in buffer. | Labeling Change Control | BD Life Sciences sent an Urgent Product notification letter dated July 18, 2016 to all affected customers via UPS and/or e-mail. The letter informed the user of the issue and states that no action is required since product performance is not impacted. The response form is to be returned to BD as instructed. Customers requiring further assistance were instructed to contact BD Customer Support at 855-236-2772.For questions regarding this recall call 408-954-6080. |  | 14089 | 14,089 | Nationwide distribution to AR, CA, CO, IL, IN, LA, MA, MD, Ml, MN, NC, NJ, NM, NY, OR, PA, TN, TX, WA | Open, Classified |  |  | 14089
Z-2374-2016 | 74681 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | Servo Humidifier 163, Model 01-06-8125-8, REF 6419365 |  |  | December 03, 2015 | 2015 | Teleflex Medical | Cracks may occur at the 15mm and 22mm ISO female connector. | Release of Material/Component prior to receiving test results | Inmed, a division of Teleflex sent an Urgent Field Safety Notice letter dated December 3, 2015 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check inventory, cease use and distribution of stock of the affected product batch and quarantine immediately.. The letter included an acknowledgment form which was to be completed and returned your MAQUET representative.For further information or support concerning this issue, contact your local MAQUET representative at fieldactions@maquet.com. For questions contact your local sales representative or Customer Service. |  | 2700 | 2,700 ea. | US Nationwide Distribution in the states of FL, KS and WV. | Terminated | April 14, 2017 | 498 | 2700
Z-2376-2016 | 74615 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Volker 2080, 2082, S 280 and S 282 series beds with lateral release button. |  |  | June 14, 2016 | 2016 | Hill-Rom, Inc. | The current construction of the side rail latch may not be able to support the full patient weight and may give way allowing the side rail to abruptly lower and the patient to fall, potentially resulting in harm. To date there has been one report of a related injury. | Device Design | Hill-Rom sent an URGENT: MEDICAL DEVICE CORRECTION letter dated June 14, 2016 to the affected US and Canadian customers via email and courier, international customers. The letter included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The letter requests the number of beds in the consignees possession so that repair kits can be provided. If you have any questions regarding this safety notice, contact Hill-Rom Technical Support at 1-800-445-3720, your distributor, your Hill-Rom representative or contact us on www.hill-rom.com. |  | 24262 | 24262 | Worldwide Distribution - US Nationwide and the countries of: Canada, AUSTRALIA, CHINA, HONG KONG, THAILAND, POLAND, RUSSIAN FEDERATION, SLOVENIA, UNITED ARAB EMIRATES, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, ICELAND, IRELAND, ISRAEL, ITALY, LUXEMBOURG, NETHERLANDS, NORWAY,SPAIN, SWITZERLAND, and UNITED KINGDOM. | Terminated | May 15, 2017 | 335 | 24262
Z-2371-2016 | 73750 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. |  |  | March 24, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host. | Software change control | On 3/24/2016 the firm mailed Customer Information letters to their customers. |  | 29 | 29 units | Product was sent to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV &amp; WY.Product was also sent to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Cuba, Denmark, Djibouti, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, RÂżunion, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 635
Z-2380-2016 | 74131 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 CT Model 728231; To produce cross-sectional images of the body. |  |  | April 29, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Software issues found in software versions v4.1.3/4.1.4/4.1.5 in the Philips Brilliance 64 products that could affect the performance of the equipment. | Software design | The customers will be sent a letter with information about, what the problems are and under what circumstances they may occur, the actions customers can take to minimize the effect of the problem and the actions planned by Philips to correct the problem. To correct these issues, a Philips field service representative will install a software update on the affected systems. Reference field change order (FCO) 72800652. |  | 150 | 150 | Product was shipped to the following states: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV &amp; WY.Product was also shipped to the following countries: Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, China, Colombia, Cuba, Estonia, Finland, France, Germany, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Myanmar, Netherlands, New Zealand, Norway, Palestine, Panama, Philippines, Poland, Portugal, RÂżunion, Russian Federation, Senegal, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Vietnam. | Open, Classified |  |  | 1403
Z-2379-2016 | 74773 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. Product was distributed under the AMICAS label, Brighton, MA, but labeling was later changed to Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | The patient name in the Halo title bar and the thumbnails do not match the name on displayed images. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/30/2016 explaining the reason for recall and instructing the customer to provide a copy of the notification to all users. A second letter was issued to nonresponders dated 5/25/2016 |  | 93 | 93 sites have the affected version | The software was distributed to medical facilities nationwide. Government distribution was made to MN. Foreign distribution was made to Australia and to the United Kingdom. There was no military distribution. | Open, Classified |  |  | 93
Z-2391-2016 | 74553 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | Coolife Sinergy Radiofrequency Kit, SIK-17-75-4 (109971504); Kit used in conjunction with a Radiofrequency Generator to create radiofrequency lesions in nervous tissue. |  |  | June 16, 2016 | 2016 | Halyard Health, Inc | Mismatch between the length of the RF electrode (probe) and the cannula (introducer), | Process control | All distributors and healthcare facilities that were directly shipped recalled devices were notified of the recall beginning 6/16/2016 by email and by FedEx. They were instructed to examine their stock and to quarantine and discontinue use of the product. The letter contained a response form which was to faxed back to Halyard. |  | 51 | 51 kits | Distributed in the states of CT, IL, MA, MI, NV, NH, NC, and TX, and the country of Canada. | Open, Classified |  |  | 51
Z-2389-2016 | 74719 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4mm Titanium Locking Screws Self-Tapping with Stardrive Recess 12mm, Synthes 2.4mm Titanium Locking Screws Self-Tapping with Stardrive Recess 12mm, Plate, Fixation, Bone, Orthopedic Devices |  |  | June 28, 2016 | 2016 | Synthes (USA) Products LLC | It was reported that the specified lots of the 2.4mm Titanium Locking Screw Self-Tapping with Stardrive Recess 12mm can break between the head and the shaft. This resulted from the neck feature being too close to the StarDrive feature leaving a thin wall between neck and drive. | Under Investigation by firm | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated July 14, 2016 to affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory and immediately remove and return all unpackaged Locking Screws Self-Tapping with Stardrive Recess 12mm in any affected kits along with any screws still in the original packaging with the part number and lot number identified in the recall notification. Customers were also instructed to call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, complete the Verification Section and indicate the number of devices found and return with product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132 by fax to 877-523-9109 or scan/email: Synthes5246@stericycle.com, even if you do not have the product. Note: if the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 41 | 41 units | US Distribution to states of: AR, CA, CO, DE, FL, GA, IA, IL, KY, LA, MN, NJ, OK, PA, SD, TN, and WA. | Open, Classified |  |  | 41
Z-2386-2016 | 74757 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Integra External Fixation System Universal Wire Fixation Bolt - 17.5 mm Catalogue #12224218 ; Integra External Fixation System Universal Wire Fixation Bolt - 23 mm Catalogue #12224219 |  |  | July 20, 2016 | 2016 | Integra LifeSciences Corp. | Through the investigation of reported complaints Integra verified that there have been instances where the Universal Wire Fixation Bolts the Slotted Post broke during use at the surgical site. | Device Design | Integra sent an Urgent - Voluntary Medical Device Recall letter dated July 20, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Integra is asking you to take the following actions:1. Please check your IntegraÂż External Fixation System inventory to determine if you have the 17.5mm or 23mm UniversalWire Fixation Bolts and/or the large or small Slotted Posts.2. If you do have 17.5mm or 23mm Universal Wire Fixation Bolts and/or large or small Slotted Posts, stop using them immediately.3. Complete the attached "Acknowledgement and Return Form" and check the box: I do have affected products on the list and record the lot number.4. Or, complete the attached "Acknowledgement and Return Form" and check the box: I do not have affected products.5. Complete the other information as indicated on this form. Keep a copy of the form for your records.Return the completed "Acknowledgement and Return Form" by email or fax indicating your receipt and review of this notification.When your form is received, if you have noted you have affected products, an Integra Representative will contact you and provide you with directions to return the product, as well as input an order to replace the quantity you indicated on the form.Receipt of this form ensures that Integra has achieved a level of effectiveness in communicating this information. We recommend you also maintain a copy of this notification and signed copy of the acknowledgement form for your records. Regulatory agencies such as the FDA perform audits of field actions of this nature to verify that our customers have been notified and understand the nature of the field action.Should you have any questions regarding these instructions, please contact Sales Operations at 888-601-0203, option 2. We apologize for any inconvenience this may cause and thank you for your cooperation in this effort.Fo |  | 495 | 495 units | US Distribution to the states of : FL, VA, OK, CO, DC and IN. | Open, Classified |  |  | 606
Z-2400-2016 | 73577 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | InvacareÂż Bariatric Bed with mattress and rails , Junction Boxes, |  |  | March 15, 2016 | 2016 | Invacare Corporation | The firm recently identified a quality issue with the Junction box used in the Invacare Bariatric Bed. They discovered that the Junction Box may omit sparks which could result in a fire in the presence of flammable materials. | Nonconforming Material/Component | Invacare sent an Urgent Medical Device Recall letter dated May 12, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions call Invacares recall support line at (877) 413-6008, Monday  Friday, 8 a.m.  5 p.m. EST, for assistance. |  | 2702 | 2,702 units | US Nationwide Distribution including DC and PR. Internationally Canada. | Open, Classified |  |  | 8947
Z-2404-2016 | 74664 | container, liquid medication, graduated | KYW | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | Medicine Cup, 1 OZ. Liquid medication dispenser to issue a measured amount of liquid medication. |  |  | May 11, 2016 | 2016 | Medline Industries Inc | There is 5 mL printed next to the increment on the medicine cups instead of 10 mL. The cup has the correct mL markings for the other gradations, but sequentially, the markings are listed as 2.5 mL, 5 mL, 7.5 mL, and 5 mL. | Process control | Medline Industries sent a an URGENT RECALL letter dated 5/11/2016 to all affected consignees via certified mail. The URGENT RECALL letter described the reason for the recall and the actions to take. |  | 903 | 903 Cases and 4,090 Sleeves (8,605,000 Medicine Cups) | Continental U.S. | Open, Classified |  |  | 903
Z-2390-2016 | 74619 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Specialty Triathlon Tibial Alignment Handle with Secondary Lock, Non-Sterile, |  |  | June 23, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | It was reported that the secondary locking mechanism, and its corresponding components, disassociated from the Specialty Triathlon Tibial Alignment Handle with Secondary Lock Assembly during surgery. Upon further investigation, it was discovered that the weld, which was intended to hold the secondary locking mechanism together, did not meet the weld size specified on the engineering drawing. | Process control | Stryker sent an Urgent Product Recall letter dated June 24, 2016, with a Business Reply Form attached to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to inform users of the Urgent Product Recall letter and forward the notice to all those individuals who need to be aware within their organization. Customers were instructed to return all affected product at their location to:Stryker C/O Stericycle2670 Executive Dr., Suite AIndianapolis, IN 46241Customers were also instructed to complete and sign the enclosed Business Reply Form and fax it to 888-912-7352 or email to Stericycle at strykerortho4582@stericycle.com. Customers with questions were advised to call (201) 831-6693.For questions regarding this recall call 201-831-5000. |  | 47 | 47 units | Nationwide Distribution to FL, IN, LA, NC, PA, TN and UT | Open, Classified |  |  | 47
Z-2525-2016 | 74534 | No | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | StratusÂż CS STAT Fluorometric Analyzer-microprocessor-controlled instrument |  |  | June 22, 2016 | 2016 | Siemens Healthcare Diagnostics Inc | Software defect, where either an Above Assay Range or an inaccurate value could potentially be reported, in the Stratus CS Acute Care Diagnostics System regarding TestPak calibration and quality control (QC) status when a TestPak is due to expire within 48 hours. | Software design | Siemens Healthcare issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on June 22, 2016 both in the United States and Outside the US for communication with affected customers. This notice informs customers of the issue and provides mitigation instructions. Customers are to contact Siemens Customer Care Center or their local Siemens technical support representative at 800-405-6473. Siemens Healthcare is currently developing a software update to address this issue and will be providing new information as it becomes available |  | 977 | 977 units | Worldwide Distribution -- USA, U.A.E., Algeria, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia Herzeg., Botswana, Canada, Colombia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Namibia, P.R. China, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., and United Kingdom. | Open, Classified |  |  | 977
Z-2396-2016 | 74743 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Orthopedic manual surgical instrument |  |  | July 01, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Malfunctioning of the Exeter Rasp/trial introducer/extractor Handle. | Device Design | Stryker notified Branches/Agencies of the Urgent Medical Device Recall by e-mail on July 1, 2016 and an "Urgent Medical Device Recall Letter and Business Reply Form" was sent to branches, agents and hospital risk management via UPS (with return receipt) on July 5, 2016. The letter identified then affected product, problem and actions to be taken. For questions call (201) 831-6693. |  | 241 | 241 units | Worldwide Distribution - US Nationwide | Terminated | June 08, 2017 | 342 | 241
Z-2405-2016 | 74724 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) |  |  |  | July 15, 2016 | 2016 | Topcon Medical Laser Systems, Inc | Topcon discovered an anomaly with titration, only when EpM is turned on during the titration process.This event may occur if the user turns on EpM, turns off the Landmark (LM) function (with any EpM supported pattern) and then enters titration mode. | Device Design | The firm, PASCALS Synthesis, sent an "Urgent Field Safety notification" letter dated July 15, 2016 to all affected sites by traceable method. The letter described the product, problem and actions to be taken. The customers were instructed to Acknowledge the receipt of the field notice via an official letter; locate all systems with Endpoint Management software feature enabled; for all affected units- inform the physician, hospital or institution customers of the addendum and additional warnings issued as a result of this notice; obtain signed confirmation from the end-user regarding receipt of the addendum and additional warnings; send the signed confirmation from ALL end-users to TMLS, no later than October 15, 2016; and complete to the following record "Products lines affected" for all affected units and either fax to (408) 235-8259 or scan and email to tmlsinfo@topcon.com.Contact Customer Care if you have questions at: 1.408.235.8200 or email: tmlsinfo@topcon.com. |  | 205 | 205 units (21 in US, 184 OUS) | Worldwide Distribution: US (nationwide) and countries of: Canada, Mexico, Japan, Netherlands, Algeria, Australia, Brazil, Chile, China, France, Guatemala, Hong Kong, Italy, New Zealand, Pakistan, Philippines, Singapore, Taiwan, and United Kingdom. | Open, Classified |  |  | 205
Z-2395-2016 | 74716 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Q AND Q.ZEN- MODULAR ANGIOGRAPHIC SYSTEM, interventional fluoroscopic x-ray system22 |  |  | July 13, 2016 | 2016 | Siemens Medical Solutions USA, Inc | The possibility exists for contaminants, in the form of biomass, to develop in the cooling system of Artis systems resulting in damage to the pump system. | Under Investigation by firm | Siemens has implemented a Customer Safety advisory Notice to all affected customer via Update Instruction AX007/15/S. Additional, a corrective action will be implemented via Update Instructions AX006/15/S and AX003/16/S. |  | 14298 | 11,111, worldwide and 3,187 in the USA | Nationwide Distribution to: AK,AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NC,ND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,PR,RI,SC,SD,TN,TX,UT,VA,WA,WI,WV,WY, Puerto Rico and Guam | Open, Classified |  |  | 14298
Z-2398-2016 | 74687 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | ZystonÂż Transform Implant Kit. Model Number PCR8200-3101. Device labeled sterile. |  |  | June 30, 2016 | 2016 | Zimmer Biomet Spine, Inc | Zimmer Bioment Spine, Inc. announces a voluntary field action for the ZystonÂż Transform Instrument Case and ZystonÂż Transform Implant Kit because the locking mechanism on the implant is not functioning as intended. | Under Investigation by firm | Verbal notification of the decision to recall the product and related instructions was provided on June 30, 2016. Additionally, written notification of the decision to recall the product and related instructions will be provided to each of the individuals responsible for locations that have affected inventory under their control. Clear instructions and materials necessary to support the accountability and retrieval of product from the field will be provided with the recall notification. In order to ensure that the locations were alerted of the recall in a timely fashion, supplemental email notifications will also be sent to each recall notification addressee before distributing the recall documentation via Federal Express. |  | 5 | 5 | Nationwide to PA, TX, TN, NV, VA, WV, and KY. No Canadian, other foreign, VA/gov/military consignees. | Open, Classified |  |  | 5
Z-2392-2016 | 74714 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi Scenaria Computed Tomography System Hitachi Medical Systems America, Twinsburg, OH 44087 |  |  | June 30, 2015 | 2015 | Hitachi Medical Systems America Inc | The firm discovered that the centrifugal force applied to the power supply may cause it to fail over time. The exam performed at the time of the failure will likely be lost, resulting in the need to rescan the patient. | Component design/selection | Hitachi sent an Device Correction Notice dated April 1, 2015, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.Customers were informed that there is a problem with the Hitachi Medical Systems America, Inc.CT system; the issue is with a power supply in the CT gantry sub-system. During fast rotation (i.e., 35 sec), the centrifugal force applied to the power supply may cause it to fail over time.For further questions, please call (330) 425-1313 ext. 3720. |  | 33 | 33 units | US Distribution to the states of : CA, FL, GA, IA, IL, IN, MD, NC, NY, OH, PA, SD, TN, TX &amp; WY., and Internationally to Brazil and Mexico. | Open, Classified |  |  | 33
Z-2535-2016 | 74799 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Straumann Bone Level Implant, Âż 4.1mm RC, SLAÂż 12mm, TiZr, NTP |  |  | July 14, 2016 | 2016 | Straumann Manufacturing, Inc. | Misalignment with the Loxim transfer piece markings, the positioning of the implant will be incorrect for optimal restoration | Under Investigation by firm | Straumann notified customers via telephone on Thursday, July 14, 2016. Customers with an e-mail address on filewere also e-mailed on Friday, July 15, 2016.Straumann notified customers were sent a hard copy letter via UPS on Saturday, July 16, 2016 for delivery on Monday, July 18, 2016.The customers have been notified not to use these implants and to return unused implants to Straumann. The customers that have used the implants have been asked to provide the restorative situation for the implants. Depending on the restorative situation, additional instructions may be necessary to restore the implants. |  | 33 | 33 pieces | Nationwide Distribution. | Open, Classified |  |  | 103
Z-2470-2016 | 73184 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Implant EP Pick-Up Coping |  |  | November 25, 2015 | 2015 | Biomet 3i, LLC | Pouches may not have been sealed during packing. | Packaging | Biomet sent an Urgent Device Recall Notice to all affected consignees the week of November 30, 2015, communicating the issue and customer responsibilities. International consignees were notified via traceable courier or email week of November 30, 2015. Customers were instructed to return affected product by calling BIOMET at 1-800-342-5454 to obtain a RMA number to return products. Customers were also asked to complete the Business Reply Form and fax to 561-514-6316.For questions regarding this recall call 561-776-6700. |  | 1 | 1, 648,273 devices (all products subject to this recall) | Worldwide Distribution Nationwide andHawaii, Guam, Puerto Rico and Virgin Islands (U.S.) .FOREIGN CONSIGNEES: Albania, Andorra, Argentina, Australia, Austria, Barbados, Belgium, Bolivia, Brazil, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Denmark, Dominican Republic, France, Germany, Great Britain, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Jamaica, Japan, Korea (South), Kuwait, Lebanon, Liechtenstein, Lithuania, Martinique, Mexico, Mongolia, Netherlands, Northern Mariana Islands, Paraguay, Peru, Poland, Portugal, Qatar, Reunion, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tabago, Turkey, United Arab Emirates, Uruguay and Viet Nam. | Open, Classified |  |  | 1648274
Z-2526-2016 | 74520 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AB Sciex API 3200MD" LC/MS/MS System. In-Vitro Diagnostic |  |  | June 16, 2016 | 2016 | Ab Sciex | MultiQuant MD software where under certain conditions a usercan be presented with incorrect quantitative results when using the Sum Multiple Ions feature. | Software design | AB SCIEX issued recall notification by letter dated June 16, 2016, to all affected customers. The letter identified the problem, provided temporary measures until a permanent software is available. A response form was to be completed and returned. questions regarding this noticecontact SCI EX at +1 289 982 2712. |  | 124 | 124 units | Worldwide Distribution - US (Nationwide) Internationally to Canada, Austria, China, Netherlands, Norway, India, Korea, Germany, France, Switzerland, Ireland, Italy, Singapore, Russia, UK, Malaysia, | Open, Classified |  |  | 275
Z-2532-2016 | 74734 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iConnect Access used with Ortho PACS software provides medical specialists with access to diagnostic quality images, reports, and various types of patient data over conventional TCP/IP (e.g., internet) networks. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Studies that are viewed in iConnect Access that originate on Merge PACS / Ortho PACS may not display the correct current patient demographics if there has been a PDE (Patient / Study Demographic Edit) performed after the images were ingested in PACS. | Software design | The recalling firm issued a letter dated 1/28/2016 via e-mail on 1/31/2016 or via certified mail if they did not have an e-mail address for the customer. |  | 146 | 146 sites may have the affected version for both products | Distribution was made to medical facilities nationwide and to one foreign medical facility in New Zealand. There was no military or government distribution. | Open, Classified |  |  | 146
Z-2530-2016 | 74533 | No | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | Perioscopy Sheath: Part number: 1008; |  |  | June 24, 2016 | 2016 | Danville Materials, Inc. | The Perioscopy Sheaths are currently provided with a label that indicates the Sheath has a 3 year shelf-life/expiry date. Danville Materials has become aware that validation data are not available to support the 3-year shelf-life/expiration dating claim. | Under Investigation by firm | Recall letters were sent on June 24 to Distributor OEM and customers by traceable delivery method. Letters informed customers of the issue and the affect part and lot numbers. Pro duct is to be quarantined and returned. Distributors are instructed to contact all customers and provide a copy of the recall letter. The response form should be completed and returned by each consignee. |  | 2123 | ~ 2123 units | US and New Zealand, Australia, Canada, Japan, Israel, United Kingdom, Italy, Mexico. | Open, Classified |  |  | 2123
Z-2533-2016 | 74625 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Synovis VASCU-GUARD Peripheral Vascular Patch in the following sizes and product codes: |  |  | June 24, 2016 | 2016 | Baxter Healthcare Corp. | Baxter is issuing a safety alert in response to postmarketing reports received for the VASC U-GUARD Peripheral Vascular Patch. Baxter has received reports for intraoperative or postoperative bleeding episodes, which required additional clinical intervention. One report involved a case with a fatal outcome. UPDATE 9/1/2016 Baxter issued a notice retrieving 6 lots of VASCU-GUARD due to the increased postmarketing reports related to intraoperative and postoperative bleeding. | Under Investigation by firm | Baxter sent a Safety Alert to customers via USPS overnight delivery beginning 6/24/2016. The letter told of the reported adverse events, asked for product that product immediately be discontinued from use, to segregate/quarantine all affected lots, and to return the enclosed customer reply form. General questions regarding the Safety Alert should be directed to Baxter Product Surveillance at 1-800-437-5176, between 8:00 am - 5:00 pm Central Time Monday thru Friday. Clinical inquiries can be directed to Baxter's Medical information at 1-800-933-0303, Monday thru Friday 8:00 am - 5:00 pm central time.UPDATE 9/1/2016 Baxter issued an Urgent Product Recall letter beginning 9/1/2016. This letter explained the issue and identified affected product that was to be removed from the consignee and returned to Baxter. Customer is to contact Baxter Healthcare Center for Service to arrange for return and credit. This can be done at 888-229-0001 Monday thru Friday 7:00 am - 6:00 pm Central Time. General questions regarding the recall can be addressed to Baxter Product Surveillance at 800-437-5179 Monday thru Friday, 8:00 am - 5:00 pm Central Time. |  | 294867 | 32,763 eaches UPDATE 9/1/2016 2,042 eaches | Nationwide including Puerto Rico, and outside the US to include: Australia, Austria, Finland, Germany, Greece, Ireland, Italy, the Netherlands, Norway, Poland, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, United Kingdom. | Terminated | June 23, 2017 | 364 | 294867
Z-2536-2016 | 74582 | No | NDN | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | Stryker AutoPlex System. Model Numbers 0607-687-000, 0605-887-000, 0605-687-000, and 0605-683-000. For bone cement. |  |  | June 14, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | The Piston Head can become separated from the Delivery Piston, blocking the injection assembly Valve resulting in a cement backflow towards the injector handle. Potential for delay in surgery if additional cement needs to be prepared for the injection procedure. | Process control | On 6/13/2016, URGENT MEDICAL DEVICE RECALL NOTIFICATION notifications were sent to the affected customers via courier and sales representatives via email. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments:Jennifer Olson269-389-2644jennifer.olson@stryker.com |  | 4841 | 4,841 boxes (9,682 units) | Worldwide distribution. US nationwide, France, Netherlands, Spain, Italy, Canada, Israel, Romania, Switzerland, and United Kingdom. | Terminated | March 27, 2017 | 286 | 4841
Z-2539-2016 | 74672 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TRYX Antibacterial Envelope Rx Only |  |  | June 29, 2016 | 2016 | TYRX Inc. | TRYX products are being recalled since the processes of spaying, welding, drying oven and polymer were not adequately validated. | Process control | Medtronic sent an Urgent Medical Device Recall letter dated June 2016 to affected customers. The letter identified the affected product problem and actions to be taken. Customers are asked to immediately remove and quarantine all unused product that remains in inventory, return unused product to Medtronic and contacted the local Medtronic representative or Customer Service at 800-848-9300 to assist with the return and credit of unused product. Medtronic will provide credit for all non-expired, unused product. The customer should complete the attached Customer Confirmation Certificate and email to RS.CFQFCA@Medtronic.com or fax to Medtronic at 651-367-0612 Attention Customer Focused Quality. |  | 3460 | 3460 units | US Nationwide Distribution | Open, Classified |  |  | 6647
Z-2538-2016 | 74683 | No | HCF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Weck EFx Classic Fascial Closure System, Rx Only, Sterile, |  |  | March 24, 2016 | 2016 | Teleflex Medical | Incorrect expiration date was printed on the product label. | Unknown/Undetermined by firm | Teleflex sent an Urgent Medical Device Recall Notification letter dated March 24, 2016, to all affected consignees. The letter requested that consignees check their stock, immediately discontinue use and quarantine any products. To return affected product consignees were instructed to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, ATTN: Customer Service or email to recalls@teleflex.com. Customers with questions were instructed to contact their local sales representative or Customer Service at 1-866-246-6990.For questions regarding this recall call 610-378-0131. |  | 45 | 45 | Nationwide Distribution to CA, CO, IA, KS, MD, MI, NY, NC, PA | Open, Classified |  |  | 45
Z-2547-2016 | 74663 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard TruGuide Disposable Coaxial Biopsy Needle  Mixed Product |  |  | July 06, 2016 | 2016 | Bard Peripheral Vascular Inc | May contain two trocar tip stylets with coaxials or two blunt tip stylets instead of one of each. | Packaging | Bard Peripheral Vascular (BVP) sent an Urgent Medical Device Recall Notification letter dated July 6, 2016 to all affected customers via FedEx with proof of delivery notification.. The letter identified the product, the problem, and the action to be taken by the customer. The notification provided customers the following instructions:1. Do not use or further distribute any affected product.2. Check all inventory locations within your institution for affected product code I lot number combination listed in the recall notice. If you have further distributed any of the product code I lot numbers, please immediately contact that location, advise them of the recall, forward these instructions and have them return the affected product to BPV.3. Please remove any identified product from shelves.4. If used the affected product, complete and return the attached Recall and Effectiveness Check Form indicating no product will be returned.Once the product affected by recall has been removed from inventory;5. Fill out the Recall and Effectiveness Check Form. Be sure to state the quantities and lot numbers of each recalled product in stock even if no longer have possession of the recalled product.6. Call the firm's Recall Coordinator Raye Seisinger at 1-800-321-4254 Option #2 Ext 2501 (M-F 6am to 3pm MST) or email at raye.seisinger@crbard.com. Once all information has been verified, the Recall Coordinator will issue a Return Authorization (XC) Number or Consignment Recall Number (XH) to facilitate the expedient return of the product. BPV will provide a credit for your returned product.7. Fax the Recall and Effectiveness Check Form to BPV at 1-800-994-6772. If cannot FAX the form, please call the BPV recall coordinator at 1-800-321-4254 Option #2 Ext 2501) and report the required information verbally.8. A mailing label is enclosed to return the affected product. Please mark the outside package as "RECALLED PRODUCT" and i |  | 1250 | 1250 units | Nationwide Distribution to the following; CA, LA, VA, IA, AR, SC, PA, MA, IL, MO.Foreign distribution to the following; Canada, Belgium, Australia, Taiwan, China, Argentina. | Terminated | November 17, 2016 | 134 | 1250
Z-2548-2016 | 74803 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Xpert MRSA |  |  | July 18, 2016 | 2016 | Cepheid | Internal investigation found the certain lots of MRSA kits have the potential for performance issues at lower positive sample concentrations. | Nonconforming Material/Component | Recall letters dated July 18, 2016 were sent to affected customers on July 19, 2016 by Federal Express. Letters advised users of the affected product and asked that any remaining inventory be destroyed. |  | 2000067041 | 68 kit for lot 1000037539 and 127 kit s for lot 1000029307. | US only: PA, DE, FL, SC, ME, CA, IL, NC, TX, WA, NY, TN, FL, OR, OK, MO, MS, GA, NM, IN, OH, MD, VA, CO, AZ, and NC. | Open, Classified |  |  | 2000067041
Z-2542-2016 | 74876 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Xario 100 Diagnostic Ultrasound System, TUS-X100; Xario 200 Diagnostic Ultrasound System, TUS-X200. |  |  | August 04, 2016 | 2016 | Toshiba American Medical Systems Inc | Toshiba American Medical Systems (TAMS) is recalling the Xario Diagnostics Ultrasound System because it may become hot because of a software error. | Software design | An Urgent Medical Device Correction letter was sent to all customers to bring to attention a potential problem of transducer heating with your Xario Ultrasound System. The letter informs the customers that the surface of the acoustic lens of a transducer may become hot because of a software error. The error occurs when a transducer is disconnected, reconnected, or another transducer is selected. The letter informs the customers of error message and the corrective action. Customers are instructed to complete the attached form and fax it to (877) 349-3054 or TAMS-RegulatoryAffairs@toshiba.com. Customer with any questions are instructed to call (714) 730-5000 or their Toshiba service representative at (800) 521-1968. |  | 93 | 93 units | US Distribution | Open, Classified |  |  | 93
Z-2546-2016 | 74796 | Yes | NPT |  |  | Cardiovascular | PMA | Medtronic EnVeo R Delivery Catheter System (DCS), Models: ENVEOR-L, ENVEOR-L-C, and ENVEOR-US. |  |  | August 03, 2016 | 2016 | Medtronic Cardiovascular Surgery-the Heart Valve Division | Potential for vascular trauma with the use of the Medtronic EnVeo R Delivery Catheter System (DCS). | Device Design | Firm initiated a recall on 8/3/16 by sending an Urgent Medical Device Correction letter to inform them of the potential for vascular trauma with the use of the Medtronic EnVeo R Delivery Catheter System (DCS). The letter informs the customers of the recommendations to reduce the rate of vascular trauma events and the procedural uses. The letter informs the customers that the Evolut R System Instructions for Use (IFU) will also be updated appropriately consistent with the Dear Doctor communication. Patients with questions can contact Medtronic Cardiovascular Patient Services at 877-526-7890 (Monday-Friday, 7:30am-5pm Central Time). Customers with any questions are instructed to contact their Medtronic Field Representative or Lifeline Technical Services at 1-877-526-7890.On 8/19/16, firm sent a copy of a customer notification letter that will be sent to their customers to inform them that of the communication letter that was released in August 2016 by Medtronic regarding potential for aortic vascular trauma with the use of the Medtronic EnVeo R Delivery Catheter System (DCS). The letter informs the customers that FDA classify the recall as a Class II recall and the recall does not require customers to return the product. The recall letter informs the customers that Medtronic has taken further action by updating all customer training materials and the Instructions for Use (IFU). |  | 7319 | 7,319 units | Worldwide Distribution -- US, Australia, Austria, Belgium, Brazil, Colombia, Costa Rica, Czech Republic, Denmark, Dominican Republic, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Kazakhstan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macedonia, Malaysia, Morocco, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | March 31, 2017 | 240 | 7319
Z-2551-2016 | 74846 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Unity Z3D software. |  |  | July 23, 2015 | 2015 | Merge Healthcare, Inc. | The software is unable to accurately determine the calcium score of scans with a slice thickness not equal to 3 mm. | Software design | The firm, Merge, sent a "canned" email on 7/23/2015 to its customers which also contained a personal message. The recalling firm issued a second letter entitled "URGENT: MEDICAL DEVICE RECALL" dated 1/27/2016 via e-mail on 1/30/2016. The letters described the product, problem and actions to be taken. The letter informed customers that an upgrade is available. The customers were instructed to call Merge Customer Service at (877) 741-5369 or send an email to support@merge.com to schedule the upgrade-Reference this recall #2016-14. If you decide to decline this maintenance release, complete the attached form and return to Merge Healthcare via recall@merge.com. Alternatively, complete the online form: http://www.merge.com/Landing-Pages/Customer-Upgrade-Decline-Form.aspx no later than February 22, 2016. Ensure that all users of the product are provided with this notification. If you have further distributed this, identify your customers and notify them at once of this product recall and provide a copy of the recall notification letter.If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 9 | 9 sites potentially have the affected version | US Distribution to states of: CA, PA, TX, and MT. | Open, Classified |  |  | 9
Z-2550-2016 | 74736 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Quattro Air FFM MED  AMER, Product Code: 62702, full face mask for use with active-exhaust-valve ventilators |  |  | July 12, 2016 | 2016 | Resmed Corporation | Masks in this lot, labelled and sold as 62702 vented Quattro Air FFM size Medium actually contain a non-vented (NV) Quattro Air FFM Large mask. | Packaging process control | An urgent product recall notification letter was sent to all customers on 7/12/16 to inform them that ResMed is recalling NV Quattro Air &amp; Quattro Air Mask due to a mislabeling Issue. The letter informs the customers that during the investigation of a product complaint, ResMed identified that a small number of masks labelled and sold as 62702 Quattro Air FFM MED  AMER, Lot 1142623 actually contain an NV Quattro Air FFM Large mask. The letter informs the customers to notify anyone they may have sold the products to. Customers with any questions are instructed to call the ResMed Customer Service team: 858-836-5968. |  | 42 | 42 units | US nationwide distribution | Open, Classified |  |  | 42
Z-2549-2016 | 74704 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DS USA Inc., Passport Series Patient Monitor intended for monitoring, displaying, reviewing, alarming, and transferring of multiple physiological parameters including ECG (3-lead, 5-lead or 12-lead selectable, arrythmia detection, ST segment analysis, and heart rate (HR), Respiration rate. |  |  | June 23, 2016 | 2016 | Mindray DS USA, Inc. dba Mindray North America | Mindray has identified a potential leakage issue with the CO2 Module's Dryline II Water Trap used with the Passport Series Patient Monitor. | Under Investigation by firm | Mindray DS USA inc sent a recall letter on 6/23/2016 to their customers. |  | 360 | 360 units | Nationwide, Canada | Open, Classified |  |  | 360
Z-2545-2016 | 74880 | No | MGZ | External vein stripper. | Cardiovascular | Cardiovascular | 510(k) | 1.5 mm HYDRO LeMaitre Valvulotome, 40 cm. Sterile |  |  | August 03, 2016 | 2016 | LeMaitre Vascular, Inc. | Wire which deploys the loop may be come detached at the handle causing the device to be inoperable | Process control | LeMaitre Vascular, Inc. sent an Urgent Medical Device Recall letter dated June 29, 2016 to customers via priority mail. The letter identified the affected product, problem and actions to be taken. Customers were requested to locate, quarantine and return the devices. Customers were instructed to complete and return the response form to LeMaitre Vascular, Inc. as a record of notification and reconciliation. Any unused product will be requested to be returned and replaced. For questions call 781-221-2266 ext. 183. |  | 50 | 50 units | Worldwide Distribution -  US Nationwide and in the countries of Australia, BELGIUM, BRAZIL, CANADA, Czech Republic, DENMARK, FINLAND, FRANCE, GERMANY, Italy, Japan, NETHERLANDS, Russia, Spain, SWEDEN, Switzerland, and UK. | Open, Classified |  |  | 2422
Z-2579-2016 | 64785 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS Workstation, versions 3.1.1.x through 3.2.1.x Picture Archiving and Communication System |  |  | November 15, 2012 | 2012 | GE Healthcare It | While merging exams in a test instance, merges carry forward in the production system, when the middle tier is configured to the wrong IP address of the Centricity Exam Manager. When a current and historical exam are opened at the same time, the system asynchronously tries to access the operation specifying how each image should be grouped, causing random image-series grouping errors. | Software design | GE sent an Urgent Medical Device Correction letter dated November 15, 2012. The software has been modified to correct these defects. The modified software was released and installed in user sites under GE Healthcare FMI #85196. |  | 498 | 498 units | Worldwide distribution including US nationwide, Puerto Rico, Australia, Austria, Belgium, Bermuda, CANADA, Denmark, Egypt, ENGLAND, France, Germany, India, Ireland, Hong Kong, Korea, Kuwait, Italy, Malta, Netherlands, Portugal, Qatar, Saudi Arabia, SCOTLAND, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, and UK. | Open, Classified |  |  | 498
Z-2580-2016 | 74828 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Immunodiagnostic Products TSH Reagent Pack, REF/Catalog Number/Product Code 191 2997, Unique Device Identifier No. 10758750000227, IVD --- Ortho-Clinical Diagnostics Pencoed Bridgend CF35 5PZ UK |  |  | July 28, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | A field correction was issued by the firm for VITROS TSH reagent lot 5040 due to a higher than expected frequency of customer calibration curves falling outside of the calibration quality parameters and therefore unusable to process samples. Customers using VITROS TSH Lot 5040 will potentially experience an unexpected increase in the unsuccessful calibrations potentially causing a delay in testing and reporting results for VITROS TSH. | Under Investigation by firm | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter on July 28, 2016, via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received VITROS Immunodiagnostic Product TSH, Lot 5040 to inform them of the issue and the required actions. Instructions were provided to customers. Customers with questions were instructed to contact Ortho Care" Technical Solutions Center at 1-800-421-3311. Foreign affiliates were informed by e-mail on July 28, 2016, of the issue and instructed to notify their consignees of the issue and the required actions.For questions regarding this recall call 908-218-8776. |  | 4144 | Domestic: 2883, Foreign: 1261 | Worldwide Distribution - US to VA, WV, GA, TN, KY, MN, SD, AR, OK and Internationally to Canada | Open, Classified |  |  | 4144
Z-2584-2016 | 74419 | No | MQN | Bone plate. | Dental | Dental | 510(k) | TI Single Vector Distractor Body with Left Foot/30mm; Catalog ID 487.965 |  |  | May 23, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is voluntarily initiating a recall due to a possibility the fastener on the Titanium Single Vector Distractor Body may become prematurely separated from the Proximal Foot Plate. | Under Investigation by firm | The firm, DePuy Synthes, mailed out an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated 5/23/2016 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory and immediately take the following actions: If you DO have any of the identified devices call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, complete the Verification Section (page 3 of the letter) indicating the number of devices that were found and return the Verification Section (page 3 of the letter) with the product to the following: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132, by fax 844-449-8498 or scan/email: Synthes8096@stericycle.com., and complete the Verification section even if you DO NOT have the identified product.Note: If the Verification Section is answered on behalf of more than one facility and/or individual clearly indicate the name and address of the facility and/or individual.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 259 | 259 units | Worldwide Distribution-US (nationwide) distribution to states of: AZ, CA, CO, DC, FL, IL, MA, MI, NJ, NY, PA, TX, UT, WI, and WV; and countries of: Australia, Canada and Switzerland. | Open, Classified |  |  | 1817
Z-2403-2016 | 74843 | [6 empty] | OraQuick Ebola Rapid Antigen Test, Product Code 3001-2807; Intended for the qualitative detection of antigens for Ebola viruses. |  |  | July 20, 2016 | 2016 | OraSure Technologies, Inc. | Failed stability testing at 8 months. This lot produced intermittent false negative results for the positive kit control with devices stored at 30C conditions. | Process control | OraSure contacted the sole customer via email and mail (letter dated July 21, 2016) to request return of all the product from the affected lot. |  | 3075 | 3,075 units | Distributed to one customer only in US state of Georgia. | Terminated | May 10, 2017 | 294 | 3075
Z-2552-2016 | 74781 | No | LTI |  |  | Gastroenterology/Urology | PMA | ReShape Integrated Dual Balloon System, Gen 1, Catalog No. RSM110, Model No. 01-0013-001 |  |  | May 04, 2016 | 2016 | Reshape Medical Inc | Due to a suspected misalignment of the connection between the proximal balloon cap of the implant and the delivery catheter. The possible misalignment is the result of an assembly error during manufacturing and may lead to leakage of saline solution during the balloon inflation procedure. | Process control | A Dear Doctor letter dated May 4, 2016, was sent to all the customers to inform them that ReShape Medical is recalling ReShape Integrated Dual Balloon System, Gen 1 due to a suspected misalignment of the connection between the proximal balloon cap and the delivery catheter. The letter informs the customers that the possible misalignment is the result of an assembly error during manufacturing and may lead to leakage of saline solution during the balloon inflation procedure. Customers are informed of the planned field correction and if customers have any questions then they're instructed to contact ReShape Medical Customer Service at (844) 937-7374.For further questions please call (949) 429-6680 ext. 106 |  | 56 | 56 units | US: Distribution to the states of : NY, OH and CA. | Terminated | August 17, 2016 | 105 | 56
Z-2565-2016 | 74602 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Valleylab Laparoscopic Wire J-Hook Electrode 36CM |  |  | July 08, 2016 | 2016 | Medtronic | Potential for compromise of the package resulting in a breach of the sterile barrier | Process control | Medtronic/Covidien sent an Urgent Medical Device Recall letter dated July 8, 2016, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. The letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Customers with questions were instructed to contact their Medtronic representative or Customer Service at 800-882-5878. |  | 349629 | 349629 | Worldwide Distribution - US (nationwide) and Internationally toAustralia, Canada, Japan, Korea, Hong Kong, Thailand, | Open, Classified |  |  | 995222
Z-2553-2016 | 74403 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | METS SMILES Total Knee Replacement |  |  | May 31, 2016 | 2016 | Stanmore Implants Worldwide Ltd. | The manufacturer has identified that the stated Femoral Plateau Plates may not fit into the intended Femoral Knee mating component and thus sit flush upon the Femoral Knee component prior to implantation. | Device Design | The recalling firm sent recall notification letters to their consignees on 05/31/2016. Follow up communciation was sent out on 7/26/2016 after a technical and medical assessment was completed. |  | 3 | 3 units | Directly distributed to TX, further distributed to AL and MS | Open, Classified |  |  | 3
Z-2593-2016 | 66448 | No | GCX | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Hospira RECEPTAL CANISTERS: 1L(1000mL), 1.5L(1500mL), 2L(2000mL) |  |  | September 10, 2013 | 2013 | Hospira Inc. | Hospira has become aware of customers using the incorrect size Receptal liners with their Receptal canisters, e.g., 1.5 L liner, 2L liner or 3L liner with a 1L canister. It is possible that if an incorrectly sized liner is used with a different size canister the lid may not fit and the vacuum/suction will not be sufficient and/or consistent. | Process control | URGENT MEDICAL DEVICE RECALL letters dated September 10, 2013 were distributed to direct accounts notifying them of the recall. The letter requests that customers to ensure that they only use liners and canisters of corresponding sizes (e.g. a 1000mL liner with a 1000mL canister) and to only use Hospira components. The letter further advises customers of an ongoing recall of the 1000mL Receptal systems. Customers were requested to complete and return the attached reply form and to notify other healthcare providers within their organization.Customers who further distributed the Receptal systems were requested to forward the letter to their customers and to request that the contact Stericycle at 1-888-943-5175 (Monday - Friday, 8:00 AM - 5:00 PM, CST) to obtain a reply form. |  | 4544224180761321 | 67,410,861 liners and 67,410,861 lids/canisters in total | *** US: Nationwide (AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY), District of Columbia, Puerto Rico; *** FOREIGN: Australia, Canada, Costa Rica, Singapore, South Africa | Open, Classified |  |  | 4544224180761321
Z-2591-2016 | 73465 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | PEDIATRIC KIDNEY BIOPSY TRAY |  |  | October 05, 2015 | 2015 | Avid Medical, Inc. | Avid Medical Inc. states a latex-containing Telfa" Non-adherent Dressing 4"x3" (10.2cm x 7.6 em) was included in a limited number of convenience kits (custom procedure trays) that were not labeled as containing latex. | Employee error | Avid Medical Inc. states a latex-containing Telfa" Non-adherent Dressing 4"x3" (10.2cm x 7.6 cm) was included in a limited number of convenience kits (custom procedure trays) that were not labeled as containing latex. |  | 20 | 20 | Recalled product was distributed to 2 states: Virginia and Massachusetts. | Terminated | January 10, 2017 | 463 | 700
Z-2587-2016 | 74595 | No | HJO | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | OPTICAL BIOMETER AL-Scan: Software Version: 1.09 and earlier. |  |  | July 01, 2016 | 2016 | Nidek Inc | Nidek Inc. received information from our manufacturer NIDEK CO. Japan that the shown axiallengths were much shorter at AL-Scan measurement. | Software in the Use Environment | Dear Doctor/Recall letters were sent to affected customers on July 7, 2016. The letters describe the issue and provide instruction should the issue occur. The letters inform users that Nidek will reach out in the next few weeks to schedule installation of new software that will solve the problem.Questions or concerns may be directed to:Customer Service 1.800.722.0219; Monday- Friday 8 am to 5pm PSTE-mail: AL-Scan_support@noritsuservice.com. |  | 17 | 17 | US Nationwide Distribution in the states of: NY, PA. GA. AZ, VA, DE, ME,. MO, TN, and NM. | Open, Classified |  |  | 17
Z-2588-2016 | 74752 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PROFEMUR(R) MODULAR NECK, HIP FEMORAL NECK MicroPort total hip systems intended for use in total hip arthroplasty for reduction or relief of pain and/or improved hip function in skeletally mature patients. |  |  | June 30, 2016 | 2016 | MicroPort Orthopedics Inc. | Profemur(R) neck, Item Number PHA01222 (Short AR VV2 Ti Neck) was labeled as Item Item Number PHA01212 (Short AR VV1 Ti Neck). | Error in labeling | The firm initiated their recall by email on 06/30/2016. |  | 14 | 14 units | Belgium, Italy, Spain, and Poland | Open, Classified |  |  | 14
Z-2606-2016 | 74800 | No | LOD | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | Cobalt HV Bone Cement |  |  | July 26, 2016 | 2016 | Encore Medical, Lp | The outer packaging was mislabeled on the box indicating "Cobalt HV with Gentamicin". The bone cement does not contain antibiotics. | Labeling mix-ups | The recalling firm sent an Urgent Field Safety Notice letter dated July 26 2016 to Sales Agents. The letter identified the affected product, problem and actions to be taken. The sales agents were instructed to pass on the notice to any organization where the potential affected product has been transferred. Agents were instructed to identify hospitals in their territory, complete a visually inspection of the affected product in the hospital's inventory, complete an Acknowledgement and Receipt form. If affected products is found contact Customer Service at 1-800-456-8696 for replacement. |  | 866 | 866 units | US Nationwide Distribution | Terminated | January 13, 2017 | 171 | 866
Z-2603-2016 | 74801 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TIGR Matrix Surgical Mesh, Model No. NSTM1015, NSTM1520, NSTM2030 |  |  | July 19, 2016 | 2016 | Novus Scientific Ab | Novus Scientific is implementing additions to the product's instructions for use concerning certain inguinal hernia repairs. | Other | An Urgent Safety Alert letter dated 7/19/16 was sent to all their customers to inform them that Novus Scientific is implementing additions to the product's instructions for use (IFU) with a new contraindication and warning related to the clinical study result. The letter informs the customers that until the new IFU is made available, surgeons using the product for hernia repair should follow the instructions in the urgent safety alert in addition to those in the current IFU.Customers are informed of the actions to be taken:" Additional contraindication: Not suitable for the repair of direct inguinal hernias." Additional warning: Because TIGR Matrix Surgical Mesh is fully resorbable, it should not be used in repairs where permanent support from the mesh is required.Customers with any questions or concerns are instructed to contact Novus Scientific AB ( +46 (0)18 700 11 50. |  | 3458 | 3458 units total (657 units in US) | US : IL, CA, FL, KY, MO, NH, VA, TX, NC, OH, NY, TN, Worldwide: Sweden, Italy, Denmark, UK, Singapore, Germany, Belgium, Spain, Austria, Israel, Malta, Switzerland, Ireland, Norway, Netherlands, Turkey, New Zealand, Bulgaria, France, Iran, Finland, Poland, South Africa | Terminated | December 23, 2016 | 157 | 3458
Z-2604-2016 | 74884 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE VC150 Vital Signs Monitor; Intended to monitor a single patient's vital signs at the site of care. |  |  | July 27, 2016 | 2016 | INNOKAS MEDICAL OY | A software error on released software versions 1.6.12, 1.6.12F and 1.6.16 may give wrong time data to measurements. | Software design | Innokas Medical has notified the one US distributor on 7/27/16 via letter of the deficiency and correction of the device. |  | 1715 | 1715 units | AR, AZ, FL, IN, LA, MA, MI, NC, NJ, NM, NY, OH, PA, TN, WI | Terminated | April 19, 2017 | 266 | 1715
Z-2595-2016 | 73607 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Adult ThermoGardÂż Dual Dispersive Electrodes (for patients &gt;15 kg.), 10' (3.05m) Cable, Catalog Number 51-7310. |  |  | May 18, 2016 | 2016 | Conmed Corporation | A version of the dual dispersive electrodes may not be compatible with some electrosurgical generators, placing patients at risk for undetected pad lift and a potential burn. | Device Design | On May 17, 2016 CONMED Corporation distributed Urgent Device Recall notification letters and Business Reply Forms to their customers via courier service. CONMED Corporation has decided to recall these ThermoGardÂż Dual Dispersive Electrodes for all customers who do not use an electrosurgical generator that employs the Contact Quality Monitoring systems listed. If your electrosurgical generator is not listed, your generator should not be used with the lot codes of the ThermoGard Dual Dispersive Electrodes identified. Customers are advised to inspect inventory for any of the recalled devices. If you use the ThermoGardÂż Dual Dispersive Electrodes in conjunction with an electrosurgical generator listed, no further action is required. However, if your electrosurgical generator and CQM system is not listed above, do not use the affected lot codes of ThermoGardÂż Dual Dispersive Electrodes. Please note that ThermoGardÂż Dual Dispersive Electrodes produced after July 5, 2015, are not subject to this recall. Customers should contact all of those departments or organizations within their facility and any other facilities that you may have supplied or given these affected products to. It is imperative that all end users of these devices receive this notice and respond immediately. If you have questions, please contact Patricia Cotter, CONMED Recall Coordinator 315-624-3237 (8AM-5PM EST, M-F) or fax to 315-624-3225 or email thermog@CONMED.com. |  | 307684 | 307,684 units | Worldwide Distribution -- USA, including AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, &amp; Puerto Rico; and, the countries of Australia, Canada, Belgium, Brunei Darussalam, China, France, Guatemala, Hong Kong, Israel, Italy, Japan, South Korea, Malaysia, Mexico, New Zealand, Norway, Portugal, Saudi Arabia, Spain, Sweden, and United Arab Emirates. | Open, Classified |  |  | 447697
Z-2360-2016 | 74665 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Alere INRatio2 PT/INR Professional Monitoring System |  |  | July 11, 2016 | 2016 | Alere San Diego, Inc. | Alere San Diego is initiating a voluntary recall for the Alere INRatioÂż/INRatioÂż2 PT/INR Monitoring Systems (Test Strips and Meters). This action is being initiated as a result of patients who had a therapeutic or near-therapeutic INR with the Alere INRatioÂż/INRatioÂż2 PT/INR Monitoring System but a significantly higher INR (outside of therapeutic range) when performed by a central laboratory. | Other | Alere issued a press release on July 11, 2016. An Urgent Medical Device Letter dated July 26, 2016 was sent to affected customers on July 29, 2016. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to consult with their healthcare provider as soon as possible to transition to an alternate method of PT/INR testing and to contact the firm regarding return or disposal of unused product and obtaining device replacement from a different manufacturer. For questions contact the Alere INRatid1 Recall Hotline at 1-866-723-2535. |  |  |  | Worldwide Distribution - US Nationwide and foreign distribution to the countries of: Argentina, Australia, Azerbaijan, Bangladesh, Belarus, Belgium, Brazil, Brunei Daruss, Canada, Chile, China, Colombia, Egypt, France, Hong Kong, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Malaysia, Maldives, Netherlands, Russia, Rwanda, Singapore, South Africa, South Korea, Spain, Taiwan, Thailand, Trinidad, Ukraine, Uruguay, and Vietnam. | Terminated | July 20, 2017 | 374
Z-2605-2016 | 74737 | No | FNM | Alternating pressure air flotation mattress. | General Hospital | General Hospital | 510(K) Exempt | Panacea Air Pro Mattress Control Unit, Model: Platinum 5000, US Model Number 5000. |  |  | July 11, 2016 | 2016 | American National Mfg Inc | The Panacea Air Pro Mattress Control Unit is being recalled due to a firmware problem, which may result in the device not having the firmness range required. | Software Manufacturing/Software Deployment | An Urgent Medical Device Recall letter, dated 7/11/16, was sent to customers to inform them of a recall involving the Panacea Air Pro Mattress Control Unit also labeled as Platinum 5000, US Model Number 5000, Serial numbers 9500051 through 9500300. The letter informs the customers that the recall has been identified due to firmware problem. The letter informs the customers of the problem identified and the actions to be taken. Customers were asked to immediately examine their inventory and quarantine the affected device. A shipping label would be provided for return and update of the affected devices. All customers that the device was further distributed to would need to be notified of the recall. |  | 31 | 31 units | Distributed to the US -- WI, OH, IL, TX, AZ, MD, IA, PA, CO, WA, ID, MN, LA, WY, and MS. | Open, Classified |  |  | 31
Z-2608-2016 | 74656 | No | HHI | Vacuum abortion system. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Conical portio adapter (acorn) small &amp; large are used for examination and diagnosis in conjunction with endoscopic accessories in gynecological laparoscopy, to determine patency of the fallopian tubes. |  |  | May 19, 2016 | 2016 | Richard Wolf Medical Instruments Corp. | The distal end of the Conical portio adapter (acorn) is composed of a small cap (head) that partially sits inside a hollow shaft (tube). These two parts are welded together to create one piece. In the impacted batches of product, the weld was improperly performed. As a result, the two pieces are not properly joined. This potential failure can result in the separation of the two pieces. The possible separation of these two components presents a potential risk to patients. | Nonconforming Material/Component | Customers impacted were notified via Urgent Medical Device Recall Letter dated 5/31/2016 via certified mail. The letter describes the reason for the recall. Richard Wolf Identified the following measures to be taken by the consignees:1. Immediately examine your inventory and quarantine product subject to recall. Richard Wolf Medical Instruments Corporation strongly recommends not using the potentially defective products. Please make sure that all users of the aforementioned products become aware of this urgent safety information. 2. If you have further distributed this product, please identify your customers and immediately notify them about this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter, but must include instructions on what customers should do with the recalled product. 3. PLEASE COMPLETE ENCLOSED ANSWER FORM {even if the product is no longer in your inventory) and return the record to Richard Wolf Medical Instruments Corporation before June 17, 2016. Your return of the answer form will: " Confirm receipt of this safety information" Trigger a replacement order. Richard Wolf Medical Instruments Corporation will provide your facility with a return authorization number and shipping instructions and will bear the cost for shipping (returns to RWMIC and replacements to facility). 4. Please immediately inform Richard Wolf Medical Instruments Corporation of any adverse events.If the consignees have any questions they are asked to contact Lisa Williams, Regulatory Affairs Assistant at Richard Wolf Medical Instruments Corporation Richard Wolf Medical Instruments Corporation.Telephone: (847) 913-1113 Ext: 256 Telephone: (847) 913-1113 Ext: 225Fax: (847) 913-0924 Fax: (847) 913-0924Email : lwilliams@richardwolfusa.comMonday through Friday, 7:30AM to 4:00 PM CST. Monday through Friday, 9:00AM to 5:30 PM |  | 16883.81 | Product code 8378.90- 94 pieces &amp; Product code 8378.91 - 32 pieces | U.S. Consignees: AK, AL, AR, AZ, CA, CT, DC, FL, GA, IA, IL, KY, MA, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OR, TX, UT and VA. | Open, Classified |  |  | 16883.81
Z-2602-2016 | 74545 | No | MHY |  |  | Neurology | PMA | Medtronic, Lead Kit for DBS Stimulation. Model Numbers 3387S-40 (p/n 3387S0007V) &amp; 3389S-40 (p/n 3389S0008V). |  |  | June 10, 2016 | 2016 | Medtronic Neuromodulation | Medtronic has voluntarily decided to retrieve unused Model 3387S-40 and 3389S-40 DBS leads from three manufacturing lot numbers.During the manufacturing process of a DBS lead component, Medtronic identified the potential for lead insulation damage. | Process design | Consignees were delivered in person by Medtronic Field Representatives a "Medical Device Removal" letter dated June 2016. The letter described the problem and the product involved in the recall. Advised consignees to return the unused product and to complete the Customer Confirmation Form. For questions contact Medtronic Technical Services at 800-707-09033 weekdays 7 am - 6 pm, or contact your Medtronic representative. |  | 70 | 70 | US: AR, CA, FL, IL MD, MO, NE, NY, OH, TN WI. | Open, Classified |  |  | 70
Z-2607-2016 | 74252 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Radiesse (+) Lidocaine, Injectable Implant 1.5cc, Rx only, P/N: 8063M0K1, Merz Aesthetics. |  |  | May 20, 2016 | 2016 | Merz North America, Inc. | Complaints reported of Radiesse (+) unable to be expelled from the syringe and/or needle. | Nonconforming Material/Component | Merz sent an Urgent Medical Device Recall Letter dated May 20, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to complete the enclosed response form and return via fax to (262) 835-3330 or email to complaints2@merz.com. For questions contact Merz at (262) 835-3300 ext. 2627. |  | 1591 | 1,591 units | US Nationwide Distribution in the states of AR, CA, CO, CT, DC, FL, GA, IL, IN, LA, MA, MD, MI, MN, MO, NC, ND, NJ, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, VT and WA. | Terminated | May 08, 2017 | 353 | 1591
Z-2600-2016 | 74855 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Large QWIX Compression screw, dia. 7.5 mm L 90 mm, Lot No. FH86, Rx only, |  |  | July 21, 2016 | 2016 | Integra LifeSciences Corp. | Labeling on the Qwix Screw product states that it is a "compression screw" instead of a "positioning screw". | Labeling mix-ups | Integra sent an Urgent Voluntary Medical Device Recall letter and acknowledgement return response form dated July 21, 2016 to their affected customers. The recall letter identifies the product and lots involved, problem and actions to be taken. All customers are asked to return the completed acknowledgement and return response form by email to FCA1@integralife.com or fax to 1-609-275-9445. Should any customer has any questions regarding the instructions, they can contact Customer Service at 1-800-654-2873. |  | 120 | 120 units from multiple lots | US Nationwide Distribution in the states of AL, GA, ID, MA, MD, MN, NC, NJ, NY, OH, OK, PA, PR, TX, VA and WA | Open, Classified |  |  | 120
Z-2601-2016 | 74564 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens RAPIDPointÂż 500 v2.2.2A Software Upgrade Kit; SMN 11066719, Software Version 2.2A |  |  | July 05, 2016 | 2016 | Siemens Healthcare Diagnostics Inc | Some v2.2.2 upgrade kits include a dialysate mode which not cleared/approved for shipment in the United States | Under Investigation by firm | Siemens sent a Customer Notification dated July 2016 to all affected domestic (US) customers via Federal Express on July 5, 2016. Customers were asked to take the following actions: Please do not use the dialysate mode. If you have enabled this feature, please disable it by following these instructions:1. At the System screen, select Setup. 2. Select Secured Options &gt; Analysis Options. 3. Unselect Dialysate. 4. Select the Continue button 3 times.Please review this letter with your Medical Director. Complete and return the Field Correction Effectiveness Check attached to this letter within 7 days.Customers with questions were instructed to contact their Siemens Customer Care Center or their local Siemens Technical Support representative.For questions regarding this recall call 781-269-3000. |  | 76 | 76 | Nationwide Distribution to AL, CA, MA, NJ, OH, TX, WA, PA | Terminated | August 02, 2017 | 393 | 76
Z-2609-2016 | 66606 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 1) The HomeChoice automated peritoneal dialysis system. Dataplate located on the device is labeled in part: Baxter. |  |  | October 07, 2013 | 2013 | Baxter Healthcare Corp. | In systems with version 10.4 software Initial Drain logic, the device will attempt to drain the patient to empty at the beginning of therapy (initial drain). The operator can stop, but cannot bypass, the active Initial Drain. This is to mitigate against Unintended Increased Intraperitoneal Volume (IIVP). This can cause serious problems in patients with unrelated a co-morbid condition of ascites. | Labeling design | An IMPORTANT PRODUCT INFORMATION letter dated October 7, 2013, was sent to all affected customers. The letter included instructions for clinicians to: 1) be aware that HomeChoice will attempt to drain the patient to empty at the beginning of therapy (Initial Drain); 2) review HomeChoice instructions for use for additional details on HomeChoice 10.4 Initial Drain Logic; 3) acknowledge your receipt of the notification by completing the attached Customer Reply Form and returning by fax to either 224-270-5457 or scanning and e-mailing it to fca@baxter.com; 3) if the recipient of the letter is a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, to forward the communication to all customers who may have received these products. Clinical questions should be addressed to Baxter's Renal Division Clinical Helpline at 888-736-2543, option 2 (Monday - Friday, 8:00 AM - 5:00PM CST). General questions about this recall should be addressed to The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 65590 | *** 1) Product Codes 5C4471 and 5C4471R: Approximately 48,600 units; *** 2) Product Codes 5C8310 and 5C8310R: Approxiamtely 16,990 units | Worldwide distribution. US nationwide including Puerto Rico; Australia, Canada, New Zealand, and Japan | Completed |  |  | 65590
Z-2610-2016 | 74010 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | BeaconÂż Tip Torcon NBÂż Advantage Catheter |  |  | April 15, 2016 | 2016 | Cook Inc. | Increase in reports of polymer degradation of the catheter tip, resulting in tip fracture and/or separation.degradation could include loss of device function, separation of a device segment leading to medical intervention, or complications resulting from a separated segment. Such complications include device fragments in the vascular system, genitourinary system, or other soft tissues. Fragments within the vascular system could result in embolization to the heart or lungs, or occluding blood flow to end organs. | Unknown/Undetermined by firm | Cook Inc. sent an URGENT: MEDICAL DEVICE RECALL NOTIFICATION letter dated April 15, 2016, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers with any medical questions or concerns, were instructed to contact Cook Medical Customer Relations at 800.457.4500 or 812.339.2235. For information regarding the recall, please contact Stericycle Expert Solutions at 866.201.9067. |  | 2662505 | 2,662,505 | Worldwide Distribution - US (Nationwide) and Internationally to Algeria, Andorra, Argentina, Armenia, Australia, Austria, Bahrain, Bangladesh, Barbados, Belgium, Bermuda, Brazil, Brunei Darussalam, Bulgaria, Canada, Canary Islands, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Djibouti, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guatemala, Guernsey, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jersey, Jordan, Kazakhstan, Kenya, Korea, Kosovo, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Macau, Malaysia, Malta, Martinique, Mauritius, Mexico, Monaco, Mongolia, Morocco, Myanmar, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, "Palestinian Territory, Occupied", Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad And Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, U.S. Virgin Islands, Yugoslavia. | Terminated | August 04, 2017 | 476 | 4146309
Z-2627-2016 | 74919 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 31, 2016 | 2016 | Merge Healthcare, Inc. | Potential for RadSuite AV viewer to skip image slices when the user presses page up or page down when the viewer displays multiple viewports and the user has mouse focused on any viewport other than the first viewport. | Software design | Merge Healthcare sent an Urgent Medical Device Recall letter dated January 28, 2016, and via email on January 31, 2013, to all affected customers. the letter identified the product the problem and the action needed to be taken by the customer.The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it. If you have any additional questions call Merge Customer Service at (877) 741-5369.For further questions, please call (262) 367-0700. |  | 25 | 25 sites potentially have the affected versions | Distribution was made to AL, AZ, CA, IL, MD, MA, MI, MO, NV, NJ, NC, OH, PA, TN, and VA. Government distribution was made to OK. There was no foreign/military distribution. | Open, Classified |  |  | 25
Z-2630-2016 | 74759 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 120 V Neptune 3 Rover, Model Number: 0703-001-000 |  |  | July 01, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker initiated a voluntary recall of the 120 V Neptune 3 Rover on 07/01/2016, due to a potential for the device to experience system errors causing the device to shutdown if high levels of electromagnetic interference are present in the operating room. | Under Investigation by firm | Stryker Instruments initiated a voluntary recall of the Stryker Neptune 3 Rover via certified mail on 07/01/2016, due to the potential for the Neptune 3 Rover to experience system errors if there are high levels of electromagnetic interference in the operating room. 1. Immediately review the Recall Notification.2. Confirm that the Rover serial numbers identified on the Business Reply Form (BRF) are at your facility.3. A Stryker representative will facilitate the repair of your affected Neptune 3 Rover(s). Please coordinate the availability of the product to assure the repair can be implemented.4. Please complete and return the Business Reply Form (BRF).Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification.5. Fax the completed, signed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com.6. Stryker will document the repair of each unit by serial number.For questions regarding this recall, please contact Stryker Instruments:Kelly Jo Davis269-389-2921kellyjo.davis@stryker.com |  | 78 | 78 | Domestic:MI, CA, ID, IAVA/DOD: NoneForeign: None | Terminated | June 02, 2017 | 336 | 78
Z-2624-2016 | 74840 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition AS with software version VA48A-SP2; Model # 10430603, computed tomography x-ray system |  |  | July 13, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Due to an internal communication error between the firmware and the software of the components, the planned CT scan executes properly, but the injector is not started. Therefore, the contrast agent is not injected and the desired examination result is not achieved. This error only affects the automatic mode or coupled mode and does not affect the manual control of the injector. | Software in the Use Environment | Siemens mailed to customers an Important Customer Safety Notice to inform about a possible malfunction when using the affected systems and what actions to take to eliminate the problem. Also, Siemens informed their customers that they were currently developing a new firmware to eliminate the internal communication error, which should be released in August 2016. Customers were asked to contact their service organization at 800-888-7436, if they had any questions. |  | 1 | 1 unit | Distributed to: MI, NY, CA, KY, ND, NE | Open, Classified |  |  | 7
Z-2628-2016 | 74808 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC UroView 2800, fluoroscopic x-ray system |  |  | July 22, 2016 | 2016 | GE OEC Medical Systems, Inc | A firmware issue in the supplied LCD workstation monitor(s). Under some circumstances, the monitor(s) may appear dark with no video displayed. | Software design (manufacturing process) | On July 22, 2016 GE Healthcare Surgery began distributing an Urgent Medical Device Correction letter to customers via mail to inform them of the issue and provide instructions for immediate mitigation of the issue. The letter was sent to 3 titles at each facility affected: 1.) Facility Administrator, 2.) Director/Manager of Radiology and Radiologists, 3.) Radiology Department. Consignees instructured that local GE Healthcare Field Engineer will contact them to coordinate replacement of the affected monitor(s), and will ensure that system meets specifications. |  | 47 | 47 total | Distribution US nationwide and Korea. | Open, Classified |  |  | 47
Z-2631-2016 | 74850 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | RESPONSIVE, KNEE: FEMORAL COMPONENT CR/CS, SIZE 3R, REF 90-SRK-311300, QTY: 1, STERILE R, |  |  | July 18, 2016 | 2016 | Medtronic Sofamor Danek USA Inc | The alpha total knee replacement components that may have not been passivated in accordance with the requirements of the job traveler. | Employee error | On 07/18/2016, the firm initiated the correction by verbally notifying (by visit and telephone) three of four (3 of 4) implanting surgeons; the fourth implanting surgeon was verbally notified on 07/25/2016. The firm followed with a letter to all four (4) surgeons on 07/28/2016. Medtronic recommended that the condition of the patients receiving the subject part be assessed for any indications of premature failure during normal post-operative follow-up visits. |  | 3 | 3 units | AR, MN,MA | Open, Classified |  |  | 16
Z-2645-2016 | 74552 | No | MHY |  |  | Neurology | PMA | DBS Pocket Adaptor, Model 64001 and Model 64002. |  |  | June 23, 2016 | 2016 | Medtronic Neuromodulation | Due to the potential for pocket adaptor conductor wire fractures involving DBS Pocket Adaptor Model 64001 and Model 64002, a safety notification was sent to healthcare professionals to reinforce device specific labeling for the handling of pocket adaptors and system integrity checking during implant procedures. | Nonconforming Material/Component | Consignees were sent on 6/23/2016 a "Medical Device Safety Notification," dated June 2016. The letter was addressed to Healthcare Professional. The letter described "Background" information on the problem and product involved in the recall. It also provided "Recommendations" which included "Implanting the pocket adaptor with the neurostimulator" and "Checking system integrity".For questions contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7 am - 6pm CST. |  | 20541 | 7,542 in the US and 12,999 outside the US. | Worldwide Distribution. | Open, Classified |  |  | 20541
Z-2668-2016 | 74607 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeamÂż Radiotherapy Delivery System and EDGE" Radiotherapy Delivery System, K140528. The TrueBeam and Edge Systems are intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body where radiation therapy is indicated for adults and pediatric patients. |  |  | June 30, 2016 | 2016 | Vision RT Ltd | Vision RT Ltd received three (3) reports from different sites that discovered the following issue, unintended changes were made to planned couch parameters, specifically the couch rotation parameter, during patient set-up for the Optical Surface Monitoring System [OSMS].1. CRM 3194, No Serious Injury, No MDR, Aware date: Dec 11, 2015 2. CRM 3321, No Patient Involved, No MDR, Aware date: Jan 19, 20163. CRM 4570, No Serious Injury, No MDR, Aware date: March 15, 2016No patient harm was reported in any of these cases. | Other | Vision RT Ltd will provide all affected AlignRTÂż Plus Users with an Urgent Medical Device Correction Notification letter, dated 2016-06-24 Subject: Unexpected Overwrite of Plan Couch Parameter. The letter will provide Users with a detailed description of the problem and provide instructions to avoid the issue. Non-responders will be notified again until we achieve 100% for North America and International consignees. Vision RT will make three (3) attempts using at least two different methods of contact to obtain proof of notification.Containment: Vision RT will provide a copy of the Urgent Medical Device Correction letter for any site accepting an AlignRT device. The letter will provide Users with a detailed description of the problem and provide instructions to avoid the issue.Product will not be returned to Vision RT.Varian Medical Systems and Vision RT are investigating development of a technical fix for the affected products. Notification of the intention to correct is provided in the Urgent Medical Device Correction Notification letter, dated 2016-06-24 Subject: Unexpected Overwrite of Plan Couch Parameter. Vision RT Ltd will contact all affected AlignRTÂż Plus users and schedule upgrades to implement the correction when available. |  | 47 | 47 | Product distributed in US.The foreign countries where the product was distributed are as follows: Canada, Germany, Japan, Sweden, Switzerland, Taiwan, and United Kingdom.No product was distributed to US government accounts. | Open, Classified |  |  | 47
Z-2639-2016 | 74013 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Ziehm Vision RFD 3D Mobile C-arm for Mobile Fluoroscopic Imaging |  |  | April 15, 2016 | 2016 | Ziehm Imaging Inc | Malfunction of the radiation-warning timer and the 3D-specific interval warning. No 5-minute timer dialogs or beeper will occur until the unit is rebooted. | Device Design | Ziehm Imaging, Inc planned action to bring these defects into compliance: 1. You will contact customers and initiate a software update to correct the defects.2. You will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.For further questions, please call (866) 949-4346. |  | 31 | 31 devices | US Distribution including Puerto Rico and to the states of :FL, GA, VA,MA. ME, MI, NY, NJ, PA, RI, IL, TX, AZ, CA, NV, UT, and WA. | Open, Classified |  |  | 62
Z-2646-2016 | 74790 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BVX-330BR Bovie Precise 360 Handpiece 33cm, blade |  |  | July 18, 2016 | 2016 | Bovie Medical Corporation | The adhesive may be inadequately applied or the assembly may be manipulated after application of the adhesive resulting in an incomplete bond to the angled tip. | Nonconforming Material/Component | Consignees were notified by written recall notice sent via certified mail or courier on July 18, 2016 to quarantine their inventory of the Precise 360 Handpiece, identify and contact their customers that have received the affected product and provide the a copy of the attached field safety notice. Customers are requested to return affected product, and return all affected product currently in their possession. All customer inventory should be reviewed, discontinue use, and return all Bovie Precise 360 handpieces to: Quality Assurance Department, Bovie Medical Corporation, 5115 Ulmerton Road, Clearwater, Florida 33760, Attn: Precise 360 Recall. Also, please write the authorization number RMA XXXX on the outside of the shipping container and use UPS account number 342562 and return via ground shipping and Fax or email the attached recall return response form to the email or fax indicated on the form as soon as possible and include a copy with the shipment. A completed recall return response form is required even if the customer has no product to return. If you have any questions about this recall or the product affected by the recall, call 727-803-8617. |  | 95 | 95 | Worldwide Distribution - US Nationwide in the states of: MD, GA, MO and FL and the country of Finland. | Open, Classified |  |  | 95
Z-2658-2016 | 74820 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core128, Model number 728323 |  |  | July 20, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. | Software Design Change | Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. ForNorth America and Canada, contact the Customer Care Solutions Center (1 -800-722-9377). |  | 3 | 3 | US Nationwide Distribution in the states of: AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,NC,ND,NE,NJ,NM,NV,NY,OH,OK,OR,PA,PR,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY | Open, Classified |  |  | 615
Z-2650-2016 | 74923 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic, Viva Quadripolar Implantable Cardioverter Defibrillators: |  |  | August 12, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | 78 Viva CRT -Ds and Evera ICDs that were manufactured with a specific subset of circuit components may experience rapid battery depletion due to a low resistance path developing within the circuit component. There have been seven (7) confirmed failures (9%) through July 16, 2016. Reported complications have included shortness of breath, pocket heating, low heart rate, and early device explant. | Nonconforming Material/Component | Consignees were sent or hand delivered on August 12, 2016 a Medtronic "Urgent Medical Device Correction" letter dated August 2016. The letter was addressed to Physician. The letter described the problem and the devices involved in the recall. The letter described a number of "Patient Management Recommendations". For questions contact your local Medtronic Representative or Medtronic Technical Services at 800-723-4636. |  | 78 | 78 (39 US, 39 OUS) | Worldwide Distribution - US Nationwide in the states of US: AZ, CA, FL, GA, IL, KY, MD, MN, NJ, NY, NC, OH, OK, PA, PR, TX. and countries of: Australia, Finland, France, Germany, Iran, Italy, Japan, Korea, Netherlands, Puerto Rico, Saudi Arabia, Spain, and United Kingdom. | Open, Classified |  |  | 78
Z-2660-2016 | 74806 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Infinite F50 instrument. |  |  | July 25, 2016 | 2016 | Tecan US, Inc. | The batteries have the potential to overheating, melt or char. | Nonconforming Material/Component | Tecan contacted all customers by Safety Announcement, via certified mail starting July 25, 2016. The letter identified the product, the problem, and the action to be taken by the customer. Because affected batteries pose a fire and burn hazard, it is extremely important that customers check whether their battery is affected. Customers were advised to go immediately to the HP Battery Recall website at: http://www.HP.com/go/batteryprogram2016 to check if their battery is affected by this recall. Customers were also advised to complete and return the Customer Response Form at QA@tecan.com or fax it to 919-287-2961.Customers with questions should contact HP at 1-888-202-4320 or Tecan Technical Service's Helpdesk at 800-832-2687, option 1. |  | 1 | 1 unit | Worldwide Distribution - US including AK, AL, AR, CA, FL, GA, HI, KY, MA, MD, NC, NY, PA, PR, SC, TN, TX, UT, VA. International account: Canada. | Open, Classified |  |  | 65
Z-2659-2016 | 74937 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NovaPACS versions 8.4.7, 8.5.3, and 8.5.6. Distributed by Novarad Corporation. Picture archiving and communication system (PACS) |  |  | August 12, 2016 | 2016 | Novarad Corporation | Novarad Corporation announces a voluntary field action for the NovaPACS versions 8.4.7, 8.5.3, and 8.5.6 due to a "pop-up screen" that was prematurely released. | Software design | Novarad sent an Urgent Medical Device Correction letter dated August 12, 2016, to all affected consignees. As a workaround, customers are advised to use the HTMLS Viewer for review purposes only and not for diagnostic purposes. The issue has been fixed in all future releases of NovaPACS software. You can eliminate the possibility of incorrect information appearing by contacting Novarad Technical Supoort to arrange for an upgrade to a corrected version of NovaPACS software. This correction is available to all affected customers at no cost and will be included din all future upgrades.Customers will need to fax, mail, or email th eincluded Acknowledgement and Receipt Form back to Novarad. Customers with questions or concerens regarding this notification or would like to request an upgrade, should contact Novarad Technical Support at 801-221-5895.For questions regarding this recall call 801-642-1001. |  | 90 | 90 | Worldwide Distribution - US (Nationwide) Canada and other foreign consignees. | Open, Classified |  |  | 90
Z-2537-2016 | 74561 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Roadrunner Uniglide Hydrophilic Wire Guide, Catalog Numbers and GPN: HPWA-18-260, G30489; HPWA-18-320, G30490; HPWA-25-260, G30491; HPWA-35-260, G30492; HPWAS-35-260, G30506; HPW-35-150, G56149; HPW-35-180, G56150; HPWS-35-150, G56152; HPWS-35-180, G56153; HPWA-35-80, G56172; HPWA-35-150, G56173; HPWA-35-180, G56174; HPWA-18-180, G56160; HPWAS-35-150, G56176; HPWAS-35-180, G56177. |  |  | June 24, 2016 | 2016 | Cook Inc. | Potential coating contamination with glass particles. | Material/Component Contamination | Cook Inc. initiated a voluntary recall of the Roadrunner UniGlide Hydrophilic Wire Guides on June 24, 2016 via courier due to the potential for the device to be contaminated with glass particles. The firm issued a press release dated August 12, 2016. Customers were asked to take the following actions: 1. Examine inventory immediately to determine if you have affected products and quarantine affected products. 2. Return affected products to Stericycle Expert Solutions (a third-party recall administration service provider). Use the enclosed label and a copy of the Acknowledgement and Receipt Form to receive a product credit.3. Even if you do not have affected products on hand, you must still complete the Acknowledgement and Receipt Form and fax it to 866.796.4780 or email it to cookmedical4502@stericycle.com.4. Report adverse events to Cook Medical Customer Relations at 800.457.4500 or 812.339.2235. Monday through Friday between 7:30 a.m. and 5:00 p.m. (Eastern Time). Or by email at CustomerRelationsNA@cookmedical.com.Should you have any medical questions or concerns, please contact Cook Medical Customer Relations at 800.457.4500 or 812.339.2235. For information regarding the recall, please contact Stericycle Expert Solutions at 866.912.9552.Cook is providing notification to distributors and informing them that this notice must be shared with appropriate personnel, including down to the user level, within your organization or to any organization where the potentially affected devices have been transferred. |  | 8750 | 8,750 devices were distributed in commerce | Worldwide distribution. US (nationwide), Australia, Austria, Belgium, Canada, China, Czech Republic, France, Germany, Hong Kong, Hungary, Italy, Jordan, Korea, Spain, South Africa, Taiwan, Turkey, and United Kingdom | Open, Classified |  |  | 8750
Z-2671-2016 | 64796 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Centricity Laboratory Core Lab System 4.1 |  |  | November 20, 2012 | 2012 | GE Healthcare It | The Centricity Laboratory removes free-text notes on a master panel when the user deletes the slave panel. The deletion of a slave panel will (auto slash) the corresponding test (item) results on the master panel. Additionally, the displayed and HL7 results show neither the appropriate abnormal flag value nor the reference range values, when using large_num.g &gt;7 to report numeric results. | Software design | GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated Nov 20, 2012 to all affected customers. The letter described the product, problems and actions to be taken. The customers were instructed to follow the safety instructions and to ensure that all potential users in their facility are made aware of this safety notification and the recommended actions. The software has been modified to correct these defects. The modified software will be released and installed in user sites through GE Healthcare FMI #86071.Customers were instructed to contact their GE Service Representative with any questions regarding this issue. For other questions, contact GE Helpdesk:North America; (905) 479-1249 or 1-888-778-3375; email: tgg.nahelpdesk@med.ge.comEurope; +44 1603 877595; email: tgg.ukhelpdesk@med.ge.comAsia Pacific; +61 3 5327 4000; email: tgg.ozhelpdesk@med.ge.com.For questions regarding this recall call 847-277-5000. |  | 17 | 17 | Worldwide Distribution - US Distribution to GA, OH, NY, TX, TN ,MN and Internationally to India, Australia, Canada, England, Qatar, and Scotland. | Open, Classified |  |  | 17
Z-2678-2016 | 74927 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Geo-Med, TURP/CYSTO PACK PGYBK REF 89-7931.07, Rx Only, STERILE EO, |  |  | August 11, 2016 | 2016 | DeRoyal Industries Inc | The firm is recalling two custom surgical kits containing the BARD(TM) Toomey Irrigation Syringe 70cc with Resectoscope Tip, Catheter Tip, Luer Adapter and Cap. The BARD(TM) component is being recalled for a possible slit defect on the package that may affect the component's sterility. | Nonconforming Material/Component | DeRoyal issued their recall on 08/11/2016 to the end-user level by sending overnight letters to its direct accounts. The customers were asked to identify the affected trays, place them in quarantine and return them to DeRoyal for replacement or credit. |  | 26 | 26 units | Distributed in LA and TN. | Terminated | June 13, 2017 | 306 | 50
Z-2670-2016 | 74616 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | UNIQ 1.0: UNIQ FD10, 722026; UNIQ FD10/10, 722027; UNIQ FD20, 722028; UNIQ FD20/10 biplane, 722029; UNIQ FD10 OR Table, 722033; UNIQ FD10/10 OR Table, 722034; UNIQ FD20 OR Table, 722035; UNIQ FD20/20, 722038; |  |  | July 07, 2016 | 2016 | Philips Electronics North America Corporation | Increase in the failure rate of certain low-voltage DC power supplies (DCPS) | Component change control | Philips sent an Urgent - Field Safety Notice dated June 20, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Philips Field Service Engineers will replace the affected DCPS free of charge.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722-9377.For further questions regarding this recall please call (978) 659-3000. |  | 1536 | 1536 devices (includes both Allura and UNIQ) | Nationwide, Canada, AlgeriaArgentinaAustraliaAustriaBahrainBangladeshBelgiumBrazilBrunei DarussalamBulgariaCHILEChinaColombiaCroatiaCubaCyprusCzech RepublicDenmarkDominican RepublicEcuadorEgyptFinlandFranceFrench PolynesiaGeorgiaGermanyGreeceHondurasHong KongHungaryIcelandIndiaIndonesiaIranIrelandIsraelItalyJapanJordanKazakhstanKenyaKorea, Republic ofLebanonLithuaniaMacedoniaMalaysiaMaltaMexicoMoroccoMyanmarNepalNetherlandsNew CaledoniaNorwayOmanPakistanPalestine, State ofPERUPhilippinesPolandPortugalQatarRÂżunionRomaniaRUSSIAN FEDERATRussian FederationSaudi ArabiaSingaporeSlovakiaSouth AfricaSpainSri LankaSwedenSwitzerlandTaiwanTanzaniaThailandTunisiaTurkeyUkraineUnited Arab EmiratesUnited KingdomUruguayUzbekistanViet Nam | Open, Classified |  |  | 1536
Z-2673-2016 | 74666 | No | ELC | Ultrasonic scaler. | Dental | Dental | 510(k) | Diamond Coated Tips and Burs are accessories to the air powered Bone drill (VARIOSURG with and without light) |  |  | July 08, 2016 | 2016 | Nakanishi Inc. | Through an error, the labeling failed to include the symbol or other text indicating that the diamond coated tips were single use only. A contributing factor to this error is that in other markets such as Japan, the diamond coated tips are multiple use and are not limited to single use only. The tips are stainless steel with a coating of fine diamond powder. The tips are sold non-sterile and are to be sterilized prior to first use. The tips were cleared as single use only.This problem was discovered during an ongoing, proactive, quality system improvement plan that was initiated by Nakanishi in October 2014. Because the correct instructions for sterilization are not available, it is possible that the devices might not be sterile prior to use. The absence of the single use statement introduces a risk of cross-contamination if the tips are reprocessed and reused. | Error in labeling | Naknishi sent out an URGENT : MEDICAL DEVICE CORRECTION letter on July 18, 2016 describing the potential safety issue regarding their Diamond Coated Products. The firm is requesting customers take the following actions: 1. Examine their inventory immediately to determine if they have any product, both used and unused, bearing any of the Nakanishi Product Codes listed in List of Products enclosed with this letter.2. Immediately dispose of any USED product according to applicable regulations.3. Segregate any UNUSED product for return to Nakanishi.4. In order for the firm to complete their records, they are requesting that customers complete the enclosed Correction Activity Sheet and FAX or Mail to: NSK America, Corp.1800 Global Parkway, Hoffman Estates, IL 60192 US Facsimile: 1-800-838-9328Customers with questions were advised to Telephone: 1-888-675-1675.NSKrecall@NSKDental.comFor questions regarding this recall call 847-843-7664. |  | 2574 | 2574 units | Nationwide Distribution to Illinois, Florida and Georgia. | Open, Classified |  |  | 37577
Z-2677-2016 | 74811 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | CURVED-TIP STAPLER 30,IS4000; Model number 470530; |  |  | July 14, 2016 | 2016 | Intuitive Surgical, Inc. | Field failures were reported related to the da Vinci Xi Surgical System EndoWrist Stapler 45 and 30instruments having an interruption in shaft rotation and / or instrument jaws not opening. | Nonconforming Material/Component | Intuitive sent a Stop Use e-mail to all affected sites on July 14, 2016 to inform users of the issue prior to the commencement of a recall. E-mail described the problem and asked that all surgeons who use the stapler be informed to immediately stop use. Formal letters will follow. Medical Device Recall letters were sent on July 21, 2016. Customers were instructed to return quarantined affected product whether used or unused to Intuitive Surgical. Customers should contact customer service to initiate the standard Return Material Authorization (RMA) process by phone or email:North and South America: 800-876-1310, option 3 (6AM to 5PM PST)Email: customersupport..servicesupport@intusurg.comEurope, Middle East, Asia and AfricaPhone: 800-821-2020 or +41 21 821 2020Email: ics@intusurg.comCustomers were also instructed to complete the attached Return Acknowledgement Form and submit via email to Regulatory Compliance at isi.compliance@intusurg.com or by fax to +1 (408) 523-0619.For questions regarding this recall call 408-523-2100.10-31-16 Update: Monthly status update states that an additional 93 instruments have been identified and an additional communication will be sent to these sites in Mid November 2016, |  | 875 | 875 instruments (827 a la carte + 48 in kits); total, all models | Worldwide Distribution - US (nationwide) Internationally to Australia, Belgium, Denmark, France, Germany. Italy. Japan, Spain, Sweden, Switzerland, Taiwan, and Turkey, | Terminated | January 03, 2017 | 173 | 875
Z-2680-2016 | 75008 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | NIO-A / NIO-P is an instant IO access device comprises housing with spring load mechanism that injects a trocar needle into the bone marrow of a patient and allows connection of syringe or IV administration set. |  |  | August 18, 2016 | 2016 | Waismed Ltd | WaisMed received a report in which during the operation of the NIO-A, the needle was not released from the device as expected. The device was placed on the floor and after several minutes the needle was released spontaneously. NIO-A and the NIO-P modelshave the same activation mechanism. However, none of the affected lots of the NIO-A are distributed in the US, and thus the company is only recalling a single lot of NIO-P in the US. | Component design/selection | The firm emailed the US Consignees on August 18, 2016 that they were conducting a recall on the NIO/A and NIO/P products. The US distributor shall contact the customers and will collect the products from each of the customers. The recalled products shall be moved to Storage quarantine within one week. The company shall keep track on the process and the recall efficacy. |  | 50 | 50 | WI, TX, CA, LA and the following countries: UK, Spain, Poland, Israel, Czech Republic, France, Belgium, Denmark, Italy, Cyprus, Austria, Greece, Slovenia, Australia, Germany, Estonia, New Zealand, Switzerland, Costa Rica, Guatemala, Canada, Malaysia, Poland, Singapore, Lithuania, Romania | Open, Classified |  |  | 50
Z-2687-2016 | 63898 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Non-DEHP Micro-Volume Extension Set, Male Luer Lock Adapter, 36 (91 cm), Volume 0.32 mL; an Rx sterile, single use nonpyrogenic fluid pathway; 50 sets per case; |  |  | December 14, 2012 | 2012 | Baxter Healthcare Corp. | Some of the individual extension set packages were improperly sealed, compromising the sterility of the product. | Unknown/Undetermined by firm | Baxter sent an Urgent Product Recall notification letter dated December 14, 2012, to all affected customers. Baxter instructed customers to locate and remove all affected product from their facility. Customers were instructed to contact Baxter Healthcare Center for Service to arrange for return of product and credit and to complete the attached Customer Reply Form and return it to Baxter or scanned email. Customers with questions should call 1-888-229-0001. For questions regarding this recall call 800-422-9837. |  | 15257 | 15,257 units | Worldwide Distribution - US including FL, GA, IL, KS, KY, MI, MO, NC, ND, OH, PA, SC, TN, TX, WV, Puerto Rico and Internationally to Canada. | Completed |  |  | 15257
Z-2685-2016 | 74924 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Interventional Radiology (IR) images are stored as JPEG2k Lossless in Merge Enterprise Archive (EA) and are not displaying correctly in RadSuite. | Software design | The firm, Merge, sent an "URGENT: MEDICAL DEVICE RECALL" letters dated 1/29/2016 and 5/25/2016 via e-mail or via certified mail if they did not have an e-mail address for the customers. The letters described the product, problem and actions to be taken. The customers were instructed to do the following: you must discontinue using the affected version of RadSuite or iConnect Enterprise Archive with GE AW Workstation; use the workaround listed in the letter; complete and return the enclosed form no later than June 17, 2016; ensure that all users of the product are provided with the notification; if you have further distributed this product, please identify your customers and notify them at once of this product recall Merge has a released fix available for this issue.If you have any additional questions, please send an email to recall@merge.com or call 800-724-5970 ext 8027. |  | 96 | 96 sites potentially have the affected versions for both products | US Distribution to: CO and OK. | Open, Classified |  |  | 96
Z-2681-2016 | 74920 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | PCR Compano Basic Reader using 100-127V as mains voltage. Model numbers: 732027, 732028, and 732040. |  |  | August 12, 2016 | 2016 | Philips Electronics North America Corporation | A defect component of the Compano reader power supply can cause the power supply to burn and some smoke to leave the housing. | Component design/selection | Philips will send a Field Safety Notice to all affected customers on August 12, 2016. The notice informs consignees of the issue and informs them that Philips will replace the power supply unit of the PCR Compano Basic reader system of all affected units free of charge (FCO73200037). |  | 388 | 388 | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC,ND, NH, NJ, NM, NV, NY, OH, OR, PA, TX, VT, WA, WI, WV, and WY; and, the countries of Canada, Brazil, Germany, Iraq, Japan, Korea, Kuwait, Libya, Mexico, Portugal, and Qatar. | Open, Classified |  |  | 388
Z-2684-2016 | 74867 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | INTEGRA Spinal Tray, Catalogue No: 3404875, Contents: 10 Each per Case, Sterile. |  |  | August 04, 2016 | 2016 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences is recalling Spinal Trays, Cat No. 3404875, because of the affected recalled Hospira 5% Lidocaine HCL that is a component of some lots of theSpinal Trays. | Other | Integra sent an Urgent Medical Device letter dated August 4, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to return unused, affected product. Confirmation of the recall notification delivery via delivery verifiable email; Return of Recall Acknowledgement form; Return &amp; disposition of unused, affected product, if the customer in their response indicates they have affected devices. Any products returned to Integra as a result of the recall will be stored in a segregated, labeled area. Recalled products that are returned will be dispositioned as scrap and destroyed. Any Lidocaine returned with the products will be dispositioned as directed in the Hospira Inc. Drug Recall Notification.If you have questions of a customer service nature or beyond the information in the Hospira Inc. URGENT DRUG PRODUCT RECALL NOTICE, please contact either your local Integra Pain Management Sales Representative or customer service at 1-800-241-2210.For further questions, please call (801) 886-9505. |  | 200 | 200 packages (20 cases) | US Distribution to the states of : OH and CA. No foreign, Govt/VA/military consignee. | Open, Classified |  |  | 200
Z-2683-2016 | 74793 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | MAQUET Servo Humidifier 163, Inmed Mfg. Sdn. Bhd. Lot PT 2577, Jalan Perusahaan, 4, 34600 Kamunting, Perak, Malaysia. |  |  | December 03, 2015 | 2015 | Teleflex Medical | Cracking may occur at the 15mm and 22mm ISO femaleconnector. | Device Design | INMED sent an Urgent - Field Safety Notice dated December 3, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to cease use and distribution of stock of the affected product batch and quarantine immediately. Customers are to return all affected product to MAQUET. An Acknowledgement form was included to be completed and returned to their MAQUET representative. If the product was distributed, they were instructed to provide the field safety notice to all of their customers who have received product in scope of the Field Action. Their customer is then required to complete the Acknowledgement form and return it to them.Should you require any further information or support concerning this issue, please contact your local MAQUET representative at fieldactions@maquet.com.INMED is committed to providing high quality, safe and effective products. W e sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service. For further questions, please call (610) 378-0131. |  | 2700 | 2,700 units | US Distribution to the states of : FL, KS, and WV. | Completed |  |  | 2700
Z-2690-2016 | 74525 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | FUJIFILM Digital Mammography System Aspire HD (FDR MS-1000) with Biopsy Positioner (FDR-1000BPY) |  |  | June 09, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | If an exposure is aborted during a stereo biopsy workflow, an error may occur which would prevent the system from continuing to operate properly. | Software design | The firm, Fujifilm, sent a "FIELD SAFETY NOTICE" letter dated June 9, 2016 via Federal Express to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to read and follow the instructions in the letter; call your local service representative, in case an error occurs, and complete and return the Customer Feedback Form via fax (203)-251-7862. Both affected systems will be corrected via on-site visits by FMSU Customer Service Engineers.If you have any questions about this matter, please contact the Product Manager, Women's Health directly at (303)-888-4549 or call your local FUJIFILM office. |  | 2 | 2 Units | US Distribution to: NC and HI. | Open, Classified |  |  | 2
Z-2704-2016 | 74830 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | ADVIA Chemistry XPT, SMN 10723034, IVD. --- This issue affects only the ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment (A1c_3) results when HbA1c values are reported in International Federation of Clinical Chemistry (IFCC) equivalent units (HbA1cR). |  |  | July 20, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens identified an issue with the ratio equation provided on the ADVIA Chemistry XPT System Software Test Definition (TDef) Version 1.0 disks (SMN: 11127343 and 11222123). This issue affects only the ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment (A1c_3) results when HbA1c values are reported in International Federation of Clinical Chemistry (IFCC) equivalent units (HbA1cR). The ratio equation used to calculate ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment (A1c_3) results in IFCC units contains an error. The error results in falsely depressed HbA1cR results. | Software design | Siemens sent an Urgent Field Safety Notice dated July 20, 2016, to all affected customers via Fed Ex and the Urgent Field Safety Notice (CHSW16-02 July 2016) was e-mailed to regional countries for implementation Outside US (OUS). Customers have been advised to stop reporting results and call Siemens immediately to schedule service. The file with the typographical error is located on ADVIA Chemistry XPT Test Definition (TDef) disk. -- Field service personnel were sent a support bulletin describing the issue and instructing them how to correct the equation. Service engineers have been instructed to remove all version 1.0 TDef disks from customer sites regardless of whether they run the ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment assay and report in IFCC units.Customers with questions were advised to contact their Siemens Customer Care Center or their local Siemens technical support representative. |  | 229 | Domestic: 32 units; Foreign: 197 units | Worldwide Distribution - US to MA, NY, MI, IL, TX, CO, AZ, CA, WA and Foreign distribution to the following countries: Australia, Austria, Bahrain, Brazil, China, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, Italy, Malaysia, Netherlands, New Zealand, Norway, Portugal, Portugal, Republic Korea, Singapore, Slovakia, Spain, Sweden, Turkey, United Kingdom, and Vietnam. | Open, Classified |  |  | 229
Z-2691-2016 | 74967 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | HexaPOD evo RT System |  |  | August 18, 2016 | 2016 | Elekta, Inc. | Potentially unrecognized incorrect position of the treatment couch in 3D workflow, i.e. the HexaPOD has not moved fully to the 3D position. | Nonconforming Material/Component | Elekta sent an Important User Notice 618-03-303-025 to affected customers on 8/18/2016. The notice identified the affected product, problem, clinical impact, solution and actions to be taken. The Notice reminds users that the Interlock Check is not possible during 3D Workflow as iGUIDE is not in use. The notice reminds users that if iGUIDE is not being used - (A) Ensure the HexaPOD is at its pre-defined 3D position before treatment, and (B) Ensure the External Inhibit LED at the EnableSwitch Board is off. For questions contact your local Elekta office. |  | 13 | 13 | Worldwide Distribution - US Nationwide in the states of LA, PA, WA, and the countries of: Australia, Denmark, France, India, Italy, and Japan. | Terminated | May 23, 2017 | 278 | 13
Z-2688-2016 | 74064 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Invivo Sentinelle Venguard Breast MRI Auxiliary Table/Tabletop with 8/16 Channel Coil Array for GE 1.5T/3T MRI Systems. |  |  | May 06, 2016 | 2016 | Invivo Corporation | A diode was inadvertently not assembled into the system, which may result in coil overheating | Under Investigation by firm |  |  | 3 | 3 | WI | Completed |  |  | 3
Z-2689-2016 | 74758 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | A7 Anesthesia Delivery System, a device used to administer to a patient, continuously or intermittently, a general inhalation anesthetic and to maintain a patient's ventilation. Part numbers 0632F-PA0000X (US domestic) and 0632B-00014 (international) |  |  | June 15, 2016 | 2016 | Mindray DS USA, Inc. dba Mindray North America | Potential for a leak to occur on the back-up O2 and air e-size cylinder yokes on the A7 Anesthesia Delivery System. | Nonconforming Material/Component | Mindray sent via certified mail with return receipt a recall letter dated June 15, 2016 to their affected customers. Mindray will replace the gasket on all e-size cylinder yokes on the affected A7 systems. Customer was instructed to contact their local Mindray Service Representative to arrange for this replacement. Customers can continue to use their A7 system while awaiting the replacement of the gasket(s). Customers can contact Ms. Diane Arpino, Director, Quality Operations and Regulatory Affairs via email to: d.arpino@mindray.com or via telephone to: (201)995-8407 |  | 167 | 167 units (165 units - US) and (2 units - International) | US Nationwide Distribution to AL, MA, MD, MN, MO, NE, NJ, OK, OR, PA and VA; and Canada | Open, Classified |  |  | 167
Z-2705-2016 | 74190 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Double Pump RF Patient Cassette; Product Number: 72204055. |  |  | May 10, 2016 | 2016 | Medical Vision Ab | Sterility of device may be compromised due to sterile package breakage. | Unknown/Undetermined by firm | The An Urgent Field Safety Notice: Medical Device Field Safety Corrective Action / Recall letter was sent via certified mail on May 10, 2016. The letter informs customers of the issue and asks them to take the following actions: 1. Locate and quarantine affected unused devices immediately; 2. Return quarantined product to your national Smith &amp; Nephew agency/distributor. 3. Complete the return slip and fax it to your national Smith &amp; Nephew agency/distributor. 4. Please make sure this safety information is passed on to all those who need to be aware of it within your organization. 5. Please maintain awareness on this notice and resulting action until the Field Safety Corrective Action is terminated to ensure effectiveness of this action. |  | 3295 | 3,295 devices | Distributed to the US, to the states of NH, OK, AL, PA, IL, MA, TX, MI, NC, CA, MN, CT, GA, NY, WI, KY, OH, VA, AL, NJ, TN, SD, and IN. | Open, Classified |  |  | 3295
Z-2692-2016 | 74875 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System, Catalog No. 775222 |  |  | August 04, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter is recalling the UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System to reduce the potential for flammable liquids coming into contact with electrical components in the Stainer module possibly causing a fire. | Under Investigation by firm | Beckman Coulter sent an Urgent Medical Device recall letter dated August 4, 2016, to all affected customers to inform them that the UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System to reduce the potential for flammable liquids coming into contact with electrical components in the Stainer module possibly causing a fire. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with any questions regarding the recall notice were instructed to contact Beckman Coulter Customer Support Center:http://www.beckmancoulter.com/customersupport/support; 800-526-7694 in the United States and Canada; and outside the United States and Canada, are instructed to contact their local Beckman Coulter Representative. |  | 759 | 759 units total (358 units in US) | Worldwide Distribution to US (nationwide) and Internationally to Australia, Belgium, Canada, China, Croatia, Czech Republic, France, Germany, Hong Kong, India, Israel, Italy, Japan, Kuwait, Malaysia, Mayotte, Mexico, New Zealand, Philippines, Poland, Portugal, Qatar, Saudi Arabia, Slovakia, South Korea, Spain, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Republic, United Kingdom, and Vietnam. | Open, Classified |  |  | 759
Z-2707-2016 | 74827 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SMR allen wrench 5 mm, product code 9013.50.210 |  |  | July 21, 2016 | 2016 | Limacorporate S.p.A | Product found to be dimensionally non-conforming (slightly over dimensioned on the diameter of the Zimmer connection) and non-functional for connecting with some T-handles. | Nonconforming Material/Component | The firm received two complaints from the US field about the inta-operative malfunction with the SMR allen wrench 5mm (product code 9013.50.210) with the T-handle. Both parts involved were manufactured with the same lot #14AA311. No adverse effects for patients reported. |  | 76 | 76 | US, Italy, Australia | Terminated | January 13, 2017 | 176 | 76
Z-2706-2016 | 74856 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Terumo Custom Cardiovascular System Cardiovascular Procedure Kits labeled containing Coated Tubing as follows: |  |  | July 25, 2016 | 2016 | Terumo Cardiovascular Systems Corp | Terumo Custom Cardiovascular Procedure Kits are labeled as Xcoating surface coated tubing, but tubing is non-coated | Process control | Terumo Cardiovascular Systems issued the Safety Alert dated July 25, 2016 and advises the user that their kit may include uncoated tubing, and further advises the consignee to continue using their kits with consideration of anticoagulation protocols during a procedure. Each consignee is asked to complete a Customer Response Form, and to return the Customer Response Form to Terumo acknowledging their receipt of the form. Consignees can return their affected product for credit/replacement; however, return of the product from the field is a customer option and is not under recommendation from Terumo Cardiovascular Systems. For further questions, please call (800) 262-3304. |  | 2207 | 2207 kits | NationwideForeign: Canada, Latin America | Open, Classified |  |  | 2207
Z-2682-2016 | 74685 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DuraDiagnost stationary X-ray system |  |  | September 17, 2015 | 2015 | Philips Electronics North America Corporation | Buttons or keys on the monitor allow the user to turn the volume setting down to 0", in which case the beep will not be heard at the conclusion of an exposure, as required 21 CFR 1020.31(j), and the Height-Adjustable Tabletop is missing the required labels [21 CFR 1010.2., 21 CFR 1010.3(a), and 21 CFR 1020.30(3)]. | Radiation Control for Health and Safety Act | Philips plans to bring the products into compliance with the Federal standard at no cost to customers by implementing a field correction.1. You will contact customers and initiate the actions to correct the failure to comply issues.2. You will attach the warning label to maintain an audible volume, and issue a supplement to the user manual that instructs the user how to adjust the volume for the local noise environment and include a warning that the user shall maintain an audible volume.3. You will attach the identification label to the affected Height-AdjustableTabletop.4. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.If you need any further information or support concerning this, please contact your local Philips representative or Philips Customer Service: 1-800-722-9377. |  | 15 | 15 | US Distribution to the states of : ME, NE, OR, CA, TX, NC, WA and OH. | Open, Classified |  |  | 15
Z-2711-2016 | 74900 | No | NQE | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | SPECULAR MICROSCOPE CEM-530; Software version 1.08 and 1.09. |  |  | August 08, 2016 | 2016 | Nidek Inc | Software version 1.08 and 1.09 for the Specular Microscope CEM 530 included a change of analysis results feature that was not reviewed and approved by the FDA. | Software change control | The firm, Nidek, sent a Dear Doctor letters entitled "RECALL NOTIFICATION" dated August 15, 2016 to affected sites by Federal Express. The letter informs users of the issue and stated that Nidek field service representatives will visit to install updated software.If you have any questions, please call Customer Service at 1-800-722-0219; Monday-Friday 8am to 5pm PST or email: cem530update@noritsuservice.com. |  | 29 | 29 affected devices | US Distribution to states of: TX, KS, FL, NV, CA, PA, KY, OR, TN, AL, VA, MO, OK, CT, NY. LA, NM, and IN. | Terminated | January 09, 2017 | 154 | 29
Z-2712-2016 | 74961 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | August 18, 2016 | 2016 | Elekta, Inc. | In a specific workflow where contours are edited (enlarged, moved, copied, etc.) and then the contour is deleted on some slices and saved in the same session, it is possible that the deleted contours are still present which would not be the intent of the planner. | Unknown/Undetermined by firm | Important User Notice 382-03-MON-003 was sent to customers on August 18, 2016. The notice informs users of the specific product and version numbers affected by the issue, and any work arounds that can be used to avoid the issue. The letter also states that the problem will be in a future Monaco release. The customers were instructed to complete and return the Important Field Safety Notice Acknowledgement form to their local Elekta as soon as possible but no longer than 30 days from receipt. Questions can be directed to the local Elekta office. |  | 3012 | 3,012 units | Worldwide Distribution -- USA, to the states of AK, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MD, MA, MI, MN, MO, MT, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, Puerto Rico, Algeria, Antigua and Barbuda, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Botswana, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Laos, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Nambia, Nepal, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, Vietnam, and Zimbabwe. | Open, Classified |  |  | 3012
Z-2709-2016 | 75027 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. The firm name on the label is Merge Healthcare, Hartland, WI. Image processing system. |  |  | February 09, 2016 | 2016 | Merge Healthcare, Inc. | A report can be confirmed with the incorrect patient demographics resulting in it being saved under the incorrect patient record in the Electronic Medical Record (EMR). | Software design | MERGE sent an Urgent Medical Device Recall letter dated February 9, 2016, to all affected customers via email or via certified mail if they did not have an e-mail address for the customer.The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it.For further questions, please call (800) 724-5970 ext. 8027. |  | 116 | 116 customers potentially have the affected versions | Nationwide Distribution | Open, Classified |  |  | 116
Z-2710-2016 | 74909 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Push Button Blood Collection Set, Sterile, Rx only, 0.6 x 19 mm x 305 mm, 23G x 3/4" x 12", REF 367342. |  |  | August 08, 2016 | 2016 | Becton Dickinson &amp; Company | Becton Dickinson is recalling various lot numbers of BD Vacutainer Push Button Blood Collection Set because of reports that specific lots exhibit separation of fron and rear barrels upon activation of the safety feature, which retracts the needle. | Under Investigation by firm | BD notified their customers via e-mail and sent an "Urgent Product Recall Letter" and a Recall Response Form dated August 8, 2016 via UPS. |  | 1788150 | 1,788,150 units | US and International | Open, Classified |  |  | 1788150
Z-2708-2016 | 74824 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Alcon AcrySof IQ IOL with ULTRASERT Delivery System |  |  | July 28, 2016 | 2016 | Alcon Research, Ltd. | The ULTRASERT Delivery System from certain lots have an inferior surface characteristic that could result in the IOL becoming lodged in the Delivery System | Process control | Alcon Research sent an Urgent Field Safety Notice letter dated July 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check their inventory and segregate any product being recalled. Customers were asked to complete and return the Field Safety Notice Response Form. For questions or concerns contact Alcon. |  | 371 | 371 units | Armenia, Belgium, Bulgaria, Croatia, Finland, France, Germany, Hungary, Italy, Lithuania, Malaysia, Netherlands, Norway, Poland, Portugal, Qatar, Romania, Slovenia, Spain, Sweden, Switzerland, Thailand, United Kingdom (Great Britain and Ireland), and United Arab Emirates. *** While this product is available in the US, these lots were not distributed to US consignees *** | Terminated | June 07, 2017 | 314 | 371
Z-2713-2016 | 74706 | No | CBT | Arterial blood sampling kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | The 9025TRU Arterial Blood Gas (ABG) Kit is used to obtain arterial blood sampling for pH, Blood Gas, Electrolyte and Metabolite Analysis. |  |  | June 27, 2016 | 2016 | Cardinal Health | Arterial Blood Gas (ABG) Kit- Missing Safety Component. The ABG Kit(s) have been identified as missing the locking clasp component. CareFusion/BD has identified a potential customer/user safety risk associated with the missing locking clasp component. | Under Investigation by firm | CareFusion/BD sent out a FIELD SAFETY NOTICE of the field correction in writing on June 27, 2016. The letter described the nature of the recall and provided actions to be performed. Carefusion/BD directed customers to take the following actions:Step #1: Inspect current inventory on-hand to identify the presence of the affected Arterial Blood Gas Sampling Kit(s). Step #2: Complete the Customer Response Form by checking all appropriate boxes and providing all accurate information as required. Step #3: Destroy all affected product(s) in-stock in accordance with their facility's destruction protocol. Step #4: Return the Customer Response Form to GMB-GLB-ALFieldActions@CareFusion.com Step #5: If the customers want to obtain credit and/or replacement of destroyed product(s) they are asked to call their CareFusion/BD Customer Support Department at 800.323.9088 (Option #1 ). For any additional questions and support concerning the voluntary Field Safety Corrective Action, customers are asked to call (847)362.9303. |  | 2709466 | Lot #884993 -16,000 units, Lot #903136 -1,900 units and Lot #903137 - 300 units. | Worldwide Distribution - US (nationwide) and Canada. | Open, Classified |  |  | 2709466
Z-2721-2016 | 75052 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Medical Linear Accelerator |  |  | July 08, 2016 | 2016 | Elekta, Inc. | There have been instances in hospitals where the joint between the drive gear and the gantry base drive wheel has become loose, leading to fatigue failure of some of the securing bolts. | Other | Important Field Safety Notice 200-01-103-081 was released July 8, 2016 for distribution to all customers with Digital Accelerators without XVI. It informs users of the specific product and version numbers affected by the issue. The modification involves an inspection and correction procedure. Should the inspection pass, the corrected in not required. All gantry base wheel bolts should be checked for presence and torque and should be replaced if found to be showing signs of fatigue. Elekta service is actively reaching out to affected customers in order to organize a time to implement the inspection and correction. The notice included an acknowledgment form to be returned. |  | 553 | 553 | AL, AZ, CA, FL, GA, KY, MI, MN, MO, NE, NJ, NY, NC, OH, OK, PA, SC, TX, WA, WV, Argentina, Austria, Australia, Bosnia and Herzegovinia, Bangladesh, Belgium, Bulgaria, Brazil, China, Columbia, Cuba, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Iran, Italy, Japan, Jordan, Lebanon, Libya, Morocco, Mexico, Nigeria, Netherlands, Norway, Pakistan, Panama, Peru, Philippines, Portugal, Serbia, Slovenia, South Africa, South Korea, Spain, Sudan, Sweden, Syria, Taiwan, Thailand, Turkey, United Kingdom, Venezuela, Vietnam, Yemen | Open, Classified |  |  | 553
Z-2731-2016 | 75081 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare SIGNA Creator, SIGNA Explorer. Diagnostic Imaging Device. |  |  | August 31, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare is conducting a recall dealing with an incorrect RF Power Monitor Function involving SIGNA CREATOR and SIGNA EXPLORER which could result in higher than expected thermal dose to the patient and higher than expected localized heating. There have been no failures of this type reported, and no injuries reported as a result of this issue. | Software design | Consignees were sent on 8/31/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 67921 dated August 31, 2016. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, and Radiologists.The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 383 | 383 (US = 24;OUS = 359) | US: AZ, CA, CT, FL, IL, MA, MS, MO, NC, PR, TX, WI.OUS: Algeria, Australia, Bangladesh, Belgium, Brazil, Canada, Colombia, Dominican republic, Egypt, France, Germany, Ghana, Guatemala, HONG KONG, Hungary, India, Indonesia, Iran, Italy, Jamaica, Japan, Korea, Kuwait, LEBANON, Lithuania, Maldives, Martinique, Mexico, MOROCCO, Netherlands, New Zealand, OMAN, Pakistan, Philippines, Poland, Portugal, Romania, SAUDI ARABIA, Singapore, South Africa, Spain,Sweden, Tanzania, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam. | Open, Classified |  |  | 383
Z-2719-2016 | 75033 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | ReliaMedÂż Silver Alginate/CMC. Each dressing is packed in a pouch (primary packaging), each pouch is packed in a carton (secondary packaging) holding multiple (depending on stock keeping unit) pouches and each carton contains one (1) Instructions For Use (IFU) leaflet. |  |  | January 15, 2015 | 2015 | Advanced Medical Solutions Ltd. | On 08 January 2015 AMS Sole Customer for the products notified AMS of the issue after discovering the misplaced IFUs during its review of packaging in its warehouse | Package design/selection | Advanced Medical Solutions' (AMS) Sole Customer has already received notification via telephone conference call, followed by emails and correspondence. |  | 47569 | 47,569 cartons | OH | Terminated | November 23, 2016 | 678 | 76773
Z-2732-2016 | 74912 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 and DxH 600 Coulter Cellular Analysis System. |  |  | August 08, 2016 | 2016 | Beckman Coulter Inc. | There is a possibility of specimen misidentification for the stand-alone DxH 800 and DxH 600 where the Primary Identifier is configured to Tube Position ID. | Software design | An Urgent Medical Device Recall letter dated 8/8/16 was sent to all customers who purchased the Unicel DxH 800 and DxH 600 Coulter Cellular Analysis System to inform them that there is a possibility of specimen misidentification for the stand-alone DxH 800 and DxH 600 where the Primary Identifier is configured to Tube Position ID. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with any questions regarding the notice are instructed to contact Customer Support Center: http://www.beckmancoulter.com/customersupport/support; 800-526-7694 in the US and Canada; outside the US and Canada, contact local Beckman Coulter representative. |  | 5282 | 5,282 units total (2,877 units in US) | Affected geographies: Andorra, Australia, Belarus, Belgium, Bosnia and Herzegovina, Brunei Darussalam, Canada, China, Columbia, Croatia, Czech Republic, Egypt, Ecuador, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Lebanon, Macao, Malaysia, Mayotte, Mexico, Morocco, Myanmar, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Republic, United Kingdom, Uruguay and Vietnam. | Terminated | April 29, 2017 | 264 | 5282
Z-2715-2016 | 74853 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge RadSuite software. Radiological image processing system. |  |  | April 09, 2015 | 2015 | Merge Healthcare, Inc. | When RadSuite is used with IPID (Issuer of Patient ID) as a part of the "Patient Identifier," it is possible in some circumstances that the demographics of one patient will be applied to a study or studies for another patient. | Software design | The recalling firm issued letters dated 4/3/2015 via email on 4/9/2015 notifying the customer of the issues. |  | 10 | 10 sites have the affected software that is configured with the specific conditions listed in Code Information | Distribution was made to medical facilities in AL, MI, MO, PA, TN, and TX. | Open, Classified |  |  | 10
Z-2717-2016 | 74975 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | NOMAD Pro2 Packaged X-Ray System, Part Number 08500021; NOMAD Pro2 Packaged X-Ray System Black, Part Number 08500022; NOMAD Pro2 Vet Packaged X-Ray System, Part Number 08500023; NOMAD Pro2 Vet Packaged X-Ray System Black, Part Number 08500024; NOMAD Pro2 Packaged X-Ray System, GBL, Part Number 08500025 and Part Number 08500026; NOMAD Pro2 Packaged X-Ray System, GLB, Vet, Part Number 08500027 and Part Number 08500028; NOMAD Pro2 Package X-Ray System, GLB, Black, Part Number 08500029 |  |  | July 18, 2016 | 2016 | Aribex Inc | Firm received customer complaints where end users reported that Nomad Pro and Pro 2 handsets had appeared to "overheat", "melt", "smoke" and/or "ignite." Investigation found that the internal lithium ion batteries had entered into a thermal runaway condition as a result of a short circuit on the battery pack PCB. | Other | The firm sent a "FIELD ACTION NOTICE" to all affected customers. |  | 16871 | 16,871 | AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA. HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY , Puerto Rico, Afghanistan, Argentina, Australia, Bahamas, Bahrain, Belgium, Bermuda, Chile, China, Columbia, Denmark, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Kuwait, Malaysia, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Saudi Arabia, Papua New Guinea, Philippines, Poland, Qatar, Saudi Arabia, Slovenia, South Africa, Sweden, Switzerland, Thailand, Trinidad and Tobago, Turkey, UAE, United Kingdom, Venezuela, Vietnam | Terminated | July 26, 2017 | 373 | 16871
Z-2714-2016 | 74862 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Acecide-C High Level Disinfectant &amp; Sterilant 875 mL. Packaging is 6 inner cartons in one master carton;1 inner carton consists of a set of two liquids (1 acid &amp; 1 buffer) banded together. |  |  | August 02, 2016 | 2016 | Best Sanitizers Inc | Certain lots of the acid solution in Acecide-C are degrading faster than normal. | Nonconforming Material/Component | Customer letters dated 8/2/16 will be sent. |  | 3023 | 3,023 cases USA; 666 cases shipped to Canada (OTC drug there) | Distributed to one (1) account in the state of PA.Distributed to one (1) account in Canada | Terminated | August 18, 2017 | 381 | 3023
Z-2733-2016 | 75028 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra Cranial Access Kits, Catalogue Nos. INSHITH; INSHITHRZN; INSHITHND |  |  | August 24, 2016 | 2016 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences has been notified of a medical device recall by Adaptive Surgical, LLC involving their Riverlon" Nylon 3-0 Sutures that are packaged in Integras Cranial Access Kits. It was found that product seals may become open during extreme distribution conditions and render the device unsterile. | Packaging | Consignees notified on August 12, 2016 via letter to identify and return product. |  | 710 | 710 | Nationwide and VA/govt/military. No foreign consignees. | Terminated | March 30, 2017 | 218 | 710
Z-2734-2016 | 74886 | [6 empty] | 2 over 2 Elevating Monitor Suspension - Omega Medical Imaging, LLC Elevating Monitor Suspension Systems. |  |  | June 21, 2016 | 2016 | Omega Medical Imaging, Inc. | Elevating Monitor Suspension separated from the lifting column bracket | Device Design | Consignees were notified by a Field Safety Advisory Notice - FSAN 16-0001 of this immediate Recall on 6/21/2016. |  | 26 | 26 units | CO, MN, GA, NC, IN, OH, KY, IL, NY, AL, MO, LA, PA, and SD Denmark | Open, Classified |  |  | 26
Z-2736-2016 | 74894 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT Model 882482, gamma camera for Single Photon Emission Computed Tomography (SPECT) |  |  | May 25, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The firm became aware of an issue where during a non-Auto Body Contouring (ABC), non-circular, rel-180 scan, the detector may come in contact with the patient. The issue was caused by a software defect in the workflow setup script. | Software design | The firm is informing the consignee of this issue through Customer Information Letter. |  | 424 | 424 units | Worldwide distribution. US nationwide, Algeria, Argentina, Armenia, Australia, Austria, Belgium, Brazil, Czech Republic, Denmark, Finland France, Gabon, Germany, Greece, India, Indonesia, Ireland, Japan, Korea, Republic of Korea, Republic of Lebanon, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Oman, Panama, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Viet Nam. | Open, Classified |  |  | 544
Z-2739-2016 | 64761 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Centricity Enterprise Archive is a software product for receiving, archiving and sending of medical data. Qualified system administrators install, monitor and maintain the system. DICOM devices (e.g. modalities, workstations) communicate with the archive using the DICOM protocol (published by ACR-NEMA). The clinical use is limited to the interaction other systems have with the Centricity Enterprise Archive. Centricity Enterprise Archive is not intended for use in patient vicinity. The intended users for Centricity Enterprise Users service personnel and System administrators, and non clinical users. As of V4.0 XDS enabled systems can communicate with the archive using the XDS and XDS-I profiles (published by IHE). |  |  | November 16, 2012 | 2012 | Ge Healthcare It | There is a potential for data loss associated with empty path names resulting from HL-7 updates when the cache disk is almost full. In an unlikely event that the cache disk is almost full (&lt;100MB) and HL-7 update events are received, the update action fails and an empty path name is written to the EA database. As a result the instance is not deleted from the storage device but no longer referenced. When trying to retrieve the instance an error is returned. | Software design | URGENT MEDICAL DEVICE CORRECTION letters dated November 16, 2012 and April 19, 2013, were sent to all affected customers.The letters requested that customers verify that all Enterprise Archive partitions have enough free space available and that the cache disk has at least 10% free space available. The letters further requested that customers whose disks are not being purged to the configured watermarks and the cache disk is almost full to contact their local GE service person. A permanent product correction (Service Pack 12 and release notes) will be provided by GE Healthcare IT. Customers with questions can contact their GE service representative or the GE Healthcare IT Remote Online Center Customer Care Center at 1-800-437-1171. |  | 1379 | 1,379 units | Nationwide Distribution | Terminated | January 11, 2017 | 1517 | 1379
Z-2723-2016 | 64261 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Infusion Pump. List Number 11971. The Plum A+ Infusion Systems are a multi-purpose family of devices that provides precise delivery of multiple therapies across the general spectrum of clinical care. The Plum A+ family infusion pumps are self-contained, microprocessor based infusion devices. |  |  | February 06, 2013 | 2013 | Hospira Inc. | The fluid shield diaphragm for the Plum A+ Infusion pumps was undersized, possibly resulting in alarms and delay or interruption of therapy. | Nonconforming Material/Component | Hospira sent an Urgent Device Field Correction letter, dated February 6, 2013, to customers. Customers were requested to return a reply form to Stericycle indicating that they have received the letter, and notified their users and/or customers. The customers were asked to supply the number of Plum A+ infusers at their facility. Product will be corrected by Hospira in the field. |  | 357778 | 357,778 total | United States nationwide distribution including PR, and USVI. Costa Rica, Turks and Caicos, and Dominican Republic. | Terminated | November 14, 2016 | 1377 | 357778
Z-2738-2016 | 73833 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | PINNACLEÂż Introducer Sheath |  |  | March 16, 2016 | 2016 | Terumo Medical Corporation | Terumo Medical is initiating this recall due to complaints of valve leakage that have been related to a manufacturing issue. If affected introducer sheaths are used, there is possible risk of incremental blood loss through the valve. | Device Design | Terumo sent an Urgent Product Recall Bulletin dated March 7, 2016, to all affected customers. All affected customers will be notified regarding the required recall actions (stop using/quarantine the affected products). TMC will strive for 100% verification to confirm that all customers have acknowledged the recall. Follow-up communication is being sent specifying the mechanisms to be used for labeling and returning the involved product from the customer level.We regret any inconvenience this action may cause your operations and we appreciate your understanding. Please contact us if you need assistance procuring additional inventory. We are committed to continuing to offer products that meet the highest quality standards that you expect from Terumo Medical Corporation. If you have questions or need further information, please call Customer Care at 1-800-888-3786 or contact your local sales representative.For further questions regarding this recall please call (410) 392-7226. |  | 1370995 | 1,370,995 eaches have been distributed to the field | Nationwide Distribution | Open, Classified |  |  | 1370995
Z-2743-2016 | 74613 | No | NOV | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Achillon Minimally Invasive Device for Achilles Tendon Rupture Treatment |  |  | June 17, 2016 | 2016 | Integra LifeSciences Corp. | Data from internal studies suggests there may be a potential that the secondary package seal integrity does not provide the level of sterility intended for Minimally Invasive Achilles Tendon Suture System Achillon. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall letter and recall acknowledgment and return form dated June 17, 2016 to their customers via traceable courier. Customers were asked to review their inventory and complete the attached form. Customers with questions should contact Customer Service at 1-800-654-2873.For questions regarding this recall call 717-840-3431. |  | 6729 | 6,729 units | Nationwide distribution | Open, Classified |  |  | 6729
Z-2744-2016 | 74710 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microtek Franz Angiography Drape with Radiation Shield, Item # RBA202, and Microtek Angiography Drape with Radiation Shield and Pouches, Item # RBA203. Sterile EO. Microtek Medical, Inc., Columbus, MS 39702. |  |  | July 14, 2016 | 2016 | Ecolab Inc | Ecolab is recalling 2 models of Angiography Drapes with Radiation Shields because the sterile packaging may contain small holes in the film of the pouch. | Packaging process control | Consignees were sent on 7/14/2016 an Ecolab "Urgent Medical Device Recall" letter dated July 14, 2016. The letter described the problem and the product involved in the recall. Advised consignees to immediately quarantine the product and to notify their customers. For questions contact Customer Service at 1-800-824-3027. |  | 6750 | 6750 eaches (675 cases) | US: OH, NY, CA, MN. | Open, Classified |  |  | 6750
Z-2747-2016 | 74895 | No | LCJ |  |  | General &amp; Plastic Surgery | 510(k) | MENTOR 2" Tall Height TE With Sutures 350cc PN: 354-9312 |  |  | August 09, 2016 | 2016 | Mentor Texas, LP. | Tissue expander product mix found: CPX4 350cc was placed in a carton for the CPX4 450cc | Labeling mix-ups | The firm, Mentor, sent an "MEDICAL DEVICE REMOVAL" letter dated 8/11/2016 to affected consignees. The letter will be followed up with a respective sales representative visit. The letter described the product, problem and actions to be taken. The consignees were instructed to immediately identify all unused units and return them per the normal return process; forward the letter to anyone your facility that needs to be informed; if you have sent these units to another facility, forward the letter to that facility; and maintain awareness of the letter until all units in your inventory have been returned.If you have any questions, please contact your local Mentor sales representative or Mentor Customer Service at 1-800-235-5731. |  | 65 | 65 units | Worldwide Distribution: US (nationwide) and countries of: South Korea, Argentina, Mexico, Canada, Australia and New Zealand. | Open, Classified |  |  | 135
Z-2741-2016 | 74986 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Radial Artery Catheterization Kit with Sharps Safety Features |  |  | August 08, 2016 | 2016 | Arrow International Inc | Shipping carton labeling error. | Process control | Teleflex sent an Urgent Medical Device Recall Notification letter dated August 23, 2016, via FED-Ex 2 Day Air. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine all affected products. Instructions for the products return were provided in the firm's recall notification letter. Customers were asked to complete the enclosed Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email it to recalls@teleflex.com. Customers with questions were instructed to call 1-866-246-6990. |  | 690 | 690 kits | Nationwide Distribution to OH, IN, TN, NC, and KY. | Open, Classified |  |  | 690
Z-2746-2016 | 74915 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The Connecting Bracket (Fixture) Device (1002.65A0) is designed as a connecting element to support ancillary head rests on a Maquet operating room table immediately before, during and after surgical interventions as well as for examination and treatment. |  |  | August 12, 2016 | 2016 | Maquet Cardiovascular Us Sales, Llc | Post Market Surveillance cases reported to Maquet showing instances in which the welded joint between the square mount and the first fixable swivel joint of Maquet's Connecting Bracket (Fixture) Device is broken. | Use error | Maquet sent an "Urgent Product Recall Medical Device Correction Letter and Medical Device Correction Response Form" dated August 12, 2016 via UPS to their affected customers. Customers were advised to inspect all devices that include the fixture for damage before every use. Customers were asked to fill out and sign the enclosed Medical Device Correction Response Form and return it via email to FieldAction@maquet.com or fax to 1-973-909-9927. Customers with questions were advised to contact their MAQUET Representative or call MAQUET Customer Service at 1-888-627-8383.For questions regarding this recall call 973-709-7487. |  | 19 | 19 units (US) and 1975 units (OUS) | Nationwide Distribution to CA, FL, IA, IN, MO, UT and WI | Open, Classified |  |  | 19
Z-2745-2016 | 74771 | No | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew Whipknot Soft Tissue Cinch (USP#5) suture. PN: 7211015 Used during ligament reconstruction. |  |  | July 19, 2016 | 2016 | Smith &amp; Nephew, Inc. | Sterility compromised due to breach in sterile packaging | Package design/selection | Smith &amp; Nephew issued recall letters via Federal Express on July 19, 2016 to the customers. Customers asked to inspect inventory and locate any unused devices, and quarantine immediately. Contact Smith &amp; Nephews Field Actions Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to return the completed Inventory Return Certification Form and obtain a return authorization (RA) number if you have product to return. |  | 34065 | 34,065 units | Worldwide distribution. US nationwide, and the countries of: AU, AT, BE, CA, CH, CR, DB, DE, ES, FI, FR, GB, GR, IC, IL, IN, IT, MY, NL, NO, PL, PR, PT,SE, SG, TH, and ZA. | Open, Classified |  |  | 34065
Z-2749-2016 | 74812 | Yes | LWP |  |  | Cardiovascular | PMA | Pacemakers: Accolade MRI DR Model L311; Accolade MRI DR EL Model L331; and Essentio MRI DR Model L111. |  |  | June 15, 2016 | 2016 | Boston Scientific Corporation | Five Accolate MRI and Essentio MRI DR pulse generators with an alternate MICS module configuration were shipped to US customers and implanted. This alternate configuration did not receive FDA approval yet. | Process control | As we considered the value of a direct physician communication we have not identified any information that would be of value to the physician because device performance and behavior remain unchanged. |  | 5 | 5 | Distributed to the states of TX, NH, MA, PA, and MN. | Open, Classified |  |  | 5
Z-2742-2016 | 74917 | No | LTJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA CentaurÂż PSA Assay; Cat No. 06574155/SMN 10310292 (100 test)/02676506/SMN 10310293 (500 test) |  |  | August 09, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | Communication provided to emphasize that the PSA values should be interpreted in accordance with current clinical guidelines for defining biochemical recurrence following radical prostatectomy. | Under Investigation by firm | An Urgent Medical Device Correction (UMDC) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on July 29, 2016 to be delivered to customers on August 1, 2016. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on July 28, 2016. These emphasize that the prostate specific antigen (PSA) values should be interpreted in accordance with current clinical guidelines for defining biochemical recurrence following radical prostatectomy (e.g., the 2013 American Urological Association (AUA) Guidelines or the 2015 European Association of Urology (EAU)). These guidelines define biochemical recurrence of prostate cancer as a detectable or rising PSA value post-radical prostatectomy that is &gt; or = 0.2 ng/mL (ug/L) with a second confirmatory level of &gt; or = 0.2 ng/mL (ug/L). The communications also provide functional sensitivity from a recent study to give an expectation of precision at the low end of the assay. |  |  |  | Nationwide AlbaniaAlgeriaAngolaArmeniaAustriaBahrainBelgiumBulgariaBurkina FasoCroatiaCyprusCzech RepublicDenmarkEgyptEstoniaFinlandFranceGeorgiaGermanyGreat BritainGreeceHungaryIcelandIraqItalyJordanKazakhstanLatviaLithuaniaLuxembourgMacedoniaMaltaNetherlandsNorwayPakistanPolandPortugalQatarRomaniaRussian Fed.SenegalSerbiaSlovakiaSloveniaSpainSwedenSwitzerlandTunisiaTurkeyUnit.Arab Emir.UzbekistanVatican | Open, Classified
Z-2760-2016 | 73196 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | The system provides the ability to transmit patient data files for storage, viewing and analysis at distributed locations via the intranet or Internet, or may function as a stand-alone device. |  |  | October 19, 2015 | 2015 | Invivo Corporation | Display of SpO2 and Non-Invasive Blood Pressure (NIBP) Measurements May Freeze | Unknown/Undetermined by firm | Phillips Medical Systems sent out a letter to their customers, dated October 19, 2015. The letter notified customers and distributors of the voluntary recall of the Xper Flex Cardio Patient Monitoring System, and identified the product issue as "Display o fSpO2 and NIBP measurements may freeze". Phillips assured customers that the recalled device can be used if each monitored patient is closely observed and not left unattended. These directions are directed within the instructions of the manual for regular use. Continued use is advised until Phillips contacts the individual customer to correct the issue with a software update. Some hardware is more dated, so software updates will require special handling. If customers have any additional questions, they are encouraged to contact their local Phillips representative. For further questions, please call (877) 468- 4861 option 1 then option 2. |  | 2929 | 2,929 | Nationwide Distribution to the states of :AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TX, TN, UT, VA, WA, WI, WV, Alaska, Hawaii, &amp; Washington, DC. Worldwide Distribution to the countries of : Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bulgaria, Canada, Cayman Islands, Chile, Columbia, Croatia, Democratic Republic of Congo, Denmark, Ecuador, Egypt, Estonia, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Islamic Republic of Iran, Ireland, Israel, Italy, Jordan, Kazakhstan, Republic of Korea, Lebanon, Libya, Malaysia, Mauritius, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, State of Palestine, Panama, Papua New Guinea, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Taiwan Province, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, United Republic of Tanzania, Uzbekistan, &amp; Vietnam. | Open, Classified |  |  | 2929
Z-2758-2016 | 74395 | No | IYX | Scintillation (gamma) camera. | Radiology | Radiology | 510(k) | e.cam and Symbia E gamma camera systems with patient beds |  |  | August 05, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | "We have received reports of patients hair being caught in the e.cam and Symbia E patient bed up/down drive. The purpose of this letter is to remind users of the patient positioning instructions for your e.cam or Symbia E system and to inform you of an update that will be performed by your local Siemens service representative." | Device Design | The firm, Siemens, sent a "Customer Advisory Notice" letter on 8/5/2016 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the "safety" sections of the user manual with your staff; continue using your system under the guidelines set forth in the letter; place advisory in the system's instructions for use; and if you have sold this equipment and it is no longer in your possession, forward the notice to the new owner and inform us about the new owner of the equipment.Siemens has designed an enhancement kit to further reduce the gap in the patient bed up/down drive. Your local Siemens service organization will be contacting you to schedule this update.If you have any questions regarding this important notice, contact your local Service representative at the following numbers: America: 1-800-888-7436; Europe, Middle East, and Africa: +49 9131 940 4000 or Asia and Australia: +86 (21) 3811 2121. |  | 5451 | 5451 | Worldwide Distribution: US (Nationwide) including states of: District of Columbia and Puerto Rico and countries of: Algeria, Argentina, Austria, Australia, Bangladesh, Bahrain, Belarus, Belgium, Brazil, Cameroon, Canada, Chile, China, Colombia, Congo-The Democratic Republic of the, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, Germany, Great Britain, India, Iran, Ireland, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Lebanon, Mauritius, Malaysia, Mexico, Morocco, Myanmar, Nepal, Netherlands, Nigeria, Norway, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Spain, Serbia (Republic of Serbia), South Africa, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Uganda, United Arab Emirates, Uzbekistan, Viet nam, and Yemen. | Open, Classified |  |  | 5451
Z-2752-2016 | 74930 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Coonrad/ Morrey Elbow Pin/ Bushing Replacement size Kit Extra Small. This product is sold sterile. The parts are placed into an inner cavity and are sealed using a nitrogen process. The inner cavity is then placed into an outer cavity and heat sealed with outer Tyvek lid on the outer cavity. The patient record label is placed on the Tyvek outerlid. The sealed cavities and package insert is placed into a folding carton, sealed and shrink wrapped. |  |  | August 16, 2016 | 2016 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a voluntary medical device recall of a single lot of the Coonrad/ Morrey Elbow Pin/ Bushing Replacement size Kit Extra Small because the lot was packaged with a missing humeral bushing. The missing bushing could result in a surgical delay or additional surgery if a replacement is not available. | Process design | Zimmer Biomet sent an Urgent Medical Device Recall Removal letter dated August 16, 2016, for their C/M Elbow Pin/ Bushing Replacement Kit via FedEx. All distributors were notified via electronic mail. Hospital risk managers, as well as distributors with product, were notified via courier. Customers were instructed to complete the following items including locating and removing the product in their territory, as well as identifying hospitals that currently have the affected product. Distributors will return on-hand product to Zimmer Biomet and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers were provided with a letter identifying the issue and their responsibilities. These responsibilities include:-Assisting the Zimmer Biomet sales representative with the quarantine of the product and completion and return of the Attachment 1 form.-Complete and return Attachment 1 (Acknowledgement of Responsibility) to the fax number or email address provided.If there are questions or concerns please call the customer call center at 1-800-348-9500 ext 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an on-call representative in the case of an emergency. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com.For questions regarding this recall call 877-946-2761. |  | 25 | 25 units | Worldwide Distribution - US including AL, AZ, CA, CT, GA, MN, NC, OH, OR and Internationally to Australia, Germany, Japan, Korea, and Taiwan. | Terminated | June 12, 2017 | 300 | 25
Z-2767-2016 | 75006 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Sterile Drape, SK6000 PAQ drape pouch (for use with the SPY Fluorescent System (SPY 2000/2001)) |  |  | August 18, 2016 | 2016 | NOVADAQ TECHNOLOGIES INC. | Novadaq is informing SK6000 PAQ users that a pinhole through the drape pouch film material was identified in some SK6000 PAQs during shipping validation testing. | Packaging | The firm, Novadaq, sent an "URGENT DEVICE CORRECTION: SK6000 PAQ Drape Pouch" letter dated 8/18/2016 to affected consignees. The Firm instructed consignee to cease distribution of affected product; accept the shipping-validated replacement PAQ; complete and sign the enclosed "Acknowledgment of Receipt" form, scan it and email to quality@novadaq.com. If you have any questions, contact Director of Regulatory Affairs at 1-905-629-3822 ext 205 or email: jpendlebury@novadaq.com.. |  | 20 | 20 units | US Distribution in states of: FL, MI, and VA. | Open, Classified |  |  | 20
Z-2768-2016 | 74974 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Certegra Workstation used in conjunction with Medrad Stellant CT Injection System |  |  | August 02, 2016 | 2016 | Bayer Healthcare | CertegraWorkstation when used in conjunction with the Medrad Stellant CT Injection System may exhibit thermal events, such as: smoke and sparks of fire. | Under Investigation by firm | Bayer Healthcare mailed an Urgent Field Safety Notification letter on August 03, 2016 to customers who were affected by the recall informing them of the issue and what actions to take if they were impacted. Customers were advised to use the enclosed Interim Instructions for Use for further information. For questions customers should call 1-877-229-3767. |  | 165 | 165 systems | Worldwide Distribution - US (nationwide) and Internationally to United Arab Emirates, Australia, Canada, Germany, Spain, and United Kingdom. | Open, Classified |  |  | 165
Z-2703-2016 | 74785 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ADVANCE(R) II MODULAR TIBIAL BASE W/SCREWHOLES, Knee Tibial Base, REF KTTIHA60, SIZE 6 STD; SURFACE POROUS, HA COATED; FIT CEMENTLESS; USE W/ SIZE 6 FEMUR &amp; SIZE 6 TIBIAL INSERT; CONTENTS, 1 EACH, STERILE R, Rx ONLY, MicroPort Knee Systems Total Knee Systems MicroPort Total Knee Systems are indicated for use in knee arthroplasty in skeletally mature patients with the following conditions: 1) noninflammatory degenerative joint disease including osteoarthritis, traumatic arthritis, or avascular necrosis; 2) inflammatory degenerative joint disease including rheumatoid arthritis; 3) correction of functional deformity; 4) revision procedures where other treatments or devices have failed; and treatment of fractures that are unmanageable using other techniques. |  |  | June 21, 2016 | 2016 | MicroPort Orthopedics Inc. | ADVANCE(R) HA Coated tibial bases are being recalled due to a higher than expected revision rate as a result of component loosening. | Use error | MicroPort sent an Urgent Field Safety Notice dated June 22, 2016, to all affected customers. The firm requested customers complete the enclosed form and return by fax 1-901-451-6032 or by email to Postmarket@ortho.microport.com. Customers were also requested to return the affected product. Non-responding consignees will be notified after 30 days.For questions regarding this recall call 901-867-4771. |  | 38 | 38 units | Worldwide Distribution only to Australia, Austria, France, Italy, Spain, and United Kingdom | Terminated | July 26, 2017 | 400 | 889
Z-2755-2016 | 74584 | No | GES | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Intermediate carton label reads in part as: "Bard-Parker***Stainless Steel Surgical Blades***#12B***QTY 150 (3 x 50)***REF 371712***LOT 0088464." |  |  | June 27, 2016 | 2016 | Aspen Surgical Puerto Rico Corp. | After receiving a complaint the firm disclosed that Conventional Stainless Steel Surgical Blade catalog number 371712 lot 0088464 was incorrectly identified in the single unit foil pack as 12 instead of 12B. | Mixed-up of materials/components | An "Urgent: Medical Device Correction" letter was sent on June 27, 2016. The letter provided background on the issue, as well as actions to be taken by the user. Customers have the alternative to either keep the product or return it for replacement. A response form was included for customers to complete and return. Customers are to maintain awareness of the incorrect identification prior to use. Questions concerning the correction can be directed to Pedro.Garcia@WelchAllyn.Com. |  | 6000 | 6,000 pieces | US distribution only. | Terminated | May 15, 2017 | 322 | 6000
Z-2759-2016 | 75020 | No | GDO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | WECKÂż Horizon" Manual-Load Ligating Clip Applier, For Medium-Large Size Horizon Ligation Clips Curved Jaw, 11 (28cm), Rx Only, REF 337111, |  |  | August 24, 2016 | 2016 | Teleflex Medical | Misbranded: Incorrect etching on the device. | Error in labeling | Teleflex sent an Urgent Medical Device Notification letter dated August 24, 2016 to customers via FedEx. The letter identified the affected product, problem and actions to be taken. For questionscontact your local sales representative or Customer Service at 1-866-246-6990. |  | 49 | 49 units | Worldwide Distribution US Nationwide in the states of FL, IL, KY, MI, NY and OK and countries of Mexico and Canada. | Open, Classified |  |  | 49
Z-2761-2016 | 74891 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Nylon Surgical Suture for human use and vet use, sold under brand name Riverlon, Securolon, Weblon. Product is sterile. Part numbers: |  |  | August 05, 2016 | 2016 | Riverpoint Medical, LLC | Nylon Surgical Sutures are recalled because the seals may become open during distribution or subsequent packaging. | Nonconforming Material/Component | Riverpoint Medical sent an Urgent Medical Device letter, dated August 5. 2016, and the Customer Response Form to their customers via email. The letter identified the affected product, problem and actions to be taken. Customers are instructed to check inventory, discontinue selling affected sutures and return affected devices to the firm. Distributors are instructed to notify their downstream customers. Consignees should submit the Customer Response Forms. For questions call Riverpoint Medical at 1-866-445-4929. |  | 2182 | 2110 boxes distributed in the US and 72 boxes distributed outside US | US Nationwide distribution in the states of ID, MA, OR, NY and in the country of Australia. | Open, Classified |  |  | 2243
Z-2764-2016 | 74893 | No | GWL | Physiological signal amplifier. | Neurology | Neurology | 510(k) | Patient Cable. Part Numbers: 4460, 6735. Sterile within Chipboard Boxes within corrugated cartons. |  |  | August 10, 2016 | 2016 | Blackrock Microsystems, LLC | Blackrock Microsystems, LLC announces a voluntary field action for the Blackrock NeuroPort Biopotential Signal Processing System because the Neural Signal Amplifier and the Patient Cable are incorrectly labeled as Type CF Applied Parts. | Labeling False and Misleading | Blackrock Microsystems sent an Urgent Medical Device letter dated August 12, 2016 via email. The letter identified the affected product, problem and actions to be taken. Customers were instructed not use the amplifier in application requiring a type CF applied part; perform a physical count of the inventory of amplifiers and patient cables with the part numbers listed. For questions regarding this Correction call 801-582- 5533, ext. 240. |  | 120 | 120 total cables | Worldwide Distribution - US Nationwide and Canada, | Open, Classified |  |  | 120
Z-2753-2016 | 74722 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Axiom REG implants Sterile |  |  | June 20, 2016 | 2016 | Anthogyr | Following a FDA inspection showing that the firm failed to validate the design and process of its implants. | Device Design | The firm, Anthogyr, contacted physicians by phone and sent an " URGENT: MEDICAL DEVICE RECALL" notification by mail, dated 06/20/2016, informing users of a recall due to inadequate validation of design and process of implants by the FDA. The notice described the product, problem and actions to be taken. The physicians were instructed to isolate the parts and return any Anthogyr implants from lots subject to recall to Business Area Manager, 247 SouthWest 28th road, Miami, FL 33129;and complete and return the RETURN RESPONSE form by fax to: +33 (0)4 50 93 78 60 or e.mail: j.duthil@anthogyr.com or mail: Anthogyr- Director of Quality &amp; Regulatory Affairs- 2237 avenue- Andre Lasquin-France. The firm stated that they will not market their products in the United States until they obtain validation through 21 CFR Part 820. If you have any questions, call 404-718-9689 or email: jduthil@anthogyr.com or mfoley@anthogyr.com. |  | 111 | 111 units | US Distribution to sates of: NC, NY, and OR. | Terminated | March 28, 2017 | 281 | 121
Z-2756-2016 | 74934 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Aortic Arch Cannula and Coronary Artery Perfusion Cannula with Ballon (part number NA-55X6). The Aortic Arch Cannulae are designed to be used in the extracorporeal circuit during cardiopulmonary bypass surgery. The cannula consists of varying lengths of non-wire reinforced (flexible) polyvinyl chloride tubing that terminates in an angled tip. These devices come in direct contact with the central circulatory system but they are not intended to control, diagnose, monitor or correct a defect. |  |  | August 12, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field for the Aortic Arch Cannula and Coronary Artery Perfusion Cannula With Ballon because the tip of the cannula should be bent at either a 45 or 90 degree angle. Samples have been returned where the angle of the tip is less than the requirement. | Process control | LivaNova issued an Urgent Safety Alert letter via certified mail on August 12, 2016 to all affected customers. The letter instructed customers to check their inventory to determine if the Cannulae belongs to the product code and lots listed. If the customer does not have any of the products described in the communication, please complete the response form below and return it per the instructions indicated on the form; otherwise, please check the tip angle to determine if it is conforming to specifications. If specification is not met, please file a complaint through your normal process whether you decide to keep it and use it or return the product to Sorin Group; if you are not willing to use it, return the affected parts to Sorin Group USA. Please contact Customer Service 1-S00-650-2623 or email CustomerService@livanova.com to have a Return Material Authorization (RMA) issued. Complete the Customer Response Form attached to this letter to acknowledge receipt of this letter. Please return this form no later than September 15, 2016.For questions regarding this recall call 303-467-6306. |  | 200 | 200 units worldwide | Worldwide Distribution - US including IL, MO, TX. Internationally to Canada and Italy. | Terminated | March 27, 2017 | 227 | 1405
Z-2750-2016 | 74760 | No | KNC | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Affinity Four Birthing Bed Model P3700 |  |  | July 15, 2016 | 2016 | Hill-Rom, Inc. |  | Nonconforming Material/Component | Hill-Rom Inc. initiated a voluntary recall of the Affinity Four Birthing Beds on July 15, 2016, with letters to customers mailed certified via FedEx. Customers were asked to do the following: During use, if you encounter a situation where the CPR mechanism fails to activate, the head of the bed can still be lowered, electrically, using the siderail articulation controls. A Hill-Rom representative will be contacting each facility to make arrangements for replacing the affected head drive.Please pass this notice on to all those who need to be aware within your organization and/or to any organization where the affected devices have been transferred.If you have any questions concerning this correction, please contact Hill-Rom Technical Support at 800-445-3720. |  | 94 | 94 beds | US Distribution to the states of : CA, FL, IN, KS, LA, MI, MO, NC, NY, OK, OR, PA, TN, TX, WA, WYForeign:BULGARIA, INDIA, UNITED ARAB EMIRATES | Terminated | March 23, 2017 | 251 | 94
Z-2751-2016 | 74798 | No | KKY | Polytetrafluoroethylene with carbon fibers composite implant material. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hard Tissue Replacement (HTR) Implant, Rx only, for use in single procedure only. Custom made. |  |  | June 13, 2016 | 2016 | Biomet Microfixation, LLC | Potential sterility deficiency. | Packaging | Zimmer Biomet Customers were sent an Urgent Product Recall letter dated June 13, 2016, via via FedEx mail. The letter notified customers of a recall of Hard Tissue Replacement Polymer Implants due to voids in sterile packaging. The firm instructed customers to locate and discontinue use of the implants, complete the Fax Response form and fax it to 904-741-9425, send back any quarantined product using the FedEx labels provided along with the recall notification, and to notify facilities where the implant was further distributed by forwarding them a copy of the provided notice. The firm instructed distributors to forward the notice to officials of their customers who are in charge of receiving recall notices. Additionally, the firm informed distributors that they are responsible for locating and returning distributed product. Questions are to be directed to 1-800-874-7711 or 904-741-4400 x9574, Monday thru Friday from 8am to 4pm ET. Her after hours phone is 812-614-9499.For questions regarding this recall call 904-741-9225. |  | 54 | 54 units | Worldwide Distribution - US including AZ, CA, FL, MO, NH, PA, TX, &amp; VA, and Internationally to Austria, Canada, England, Germany, Greece, Italy, Netherlands, &amp; South Africa. | Open, Classified |  |  | 54
Z-2769-2016 | 74649 | No | BZC | Pulmonary-function data calculator. | Anesthesiology | Anesthesiology | 510(k) | Innocor Cardiopulmonary Exercise and Lung Function Testing System |  |  | May 13, 2016 | 2016 | Innovision A/S | The purpose of the recall is to replace data transmission cable (USB-cable) connecting the lnnocor device with an external computer.The replacement USB cable will provide galvanic isolation of the lnnocor from the external computer via opto-couplers, as opposed to the grounding protection currently in use.The replacement is intended to prevent risk of electric shock. | Device Design | Recall notifications described the issue and actions to be taken by the firm. The notice also stated that replacement costs would be carried by Innovision ApS. Acknowledgment of the notice can be made by contacting service manager Knud Pedersen at kp@innovision.dk or (+45) 65 95 91 00) to schedule the service activity and to answer any questions. |  | 8 | 8 | Worldwide Distribution -- US, to Massachusetts and Washington, DC; and, the countries of Canada, Germany, and United Kingdom. | Open, Classified |  |  | 8
Z-2766-2016 | 74839 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SureSight Obturator Introducer Localization Systems (ILS) for ATEC in MRI. |  |  | August 02, 2016 | 2016 | Hologic, Inc | Tip Breakage at the glue joint between the obturator body and tip | Process control | Hologic sent an Important Product Recall letter to customers in the United States through Certified letter. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine their inventory and cease use of all affected recalled product. Customers were asked to complete the enclosed Recall Response Card indicating whether you have any of the recalled product on hand, as well as the quantity of product that you have in stock. |  | 535 | 535 units | US Nationwide Distribution | Terminated | March 30, 2017 | 240 | 1975
Z-2777-2016 | 73689 | No | MDD |  |  | General &amp; Plastic Surgery | PMA | Integra Meshed Dermal Regeneration Template 5 cmx 5 cm (2 in x 2in) Rx Only |  |  | March 11, 2016 | 2016 | Integra LifeSciences Corporation | Integra's post QA release review of historical product release test results for Meshed IDRT products identified the Peel Strength test average result was incorrectly calculated for a single Lot (Lot 105A00324750). | Under Investigation by firm | Integra LifeSciences Inc. sent an urgent voluntary medical device recall letter/recall acknowledgement and return form dated March 11. 2016. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for the affected product and immediately stop using and remove from service. Customers were asked to complete the attached form and return by email or fax as indicated on the form. Customers with questions were advised to contact Customer Service at 1-888-601-0203.For questions regarding this recall call 609-275-2700. |  | 14 | 14 packages | Nationwide Distribution to PA | Open, Classified |  |  | 14
Z-2770-2016 | 74654 | stains, hematology | KQC | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Fluorocell PLT reagent, Catalog #CD994563 |  |  | July 01, 2016 | 2016 | Sysmex America Inc | There is a potential that fluorescent platelet (PLT-F) and immature platelet fraction (IPF%) results performed using the recalled lots may exhibit lower than expected PLT-F and IPF values. The issue may be recognized by a large discrepancy between the Impedance platelet (PLT-I) value and a false low PLT-F value. In most cases, the falsely decreased PLT-F results displayed a "PLT Abn Scattergram" message with asterisks beside the results indicating the data is unreliable. In rare instances, platelet flags may be absent. The PLT-I counts are unaffected.The PLT-F is a reflex test used to confirm PLT-I results due to flagging or low values. As such, the likely impact to patient results is low. Affected results display a discrepancy between the PLT-I and PLT-F, with lower than expected PLT-F results that do not match the patient's clinical picture. | Under Investigation by firm | Sysmex sent a Product Notification Letter dated August 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. . Customers were instructed to cease use and discard the affected lots, and reply to Sysmex America, Inc. using the provided fax-back form.An updated Product Notification (Attachment 1) was mailed to all XN-Series customers on August 12, 2016 informing them of the additional recalled lots. This notification included the phone number for the Technical Assistance Center at 1-888-879-7639 in the U.S., and 1-888-679-7639 in Canada. |  | 3868 | 3868 units | Worldwide Distribution - US (Nationwide, Washington DC) Canada; Cuba; Chile; and Dominican Republic. | Open, Classified |  |  | 3868
Z-2771-2016 | 74898 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Catheter, Intravascular, Therapeutic, Long Term Greater than 30 days. |  |  | February 17, 2016 | 2016 | Arrow International Inc | Notification of component molding abnormality by supplier used to package certain Arrow Select Kit (ASK) products. | Unknown/Undetermined by firm | The firm, Arrow International, sent an "Urgent Medical Device Recall Notification" letter dated 02/19/2016 via FedEx 2-day delivery to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any of the affected products found; return product-a customer service representative will contact you with a RGA number and instructions; and complete and return the enclosed Recall Acknowledgement Form via fax to 1-855-419-8507 Attn: Customer Service or email to recalls@teleflex.com, even if you do not have no affected product.If you have any questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 6121 | 6,121 units in total | US Distribution to states of: NC, OH, FL, and MD. | Open, Classified |  |  | 6121
Z-2776-2016 | 75012 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 EP Navigation System and Accessories, Product Codes: FG540000 and FG540000C, Software Versions: v 3.2 , v 4.3 and v 2.3. |  |  | August 15, 2016 | 2016 | Biosense Webster, Inc. | Biosense Webster, Inc. is initiating a voluntary correction to re-install identified CARTO 3 Systems that were moved and installed in China without testing and approval by Biosense Webster Technical Services, as described in the Instructions for Use (IFU). | Under Investigation by firm | Biosense Webster sent a Customer Notification letter dated August 23, 2016, to all affected customers to inform them that Biosense Webster, Inc. recently became aware that some systems were improperly transported and installed at hospitals in China by Biosense Webster, Inc. Clinical Specialists. The letter informs the customers that according to the CARTO 3 System Instructions for Use (IFU), only the Biosense Webster, Inc. Technical Services team members are authorized to properly install the systems. Customers are informed that each impacted CARTO 3 System must be fully tested and approved by the Biosense Webster, Inc.Technical Services Team. The letter informs the customers of the actions to be taken. Actions Requested on Your Part:1. Read this Medical Device Correction notification carefully.2. If a CARTOÂż 3 System currently resides within your hospital, the system has been corrected, and you should ensure that you have received a Field Service Report that indicates the system is in order .For questions regarding this recall call 909-839-8500.3. Review, complete, sign, and return the Acknowledgement Form in accordance with the instructions listed on the form.4. In the future, if a CARTOÂż 3 System is installed in your facility, you should only conduct a clinical procedure after you have received a Field Safety Report dated after the installation of the device indicating that the system is in order .5. Maintain awareness of this notice and pass this notice on to anyone in your facility that needs to be informed.Customers with any questions are instructed to contact their Biosense Webster, Inc. sales representative. |  | 338 | 338 units | International Distribution to China, Hong Kong and Brazil. | Open, Classified |  |  | 338
Z-2775-2016 | 74837 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | Endura MR Mass Spectrometer u using software versions Endura MD Software 1.0 and 1.0 SP1. |  |  | July 19, 2016 | 2016 | Thermo Finnigan LLC | Thermo Fisher has determined that the Endura MD mass spectrometer instrument control software versions 1.0 and 1.0 SP! have a software defect which affect data accuracy. | Software design | ThermoFisher Scientific sent an Urgent Medical Device Field Correction letter dated July 19, 2016, to all affected customers via FedEx. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to discontinue use of the affected product until a company authorized Field Service Engineer installs corrected software to the instrument system. Customers were also instructed to communicate to all who operate the instrument and to complete the provided Acknowledgment Form, sign, and return it via email FieldActionTeam@thermofisher.com. Customers with questions should contact the Field Action Team at FieldAction@thermofisher.com. |  | 6 | 6 (5 in Us distribuion and 1 in Japan) | Worldwide distribution - US including FL, NY and Internationally to Japan. | Open, Classified |  |  | 6
Z-2779-2016 | 73989 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VIDAS NT-pro-BNP |  |  | April 15, 2016 | 2016 | Biomerieux France Chemin De L'Or | Due to a packaging error; The kit does not contain the correct number of C1 or C2 control vials. | Packaging | bioMerieux. sent an Urgent Product Removal Notice letter dated April 26, 2016 to customers via FedEx. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check for affected products in their inventory and discard the incomplete products. Customers were asked to complete the attached Acknowledgement Form and return it via fax to bioMerieux. For questions contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666. |  | 229 | 229 kits | US Nationwide Distribution in the states of: AL, AZ, CA, FL, GA, IA, KS, KY, MA, MI, MN, NC, NE, NJ, NY, OH, SC, TX, VA and WV. | Terminated | July 21, 2017 | 462 | 229
Z-2778-2016 | 74940 | No | LBS | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | LDX Optical Check Cassettes (OCC) |  |  | August 02, 2016 | 2016 | Alere San Diego, Inc. | Incorrect expiration date on the magnetic strip that does not match the labeled expiration date. | Packaging | Alere issued a recall notification to customers via telephone and email. Customers was informed of the affected product, problem and actions to be taken. Customers were instructed if they have the affected lot contact the Alere technical service department at 877-308-8289 to report this lot and receive a free replacement. For questions call 877-308-8289 and reference their specific case number. |  | 12951 | 12,951 | Worldwide Distribution - U.S Nationwide in the states of: TN, MA, GA, OK, VA, FL, TN, MO, AL, TX, IL, NY, VT, RI, KY, WI, AL, NC, LA, SD, TX, VA, MI, MN, IN, IA, CA, AR, PA, HI, WV, KS, CO, CT, UT and the countries of Japan, England, Canada | Open, Classified |  |  | 12951
Z-2781-2016 | 74899 | No | LCJ |  |  | General &amp; Plastic Surgery | 510(k) | CPX4 Breast Tissue Expander 650 cc |  |  | August 09, 2016 | 2016 | Mentor Texas, LP. | The CPX4 Tissue Expander may have been packaged in the Artoura Tissue Expander box. | Employee error | Mentor Texas sent an Urgent Field Safety Notice dated August 11, 2016 to customers.via letter and a follow-up with respective Sales Representative visit. The letter identified the affected product, problems and actions to be taken. Customers were informed the firm will credit/replace returned affected products. For questions contact your Sales Representative . |  | 30 | 30 units | Worldwide Distribution - US Nationwide and the countries of: Germany, UK, and Belgium | Open, Classified |  |  | 78
Z-2785-2016 | 74851 | No | LDH |  |  | General Hospital | 510(k) | GREER 2 ML Sterile Empty Vials 2-13 vial 25 ea - Silver Cap |  |  | February 03, 2015 | 2015 | Allergy Laboratories, Inc. | Potential for glass particles within the vials | Process control | Allergy Laboratories notified the sole consignee via phone call on February 3, 2015. For further information, please call (405) 235-1451. |  | 81625 | 81,625 vials (3,265 packs) | US Distribution to the state of : NC | Terminated | March 13, 2017 | 769 | 81625
Z-2786-2016 | 74931 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | AccuLIF PL 6-9mm x 11mm x 27mm x 8 Cage, Rx only, Sterile R, Legal The AccuLIF PL Cages are indicated for intervertebral body fusion with autograft and/or allogeneic bone graft comprised of cancellous and/or corticocancellous bone graft when the subject device is used as an adjunct to fusion in patients with degenerative disc disease (DDD) at one level or two contiguous levels from L2 to S1. DDD is defined as back pain of discogenic origin with degeneration of the disc confirmed by history and radiographic studies. Additionally, the AccuLIF PL Cages can be used as in adjunct to fusion in patients diagnosed with degenerative scoliosis. The AccuLIF Pl Cages are always to be used with supplemental internal spinal fixation. Additionally, the AccuLIF PL Cages are to be used with autograft and/or allogeneic bone graft comprised of cancellous and/or corticocancellous bone graft when the subject device is used as an adjunct to fusion. |  |  | July 26, 2016 | 2016 | Howmedica Osteonics Corp. | Since Stryker acquired the AccuLIF product in March of 2014, there have been Product Inquires (PIs) related to the PL implant confirming to have a reduction of height &gt; 1 mm. The current occurrence rate for reported post-operative reduction of cage height (&gt; 1mm) of AccuLIF PL Expandable Cages is approximately 0.47% and has demonstrated a slight upward trend since September 2015. | Unknown/Undetermined by firm | Stryker Spine sent an "Urgent Recall Letter and Product Accountability Form" dated July 26, 2016 to all affected customers. For patients who have had an AccuLIF Posterior Lumbar (PL) Expandable Interbody implant, Stryker Spine is recommending routine clinical and radiographic post-operated evaluation. Should the patient report any change in or develop near-onset symptoms, more urgent clinical and radiographic evaluation should be completed. Customers with questions should call 201-749-8389 or 201-749-8346. |  |  |  | NationwideDistribution | Open, Classified
Z-2782-2016 | 74932 | No | PCL | Rubeola (measles) virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Trinity Biotech Captia Measles IgM, REF/Product Code 2326060,Test Kit for 96 Determinations, IVD |  |  | July 11, 2016 | 2016 | Clark Laboratories, Inc. (dba,Trinity Biotech USA) | The product lot failed to meet QC acceptance criteria during stability testing. The Calibrator OD was dropping causing the positive control to be invalid (running high out of range). | Nonconforming Material/Component | Recall notification letters (dated 7/11/2016) were sent by FedEx on 7/11/2016. Customers were instructed to discontinue use of any remaining inventory (discard on site if any remaining). Customers were also advised that if the test resulted in a valid run, such that all Controls and Calibrator met the Quality Control Criteria as stated in the package insert, the patient results can be reported and a re-evaluation of patient results is not necessary. For questions, please contact Trinity Biotech's Help Desk at 1-800-325-3424 Option #3. |  | 56 | US: 6 kits, Foreign: 50 kits | Distribution to US nationwide including California, and China (Hong Kong). | Terminated | May 09, 2017 | 302 | 56
Z-2791-2016 | 74727 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Supracondylar Nail T2 SCN 09x300 mm |  |  | June 30, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Stryker discovered that a Tibial Nail and femoral Nail were incorrectly labeled and distributed to the field. | Labeling mix-ups | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 30, 2016 by e-mail to its customer requesting that they quarantine affected devices and sent the letter via UPS (with return receipt) on July 1, 2016. The letter described the product, problem and actions to be taken. The customer was instructed to inform users of the Medical Device Removal and forward the notice to all who need to be aware; inspect your inventory and return affected products to Stryker Orthopaedics/PFA Product Returns, Attn: Distribution Inventory Team; 325 Corporate Drive; Dock M-East , Mahwah, NJ 07431 REF: PFA #2016-094; complete and return the Business Reply Form via email to SO_M_PRODUCT_FIELD_ACTION_RESPONSE@stryker.com within 5 days; and keep a copy of the completed and executed Business Reply Form for your records. If you have any questions, feel free to contact the Mgr., Divisional Regulatory Compliance, Marketing Quality at (201) 831-6693. |  | 1 | 1 unit in total | US distribution to state of: CA. | Terminated | June 21, 2017 | 356 | 1
Z-2789-2016 | 75015 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Patella Assembly Instrument, REF 3182-1000, Non-Sterile, NDC (01) 07613327215854 (10) |  |  | August 01, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Upon failure of the press-fit, the clamping sub components and pins could disassociate from the Scorpio Patella Assembly Instrument. | Nonconforming Material/Component | Stryker sent an "Urgent Medical Device Recall Notification" Letter and Recall Business Reply Form dated August 1, 2016 to their affected customers via email and also via UPS on August 1, 2016.Customers are asked to quarantine the product subject to this recall. Branches/Agencies are asked to return all affected products available at their location to: Stryker C/O Stericycle, Attn: RA 2016-098 - Event 8296. Hospital/Branches: Customers are asked to complete and sign the Business Reply Form and fax a copy within 5 business days to: 877-272-6141 or email to Stericycle strykerortho8296@stericycle.com. Any questions, customers can call (201) 831-6693. |  | 2039 | 2,039 units | US Nationwide and Internationally | Terminated | June 08, 2017 | 311 | 2039
Z-2790-2016 | 74903 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira MedNet Medication Management Suite software, version 6.1, List Numbers 16037-64-02 and 16037-64-03, in combination with the SapphirePlus 13.1x infusion pump |  |  | July 28, 2016 | 2016 | Hospira Inc., A Pfizer Company | Hospira MedNet 6.1 software, in combination with the SapphirePlus 13.1x infusion pump, may result in incorrect bolus amount calculations for drugs whose concentrations are listed in Million Units. | Software design | Hospira sent an Urgent Medical Device Correction letter dated July 28, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter provided instructions for step-by-step process to mitigate the issue. Customers with questions were instructed to contact Hospira Global Complaint Management at 1-800-441-4100 or Hospira Technical Service Center at 1-800-241-4002, option 4. |  | 20 | 20 | Worldwide Distribution - US including ND, TX, MI, MN, WI, IA, ME, HI, PA, MA, NJ, NY, CT, and Internationally to Canada. | Open, Classified |  |  | 20
Z-2793-2016 | 75125 | No | LLG | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Syngo Plaza picture archiving and communication system (PAC) (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images. |  |  | August 02, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is releasing a Customer Advisory Notice to explain meaning of a suffix o.p. that shows on some DICOM images on area, distance or perimeter values. This suffix stands for on projection and indicates that a shown measurement value may be different from the real value. | Error in labeling | Siemens mailed to customers on August 04, 2016 a Customer Advisory Notice that provided interpretations about the meaning of measurement values like area, distance and perimeter shown as "o.p." by the syngo.plaza. |  | 101 | 101 Systems | Nationwide Distribution | Terminated | March 20, 2017 | 230 | 101
Z-2795-2016 | 74732 | No | PII |  | Microbiology | Microbiology | 510(k) | T2CandidaÂżCartridges; Part Number: 90-01794 (box of 12) and Part Number: 91-01534 (individual cartridge base). |  |  | July 07, 2016 | 2016 | T2 Biosystems Inc | Elevated levels of invalid results. | Under Investigation by firm | Customers were notified of the recall vie e-mail and certified mail on July 7, 2016, with an Urgent Medical Device Recall letter. Customers were asked to take the following actions:1. Please examine your current inventory and determine whether you have any of the recalled lots listed. (a.) If you DO have cartridges from the recalled lots, please note the quantity of cartridges you have of each lot number on the attached form entitled T2Candida Cartridge Recall Customer Response Form (CRF). Please also immediately quarantine any cartridges from the indicated lots and do not use them. (b.) If you DO NOT have any cartridges from the recalled lots, please also note this in the space provided on the form2. Send the completed CRF to T2Biosystems via e-mail to orders@t2biosystems.com or via fax to 781-240-0541. Your primary contact at T2 for all recall-related correspondence or questions is Georgette Yannios (gyannios@t2biosystems.com or 877-504-8282 Option 3). All product will be replaced at no charge to you as soon as possible.3. Once T2 receives the CRF, we will issue a Return Material Authorization (RMA) number to you along with a FedEx shipping label. After you receive this RMA number and label, please follow the instructions provided with the RMA for shipment of cartridges back to T2 (please do not ship reagent trays as they are not subject to this action). |  | 6864 | 6,864 | Distributed to the states of UT, FL, NH, PA, GA, MI, AL, IN, NJ, CA, NY, LA, RI, TN, TX, WI, and CT. | Open, Classified |  |  | 6864
Z-2796-2016 | 74815 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | May 06, 2013 | 2013 | Merge Healthcare, Inc. | Multiple studies were missing at a customer location that never made it to the archive before watermarking | Software design | The recalling firm first issued a letter via email dated May 6, 2013. A letter dated January 28, 2016, was also issued via email. Both letters informed the customer a software upgrade was available and of the issue and makes recommendations about using the software. They are also informed an upgrade is available to correct the issue. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it.For further questions please call (800) 724-5970 Ext. 8027. |  | 45 | 45 sites potentially have the affected versions | Distribution was nationwide to medical facilities. Government distribution was made to LA and OK. | Open, Classified |  |  | 45
Z-2802-2016 | 74910 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | RAPIDLab 1260 Blood Gas Analyzer |  |  | August 11, 2016 | 2016 | Siemens Healthcare Diagnostics Inc | There is a potential for the first and/or last name of one patient to be printed with Patient ID and result data from a different patient, even though those fields have been turned Off for the system | Software design | Siemens Healthcare issued Urgent Field Safety Notice to (US) customers via Federal Express. . All affected customers outside the US provided a copy of an Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each countries local regulations and procedures.This notice informs customers of the issue, Actions to be Taken by the Customer and to return the response form.Questions contact your Siemens Customer Care Center or your local Siemens technical support representative |  | 114 | 114 units | NationwideForeign:AlbaniaAlgeriaAngolaArgentinaArmeniaAustraliaAustriaBahrainBangladeshBelarusBelgiumBosnia Herzeg.BotswanaBrazilBruneiBulgariaCanadaChileColombiaCroatiaCzech RepublicDenmarkEcuadorEgyptEstoniaFinlandFranceFren.PolynesiaGeorgiaGermanyGreeceHong KongHungaryIcelandIndiaIndonesiaIraqIrelandIsraelItalyIvory CoastJapanKazakhstanKenyaKuwaitLatviaLebanonLesothoLithuaniaLuxembourgMacedoniaMalaysiaMexicoNetherlandsNew CaledoniaNew ZealandNorwayP.R. ChinaPeruPhilippinesPolandPortugalQatarRep. of YemenRepublic KoreaRomaniaRussian Fed.Saudi ArabiaSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTaiwanThailandTunisiaTurkeyTurkmenistanU.A.E.United KingdomUruguayVatikancityVietnam | Open, Classified |  |  | 13245
Z-2830-2016 | 74712 | No | EZL | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Option-vf Urinary Catheter. Female use only. Latex Free. Sterile, Rx only. |  |  | July 11, 2016 | 2016 | C.R. Bard, Inc. | During an FDA inspection it was found out that the Practical Foley Catheters to be potentially nonsterile. | Device Design | C.R. Bard sent an Urgent - Medical Device Product Recall letter dated July 8, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm notifies customers of the recall, provides a clinical risk statement, and gives instructions regarding product disposition. Facilities are instructed to examine their inventory and quarantine any recalled product. The firm requested that customers complete the Recall &amp; Effectiveness Check Form if product is or is not in inventory. If product was further distributed, customers should be forwarded the recall notification letter and Recall&amp; Effectiveness Check Form.If you or the patient using these catheters has had an adverse event related to the recalled catheters, please contact Bard Medical Division Field Assurance at 1-800-526-4455 (option 5, then option 4) or via email at BMD.FieldAssurance@crbard. com. |  | 274 | 274 units | US Distribution to the states of : AL, CA, CO, FL, IL, IN, MA, MD, NJ, NY, OR, PA, TN, WA, and WI. | Terminated | March 22, 2017 | 254 | 274
Z-2804-2016 | 74009 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The VITROSÂż Chemistry Products Calibrator Kit 31, Lot 3155 used in conjunction with: VITROSÂż Chemistry Products HbA1c Reagent Kit, Generation 6 (GEN 6) and Assay Data Disk (ADD), Data Release Versions (DRV) 5873, 5874, 5875, or 5876, is used to calibrate the VITROS 5,1 FS Chemistry System, VITROS 4600 Chemistry System and the VITROS 5600 Integrated System for the determination of percent glycated hemoglobin (HbA1c) in human whole blood. |  |  | April 19, 2016 | 2016 | Ortho-Clinical Diagnostics | There is an incorrect value (data/calibration mathematics) on ADDs. This incorrect value will prevent a successful calibration of the assay. | Employee error | On 19 April 2016, Ortho Clinical Diagnostics distributed Correction Notification letters (Ref. CL2016-087) and ADD DRV 5877 notices to their US consignees via courier service and to their foreign affiliates via email. Customers were advised to install the Assay Data Diskette DRV 5877 or above on their VITROS system. Following a successful calibration and quality control assessment, it is acceptable to use VITROS HbA1c Reagent Kit, GEN 06. Customers who further distributed the product should also provide a copy of the enclosed letter to all customers who received the product. It is suggested that customers complete and return the Confirmation of Receipt form to indicate that you have been informed of this issue. Customers with questions can contact the Technical Solutions Center at 1-800-421-3311. |  | 520 | 520 Units | Worldwide Distribution - US Nationwide in the states of: AL, AR, AZ, CA, CT, DC, FL, GA, IL, IN, KY, LA, MD, MI, MO, MS, NC, NY, OH, OK, OR, PA, SC, TN, TX, WA, WV &amp; Puerto Rico;** Foreign: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy &amp; Spain. | Open, Classified |  |  | 520
Z-2832-2016 | 74991 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | The Bigger-Better-Bladder, TM 2 With 3/8"ID tubing, 72" long tubing, ITEM BBB38-72 |  |  | July 01, 2016 | 2016 | Circulatory Technology Inc | The seal between the balloon and the housing may leak. | Device Design | The firm, Circulatory Technology Inc., sent an amended "Urgent: Expanded Device Recall" letter dated July 1, 2016, and response form to their customers via email. The letter described the product, problem and actions to be taken. The customers were instructed to immediately inspect their inventory, quarantine and prepare to ship the product back to Circulatory Technology Inc. Customers were also instructed to complete and return the response form via email to ytamari@cirtec.com. Customers with questions can call 516-624-2424, Monday through Friday 9:00 AM - 7:00 PM, EST. |  | 20 | 20 units | US Distribution to states of: LA, MI, NY, OH, OR &amp; TX. | Open, Classified |  |  | 20
Z-2834-2016 | 74825 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | LAUREATE World Phaco System |  |  | July 28, 2016 | 2016 | Alcon Research, Ltd. | Steps 16-25 are missing from the LAUREATE Multipak FMS Directions for Use (DFU) | Labeling Change Control | All remaining inventory of the identified affected products within Alcon DCs will be supplied corrected directions for use material prior to any further shipments. |  | 8646 | 8,646 UNITS | Burundi, China, Iran, South Korea, Taiwan, Thailand, USA. *** The recalled product is not sold in the USA or its territories. However, donated units are in use by a reduced number of USA clients for Medical Missions outside the country. | Terminated | June 07, 2017 | 314 | 8646
Z-2809-2016 | 74966 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Transpac IV Monitoring Kit with 2 SafeSet Reservoir and 2 Blood Sampling Ports, 84" Tubing, Disposable Transducer, 3 mL Squeeze Flush, Macrodrip (Pole Mount), Item No. 42646-06 |  |  | August 10, 2016 | 2016 | ICU Medical, Inc. | ICU Medical Inc. has identified a potential for leak with certain SafeSet Reservoir and blood sampling kitscontaining the Transpac IV transducers. | Process control | An Urgent Medical Device Recall Notification letter dated 8/10/16 was sent to all their customers who purchased the Transpac IV with SafeSet reservoir kit and without SafeSet Reservoir, and with Transpac IV with marvelous valve without SafeSet reservoir. The letter informs the customers that ICU Medical Inc. has identified a potential for leak with certain SafeSet Reservoir and blood sampling kits containing the Transpac IV transducers. The recall letter list the affected products and lot numbers. Customers are informed of the actions to be taken. Customers with questions are instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address: customerservice@icumed.com. Customers are instructed to complete the attached Medical Device Recall Response Form and fax to (801) 264-1755 or e-mail to recall@icumed.com. |  | 4180 | 4180 units | US Nationwide Distribution | Terminated | July 10, 2017 | 334 | 9205
Z-2829-2016 | 74963 | No | LQK | Home uterine activity monitor. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | AirStrip Sense4Baby Model C Fetal Monitor |  |  | August 04, 2016 | 2016 | AirStrip Technologies, Inc. | Belt clip may become detached causing electronic components to be exposed. | Nonconforming Material/Component | The firm sent out customer notification letters on August 4, 2016. Customers were instructed to cease distribution and use of the device and notify any customers in possession of it. Customers were informed that Airstrip would make arrangements to return the affected Central Units to AirStrip for correction at AirStrips expense. If customers have not been contacted by AirStrip to make these arrangements call AirStrip tollfree at 1 (877) 258-5869.For further questions, please call (210) 805-0444. |  | 42 | 42 | U.S. distribution to the following; AR, LA.Foreign distribution to the following; South Africa, Netherlands. | Terminated | November 30, 2016 | 118 | 42
Z-2841-2016 | 73995 | No | CFL | Human growth hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IG1, IGF-I reagent, Catalog # LKGF1; For the quantitative measurement of insulin-like growth factor I (IGF-I) in serum or heparinized plasma. |  |  | April 15, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | 24 minutes of incubation is required prior to processing patient samples and for the pretreatment solution to reach full equilibration when using pretreatment solution when using the pretreatment solution (LGFA) contained in the kit. | Unknown/Undetermined by firm | On April 15, 2016, Siemens Healthcare Diagnostics Inc. distributed Urgent Medical Device Correction notices to their domestic customers and Urgent Field Safety Notices to their internationally customers via courier service. Customers are advised to take the following four actions: 1) To continue using IGF-I kit lot 411 and the associated pretreatment solution lot 055 to report patient results, ensure the IGF-I pretreatment solution is allowed to incubate with the patient sample for a minimum of 24 minutes prior to processing on the IMMULITE/IMMULITE 1000 systems. 2) Complete and return the Effectiveness Check Form attached to the letter within 30 days. 3) Please review the letter with your Medical Director. 4) If you have received any complaints of illness or adverse events associated with the product, immediately contact your local Siemens Customer Care Center or your local Siemens technical support representative. If customers have further distributed the product, please retain the letter with your laboratory records, and forward to those who may have received this product. |  | 525 | 525 units | Domestic: IL, FL, NY, OH, LA, MO, MA; International: Canada, others will be inserted | Terminated | April 20, 2017 | 370 | 525
Z-2836-2016 | 75114 | No | KIF | Root canal filling resin. | Dental | Dental | 510(k) | Sealapex Xpress, Part No. 33639 |  |  | August 31, 2016 | 2016 | Ormco/Sybronendo | SybronEndo is recalling the Sealapex Xpress because it was labeled with an incorrect expiration date of 2017-11 while the syringe was labeled with the correct expiration date of 2016-11. | Employee error | The firm, Kerr, sent an "Urgent: Medical Device Recall" letter dated 8/31/16 to its customers. The letter describes the product, problem and actions to be taken. The customer were instructed to contact SybronEndo Customer Care at (800) 537-7123 to receive a RMA number if they have any of the affected product; and to complete and email back the enclosed acknowledgement form to Sealapex.Response@kerrdental.com , even if they do not have any of the affected product.If you have any questions, call 714-516-7802 or email to: mark.dzendzel@syndrondental.com. |  | 179 | 179 units | Worldwide Distribution: US (nationwide) and countries of: Canada, Australia, Mexico, Belgium, Denmark, Germany, Great Britain, and Sweden. | Open, Classified |  |  | 179
Z-2835-2016 | 74849 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | FPS 35mm x 3.5 mm non-locking screw |  |  | July 21, 2016 | 2016 | Ortho Solutions Inc | Mislabeling of a single batch of FPS 3.5 x 35 mm Locking Screw as being Non-Locking Screws. | Labeling mix-ups | Ortho Solutions sent a "Field Safety Notice" letter dated July 21, 2016 to their affected customers.The letter identify the Solutions Business Development Manager will contact the customer to arrange for the isolation and removal of the unused product from their inventory. Replacement product will be provided. The customer was advised to passed on the "Field Safety Notice" to all who need to be aware within their organization and to any organization to which their have further distributed or transferred the affected screw. Upon removal of the affected device, customers are asked to please sign and return the "Field Safety Corrective Action Declaration Form", by which to confirm they have received the "Field Safety Notice" and that they declare that the required recall has been completed at their hospital. For further questions, call (201) 258-7305. |  | 15 | 15 | US Distribution | Open, Classified |  |  | 15
Z-2839-2016 | 75121 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | G1 Dissection Tool - 3mm FLUTED BALL; For cutting and shaping bone including spine and cranium. |  |  | August 26, 2016 | 2016 | The Anspach Effort, Inc. | The firm discovered on March 4, 2016 that several product codes were inadvertently left off of Hold 1503. | Device Design | On April 26, 2016 DePuy Synthes sent a letter informing their customers of the recall. ACTION REQUIREDOur records show that your facility has received one or more of product(s) subject to the recall (correction). The AnspachEffort, Inc. asks that you:1) There is no requirement for retaining your inventory of Gl Dissection Tools.2) To ensure Gl Cutting Tool &amp; Attachment compatibility, please reference Table 1 to review which Attachments arecompatible. The Cutting TooVAttachment combination noted with a check mark in Table I will function as intended.The Cutting Tool/Attachment combinations noted with an "X" in Table 1 will not secure effectively and should not beused.3) Review, complete, sign and return the attached reply form confirming your receipt of this letter by fax or email detailsprovided on the form. Returning the form promptly will prevent you from receiving repeat notices.1. Share this letter with others in your facility that needs to be made aware of this recall (correction).2. If you distribute any of the product(s) subject to the recall (correction) to other services or facilities, please forwardthis letter as appropriate immediately.3. Keep a copy of this notice with the product(s) subject to the recall (correction).If you have any questions regarding this correction, please contact the Complaint Handling Unit Manager at 561-494-3673or contact your Anspach Sales Representative. |  | 63 | 63 | AZ, DC, CA, MD, MI, and ORSwitzerland | Open, Classified |  |  | 397
Z-2843-2016 | 74928 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5mm Bowed Locking Compression Femur Plates, 14 hole |  |  | July 27, 2016 | 2016 | OrthoPediatrics Corp | OrthoPediatrics Corp.is voluntarily recalling multiple lots of the 3.5mm Bowed Locking Compression Femur 14-hole Plate and 18-hole Plate due to the devices being mislabeled. | Labeling mix-ups | Orthopediatrics initiated the recall of 3.5mm Bowed Locking Compression Femur Plates by distributing e-mails on 0July 27, 2016, followed by telephone calls to all customers on July 28, 2016. All distributors and direct representatives were sent an e-mail to provide them with a written notification of the voluntary recall and instruct them to quarantine all identified devices from the affected lot numbers immediately. Customers were instructed to quarantine recalled product and response via a hard copy Reply Form that was included in a hard copy of the Recall Notice. Each sales representative was also asked to contact Orthopediatrics Logistics Department to receive a Return Authorization Number for returning recalled product.in order to track their affected plates. If customers have any questions about this recall, they were instructed to reach out to the firm via phone (office: 574-267-0865) or email (acargill@orthopediatrics.com). |  | 41 | 41 | Nationwide Distribution to AZ, CA, CO, DE, FL, GA, KY, MI, MN, MO, KY, RI, NY, OH, OR, RI, TN, TX, VA, WAForeign: NoneVA/DOD: None | Terminated | March 24, 2017 | 240 | 41
Z-2842-2016 | 75041 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator |  |  | August 26, 2016 | 2016 | Philips Electronics North America Corporation | The handle can separate from the MRx housing due to breakage of mounts on the rear case. | Device Design | Philips Healthcare issued customer an Urgent Medical Device Correction notification/Field Safety Notice on 11/19/14. The Urgent Medical Device Correction notification/Field Safety Notice FSN86100158A informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer, and identifies what action Philips plans to take to remedy the issue.You can continue to use your MRx prior to receiving the hardware upgrade. If one or both sides of the handle are found to have separated from the MRx housing, continue to treat your patient per existing protocols. Do not use the handle to carry the MRx. The MRx can be carried in another manner. For example, the MRx shoulder strap or bed rail hook can be used if available on the unitA hardware upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Contact your local Philips representative or call 1-800-722-9377. |  | 394 | 394 units | Worldwide distribution. US Nationwide, Canada, AUSTRALIA, FRANCE, GERMANY, INDIA, JAPAN, NETHERLANDS, SOUTH AFRICA, SPAIN, SWEDEN, THAILAND, and UNITED KINGDOM. | Open, Classified |  |  | 394
Z-2845-2016 | 74896 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | 3I OSSEOTITE CERTAIN DENTAL IMPLANTS, 4x11.5 mm, Sterile, Rx only. |  |  | August 08, 2016 | 2016 | Biomet 3i, LLC | Product was incorrectly packaged. | Packaging process control | An Urgent Medical Device Removal Notice is being sent to Customers August 12, 2016, notifying them of the packaging of OSSEOTITE Certain Implant (4x11.5mm) in T3 Implant boxes. This recall is specific to Lot No. 2016031461. The material shipped within the box was described as being correct ,but only the box it was shipped in was incorrect. Customers were notified that there are no immediate or long term health consequences. In the notification letter, customers are instructed to complete the Certificate of Acknowledgement, and to return it via email (postmarket@zimmerbiomet.com) if no product is on hand or to return the certificate with product by shipping materials to: Biomet 3i - Post-Market Returns. Replacement product will be sent to customers upon return. Customers were also advised to keep a copy of their Certificate of Acknowledgement. Further questions should be directed to 1-800 443-8166 between the hours of 8:00 AM and 6:00 PM, Monday thru Friday. For further questions, please call 1-800 342-5454. |  | 92 | 92 units | US Distributions to the states of : CA and FL., and Worldwide to : Colombia, France, Germany, Italy, and Spain. | Open, Classified |  |  | 92
Z-2849-2016 | 75080 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TriMed Semi-Tubular Plate, 10-Hole, Model No. SMTP-10 |  |  | August 29, 2016 | 2016 | Trimed Inc | TriMed is recalling the Semi-Tubular Plates with 10 holes (SMTP-10) because a non-locking screw passed through a hole in a SMTP-10 plate being implanted. | Device Design | The firm, TriMed, sent a "Medical Device Recall Notification" Letter dated 8/29/16 to customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory and quarantine the product subject to this recall; notify your customers at once of this product recall; discontinue use; return product, and complete and return the enclosed MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement and Receipt Form as soon as possible by Fax to: 1-661-254-8485 or email to: davidmedoff@trimedortho.com. Customers with questions are instructed to contact Director of Operations at (800) 633-7221, davidmedoff@trimedortho.com, M-F, 9am-5pm, Pacific Standard. |  | 30 | 30 units | Worldwide Distribution-US (nationwide) including states of: Alabama, Florida, Indiana, Connecticut, California, Wisconsin, Missouri, Florida, Hawaii; and countries of: Sweden and South Africa. | Open, Classified |  |  | 30
Z-2850-2016 | 75094 | No | MQL | Reproductive media and supplements. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sage In-Vitro Fertilization- CSC (Choline Substituted Cryopreservation) Freezing Medium . Product Number: ART -8017 |  |  | August 25, 2016 | 2016 | CooperSurgical, Inc. | Products contain international package insert and notthe correct USA package insert | Packaging change control | CooperSurgical sent an Advisory Notice dated August 25, 2016, to affected customers via UPS with confirmed delivery receipt. The letter identified the affected product and detailed the necessary actions to be taken by customers who carried the affected lots in their inventory. The advisory notification asked customers to replace the previously inserted incorrect Instructions for use (IFU) with the correct IFU sent with the advisory notification.For questions regarding this recall call 203-601-5200. |  | 1 | 1 unit | Nationwide Distribution to CA, GA, and NY | Open, Classified |  |  | 5
Z-2846-2016 | 74647 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | Broselow/Hinkle, PEDIATRIC EMERGENCY SYSTEM, PINK/RED Oxygen Delivery Module 2/PACK, Product Code 7700RAW2 |  |  | June 29, 2016 | 2016 | Carefusion 2200 Inc | The Oral Airway pouch within the PINK/RED Oxygen Delivery Module, 7700RAW, was incorrectly packaged with the 8 cm Berman Oral Airway. The pouch contained within the module is correctly labeled as 5 cm but actually contains an 8 cm Berman Oral Airway component. Therefore, the component within the pouch is incorrect.The PINK/RED Oxygen Delivery Module, 7700RAW, is intended to be used within the PINK/RED Pouch of the BROSELOW pediatric safety system.Three (3) lots (lot# 914873, lot# 914876 and lot #914878) of the PINK/RED Oxygen Delivery Module, 7700RAW, are affected by this voluntary Field Safety Corrective Action. | Under Investigation by firm | Recall notices were sent 6/29/2016. |  | 544 | 544 (2/PACKS) | Illinois | Open, Classified |  |  | 594
Z-2848-2016 | 74885 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | RENASYS Sterile Transparent Film and all the RENASYS Dressing kits that contains the RENASYS Sterile Transparent Film |  |  | July 21, 2016 | 2016 | Smith &amp; Nephew, Inc. | Sterility seal integrity failures on the Renasys Sterile Transparent Film | Unknown/Undetermined by firm | Smith &amp; Nephew sent an Urgent Medical Device Removal Notice dated July 21, 2016, to all consignees informing them of a voluntary medical device recall involving Sterile RENASYS Transparent Film and the RENASYS Dressing Kits that contain the folded RENASYS Transparent Film. Smith &amp; Nephew have identified that the process of folding the Sterile RENASYS" Transparent Film pouch during its packaging has the potential to damage the pouch seals resulting in a loss of sterility of the product. The loss of sterility has the potential to contribute to a local infection. Actions to be taken by customers and distributors:All affected Sterile RENASYS Transparent Film and the RENASYS Dressing Kits that contain the folded RENASYS Transparent Film must be returned to Smith &amp; Nephew.For questions regarding this recall call 727-642-5849.Users must complete the Acknowledgment, Receipt and Confirmation of Return Form attached with this notice and return to fieldactions@smith-nephew.com. This will trigger the return authorization process if you have product to return. An Acknowledgement, Receipt and Confirmation of Return Form must be returned even if you do not have any affected product in your inventory. Device users are reminded that reports of adverse events experienced with medical devices should be reported to FDAs MedWatch Adverse Event Reporting program. Reports may be made online, by fax or by mail. For questions, or additional information please contact Customer.Large@smith-nephew.com. |  | 45455 | 45,455 devices (in the USA) | Nationwide Distribution to AK, AL, AR, AZ, CA, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TX, UT, VA, WA, WI, Hawaii, and Puerto RicoArgentina, Australia, Bermuda, Brazil, Brunai, Canada, Curacao, Germany, Great Britain, Hong Kong, Ireland, Jordan, Malaysia, New Zeland, Peru, Saudi Arabia, Serbia, Singapore, South Africa, SriLanka, Tunisia, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 45455
Z-2869-2016 | 74989 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sertera Biopsy Device, Sterile; Used primarily in ultrasound-guided breast biopsies. |  |  | August 22, 2016 | 2016 | Hologic, Inc | Incorrect label on the Sertera Biopsy Kit | Labeling Change Control | Hologic notified customers in the United States through Certified letter on 8/22/16. The letter contains instructions of returning the product, provides contact information, and includes a postage-paid return card for the consignee to complete and return to Hologic. Contact technical support at 508-263-2900 or l-800-442:9892 Press Option 6 ). |  | 570 | 570 units | Nationwide. | Open, Classified |  |  | 570
Z-2867-2016 | 74296 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Expression Information Portal (IP5) Model 865471. The IP5 is intended to be used as a secondary display control unit for the Expression MRI Patient Monitoring System and the Expression MR200 MRI Patient Monitoring System. |  |  | September 14, 2014 | 2014 | Invivo Corporation | Frozen Display Numerics and Disabled Menu Keys after extended run time. This customer notification was sent August 22, 2014. | Software in the Use Environment | A recall notification letter was sent to hospitals who purchased the recalled product via UPS March 2014. The letter instructed users to power cycle the IP5 unit every day. This is done by pressing the power button on the side of the unit. Customers were instructed to communicate the recall to all those subject and to attach an addendum to the IP5 Instructions for Use. The firm advised customers to call their Philips representative, option 1, then option 2, if they had any questions or concerns. |  | 639 | 639 units | Worldwide distribution. US nationwide, Australia, Austria, Denmark, France, Germany, Iceland, Ireland, Italy, Netherlands, Norway, Poland, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, United Arab Emirates, and England | Completed |  |  | 639
Z-2854-2016 | 75046 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | EkoSonic Endovascular Devices |  |  | August 09, 2016 | 2016 | EKOS Corporation | The recalling firm identified ten devices which were distributed and could potentially be non-compliant. These devices potentially underwent a routine manufacturing test with non-compliant parameters, which may result in a small blister in the epoxy over transducers in the treatment zone. It is not known at the time the recall was initiated whether the non-compliant test parameters, or presence of the blisters, adversely affect operation of the device. | Storage | All affected consignees were contacted by e-mail on August 9, 2016, and asked to sequester implicated devices. |  | 10 | 10 units | Three international consignees - Germany, New Zealand and Taiwan | Completed |  |  | 10
Z-2868-2016 | 75029 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | PHS Spinal Tray 24G Gertie Marx, Catalogue No. 899379. |  |  | August 24, 2016 | 2016 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences has recently identified that one lot of Spinal Tray 24 G Gertie Marx Catalogue # 899379 was shipped containing the 24 G Gertie Marx needle without an introducer. | Employee error | Consignees notified on 8/24/16 via an Urgent - Field Corrective Action letter e-mail. The letter identified the affected device and the reason for the recall. Customers were asked to review their inventory of the Spinal Tray 24 G Gertie Marx to confirm the quantity received. The attached form should be completed and returned by e-mail or fax as indicated. Once the form is received, Customer Service will place an order for the customer to receive replacement devices. Questions should be directed to Customer Service at 1-800-241-2210. |  | 320 | 320 packages | Distributed to CA. | Open, Classified |  |  | 320
Z-2866-2016 | 75036 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | StÂżckert S5 System, S5 Heart-lung machine, Cardiopulmonary bypass heart-lung machine console |  |  | August 25, 2016 | 2016 | Sorin Group USA, Inc. | Sorin/LivaNova is initiating a field correction on the S5 Heart-lung machine because of the potential for failure of the Uninterruptible Power Supply. | Employee error | Sorin Group Deutschland GmbH created Field Correction Order (FCO) 2016-003 on August 11, 2016 to replace the affected power supply with a new unit. The customers will be contacted by the Field Service Representatives to schedule the replacement at the customer facilities. No direct action is required by the customer. |  | 5264 | 5264 Units Worldwide | Worldwide Distribution - US Nationwide in the states of AZ, CA, CO, FL, HI, KS, MA, NC, TX, UT and the countries of: Great Brittan, Peru, Russia, India, Hungary, Thailand, Germany, Azerbaijan, Austria, Italy, Slovakia, Bahrain, and Malaysia. | Terminated | May 26, 2017 | 274 | 5264
Z-2852-2016 | 75143 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.x picture archiving and communication system |  |  | August 05, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software error. Incorrect values for the volume calculation from a freehand VOI at the customer site. In volume calculations of prostate as well as in volume calculations of liver were too high. | Software design | Siemens sent an Customer Letter dated August 5, 2016, to all affected customers. The letter identified the product the problem and the action need to be taken by the customer. The letter informed customers of a software change VB10B_HF02 that they have implemented to address several issues.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions please call (610) 219-4834. |  | 338 | 338 systems | Distributed Nationwide | Open, Classified |  |  | 338
Z-2865-2016 | 75034 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Patient information in the header is only found on the first page of the report and not on the subsequent pages. | Software design | Merge sent an Urgent Medical Device Recall letter dated January 28, 2016 to affected customers via e-mail or via certified mail if they did not have an e-mail address. The letter identified the affected product, problem and actions to be taken. For questions call Merge Customer Service at (877)741-5369 or support@merge.com. |  | 42 | 42 sites potentially have the affected version | US Nationwide Distribution | Open, Classified |  |  | 42
Z-2857-2016 | 75051 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | TECNIS Symfony Extended Range of Vision IOL, Model No. ZXR00 |  |  | August 26, 2016 | 2016 | Abbott Medical Optics Inc. (AMO) | Abbott Medical Optics is recalling Intraocular Lenses (IOLs) because they were not being checked prior to release and this may result in the release of mislabeled IOLs. | Process control | Abbott Medical Optics, Inc. sent an Urgent Field Safety Notice to all affected customers to inform them that Abbott Medical Optics (AMO) is recalling 737 IOLs due to detection of an inspection equipment malfunction that resulted in a total of 737 distributed lenses not being fully checked prior to release. The letter informs customers that this malfunction may result in the release of mislabeled IOLs. The letter informs the customers of the actions to be taken. Customers are instructed to complete Customer Reply Form and fax it to AMO Quality Assurance at 714-247-4510 or email to RegCompliOne@abbott.com. Customers with any questions related to the IOL return process are instructed to contact an AMO Customer Service Representative at 877-266-4543.Customers with product complaints or adverse events are instructed to inform AMO by calling 877-266-4543. |  | 10 | 10 units | Worldwide Distribution - US nationwide, PR, and Internationally to Canada, Latin America, Europe, Africa, Middle East, and Asia Pacific. | Terminated | December 23, 2016 | 119 | 737
Z-2873-2016 | 64230 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeCare PCA Infusion System with Hospira MedNet Software, PCA Serial List Number 20709 |  |  | March 08, 2013 | 2013 | Hospira Inc. | Screw rotation error occurring when the pump detects a mismatch between the number of steps given to the motor driving delivery and the number of steps the motor actually took. When this malfunction occurs, the pump will alarm and infusion stops. | Nonconforming Material/Component | Hospira will contract with Stericycle, Inc. in Indianapolis, IN to send letter and a reply form to Hospira, Inc. Receipt of the letters will be confirmed by UPS tracking. Customers will be requested to return a reply form to Stericycle indicating the number of pumps on hand and that they have received the letter and notified their users.In addition, letters will be sent by email to ECRI, NRAC, Noblis and ISMP for notification of their members and posting on their websites. The letter will also be posted to the Hospira website. Letters will be sent to Hospira customers of record who purchased the devices directly from Hospira. Customers will be directed to notify their consignees if they further distributed the product and confirm they have done so.Telephone attempts will be made to contact customers that do not respond in order to confirm receipt of the letter. Final effectiveness will be done as the devices are upgraded with new plastic motors. |  | 8794 | 8,794 | United States (Nationwide) Distribution including Puerto Rico and American Virgin Islands. | Open, Classified |  |  | 8863
Z-2871-2016 | 75158 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CADstream software. |  |  | November 01, 2012 | 2012 | Merge Healthcare, Inc. | An incorrect biopsy or missed target could result if the incorrect grid is selected within the application. | Error in labeling | Merge sent Urgent Medical Device Recall dated January 28, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 1349 | 1,349 sites potentially have the affected versions | US Nationwide Distribution | Open, Classified |  |  | 1349
Z-2875-2016 | 74954 | No | EOQ | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Pentax Video Bronchoscope |  |  | May 31, 2016 | 2016 | Pentax of America Inc | Pentax America Inc. is recalling various bronchoscopes which contain an incorrect Operation and Reprocessing IFUS. | Packaging process control | Pentax Medical sent a U.S. Urgent Field Correction notice to all affected customers with a recall letter/return response form on May 31, 2016 via Fed Ex.Customers with questions were instructed to call 800-431-5880, ext. 2064. |  | 1719 | 1719 units | Nationwide Distribution | Open, Classified |  |  | 1719
Z-2870-2016 | 74952 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | MED 5223 - VacuFlow Safe, Multi-Sample Blood Collection Set With Luer Adapter Tube 12", Sterile, Rx only, Distributed by Tanner Scientific, Sarasota, FGL 34243  Made in China. |  |  | August 09, 2016 | 2016 | Global Healthcare Inc | The multi-sample needle retractable sheath may not fully retract over the needle during the changes of the blood collection tubes. | Under Investigation by firm | The distributors were first notified on August 9, 2016 via phone to place product on hold. They were notified of the recall on August 10, 20l6 by email. |  | 282 | 282 cartons (1000 units per carton) | FL | Open, Classified |  |  | 282
Z-2872-2016 | 75194 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology, versions 12.1.1 and 12.2 |  |  | July 12, 2016 | 2016 | Mckesson Medical Imaging Group | The STAT priority icon was not displayed in the Study List when the study was opened. | Software design | McKesson sent an Urgent Field Safety Notice on July 12, 2016 to customers with potentially affected software versions. The letter requested that they contact McKesson to acknowledge that they read and understood the Field Safety Notice and to arrange installation of software update/upgrade. Customers with questions were instructed call the McKesson Support department and ask to speak to their Support Manager.For questions regarding this recall call 404-338-3556. |  | 164 | 164 devices | Worldwide Distribution - US including AL, AK, AZ, AR, CA, CO, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OK, PA, RI, SC, SD, TN, TX, VA, WA WV, WI, WY, and Internationally to Canada, France, Great Britain and Ireland | Open, Classified |  |  | 164
Z-2878-2016 | 75113 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Syngo RT Therapist, Accelerator, Linear, Medical |  |  | July 26, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software patch installation to address several safety issues. | Software design | Siemens sent a Field Safety Notice to all affected customers informing them of a new software patch that they plan to install to address software issues. Customers were advised to notify anyone at their organization who should be aware of the information contained in the Notice. Customers were also advised to include the Field Safety Notice in their System Owner Manual, chapter "Safety Advisory Letters" where it should remain. For questions regarding this recall call 610-219-4834. |  | 103 | 103 accelerators | Nationwide distribution | Open, Classified |  |  | 103
Z-2881-2016 | 74805 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | TRUEread Blood Glucose Test Strips (Glucose Oxidase) are used with TRUEresult and TRUE2go meters. TRUEread Blood Glucose Test Strips are available in vials with 10 ct., 25 ct., 50 ct., and sold as 10 ct. with meters. |  |  | June 28, 2016 | 2016 | Trividia Health | Product gives incorrect low blood glucose levels. | Nonconforming Material/Component | The firm, Trividia, issued a press release on 06/28/2016 that instructed users to locate the test strip lot number on product on hand in order to determine if they possess affected product. Customers were instructed to contact Stericycle at 1-888-835-2723 or check online at www.trividiahealth.com/product-notice to see if their inventory was affected by the recall. Any strips under recall will be returned for disposal and replaced with no charge. In addition, the firm sent an "Urgent Medical Device Voluntary Recall" notice to customers addressing the recall and steps to follow. The customers were instructed to identify affected product using the same method: contacting Stericycle at 1-888-835-2723 or checking www.trividiahealth.com/product-notice to see if the product in possession was from affected lots; to notify patients that they may continue use of TRUEread, TRUEtrack, TRUEbalance, and TRUEtest strips not under recall, and to use them with the appropriate meter. Pharmacists/DME providers were instructed to call Stericycle if health care professionals contacted them regarding the recall. Distributors were requested to check their meter and strip inventories to see if they had any affected lots. Distributors were instructed to notify their own customers of the recall and collect product that is affected by the recall. Once all affected product is received by customers, distributors need to call Stericycle to obtain prepaid return labels. All returned product will be replaced at no charge.If you have any questions, contact Vice President of Marketing, at (800) 342-7226 ext. 2183 or email: LNardi@trividiahealth.com. |  | 22416 | 5,527,921 In total =4,621,639 strips &amp; 393,808 kits domestically; 500,058 strips &amp; 22,416 kits internationally | Worldwide Distribution-US (nationwide) including states of: AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV; and Puerto Rico and countries of: Austria, Bangladesh, Cambodia, Canada, Pakistan, Sri Lanka, Thailand, Turkey, United Kingdom, and Vietnam. | Open, Classified |  |  | 22416
Z-2879-2016 | 74508 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Syringe Module Model 8110 with software version 9.15 |  |  | August 08, 2016 | 2016 | CareFusion 303, Inc. | A software anomaly with the Alaris Syringe module software version 9.15 may cause an infusion to unexpectedly stop when the Syringe module is transitioning from one rate to another. | Software design | The firm, BD, sent an "Urgent: Medical Device Recall Notification" letter dated 8/8/16 to customers to inform them that BD has identified a software anomaly associated with the Alaris Syringe module software version 9.15. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the instructions and promptly complete and return the Customer Response Card to expedite the corrective action process.Customers with recall related questions are instructed to contact the BD Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact customer advocacy at (888) 812-3266. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3229. |  | 12000 | 12,000 units | Worldwide Distribution: US(nationwide) and countries of: Canada, Australia and United Arab Emirates. | Open, Classified |  |  | 12000
Z-2884-2016 | 74861 | No | DKJ | Salicylate test system. | Toxicology | Toxicology | 510(k) | DRI Salicylate Serum Tox Assay, Catalog number 0977. |  |  | July 25, 2016 | 2016 | Microgenics Corporation | Some lots of DRI Salicylate Serum Tox Assay negative patient samples are recovering higher than the product sensitivity claim of 4.4 mg/dL. | Under Investigation by firm | The firm, Thermo Fisher Scientific, sent an "URGENT MEDICAL DEVICE FIELD ACTION" letter dated July 25, 2016 by regular mail to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to determine if they are using or have inventory of any of the affected lots; discontinue use and destroy any remaining inventory of affected lots per local wastes ordinances; retain a copy of the letter for laboratory records; if you have forwarded kits of the affected lots, provide a copy of letter to them; and complete and return attached Medical Device Field Action Response Form within 5 days to Thermo Fisher Scientific Technical Service via fax at 1-888-527-8001 (USA toll free), 510-979-5420 (outside of USA) or email to techservice.mgc@thermofisher.com.If you have any questions contact Technical Services at 1-800-232-3342 (USA only) or 510-969-5000 (outside the USA) press option 2 then option 3. |  | 1708 | 1708 kits | Worldwide Distribution: US (nationwide) to states of: AZ, CA, MO, TX, and PR and countries of: Canada, Germany, Hong Kong, Italy, Malaysia, and Taiwan. | Terminated | November 21, 2016 | 119 | 1708
Z-2887-2016 | 75016 | No | JNM | Transferrin immunological test system. | Immunology | Immunology | 510(k) | R&D; Systems, QuantikineÂż IVDÂżsTfR Immunoassay Kit, for the quantitative determination of soluble transferrin receptor (sTfR) concentrations in human serum and plasma For in Vitro Diagnostic Use, REF DTFR1. |  |  | August 24, 2016 | 2016 | R &amp; D Systems, Inc. | R&D; Systems, Inc. received two customer complaints that results of the Controls provided with the QuantikineÂż IVDÂż Human sTfR Immunoassay kit were out of the range high.Internal testing within R&D; Systems confirmed a problem with Control results being out of range high. | Nonconforming Material/Component | Consignees were sent via US mail on 8/24/2016 a R&D; systems "Urgent Medical Device Recall" letter dated August 23, 2016. The letter described the problem and the product involved in the recall. Advised consignees to stop using the kits and to discard all unused kits and complete and return the enclosed Acknowledgement and Receipt Form. For questions contact Bio-Technes Technical Service Department at 1-612-379-2956 or techserv@bio-techne.com. |  | 725 | 725 kits (502 US, 223 OUS) | US: CA, CO, KY, LA, MD, MA, NJ, OH, RI, UT, VA, WA.OUS: AUSTRALIA, China, INDIA, INDONESIA, ISRAEL, JAPAN, MEXICO, SOUTH KOREA, UNITED KINGDOM. | Terminated | July 26, 2017 | 336 | 725
Z-2886-2016 | 75078 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare Xhibit Telemetry Receiver (XTR), Model 96280. |  |  | August 25, 2016 | 2016 | Spacelabs Healthcare Inc | The firm received multiple reports of Xhibit Telemetry System going offline or locking up. In addition, short-duration asystole alarms may not display an audio and visual alarm as reported from a single Canadian hospital. | Software design | Spacelabs sent a letter dated August 25, 2016 to affected US consignees informing them of the going offline or locking up issue when Xhibit Central Station is used as part of an Xhibit Telemetry System. On September 1, 2016, a customer letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record.Spacelabs sent the second Urgent - Medical Device Correction letter, dated September 15, 2016, to affected US consignees informing them that they have received reports of certain short-duration asystole alarms generated by the Xhibit Telemetry Receiver (XTR) where there may be no audible or visual alarm. On September 22, 2016, a second letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record.Spacelabs Healthcare will contact consignees to schedule a convenient time for Spacelabs to update the recalled Xhibit Telemetry System at no cost. Customers with Xhibit Telemetry Receivers will get all Xhibit Central Stations updated regardless if connected to an Xhibit Telemetry Receiver.Consignees are directed to technical support for additional information or technical assistance:Technical SupportSpacelabs Healthcare, Inc35301 SE Center St.Snoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 316 | 316 XTR units total (248 units in the US and 68 units outside US) | Worldwide distribution. US nationwide including Puerto Rico, Australia, Canada, China, France, Germany, Ireland, Italy, Netherlands, Philippines, Poland, Romania, Switzerland, and United Kingdom. | Terminated | March 13, 2017 | 200 | 316
Z-2888-2016 | 75160 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Spiegelberg Brain Pressure Monitoring System |  |  | December 29, 2009 | 2009 | Spiegelberg Gmbh &amp; Co. KG | In the vast majority of products specified, a shrinkage of the air chamber during sterilization has occurred. Because of these false high readings were displayed in the lower measuring range from 0 to 20 mmHg.The observed measurement error was particularly at the lowest end of the range (to 0 mmHg) up to 8 mmHg. | Packaging | Spiegelberg Gmbh &amp; Co. KG sent an Urgent Medical Device Recall letter dated December 29, 2009, to all affected customers. Customers were instructed to quarantine and return stored products.Furthermore, they were asked to inform all their customers which have received the affected products and recall the products accordingly. Furthermore, they were asked to tell us about the actions taken by them. |  | 22 | 22 units | Nationwide Distribution to PA only | Completed |  |  | 22
Z-2897-2016 | 75082 | No | EIH | Porcelain powder for clinical use. | Dental | Dental | 510(k) | BruxZir Shaded, Model No. 70-1167-NOW0001, 70-1167-NOW0002, 70-1167-NOW0003, 70-1167-NOW0005, 70-1167-NOW0006 |  |  | August 15, 2016 | 2016 | Prismatik Dentalcraft, Inc | Prismatik Dentalcraft is recalling BruxZir Shaded because the new design was manufactured at risk and were to be quarantined and held until the validation was completed. Prismatik failed to properly identify the lots as quarantined product. | Release of Material/Component prior to receiving test results | Prismatik DentalCraft, Inc. sent an Urgent Medical Device Recall letter dated August 15, 2016, to all their customers to inform them that Prismatik Dentalcraft is recalling BruxZir NOW product because the new design was manufactured at risk and was supposed to be quarantined and held until validation was completed. Prismatik failed to properly identify the implicated lots as quarantined product. Customers are informed that the BruxZir Now with the thicker sprue design was packaged and shipped to customers in error. Customers are informed of that there is no risk to the patient and of the actions to be taken for the customers and from Prismatic Dentalcraft Inc. For questions regarding this recall call 949-440-2600. |  | 150 | 150 units | Nationwide distribution to SC, MN, NC, IN, AZ, OR, WA, CA, WY | Open, Classified |  |  | 150
Z-2890-2016 | 75100 | No | FOX | Infusion stand. | General Hospital | General Hospital | 510(K) Exempt | The EKOS cart has the catalog number 700-60101. |  |  | August 22, 2016 | 2016 | EKOS Corporation | Casters on the EKOS carts may be missing washers which can cause premature wear and dislodgement of the casters/wheels leading to failure. | Device Design | The firm, EKOS, sent an "URGENT FIELD SAFETY ACTION" letter dated August 22, 2016 to its customers. The letters were sent to customers informing them of the field safety action, how to identify the affected carts at their facility and instructing customers locate the cart, quarantine it and do not use the cart for any procedures. A representative from EKOS Corporation will be contacting customers in the near future to arrange a date and time when the cart at a facility can be inspected and any affected caster(s) replaced. As the cart is inspected, it will be tagged to easily identify it and confirm it has been inspected and repaired as necessary. Customers will be provided with a copy of the inspection/repair form for the cart they inspect. Please retain that form for records.If customers have any questions, please contact EKOS Customer Service at 1-888-400-3567 or your local EKOS representative. |  | 455 | 455 units | US Nationwide Distribution in states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, and WY. | Terminated | December 30, 2016 | 130 | 455
Z-2895-2016 | 74892 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Zyno Medical Administration Set, Sterile. RX |  |  | July 29, 2016 | 2016 | Zyno Medical LLC | Administration Set potential filter leaking | Process control | Zyno issued Recall letter dated July 26th, 2016 to affected customers. Customers who received affected products were contacted between 7/29/2016 and 8/1/2016. Customers were asked to return unused product for replacement.For questions contact Customer Service Representative, at 508-907-7796, or laryloring@zynomed.com.Expanded Recall: Two recall letters issued on 11/3/2016 to address the newly identified lots. One letter for customers who received the initial recall letter and the other letter for newly affected customers who did not receive the initial recall letter. |  | 78063 | 78063 sets | US Nationwide Distribution | Terminated | July 11, 2017 | 347 | 318202
Z-2783-2016 | 74835 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Light Adaptor for Small Battery Drive and Small Battery Drive II. |  |  | June 24, 2016 | 2016 | The Anspach Effort, Inc. | There was a potential for the Adaptor and Light Adaptor for Small Battery Drive and Small Battery Drive II to generate excessive internal pressure that may cause the products to burst. | Device Design | An Urgent Notice for this Recall was provided to the consignees on 6/24/2016. They were to review their inventory and return the products with instructions. IF YOU HAVE ANY PRODUCT: To immediately quarantine product so it's not used; ensure all in their facility read the Urgent Notice for the Recall; Call the DePuy Synthes Customer Support to obtain a Return Materials Authorization (RMA) Number; to complete the Verification Section indicating the product subject to recall was located; to return the Verification Section with the product to Product Returns, The Anspach Effort Inc., 4500 Riverside Drive, Palm Beach Gardens, FL 33410; Send a copy of the completed Verification Section to DePuy Synthes, Customer Quality Dept by: FAX (561) 627-2682 or Scan/email: DPYUS-PowerToolsFieldActions@its.jnj.com;Keep the notice visibly posted for awareness until all products have been returned; maintain a copy of the notice subject to this action and keep a copy for your records. The return documentation acknowledges your receipt of the medical device removal information.IF YOU DO NOT HAVE the identified product: complete the Verification Section of the recall, by checking the box indicating no affected product located; Include your name title, address, telephone #, signature and date; Send a copy of the completed Verification Section to DePuy Synthes Customer Quality Dept by: Fax: (561) 627-2682, or Scan/e-mail: DPYUS-PowerToolsFieldActions@its.jnj.com; This return document acknowledges your receipt of medical device removal information;ADVISORY: For Alternative products: This product does not have any inventory for replacement. Since the recall of this device may have an impact on your surgical schedule, DePuy Synthes Power Tools would like to propose alternative products to assist you with a replacement for the Recalled Adaptor (05.100.024) or Light Adaptor (05.001.108) . The products are the Small Battery Drive II Hand-piece (532.110) can also be operated |  | 1833 | 1,833 devices (05.001.024 and 05.001.108) | Worldwide Distribution - US Nationwide and including the countries of Canada, Austria, Australia, Belgium, Brazil, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Libya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, Slovakia, Spain, South Africa, South Korea, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Vietnam, and areas of Middle East, Africa, Asia &amp; Pacific. | Open, Classified |  |  | 1833
Z-2889-2016 | 74901 | No | KDC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ETHICON MERSILENE" Tape |  |  | August 01, 2016 | 2016 | Ethicon, Inc. | ETHICON MERSILENE" Tape (also referred to as MERSILENETM Polyester Fiber Strip or MERSILENETM Fiber Ligature) products are being recalled since a group of MERSILENE" Tape product codes are being supplied with an incorrect Instructions for Use (IFU) insert. | Other | Ethicon Inc. sent an Urgent - Medical Device Correction Notification dated August 1, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ACTION REQUIRED:1. Examine your inventory immediately to determine if you have product subject to this correction on hand and locate such product(s). Refer to Attachment 2 for the Product Identification Tool to identify the products that are subject to this correction by using package labels.2. Please read Attachment 1 included within this letter thoroughly - MERSILENE" Tape Instructions for Use (IFU) .3. Ensure all users are aware of the issue by distributing this notice to all users of the Ethicon MERSILENE" Tape devices in your facility.4. Replace the existing incorrect MERSILENE" Suture Instructions for Use (IFU) with the MERSILENE"Tape Instructions for Use (IFU) provided in Attachment 1 and dispose of the incorrect MERSILENE" Suture Instructions for Use (IFU) per your internal procedures.5. Complete the Business Reply Form (BRF) (Attachment 3) confirming receipt of this notice and fax or email it to Stericycle at 1- (866) 723-2836 or Ethicon8698@stericycle.com within three (3) business days. Please return the BRF even if you do not have product subject to this correction.6. Keep this notice visibly posted with product subject to this correction in your facility for awareness.For further questions, please call (908) 218-0707. |  | 396024 | 396,024 eaches | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico. | Open, Classified |  |  | 396024
Z-2892-2016 | 75171 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza VB10A, Picture Archiving and Communication System |  |  | September 06, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software upgrade to eliminate several issues | Software design | Siemens notified customers via mail on 09/07/2016 informing them of the latest syngo.plaza software change that was available to address several issues. |  | 47 | 47 systems | Distributed Nationwide | Terminated | May 08, 2017 | 244 | 47
Z-2898-2016 | 74955 | No | PHV |  | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CareLink iPro Version 1.10, Catalog No. MMT-7340 |  |  | August 05, 2016 | 2016 | Medtronic Inc. | Medtronic MiniMed is recalling the CareLink iPro Therapy Management Software due to a time stamp error. | Software design | Medtronic sent an Urgent Field Corrective Action letter dated August 5, 2016, to affected customers who received the CareLink iPro Therapy Management Software. The letter informed the customers that the recall affects the users of CareLink iPro software who have downloaded reports that include data that was uploaded from certain models of Roche BG meters. Customers were informed that the recall is due to time stamp error and data from certain Roche meters uploaded to CareLink iPro may resulted in incorrect reports. Customers were informed that if they manually enter Logbook data from the Roche ACCU-CHEK meters or upload data from a blood glucose meter other than the Roche meters are not affected by the recall situation. Customers with any questions regarding the recall letter are instructed to contact the Medtronic 24-Hour HelpLine at 1.800.646.4633.For questions regarding this recall call 818-576-4700. |  | 33 | 33 units | Nationwide Distribution to TN, MN, and WA | Open, Classified |  |  | 33
Z-2899-2016 | 74951 | No | DDR | Myoglobin immunological test system. | Immunology | Immunology | 510(k) | Siemens ADVIA CentaurÂż Calibrator U, |  |  | August 08, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | Failure of analytical sensitivity with Calibrator U kit lots ending in 63 and 64 and a drift from the internal standardization | Nonconforming Material/Component | Siemens Healthcare Diagnostics issued an Urgent Medical Device Correction (UMDC) to customers in the US on August 8, 2016, to be delivered to customers on August 9, 2016. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on August 8, 2016. The communications notify customers of the failure of analytical sensitivity with Calibrator U kit lots ending in 63 and 64 and a drift from the internal standardization. The communications also inform customers that analytical sensitivity and alignment to the internal standardization will be restored with the release of Calibrator U kit lots ending in 65. Customers may continue to use Calibrator U lots CU63 and CU64 with any indate reagent kit lot, current QC ranges and current Master Curve Material lot 45719. Revised QC targets and ranges are located on the Bio-Rad website at QCnet.com for use with CU65 and higher. For further questions, please call (508) 668-5000. |  | 2248 | 2,248 kits | Worldwide Distribution - Nationwide Distribution including the countries of:Argentina Australia Brazil Canada Chile China Ecuador India Japan Korea Mexico Malaysia Peru Philippines Saudi Arabia Singapore Taiwan Vietnam South Africa | Open, Classified |  |  | 2248
Z-2900-2016 | 75111 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | AQUIOS CL Flow Cytometer, Catalog No. B30166; Intended for use with in vitro diagnostic flow cytometry applications using up to four florescent detection channels. |  |  | September 06, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter is recalling the AQUIOS CL System because the use of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe. | Software Design Change | An Urgent Medical Device Recall letter dated 9/6/16 was sent to customers to inform them that the use of Positions 9 and 10 of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe. The letter informs the customers of the impact of the recall, action to be taken, and the resolution by Beckman Coulter. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding this notice, are instructed to contact Customer Support Center: http://www.beckmancoulter.com/customersupport/support; call 1-800-369-0333 in the US and Canada; outside the US and Canada, contact their local Beckman Coulter Representative. |  | 148 | 148 units total (12 units in US) | US: IL, PA, NY, GA, WA, FL, SC.Affected geographies: Angola, Australia, Belgium, Botswana, Canada, Chile, Columbia, Czech Republic, Denmark, Finland, France, Germany, Ghana, Israel, Italy, Malaysia, Nambia, Netherlands, Panama, Poland, Romania, Slovakia, South Africa, Spain, Sweden, Switzerland, Uganda, United Kingdom and Zambia | Terminated | August 07, 2017 | 335 | 148
Z-2917-2016 | 75019 | No | OEA | Surgical apparel. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | BLUE IMPERVIOUS GOWN WITH THUMBHOOKS, X-LARGE. Non-Sterile, Single Use Only, Plastic Non-Surgical Isolation Gown Non-Sterile, Single Use Only, Plastic Non-Surgical Isolation Gown. Serves as physical barrier to the transfer of microorganisms, body fluids and particulate material. Level of barrier protection is nonspecific. |  |  | July 20, 2016 | 2016 | Cardinal Health | Blue Impervious Gown with thumbhooks, X-large. All lots do not meet design specifications, specifically 16 CFR 1611 - Standard for the Flammability of Vinyl Plastic Film. The standard defines a specific test method that quantifies the rate of flame propagation against a minimum value using an engineered test fixture. The method requires that the test be performed on different orientations of the film (machine direction and cross direction). The machine direction is defined by the manufacturing process generating the film. The formed film exits the machine moving in the machine direction. The cross direction is at 90 degrees to the machine direction. The more rapid burning direction defines the results of the test (either a pass or fail). The products being recalled failed when tested in the machine direction with a result of 1.9-2.0 inches/second versus a limit of 1.2 inches/second. They passed by self-extinguishing when tested in the cross direction. | Process control | The firm sent an URGENT: PRODUCT RECALL dated 7/20/2016 describing the reason for the recall. The letter also provides instructions for the user on how to address this issue.Cardinal Health anticipates replacement inventory of 5213PG will be available by August 22, 2016 and V5213PG by August 29. Until that time, the recommended substitute product is 5210PG, which shares all the design characteristics of 5213PG (blue polyethylene film material, over-the-head donning/tear-away removal, half-back, thumb hook wrists) and comes in similar packaging (15 individually packed gowns per box, 5 boxes per case for a total of 75 gowns per case). The firm advises that the 5210PG is "Universal" size, which is smaller than 5213PG (XL size), so it may not work for all users.The firm is requesting all customers to take the following actions:1. Immediately stop distributing affected product.2. Return the enclosed acknowledgment form via fax to 847.689.9101, whether ornot they have affected product, as Cardinal Health is required to confirm receipt ofthis notification.3. Contact the Distributor Management Customer Service team at 800.635.6021 toarrange for credit and return of any affected product they may have.4. Customers do not need to notify any customers to whom they may have distributed product affected by this recall.If they have any questions relating to this recall, please contact Cardinal Health Customer Advocacy at 800.292.9332. |  | 4062511 | 4,062,511 eaches | Nationwide to U.S. consignees located in the following states:AL ,AR ,AZ ,CA ,CT ,FL ,GA ,IL ,IN ,KY ,LA ,MA ,MD ,MI ,MO ,NC ,NJ ,NM ,NV ,NY ,OH ,PA ,SC ,TX ,UT ,VA ,WA and WV. | Open, Classified |  |  | 4062511
Z-2906-2016 | 75065 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | CUSA Dissectron Fingertip Laparoscopic Cannula, REF F75221, Sterile EO, Rx only, packaged in a Tyvek pouch and secondary packaged in Tyvek pouch, Five (5) sealed pouches are packaged in a cardboard box. |  |  | August 23, 2016 | 2016 | Integra LifeSciences Corp. | During final sampling inspection, one of the lots was discovered with a total incomplete packaging seal of the sterile product. | Process control | Integra sent an "Urgent Voluntary Medical Device Recall" Letter and Recall Acknowledgement and Return Form dated August 23, 2016 to their affected customer. The letter identified the product, the problem, and the action to be taken by the customer. The customer was advised if they have distributed the product to identify those customers and forward a copy of the enclosed Integra Medical Device Voluntary Recall Notification to those customers. The customer was also asked to review their inventory for the affected product and complete the attached form indicating whether or not they have any of the affected product in their inventory. Customers with questions were instructed to contact Customer Service at 1-800-654-2873. |  | 9 | 9 | Nationwide Distribution to KS only | Open, Classified |  |  | 53
Z-2914-2016 | 75048 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CT3Plus (a.k.a. CT3PZ); |  |  | September 19, 2016 | 2016 | New Star Lasers, Inc. | FDA inspection found that customers who had received devices with promotional material with outside use indications were not notified when labeling was revised to remove these indications. | Labeling False and Misleading | Letters sent on September 19, 2016 via US mail advised customers of the unapproved claims that may be in their literature and advised of the approved uses for each device. |  | 97 | 44 systems CT3Plus; 47 systems CT3PZ | US distribution only. | Open, Classified |  |  | 119
Z-2911-2016 | 75014 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Walrus Arterial Embolectomy Catheter |  |  | July 07, 2016 | 2016 | Arrow International Inc | Arrow is recalling due to incorrect labeling of products. | Component design/selection | The firm, Arrow International, Inc., sent an "Urgent Medical Device Recall Notification" letter dated July 07, 2016 via FedEx to its customers. The letter described the product, problems and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any products they may have on hand; return product, and complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com, even if you have no affected stock.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 2425 | 2,425 catheters | US Nationwide Distribution | Open, Classified |  |  | 2425
Z-2918-2016 | 75050 | No | HTN | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Arthrex Suture Washer |  |  | August 12, 2016 | 2016 | Arthrex, Inc. | The device was assembled incorrectly. Suture and washer were not assembled; components were placed loose in the package. | Packaging process control | Consignees were notified 8/12/2016 of the potentially affected devices; Product/Batch Numbers: AR-7000-818T Batch #10028617. For Healthcare Professionals: immediately discontinue use of all affected products with stated numbers in their possession. Call (855) 552-5056 or email returns@arthrex.com with the product number, lot, and quantities so they may ship replacement products and expedite return of recalled product. To please complete the "Recall Acknowledgement of Receipt: postal card and return to Arthrex, Inc.Distributing Agencies: Immediately identify and hold all stated products in your inventory. Call 888-519-7255 ext. 72067 or e-mail AIM@arthrex.com Attn: Christi Zimmer. Have product number, lot number and quantities you may have so replacement product may be shipped and expedite the return of the recalled product. Please complete the "Recall Acknowledgement of Receipt" postal card and return to Arthrex. Inc. Any questions about the return of products to call 866-267-9138 and ask for Ellen Barry. Also by email to complaints@arthrez.com. |  | 45 | 45 devices | OR, NY, MN, KY, CA, ID, CO, MA, and FL.Thailand. | Open, Classified |  |  | 45
Z-2833-2016 | 74838 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TriFoillmaging Triumph/Triumph II Research CT scanners. |  |  | May 24, 2016 | 2016 | Northridge Tri-Modality Imaging, | The Firm has determined that several Triumph/Triumph II Research CT scanners produced by Northridge Trimodality Imaging, inc dba TriFoillmaging (or by Gamma Medica-Ideas, Inc and field upgraded with a replacement x-ray tube) may be out of compliance with 21 CFR 1020.40 (Cabinet X-ray standards) | Radiation Control for Health and Safety Act | TriFoil lmaging planned action to bring defect into Compliance:1. All repairs were performed at customer facilities. Customers were contacted by phone to arrange a convenient repair time.2. TriFoil service engineers installed a collimator on each of the affected cabinet x-ray systems to bring their emission rates into compliance with the emission limit in the performance standard.3. Full radiation surveys were performed, at maximum exposure settings, to verify each repair was effective.4. All costs associated with the repair and follow-up testing were paid directly by TriFoil Imaging.CDRH approves the CAP and understands that the CAP was completed onAugust 29, 2016.For further questions please call (818) 709-2468. |  | 3 | 3 | USA Distribution | Open, Classified |  |  | 3
Z-2913-2016 | 75108 | No | JFS | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens RAPIDPoint 405 Measurement Cartridge |  |  | September 02, 2016 | 2016 | Siemens Healthcare Diagnostics Inc | Error in the value assignment for the Chloride calibrators in the RAPIDPoint 400 and RAPIDPoint 405 Measurement Cartridge | Under Investigation by firm | Siemens Healthcare issued An Urgent Field Safety Notice via e-mail to affected Siemens Healthcare Diagnostics Regional offices on September 2, 2016 both in the United States and outside of the United States, for communication with affected customers e: the bias of the chloride values. US accounts sent hard copy via Federal Express and requested to:Complete and return the Field Correction Effectiveness Check form attached to this letter within 7 days. Review this letter with your Medical Director.Questions:Contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 6638 | 6638 cartridges | Nationwide:Foreign:Canada, Bahamas,Germany, Japan, Mexico.Belgium, Italy,Spain, Korea, France, Austria, Korea, France, Poland, Australia, china, AE, Bangladesh, Columbia, Czechkoslovakia, Bosnia | Terminated | August 07, 2017 | 339 | 7327
Z-0008-2017 | 74982 | No | OUF | Sirolimus test system. | Toxicology | Toxicology | 510(k) | Thermo Scientific QMS Everolimus Assay; |  |  | August 15, 2016 | 2016 | Microgenics Corporation | A reagent stability issue (still under investigation) causes less separation in the calibration curve, which affects results across the entire range. For controls, maximum bias of +/- 35% for low levels, +/- 33% for mid-range levels, and +/- 47% for high levels is observed. Bias in patient samples ranges from -15.7% to +5.2% with an average bias of -5.6% when comparing the QMS everolimus assay to an everolimus LC-MS reference method. Patient samples across the measuring range  sub-therapeutic, therapeutic, and supra-therapeutic  could be affected. | Under Investigation by firm | Thermo Fisher sent an Urgent Medical Device Field Action letter dated on August 15, 2016 to customers via regular mail. The letter identified the affected product, problem and actions to be taken. Customers are advised to Discontinue use and destroy any remaining inventory of the affected lots per your local waste ordinances. Customers were instructed to cmplete the attached Medical Device Field Action Response Form and return theform to Thermo Fisher Scientific Technical Service. For questions contact Technical Service at 1-800-232-3342 (USA only) or 510-979-5000 (outside the USA) then press option 2, then option 3. |  | 1663 | 163 kits in US, 669 kits worldwide - toatal 831. | Worldwide Distribution - US Nationwide in the states of KY, CA, NY, DC, NC, TX, IL, OH, IL, FL, IK, IA, MS and countries of Mexico, Taiwan, Guatemala, Malaysia, Argentina, Dominican Republic, Philippines, Germany, Australia. | Terminated | January 04, 2017 | 142 | 1663
Z-0003-2017 | 75235 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard 360 Revision Knee System |  |  | August 08, 2016 | 2016 | Zimmer Biomet, Inc. | The drill bit used to drill the hex pre-drilled hole in the screw head of the Tibial Augment Bolts included in the Vanguard 360 Revision Knee System was larger than allowable per the process specification. As a result, the hex geometry can still be visible, but there are radial cuts into the hex flat. Therefore, when the bolt is inserted into the augment, the screwdriver is unable to catch on the hex geometry, or the screwdriver strips the remaining hex geometry, and the bolt is unable to be fully tightened.The bolt stripping or the bolt not being fully tightened will be noticed by the surgeon intraoperatively when he or she is unable to screw the augment bolts into the augment. This possible event can result in an intraoperative delay resulting in increased time under anesthesia. | Process control | Zimmer Biomet sent an URGENT MEDICAL DEVICE RECALL REMOVAL notifications dated August 8, 2016, to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notice you have further questions or concerns please call 1-800-348-9500 ext. 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of business hours will receive a prompt to record a voicemail and will be returned upon receipt. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. For further questions, please call (800) 348-2759. |  | 454 | 454 | Worldwide Distribution - US Distribution to the states of : VA, UT, TX. OR, MN, AL, PA, CA and KY., and to the countries of : AUSTRALIA, China, Italy, Netherlands and Japan. | Terminated | June 05, 2017 | 301 | 454
Z-0002-2017 | 75127 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Merit Inflation Syringe Kit. Catalog Number KOS-02857 |  |  | August 25, 2016 | 2016 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. announces a voluntary field action for Merit Inflation Syring Kit due to unit labels missing required device labeling information. | Labeling design | Merit Medical Inc. notified all affected customers on August 25, 2016, via verbal and written notification to return the affected units. The returned units will be re-labeled.For further questions, please call (801) 316-3690. |  | 1297 | 1,297 total units | One consignee located in CA. No Canadian, other foreign or VA/govt/military consignees. | Open, Classified |  |  | 1297
Z-0001-2017 | 75118 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDIÂż H/S Cuvette |  |  | December 11, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received complaints of the CDI System 500 monitor displaying the H/S DISCONNECT AT CUVETTE error message when the CDI H/S Cuvette does not make a proper connection to the CDI H/S Probe. In this situation, blood parameter values for HCT, Hgb, and SO2 do not display. | Other | Terumo (Planned Action) plans to execute the following steps as part of the correction activity:Notify all customers of the potential that a cuvette may disengage from the circuit (probe) leading to a "disconnect" message on the monitor. This notification will be performed via a Customer Notification Letter. For further questions, please call (410) 392-7277. |  | 140000 | 140,000 units | Worldwide Distribution - USA (nationwide) Distribution. | Completed |  |  | 140000
Z-0017-2017 | 74320 | No | EIH | Porcelain powder for clinical use. | Dental | Dental | 510(k) | Zenostar MT Color Impulse blue, Art. no. 681165 |  |  | May 27, 2016 | 2016 | Ivoclar Vivadent, Inc. | The primary packaging of these products may show leaks in individual cases. As a consequence, the liquid may escape inadvertently and direct contact with this liquid may lead to skin irritation. | Process design | Wieland sent an Urgent Medical Device Recall Notification letters dated May 27, 2016, and Response Forms to their customers via courier service to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to Segregate any affected materials immediately. Customers were asked to complete the attached Acknowledgment form and return it to Ivoclar Vivadent, Inc. by mail, fax or email as indicated on the form. For questions contact Ivoclar Vivadent Technical Customer Service at 800-533-6825. |  | 79 | 79 units | US Nationwide Distribution in the states of: AK, AL, AZ, CA, CO, CT, FL, IA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NE, NJ, NY, OH, OR, PA, TN, TX, UT, VA, WA, and WI. | Open, Classified |  |  | 654
Z-0038-2017 | 75173 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation 2.5, 3.0, 3.5, 4.0, 4.5, 4.7, 5.0 and 4.3 (InverseArc 1.0) -- Radiation Therapy Treatment Planning System |  |  | September 01, 2016 | 2016 | RAYSEARCH LABORATORIES AB | RaySearch became aware of the problem as it was discovered in cooperation with a customer experimenting with the dose calculation on a phantom. The correction concerns two issues found with the dose calculation when using a region of interest (ROI) of type Fixation or Support with material override within the patient outline (External ROI) in RayStation 2.5, 3.0, 3.5, 4.0, 4.5, 4.7, 5.0 and 4.3 (InverseArc 1.0). | Under Investigation by firm | RaySearch sent a Field Safety Notice, on September 1, 2016, (RSL-D-61-301, dated 9/5/2016), was sent by e-mail to all affected customers. ACTIONS TO BE TAKEN BY THE USERDo not use ROIs of type Fixation or Support within the patient outline. Always use other ROI types formaterial override within the External ROI. A small overlap with the External ROI is acceptable if the Fixationor Support ROI is mainly outside the External ROI.Please educate planning staff and all users about this workaround.Inspect your product and identify all installed units with the above software version number, then confirmyou have read and understood this notice (contact information below).SOLUTIONA correction of these issues will be available in the next version of RayStation, scheduled for market releasein December 2016. In the meantime, this field safety notice is distributed to all customers. Until a correctedversion has been installed, all affected users must maintain awareness of this field safety notice.TRANSMISSION OF THIS FIELD SAFETY NOTICEThis notice needs to be passed on to all those who need to be aware within your organization. Pleasemaintain awareness of this notice as long as this version of RayStation is in use to ensure effectiveness ofthe workaround. |  | 1368 | USA: 522 units, OUS 846 units | Nationwide distribution | Open, Classified |  |  | 1368
Z-0039-2017 | 74987 | No | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension VistaÂż Calcium (CA) FlexÂż reagent cartridge |  |  | August 03, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics determined that the Dimension Vista Calcium (CA) Flex reagent cartridge lot # 16060BB may produce erroneously low or high results from specific well sets. | Under Investigation by firm | Siemens mailed on August 09, 2016, an Urgent Medical Device Recall letter to all customers who purchased the CA lot 16060BB. The notification letter requested that customers discontinue use and discard the affected kit. |  | 5991 | 5,991 cartridges | Distributed Nationwide and the countries of : Austria, Bahamas, Belgium, Canada, Czech Republic, Denmark, Germany, Italy, Japan, Netherlands, Portugal, Saudi Arabia, South Korea, Slovenia, and Spain. | Terminated | May 11, 2017 | 281 | 5991
Z-0016-2017 | 75017 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Panoramic X-ray Model PC-1000, Laser-1000 packed in a crate |  |  | August 19, 2016 | 2016 | Panoramic Rental Corp. | Panoramic Corporation is performing a voluntary Medical Device Correction on certain PC-1000 X-ray systems due to the potential for a safety nut to be ineffective, which could result in blunt force trauma that may impact the head, neck or shoulders. | Device Design | Panoramic Corporation sent an Urgent Medical Device Correction letter dated August 19, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached form and return to Panoramic. For questions contact Panoramic at correction@pancorp.com or 888-381-2577 Monday through Thursday 8:00 AM to 6:00 PM and Friday 8:00 AM to 5:00 PM EST in order to verify the effectiveness of the safety component. |  | 13340 | 13,340 | US Nationwide Distribution including Puerto Rico | Open, Classified |  |  | 13340
Z-0010-2017 | 75157 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core Model Number 728321 |  |  | August 24, 2016 | 2016 | Philips Medical Systems | The infant performance test did not meet the test specification for the Infant Head Uniformity test with High Resolution. As a result, when the user performs High Resolution head scans, the reconstructed images may exhibit degraded image quality manifested as non-uniformity of the image, CT number shift greater than 5 Hounsfield units, and reduction in gray/white matter differentiation. The issue is most sever in the High Resolution mode, and the factory reference pediatric scan protocols are set to High Resolution default. Adults head scans have a factory reference protocol that defaults to standard resolution which doesn't experience the issue. | Software design | Philips sent an Urgent Field Safety Notice (FSN) CLE 16-045 letter dated August 26, 2016, via a certified letter to customers. The letter identified the affected product, problem, actions to be taken by customer/user and actions planned by Philips to correct the problem. The letter informed customers a Philips field service representative will install a software update on the affected systems. For questions contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, follow the prompts).. |  | 358 | 358 | Worldwide Distribution - US Nationwide in the states of:: AK, AL, AR, AZ, CA, CO, CT, DE, GA, IA, IL, IN, FL, KS, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, ND, NE, NJ, NH, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY. and the countries of Algeria, Argentina, Australia, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, China, COlumbia, Cuba, Denmark, Dijbouti, Dominican REpublic, Egypt, Estonia, Finland, France, Germany, Gilbraltar, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Martinique, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Seitzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uniter Arab Emirates, United Kingdom, Uzbekistan, and Vietnam. | Open, Classified |  |  | 1302
Z-0013-2017 | 75126 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Multix Fusion Stationary X-Ray System |  |  | August 01, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential issue with the patient table for Ysio Max and Multix Fusion systems. There is an unlikely risk that an internal part of the table may detach causing the table top to tilt to the left side which could result in an injury to a patient on the table. | Process control | Siemens sent a Customer Safety Advisory Notice letter dated August 4, 2016 to their customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to discontinue use of the tables until one of their service technicians can inspect the affected part of each table. |  | 37 | 37 Systems | US Nationwide Distribution | Open, Classified |  |  | 37
Z-0045-2017 | 75068 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ROI-A Anterior Delivery Device. Orthopedic manual surgical instrument. Model number: SI-ROIA-0023 |  |  | August 29, 2016 | 2016 | LDR Spine USA, Inc. | The ROI-A Anterior Delivery Device T-Handle, part number SI-ROIA-0023 has experienced binding of the rotation of the t-handle. There have been no reports of patient injury | Device Design | The firm, LDR, contacted affected consignees via phone and then sent a follow up email dated 9/8/16. The phone script and email described the product, problem and actions to be taken. The consignees were instructed to promptly return the instrument using the RMA#1180. The firm will perform a rework that prevents the binding to occur.If you have any question, contact the Quality Engineering Manager by email: ron.musselman@zimmerbiomet.com, phone: 512-344-3436 or Field Inventory Manager - US Supply Chain or phone: 512-344-3436. |  | 24 | 24 units | US Distribution to states of: GA, OR, FL, KY, NC, AR and PA. | Terminated | January 13, 2017 | 137 | 24
Z-0043-2017 | 75115 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Operating Table TruSystem 7000 U, Material Number 1604788; Operating Table TruSystgem 7000 U (MB), Material Number 1604786 |  |  | July 10, 2016 | 2016 | Trumpf Medical Systems, Inc. | Three complaints describing TruSystem tables unexpectedly descending approximately 20 centimeters. | Nonconforming Material/Component | Trumpf Medical sent an Urgent Field Safety Notice dated July 8, 2015, to all affected consignees. The letter indicated that the tables should not be used until an authorized service engineer has exchanged the spindles of the affected operating tables. The letter included an acknowledgement form which is to be returned. The letter also requested a sub-recall if the product had been further distributed. For questions regarding this recall call 843-329-0543. |  | 84 | 84 units | Worldwide Distribution - US including CA, CT, GA, IL, NC, OR, SC, VA, WI, WY, and Internationally to Australia, Azerbaijan, Brazil, Finland, France, Germany, Hungary, Italy, Luxembourg, Norway, Poland, Russia, Slovakia, and Switzerland | Open, Classified |  |  | 84
Z-0046-2017 | 74871 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems, Inc Aquilion Lightning TSX-035A CT Scanner. |  |  | April 25, 2016 | 2016 | Toshiba American Medical Systems Inc | It has been found that if the scan range extends beyond the maximum Field of View (FOV), a reconstruction operation error may occur. As a result, scanning may be interrupted and rebooting of the system may be necessary. Please note that this issue has not occurred in the U.S. | Software design | Toshiba America Medical System, Inc. planned action to bring the defect into Compliance:1. A notification will be sent to customers with affected systems in compliance with 21 CFR 1003.21 upon approval by the Secretary. The notification letter includes a statement by the correction will be made at no cost to the customer.2. To correct this issue, modified software will be installed on the systems to prevent this occurrence from happening. The consignee will be contacted by their Toshiba service representative for an appointment when the revised software becomes available.3. To avoid this problem in scan planning, the consignee will be advised to do the following until the modified software is installed:a. Confirm that the scan range does not extend beyond the maximumFOV.b. If the ROI for the scan range (indicated by the yellow solid lines) is set as the same size as the maximum FOV (indicated by green dashed lines), the ROI position must be changed. Instructions for changing ROI will be provided as attachments to the notification letter.4. The modification to the software will be completed within 90 days from the date that the letters are submitted to Toshibas customers.5. Progress letters will be provided to the Secretary, including the number of electronic products repaired. For further questions, please call (714) 730-5000. |  | 165 | 165 | US Distribution to the states of: PA, IA, MA, SD, MI, LA, KY, NC, NY, AR, IN, WI, OH, TY, FL, WV, WA, CT, MS, AL, ND, CO, HI, NO, MN, ID, OH, GA, MT, LA, NM, VT, TN, CA, NH, NJ, NE, SC and KS. | Terminated | July 13, 2017 | 444 | 165
Z-0050-2017 | 75208 | No | DGC | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | IgE Calibrator Series |  |  | September 07, 2016 | 2016 | Randox Laboratories Ltd. | According to the firm, the IgE Calibrator Series IE2492 (lot 387-392IE) value assignment has now been aligned to the 3rd International Standard for human serum IgE 11/234.Turbidimetric (Non-IFCC Cal.) Quality Control Targets and Ranges now apply. The calibrator value sheet is to be updated with the new values. | Under Investigation by firm | The firm plans on contacting each consignee about the recall. |  | 4 | 4 | All consignees are in the US. | Open, Classified |  |  | 4
Z-0007-2017 | 75031 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 0.2 Micron Filter, 50 mm |  |  | August 24, 2016 | 2016 | Baxter Corporation Englewood | Baxter Healthcare Corporation is issuing a voluntary product recall for all unexpired lots of the 50mm 0.2 Micron Filter (Product Code: H93835) due to the potential for the filter membrane layer to be missing and for particulate matter to be present. | Nonconforming Material/Component | The firm, Baxter, sent an "URGENT PRODUCT RECALL" letter dated 8/24/2016 to its Consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) locate and remove all affected product; 2) contact Baxter Health Care Center for Service to arrange for return of product at 888-229-0001 between the hours of 7am-6pm Central Time, Monday-Friday; 3) Complete and return the Customer Reply form to Baxter via fax to 224-270-5457 or scan and email to fca@baxter.com, and 4) if you distributed this product to others facilities or department within your institutions forward copy of the communication.For general questions, contact Baxter Product Surveillance at 800-437-5176 between the hours of 8am-5pm Central Time, Monday through Friday. |  | 202647 | 202,647 | Worldwide Distribution - US (Nationwide) and countries of: Austria, Belgium, Czech Republic, Denmark, Finland, Germany, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | July 26, 2017 | 336 | 202647
Z-0057-2017 | 74614 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | The Better-Bladder With 14" ID tubing, ITEM BB14, STERILE, |  |  | July 01, 2016 | 2016 | Circulatory Technology Inc | Twisted or wrinkled balloon components which are considered a cosmetic defect by the firm. | Use error | The firm, Circulatory Technology Inc., sent an "Urgent: Device Recall" letter dated July 1, 2016 and response forms to their customers via email. The letter described the product, problem and actions to be taken. The letter instructed customers to immediately inspect their inventory visually for units with wrinkled or twisted balloons and quarantine. Customers were also advised to complete and return the response form via email to ytamari@cirtec.com. Customers with questions are asked to contact Yehuda Tamari at (516) 624-2424, Monday to Friday, 9:00 AM - 7:00 PM, EST. |  | 958 | 958 units | US Nationwide Distribution to states of: AL, AR, CA, Washington, D.C., FL, GA, IA, IL, IN, KY, LA, MI, MS, NC, NJ, NY, OH, OK, PA, SC, TN, TX &amp; VA. | Terminated | April 14, 2017 | 287 | 958
Z-0051-2017 | 75213 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | MicroStic Plus w/SS Wire 5F and MicroStic Plus w/NiTi Wire &amp; Echo Needle 4F. |  |  | September 08, 2016 | 2016 | Argon Medical Devices, Inc | The supplier initiated a recall because affected lots of the micro-introducer have an outer diameter of the dilator hub that is slightly smaller than is capable of securing the dilator to the sheath | Nonconforming Material/Component | An Urgent - Product Recall Notice was sent to consignees identifying the affected device and the reason for the recall. Customers were asked to return all affected devices. The firm will send replacement devices upon receipt of the returned affected product and once assembly of a cleared lot that does not contain recalled devices has been completed. If there are questions regarding the letter or recall action, customers can contact Beckie.Ellis@argonmedical.com. |  | 200 | 200 units | Worldwide Distribution -- US, to the states of TN and MS; and the country of Canada. | Terminated | June 07, 2017 | 272 | 200
Z-0058-2017 | 75130 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Teleflex Percuaneous Solutions, Gripper Grasper, Ref ML003EE, Rx Only Sterile Manufactured for: Teleflex Medical, Park, NC 27709 |  |  | September 09, 2016 | 2016 | Teleflex Medical | The product is being recalled due to a labeling error. The Gripper Graspers are labeled as Johans Graspers and the Johans Graspers are labeled as Gripper Graspers. | Process control | The firm, Teleflex Medical, sent an "Urgent Medical Device Recall Notification" Letter dated 9/9/2016 to consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any products listed; to return product- complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com even if you do not have the affected product.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 160 | 160 ea. in total | US Distribution to states of: AZ, CA, FL, NY, NC, and TX. | Open, Classified |  |  | 160
Z-0056-2017 | 74627 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | The Medix NatalCare LX Intensive Care Incubator. Neonatal incubator. |  |  | August 25, 2016 | 2016 | Natus Medical Incorporated | NatalCare LX incubators may have a defective base where the weld holding the wheel bushing to the tube is inadequate. There were several complaints of wheels bending out of alignment. | Unknown/Undetermined by firm | Natus sent an Urgent Notification Field Safety Notice on August 25, 2016 to all affected customers. Customers are asked to inspect the quality of the welds and wheel alignment for the Natal Care LX Incubators at their facility per attached Inspection Instructions for Assessing the Natal Care LX Wheeled-Based Joint and Weld.Customers are provided with the Natal Care LX Rolling Base Verification Form (DOC-019727) and are requested to verify the weld and gap status for each of the four wheels on the incubator and to note the results on the verification form. After the Verification form is completed and returned to Natus, Natus will arrange to ship a repair kit for each affected wheel, along with instructions and simple tools to facilitate the repair of the Natal Care LX base.Natus Technical Service will follow-up with phone calls to all non-responding customers and again to verify that the repairs have been performed.Consignees who have any questions, please call Technical Service at 888-496-2887 or send email to technical_service@natus.com |  | 104 | 104 units | US Distribution to the states of : AL, CA, FL, IL, IN, KY, LA, MA, MS, NJ, NY, OK, PA, SC and VA, and Internationally to Pakistan. | Open, Classified |  |  | 104
Z-0060-2017 | 75077 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Universal Acetabular Cup Positioner/Impactor, REF 2101-0200, Non-Sterile, Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430, USA, A Subsidiary of Stryker Corp. |  |  | August 18, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has received reports of the thread length protruding past the dome of the acetabular trial or implant for the Trident Universal Impactor/Positioner. | Nonconforming Material/Component | Stryker sent an Urgent Medical Device Recall Notification dated August 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions Needed: 1. Please inform users of this Urgent Medical Device Recall and forward this notice to all those individuals who need to be aware within your organization.2. Branches/Agencies: Return all affected products available at your location toStryker C/0 Stericycle3. Hospitals/Branches: Complete and sign the enclosed Business Reply Form and fax a copy to 877-546-0444 or email to Stericycle strvkerortho6536@stericycle.com. Our records indicate that you have received the above referenced instrument. It is our responsibility to ensure that customers who may have received this affected instrument also receive this important communication. Please assist us in meeting our regulatory obligation by faxing back the attached Business Reply Form within 5 days.We regret any inconvenience this action may cause you and if you have any questions, feel free to contact me at (201) 831-6693.For further questions regarding this recall, please call (201) 831-5000. |  | 3768 | 3,768 units | US and Internationally | Open, Classified |  |  | 3768
Z-0062-2017 | 66544 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia T Series SPECT + CT |  |  | August 16, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | The detectors on the Symbia S and Symbia T camera systems are supported and moved horizontally and vertically by four linear bearing cars. There is a potential for the linear bearing car to fail allowing the ball bearings to fall out. | Component design/selection | Customers were contacted by a service organization that performed product inspections at their sites. |  | 1109 | 1109 units | Worldwide Distribution - US: Nationwide, including Washington D.C., and Puerto Rico. and the countries of: Poland, France, Austria, Czech Republic, Ireland, Germany, Italy, Japan, Canada, United Arab Emirates, Turkey, Netherlands, United Kingdom, Brazil, South Africa, Kuwait, Belgium, Russia, Taiwan, Australia, Bangladesh, Bulgaria, Brazil, Switzerland, Chile, China, Colombia, Denmark, Algeria, Ecuador, Estonia, Egypt, Spain, Finland, Hong Kong, Croatia, Hungary, Indonesia, Iran, Israel, India, South Korea, Kuwait, Kazakhstan, Lithuania, Latvia, Morocco, Mexico, Norway, New Zealand, Peru, Phillipines, Pakistan, Portugal, Qatar, Romania, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, Thailand, Tunisia, Taiwan, Venezuela, and Vietnam. | Completed |  |  | 1841
Z-0063-2017 | 75177 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Blood Pressure Monitor UA-767PBT-Ci for small cuff. Designed to be used by end users who are eighteen (18) years and older at home to monitor their blood pressure (systolic and diastolic) and pulse rate. |  |  | September 12, 2016 | 2016 | A &amp; D Enginering Inc | A&D; may have shipped blood pressure monitors that were not labeled for use with small cuff. | Labeling Change Control | Notification letters were sent on September 12, 2016 to all affected customers. Letters identify the product and reason for the recall. Letters include a serial number list that is customer specific and ask that inventory be reviewed and that other organizations that may be impacted by this issue be notified. |  | 378 | 378 units | to distributors in NY, MA, VA. TX, | Open, Classified |  |  | 378
Z-0064-2017 | 74922 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | TransMotion Medical, TMM3 Series Video-Fluoroscopy Swallow Study Stretcher-Chair. |  |  | June 29, 2016 | 2016 | Winco Mfg., LLC | Non-conforming component caused arms of TMM4 Stretcher chair to no longer meet design specifications. | Nonconforming Material/Component | Customers were sent an Urgent: Medical Device Recall notification on 7/12/2016 for the specific models of the TransMotion (TMM) Stretcher chairs. The letter explained the reason for the recall and asked customers to take action. Customers are asked to follow the instruction in the Operator Manual for proper egress. Patients are not allowed to egress from the chair without assistance nor should they be left unattended. The chairs will be repaired on site; and, removal of the chair from service is not required with proper supervision and use. Customers are asked to complete the instructions for the acknowledgement form and return. The firm will contact the customers to arrange for repair. Questions regarding the field correction should be directed to the Customer Care department at 800-237-3377 or 352-854-2929 or e-mail QA-RA@wincomfg.com |  | 264 | 264 Stretcher Chairs | Worldwide Distribution -- USA, including AR, CA, CO, CT, FL, GA, IL, KS, LA, MD, MI, MN, MO, MS, NC, NJ, NM, NV, NY, OH, OK, PA, RI, TN, TX, VA, WA, WI, and Puerto Rico; and the countries of Canada, Australia, Italy, Panama, Switzerland and Taiwan. | Open, Classified |  |  | 264
Z-0069-2017 | 75297 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | V. Mueller Neuro/Spine, SHADOW-LINE, SPINAL DISTRACTION SCREW, STERILE, various sizes (12mm, 14mm, 16mm and 18mm) 5 PER BOX. |  |  | September 23, 2016 | 2016 | Cardinal Health 200, LLC | Potential risk associated with corrosion demonstrated on the distraction screw surface. | Package design/selection | On September 23, 2016, Becton Dickinson sent a Field Safety Notice to all affected customers describing the details of the recall and necessary action to take. ACTIONS TO BE TAKEN BY BD:A credit will be applied to your account for the affected product upon receipt of the affected unit(s) or receipt of certification of destruction documented on the customer response form.Complete the Recall Response form certifying that all affected product in your inventory has been destroyed and mail/email a copy of the response form.: The firm requests customers to promptly return the enclosed response form to expedite the correction process and acknowledge receipt of this notification. Or email: GMB-GLBÂżVMFieldActions@carefusion.com For further questions, please call (847) 473-1500. |  | 72076 | 72,076 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Germany, New Zealand, Saudi Arabia, South Africa, Switzerland and Thailand. | Open, Classified |  |  | 72076
Z-0067-2017 | 75148 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | The Introes Pocket Bougie is an endotracheal tube introducer. The device is an endotracheal tube introducer and it is used as a guide during intubation. Anesthesiology. |  |  | August 19, 2016 | 2016 | Brathwaites Oliver Medical | The Introes Pocket Bougie an endotracheal tube introducer size 14 Fr (4.7mm) x 60 cm may break during intubation. | Nonconforming Material/Component | On August 19, 2016 the firm sent a MEDICAL DEVICE RECALL NOTICE &amp; CORRECTIVE ACTION letter to affected customers. The letter described the reason for the recall and necessary actions to be taken by the consignees. These included:BOMImed is requesting bougies with lot# 101615 to be returned for evaluation due a potential manufacturing concern.Complete the included Medical Device Recall Customer Letter and follow the instructions to obtain a Return Material Authorization number andapplicable shipping instructions to return the affected products.BOMImed Regulatory department also requires the below CustomerConfirmation section to be completed and return toregulatory@bomimed.com to confirm this notification was received. |  | 760 | 760 units | U.S. Consignee located in the following states: IL, WI, WE, OR, NY, UT, IN, NE, PA, MO, OH, AZ, MA, CT and NC.*** Foreign Consignee: Canada. | Open, Classified |  |  | 760
Z-0071-2017 | 75330 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC Brivo 715 Prime, OEC Brivo 785 Essential, OEC Brivo 865 Advance |  |  | September 23, 2016 | 2016 | GE OEC Medical Systems, Inc | GE Healthcare Surgery announces a voluntary field action for the OEC Brivo 715 Prime, OEC Brivo 685 Essential, and OEC Brivo 865 Advance because the circuit board that controls vertical motion of the C-arm may fail, resulting in the C-Arm moving up or down without command. | Nonconforming Material/Component | GE Healthcare sent an Urgent Medical Device Correction letter dated September 23, 2016 to customers. The letter informed customers of the affected product, safety issue, and safety instructions for immediate mitigation of the issue. GE Healthcare will correct all affected products at no cost to customers. A GE Healthcare representative will the customers to arrange for the correction. Please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 148 | 148 total | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, FL, KS, MS, NJ, NV, NY, PA, TX, UT, and WV. No Candaian or Va/govt/military consignees. There are 133 other foreign consignees. | Terminated | April 24, 2017 | 213 | 148
Z-0068-2017 | 75273 | No | OMT | Absorbable lung biopsy plug. | General &amp; Plastic Surgery | Anesthesiology | 510(k) | BioSentry Tract Sealant System, Product Code: 768022019S; Indicated to provide accuracy in marking a biopsy location for visualization during surgical resection. |  |  | September 24, 2016 | 2016 | Surgical Specialties Mexico S DE RL DE CV | Surgical Specialties Mexico is recalling the BioSentry Track Sealant System due to the potential for an open seal on the product package. | Process control | An Urgent Product Recall Notification Letter dated 9/24/16 was sent to customers to inform them that Surgical Specialties Mexico is recalling the BioSentry Track Sealant System due to the potential for an open seal on the product package. The letter informs the customers to complete the customer acknowledgement form and return via fax to their quality assurance department at (610) 404-4010 or return via email to kknappenberger@surgicalspecialties.com. Customers with questions are instructed to contact Kelly Knappenberger at (484) 220-2402 or Kris Liberacki at (262) 994-8083. |  | 512 | 512 boxes (2560 units) | US: OH, PA, CO, MA, DE, CA, NY, VA, WA, NJ, CT, FL, TX, NH, ID, OK, AZ, GA, PA, NV, VT.Worldwide: Spain, Italy, United Kingdom, Portugal, Turkey, Germany, Denmark, | Open, Classified |  |  | 512
Z-0070-2017 | 75279 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic Navigation O-arm O2 Surgical Imaging System Product Catalog Number: BI-700-02000 |  |  | September 22, 2016 | 2016 | Medtronic Navigation, Inc.-Littleton | O-arm O2 Surgical Imaging System Spatial calibration may be erroneous in Stealth Station navigated images | Software design | Medtronic Navigation contacted customers via telephone beginning September 23, 2016. Customers were informed of the affected product, problem and actions to be taken. Medtronic requests that users immediately discontinue the use of the O-Arm O2 Surgical Imaging system for automatic registration in Stealth Station navigated surgical procedures until further notice.Medtronic Service personnel have been scheduled to begin the necessary inspections and re-calibration activities on customer systems beginning on September 23, 2016. |  | 14 | 14 units | Worldwide Distribution - US Nationwide in the states of CO, KA, NE, MI, MO, OH, WA and country of Switzerland | Terminated | May 12, 2017 | 232 | 14
Z-0073-2017 | 75192 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Fujifilm Synapse PACS, Software versions: 4.0.xxx, 4.1.xxx, version 4.2.xxx, version 4.3.xxx |  |  | September 15, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | Image data for a patients image may not be correct | Software design | The firm, FMSU, sent a "Medical Device Correction" letter dated September 15, 2016 to its customers. The letter describes the product, problem and actions to be taken. The firm will install a patch that corrects the issue (Synapse IPSS v1.3.1.0 Patch).Customers with Version 4.0 may upgrade to version 4.4 (will not require the patch) or upgrade to 4.1, 4.2, or 4.3 and install the patch. FMSU will install the patch on the server remotely if your site uses Active Monitoring. Your notification that the patch was installed will be your service ticket. For sites not using Active Monitoring, FMSU will schedule a time to install the patch on your server.Questions or concerns, contact the FUJIFILM Customer Support Center at 1-888-FUJIMED (1-800-385-4633) Option #1. International callers can access the FUJIFILM Customer Support Center by calling 1-203-602-3580. |  | 251 | 251 units | US Nationwide Distribution. | Open, Classified |  |  | 251
Z-0072-2017 | 74162 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur XPT Immunoassay System |  |  | April 28, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | The ADVIA CentaurÂż XPT default setting for the Daily Maintenance Task (Daily Cleaning Procedure) frequency may have the Daily Maintenance Task frequency set to, As needed instead of Daily. Not performing the daily maintenance may impact any assay and can be detected through monitoring of quality control and calibration results. Systems running in a language other than English are not impacted. | Under Investigation by firm | Siemens sent an Urgent Medical Device Correction letter dated April 28, 2016, to all affected customers. The letter advised customers not to release results generated by the system if quality control results are not within acceptable ranges. Customers were advised to perform the following steps to check the Setting for the Daily Maintenance Task frequency on your instrument: *Open the Maintenance Tab, * Select Daily Maintenance task and * Verify that the Frequency Column is set to "Daily at..". The Daily Maintenance Task frequency setting can only be reset by a Siemens Customer Service engineer. If your instrument is set to "As Needed": Daily Maintenance Task Frequency Setting May be Incorrect. Customers should contact their local Technical Support Center to schedule a visit by your Siemens Customer Service engineer. Customers should also perform the daily maintenance by selecting the Daily Maintenance Task and then selecting Perform on the right bar. In addition, all customers were asked to complete and return the Field Correction Effectiveness Check within 30 days.For questions regarding this recall call 914-631-8000. |  | 292 | 292 units Total (3 domestically &amp; 289 internationally) | Worldwide Distribution - US including MD, MN, &amp; MS; **Internationally** AE, AT, AU, BE, BR, CA, CN, DE, DK, EG, ES, FI, FR, GB, GP, HU, IN, IT, JO, JP, KR, MY, NO, NZ, PH, PT, QA, SA, SE, SG, TR &amp; VN | Open, Classified |  |  | 292
Z-0078-2017 | 75350 | No | NAW |  |  | Radiology | PMA | TheraSphere(R) YTTRIUM-90 CLASS MICROSPHERE INJECTION |  |  | September 21, 2015 | 2015 | Biocompatibles U.K., Ltd. | One mislabeled dose of TheraSphereÂż Y-90 Glass Microspheres, which was shipped to Hong Kong (total radioactivity incorrect). The product ws labeled as 10 GBq and the actual does was 5 GBq. | Packaging | Biocompatibles UK Ltd sent an Urgent Field Safety Notice dated September 21, 2015, to all affected customers. The firm did not request return or destruction of the product. Customers with questions were instructed to call 613-410-1870.For questions regarding this recall call 613-801-1848. |  | 64 | 64 doses | Worldwide Distribution - US to CA, FL, IL, WI, IN, OR, PA, MA, NY, KY, TX, MI, MD, NV, WA, OH, MN, and Internationally to Germany, Turkey, Hong Kong, Italy, Portugal | Completed |  |  | 64
Z-0076-2017 | 75303 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | September 29, 2016 | 2016 | Elekta, Inc. | When creating 3D plans using either MU or Dose weighting modes, if the user changes the Physician's Intent Rx Dose and/or the number of fractions, and then modifies the wedge angle, the MU value is scaled incorrectly. | Device Design | Elekta sent an Important Field Safety Notice 382-01-MON-005 was sent to all affected customers on 9/29/2016. The notice informs users of the specific product and version numbers affected by the issue,and any work arounds that can be used to avoid the issue. The notice included an acknowledgment form which is to be returned.For questions regarding this recall call 314-993-0003 or 800-878-4287. |  | 365 | 365 | Worldwide Distribution - US to CA, FL, KY, MD, MO, MT, NJ, NC, PA, WA, WI. Internationally to Algeria. Antiqua &amp; Barbuda, Argentina, Australia, Austria, Belarus, Botswana, Brazil, Bulgaria, China, Columbia, Croatia, Czech Republic, Ecuador Estonia, France, Germany, India, Iran, Israel, Italy, Japan, Malaysia, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Portugal, Romania, Russia, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, and Zimbabwe. | Open, Classified |  |  | 365
Z-0077-2017 | 75296 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Head Frame - Model number HHF -ALL Head Frame Posts (PN1008016 ) and Post Kits (PN1008016). |  |  | September 13, 2016 | 2016 | Varian Medical Systems, Inc. | Varian Medical Systems has received a report that a user was able to easily rotate the head frame posts when attached to the mounting cam on the metal head ring. There was no report of serious injury due to this issue | Device Design | Varian Medical Systems sent an Urgent Medical Device Removal letter dated September 14, 2016, to all affected sites. Each site was also contacted by telephone to advice users of the issue and request of return of the head frame posts on September 14 and 15th. Calls enabled Varian to prioritize replacement parts for those with patients scheduled in the immediate future.Recommended User Actions:1.CEASE USE of the Varian Head Frame Assembly and QUARANTINE the Varian Head Frame Assembly and components.2.Collect, Clean, Disinfect all Head Frame Posts [PN1008016] if used previously.3.Contact Varian 1.888.827.4265 for Return Material Authorization number, and return all Head Frame Posts in accordance with instructions provided with the enclosed Recall Return Material Form.We sincerely apologize for any inconvenience and thank you in advance for your cooperation. If you require further clarification, please feel free to contact your local Varian Medical Systems Customer Support District or Regional Manager. This notice has been provided to the appropriate Regulatory Authorities. For further questions, please call (650) 424-6833. |  | 75 | 75 devices | US Distribution and Internationally to 1 site in Canada, India and Morocco | Open, Classified |  |  | 75
Z-0082-2017 | 75037 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | SIEMENS IMMULITE/IMMULITE 2000 Systems Rubella IgM (RUM), REF/Catalog Number for US: L2KRM2(D), Siemens Material Number (SMN) for US: 10374021, REF/Catalog # for OUS: L2KRM2, SMN for OUS: 10381327; Unique Device Identification Number (UDI) 00630414961996; IVD --- |  |  | August 24, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens received multiple customer complaints indicating an increase in false positive and indeterminate patient sample results and increased imprecision for several lots of Rubella IgM assay on the IMMMULITE 2000/IMMULITE 2000 XPi platforms. | Under Investigation by firm | Siemens sent an Urgent Medical Device Recall Letter IMC16-20.A.US , dated August 24, 2016, to the US consignees via FedEx, and an Urgent Field Safety Notice IMC16-20.A.OUS (dated August 2016) was sent to foreign representatives for distribution to foreign consignees. Customers were instructed to discontinue use of and discard the affected kit lots. Customers were also requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Recall letter within 30 days. Field service personnel will be sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. For questions, customers were instructed to contact their Siemens Customer Care Center or their local Siemens technical support representative. |  | 3612 | Domestic: 177 Kits; Foreign: 3435 Kits | Worldwide Distribution n- US (Nationwide) and Foreign distribution to the following countries: Argentina, Bangladesh, Brazil, Bulgaria, China, Colombia, Costa Rica, Cyprus, Egypt, France, Germany, Greece, Guatemala, India, Indonesia, Israel, Italy, Kazakhstan, Macedonia, Mexico, Paraguay, Poland, Portugal, Romania, Russian Federation, Serbia, Singapore, Slovakia, South Africa, Syria, United Arab Emirates, and Uruguay. | Open, Classified |  |  | 3971
Z-0006-2017 | 74999 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD), US product: Model No.: 1103; and International Product: Model No.: 1104XX. |  |  | August 17, 2016 | 2016 | HeartWare, Inc | Foreign material found within the driveline connector of the HVAD Pump that could cause increased impedance resulting in electrical faults. | Component design/selection | HeartWare notified their customers with an Urgent Medical Device Recall notice on 8/17/2016 of the HVAD Pump, listing the US Product: Model No. 1103 and the International Product: Model No.: 1104. The letter discussed the reason for the recall as well as the risk to health. Customers were to identify affected product in their inventory; and, complete and return the acknowledgement form. Acknowledgement should be provided no later than 30 days from date of the letter either to the HeartWare representative; or, via e-mail to FSCA @Heartware.com; or, FAX the signed form to (305) 364-2665. Affected product should be returned upon receipt of replacement product; and, the completion form should also be completed and returned. Questions or concerns should be directed to a local HeartWare representative. |  | 344 | 344 pumps | Worldwide Distribution - US, in the states of AL, AR, AZ, CA, DC, FL, IL, IN, KY, LA, MA, MI, MN, MO, NC, NE, NJ, NY, OH, OK, PA, SC, TX, UT, VA, and WI; and, the countries of Argentina, Belarus, Bosnia, Herzegovina, Brazil, Canada, Chile, Colombia, Hong Kong, India, Israel, Kazakhstan, Kuwait, Lebanon, Malaysia, New Zealand, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Taiwan, Turkey, and Vietnam. | Open, Classified |  |  | 344
Z-0075-2017 | 75274 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Banyan Medical 5 x 55mm Pyramidal Tip Trocar w/ 2X Cannula w/ |  |  | September 07, 2016 | 2016 | Banyan Medical, LLC | Potential lack of assurance of sterility | Process control | Banyan Medical issued Urgent Voluntary Recall Notification on 9/7/16 to Distributors to remove product from inventory and contact hospital to examine and return product. Questions contact your local Banyan Sales Representative |  | 320 | 64 boxes of 5 | IN, TN | Open, Classified |  |  | 415
Z-0080-2017 | 75240 | No | EAJ | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | Zero-Gravity Radiation Protection System |  |  | September 09, 2016 | 2016 | Contour Fabricators Inc | Two customer complaintssuggest that excessive rotation of the cable in one direction (clockwise or counter-clockwise)without allowing the cable to unwind to its normal condition may cause the cable to breaksuddenly and the body shield to fall. While no injuries or deaths occurred in either failure, theZero-GravityTM body shield weighs 55 pounds, and in the event of a cable failure, the bodyshield will fall immediately. A fall of this nature places the user, associated staff and the patientat risk for a potential injury. | Labeling design | On 9/9/2016, URGENT: MEDICAL DEVICE CORRECTION letters dated 9/8/2016 were sent to the affected consignees via certified mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. TIDI Products is taking the following action to correct this issue:1. Revising instructions for use to include the step to allow the cable to unwind after useand prior to storing the body shield in a locked or secured position.2. Modifying user and staff training protocols to include the requirement to allow the cableto unwind after each use.3. Adding additional warning labels to the system.4. Replacing all cables in the field to ensure that any potentially stressed cable is removed.A representative from TIDI Products will be contacting you to schedule a visit from a technicianto implement this corrective action.Customer ActionsCables should be carefully inspected for signs of wear or rotational tension, which will create abend or buckle in the cable. Practitioners and medical institutions should review their routineuse of the Zero-Gravity Radiation Protection System to determine if the body shield undergoesrepeated rotation without allowing the cable to relax back to its normal condition. If the cablehas been consistently rotated in one direction without being allowed to relax back to its normalcondition or if there is any doubt about excessive rotation in one direction, TIDI recommendsthat practitioners and medical institutions discontinue use until such time as corrective actioncan be implemented by TIDI Products.If facilities continue to use their Zero-Gravity" Radiation Protection System until correction andinstall of the replacement cable can be performed, they should make sure that they are allowingthe unit to go back to its relaxed state.Please complete and return the enclosed response for |  | 447 | 447 | Nationwide. Canada.AUSTRALIAAUSTRIABELGIUMBRUNEICHINAFRANCEGERMANYINDAINDONESIAISRAELITALYKSAMALAYSIANETHERLANDSPOLANDQATARQUEBECSAUDI ARABIA SINGAPORESOUTH AFRICATAIWAN (R.O.C.) | Open, Classified |  |  | 447
Z-0079-2017 | 74750 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System, Model 5.0 |  |  | May 30, 2016 | 2016 | RAYSEARCH LABORATORIES AB | For a treatment plan consisting of multiple beam sets, the table for ROI plan dose statistics in the report may show the statistics for a beam set dose. The error can only be triggered when using a report template where statistics for a beam set dose is included as the last dose statistics prior to the plan dose statistics. | Software design | On May 19th, 2016, RaySearch Laboratories distributed Field Safety Notice, Medical Device Correction notices and reply forms to their customers via email. Actions to be taken by the customer include: 1) Inspect all report templates that are intended to be used for plans with multiple beam sets. 2) Make sure that POI Dose statistics (Plan) is included immediately before ROI Dose statistics [Plan dose] in the Report content area. 3) Please educate planning staff and all users about this workaround. 4) Inspect your product and identify all installed units with the above software version number. Finally, complete and return the reply form via email to: freddie.cardel@raysearchlabs.com or by fax to: 888-501-7195 to confirm that you have read and understand the information. Customers with questions can contact David Hedfors via email: david.hedfors@raysearchlabs.com |  | 491 | 491 units (187 domestically &amp; 304 internationally) | AZ, CA, FL, MT, NC NY, TX, OH &amp; WA | Open, Classified |  |  | 491
Z-0087-2017 | 75229 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Uniflex Humeral Nail: 467265 G-PIN 2.8BL 3.2PL 60CM 2PK |  |  | August 02, 2016 | 2016 | Zimmer Biomet, Inc. | The G-PIN 2.8BL 3.2PL 60 CM 2PK is intended to contain one (1) 467265-01 Ball Tip Guide Pin and one (1) 467255-02 Plain Tip Guide Pin. The G-PIN 2.8BL 3.2PL 60 CM 2PK may have been packaged incorrectly either with two (2) 467265-01 Ball Tip Guide Pins and zero (0) 467255-02 Plain Tip Guide Pins or with zero (0) 467265-01 Ball Tip Guide Pins and two (2) 467255-02 Plain Tip Guide Pins.If the G-PIN 2.8BL 3.2PL 60 CM 2PK contains the incorrect combination of Plain Tip Guide Pins and Ball Tip Guide Pins, the surgeon would either have to implant the nail without the Plain Tip Guide Pin as a guide, ream with a Plain Tip Guide Pin, or choose to not ream the humeral canal. The risk associated with not having the correct combination can include a delay in surgery of less than 30 minutes, the reamer tip remaining in the distal humeral canal, or an iatrogenic fracture. | Mixed-up of materials/components | The firm, Zimmer Biomet, sent an " URGENT MEDICAL DEVICE RECALL REMOVAL" notification dated 8/2/2016 affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The consignees were instructed to review the notification and ensure affected team members are aware of the contents; immediately locate and quarantine affected product in your inventory; complete and return the Certification of Acknowledgment portion of Attachment-1 Inventory Return Certification Form via digital copy within three (3) days via email to CPWARFieldAction@zimmerbiomet.com or via fax to 574-372-1683;Immediately return all affected product from your distributorship and affected hospitals within your territory; and retain a copy of your recall acknowledgement and product return form.If after reviewing this notice you have further questions or concerns please call 1-800-348-9500 ext. 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of business hours will receive a prompt to record a voicemail and will be returned upon receipt. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. |  | 22 | 22 units | US Distribution to states of: MD,TX, FL, GA, VA and PA. | Terminated | March 24, 2017 | 234 | 22
Z-0004-2017 | 75181 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Skintact Electrodes for Defibrillation, DF29N. |  |  | September 01, 2016 | 2016 | Leonhard Lang Medizintechnik GmbH | There is a risk that defibrillation electrodes model DF29N will be connected with the defibrillator Welch Allyn AED 10 only with delay or not at all. This may cause a situation in which a patient, who is in a life threatening condition and requires adefibrillation shock, cannot be treated in good time. | Device Design | Consignees was sent via e-mail a Leonhard Lang "Important Safety Notice" dated September 1, 2016. The letter described the product being recalled, Description of the defect, Actions and time frame of the recall, &amp; Compensation for the recalled electrodes. Advised consignees to inform all users within their organizations of the recall; and, secure and destroy the unused electrodes in their inventory. Confirmation of destruction can be done by completing and returning the "Confirmation of Destruction / Consumption" form by October 14, 2016. If the product was further distributed, they were to forward a copy of the Safety Notice. For questions contact sales staff. |  | 11110 | 11,110 ( US 8,040; OUS 3,070) | Worldwide Distribution - US, to the state of Florida; and, the countries of South Africa, Germany, Slovenia, France, United Arab Emirates, Great Britain, Italy, Israel, Serbia, Thailand, Thailand, South Africa, Lebanon, and Poland. | Open, Classified |  |  | 11110
Z-0084-2017 | 75234 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Phoenix Recon Targeting Arm |  |  | August 09, 2016 | 2016 | Zimmer Biomet, Inc. | Old revisions and new revisions of the soft tissue sleeve and recon targeting arm are not interchangeable, as the new revision tissue sleeve will not fit into the old revision targeting arm. The old revision tissue sleeve has too much clearance with new revision targeting arm.Risks of encountering the affected product include: the correct instrument may not be readily available for the procedure causing the surgeon to complete the procedure by hand; or a delay in surgery greater than 30 minutes may occur while another Recon Targeting Arm or Recon Soft tissue Sleeve is located or while the surgeon completes the procedure by hand. | Finished device change control | On 8/9/2016, URGENT MEDICAL DEVICE RECALL REMOVAL notifications were sent to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions or concerns call the customer call center at 1-800-348-9500 ext 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. |  | 31 | 31 | US Nationwide Distribution in the states of: MI, UT, SD, OH, WA, KS, MO, NJ, GA, and TX. | Terminated | March 23, 2017 | 226 | 31
Z-0085-2017 | 75056 | No | JOS | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CONCEPT HEATWAVE Electrode 130 mm x 2.3 mm diameter, REF/Catalog Number CS-023, STERILE, Rx Only |  |  | August 29, 2016 | 2016 | ConMed Corporation | CONMED received complaints of damage to the CS-023 electrodes shaft insulation which could result in burn marks on the shaft. If the shaft insulation is damaged in any way, it could increase the risk of patient burn. | Process change control | Conmed sent an URGENT DEVICE RECALL letter dated August 31, 2016 and Business Reply Form were sent to the domestic consignees on August 29, 2016 via USPS First Class Mail. The letter identified the affected product, problem and actions to be taken. For questions or requests contact the Field Action Support Team at 1-800-535-8536, fax to 315-624-3225, or email largheatwv@conmed.com. |  | 4770 | Domestic: 1740; Foreign: 4770 | Worldwide Distribution - US Nationwide in the states of: CA, HI, IL, IN, KS, LA, MA, NC, NM, NY, OH, OK, PA, TN, TX, UT, WA and WI; plus Foreign distribution to Australia, Canada, France, Germany, Italy, Spain, Czech Republic, Lithuania, Netherlands, Poland, Korea,Morocco, Malaysia, Mexico, Taiwan, and South Africa. | Open, Classified |  |  | 4770
Z-0005-2017 | 72001 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System Controller |  |  | June 08, 2016 | 2016 | HeartWare Inc | Loose connector ports in controller body. | Under Investigation by firm | HeartWare sent an Urgent Medical Device Correction letter dated June 8, 2016, to all affected customers, HeartWare advised clinicians to remind patients to follow unit instructions and water avoidance, and to regularly schedule appointments to check the integrity of the unit and to be cautious with connectors when inspecting connectors. Clinicians were also instructed to send back the enclosed "Acknowledgement Form" within 30 days and to forward the message to necessary individuals within their organization.Customers with questions were instructed to contact their local HeartWare representative.For questions regarding this recall call 305-364-1575. |  | 13363 | 8799 Worldwide, 4,564 in US | Worldwide Distribution - US Nationwide in the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI and countries of Australia, Austria, Belgium, Canada, France, Germany, Lebanon, Malaysia, Sweden, Switzerland, &amp; United Kingdom. | Open, Classified |  |  | 13363
Z-0090-2017 | 74950 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products LAC Slides, REF/Product Codes 843 3880 (300 slides per sales unit) and 815 0112 (90 slides per sales unit), Unique Device Identifier Nos. 10758750004911 and 10758750004454, Rx ONLY, IVD --- |  |  | August 11, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics confirmed that N-Acetylcysteine (NAC), a drug used for acetaminophen overdose, mucolytic therapy and is used as nutritional supplement, can interfere with VITROS GLU, LAC, TRIG, and URIC Slides assays. | Under Investigation by firm | Ortho-Clinical Diagnostics sent an URGENT PRODUCT CORRECTION Letter (Ref. CL2016-152) dated August 11, 2016 via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to all customers who could order and use affected product to inform them that new interferant were identified for GLU, LAC, TRIG and URIC Slides which may cause biased results. Foreign affiliates were informed by e-mail on August 11, 2016, of the issue and instructed to notify their consignees of the issue. For questions, please contact our Ortho CareTM Technical Solutions Center at 1-800-421-3311.For questions regarding this recall call 908-218-8776. |  | 62013 | US: 32,081 sales units; Foreign: 29,932 sales units | Worldwide Distribution - US (Nationwide) Puerto Rico and Bermuda), Internationally to Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Poland, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 805842
Z-0093-2017 | 75163 | No | ORZ | Eyelid thermal pulsation system. | Ophthalmic | Ophthalmic | 510(k) | LipiFlow thermal Pulsation System, consisting of Console and Activator, Model LFD-1100 |  |  | September 13, 2016 | 2016 | Tearscience, Inc | Complaints related to an error condition occurring during patient treatment where the FipiFlow System automatically stopped treatment after detecting a problem with the Activator. | Device Design | The firm, TearScience, sent a "Medical Device Recall" letter dated September 9, 2016 to its consignees/customers on 9/13/2016. The letter described the product, problem and actions to be taken. The customers were instructed to complete and return the enclosed Acknowledgement and Receipt form by fax to (919) 882-9603, email to: Customerservice@tearscience.com or mail to TearScience, ATTN: CUSTOMER SERVICE RCL001, 5151 McCrimmon Parkway, Ste 250, Morrisville, NC 27560; to discontinue use of the product; to remove all a product with lot #201615220020 from your inventory; and call TearScience Customer Service at 919-459-4891 to return the product for above mentioned lot# and request new product for replacement.For any questions regarding this letter, please call (919) 459-4891 Monday through Friday, 8:30AM to 5:30PM, Easter Time. |  | 12000 | 1,200 Activators (120 boxes of 10 units) | US Distribution including states of: AK, AR, CA, CO, FL, GA, ID, IL, IN, LA, MD, MA, MN, MO, NV, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA , WA, WI and Puerto Rico. | Open, Classified |  |  | 12000
Z-0088-2017 | 75257 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | SDMS (SenTec Digital Monitoring System) |  |  | September 15, 2016 | 2016 | SenTec AG | Sentec AG found that with one batch of an internal SenTec Monitor component (Docking Station Module) there is a risk that a fastening screw could become loose under certaincircumstances. A loose screw within the monitor can create uncontrolled electrical connections, including short circuits, with the risk of smolder or fire and electrical shock (patient and/or operator). | Component change control | Distributors were sent via e-mail a SenTec 'Mandatory initiated field correction action" letter dated September 15, 2016 . The letter provided a "Description of the issue, Affected products, Identify and localize affected products, Removal from use, Replacement of affected products at end customer &amp; distributor / sub-distributor level". It also provided "action information for sub-distributors and affiliates, and end customers". For questions contact SenTec Inc. by e-mail (customersevice.usa@sentec.com) or phone (636-343-0000). |  | 21 | 75 (54 US; 21 =OUS) | Worldwide Distribution - US Distribution to the states of : LA, MI, MO, OH, TN. and to the countries of : United Kingdom, The Netherlands and Australia. | Terminated | June 23, 2017 | 281 | 21
Z-0118-2017 | 75161 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CADstream software |  |  | October 19, 2011 | 2011 | Merge Healthcare, Inc. | Customers may experience an issue with the software study preferences when changes are made to the study protocol, resulting in incorrect patient follow-up. | Software design | Merge sent an Urgent Medical Device Recall dated January 28, 2016 to affected customers via email and certified mail if no email address was available. The letter identified the affected product, problem and actions to be taken. For questions Customer Service at 877-741-5369 or support@merge.com. |  | 844 | 844 sites potentially have the affected versions | US Nationwide Distribution | Open, Classified |  |  | 844
Z-0113-2017 | 75119 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Cytoplast Titanium-Reinforced Posterior Singles Membrane |  |  | September 07, 2016 | 2016 | Osteogenics Biomedical, Inc. | The device has the potential to be packaged in the wrong product box. | Labeling mix-ups | Osteogenics Biomedical sent an Urgent Customer Notification letter dated September 7, 2016, to all affected customers. This recall has been extended to the user level (clinician). All units sold domestically were to the end users (clinicians). End users were instructed to inspect inventory and return any affected product.For further questions, please call 1-888- 796-1923. |  | 32 | 32 boxes | Worldwide Distribution - US Distribution and to the countries of : Colombia, Venezuela/Ecuador, Japan, India, Canada, and UK. | Open, Classified |  |  | 32
Z-0136-2017 | 75255 | No | JEC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Agee-WristJackÂż Fracture Reduction System; |  |  | September 14, 2016 | 2016 | Hand Biomechanics Lab Inc | A number of Agee WristJackÂż Fracture Reduction Systems have been manufactured with a molded beam slightly larger than optimal. This may result in the surgeon experiencing difficulty fine-tuning the length (red) and radial/ulnar alignment (blue) adjustments on the device. | Under Investigation by firm | Hand Biomechanics Lab, Inc. sent a Voluntary Medical Device Recall letter to all affected customers beginning on September 14, 2016 and sent by phone/fax numbers - or e-mail. Letters request return of devices which have not yet been used. Customers with questions were instructed to contact Customer Service at (800) 522-5778.For questions regarding this recall call 916-923-5073. |  | 65 | 65 units | Nationwide Distribution | Open, Classified |  |  | 65
Z-0135-2017 | 74774 | No | GBW | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | COREÂż Trumpet Handpiece only, Catalog Number CD8450 |  |  | June 27, 2016 | 2016 | ConMed Corporation | The packaging seal may contain a crease. The crease may result in an open channel. These channels may compromise sterility of the product. | Under Investigation by firm | ConMed sent an Urgent Device Recall letter dated June 27, 2016, to all affected customers with reply forms via USP First Class Mail. International customers received a notice dated July 11, 2016 by courier service. Customers were advised to inspect their inventory for any of the devices subject to recall. Customers were urged to contact all departments within their facility and any other facilities within their organization that may have received affected products. It is imperative that all end users of these devices receive the notification. Customers who have received any of the affected product are asked to complete and return the reply form along with the device to: CONMED Corporation 525 French Road Utica, NY 13502 USA Attn: Ed Kovac. Please do not return used devices. Replacement product will be issued unless otherwise requested. If customers do not have any affected devices to return, complete and return the business reply form indicating they have no devices and return by one of the means listed:1. Email to: coresihp@conmed.com2. Fax to: Field Action Support Team at 315-624-3225. 3. US mail, using enclosed postage-paid envelope, to: CONMED, 525 French Road, Utica, NY 13502; Attn: Field Action Support Team. Customers with questions or requests, should contact the Field Action Support Team at 315-624-3237, fax to 315-624-3225, or email coresihp@conmed.com. |  |  |  | Worldwide Distribution - US (Nationwide) Internationally: to Australia, Austria, Uruguay, Denmark, Germany, Ireland, Italy, United Kingdom, Guatemala, Czech Republic &amp; Canada. | Open, Classified
Z-0123-2017 | 75207 | No | CDT | Triglyceride test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Randox Lipase (Colorimetric) |  |  | August 02, 2016 | 2016 | Randox Laboratories Ltd. | According to the firm carry over was observed when the Lipase test is run directly before or after the LDL and Triglycerides assays. Customers are instructed to amend the running order for all instruments using these reagents to ensure Lipase is not place directly before or after LDL or Triglycerides. A correction was made to the IFU for the Lipase assay and all customers who purchased the Lipase kit in the USA were contacted with the updated instruction. | Nonconforming Material/Component | Randex Laboratories sent an Urgent Field Safety Notice letter dated June 17, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the response section of this form technical.services@randox.com. For questions contact Randex Technical Service. |  | 18 | 18 | US Nationwide Distribution | Open, Classified |  |  | 18
Z-0125-2017 | 74397 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Calix P PEEK Lumbar System, PLIF Trial |  |  | May 17, 2016 | 2016 | X Spine Systems Inc | The trials and rasps used to prepare the surgical site for placement of the implant may become detached from the inserter assembly part and become lodged in the intervertebral space. | Component design/selection | Consignees will be notified via email and courier and will be instructed to return the affected product. Upon receipt of the recalled trays, the trials and rasps will be removed from the tray and quarantined for evenutal destruction. |  | 1504 | 1504 units | Product was shipped to the following states: AZ, CA, CO, FL, GA, ID, IN, KY, LA, MD, MI, MO, NC, NV, NY, PA, PR, TX &amp; WV.Product was also shipped to the following countries: Australia, England, Germany, Italy, Mexico, Portugal &amp; United Kingdom. | Open, Classified |  |  | 3067
Z-0112-2017 | 75018 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Sagittal Blade 18.0X0.97X90MM, Model Number: 6118-097-090 |  |  | August 17, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the Sagittal Blade 18.0X.97X90MM due to the potential for a sterile barrier breach in the product's packaging. | Packaging | Stryker Instruments is voluntarily recalling the Sagittal Blade 18.0X.97X90MM, product number 6118-097-090, lot number 16029017, due to a breach in the packaging causing a breach in sterility. There is a potential for a bone or soft tissue infection which may require medical or surgical intervention. Customer notices were issued via certified mail on 08/17/2016. Actions to be taken by the Customer/User:1.Immediately review this Recall Notification. 2.Check all stock areas and/or operating room storage to determine how many Sagittal Blades from the affected lot (listed above) are at your facility. Quarantine and discontinue use of the recalled Sagittal Blades.3.Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if customers dont have any affected product on hand. Note: Customer signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification.4.If customers have further distributed this product, please forward the recall letters and the attached BRF to all affected locations. Please indicate each location on the BRF. 5.Fax the completed BRF to Stryker Instruments Regulatory Department at 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com.6.If the BRF for customer facilities indicates that recalled product is currently on hand, a FedEx label will be emailed to you. This shipping label should be used to return recalled product. 7.Upon receipt of the recalled product, replacement product will be shipped to customer accounts.For questions regarding this recall, please contact Stryker Instruments: Kelly Jo Davis, 269-389-2921, kellyjo.davis@stryker.com |  | 225 | 225 | US Nationwide in the states of RI, WA, CA, CO, WI,FL, ID, ND, MD,TXForeign: Australia, New Zealand, Netherlands, France, Italy | Terminated | March 29, 2017 | 224 | 225
Z-0122-2017 | 74918 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | 25 gauge VersaVit 2.0 Vitrectomy Cutter, Product/Part Number 70935, packaged in a Tyvek/Mylar double pouch configuration. |  |  | August 11, 2016 | 2016 | Synergetics Inc | Vitrectomy cutters have been found not to be operating within the correct pressure pulse per specification. | Other | A customer notification letter was sent out on 8/11/2016 to the hospital/surgical center level. Customers were asked to fill out and return a recall acknowledgment form and send back any affected devices. |  | 9 | 9 units | Devices were distributed to KY, VA, TN, OH, IN, CA, NJ, NY, TX, SC, AL, MS, MI, and UT. Foreign distribution to the UK, France, Mexico, and India. | Terminated | April 27, 2017 | 259 | 408
Z-0111-2017 | 75237 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Distal Femoral Augment 75 X 5 RL/LM |  |  | August 09, 2016 | 2016 | Zimmer Biomet, Inc. | Contents of the package do not match the product labeling.Product is labeled: Vanguard Distal Femoral Augment 75 X 5 RL/LMCarton contains: Vanguard Distal Femoral Augment 75 X 5 LL/RMRisks:Adverse health outcomes that could occur as a result of this event are a delay less than 30 minutes and a different size augment being used to complete the procedure.As the augment will not assemble with femoral component due to its shape, it would be easily identified during surgery. | Labeling mix-ups | On 8/9/2016, URGENT MEDICAL DEVICE RECALL REMOVAL notifications were sent to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notice you have further questions or concerns please call 1-800-348-9500 ext. 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of operating hours will receive a call back the next business day. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. |  | 5 | 5 | Worldwide Distribution - US Nationwide in the states of NJ, MI, FL and the countries of Sweden and Canada. | Terminated | May 17, 2017 | 281 | 5
Z-0128-2017 | 75250 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | CD79b PE (ASR) , Catalog number 335816 ; |  |  | September 21, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | CD79b (SN8) PE is contaminated with a FITC conjugate of unknown antigenic specificity | Under Investigation by firm | BD Life Sciences sent an Urgent Product recall letter dated September 2016 on September 21, 2016 ,to all affected customers via UPS and/or e-mail. Customers were instructed to inspect their inventory to determine if they have any of the affected lots and to discard the affected lots, and BD will issue replacement product. Customers were also instructed to complete the attached Customer Response Form whether or not they have any affected product. Customers needing further assistance should contact BD Customer Support at 855-236-2772 (prompt 3). For customers outside the US, contact their local BD Biosciences representative or distributor.For questions regarding this recall call 408-954-6307. |  | 14 | 14 units | Nationwide Distribution to CA, MN, OH and TX. | Open, Classified |  |  | 14
Z-0044-2017 | 75053 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Willy Rusch GmbH Tracheostomy Tube Set |  |  | August 26, 2016 | 2016 | Teleflex Medical | The connector may disconnect from the tracheostomy tube during use. | Use error | Teleflex sent an Urgent Field Safety Notice dated August 26, 2016, to all affected consignees. The letter requested that consignees cease use and distribution of stock, quarantine immediately, and return the product. Also, the letter requested a sub-recall if the product had been further distributed. The letter included an Acknowledgement Form which is to be returned. Customers with questions were instructed to contact their local sales representative or Customer Service.For questions regarding this recall call 610-378-0131. |  | 1095 | 1095 units | Worldwide Distribution: US (nationwide) to CA and countries of: Austria, Belgium, France, Germany, Italy, Japan, Poland, and Turkey. | Terminated | March 21, 2017 | 207 | 1095
Z-0138-2017 | 74542 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | OER-Pro Endoscope Reprocessor, Endoscope washer/disinfector |  |  | June 23, 2016 | 2016 | Olympus Corporation of the Americas | New warning requiring Users to conduct duodenoscope precleaning and manual cleaning even when using an Automated Endoscope Reprocessor (AER) that has instructions that may indicate a user could forego certain steps in precleaning and manual cleaning of the endoscopes. | Under Investigation by firm | Olympus America Inc. (OAI) mailed a letter to customers informing them that they are issuing an updated Operation Manual for the OER-Pro Automated Endoscope Reprocessor.The following actions were asked to be taken:-Olympus has discontinued previously distributed copies of the OER-Pro Operation Manual. - Inspect your inventory of Operation Manuals and discard any existing inventory of OER-Pro Operation Manuals. - Implement use of the enclosed OER-Pro Operation Manual, which contains the new Warning on pages 73 and 80.- Ensure all reprocessing personnel are completely knowledgeable and thoroughly trained on the requirement to perform precleaning and manual cleaning of Olympus duodenoscopes prior to placing the duodenoscopes in the OER-Pro for high level disinfection. - Additional copies of the new OER-Pro Operation Manual can be obtained by contacting our Technical Assistance Center at 1-800-848-9024, option 1, or by indicating on the enclosed questionnaire. Additional Operation Manuals will be mailed to your facility - Please indicate on the enclosed questionnaire that you have received this notification. Fax the completed form to (484) 896-7128. |  | 2686 | 2, 686 units | Distributed to: AK,AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NC,ND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,RI,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY | Open, Classified |  |  | 2686
Z-0289-2017 | 74590 | No | FBK | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Vari-Safe Injection Needle US endoscopy 5976 Heisley Road Mentor, Ohio 44060 |  |  | May 18, 2016 | 2016 | US Endoscopy Group Inc | US Endoscopy is conducting a voluntary product recall of one lot of the Vari-Safe Injection Needle as it was identified that the lot was incorrectly distributed. The lot was assembled with internal catheter components which differ from the original Vari-Safe needle internal catheter components. All components are dimensionally similar to the original Vari-Safe needle internal catheter components and the components have history of use on other US Endoscopy needle products. | Other | US Endoscopy sent an Urgent- (Vari Safe) injection needle recall letter dated July 6, 2016.US Endoscopy requests that you:1.Quarantine the affected Vari-Safe Injection devices in your stock.2. Destroy each affected device by opening the device package, cutting the device catheter with a scissors, and then discarding the device and package.3.Complete the attached Customer Response Card and return to the attention of US Endoscopy Customer Service, via fax 1-440-639-4495 or email orders@us endos copy.com .A no charge replacement will be provided for each unused Vari-Safe Injection Needle device from the affected lot. We apologize for any inconvenience this matter may cause you, and as always, US Endoscopy is dedicated to supporting our products and valued Customers. For further questions please call (800) 769-8226. |  | 36 | 36 units | US Distribution to the states of : OH, OR, WA and Internationally to the country of Italy. | Terminated | May 10, 2017 | 357 | 36
Z-0137-2017 | 75239 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CADstream software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | The software was not presenting the PURE filter user interface message on DICOM transfer studies. If a site starts using PURE filter, the user interface message will not display. This may result in incorrect patient follow-up. | Software design | Merge Healthcare Inc. sent an Urgent Medical Device Recall letter dated January 28, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it.Merge Healthcare is committed to improve efficiencies and enhancing the quality of healthcare worldwide.If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 844 | 844 sites potentially have the affected versions | Worldwide Distribution - USA (nationwide) Distribution. Foreign/military/government distribution was also made. | Open, Classified |  |  | 844
Z-0262-2017 | 71968 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Plastic Kit, Kit number CMP1084(B |  |  | April 22, 2015 | 2015 | Windstone Medical Packaging, Inc. | The Medtronic COVER LIGHT HANDLE DISPOSABLE GLOVES, Part # 13140, of the custom AMS kit were recalled because gloves may contain splits or holes compromising sterility. | Unknown/Undetermined by firm | Aligned Medical Solutions (AMS) sent the Medtronic Urgent Medical Device for Devon Light Glove letter, dated April 23, 2015 to all affected consignees.Consignees were informed that the sterile kits assembled and distributed by AMS contain the recalled Devon COVER LIGHT HANDLE DISPOSABLE Glove (AMS part #13140 or Medtronic part # 571711) due to splits or holes. Consignees were advised to identify any remaining affected kits and follow the over-labeling instruction. At the time the kit is opened for use, any light gloves should be identified and disposed of pursuant to the medical waste policies in effect at your institution. AMS can provide replacement product if needed or you can contact your local Medtronic sales rep for sterile replacements as well.Due to the manufacturing process and finished sterilized packs in inventory, or at your facility, over labeling of inventory and/or future shipment affected by this recall will be labeled with a sticker attached to each pack. This sticker will read:RECALL NOTICEMedtronic announced the recall l of Devon Light Glove contained in this kitafter it was sealed and sterilized. All other components in the kit are notaffected by this recall.1. At the time the kit is opened for use any Light Glove manufactured by Medtronic should be identified and set aside2. The recalled light glove should be rendered unusable to protect against inadvertent use and disposed of pursuant to the medical waste policies in effect at your institution.Consignees should complete the attached Reply Form and fax to AMS at 406-256-9875.For any questions about this recall, please call 321-527-7714.For questions regarding this recall call 404-259-6387. |  | 1164 | 1164 Kits | Nationwide Distribution to California, Florida, Georgia, Illinois, Kansas, Louisiana, Michigan, New Jersey, Nevada, Texas, Virginia, and Wyoming. | Completed |  |  | 104187
Z-0290-2017 | 75242 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | RTI Surgical Streamline OCT Occipito-Cervico-Thoracic Spinal Rod, CoCr, 3.2mm (D) x 240mm (L) |  |  | September 16, 2016 | 2016 | RTI Surgical, Inc. (dba Pioneer Surgical Technology, Inc.) | RTI Surgical Inc. (RTI) is conducting a voluntary recall on two specific lots of the StreamlineÂż OCT StraightRod, CoCr. 3.2mm (D) x 240mm (L) due to the potential for small cracks to be present on the rods. The small cracks on the rods used in implants could cause reduced strength of these affected implants, and lead to patient injury. | Process control | RTI issued e-mail cutsomer notices for their voluntary recall of lots 234829 and 234829 of the StreamlineÂż OCT StraightRod, CoCr, 3.2mm (D) x 240mm (L) on 09/16/2016. RTI Surgical has produced replacement implants that have been verified as not containing any cracks. These implants will be sent to customers once they have identified how many replacementsare needed. Please follow the below instructions.1. Review inventory located in your facility and any other point of use locations that you may be responsible for. Immediately quarantine any implants listed in this recall notice. Please refer to the pictures included in customer notices for the etch location of the lot numbers for these implants:2. Contact the RTI Surgical Field Inventory representative Jennifer Mannor so that replacements can besent and an RGA and return shipping labels can be created if needed.3. Review, complete and return the attached recall reply form to Dan Nelson (dnelson@rtix.com) in accordance with the directions on the business reply form within 5 business days of receipt of this notification.Should customers have any questions please do not hesitate to contact RTI Surgical Field Inventory representative Jennifer Mannor (Email: jmannor@rtix.com Phone: 906-225-5872. |  | 66 | 66 | US Nationwide Distribution in the states of: AL, CA, CO, FL, IN, IL, MI, MO, MS, NC, NY, OK, TN, TX, and VA,. | Terminated | June 01, 2017 | 258 | 66
Z-0296-2017 | 74701 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Endoscopic Articulating Linear Cutter, 38ATW35. VASCULAR/THIN 35MM |  |  | July 15, 2016 | 2016 | Ethicon Endo-Surgery Inc | Quality control identified a component issue where the pinion gear in the device could fail under extreme use cases. If this condition occurs, staples will be formed past the cut line and the device can be opened and removed from the patient. The firing stroke may be interrupted and the knife may not fully return to the home position potentially exposing the healthcare professional to a sharps injury. | Device Design | Ethicon sent an Urgent Medical Device Recall letter dated July 15, 2016, to all affected consignees via UPS next day mail. The letter identified the product, the problem, and the action to be taken by the customer. Consignees were instructed to examine their inventory immediately and quarantine any affected product and return to ENDOPATH. If product subject to the recall was further distributed consignees were instructed to notify the other facility and arrange for return of the affected product. Customer were instructed to complete the Business Reply Form confirming receipt of this letter and fax to Stericycle at 1-888-228-5719 or email to ethicon5182@stericycle.com.Customers with questions were instructed to contact the Ethicon Customer Support Center at 1-877-384-4266.For questions regarding this recall call 513-337-3582. |  | 213 | 213 | Worldwide Distribution - US including AL, AR, AZ, CA, CO,CT, DC, FL,GA, IA,IL,KS,KY,LA,MA,MD, ME, MI,MN,MO, MS, MT,NC, ND, NH,NJ, NM, NV, NY, OH, RI, TN, TX, VA, VT, WA, WI and Internationally to Brazil and Canada. | Open, Classified |  |  | 318
Z-0055-2017 | 75243 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions Twin-Pass .023" Dual Access Catheter, model 5230. Sterilized with Ethylene Oxide. |  |  | September 16, 2016 | 2016 | Vascular Solutions, Inc. | Vascular Solutions is recalling Twin-Pass dual access catheters because there is a potential for excess manufacturing material to remain at the tip of the catheter or within the distal portion of the rapid exchange lumen. | Process control | Vascular Solutions sent an Urgent Medical Device Recall letter via first class mail beginning 9/16/2016. The letter identified affected product, stated the issue, and instructed that product should not be used and to be returned to Vascular Solutions. A response form was requested to be returned. Questions can be directed to Customer service by e-mail at customerservice@vasc.com. |  | 11670 | 11,670 units total | Worldwide Distribution: US (nationwide) including Puerto Rico and country of Italy. | Open, Classified |  |  | 11670
Z-0291-2017 | 75209 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue MX40 Patient Monitor: IntelliVue MX40 WMTS Smart-hopping (1.4GHz) |  |  | September 19, 2016 | 2016 | Philips Electronics North America Corporation | Philips IntelliVue MX40 Patient Wearable Monitor Configuration Setting Disables Generation and Delivery of ECG Alarms to PIIC iX | Software design | The firm, Philips Healthcare, sent a "Field Safety Notice (FSN)" dated 9/14/2016 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to make sure that the MX40 Patient Wearable Monitor local configuration setting 'Alarms On' is configured to its factory default until your software is upgraded; and follow the instructions listed in the letter. A Philips Healthcare representative will contact customers with affected devices to arrange for the installation of the software upgrade.If you need any further information or support concerning this issue, please contact Philips Customer Support at 1-800-722-9377. |  | 9804 | 9804 units | Worldwide Distribution-US (Nationwide) and countries of: Aruba, Australia, Austria, Belgium, Bulgaria, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Ghana, Greece, Hong Kong, Iceland, India, Iraq, Ireland, Italy,Japan, Korea, Republic of, Malaysia, Mexico, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Singapore, Slovenia, South Africa, SpainSweden, Switzerland, Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 13840
Z-0294-2017 | 75425 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Fusion Workstation.; Indicated for the transmission and review of radiological images. |  |  | October 28, 2011 | 2011 | Merge Healthcare, Inc. | After a period of time running Fusion Workstation, the Hounsfield measurement tool will report incorrect values. | Software design | The recalling firm initiated phone calls to affected sites beginning 1/30/2016. |  | 13 | 13 sites potentially hae the affected version | Distribution was made to AZ, CA, ID, IL, IN, MI, MN, and NJ. | Open, Classified |  |  | 13
Z-0298-2017 | 74809 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Neptune Plus (Hemostatic Pad), Part Numbers: 8870-03 |  |  | July 07, 2016 | 2016 | TZ Medical, Inc. | Neptune Plus was mis-labeled and included the statement WITH ANTIMICROBIAL BARRIER . This statement should not appear on label. | Labeling Change Control | TZ Medical notified their consignees by telephone on July 7, 2016.Customers with the boxes with the identified statement, "With Antimicrobial barrier, the firm plans to overlabel the statement on-site. If customer does not have the mis-labeled box upon contact, no action is planned. The firm also determined to overlabel devices on-hand.For questions regarding this recall call 503-639-0282. |  | 1410 | 1410 (141 boxes of 10) | Nationwide Distribution to AL, CA, FL, GA, IL, IN, KY, MO, MS, NC, NV, NY, PA, SC, TN, TX, and WV. | Open, Classified |  |  | 1410
Z-0297-2017 | 75376 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | The PAXWIRE Occlusion Balloon System consists of an occlusion catheter, inflation handle and syringes. It is designed for use in the iliofemoral artery to provide temporary occlusion as well as a path for dilator and/or sheath introduction, removal or repositioning during catheterization procedures. |  |  | October 06, 2016 | 2016 | AccessClosure, Inc., A Cardinal Health Company | AccessClosure, Inc. (ACI) is voluntarily recalling the PAXWIRE Occlusion Balloon System device because it was made aware from customers that the device may be removed in the absence of vascular sheath across the arteriotomy. This is a deviation from the firm's standard catheterization technique. | Device Design | Cardinal Health sent an Urgent Medical Device Field Safety Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed complete the attached Field Action Return Response Form as soon as possible indicating the number of devices that have been used as their location and the number that will be returned. Customers should discuss the reimbursement for their purchased and return product with their sales representative, who will coordinate the process. Customers with questions were instructed to call 408-610-6570.For questions regarding this recall call 408-610-6500. |  | 6 | 6 in USA | US distribution to FL and NJ | Open, Classified |  |  | 6
Z-0117-2017 | 75393 | Yes | NIK |  |  | Cardiovascular | PMA | Quadra Assura, Model No. CD3265-40, CD3265-40Q, CD3365-40C, CD3365-40Q. |  |  | October 10, 2016 | 2016 | St Jude Medical Inc. | St Jude Medical is recalling implantable cardiac defibrillator (ICD) and cardiac resynchronizationtherapy defibrillator (CRT-D) devices because they are subject to premature battery depletion due to lithium cluster with the battery causing a short circuit. | Device Design | St. Jude Medical, Inc., sent a customer letter to all affected customers to inform them that St. Jude Medical recently announced a global medical device advisory for a subset of their company's Fortify, Fortify Assura, Quadra Assura, Unify, Unify Assura and Unify Quadra implantable cardiac defibrillator (ICD) and cardiac resynchronization therapy defibrillator (CRT-D) devices. The letter informs customers that the cause for Premature Battery Depletion is attributed to the formation of lithium clusters within the battery causing a short circuit. There is a possibility that affected devices may lose battery power within days. Thus far, there have been two (2) deaths and ten (10) serious events (fainting) that may have been associated with premature battery depletion. Customers with questions were instructed to contact www.sjm.com/batteryadvisory, (866) 915-5065. |  | 193954 | 193,954 units | Worldwide Distribution - US (nationwide) and OUS | Open, Classified |  |  | 603524
Z-0299-2017 | 75224 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Keystone Dental PrimaConnexÂż Straight Implant TC, RD Internal Connection 4.0 x 11.5mm. |  |  | September 15, 2016 | 2016 | Keystone Dental Inc | Small diameter (SD) cover screw was packaged with the regular diameter implant and do not match the implant diameter (RD) | Packaging process control | Keystone Dental initially called all customers on September 15, 2016. Forms were completed to document these calls. In addition a letter has been sent via Federal Express/or hand delivered on September 19, 2016, to all distributors and customers. Customers with questions were instructed to contact Customer Relations.For questions regarding this recall call 781-328-3390. |  | 55 | 55 units | Worldwide distribution - US (Nationwide Distribution to CA, CO, IL, MA, MN, ND, NY, NV, PA, and VA.) and Italy | Terminated | January 03, 2017 | 110 | 55
Z-0300-2017 | 75167 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation System (Product Code ENGEN) configured with Thermo Scientific Centrifuge Module and TCAutomation Software Version 3.6.1 and Below; IVD. --- Thermo Scientific Centrifuge Module has Product Code |  |  | September 07, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho confirmed that it is possible for the mis-association of sample identification and assay results due to an error with Thermo Fishers centrifuge module. It is possible for mis-associated assay results to be reported from the laboratory prior to the identification of the Cross Check error on the enGen Laboratory Automation System. Assay results associated with the wrong patient and reported out of the laboratory could lead to inappropriate intervention with the potential for serious injury to the patient. | Software design | The firm, Ortho Clinical Diagnostics, sent an "URGENT PRODUCT CORRECTION NOTIFICATION" Letter (Ref. CL2016-183, dated 9/7/2016) was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all enGen customers with centrifuge modules to inform them of the issue and advise them to check samples associated with cross check errors as instructed in the operator manual, and perform the preventative maintenance and cleaning procedures for that centrifuge module as provided in the maintenance manual and user guide. The customers were also instructed to Complete and return the Confirmation of Receipt form by 16 September 2016 via Fax:1.888.557.3759 or 1.585.453.4110 or email:RA-OCDUS-CONFIRMAD@ORTHOCLINICALDIAGNOSTICS.COM. For questions regarding this notification, please call the Ortho Care Solutions Center any time at 1-800-421-3311. Foreign affiliates were informed by e-mail on 9/7/2016 of the issue and instructed to notify their consignees of the issue and appropriate actions to take. |  | 60 | 60 (US: 24; Foreign: 36) | Worldwide Distribution-(Nationwide) and foreign distribution to the following countries: Belgium, Bermuda, Canada, Chile, Denmark, France, Italy, Mexico, Norway, Portugal, Sweden, Taiwan and United Kingdom. | Open, Classified |  |  | 60
Z-0301-2017 | 75384 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FemTouch delivery system includes the FemTouch Kit (AC-1007415) and FemTouch" Tip (SPSA-20002590) |  |  | September 25, 2016 | 2016 | Lumenis Limited | The cleaning instructions provided in FemTouch Operator's Manual might be ineffective for a fully assembled device and may lead to a potential for cross-contamination. | Software Manufacturing/Software Deployment | On 9/25/16 consignees were notified via an Urgent - Field Safety Corrective Action letter with a FemTouch disassembly tool kit as well as FemTouch disassembly instructions. The instructions describe the simple disassembly process to be performed as part of the Pre-cleaning phase and the process for reassembling the device prior to re-use. It is recommended that customers clean and sterilize all of their handpieces according to the pre-cleaning disassembly process before each use. Consignees were asked to complete and return a verification form. Questions about the FSCA, should be directed to the Lumenis field correction administrator, Rafi Barel at Rafi.Barel@lumenis.com. |  | 69 | 69 units | Worldwide Distribution -- USA, Italy, Russian Fed., Australia, United Arab Emirates, Spain, United Kingdom, India, Kuwait, and Belarus. | Terminated | April 11, 2017 | 198 | 69
Z-0302-2017 | 75265 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | ex-TEM Assay for ROTEM delta Thromboelastometry System. Art. no.: 503-05 US, ex-tem. |  |  | September 13, 2016 | 2016 | TEM Systems Inc | Customers complained of low/empty reagent vials. | Nonconforming Material/Component | Customers were notified of the recall by an Urgent Important Customer Information letter on 9/13/2016. The letter described the device and the problem. Customers were told to stop using the product and to call for replacement product. The letter included a Confirmation of Compliance Letter which was to be completed and returned. |  | 647 | 637 boxes x 10 vials | Nationwide Distribution -- AR, CA, DC, GA, HI, ID, IL, IN, LA, ME, MD, MA, MI, MN, MO, NE, NJ, NM, NY, NC, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, and WV. | Open, Classified |  |  | 647
Z-0305-2017 | 63994 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq One Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids |  |  | February 04, 2013 | 2013 | Hospira Inc. | Symbiq Infusers have the potential to experience a white screen during titration of a Keep Vein Open (KVO) delivery. This can allow the clinician to select the "NEXT" button prior to completion of the programming. If this error occurs, a delay in therapy could occur and potentially result in significant injury. | Software design | Hospira sent an Urgent Device Field Correction notice dated February 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the steps provided in the letter while Hospira is developing a software change. Customers were asked to complete the attached reply form and return it via fax to the number on the form. Customers with questions were instructed to contact Stericycle at 1-888-201-9332.For questions regarding this recall call 224-212-4892. |  | 35596 | 35,596 devices in total | Worldwide Distribution - USA (nationwide) and Canada | Open, Classified |  |  | 35596
Z-0304-2017 | 75408 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | 4205 Journey Rolling Walker |  |  | October 01, 2016 | 2016 | Nova Ortho-Med Inc | Nova Medical is recalling the Journey Rolling Walker because the screws used to hold the black side support pieces to the frame can become loose over time. This can result in the locking nut coming off and the screw falling out. | Component design/selection | A customer notification letter dated 10/1/16 was sent to customers to inform them that the screws used to hold the black side support pieces to the frame can become loose over time. This can result in the locking nut coming off and the screw falling out. This allows the frame support piece to detach from the walker and causes instability resulting in possible injury. The letter informs the customers of the root cause and the solution. Customers are instructed to record the serial #s of the product that have been corrected and email them to NOVA Customer Service (patricias@novajoy.com). Customers can also 1-800-557-6682 x131. |  | 1380 | 1,380 units | US Nationwide Distribution in the states of: CA, IN, HI, CT, UT, WA, FL, VA, WI, OH, MN, KS, MD, IL, MO, WY, MT, ME, PA, TN, MI, OR, NY, NJ, ID, NC, IA, OK, AR, VT, SC, TX | Open, Classified |  |  | 1380
Z-0303-2017 | 75404 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | ReTurn 7500/7500 |  |  | April 15, 2014 | 2014 | Handicare USA, Inc. | During the period of April 2009  May 2011, Handicare AB used a wing handle in which the screw and handle, after repeated tightening, may slide apart. This causes a small amount of play in the attachment of the ladder, which in turn results in high load at the attachment holes of the ladder. Eventually, this may cause mechanical fatigue in the material of the ladder. | Nonconforming Material/Component | Handicare sent a Field Safety notice dated April 21, 2014, to all affected customers. The notification instructed customers to examine the affected ReTurn units for signs of wear and to complete the response form that was mailed with the customer notification letter indicating whether additional parts were required to complete the field correction. For further questions, please call (717) 733-1444. |  | 76 | 29 units , plus 47 wing handles were shipped as spare parts | US Distribution to the states of : NY, NJ, NV, NE, IL, CT, PA and ME | Completed |  |  | 76
Z-0094-2017 | 75186 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche/Hitachi MODULAR Analytics Combination Systems |  |  | July 15, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Roche has received four reports of injury caused by the sharp edge of the center guide rail of ananalyzers sample rack tray. The cleaning procedure for the sample rack tray has been revised to clarify the cleaning process. | Device Design | On 7/15/2016, URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, an updated cleaning procedure, and instructions for responding to the formal recall notification. Please contact the Roche Support Network Customer Support Center at 1-800-428-2336 if you have questions about the information contained in this UMDC.On 10/5/2016, URGENT MEDICAL DEVICE CORRECTION UPDATE notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The notification updated the 7/15/16 notification to include additional affected systems and to provide updated cleaning instructions for all systems via the enclosed Operators Manual addendum. |  | 9014 | 9014 in total | US Nationwide Distribution | Open, Classified |  |  | 9014
Z-0311-2017 | 75375 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | PhoenixSpec Calibrator &amp; Kit is used to calibrate the PhoenixSpec &#38; Phoenix AP Instruments. The instruments are then used to measure the turbidity of microbial suspensions equivalent to McFarland standards 0.10 to 4.50. |  |  | June 01, 2016 | 2016 | Becton Dickinson &amp; Co. | BD has determined that the BBL PhoenixSpec Calibrator kits referenced in this recall from the following lots (5364776, 5345796 ) are labeled with an incorrect expiration date (year). The expiration date printed on the individual calibration tubes within the kit are correct but difficult to read on a portion of the tubes. | Process control | All US customers were contacted via letters sent via UPS. Ex-US BD sites will contact customers in their local regions. The firm will use an acknowledgement form included with the follow up letter that is mailed or faxed back to BD will be reconciled with the customer list. BD will contact 100% of our customers that have not responded to the communication to verify acknowledgement |  | 20 | 20 | BD has contacted all 31 US customers and 22 Ex-US sites which received product from the associated catalog numbers.US: AL, CA, CO, KY, MD, MI, MS, NE, NJ, OH, PA, SC, SD, TX, VA, WA and WI.Ex-US: Argentina, Europe, Canada, China, Ecuador, India, Japan, New Zealand, Peru, Singapore, Thailand and Vietnam4 US government sites were contacted | Completed |  |  | 114
Z-0313-2017 | 74772 | No | MJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Microbiology | 510(K) Exempt | VITROS Immunodiagnostic Products HBsAg Controls, Product Code 6800598 |  |  | June 22, 2016 | 2016 | Ortho-Clinical Diagnostics | The signal/cutoff (s/c) results may increase throughout the open-reconstituted timeframe of the allowed 5 days. Increased results may occur on both the negative and positive Controls. The positive Control, although showing elevated results, was still within acceptable performance. | Under Investigation by firm | The firm, Ortho Clinical Diagnostics, distributed "Important Product Correction Notification" letters (Ref. CL2016-129) dated 6/22/2016 and Confirmation of Receipts to their customers via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only). Foreign affiliates were informed by email (on 22 June 2016) of the issue and instructed to notify their consignees of the issue and required actions. All customers who received the affected product were instructed to immediately discontinue using and discard the affected lots. Forward a copy of the notification if you have provided the affected lot(s) outside of your facility. Post a copy of the notification by each VITROS System that processes VITROS HBsAg Controls and Complete and return the Confirmation of Receipt-Response Required form via fax to: 1.888.557.3759 or 1.585.453.4110 or scan to PDF and email to: RA-OCDUS-CONFIRMAD@ORTHOCLINICALDIAGNOSTICS.COM no later 7/1/2016. (Note: If you DO NOT have an alternate lot, it is acceptable to continue using the affected lot(s) following these instructions: 1) Control results must be within acceptable limits on the day of reconstitution (Day 1). 2) Store open-reconstituted vials frozen for up to 4 weeks. The firm recommends that customers aliquot the open-reconstituted material prior to freezing so that multiple aliquots of the same vial can be used. 3) Upon receipt of your replacement order, discard all remaining inventory. Customers with questions please call the Ortho Technical Solutions Center at 1-800-421-3311. |  | 1621 | 1,621 units | Worldwide Distribution-USA (nationwide) and countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain and Poland. | Open, Classified |  |  | 1621
Z-0316-2017 | 75142 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Safil QUI + UNDY 5/0 (1) 45CM DSMP13 Suture |  |  | August 01, 2016 | 2016 | Aesculap Implant Systems LLC | There is a potential for the needle to detach from the suture thread during surgery. If the needle becomes detached the defective suture will need to be replaced causing the wound to be sutured a second time. | Process control | An Urgent Medical Device Recall Notification Letter was mailed to customers on August 11, 2016 asking them to remove and return the affected product immediately. |  | 180 | 180 units | Distributed in Arizona | Open, Classified |  |  | 180
Z-0317-2017 | 75371 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ZimmerÂż Gender Solutions" Patellofemoral Joint Prosthesis |  |  | September 28, 2016 | 2016 | Zimmer Biomet, Inc. | Complaints have been reported that the handpiece is inoperable. The investigation found that the complaints were related to the lack of preventative maintenance contrary to what is recommended in the instructions for use included with the handpiece. An inoperable device may lead to a delay in surgery | Labeling design | On 9/28/2016, URGENT MEDICAL DEVICE FIELD REMOVAL notifications were sent to the affected distributors and Risk Managers / Surgeons via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Risk Managers and SurgeonsRisk Manager Responsibilities: 1. Review this notification and ensure affected personnel are aware of the contents. 2. Assist your Zimmer Biomet sales representative with quarantining all affected product. 3. Your Zimmer Biomet sales representative will remove the affected product from your facility for the initial preventative maintenance and arrange for a loaner device. It is estimated that your handpieces will undergo preventative maintenance sometime between November 2016 and March 2017. 4. Complete Attachment 3  Certificate of Acknowledgement. a. Return a digital copy via email to CPWARFieldAction@zimmerbiomet.com or via fax to 574-372-1683 within three (3) days. b. Retain a copy of the Acknowledgement Form with your field removal records in the event of a compliance audit of your facilities documentation. 5. If after reviewing this notice you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an on-call representative in the case of an emergency. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. Surgeon Responsibilities: 1. Review this notification for awareness of the contents. 2. There are no specific patient monitoring instructions related to this field removal that are recommended beyond your existing follow up schedule. 3. Complete Attachment 3  Certificate of Acknowledgement. a. Return a digital copy via email to |  | 703 | 703 | Nationwide.CANADAHONG KONGKOREASINGAPORECHINAVENEZUELAAUSTRALIAJAPANNETHERLANDSGERMANYINDIATAIWAN | Open, Classified |  |  | 703
Z-0315-2017 | 75473 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da VinciÂż Xi" Surgical System, model number IS4000, A70_P5x with P5 Software; |  |  | October 05, 2016 | 2016 | Intuitive Surgical, Inc. | Intuitive Surgical has identified a software anomaly in the da Vinci Xi P5 software that can result inunexpected master movement and potential instrument tip movement under certain circumstances | Software design | Urgent Medical Device Correction letters were sent on October 12, 2016 by traceable method to all affected customers. Letters notify customers of the issue and provided instructions to prevent it from happening until ISI software can install new software upgrade in 1st Q 2017. |  | 677 | 677 devices | Australia, Austria, Belgium, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Monaco, Netherlands, Norway, Portugal, Puerto Rico, Qatar, Romania, SouthKorea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, United States | Terminated | May 01, 2017 | 208 | 677
Z-0318-2017 | 75332 | No | FDS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Pentax Video Gastroscope is intended to be used with a Pentax video processor (including light source), documentation equipment, monitor, EndoTherapy Device such as a Biopsy Forceps, and other ancillary equipment for endoscopy and endoscopic surgery within the upper digestive tract including the esophagus, stomach and duodenum. |  |  | August 15, 2016 | 2016 | Pentax of America Inc | PENTAX Medical is initiating this field action to provide customers that have purchased the affected endoscopes with the most recent Operation and Reprocessing IFUs. | Labeling Change Control | Pentax sent as "Urgent Field Correction" and Field Correction Response Form dated August 15, 2016 to their affected customers. |  | 3436 | 3,436 (US) and 14,488 (OUS) | US Nationwide and Japan | Open, Classified |  |  | 3436
Z-0319-2017 | 75090 | No | LKQ | Cytomegalovirus serological reagents. | Microbiology | Microbiology | 510(k) | IMMULITE/ IMMULITE 1000 Systems CMM CMV IgM Assay, REF/Catalog Number LKCM1(D) - 100 tests (US), Siemens Material Number (SMN) 10370301 (US), Unique Device Identification Number (UDI) 00630414985718 (US); and REF/Catalog Number LKCM1 - 100 tests (OUS); SMN 10381296 (OUS), UDI 00630414963747 (OUS); IVD --- Made in UK Siemens Healthcare Diagnostics Products Ltd., Glyn Rhonwy, Llanberis, Gwynedd, LL55 4EL UK |  |  | August 30, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Multiple customers reported seeing an increase in imprecision with some patients resulting in high % CVs with the IMMULITE 2000/2000XPi CMV (IgM) assay starting with reagent lots 259 and above. | Under Investigation by firm | The firm, Siemens, issued an "Urgent Medical Device Recall" Letter (IMC16-22.A.US, dated 8/30/2016) in the US and an "Urgent Field Safety Notice" (IMC16-22.A.OUS, dated August 2016) Outside the US (OUS) starting on 8/30/2016 The notices instructed customers to discontinue use of and discard the affected kit lots listed in the customer notice; review your inventory of the products to determine your laboratory's replacement needs; review the UMDR and or the UFSN with their Medical Director, In the US, the method of distribution is FedEx and Outside the US the distribution is determined at the country level; complete and return the Field Correction Effectiveness Check Form via Fax to: Customer Care Center at (312) 275-7795 within 30 days, and retain this letter with your laboratory records, and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 377 | US: 5 kits; Foreign: 372 kits | Worldwide distribution- US (Nationwide) and countries of: United Arab Emirates, Afghanistan, Angola, Argentina, Australia, Bangladesh, Belgium, Bulgaria, Bahrain, Bolivia, Brazil, Belarus, Switzerland, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Germany, Dominican Republic, Algeria, Ecuador, Egypt, Spain, France, Great Britain, Greece, Guatemala, Honduras, Hungary, India, Iraq, Italy, Jordan, Canary Islands, Kyrgyzstan, Rep. Korea (S), Kazakhstan, Lebanon, Lithuania, Luxembourg, Latvia, Macedonia, Mexico, Malaysia, Nicaragua, Netherlands, Norway, Nepal, Peru, Pakistan, Poland, Portugal, Paraguay, Qatar, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, El Salvador, Syria, Turkmenistan, Tunisia, Turkey, Uruguay, Uzbekistan, Vietnam, Serbia,and South Africa. | Open, Classified |  |  | 10336
Z-0314-2017 | 74786 | No | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | SimpliRED D-Dimer, Product Code DSRK4 |  |  | June 01, 2016 | 2016 | BBI SOLUTIONS OEM LTD | A component of the kit is showing a sensitivity performance detected in stability testing at 15 months of an 18-month shelf life. Qualitative results may be affected. False negative results are possible which may incorrectly influence the clinical diagnostic decision. | Nonconforming Material/Component | BBI Solutions sent an Urgent Medical Device Recall letter dated June 30, 2016, to all affected customers. Customers were advised to immediately discontinue use of the product, consider the requirement for review of reported test results, complete and return the customer acknowledgement form to BBI Solutions and to complete and return the customer reconciliation form to the distributor. Customers who have forwarded product to other sites, should inform those customers of the product recall and provide them with a copy of the notice. Customers with questions were instructed to call Tel: +44(0)1382 564093. |  | 420 | 420 units (183 units quarantined) | US distribution to CT and IN | Open, Classified |  |  | 420
Z-0114-2017 | 75211 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | VIA 27 Microcatheter, Part No. FG27-154-01, Model No. VIA-27-154-01 |  |  | September 23, 2016 | 2016 | Sequent Medical Inc | Sequent Medical is recalling the VIA 27 Microcatheter because it may not adequately re-sheath the interventional device into the catheter as designed. | Under Investigation by firm | Sequent Medical sent an Urgent Medical Device Recall letter dated September 23, 2016, to all affected customers. The letter informed customers that Sequent Medical is recalling the VIA 27 Microcatheter because it may not adequately re-sheath the interventional device into the catheter as designed. The letter informs the customers of the risk to health and actions to be taken. Customers were instructed to send all confirmation forms and questions to www.customerserviceus@sequentmedical.com. Customers with questions should call (949) 830-9600x130.For questions regarding this recall call 949-923-9400. |  | 618 | 618 units | Worldwide Distribution - US including NY, MN, FL, MO, MD, TN, NC, OH, CA, KS, TN, WV and Internationally to Czech Republic, Germany, UK, France, Netherlands, Italy, Spain, Finland, Switzerland, Slovenia, Hungary, Northern Ireland, Turkey, and Denmark, | Terminated | May 01, 2017 | 220 | 618
Z-0321-2017 | 75062 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | TD60 display Telemetry Transmitter, SHENZHEN Mindray Bio-Medical Electronics Co., LTD, Mindray Building, Keji 12th Road South, Hi-Tech Industrial Partk, Nanshan, Shenzhen, 518057 P.R. China |  |  | September 08, 2016 | 2016 | Mindray DS USA, Inc. dba Mindray North America | Mindray has become aware that the TD60 Display Telemetry Transmitter used with the BeneVision Central Station may crack on the bezel surrounding the display and also the transmitter may not power on or will power on but display an error message. | Device Design | Mindray sent a "Medical Device Action" letter dated September 8, 2016 via UPS to their affected customers. |  | 1278 | 1,278 units | US and Canada | Open, Classified |  |  | 1278
Z-0322-2017 | 75124 | No | GAH | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Visualase Visualization Stylets PN 020-2301. The individual components within the Visualase Body Accessory Kit are packaged in Tyvek pouches. The packaged devices are then placed into a cardboard box which is labeled with the Body Accessory Kit label. Labeled sterile. |  |  | September 01, 2016 | 2016 | Medtronic Navigation, Inc. | Medtronic Navigation, Inc. announces a voluntary field action for the Visualase Visualization Stylets and Visualase Body Accessory Kit because when the Visualization Stylet is placed in the magnetic field of a magnetic resonance imaging system it pulls and/or bends toward the field. | Mixed-up of materials/components | Medtronic sent an Urgent Medical Device Recall letter dated September 2, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their stock for the affected product and if any are found contact Medtronic to arrange for a no charge replacement. Customers were asked to complete the Customer Confirmation form attached to the letter and return it to Medtronic at RS.NavFCA@medtronic.com or by Fax to 651-367-7075. Customers with questions were instructed to contact Medtronic Technical Services at 1-800-595-9709.For questions regarding this recall call 720-890-3200. |  | 100 | 100 total units | Worldwide Distribution - US including CA, GA, FL, and TX. Internationally to Australia. | Terminated | March 27, 2017 | 207 | 100
Z-0325-2017 | 75214 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Empowr PS Knee System Box Cut Guide; Used to make the housing resection for the post of the posterior stabilized implant. |  |  | September 19, 2016 | 2016 | Encore Medical, Lp | It has been discovered that during Empowr PS Total Knee replacement, while chiseling for bone removal in the box area, femoral chipping has occurred in surgery. The current design of the captured box cut guides can create an instance where the surgeon is not fully completing the box cut on the side. | Under Investigation by firm | The recalling firm sent a recall notification letter to affected consignees on 9/19/16. The field safety notice provided a newly revised Surgical Technique to educate users on proper use of the device. |  | 436 | 436 units | US | Open, Classified |  |  | 436
Z-0324-2017 | 75221 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Reprocessed Pressure Tourniquet Cuff (PTC) |  |  | September 19, 2016 | 2016 | Stryker Sustainability Solutions | Issue with the process for detection of leaking Pressure Tourniquet Cuffs (PTC). Leaking PTC devices may result in excessive blood loss requiring surgical or medical intervention or systemic leakage of local anesthetic during Bier Block requiring medical intervention. | Device Design | Stryker Sustainability Solutions sent an Customer Notification letter September 19, 2016. Customers were instructed to discontinue use of the affected products. The letter stated to follow the link below to complete the Recall Effectiveness Check Form and indicate if any devices containing expiration dates from May 2016-May 2018 remain in inventory. This form must be completed even if no affected product is found.Use the following reference number as you complete the form: [Ship to Account #]If the form indicates that any affected devices remain in inventory, a prepaid shipping label will be issued for the return of the product. Credit for all affected devices returned will be issued to the facility..Adverse reactions or quality problems experienced with the use of this product may be reported to:o Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433 X5555o http://www.stryker.com/productexperience/o The FDAs MedWatch Adverse Event Reporting program either online, byregular mail, or by fax. |  | 1695760 | 1,695,760 | U.S. distribution nationwide.No foreign distribution | Open, Classified |  |  | 1695760
Z-0326-2017 | 75197 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Product Number 3102071 Ceiling Bracket, Arch 71 White with use on ceiling |  |  | August 09, 2016 | 2016 | Hill-Rom, Inc. | Ceiling Bracket 71 wasshipped with a larger center hole. A potentiallyhazardous situation exists that if the Ceiling Bracket 71with a larger center hole is used for a concrete ceilingUltra System installation with the M8 screw and the safeworking load is exceeded, the system could fall. | Mixed-up of materials/components | On 8/9/2016, URGENT: MEDICAL DEVICE CORRECTION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any questions regarding this Field Safety Correction, please contact Hill-Rom Technical Support at 1-800-445-3720, your distributor or your Hill-Rom representative. |  | 188 | 188 | WANCOHPAMIAustraliaSweden | Terminated | March 24, 2017 | 227 | 188
Z-0086-2017 | 74018 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | 15 g R.K. Epidural Needle (TW), Catalog #/REF #100-1415 |  |  | April 22, 2016 | 2016 | Epimed International, Inc. | There is a possibility that the needles may have the potential to skive a catheters outer coating. If this were to occur, there exists the risk that a portion of the catheter could remain in the patient, which may require otherwise unnecessary medical invention. | Employee error | On April 22, 2016, Epimed sent an Urgent Medical Device Recall notices and Recall Return Response Forms to their customers via courier service. Customers were advised to examine all inventory and quarantine product subject to recall immediately. Once the product has been quarantined, please complete and return the enclosed response form as soon as possible by fax to: 518-725-0207, Attn: Quality Assurance. In addition, if you find any product that corresponds to the lot numbers given, please contact the Customer Service Department so that Epimed can arrange for replacement/credit (518-725-0209 or 518-514-1891).Further questions, please call (518) 725-0209. |  | 576 | 576 units | US Distribution to the states of : AL, AK, FL, IN, TN &amp; TX | Open, Classified |  |  | 576
Z-0327-2017 | 74879 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Affected Pain Control Tray Component (5% Lidocaine Hydrochloride and 7.5% DExtrose Injection, USP) |  |  | July 11, 2016 | 2016 | B. Braun Medical, Inc. | B. Braun Medical Inc. is voluntarily recalling specific lots of their Custom Spinal Anesthesia Tray which contains a drug component, 5% Lidocaine Hydrochloride and 7.5 % Dextrose Injection, USP that Hospira Inc. has recalled because the drug product does not meet the specification for color throughout shelf life. | Under Investigation by firm | The firm, B.Braun, mailed to customers a "VOLUNTARY DRUG RECALL NOTIFICATION" letter dated July 12, 2016. The letter described the product, problem and actions to be taken and to inform the customers of a drug recall issued by Hospira, Inc. The Customers were instructed to determine their current inventory of the affected lots Do not destroy any affected product), complete and return "Product Removal Acknowledgement" form via fax to: B.Braun Medical Inc., Quality Assurance department at (610) 849-1197 or email to PA_QualityAssurance.BBMUS_Service@bbraun.com within two weeks of receipt, even if you have no inventory. A BBMI Customer Service Representative will contact you to provide instructions for handling the affected product and arrange for return to BBMI.Should you have any questions or concerns regarding the attached information, please contact our Customer Support Department at (800) 227-2862. |  | 1704 | 1704 units | US Distribution to: AL, IL, OK, MN, MA and WI. | Terminated | March 15, 2017 | 247 | 1704
Z-0329-2017 | 75066 | No | MRV | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Glucose Tolerance Beverage- Orange 50 gram Concentration, Liquid 4/Pack of 6 bottles, 24/case |  |  | August 24, 2016 | 2016 | Azer Scientific, Inc. | Out of Specification Microbiological test results | Unknown/Undetermined by firm | Azer Scientific mailed an Urgent Medical Device Recall letter dated August 30, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to discontinue use of product and discard to avoid further distribution or use and to notify their customers/end users of the recall. Customers were asked to acknowledge receipt of the letter by completing the attached form and faxing it to (610) 901-3046. Customers with questions were instructed to call (610) 524-5810. |  | 426 | 426 cases | Nationwide Distribution | Open, Classified |  |  | 1567
Z-0332-2017 | 75477 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Xhibit Central Station, Model 96102. |  |  | October 18, 2016 | 2016 | Spacelabs Healthcare Inc | The firm received reports of telemetry SpO2 numerics dropping off the Xhibit Central display. Desaturation, high, and low limit alarms work normally. | Software design | The firm sent the Urgent - Medical Device Correction letter, dated 18 October 2016, via priority service, return receipt requested, to all U.S. customers on 10/18/16. The firm will e-mail a customer letter (translated as necessary) to all international subsidiaries and distributors of record on October 25, 2016.Xhibit Central Stations used without Xhibit Telemetry Receivers, are not affected.Customers were asked to circulate the letter to all person affected and place a copy with the appropriate monitor user manual until corrections have been implemented. Spacelabs Healthcare will contact affected consignees to schedule a convenient time for Spacelabs to update the affected Xhibit Central Station(s) with a new software version 1.1.6 at no cost.Customers with questions can call Technical Support at 1-800-522-7025 and select 2. |  | 348 | 348 units (292 units in the US and 56 units outside the US) | Worldwide Distribution -- USA, to AL, CA, CT, FL, GA, IA, ID, IL, IN, KY, MI, MN, MS, NC, NE, NJ, NY, OH, OK, PA, RI, SC, SD, TX, VA, WA, WY, and Puerto Rico; and, to the countries of AUSTRALIA, CANADA, CHILE, CZECH REPUBLIC, FRANCE, MEXICO, NETHERLANDS, PANAMA, SAUDI ARABIA, SWITZERLAND, TAIWAN, and UNITED KINGDOM. | Terminated | April 03, 2017 | 167 | 348
Z-0331-2017 | 75372 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | HANDS-OFF Infusion Port Thermodilution Catheter consists of the Arrow IPTD thermodilution catheter enclosed in a contamination shield (Arrow Cath-Gard) with integral flushing/balloon test chamber, enabling the practitioner to prepare, test, and insert the catheter without exposing it to external contamination. |  |  | September 02, 2016 | 2016 | Arrow International Inc | Labeling inconsistency | Labeling mix-ups | Arrow sent an Urgent Medical Device Recall Notification letter dated September 20, 2016, to all affected customers via FedEx 2-day air. The letter identified the problem and provided instructions to immediately discontinue use and quarantine any products with the associated lot numbers indicated in the letter. If product was found, customers were asked to complete the Recall Acknowledgement Form and a Customer Service Rep will issue a Return Goods Authorization (RGA) Number for the product's return. Disposition of recalled product will be scrapped. For further questions, please call (610) 378-0131 |  | 1361 | 330 units in US and 1,031 units OUS | Nationwide distribution | Open, Classified |  |  | 1361
Z-0333-2017 | 75447 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Guider Softip XF; |  |  | October 05, 2016 | 2016 | Stryker Neurovascular | Firm became aware of potentially defective Guider Guide Catheters due to a manufacturing non-conformance that could lead to hub leaks in the resulting product. | Employee error | Stryker Neurovascular sent an Urgent Medical Device Voluntary Recall letter dated October 5, 2016, to all affected customers by Federal Express. Letters advised of the reason for recall and models and devices affected and requested that all affected devices in inventory be returned to Stryker. Customers were advised to circulate the Field Safety Notice internally to all interested/affected parties and to complete the attached customer response form and return the form to their nominated Stryker Representative or to NVFieldActions@stryker.com.For questions regarding this recall call 510-413-2106. |  | 1478 | 1478 devices, all models and lot numbers. | Worldwide Distribution - US including AK, IL, OH and Internationally to:ARGENTINAAUSTRALIAAUSTRIABELGIUMBRAZILCANADACHILECZECH REPUBLICECUADOREGYPTFRANCEGERMANYGREECEHUNGARYINDIAIRANISRAELITALIAJAPANMEXICONETHERLANDSNORWAYPERUPHILIPPINESPOLANDRUSSIASERBIASINGAPORESLOVAKIASOUTH AFRICASOUTH KOREASPAINSWEDENSWITZERLANDTHAILANDTUNISIATURKEYUKURUGUAYVIETNAM. | Terminated | March 23, 2017 | 169 | 1478
Z-0354-2017 | 63993 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq Two Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids; Product List Number 16027 |  |  | February 05, 2013 | 2013 | Hospira Inc. | Hospira has received reports of Malfunction S205 Backup Battery Failure alarms in Symbiq v3.13 pumps. This issue to an incorrect installation of components on some of the Power Supply Controller (PSC) Boards distributed prior to September 2012. On pumps where incorrect installation of components has occurred, the backup battery power discharges and subsequently, the battery cannot be charged. Because of the battery depletion, or lack of charge, a visual and audible warning alarm, S205, is displayed. An S205 alarm is intended to notify the customer that the coin call battery, powering the backup buzzer, may need to be charged or replaced. | Process control | Hospira sent an Urgent Field Device Correction letter dated Feburay 5, 2013, to all affected customers.Hospira has implemented a correction to the manufacturing process to address this issue. Symbiq v3.13 pumps that have been shipped from Hospira since September 2012 and pumps that were upgraded to v3.13 in the Field after September 2012 are not impacted by this issue. If you are unsure of when you were shipped Symbiq v3.13 pumps or when your devices were upgraded, contact Hospira.Please complete the attached reply form and return it via fax number or e-mail address on the form, even if you do not currently have the impacted infuser. lease contact Stericycle at 1-866-201-9072 (M-F, Bam - 5pm EST) to obtain additional copies of the reply form.If you have further distributed these devices, please notify your accounts who may have received these infusers from you and ask them to contact Stericycle at 1-866-201-9072 (M-F, 8am- 5pm EST) to receive a reply form. |  | 15285 | 15,285 | Nationwide Distribution and Canada | Open, Classified |  |  | 35596
Z-0352-2017 | 75128 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Rigid Suction Canister 3000 CC and 2000 CC. |  |  | August 10, 2016 | 2016 | Medline Industries Inc | Risk of implosion when the Rigid Suction Canister is subjected to excessive negative pressure during use. | Process control | Medline Industries, Inc. sent an Urgent Recall Immediate Action Required letter dated August, 23, 2016, and destruction form via first class mail to affected customers on August 17, 2016 and August 23, 2016. The firm recommended the following actions: REQUIRED ACTION:1. Immediately check your stock for item number OR220 and OR230 with lot numbers beginning with 9961601xxxx through 9961604xxxx (listed on the original case packaging). Destroy all affected products.If the rigid canisters are NOT in the original case packaging, we are instructing you to destroy them and we will send replacements.2.Please return the completed enclosed destruction form listing the quantity of product(s) destroyed. Even if you do not have any of this product please complete and return the form, as indicated on the destruction form.When we receive your completed destruction form your account will receive replacements.3.If you are a distributor, promptly notify any of your customers that may have received affected product.If you have any questions, please contact 866-359-1704. |  | 7680 | 7,680 units | US Distribution to the following states: CA , CO , FL , GA, GA , IL , MA, MA , MD , MI, MO , NC , NE , NJ , NY, NY , OH, OK , PA, TX and WA. | Open, Classified |  |  | 7680
Z-0349-2017 | 73286 | No | IRT | Powered heating pad. | Physical Medicine | Physical Medicine | 510(K) Exempt | Model 055: Thermophore Classic, Professional Quality, Deep-Heat Therapy, Model 055 Large 14' x 27", |  |  | September 15, 2016 | 2016 | Battle Creek Equipment Co. | Products manufactured with the SPT cord have a known failure mode of sparking,causing smoke, fire and burns. | Device Design | All consumers for which Battle Creek has contact information and that purchased a recalled pad will be contacted via first class mail. Consumers will be asked to contact us by phone after verifying that the model number and lot code is part of the recall. Battle Creek plans to add information about the recall on its web site including a pdf of the consumer letter. All customer consignees (non-consumer) will be informed of the recall. Email will be the most effective method , and mail and telephone calls will be used as needed. |  | 99060 | 99,060 | Distributed within the United States, Canada, Mexico, Puerto Rico and Columbia | Open, Classified |  |  | 99060
Z-0360-2017 | 74242 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SureT G29 6mm, 18", Paradigm; SureT G29 6mm 23"; SureT G29 6mm, 23", Paradigm; SureT G29 6mm 32"' SureT G29 6mm, 32", Paradigm; SureT G29 8mm, 23"; SureT G29 8mm, 23", Paradigm; SureT G29 8mm, 32"; SureT G29 8mm, 32", Paradigm; SureT G29 10mm, 32", Paradigm |  |  | November 07, 2014 | 2014 | Unomedical As | Tubing detachment | Device Design | ConvaTec sent an Urgent Medical Device Correction letter to all affected consignees on November 7, 2014. The letter notified the consignees of what action they need to take if they experience a detached tubing. You can continue to use your infusion sets by following the advice below.1. When changing your infusion set, closely follow the instructions for use included in the product box. Check the tubing at the connect/disconnect location identified in the drawings above to make sure it is not loose.2. As always, it is essential to monitor your blood sugar levels frequently using your blood glucose meter. Proactively check your tubing connections occasionally throughout the day to ensure tubing is secure. It is especially important to check your blood sugar and tubing connections at bedtime to confirm insulin delivery is occurring.3. If you experience a high blood sugar, check your tubing connections and infusion site closely to ensure your tubing is secure.If you discover the tubing is detached:1. Do not attempt to reattach the tubing. Replace the infusion set immediately.2. Treat any high blood sugar based on guidelines provided by your healthcare professional.3. Call your distributor's Helpline to report the issue. They will give you instructions on how to return the affected infusion set.The letter included a reply card to be returned. |  | 4748000 | 4,748,000 units | Worldwide Distribution - US including CA, IN, PA, and Internationally to Canada, Korea, and Japan. | Completed |  |  | 17165789
Z-0348-2017 | 75271 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Heater Cooler Unit |  |  | September 27, 2016 | 2016 | Maquet Cardiovascular Us Sales, Llc | There is the potential for contaminated water to enter parts of the Heater-Cooler Unit circuits. There is a potential risk of bacterial contamination including Mycobacteria species. | Device Design | Maquet issued an Urgent Medical Device Recall letter/Return Response form dated September 27, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customer were instructed to complete the return response form via fax to 1-973-629-1518. For questions contact Maquet Customer Service at 1-888-627-8383 option 2 followed by option 2 Monday through Friday between 8 amd and 6 pm. EST. |  | 305 | 305 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 305
Z-0350-2017 | 75283 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TANDEM (TM) COCR SHELL UHMWPE LINER BIPOLAR, |  |  | August 29, 2016 | 2016 | Smith &amp; Nephew, Inc. | Some Bipolar shells were manufactured with an out of specification retainer groove. | Process control | The firm initiated their recall by FedEx and Email on 08/29/2016 to hospitals and distributors and requested the return of any in stock product. International consignees were notified by Email on 08/29/2016. Additionally, the firm notified the US physician of the issue by overnight letter on 09/08/2016. |  | 2352 | 2352 units | Nationwide and Australia, Belgium, Canada, China, Cyprus, France, Germany, India, Italy, Japan, Korea, Mexico, Netherlands, Poland, Portugal, Puerto Rico, Singapore, South Africa, Spain, Switzerland, Thailand, Turkey, United Arab Emirates | Open, Classified |  |  | 8158
Z-0361-2017 | 64005 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq One Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids. List number 16026. Note that Hospira retired this pump June 30, 2015. |  |  | March 16, 2011 | 2011 | Hospira Inc. | May go to a white screen due to a problem with the SOM2 module. A white screen is a diagnostic screen displayed to the user as part of the safety shutdown procedure of the device. The infusion and the processor stop, sound a 3 minute audible alarm from the backup buzzer and deactivate the couch screen and the automated loading/unloading of cassettes. Delay or interruption of therapy. | Device Design | Hospira sent an Urgent Device Recall Letters, dated February 5, 2013, to all affected customers. Customers were directed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form. Healthcare professionals were to contact Stericycle. Hospira discontinued support of these devices as of June 30, 2015. Customers with questions were instructed to call 1-800-201-9064.For questions regarding this recall call 224-212-4892. |  | 10337 | 10,337 | Worldwide Distribution - US (nationwide) and Canada | Open, Classified |  |  | 21234
Z-0363-2017 | 75452 | No | LJE |  |  | Gastroenterology/Urology | 510(k) | X-Force Nephrostomy Balloon Dilation Catheter Kit with Inflation Device, PTFE Sheath, 8 mm x 15 cm, Product Code 996081 |  |  | October 12, 2016 | 2016 | C.R. Bard, Inc. | Contract manufacturer packaged an incorrect sheath with the product. Packaged 30Fr instead of 24Fr. | Unknown/Undetermined by firm | The firm, Bard Medical, sent an "URGENT: MEDICAL DEVICE PRODUCT RECALL" letter dated 10/13/2016 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine product subject to recall; do not use or further distribute any affected product; complete and return the attached Recall &amp; Effectiveness Check Form regardless of whether or not you have any of the affect product by fax to 1-770-784-6469 or email a scanned copy to BMD.FieldAction@crbard.com by October 24, 2016, and if you have further distributed any units of the affected lot, identify your customers and notify them at once of the product recall.If you have any questions, contact BMD Recall Coordinator at 1-800-793-8110. |  | 192 | 192 units | Worldwide distribution: US distribution to: CA, KS, MO, NY, VA, WA, and to countries of: China, Italy, Sweden and United Kingdom. | Open, Classified |  |  | 192
Z-0368-2017 | 75483 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ZimmerÂż Air Dermatome II Handpiece Loaner, item number 01-8851-001-00. Device is packaged inside of a padded-foam corrugate box. A product label is applied to the lid of the box. The width plates, when sold individually, are placed in a labeled poly zipper bag, which is then placed into a corrugated shipper carton or padded envelope. |  |  | July 04, 2016 | 2016 | Zimmer Surgical Inc | Zimmer Biomet Surgical has received product complaints reporting that the surface coating applied to the device is blistering, peeling and discoloring after usage over time (see Image 1). There have been no injuries associated with the reported complaints. In addition, an independent assessment determined that the blistering, peeling and potentially flaking of the surface coating is not expected to result in any long range health consequences. Given that the frequency of this surface coating condition can worsen with age, Zimmer Biomet Surgical is currently making preparations for replacement activities to follow in order to prevent future complaints. This document is provided to alert all users of the potential for the surface coating to exhibit blistering, peeling, and discoloring after usage over time. A separate recall removal notice will be issued to facilities, in phases, with detailed instructions beginning in August 2016. Your sales representatives will be notified when a replacement is ready. Do not return product at this time as part of this action.Image 1. | Process design | The firm, Zimmer Biomet Surgical, sent an "URGENT MEDICAL DEVICE RECALL CORRECTION NOTICE" dated 7/6/2016 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to do the following:1. Carefully review this letter and ensure all users of the Zimmer Air Dermatome II at your facility have been informed of this notice.2. Inspect the Zimmer Air Dermatome II devices within your facility and indicate on the attached Certificate of Acknowledgment the condition of the device. Complete and sign the attached Certificate of Acknowledgment and return it to corporatequality.postmarket@zimmerbiomet.com.3. Before each use, continue to examine the surface coating for blistering/peeling. If the blistering/peeling characteristics are notpresent, you may continue to use. Immediately after each use, re-examine the surface coating for blistering/peeling. If the blistering/peeling characteristics are present, discontinue use. In the event that alternate devices or therapies are not available, surgeons may, at their discretion, determine the use of the device as a medical necessity and continue to use. After use, follow your facilitys wound irrigation protocol or Zimmer Biomet recommends copious irrigation with a normal saline solution or immersion of the graft in a bowl of normal saline. Zimmer Biomet recommends irrigating the donor site with saline as well. Carefully examine the skin graft and donor site for gray coating flakes and continue to irrigate as necessary. Report this as an adverse event to Zimmer Biomet Surgical at SurgicalRegulatoryReporting@zimmerbiomet.com.If after reviewing this notification you have further questions or concerns please call 1-330-364-0989 between 8:00 am and 5:00 pm EST, Monday through Friday. |  | 1030 | 1030 units | US Nationwide Distribution including states of: AL, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, and WI. | Terminated | June 21, 2017 | 352 | 1030
Z-0369-2017 | 75510 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | TBS iNsight Version v.3.0.1 |  |  | October 03, 2016 | 2016 | Medimaps Group | The FRAX adjusted for TBS values are not correct when: The FRAX feature is activated in TBS iNsight; and TBS has been computed from a spine scan where some vertebrae were excluded. | Under Investigation by firm | Medimaps sent an Urgent Field Safety Notice letter dated October 3, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. |  | 15 | 15 units in US | Worldwide Distribution - US Nationwide - in the states of: IL, NC, CA, DE, NY, MD, WI, MN. And the following countries: France,Thailand, Croatia, Canada, Italy, Spain, Russian Federation, Ukraine, Lithuania, South Africa, Brazil, Lebanon, Saudi Arabia, Korea, China Portugalm Iraq, Syrian Arab Republic, Czech Republic, Australia, Belgium, Austria, Romania, New Zealand, Switzerland, Hong Kong, Sweden, Mexico, Islamic Republic of Iran, Serbia, Netherlands, Denmark, and UK. | Terminated | June 07, 2017 | 247 | 15
Z-0365-2017 | 75302 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Covidien Devon Light Glove sub-assembly non-sterile |  |  | September 29, 2016 | 2016 | Covidien LLC | Added Instructions for Use: After application, inspect the Light Glove for barrier integrity | Process control | Medtronic issued a Correction Notice on September 29, 2016. Customers were notified via Federal Express and the letter informed customers of the addition of the following statement to the Instructions for Use (IFU): After application, inspect the Light Glove for barrier integrity. Questions or concerns, contact your Medtronic representative or Customer Service at (800) 962-9888 |  | 118800 | 118,800 | Worldwide Distribution - US NationwideForeign:Canada:AndorraAustriaBelgiumCzech RepublicDenmarkFinlandFranceGermanyGreeceIrelandIsraelItalyLatviaLuxembourgMonacoNetherlandsNorwayPolandPortugalRussian Federation Saudi ArabiaSloveniaSouth AfricaSpainSwedenSwitzerlandUnited Arab EmiratesUnited Kingdom | Open, Classified |  |  | 123750
Z-0374-2017 | 75509 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition Flash |  |  | October 19, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Hardware issue that may expose the user to a potential risk for serious injury due to exposure to rotating or electrical parts. | Nonconforming Material/Component | Siemens mailed a Customer Safety Advisory Notice to affected customers on October 19, 2016, to inform them about the hardware issue. The notice provided guidance on risk prevention and actions they are taking to develop a solution to correct the issue. For further questions, please call (610) 448-6471. |  | 57 | 57 systems distributed in USA | Nationwide Distribution | Terminated | March 13, 2017 | 145 | 57
Z-0370-2017 | 75414 | No | KDL | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | GAMBRO Capillary Dialyzer, Polyflux(R) Revaclear MAX, STERILE, Product code 110634 |  |  | October 12, 2016 | 2016 | Baxter Healthcare Corp. | Potential presence of particulate matter on the blood side of the dialyzer | Process control | Baxter sent an Urgent Product Recall letter dated October 12, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to:I. Locate and remove all affected product lots from their facility.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.3. Complete the enclosed Baxter Customer Reply Form and return it to Baxter by either fax or scanned email.For questions contact Baxter Product Surveillance at 800-437-5176. |  | 22656 | 22,656 units | Worldwide Distribution - US Nationwide and the countries of Canada and Mexico. | Terminated | August 19, 2017 | 311 | 22656
Z-0372-2017 | 75288 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Replacement floor locking cylinders for the Schaerer model 7300 surgical tables. Replacement cylinders are part number 85118. |  |  | September 27, 2016 | 2016 | Schaerer Medical USA Inc | The floor locking cylinders are contained in the base of the Schaerer 7300 model surgical tables. The cylinders were manufactured with the incorrect "O" ring which may cause the cylinder to leak. | Nonconforming Material/Component | Schaerer Medical sent a Field action letters dated September 27, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The customer is required to complete and return a response form with the number of affected tables and cylinders. All affected cylinders will be replaced in the field and returned to the manufacturer. For questions Quality Assurance at 513-272 -9116. |  | 10 | 10 replacement cylinders. | Worldwide Distribution - US Nationwide in the states of: CA, CT, FL, IA, IN, MI, MN, MO, OH, PA, and VAReplcement cylinders wre distributed | Open, Classified |  |  | 10
Z-0309-2017 | 75358 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Marathon Flow Directed Micro Catheter, UltraFlow HPC Flow Directed Micro Catheter |  |  | October 05, 2016 | 2016 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | Delamination and detachment of the PTFE (polytetrafluoroethylene) coating material may lead to PTFE coating in the blood stream. | Device Design | Medtronic sent an Urgent Medical Device Recall letter dated October 5, 2016, to all affected customers to inform them that Medtronic has identified the potential for an issue with a specific set of lot numbers of Medtronic Neurovascular products as noted where the PTFE (polytetrafluoroethylene) coating could delaminate and separate from the delivery wire. Customers are instructed to complete the attached customer confirmation certificate and fax it to Medtronic at (949) 434-5020 to the attention of Neurovascular Quality.On October 14, 2016, Medtronic announced that it has notified customers of a voluntary recall of certain lots of its Pipeline embolization device, Alligator retrieval device and X-Celerator hydrophilic guidewire. The recall also includes the stylet containing UltraFlow flow directed micro catheters and Marathon flow directed micro catheters. This voluntary recall is being conducted due to the potential separation and detachment of the polytetrafluoroethylene (PTFE) coating on parts of these devices. Should the PTFE separate from the delivery wire or stylets, PTFE particulate could enter the blood stream of the patient. PTFE in the blood stream, based on the size and quantity, could lead to a thromboembolic event.Customers were asked to complete the attached Customer Confirmation Certificate and fax it to Medtronic at 1-949-434-5020 to the attention of Neurovascular Quality. Customers with questions were instructed to contact their Medtronic representative. |  | 84278 | 84,278 units total (17,811 in US) | Worldwide Distribution - US (nationwide) and Internationally to Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Iran, Ireland, Israel, Italy, Japan, Korea, Latvia, Lithuania, Malaysia, Mexico, Nepal, Netherlands, Nicaragua, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Russian Federation, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Open, Classified |  |  | 84278
Z-0378-2017 | 75246 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LFIT Anatomic V40 Femoral Head, Low Friction Ion Treatment, Sterile, 36 mm, REF 6260-9-236; Modular components designed to be locked onto a femoral hip stem trunnion during surgery for total hip replacement. |  |  | August 29, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Stryker received several complaints describing incidence of harm secondary to taper lock failure for specific lots of numerous catalog numbers of LFIT Anatomic CoCr V40 Femoral Heads. | Under Investigation by firm | Stryker notified their Branches/Agencies of this recall by e-mail on August 29, 2016 and they were asked to quarantine the affected devices. A Recall Notification Letter and Product Accountability Form was also sent on August 29, 2016 via UPS (with return receipt) to their Branches/Agencies/Hospital Risk Management and Surgeons. On October 11, 2016, Stryker sent an updated recall notification via UPS with return receipt to their affected customers because additional customers and lot numbers were identified. |  | 42519 | 42,519 units (total Catalog numbers) | US Nationwide and Internationally | Open, Classified |  |  | 42519
Z-0328-2017 | 75395 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | FindrWIRZ Guidewire System Catalog# 40-05 |  |  | September 29, 2016 | 2016 | SentreHeart Inc | A complaint was received on September 27th, 2016 at a medical center in Germany for delamination and flaking of the PTFE lubricious coating on threeFindrWIRZ Guidewire System .025 guidewires removed from their packaging hoops. | Under Investigation by firm | SentreHeart sent an Urgent Medical Device Voluntary Recall letter dated September 29, 2016, to all affected customers by traceable delivery method. Letters explained the issue and requested that any devices from the two affected lots be returned to SentreHEART immediately. Customers were asked to complete and return the enclosed response form as soon as possible via fax at 650-354-1204 or email to customerservice@sentreheart.com. Customers with questions were instructed to call 1-855-256-7330.For questions regarding this recall call 650-241-6008. |  | 132 | 132 worldwide distribution (98) of those units in US | Worldwide Distribution - US including MO, CA, OR, GA. NH, NY, MI, TX, MD, PA, IL, KS, NE, CT, AZ and Internationally to United Kingdom and Germany | Terminated | March 01, 2017 | 153 | 132
Z-0379-2017 | 75448 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Lactic Acid Part # 09D891T21 |  |  | October 11, 2016 | 2016 | Abbott Laboratories | Abbott has identified negative interference from the drug N-Acetyl Cysteine (NAC) with the ARCHITECT Lactic Acid reagent (LN 9D89-21). | Under Investigation by firm | Abbott Laboratories sent an Product Correction Immediate Action Required letter dated October 10, 2016, to all affected customers. Please review this letter with your Medical DirectorTaking the above information into consideration, you may continue to use ARCHITECT Lactic Acid reagent (LN 9D89-21)Please complete and return the Customer Reply FormIf you have forwarded the product listed above to other laboratories, please inform them of this Product Correction and provide to them a copy of this letterPlease retain this letter for your laboratory records.If you or any of the health care providers you serve have any questions regarding this information, U.S. Customers please contact Customer Service at 1-877-4ABBOTT (available 24 hours a day, 7 days a week). Customers outside the U.S., please contact your local area Customer Service.For further questions please call (847) 937-2550. |  | 26837 | 26, 837 kits | Nationwide and Worldwide Distribution. | Terminated | May 17, 2017 | 218 | 26837
Z-0380-2017 | 75409 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Reprocessed Daig Steerable and Daig Supreme Fixed Curve Diagnostic Electrophysiology Catheters |  |  | October 06, 2016 | 2016 | Stryker Sustainability Solutions | The EP Catheters may be mislabeled for French size during reprocessing. | Under Investigation by firm | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 10/6/2016 to its customers. The letter described the product, problem and actions to be taken.Customers were instructed to discontinue use of the recalled product. Other EP Catheters reprocessed by Stryker Sustainability Solutions should be considered as alternatives. To complete and Return the Recall Effectiveness Check Form, even if no affected product is found, to your local Stryker Sustainability Sales Representative or email to SSSPFA@stryker.com or mail to:Stryker Sustainability Solutions1810 West Drake DriveTempe, AZ, 85283Attn: Jodie RueckertIf you have any questions, contact Senior Director RA at 1(888)888-3433. |  | 4 | 4 devices | U.S. distribution to the following; AZ, MN and MO. | Open, Classified |  |  | 4
Z-0386-2017 | 75554 | No | CCQ | Medicinal nonventilatory nebulizer (atomizer). | Anesthesiology | Anesthesiology | 510(K) Exempt | LMA MADdy, Rx only, |  |  | October 27, 2016 | 2016 | Teleflex Medical | Product Defect; These products may produce a straight stream instead of a fully atomized plume of medication. | Process control | Teleflex sent an Urgent Medical Device Recall Notification letter on/about October 27, 2016. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and quarantine affected products. Customers were asked to complete the enclosed Recall Acknowledgement Form and fax it to Teleflex Customer Service or email to recalls@teleflex.com. For questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. Teleflex sent an Amended Urgent Medical Device Recall Notification letter dated November 22, 2016, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to continue to return affected product for a full refund per the procedure set forth in the letter of October 27, 2016. Some customers have indicated that, due to medical necessity, they plan to continue using the affected lots rather than return them. Teleflex is advising all such customers to follow the supplemental instructions for Use included as Attachment B. These supplemental instructions allow non-destructive testing of each unit prior to the procedure to determine if it is defective. Institutions that intend to continue using affected lots with this supplemental testing, please send the acknowledgment form to recall@teleflex.com with their name and contact information and the lots that they will continue to use, or call 866-246-6990. Customers with questions should contact their local sales representative or Customer Service at 1-866-246-6990. |  | 2674650 | 2,674,650 units in total | Worldwide distribution. U.S. (Nationwide) and Internationally to Canada, Belgium, Singapore and Thailand. | Open, Classified |  |  | 2674650
Z-0377-2017 | 75178 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AtriClip LAA Exclusion System with preloaded Gillinov-Cosgrove Clip |  |  | September 22, 2016 | 2016 | AtriCure, Inc. | There is a potential for a component failure of the device which could necessitate medical intervention. | Device Design | AtriCure sent an Urgent Advisory Notice dated September 21, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and determine if they have any affected product. If so customers should ensure that all users of the affected devices are aware of the Advisory Notice. If customers further distributed the product, they should identify their customers and notify them at once of the Advisory Notice. Any units from the affected lots should be returned to AtriCure, Inc. Customers should contact AtriCure customer service at 1-866-349-2342, and select option 6 to return any affected product(s). Customers were asked to complete and return the enclosed Device Notification Acknowledgement Form as soon as possible. Customers with questions should call 513-755-4563. |  | 298 | 298 | Nationwide Distribution to the follwoing states: AL, AR, AZ, CA, CT, FL, GA, IL, KS, MD, MI, MN, MO, MS, NC, NM, NY, OH, OK, PA, SC, TN, TX, VA, VT &amp; WI.Product was also shipped to the following countries: Poland and Switzerland | Open, Classified |  |  | 298
Z-0394-2017 | 75150 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | catalog # 405612 |  |  | August 02, 2016 | 2016 | Becton Dickinson &amp; Company | A recall was initiated by Hospira, Inc. for several lots of 5% Lidocaine HCI and 7.5% Dextrose Injection, USP, 2mL Ampul, NDC # 0409-4712-01. Some BD" Spinal Trays contain the 5% Lidocaine Hydrochloride and 7.5% Dextrose Injection, USP, 2mL Ampul (NDC Number 0409-4712-01) being recalled by Hospira. As a result, BD is initiating a recall for the affected BD Spinal Trays. | Nonconforming Material/Component | BD sent recall letters/business reply card was sent via UPS/email on 8/2/2016.Stericycle is handling returns. |  | 39130 | 39,130 trays | Nationwide. | Open, Classified |  |  | 39130
Z-0387-2017 | 75357 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS INTEGRA c111 |  |  | September 27, 2016 | 2016 | Roche Diagnostics Operations, Inc. | cobas c 111 analyzers (catalog numbers 04777433001 and04528778001) with software versions up to and including 4.20 may encounter the following alarm: 7002: 108000572, A software error occurred. This alarm is generated due to a measurement timingerror. Under very rare conditions, the instrument may process two tests in the same cuvette if the runrestarts. If a used cuvette is used again result of the test(s) will be erroneous.These erroneous results may not be flagged.Falsely low or high patient results may lead to incorrect diagnostic measures and medical therapeuticdecisions. The medical risk depends on the parameter. | Software design | Roche sent an URGENT MEDICAL DEVICE CORRECTION letter dated September 27, 2016 to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were asked to complete the attached fax form and fax it to 1-866-808-1159. For questions contact the Roche Support Network Customer Support Center at 1-800-428-2336. |  | 139 | 139 | US Nationwide Distribution | Terminated | July 21, 2017 | 297 | 139
Z-0388-2017 | 75388 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SB Knife -Electrosurgical Knife, Sterile |  |  | September 09, 2016 | 2016 | Akita Sumitomo Bakelite Co., Ltd. | Mislabeled product: Sterilization pouch and carton box labeled as Standard type instead of Jr type-Long | Error in labeling | Sumitomo Bakelite on September 9, 2016 initiated removal of all SB Knife from US accounts.The information was emailed to the accounts and Sumitomo Bakelite requested all customers not to use the mislabeled SB Knife and return product for replacement. |  | 15 | 15 units | US in the states of MA | Terminated | November 10, 2016 | 62 | 23
Z-0392-2017 | 75475 | No | LTI |  |  | Gastroenterology/Urology | PMA | ReShape Integrated Dual Balloon System, Gen 1, Model No. 01-0013-001, Catalog No. RSM110 |  |  | October 14, 2016 | 2016 | Reshape Medical Inc | ReShape Medical is recalling the Integrated Dual Balloon System Gen 1 due to the potential leak of distillation fluid during balloon inflation. | Device Design | ReShape sent a customer notification letter dated October 14, 2016 was sent to customers to inform them that ReShape Medical is recalling the Integrated Dual Balloon System due to customer reports of instillation fluid leaking during the device inflation procedure. The letter informs the customers that any instillation fluid leakage during the inflation process will cause an incomplete balloon fill and require immediate removal of the device. Customers are instructed to contact ReShape Medical Customer Service at (844) 9377374 for any questions related to the product returns. |  | 114 | 114 units | US Nationwide Distribution in the states of: MO, TX, AZ, CA, NY, TN, GA, LA, FL, MA, OH, IN, IL | Open, Classified |  |  | 114
Z-0390-2017 | 75386 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Accessory Adapter Part number 600525A0 |  |  | October 05, 2016 | 2016 | Maquet Medical Systems USA | Maquet Inc. is initiating a voluntary field action on the Accessory Adapter part number 600525A0 due to a complaint in which a head rest mounted to the accessory adapter slipped out of the mountain rail. | Nonconforming Material/Component | Maquet Medical Systems USA sent an Urgent Product Recall Medical Device Removal letter dated October 5, 2016 to the affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to:1. Read this Medical Device Removal Notice and disseminate the informationprovided to all the users of the device. 2. Examine your inventory immediately to determine if you have the Accessory Adapter(s). If so, remove the affected products, quarantine them and keep them in a secure location for shipment back to Maquet. 3. Please fill out and sign the enclosed Medical Device Removal Response Form on page 4 and return it either via email to FieldActions@maquet.com or fax to 1-973-909-9927. For question contact your local MAQUET Representative or call MAQUET Customer Service at 1-888-627-8383 (Press option 2, then followed by option 3) Monday through Friday, between the hours of 8:00 a.m. and 5:00 p.m. EDT. |  | 9 | 9 Units (US) 141 Units (OUS) | Worldwide Distribution - US Nationwide in the states of AZ, CA, DC, MI, MO, and PA | Open, Classified |  |  | 9
Z-0391-2017 | 75412 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OR Table, Top U26 H V, Model # 1909793; OR Table Top U26 H V U, Model # 1909794; OR Table Top U24 H V, Model # 1909796; OR Table Top U24 H V U, Model # 1909797; OR Table Top U24 H V W, Model # 1909798; OR Table Top U14 H V, Model # 1909799 |  |  | October 25, 2016 | 2016 | Trumpf Medical Systems, Inc. | Trumpf has identified six incidents of pinched cables. | Process control | Trumpf Medical sent an Urgent Medical Device Correction letter dated October 7, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken The letter indicated that a authorized service provider will inspect the table tops and make necessary corrections. A confirmation notice was included with the letter which is to be returned. For contact Trupmpf Medical Systems, Inc. at 888-Âż474-Âż9359. |  | 4 | 4 (US) &amp; 229 (Foreign) | Worldwide Distribution - US Nationwide in the states of AZ and SC and the countries of Australia, Austria, Belgium, China, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Iceland, Italy, Norway, Qatar, Slovenia, Spain, Sweden, Switzerland | Open, Classified |  |  | 4
Z-0393-2017 | 75474 | No | DXF | Septostomy catheter. | Cardiovascular | Cardiovascular | 510(k) | NRG Transseptal Needle and NRG RF Transseptal Kit. |  |  | October 11, 2016 | 2016 | Baylis Medical Corp * | The firm is voluntarily recalling lots of the NRG Transseptal Needle and NRG RF Transseptal Kit due to the possibility for an integrity defect to the Tyvek pouch of the NRGTransseptal Needle. | Packaging process control | A recall notification, marked URGENT Voluntary Medical Device Recall was sent by Baylis Medical Company via overnight courier to all customers affected by the recall with an Acknowledgement Form. Customers were asked to examine their inventories, quarantine any affected lots, and complete the attached Acknowledgement Form. The form should be returned to Baylis Medical Company via the instructions provided. Upon receipt, the firm will contact customers with information on how to return affected product. Questions regarding the letter can be directed to the Customer Service Department at 1-800-276-4416. |  | 28093 | 28,093 units | Worldwide Distribution -- US, Canada, Chile, France, Germany, Hong Kong, Ireland, Italy, Mexico, Netherlands, Spain, UAE, and UK. | Terminated | July 10, 2017 | 272 | 28093
Z-0399-2017 | 75419 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Merge LIS software. |  |  | March 05, 2015 | 2015 | Merge Healthcare, Inc. | There is a potential for duplicate container numbers to be created for patients. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/30/2016. If no email address was available, the letters were issued via certified mail. Letters to nonresponders were issued 5/26/2016. |  | 413 | 413 sites potentially have the affected software | Distribution was made nationwide and to the Virgin Islands. Foreign distribution was made to the Bahamas. | Open, Classified |  |  | 413
Z-0401-2017 | 75256 | No | JIY | Iron (non-heme) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista IRON Flex reagent cartridges |  |  | August 25, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has become aware of deferoxamine interference with Dimension and Dimension Vista IRON assays at a different concentration than listed in the Instructions for Use (IFU). | Error in labeling | Siemens mailed an Urgent Medical Device Correction letter to all customers who received the Dimension and Dimension Vista IRON Flex reagent with an update to the Limitation of Procedure section for the Dimension and Dimension Visa IRON assays Instructions For Use (IFU) to indicate that: patients treated with metal-binding drugs (e.g. deferoxamine) may have depressed iron values, as chelated iron may not properly react in the iron assay. Customers were told that the information provided in the letter, related to deferoxamine, supersedes the information related to deferoxamine int he current IRON IFU's for Dimension and Dimension Vista products until the IFU's are updated. |  | 30247 | 30,247 units | Distributed Nationwide | Open, Classified |  |  | 68477
Z-0398-2017 | 75394 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA CentaurÂż Vitamin D Total Assay |  |  | October 05, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | Vitamin D Total- Change in Correlation Between Serum and Plasma Specimen Tubes- negative bias with plasma vs. serum samples | Device Design | Siemens issued an Urgent Medical Device Correction (UMDC) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on October 5, 2016 to be delivered to customers on October 6th, 201 6. An Urgent Field Safety Notice ~UFSN) was sent out for distribution to all customers outside the United States on October 5, 2016. These notices inform the customer ofthe change in correlation between serum and plasma specimen tube results generated on the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT Vitamin D Total assay. A revised letter was issued on November 7, 2016 by Fedx for delivery on November 8, 2016 and advises users to: In the interim, Siemens requires that customers discontinue use of plasma tubes for specimen collection for the ADVIA Centaur Vitamin D Total assay until further notice.Customers can continue to use the ADVIA Centaur Vitamin D Total assay with serum specimen collection tubes.Accounts to Complete and return the Field Correction Effectiveness Check Form attached within 30 days. Review the letter with your Medical Director.and retain this letter with your laboratory records, and forward this letter to those who may have received this product.Question contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 37437 | 37437 kits | NationwideForeign:Afghanistan, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil,Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt,Finland, France, French Polynesia, Georgia, Germany, Great Britain, Greece, Guadeloupe,Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Italy, Japan, Jordan, Latvia, Lebanon,Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Nepal, Netherlands,New Zealand, Norway, Oman, Pakistan, Paraguay, Philippines, Poland, Portugal, Qatar,Romania, Russian Federation, Saudi Arabia, Singapore, Serbia, Slovakia, South Africa, SouthKorea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates,Uruguay, Vietnam | Open, Classified |  |  | 107218
Z-0400-2017 | 75387 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | PET Infusion Pump |  |  | September 30, 2016 | 2016 | Bayer Healthcare | The Source Administration Sets (SAS) used with the Medrad Intego PET Infusion System do not meet pressure testing specifications in an occlusion situation when the product is more than one year past manufacturing date. | Process control | Bayer mailed an Urgent Medical Device Customer Notification letter to the affected consignees on September 30, 2016 asking that they review their current SAS inventory and determine if you have any product that is more than one year past its manufacturing date, Quarantine unused units and call Bayer Customer Care line to receive a Return Goods Authorization (RGA) number for the product's return. |  | 165127 | 165,127 (79,714 units in U.S.) | Distributed Nationwide and internationally to : Taiwan, Ukraine, Turkey, Thailand, Slovania, Singapore, Sweden, Saudi Arabia, Serbia, Romania, Qatar, Portugal, Poland, Panama, Oman, Norway, The Netherlands, Mexico, Malta, Morraco, Libyan, Lithuania, Lebanon, Kuwait, Italy, Israel, Hungray, United Kingdom, France, Finland, Spain, Denmark, Germany, Czech Republic, Chile, Switzerland, Belguim, Bosnia, Austria, Australia, | Terminated | March 24, 2017 | 175 | 165127
Z-0395-2017 | 75366 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Tina-Quant Hemoglobin A1c Gen. 2 |  |  | October 04, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Roche Diagnostics has confirmed elevated QC and patient sample recovery for individualTina-quantÂż Hemoglobin A1c Gen.2 and Tina-quant Hemoglobin A1cDx Gen.2 cassettes/bottles.This issue occurred over multiple lots. Single cassettes/bottles of the reagents are affected.The issue can lead to erroneously high HbA1c results. If a patient with known diabetes is affected,an erroneously high HbA1c result might lead to therapeutic consequences, such as therapy escalation(introduction of another oral antidiabetic medication or insulin) or increase of the dosage. This canfurther lead to an increased risk of hypoglycemia. Considering the unreliable detectability of the issue,relevant medical risk for the patient cannot entirely be excluded. | Storage | On 10/4/2016, URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected Direct Consignees and Distributors via courier. The recall notifications included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Direct ConsigneeActions Required" Follow the instructions in the Customer Actions for Prevention/Detection of the Issue section of this Urgent Medical Device Correction (UMDC)." Complete the attached fax form (7184-00-1016) and fax it to 1-844-294-7181." File this Urgent Medical Device Correction (UMDC) for future reference.QuestionsPlease contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a weekat 1-800-428-2336 if you have questions about the information contained in this UMDC.DistributorActions Required" Ensure proper storage conditions (2-8ÂżC) are maintained for the reagents listed in the table above,per their respective method sheet." If you have not directly shipped any of the products in the table above, you do not need to distributethe enclosed UMDC to your customers." If you have directly shipped any of the products in the table above, please distribute the enclosedUMDC 16-157 to those customers." Complete the attached fax form (7185-00-1016) and fax it to 1-844-294-7181.Questions Please contact the Roche Support Network Customer Support Center at 1-800-428-2336 if you havequestions about the information contained in this distributor letter or the enclosed UMDC 16-157. |  | 4976 | 4976 | Nationwide. | Terminated | June 05, 2017 | 244 | 4976
Z-0403-2017 | 75492 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Fluoro-4 Silicone Ureteral Stent, BardÂż Fluoro-4", Silicone Ureteral Coil Stent; BardÂż Silicone Uretera l Coil Stent with Figure Four, End, Manufacturer: C.R. Bard, Inc., Covington, GA 30014. |  |  | October 13, 2016 | 2016 | C.R. Bard, Inc. | Labeling: Statement was added to the Warnings/Precautions section requesting end users to consider the presence of a knot if significant resistance is encountered during attempts at removal. | Labeling Change Control | Bard Medical sent an Urgent: Medical Device Communication Letter dated October 13, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action required:Examine your inventory and identify any product subject to this notification. If you have any remaining inventory, include a copy of this communication with the unit(s) for future use.If you may have further distributed or transferred this product, please identify the respective organizations and notify them at once of this communication. Your notification to these organizations may be enhanced by including a copy of this product communication letter.This notification should be carried out to the user level.This product notification is being made with the knowledge of the Food and Drug administration.We appreciate your cooperation and assistance in dealing with this matter and sincerely apologize for any inconvenience that may result from this action.For further questions, please call (770) 784-6220. |  | 75089 | 75,089 units | U.S. Nationwide. | Open, Classified |  |  | 75089
Z-0414-2017 | 75491 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | 4 Port "ON" Manifold (600 psi) w/4 ft. TranspacÂż IV, Item No. 46300-14NS |  |  | September 27, 2016 | 2016 | ICU Medical, Inc. | ICU Medical Inc. has identified a potential for leak with certain SafeSet Reservoir and blood sampling kits containing the Transpac IV transducers. | Device Design | ICU Medical sent an Urgent Medical Device Recall Notification letter dated September 27, 2016, to all affected customers to inform them of the expanded recall of the monitoring kits with Transpac IV transducers. Customers were informed of updated communications from August 10, 2016, when ICU Medical sent a communication indicating a potential leak with certain SafeSet Reservoir and blood sampling kits containing the Transpac IV transducers. Based on continued evaluation, ICU Medical is expanding the notification to now include certain other monitoring kits that contain the Transpac IV transducers. A small number of these devices may have the potential to leak or potentially introduce small amounts of air into the line. Customers with questions were instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 2, or e-mail the following address: productreturns@icumed.com. Customers are instructed to complete the attached Medical Device Recall Response Form and fax to (801) 264-1755 or e-mail to recall@icumed.com. Customers with questions were instructed to call 866-829-9025. |  | 20 | 20 units | Nationwide Distribution | Terminated | July 11, 2017 | 287 | 210737
Z-0411-2017 | 75494 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 |  |  | October 20, 2016 | 2016 | Mevion Medical Systems, Inc. | Software defect:2D projection of contours in Verity. The defect causes structures to appear off from their true position in the DRR and the radiograph with an error that increases with the distance of the structure from isocenter | Software design | Mevion issued an Important Safety Notice letter on October 14, 2016 to their customers. The letter identified the affected product, problem, containment, mitigation and actions taken. Customers were informed that Mevion will be updating software when it is released. For further Information, contact: srosenthal@mevion.com. |  | 6 | 6 units | US Nationwide Distribution in the states of: FL, MO, NJ, OH, OK | Terminated | May 15, 2017 | 207 | 6
Z-0409-2017 | 75496 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 |  |  | October 31, 2016 | 2016 | Mevion Medical Systems, Inc. | Completing Setup Without Program Move Complete Can Cause An Isocenter Error | Software design | Mevion issued an Important Safety Notice letter on October 31, 2016 to customers. The identified the affected product, problem, mitigation and advisory information. The letter informed customers that Mevion will release and install a software patch that removes the reliance on tracking isocenter changes for the correct reporting of the Couch position to the user displays and to the OIS. For further information mailto:srosenthal@mevion.com |  | 6 | 6 units | US Nationwide in the states of FL, MO, NJ and OK | Terminated | June 29, 2017 | 241 | 6
Z-0410-2017 | 75498 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 |  |  | October 21, 2016 | 2016 | Mevion Medical Systems, Inc. | Recovery of Incorrect Isocenter -operational change | Software Design Change | Mevion issued an User Notice letter on October 24, 2016 to their customers. The letter identified the affected product, problem and actions taken. Customers were informed that Mevion will be updating software when it is released. For further clarification or assistance, please contact Rob Cessac by email at rcessac@mevion.com or by phone at (978) 540-1501. |  | 6 | 6 units | US Nationwide Distribution in the states of FL, MO, NJ, OH and OK, | Terminated | May 10, 2017 | 201 | 6
Z-0407-2017 | 75526 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Ostase Calibrators, Catalog No. 37305, B83876 |  |  | October 24, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter has determined that the Access Ostase Calibrator and Access Ostase QC lots do not meet their expiration date claims within the 10% criteria. | Under Investigation by firm | Beckman Coulter sent an Urgent Medical Device Recall letter dated October 24, 2016 was sent to all their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions regarding the notice are instructed to contact Customer Technical Support Center at http://www.beckmancoulter.com, (800) 854-3633 in US and Canada, and to contact local Beckman Coulter representative if outside of US and Canada. |  | 357 | 357 kits total (94 kits in US) | US Nationwide Distribution in the states of: UT, MI, MA, LA, IN, NC, CA, TX, NJ, IL, WI, MN, FL, KA, NY, CO, WA, KY, VA, and AL.Affected geographies: Belgium, Canada, China, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Italy, Japan, Mexico, Philippines, Portugal, Romania, Russian Federation, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Taiwan. | Open, Classified |  |  | 357
Z-0413-2017 | 75456 | No | KRG | Pacemaker programmers. | Cardiovascular | Cardiovascular | PMA | The ORCHESTRA PLUS Programmer |  |  | September 29, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field action for the ORCHESTRA Programmer and ORCHESTRA PLUSProgrammer because the residual longevity estimate may be less accurate. | Software Design Change | The firm, LivaNova, sent an "Urgent Field Safety Notice" dated 9/29/2016 to its customers. The notice described the product, problem, actions to be taken. The notice states that a Sorin representative will contact customers to upgrade the software in the facility's programmer.If customers have any questions, they are urged to call their LivaNova representative or contact LivaNova Technical Services at 877.663.7674. |  | 586 | 586 programmers | Programmers were distributed nationwide, to VA/govt/military, Canadian, and other foreign consignees. | Open, Classified |  |  | 1433
Z-0589-2017 | 75487 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Advia Chemistry Calibrator |  |  | October 07, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the DBIL_2 Assigned Calibrator Values (ACV) currently assigned to the affected lots may cause, on average, a positive proportional bias of approximately 41% on patient samples compared to the reference method for Direct Bilirubin measurement. | Under Investigation by firm | Siemens issued an Urgent Medical Device Correction in the US and an Urgent Field Safety Notice Outside the US beginning on October 6, 2016. Customers were notified of the issue and instructed to use the reassigned DBIL_2 calibrator value for Chemistry Calibrator Lot Numbers 680726, 680726A, 680726B, 680726C and 680726D. Customers were also requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within 30 days by faxing to the Customer Care Center at 312-275-7795.For questions regarding this recall call 914-631-8000. |  | 10614 | 10,614 units | Worldwide Distribution - US (Nationwide) Internationally to Canada, Mexico, Algeria, Australia, Austria, Bangladesh, Belgium, Brazil, Chile, Cyprus, Denmark, Egypt, Ecuador, Finland, France, Germany, Great Britain, Hong Kong, Hungary, Italy, Japan, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Paraguay, Peru, Poland, Portugal, Russian Fed. Slovakia, Slovenia, Spain, Sweden, Switzerland, Thailand, Turkey, Vatican City and Vietnam | Open, Classified |  |  | 10614
Z-0584-2017 | 75485 | No | GXD | Radiofrequency lesion generator. | Neurology | Neurology | 510(k) | RF MultiGenÂż, 0406-900-000; Refurb RF MultiGenÂż, 0406-900-000U |  |  | October 11, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the Care Instructions/Instructions for Use packaged with the Radiofrequency (RF) MultiGenÂż Generator, because the instructions for use incorrectly contained instructions supporting the sterilization of the MultiGenÂż cables. | Labeling Change Control | Stryker Instruments initiated a voluntary recall of the RF MultiGen, and Refurb RF MultiGen because the instructions for use contained incorrect instructions on sterilizing the device. Stryker distributed customer notices via certified mail on 10/11/2016. Customer's were instructed to do the following: Immediately review this recall notification.2. Enclosed was a revised version of the IVS Care Instructions. Please distribute the revised version to the appropriate person or department.3. Remove and discard any previously distributed IVS Care Instructions or Instructions for Use.4. Please complete, sign, and return the enclosed Business Reply Form (BRF) acknowledging your receipt and understanding of this notification.5. Fax the completed BRF to Stryker Instruments Regulatory Department, 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com. Customers with questions can contact: 269 323 7700 . |  | 2675 | 2,675 | Domestic:AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, WYForeign:France, Germany, Canada, UK, Lebanon, South Africa, Chile, Hong Kong, Poland, Argentina, Spain, Greece, Italy, India, Brazil, Switzerland, Netherlands, SwedenVA/DOD: | Terminated | July 25, 2017 | 287 | 2675
Z-0585-2017 | 75428 | No | FAS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | C-Max" Cutting Loop Electrodes, Storz One-Stem, 24 Fr, .014in., .35 mm, Sterile, Rx only, Manufactured by: C.R. Bard, Inc., Covington, GA 30014. |  |  | October 12, 2016 | 2016 | C.R. Bard, Inc. | Defect; The contact part could possibly separate from the electrode affecting to functionality of the device. | Other | Consignees were notified via letter on 10/12/2016. |  | 70 | 70 units | FL, GA, LA, OH, VA and WA. | Open, Classified |  |  | 70
Z-0592-2017 | 75337 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Arsenal Spinal Fixation System, Set Screw, Part No. 47127 |  |  | October 05, 2016 | 2016 | Alphatec Spine, Inc. | Alphatec Spine is recalling the Arsenal Spinal Fixation System Set Screw due to a trend in set screw postoperative disengagement from the screw body. | Under Investigation by firm | Alphatae Spine sent an Urgent Medical Device Recall notification letter dated September 29, 2016 to customers. The letter identified the affected product, problem and the actions to be taken. Customers are instructed to fill out and return the last page of recall letter using one of the methods identified in the letter. |  | 16972 | 16,972 units | Worldwide Distribution - US Nationwide in the countries of Japan, UK, NSW, Spain. | Open, Classified |  |  | 16972
Z-0591-2017 | 75417 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | FlexLab Automation Modules |  |  | September 29, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. |  | Under Investigation by firm | Siemens mailed to customers an Urgent Field Safety Notification letter warning them of the problem and to inform them that Siemens Customer Service Engineers shall be visiting customer sites to replace the power cords. |  | 448 | 448 units | Distributed Nationwide and to the following countries:Australia, Azerbaijan, Bahrain, Bangladesh, Belgium, Brazil, Canada, Czech Republic, Denmark, Egypt, Germany, Greece, Indonesia, Italy, Kingdom Saudi Arabia, Lebanon, Malaysia, P.R. China, Philippines, Portugal, Qatar, Republic Korea, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., United Kingdom, Vietnam | Open, Classified |  |  | 448
Z-0582-2017 | 75347 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Atrium Medical 28 Fr Trocar catheter, Sterile |  |  | October 26, 2016 | 2016 | Atrium Medical Corporation | Trocar Catheters are packaged with a label containing an icon that shows more holes (eyelets) than the product actually contains. | Labeling design | Atrium Medical issued an "Urgent Medical Device Recall (Removal)" communication to customers on 10/26/16 via FedEx 2 day delivery. The recall identifies the problem, product, and risk factors. In addition, it provides instructions on how to handle product that may be affected. Accounts are to read the letter completely and share with all users and staff. They have been requested to complete the enclosed Recall Reply Form, acknowledging receipt of the notice; and, remove and quarantine any affected product for return. Additionally, obtain product from an alternate supplier because replacement product is not available as all Trocar catheters have been discontinued. Contact Atrium Medical Customer Service at 1-800-370-7899, Monday through Friday between 9:00 am to 5:00 pm, for a Return Goods Authorization to return the product and receive credit. |  | 7043 | 7043 cases (10/cs) | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Costa Rica, Dominican Republic, Ecuador, Finland, France, Germany, Honduras, Hong Kong, Ireland, Jamaica, Kuwait, Mexico, Netherlands, New Zealand, Panama, Peru, Portugal, Puerto Rico, Qatar, Reunion, Romania, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, and Venezuela. | Open, Classified |  |  | 31247
Z-0587-2017 | 75513 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FlexLab, Accelerator a3600 and Aptio Automation are modular systems designed to automate Pre-Analytical and Post-Analytical processing, sample handling in order to automate sample processing in the Laboratory. |  |  | October 04, 2016 | 2016 | Inpeco S.A. | Inpeco received several notifications from the field about burnt schuko sockets. The customers reported that some smoke was seen coming from the damaged sockets. | Employee error | Inpeco sent an Urgent Field Safety Notice dated October 2016 for the Aptio Automation, the Accelerator a3600 and the FlexLab Shucko power cord cables to all affected US and foreign customers. The Urgent Field Safety Notice described the reason for the recalls and actions that should be taken.The firm is requesting all customers to complete and return the Field Correction Effectiveness Check form attached to the Urgent Field Safety Notice letter within 7 days. The Urgent Field Safety Notice letter should be forwarded to affected departments or consignees.To mitigate any possible risks to the laboratory users' health conditions due to the defective power cords Inpeco released a Technical Service Bulletin (TSB) addressed only to Field Service Engineer with the procedure which has to be performed in order to prevent other cases of damaged sockets. The FSE has to check visually the appearance of the Schuko female socket. If it is intact, without any detectable deformation nor melting the FSE has to verify that the terminals are correctly crimped into each single wire. If the female socket presents alterations, the socket has to be replaced with a new one. The affected sockets will be destroyed.Customers with questions were instructed to call (+41) 91 9118 258. |  | 188 | 188 units | Worldwide Distribution - US including Texas and New York. Foreign Consignees: Germany, Italy, France, Austria and Spain. | Open, Classified |  |  | 188
Z-0586-2017 | 75241 | No | NDC | Predictive pulmonary-function value calculator. | Anesthesiology | General &amp; Plastic Surgery | 510(k) | Version 1.2.0 of Accu-Chek Connect Diabetes Management App (iOS) released on July 11, 2016 |  |  | September 14, 2016 | 2016 | Roche Diabetes Care, Inc. | Roche Diabetes Care, Inc. initiated a voluntary recall for the version 1.2.0 of Accu-Chek Connect Diabetes Management App (iOS), released on July 11, 2016, due to the software containing a programming error (bug) in the Bolus Advisor feature, which could lead to an incorrect insulin bolus recommendation. | Software design (manufacturing process) | Roche Diabetes Care, Inc., initiated a voluntary recall on September 14, 2016, of the version 1.2.0 of the Accu-Chek Connect app on iOS products (iPhone, iPad). The Prescription Activation Service (PAS) which allows the users to activate the bolus advice feature was placed on product hold immediately. This measure prevented any new bolus advice activations. An updated version (V 1.2.2) of the Accu-Chek Connect Diabetes Management App for iOS that contains the fix for the issue was released on August 29, 2016. Patients were instructed to:Âż If you use the Accu-Chek Connect App on an IOS product (iPhone, iPad), please verify that you have installed the newest version of the app.Âż Contact Accu-Chek Customer Care at 1-800-628-3346 if you have additional questions.HCPs were instructed to provide a copy of the UMDC to those patients to whom they have prescribed the Bolus Advisor on the Accu-Chek Connect App. No product returns are expected. Customers with questions may contact Accu-Chek Customer Care at 1-800-628-3346. For further questions please call (317) 521-4312, |  | 7909 | 7,909 downloads of iOS version 1.2.0 of the Accu-Chek Connect Diabetes Management App in the US and 8,775 downloads Internationally | US Distribution to the states of :Alabama, Arizona, Arkansas, California, Colorado, Connecticut, DC, Delaware, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Nebraska, Nevada, New Hampshire, New Jersey, New Mexico, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Virginia, Washington, West Virginia, WisconsinForeign: None provided | Terminated | June 06, 2017 | 265 | 7909
Z-0588-2017 | 75443 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Triage Total 5 Control Level 1, Model No. 88753 |  |  | October 12, 2016 | 2016 | Alere San Diego, Inc. | Encoded with incorrect ranges that are not consistent with the ranges printed on the Expected Value Card. | Employee error | The firm, Alere, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated 10/2016 to customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of and discard the CODE Chip contained within the Alere Triage Total 5 Control Level 1 PN 88753 Lot C3233A; install the replacement QC Sample CODE Chip for Lot C3233AR; share this information with your laboratory staff and retain this notification as part of your laboratory QS documentation; if you have forwarded the product to another laboratory, provide copy of this notification; and complete and return the Customer/Distributor Verification Form via Fax to Alere Technical Services at +1 858 805 8457 within 10 business days. Customer with any questions about the information contained in the notification are instructed to contact:Alere San Diego, Inc.9975 Summers Ridge RoadSan Diego, CA 92121, U.S.A.Phone: 844-320-5124FAX: 858-805-8457E-mail: Responses.ts@alere.comOn 11/10/16 a revised Urgent Medical Device Correction letter was sent to their customers to inform them of the recall. |  | 513 | 513 kits | US Nationwide Distribution | Terminated | July 24, 2017 | 285 | 513
Z-0597-2017 | 75516 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core Computed Tomography X-ray system, Model 728321 |  |  | October 13, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The product label does not include the correct current rating. | Labeling Change Control | The firm will send a customer letter via certified mail and will have the receipts returned for tracking purposes. Also, a Field Service Engineer will implement the Field Change Order on affected systems to resolve the issue. Implementation of the Field Change Order will be monitored. |  | 2 | 2 | Worldwide distribution. The products were shipped to the following states: AL, AZ, CO, GA, IN, KY, LA, MA, MN, NY, OH, PA, SC, TX, and UT.There are five (5) government accounts for this recall: DEPARTMENT OF VETERANS AF, VETERANS WAY AND LAMONT ST, Mountain Home, TN 37684; US ARMY EVANS COMMUNITY H, 1650 Cochrane Cir, Colorado Springs, CO 80913; US VETERANS ADMINISTRATION, 3710 SW US VETERANS HOSPITAL, PORTLAND, OR 97239; VA MEDICAL CENTER, 4100 W 3RD ST, DAYTON, OH 45428; VAMC - PHILADELPHIA, UNIVERSITY &amp; WOODLAND AVE, PHILADELPHIA, PA 19104.There are five (5) Canadian accounts for this recall: FIVE HILLS HEALTH REGION, 55 DIEFENBAKER DR, MOOSE JAW, SASKATCHEWAN S6J 0C2; INSTITUT UNIV CARDIOLOGIE, 2725 CH STE FOY, SAINTE FOY, QUEBEC G1V 4G5; LAKERIDGE HEALTH, 1 HOSPITAL CRT, OSHAWA, ONTARIO L1G 2B9; LAKERIDGE HEALTH BOWMANVI, 47 LIBERTY ST S, BOWMANVILLE, ONTARIO L1C 2N4; MCGILL UNIVERSITY HEALTH, 1051 BOUL DÂżCARIE, MONTRÂżAL, QUEBEC H4A 0A5.There were zero (0) Mexican accounts for this recall.The recalled products were shipped to the following countries: Argentina, Australia, Belgium, China, Egypt, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Republic of Korea, Lebanon, Netherlands, New Zealand, Norway, Russian Federation, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 117
Z-0611-2017 | 75497 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iConnect Access software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | May 25, 2016 | 2016 | Merge Healthcare, Inc. | Software displayed incorrect prior reports in the viewport area, only when more than one prior study (2 or more) was viewed. | Software design | The recalling firm issued a notification letter dated 1/27/2016 via email or via certified mail if the firm did not have an email address for the customer. Nonresponders were issued a second letter dated 5/25/2016 on 5/26/2016. |  | 60 | 60 sites potentially have the affected versions | Distribution was made to medical facilities nationwide. Foreign distribution was also made. | Open, Classified |  |  | 60
Z-0604-2017 | 75555 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition AS, Computed tomography system |  |  | October 19, 2016 | 2016 | Siemens Medical Solutions USA, Inc | A problem can occur with the small light marker windows, used for the positioning laser and are integrated in the front cover of the SOMATOM Definition AS and Definition Flash systems, by either becoming loose or potentially dropping out. There is a risk that a person could touch rotating or electrical parts of the gantry if they were to reach through the window opening. | Nonconforming Material/Component | Siemens' sent a Customer Safety Advisory Notice letter dated October 19, 2016 to affected customers. The letter identified the affected product, problem, potential risks, how to avoid the potential risk related to this issue and actions to be taken. |  | 762 | 762 systems distrtibuted in U.S. | US Nationwide Distribution | Terminated | March 13, 2017 | 145 | 4599
Z-0610-2017 | 75499 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | OMNIBotics Tracker Kit -Model: RM-10000 |  |  | October 14, 2016 | 2016 | OMNI LIFE SCIENCE | Incorrect lot number on outer kit | Error in labeling | The firm, OMNIlife Science, Inc., issued an "URGENT MEDICAL DEVICE ADVISORY NOTICE" dated 10/14/16 by email to its consignees/customers. U.S. and foreign agents were also sent the notification packet via overnight mail. Acknowledgement of receipt is requested. The notice described the product, problem and actions to be taken. The customers were instructed to immediately quarantine lot 19-09-021 for OMNIBotics Tracker Kit RM-1000; overlabel the outside packaging with the label provided following the provided training paperwork and rework instructions; note advisory notice only includes listed product, it does not affect other lots; and complete and return the Acknowledgement and Receipt Form and accompanying FN-102 by Fax to 508-519-3390.Accounts are to hold item RM-10000 with the lot # 19-09-021 until a corrected label is provided for the outer box. The patient labels inside the box are correct.For further information , Regulatory Affairs Specialist at 774-226-1847 or 800-448-6664 Monday through Friday, 8:00 AM to 4:30 PM, Easter Time. |  | 15 | 15 units | Worldwide Distribution: US distribution to OK, MD, TX and country of: Australia. | Open, Classified |  |  | 15
Z-0600-2017 | 75365 | No | MJC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Magic3 Antibacterial Hydrophilic Intermittent Catheter, Male 14 Fr. (30/Box), Sterile |  |  | October 05, 2016 | 2016 | C.R. Bard, Inc. | Misbranding; the product labeled as an Antibacterial Hydrophilic Intermittent Catheter, is instead an Antibacterial Intermittent Catheter (no Hydrophilic Coating). | Mixed-up of materials/components | Bard Medical sent an Urgent Medical Device Product Recall letter dated October 5, 2016 to customers The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine and quarantine and return affected inventory. Customers were asked to complete and return the accompanying Recall &amp; Effectiveness Check Form attached. |  | 10290 | 10,290 units | US Nationwide Distribution in the states of: CA, FL, IL, IN, MN, MO, NC, NJ, NY, OH, OK, PA, TX and UT. | Open, Classified |  |  | 10290
Z-0608-2017 | 75677 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition Flash, Computed tomography x-ray system |  |  | November 02, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is releasing a software update that provides bug-fixes to improve system performance for customers with SOMATOM Definition AS, SOMATOM Definition Edge, SOMATOM Definition Flash Systems with software version VA48A_SP2 and Care Contrast license correction of potential safety issues. | Software design | Siemens mailed a Customer Safety Notice on November 9, 2016, to customers who were affected by the recall to inform them of a possible malfunction with the systems. The notification explained the underlying issue, the effect if has on the operation of the system, potential risks involved; and the actions they plan to implement to correct the issues. Customers with questions were instructed to call 1-800-888-7436.For questions regarding this recall call 610-448-6471. |  | 7 | 7 systems | Nationwide distribution to MI, NY, CA, KY, ND, and NE. | Open, Classified |  |  | 7
Z-0609-2017 | 75187 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Model 208-70-110, Long Thread Tiger Large Cannulated Headless Screw, 7 x 110 mm labeled as "7.0 x 110mm LTTHCS REF 208-70-110 QTY 30 LOT TSL003721 2016-03-29 Trilliant Surgical, LTD" |  |  | September 02, 2016 | 2016 | Trilliant Surgical Ltd. | Screws were identified as having the incorrect part number and lot number marked upon them. The part number incorrectly indicated a length of 115 mm rather than 110 mm. The lot number incorrectly indicated lot number of TSL003765 rather than TSL003721. | Nonconforming Material/Component | The firm sent out written notifcation of the recall on 9/2/16; the letter instructed the consignees to discontinue use of the affected lot, quarantine, and return product of the affected lot in accordance with an included acknowledgement sheet. |  | 11 | 11 | CA, AZ, WI, FL, NJ, OH | Terminated | May 19, 2017 | 259 | 11
Z-0599-2017 | 75629 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | The software did not show unviewed images when the last view was skipped. The physician may potentially misdiagnose when not all images are available for viewing. | Software design | The recalling firm issued a letter dated March 28, 2016 via e-mail or via certified mail to customers. The letter identified the affected product, problem and actions to be taken. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. |  | 534 | 534 sites potentially have the affected versions | Worldwide Distribution - US Nationwide including Canada. | Open, Classified |  |  | 534
Z-0603-2017 | 75606 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis, Angiographic x-ray system |  |  | November 01, 2016 | 2016 | Siemens Medical Solutions USA, Inc | The possibility exists that the semi-automatic focus switch may not work as intended when pressing the footswitch multiple times. | Software in the Use Environment | Siemens mailed a Customer Safety Advisory Notice to affected customers to inform them of a corrective action they plan to perform to prevent a possible hazard to persons and equipment. The notice identified the associated risks, underlying issues, impact on system operation and actions to be taken. |  | 4657 | 4,657 systems Worldwide | US Nationwide Distribution. | Terminated | June 21, 2017 | 232 | 4657
Z-0612-2017 | 75523 | No | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur CA 19-9 Assay (50 test kit)-for use, with the ADVIA Centaur and ADVIA Centaur XP, ADVIA Centaur XPT and ADVIA Centaur CP systems. |  |  | October 19, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | CA 19-9 assay on the ADVIA Centaur CP does not meet the correlation to the CA 19-9 assay on the ADVIA Centaur/XP/XPTsystems as stated in the Instructions for Use. | Process design | Siemens Healthcare issued Urgent Medical Device Correction (UMDC) to Siemens Healthcare Diagnostics customers in the United States on October 19, 2016 to be delivered to customers on October 20, 2016. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on October 18, 2016. The notice informs the customer of the potential to observe a bias between CA 19-9 results generated on the ADVIA Centaur CP and the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT systems. Customers may continueto use the ADVIA Centaur CA 19-9 assay to report patient results on the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT and ADVIA Centaur CP systems. Customers should not use the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT systems interchangeably with the ADVIA Centaur CP when generating CA 19-9 results.For customers using the ADVIA Centaur CP and ADVIA Centaur/XP/XPT for CA 19-9 interchangeably, a look back of patient results over the previous three months is recommended.A Field Correction Effectiveness Check is to be completed, signed and returned Siemens. Questions, contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 18733 | 18,733 kits | NationwideForeign:Albania, Algeria, Angola, Argentina, Armenia, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, Cambodia, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Germany, Great Britain, Greece, Guadeloupe, Hong Kong, Hungary, India, Indonesia, Iraq, Israel, Italy,Japan, Jordan, Kazakhstan, Latvia, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mexico, Myanmar, Nepal, Netherlands, New Zealand, Nicaragua, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Uruguay, Vatican City, Vietnam | Open, Classified |  |  | 48955
Z-0622-2017 | 75661 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion Prime TSX-303A |  |  | October 25, 2016 | 2016 | Toshiba American Medical Systems Inc | Toshiba American Medical Systems (TAMS) is recalling the Aquilion CT system because ithas been reported that the cooling oil in the x-ray tube heat exchanger may leak during system operation. | Device Design | Toshiba American Medical Systems, Inc. sent an Urgent Medical Device Correction letter dated October 25, 2016, to all affected customers who received the Aquilion CT system to inform them that Toshiba America Medical System has received reports that the cooling oil in the x-ray tube heat exchanger may leak during system operation. The letter informs the customers of the description found and corrective action. Customers were instructed to complete the attached customer reply form and fax it to (877) 349-3054 or send it to raffairs@tams.com. Customers with questions are instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968.For questions regarding this recall call 714-730-5000. |  | 291 | 291 units | Nationwide Distribution | Terminated | August 19, 2017 | 298 | 291
Z-0623-2017 | 75145 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Greatbatch Stiffer Coaxial Micro-Introducer. Model #/ Name: 10904-001 Stiffer Coaxial Micro-Introducer 4FR; 10904-002 Stiffer Coaxial Micro-Introducer 5FR. |  |  | September 08, 2016 | 2016 | Greatbatch Medical | Greatbatch Medical has identified an issue with the Stiffer Coaxial Micro-Introducer 4FR and 5FR after a field complaint reported that the rotating luer detached from the dilator hub during use. | Process design | Greatbatch sent a "Urgent Medical Device Recall" letter dated September 7, 2016 to all affected customers .The letter described the product, "Reason for Field Correction Action", "Risk to Health" , "Replacement Devices", &amp; "Instructions to Customers" . Advised customers to quarantine the product and contact end users and share the recall information with them. Requested consignees to complete &amp; return the Field Recall Response Form. For questions contact your local Greatbatch Medical representative, contact us at 1-763-951-8235, or e-mail us at FieldActionCenter@Greatbatch.com. |  | 23075 | 23,075 | US: NY, TX. | Open, Classified |  |  | 23075
Z-0614-2017 | 75267 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge RadSuite,f/k/a: Emageon Advanced Visualization , RadSuite versions 8.30.7.8, 8.30.7.9, 8.30.8, and 8.31. |  |  | July 14, 2016 | 2016 | Merge Healthcare, Inc. | It was reported by a customer that RadSuite images are not appearing as readable images. This product malfunction has the potential to result in delay in care and a possible re-exposure of the patient to radiation (x-ray). | Software design | Merge sent an Urgent Medical Device Recall letter dated October 19, 2016,(Recall #2016-075) to all affected customers. Initially, before it was considered a recall, the firm provided customers with a software upgrade to fix the issue. The letter described the Product, Issue, Potential Harm, Containment by the Customer/User, Action by Merge and Actions by Customers. Requested consignees to complete and return the response form. For questions email recall@merge.com. Customers with questions were instructed to send an email to recall@merge.com.For questions regarding this recall call 262-367-0700. |  | 29 | 29 | Nationwide Distribution including AL, CA, FL,IN, MI, MO, NY, PA, TN, and WI. | Open, Classified |  |  | 29
Z-0633-2017 | 75367 | No | OJV | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Waste Management Kit. Catalog Number K10-04381AP |  |  | September 23, 2016 | 2016 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. announces a voluntary field action for various Merit Medical Procedure Kits due to a compromised sterile barrier. | Process control | Consignees notified via letter on 09/23/16 to quarantine any devices and discontinue use, ensure all personnel to whom the devices were distributed are made aware of the field action, and return affected devices to Merit Medical. |  | 30 | 30 units | Nationwide to VA, UT, and NC. Other foreign consignee Japan. No Canadian and VA/govt/military consignees. | Terminated | March 27, 2017 | 185 | 431
Z-0636-2017 | 75306 | No | FDF | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Pentax Video Colonoscope is intended to be used with a PENTAX video processor (including light source), documentation equipment, Monitor, EndoTherapy Device such as a Biopsy Forceps, and other ancillary equipment for endoscopy and endoscopic surgery within the lower digestive tract including the anus, rectum, sigmoid colon, colon and ileocecal valve. |  |  | August 15, 2016 | 2016 | Pentax of America Inc | Pentax is initiating the product correction of several model numbers of Pentax's Colonoscope that were shipped between the clearance date of April 2014 and September 2015 because they contain the operation and reprocessing IFU's that were in use prior to the 2014 clearances and therefore did not include the cleared operation IFU, the updated reprocessing IFU or the dedicated reprocessing adaptor (for applicable models). | Other | Pentax sent as "Urgent Field Correction" and Field Correction Response Form dated August 15, 2016 to their affected customers. |  | 5090 | 5,090 (US) and 14,209 (OUS) | US Nationwide and Japan | Open, Classified |  |  | 5090
Z-0627-2017 | 75454 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | DePuy Spine SKYLINE ANTERIOR CERVICAL PLATE SYSTEM VARIABLE SELF-TAPPING SCREW 4.0 x 26mm |  |  | October 12, 2016 | 2016 | DePuy Spine, Inc. | SKYLINE Variable Ti screws may have contained SKYLINE Constrained Ti screws within the packaging. | Under Investigation by firm | The firm, DePuy Synthes Spine, sent an "URGENT! VOLUNTARY PRODUCT CORRECTION NOTIFICATION" letter dated 10/12/2016 to its Customers. Sales consultants supporting these customers received an email notification with required reply forms. The letter described the product, problem and actions to be taken. A DePuy Synthes Spine Sales consultant has been notified and will be reviewing inventory to ensure correct screw placement within your set. The customers were instructed to review, complete, sign and return the attached business reply form to DePuy Synthes Spine via email to: DPYUS-SpineFieldActions@its.jnj.com within 5 business days of receipt of this notification; forward the notice to anyone in your facility that needs to be informed; if devices have been forwarded to another facility, contact that facility and provide them with notice, and keep a copy of the notice.Questions about this recall, please contact Recall Coordinator, DePuy Synthes Spine (via telephone: 508.828.3647 or via email: DPYUS- SpineFieldActions@its.jnj.com). |  | 6 | 6 units | US Distribution to states of: CA, ID, IL, IN, KS, KY, LA, MO, NC, OR, PA, TX, VA, WA, and WV. | Open, Classified |  |  | 72
Z-0637-2017 | 75502 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 2.5 mm Reaming Rod with Ball Tip &amp; Extension/950 mm, Sterile |  |  | October 03, 2016 | 2016 | Synthes (USA) Products LLC | The peel pouches for the affected reaming rods and extraction hooks are delaminating. | Nonconforming Material/Component | DePuy Synthes mailed a Customer Safety Advisory Notice on October 19, 2016 to affected customers advising them of the malfunction, potential risks and the actions needed to be taken to avoid injury, as they work on a solution to correct the problem. |  | 1951 | 1,951 units | Distributed Nationwide | Terminated | July 06, 2017 | 276 | 11305
Z-0650-2017 | 75540 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ACTIS Flex Reamers SZ 0/1; 2/3; 4/5; 6/7; 8/9; 10/11; 12, Product is sold non-sterile, these instruments are packed non-sterile, within a rigid tube with foam protectors, with label and IFU |  |  | October 25, 2016 | 2016 | Depuy Orthopaedics Inc. | DePuy Orthopaedics, Inc. is issuing a voluntary recall for all lots of the ACTIS Flex Reamers due to the instruments breaking and potentially leaving pieces in the patient. | Under Investigation by firm | Depuy Orthopaedics initiated a recall on October 25, 2016. Notices were sent to the US Distributors via email on October 25, 2016. Customers were instructed to do the following: 1. Immediately Inspect Inventory, 2. Immediately Return US Distributor Inventory: If any affected instruments are found in a US Distributors inventory, return to:DePuy Synthes Joint Reconstruction, a division of DePuy Orthopaedics, Inc.3. To expedite the return and credit process: Be sure to reference H16-14 on all return paperwork and/or online return forms and on the outside of the box when returning recalled lots of the instrument.4. Upon receipt of the affected instruments, affected instruments will be credited against the US Distributors B&R; budget.5. Within 5-Business Days: Complete the Distributor Card or for the Reconciliation Form: Within 5-business days of initiation, the completed Reconciliation Form should be returned to DePuy Orthopaedics, Inc.6. Retain copies of all field action documents in customer filesFor further questions please call (574) 371-4917. |  | 137 | 137 | US Distribution to the states of: CA, KY, LA, MD, ME, NC, NH, NY, TX, VA, WAForeign: Austria, Ireland, and JapanVA/DOD: | Terminated | June 06, 2017 | 224 | 137
Z-0649-2017 | 75011 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Custom Fluid Management Set, Sterile, EO, Rx Only |  |  | August 12, 2016 | 2016 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. announces a voluntary field action for Fluid Administration Sets due to a molding defect in the drip chamber. | Component design/selection | Consignees were notified via an Urgent Product Recall Notice on 8/12/16 to: 1. Please immediately determine if any of the devices are within your facility, quarantine them, and discontinue use. 2. Ensure that all personnel to whom the devices were distributed are made aware of this field action. 3. Complete the Customer Response Form and return it to Merit Medical Systems with any affected products. If you have no affected products, please complete the Customer Response Form. 4. E-mail the completed Customer Response Form to Customer Service at RESPONSE@merit.com, or fax to (801) 316-4880. 5. If you have any questions concerning this communication or for assistance arranging product return, please dont hesitate to contact your Merit Sales Representative or Vicki Biehn at vbiehn@merit.com or at (801) 208-4260. |  | 107341 | 107,341 total devices | Nationwide Distribution. | Terminated | May 09, 2017 | 270 | 107341
Z-0648-2017 | 75503 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | AQUIOS Lyse Reagent Kit, Catalog No. B23538 |  |  | October 19, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter is recalling the AQUIOS Lyse Reagent Kit because it may lead to an increased generation of notifications and flags when used on the AQUIOS CL when running the AQUIOS Tetra application. The issue could lead to an increase in the number of samples requiring review and re-runs with a potential delay to the reporting of results. Depending on the specimen age, a new blood specimen may be required. | Under Investigation by firm | Beckman Coulter sent an Urgent Medical Device Recall letter dated October 19, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding the notice are instructed to contact our Customer Support Centerhttp://www.beckmancoulter.com/customersupport/support; 1-800-369-0333 in the US and Canada; outside the US and Canada contact local Beckman Coulter Representative. |  | 2367 | 2367 units total (207 units in US) | Worldwide Distribution - US Nationwide and the countries of Belgium, Botswana, Canada, Chile, Columbia, Czech Republic, Denmark, Finland, France, Germany, Ghana, Italy, Malaysia, Netherlands, Poland, Romania, Slovakia, South Africa, Spain, Sweden, Switzerland, United Kingdom and Zambia | Terminated | February 24, 2017 | 128 | 2367
Z-0645-2017 | 75478 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | LIAISON Estradiol II Gen, 310680, 310680-CN |  |  | July 13, 2016 | 2016 | Diasorin Inc. | DiaSorin is issuing an urgent medical device correction for the LIAISONÂż Estradiol II Gen assay. Due to the risk of a recently identified cross reactivity, the LIAISONÂż Estradiol II Gena assay should not be used to test patients being treated with fulvestrant (Faslodex). | Device Design | DiaSorin Inc. sent an Customer Notification letter dated July 13, 2016, to all affected customers. The letter described the problem and the product involved in the recall. The letter listed the "Recommended Actions" and requested customers to complete and return the response portion. For questions contact Product Support at 1-800-328-1482. |  | 10739 | 10,739 | USA (nationwide) Distribution and Worldwide Distribution including the states of : AR, AZ, CA, FL, ID, NJ, NV, NY and PA.., and to the countries of : Saudi Arabia, Colombia, Mauritius, Iraq, Mongolia, Greece, Albania, Austria, Bosnia, Bosnia Herzegovina, Belarus, Bulgaria, Macedonia, Poland, Serbia, Dominican Rep., Tunisia, Paraguay, Hungary, Jordan, Egypt, Belgium, Caribbean, Germany, Brazil, Mexico, Italy, India, South Korea, Cyprus, Bangladesh, Morocco, Chile, Bahrain, Palestine, Utd.Arab.Emir., Kosovo, Uruguay, Venezuela, Hungary, Australia, Vietnam, Sri Lanka, Turkey, Vietnam, Romania, Mauritius, Ecuador, Lebanon, Nepal, Argentina, Hong Kong, China and Canada. | Open, Classified |  |  | 10739
Z-0646-2017 | 75511 | No | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Pathfast D-Dimer, Reference number: PF1051-KUS |  |  | August 16, 2016 | 2016 | Lsi Medience Corporation | Two complaints were received about decrease in sensitivity (control value and/or calibration value) of the product. | Nonconforming Material/Component | Customers were notified by letter sent via email on/about August 1, 2016.They were instructed to stop using and destroy the affected kits. They were to perform a retrospective analysis of the patients tested on the batch and investigate further. A Destruction Certificate was attached to be completed and returned to Verbatim Americas, LLC.For further questions please call (704) 547-6517. |  | 2520 | 42 boxes (each containing 60 test cartridges/one kit contains 10 reagent cartridge trays and one tray has 6 cartridges). | US Distribution to the states of : AL, FL, ID and MA. | Open, Classified |  |  | 2520
Z-0651-2017 | 75324 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | PeroxiClear Peroxide Solution, (3oz. and 12 oz., 12 oz. Twin Pack, 2x12 oz. Twin Pack for Target stores, 12 oz. for Walmart Vision Center and 3 oz. Starter Kit US), Manufactured by Bausch + Lomb, Rochester, NY 14609. |  |  | September 19, 2016 | 2016 | Bausch &amp; Lomb Inc Irb | Product was found to contain excess residual levels of hydrogen peroxide. | Under Investigation by firm | The firm, Bausch &amp; Lomb, sent an "URGENT MEDICAL DEVICE VOLUNTARY RECALL" letter dated 9/29/2016 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to follow the steps outlined in the letter: review your inventory including floor displays; contact Stericycle, the firm conducting the product recall on behalf of Bausch &amp; Lomb, to arrange for return of the product; and complete and return the Recall Acknowledgement Form to Stericycle Inc,, even if you do not have any of the product in your inventory, via Fax to 1-877-779-7814, email: BauschandLomb5833@Stericycle.com or phone: 1-877-877-0481. If you have any questions about this voluntary recall process or would like additional information, please contact Stericycle, Inc., the firm conducting the recall on behalf of Bausch &amp; Lomb at 1-877-877-0481. |  | 4652822 | 4,652,822 units | Worldwide Distribution-U.S. Nationwide and country of: Canada. | Open, Classified |  |  | 4652822
Z-0375-2017 | 75551 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare System X-ray Tomography Computed : |  |  | March 15, 2016 | 2016 | Philips Electronics North America Corporation | It was discovered the patient support tabletop may fail to move in the horizontal direction when it is commanded to do so during CT scans due to a potential mechanical linkage failure. | Radiation Control for Health and Safety Act | Philips planned action: 1. Philips issued a customer notification letter on October 13, 2016. The letter includes a statement that the corrections will be made free of charge.2. Philips will initiate a field corrective action to address the issue.3. Philips will contact the customer and schedule an inspection to determine if the unit requires repair.4. If repair is required, the couch horizontal linkage assembly will be replaced.For further questions please call (978) 659-3000. |  | 1363 | 1363 | USA(nationwide) Distribution to the states of : AZ, NY, GA, UT, TN, TX, AR, NV, Il, FL, MS, LA, MD, MO, MA, OK, KY, WY, SC, ND, NE, DE, MI, WA, CT, RI, KS, OR, IN, PA, CA, NJ, NC, MN, AL and IA. | Open, Classified |  |  | 1363
Z-0653-2017 | 75373 | No | OCL | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | COBRA Fusion 50 Ablation System, COBRA Fusion 150 Ablation System, and COBRA Fusion Magnetic Retriever System. Model Numbers: 700-001, 700-001S, 700-001MI, 700-002, and 700-003. |  |  | September 22, 2016 | 2016 | AtriCure, Inc. | Due to product design and process control, part of the device can separate and remain behind when the device is removed that can potentially cause permanent impairment or damage to body function or structure. | Component design/selection | An Urgent: COBRA Fusion Ablation System Advisory Notification Letter with Acknowledgement Form was sent to all users. The letter described the issue and identified the affected device. Customers were asked to examine their inventory and determine if any affected product was on hand. All users, including those that the device may have been further distributed should be notified with a copy of the letter. Customers are to contact AtriCure customer service at 1-866-349-2342, and select option 6 to return any affected devices. The Device Notification Acknowledgement Form should be completed and returned as soon as possible. |  | 5263 | 5,263 | Worldwide Distribution -- USA, in the following states AL, AR, AZ, CA, CO, CT, GA, FL, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, WA, WI, &amp; VA; and, to the countries of Albania, Australia, Belgium, France, Germany, Ireland, Italy, Norway, Netherlands, Poland, Sweden, Turkey, United Arab Emirates, &amp; United Kingdom. | Open, Classified |  |  | 5263
Z-0654-2017 | 75641 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | The Hemotherm 400CE 115v, containing power board p/n 39945 |  |  | November 07, 2016 | 2016 | Cincinnati Sub-Zero Products Inc | The device may not maintain water flow or temperature control. | Material/Component Contamination | Cincinnati Sub-Zero sent an Urgent Medical Device Correction letter date November 7, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm has notified all customers to discontinue use and contact CSZ. Upon receipt of the new board, immediately remove and discard the affected power board and replace it with the newly provided power board. The complete and return the customer response form. For further questions, please call (513) 772-8810 |  | 50 | 50 | US Distribution to the states of : CA, FL, GA, MA, MN, NC, NJ, OH, PA and SC. | Open, Classified |  |  | 50
Z-0652-2017 | 75484 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista Chemistry 1 Calibrator |  |  | September 15, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens determined that discrepant low calcium recovery of QC and patient samples can occur with Dimension VistaÂż CHEM 1 CAL lots 5GM081 and 5GM082. | Under Investigation by firm | Siemens mailed an Urgent Medical Device recall notification letter on September 19, 2016 to notify affected customers who purchased the product and to make them aware of the potential risk to health and the actions to be taken. For questions regarding tis recall. For questions regarding this recall call 1-800-441-9250. |  | 9935 | 9,935 units | Worldwide Distribution - US (Nationwide) Internationally to Australia, Austria, Bahamas, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Italy, Japan, Netherlands, New Zealand, Norway, Portugal, Qatar, Saudi Arabia, South Korea, Slovakia, Slovenia, Spain, and Switzerland. | Open, Classified |  |  | 9935
Z-0655-2017 | 75558 | No | HHM | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | SURESOUND22US Multipack (US Distribution) containing 2 packs of six devices. |  |  | October 28, 2016 | 2016 | Hologic, Inc | Breach in sterile barrier compromising product sterility | Package design/selection | Hologic notified customers in the United States through Certified letter on 10/28/16. The letter contains instructions with the option of returning or destroying the product, provides contact information, and includes a postage paid return card for the consignee to complete and return to Hologic.Questions or require clarification of the response process please contact 1- 800-442-9892 and selection option 6 and then option 2 for surgical product technical support.For questions regarding this recall call 508-263-8920. |  | 600 | 600 units (US) ; 24 units (OUS) | Worldwide Distribution - US (Nationwide) and Internationally to Finland and Austria. | Terminated | August 18, 2017 | 294 | 600
Z-0656-2017 | 75580 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Flexible Shaft, Orthopedic manual surgical instrument |  |  | October 31, 2016 | 2016 | Synthes (USA) Products LLC | Non-conforming material used. | Under Investigation by firm | The firm, DePuy Synthes, mailed an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated 11/2016 to customers who were affected by the recall. The letter described the product, problem and actions to be taken. The customers were instructed to immediately review your inventory to identify and quarantine all affected products; if you have product to return call DePuy Synthes at 1-800-479-6329 to obtain an RA; complete and return the Verification Section with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132 within 30 business days; send a copy of the completed Verification section via fax to 888-257-7905 or scan/email: Synthes3503@stericycle.com, even if you have no product to return; forward notice to anyone in your facility that needs to be informed; if any of the affected product has been forwarded to another facility, contact that facility to arrange return and keep a copy of the notice.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 198 | 198 units | US Nationwide Distribution. | Terminated | June 12, 2017 | 224 | 198
Z-0660-2017 | 75664 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System; Used to make treatment plans for patients with prescriptions for external beam radiation therapy. |  |  | November 03, 2016 | 2016 | Elekta, Inc. | Incorrect Dose when using the reset function. | Software design | Important Field Safety Notice 382-01-MON-006 was sent to all affected customers on 11/3/2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. An acknowledgement form was included which is to be returned to Elekta. |  | 671 | 671 | AK, AZ, CA, CO, FL, GA, IL, IN, MA, MI, MN, MO, NV, NH, NY, NC, OH, OK, OR, PA, SD, TX, UT, VA, WA, WI, Algeria, Australia, Austria, Brazil, Canada, China, Cuba, Cyprus, Czech Republic, Ecuador, Egypt, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Laos, Malaysia, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Russia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, Venezuela, | Open, Classified |  |  | 671
Z-0658-2017 | 75553 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ConforMIS iTotal Cruciate Retaining (CR) Knee Replacement System, Sterile |  |  | October 18, 2016 | 2016 | ConforMIS, Inc. | Mislabeled: incorrect serial number on the F4 A/P cut guide iJig. | Labeling Change Control | ConforMIS notified the sales representative on 10/18/16 after it became aware of the device mix-up. The sales representative was requested to return any remaining components from iTotal CR kit S/N: 0381042. |  | 2 | 2 units | Distributed to the states of CO and NV. | Completed |  |  | 2
Z-0663-2017 | 75617 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Bridge Occlusion Balloon. Model: 590-001 |  |  | October 31, 2016 | 2016 | Spectranetics Corporation | Spectranetics Corporation announces a voluntary field action for the Turbo Elite Atherectomy Catheter, ELCA Coronary Atherectomy Catheter, and Bridge Occlusion Balloon due to potentially compromised integrity of the outer sterile packaging. | Packaging | Spectranectics sent an Urgent Medical Device Recall letter dated October 31, 2016, to all affected consignees. Consignees were instructed to discontinue use of the potentially impacted lots and return the product for replacement. Consignees were also instructed to complete the attached Acknowledgement and Receipt Form to facilitate the product return of any remaining product in inventory. Customers with questions were instructed to call 1-800-231-0978, Option 2. For questions regarding this recall call 719-651-8517. |  | 873 | 873 units | Worldwide Distribution - US (nationwide and PR) Internationally to Austria, Belgium, Cyprus, Denmark, France, Germany, Hungary, Italy, Netherlands, Norway, Poland, Spain, Sweden, Switzerland, United Kingdom, Australia, Bahrain, Brazil, Canada, Chile, Colombia, Dominican Republic, Israel, Japan, Kuwait, Taiwan, Thailand. | Terminated | April 18, 2017 | 169 | 37276
Z-0668-2017 | 75718 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco TRP System |  |  | November 17, 2016 | 2016 | Elekta, Inc. | Contours are not handled correctly if the CT dataset contains slices with DICOM Z coordinates of xx.x50mm. | Software design | Elekta sent an Important Field Safety Notice 382-01-MON-007 was sent to all affected customers on November 17, 2016, to all affected customers. The notice informs users of the specific product version numbers affected by the issue, and any work arounds that can be used to avoid the issue. The notice included a reply form which is to be returned. |  | 2148 | 2148 | Worldwide Distribution - US including AK, AZ, CA, CO, FL, GA, IL, IN, KY, LA, MD, MA, MI, MN, MO, MT, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, SD, TX, UT, VA, WA, WI, Puerto Rico and Internationally to Algeria, Antiqua &amp; Barbuda, Argentina, Australia, Austria, Bahrain, Belarus, Botswana, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cuba, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Laos, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Namibia, Nepal, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Thailand, Tunisia, Turkey, Unite Kingdom, Ukraine, Venezuela, Vietnam, and Zimbabwe. | Open, Classified |  |  | 2148
Z-0666-2017 | 75336 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-MR1 Ventilator: Catalog# 161010 |  |  | September 26, 2016 | 2016 | Hamilton Medical, Inc. | Oxygen tubing and the oxygen connector of the Hamilton-MR1 could potentially become loose during the preparation for ventilation. | Device Design | Hamilton Medical sent a letter dated September 6, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Hamilton Medical will update each affected Hamilton-MR1 ventilators by installing a new oxygen mixer mounting plate with the correct screws at no cost to each facility. A Hamilton Medical Field Service Technician will contact customers soon to schedule a convenient time to install the hardware upgrade. Customers may also call 1-800-426-6331, option #2 to be issued an RGA for return to the service center for the upgrade kit to be installed. Customers with questions should call 817-909-0308. |  | 44 | 44 units | Nationwide Distribution | Open, Classified |  |  | 44
Z-0667-2017 | 75639 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Z6Ms transesophageal transducers used on ACUSON SC2000 ultrasound systems, Z6Ms model 10436113. |  |  | November 07, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | Firm received reports about "acquisition errors" and the deterioration of material covering the articulating section of the Z6Ms transesophageal transducer. | Device Design | Siemens sent a Customer Safety Advisory Notification letter to all affected customers on November 7, 2016 by certified mail advising them of the issue and identified steps that can be taken to avoid potential risk of this issue. This letter has been released first in English while translations are being completed. Once the translations are completecustomers in the countries who require translations will receive the letter.Prior to this mailing, ECO #641310 was released on July 5, 2016 and implemented into forward production on July 28, 2016 (SN 63028009) to correct the susceptibility to Gastro-flex circuit corrosion and electro migration. The ECO changed standard copper traces and traditionally applied soldermask on the circuit to fully gold plated traces with a dry film solder mask.ECO #645821 was released on September 8, 2016 date and implemented into forward production on September 8, 2016 (SN 63530013) to correct susceptibility to articulation sleeve damage. The ECO changed the articulation sleeve material from C-Flex to Viton.For questions regarding this recall call 650-694-5398. |  | 33 | 33 | Worldwide Distribution - US (nationwide) Internationally to AUSTRIA, GERMANY, SOUTH AFRICA, NETHERLANDS, MAURITIUS, MEXICO, and ITALY,JAPAN, SWEDEN | Terminated | May 26, 2017 | 200 | 33
Z-0664-2017 | 75334 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Merge FlexConnect software, a component of Merge LIS. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Communication protocols interfacing with the affected software version with some select instruments were not properly handled, preventing data captured by these instructions to be processed by the rest of the LIS software. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/30/2016. Nonresponders were issued a second letter dated 5/25/2016 via email on 5/26/2016. |  | 324 | 324 sites potentially have the affected version | Nationwide Distribution | Open, Classified |  |  | 324
Z-0665-2017 | 75564 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo software. |  |  | April 30, 2014 | 2014 | Merge Healthcare, Inc. | The application may crash during the cath lab procedure. | Software design | Merge sent an Important Alert letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to contact Merge Support at 866-261-5761 or support@merge.com. For questions regarding this recall call 877-741-5369. |  | 192 | 192 sites potentially have the affected software | Nationwide Distribution | Open, Classified |  |  | 192
Z-0673-2017 | 75569 | No | CDS | Urea nitrogen test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | MSS CASSETTE GLU/LAC/UREA (BUN) as a part of the |  |  | October 31, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Incorrect Urea results may be generated sporadically on cobas b 221&lt;6&gt;Roche OMNI S6 system and cannot be detected by QC. QC measurement results can be affected as well. This issue is only detectable if comparison result of a second measurement is available. The cause for this issue is an error in the software measurement procedure.In children, decreased urea results can point at inborn disorders or urea cycle. Therefore, false normal urea values could lead to delayed/incorrect diagnosis of such a disorder.Besides the patients most at risk, the risk for the overall patient population with falsely low or high values is considered remote. No severe adverse health effects are expected; but, unnecessary diagnostic measures may be initiated.Falsely normal results may result in further necessary examination not being carried out, with the possibility of a delay in the diagnosis and treatment. A medical risk cannot entirely be excluded. | Software design | Roche sent an Urgent Medical Correction letter dated October 31, 2016, to all affected customers. The recall notifications included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Actions Required" Use the workaround outlined in this Urgent Medical Device Correction (UMDC) for Ureameasurement until the updated software version is available." If your facility has distributed the affected product to another site, please ensure this UMDC isprovided to that site." Complete the attached fax form and fax it to 1-877-766-7471." File this UMDC for future reference.QuestionsPlease contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-2272 if you have questions about the information contained in this UMDC. |  | 97 | 97 | Nationwide Distribution including ALCA CO CT FL IL IN KS MD MI MO NC NE NJ NM NY OH OK PA PR SC TX VA and WI | Open, Classified |  |  | 97
Z-0669-2017 | 75370 | Yes | NIP |  |  | Cardiovascular | PMA | MISAGO RX Self Expanding Peripheral Stem |  |  | August 04, 2016 | 2016 | Terumo Medical Corp | During recent routine testing, the MISAGO Peripheral Stent was found not to conform to product specification related to stent diameter at the proximal/distal ends and/or the shape of the stent. | Device Design | Terumo Medical Corporation issued an Urgent Product Recall Bulletin dated August 4, 2016, via Fed Ex.Customers must immediately discontinue use and quarantine any available stock of the Misago peripheral stent to prevent further use. The customers should complete the enclosed Inventory Confirmation form as soon as possible and email or fax the completed formto misago.rc@terumomedical.com or (734) 330-2034, Attn: MISAGO Recall.Customers are encourage you to contact Terumo Medical Corporation with any questions or concerns-Terumo Medical Corporation Quality Department: 1.800.283.7866, Press 4 for list of departments, Press 7 for Quality Assurance; Monday  Friday, 8 a.m.  5 p.m. ET- Email Quality Assurance: misago.rc@terumomedical.com- Customer Care: 1.800.888.3786- Contact your local Terumo representative |  | 7965 | 7,965 units | Nationwide | Open, Classified |  |  | 7965
Z-0671-2017 | 75311 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter SIGMA Spectrum INFUSION PUMP with Master Drug Library, V8 Spectrum Infusion Pump, Product Code 35700BAX2, Unique Device Identifier GTIN 00085412091570, Rx ONLY -- Baxter Healthcare Corporation, Medina, NY 14103 --- NOTE: The speakers are a part of the rear case assembly located in the back of the infusion pump. The speaker is intended to generate audible tones during each keystroke, and during alarm conditions. However, some speakers have a variation in their internal wiring that increases the potential for the speaker to fail prematurely. Spectrum Infusion Pumps with non-functional speakers will continue to display visual alarm indications, and will otherwise continue to operate as intended. |  |  | September 28, 2016 | 2016 | Baxter Healthcare Corp. | A correction was issued for the SIGMA Spectrum Infusion Pump with Master Drug Library in order to replace speakers installed in pumps or included in spare part rear case assemblies that have an increased potential to fail prematurely. | Process control | Baxter sent an Urgent Medical Device Correction letter dated September 28, 2016, to all affected customers. On 10/27/2016, Baxter Healthcare sent a f/u UMDC Letter (dated 10/27/2016) to customers. The f/u letter contains the device correction plan which states that Baxter will be replacing at no charge the rear case assembly on all affected Sigma Spectrum Infusion Pumps, as well as affected stand-alone rear case assemblies shipped to customers as spare parts.For questions, please contact your Baxter sales representative, or Baxter Healthcare Medina at 800-356-3454 (choose option 1) Monday through Friday, between 7:00 am and 7:00 pm ET. |  | 15517 | 15,517 units (US: 9560; Canada: 5957) | Nationwide Distribution, Puerto Rico and Canada | Open, Classified |  |  | 21932
Z-0367-2017 | 75043 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Greatbatch, Offset Cup Impactor with the following Parts and Catalogue #: |  |  | August 02, 2016 | 2016 | Greatbatch Medical | On December 17, 2015, Greatbatch discovered that the Standard Offset Cup Impactor with a POM-C handle grip failed sterility testing when housed in a dedicated instrument case using the sterilization parameters identified in the Standard Offset Cup Impactors Instructions for Use (IFU). | Under Investigation by firm | Greatbatch sent an Urgent Medical Device letter dated August 2, 2016, via FedEx to all affected consignees. The letter described the Reason for Field Correction Action, Risk to Health, and Instructions to Customers. Advised consignees to" immediately inform those responsible for sterilization of the devices to utilize the updated sterilization instructions, and post the attached notice where it is clearly visible to personnel who perform sterilization activities". Also instructed consignees to notify their customers and to complete and return the Field Correction Action Response Form. For questions contact your local Greatbatch Medical representative or 1-763-951-8235 or e-mail FieldActionCenter @Greatbatch.com. |  | 4478 | 4478 (2906 US, 1572 OUS) | Worldwide Distribution - US including CA, IN, KS,NJ, NC, TN, TX, and UT and Internationally to ITALY, JAPAN, FRANCE, SPAIN, GERMANY, NETHERLANDS, GREECE, LUXEMBOURG, SOUTH KOREA, IRELAND, UNITED KINGDOM, AUSTRALIA, AUSTRIA, BELGIUM, SWITZERLAND, NORWAY, SWEDEN, and TAIWAN. | Terminated | July 27, 2017 | 359 | 4478
Z-0676-2017 | 75637 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ARROW OnControl Bone Marrow Biopsy System Comprehensive Tray, Sterile, The Arrow OnControl Bone Marrow Aspiration System is intended for bone marrow aspiration of the iliac crest of adult and pediatric patients. The ArrowÂż OnControl Bone Marrow Biopsy System is intended for bone marrow core biopsy of the anterior or posterior iliac crest of adult patients |  |  | November 08, 2016 | 2016 | Teleflex Medical | Sterility: Due to a potential incomplete seal on the outer sterile package. | Device Design | Teleflex sent an Urgent Medical Device Recall Notification letter dated November 8, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product, complete the enclosed Recall Acknowledgement Form and fax to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. Customers with questions were instructed to call their local sales representative or Customer Service at 1-866-246-6990.For questions regarding this recall call 919-433-4940. |  |  |  | Worldwide Distribution - US (nationwide) Internationally to Brazil, Canada, Ecuador, Guatemala, Mexico and Peru. | Open, Classified
Z-0674-2017 | 75589 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Adiprep Adipose Transfer System. Product Codes: 51431; 51432; 51433; 514001-347. |  |  | November 03, 2016 | 2016 | Terumo BCT, Inc. | Terumo BCT announces a voluntary field action for the Harvest : Adiprep Adipose Transfer System because the affected product is incorrectly labeled and has the potential for operator confusion. | Error in labeling | Consignees will be notified on 11/07/16 via letter to contact the North America Terumo BCT support center at +1.877.842.7837 or +1.303.542.5050 if they have any affected product to issue a Return Goods Authorization for product return. |  | 660 | 660 | Nationwide, Canadian, and other foreign consignees. No VA/govt/military consignees. | Open, Classified |  |  | 660
Z-0683-2017 | 75063 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Lineage Liners: |  |  | June 21, 2016 | 2016 | MicroPort Orthopedics Inc. | Revision rate trends from within the MicroPort Orthopedics Inc. database evaluated for the modular head Metal on Metal THA systems show an increasing overall trend from 2009 to present and it was found that there was a specific hazard/harm for suspected tissue reaction to metal debris over one percent for modular head CONSERVE(R) systems and DYNASTY(R)metal acetabular liner systems. | Device Design | MicroPort sent an Urgent Field Safety Notice dated June 22, 2016, to all affected customers notifying them of the change in the instructions for use on July 7th, 2016 by email. The firm also notified US hospitals where the implantation surgeries occurred within the last twelve (12) months by letter on July 13th, 2016 delivered by FedEx Priority. Customers were asked to complete and return the enclosed form even if they do not have any affected product to return by fax to 1-901-451-6032 or PostMarket@ortho.microport.com.For questions regarding this recall call 901-867-4771. |  | 5009 | 5,009 units | Worldwide Distribution - US (nationwide) Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Cyprus, Czech Republic, Denmark, Finland, France, Georgia, Germany, Greece, Iran, Italy, Jamaica, Japan, Lithuania, Malaysia, Philippines, Poland, Portugal, Russia, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Netherlands, Turkey, Arab Emirates, United Kingdom, and Vietnam. | Terminated | July 01, 2017 | 375 | 513126
Z-0690-2017 | 75627 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FMsealer Open Shears. Model FM3001 |  |  | October 24, 2016 | 2016 | Domain Surgical, Inc. | Domain Surgical, Inc. announces a voluntary field action for the FMsealer Open Shears due to fixed jaw fractures during clincal usage. | Other | Consignees were notified on 10/24/2016 via letter to return affected products to Domain Surgical, Inc. |  | 245 | 70 units US; 175 units OUS | Nationwide and other foreign consignees. No Canadian or VA/govt/military consignees. | Terminated | April 11, 2017 | 169 | 245
Z-0682-2017 | 75573 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | STÂżCKERT HeaterCooler System 3T; Item Number 160281, HeaterCooler 3T, 240V/60Hz; Item Number 160282, HeaterCooler 3T, 208V/60Hz; Item Number 160285, HeaterCooler 3T, 120V/ 60Hz |  |  | August 01, 2014 | 2014 | Sorin Group USA, Inc. | On August 1, 2014 Sorin Group USA, Inc. informed consignees via letter to review the disinfection and monitoring practices and to strictly adhere to the instructions in the Instructions for Use ( IFU ). | Material/Component Contamination | Sorin Group sent an Important Information letter dated August 1, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by then consignee. Consignees were instructed to review the disinfection and monitoring practices and to strictly adhere to the instructions in the Instructions for Use ( IFU ). Consignees with questions were instructed to call 303-467-6527 or 303-467-6503.For questions regarding this recall call 281-228-7260. |  | 208722 | 1,618 units in the US; 129 units outside of US | Worldwide Distribution - US (Nationwide and Puerto Rico) and Internationally to Canada, Colombia, Ecuador, El Salvador, Germany, Honduras, Mexico, Saudi Arabia, Trinidad, and Tobago. | Terminated | April 20, 2017 | 993 | 208722
Z-0681-2017 | 75694 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes 0.76 mm Drill Bit/Stryker J-Latch/with 5mm Stop/44.5 mm |  |  | November 01, 2016 | 2016 | Synthes (USA) Products LLC | Package may be incorrectly labeled as containing 0.76 mm Drill Bit, but in fact contain a 1.1 mm diameter Drill Bit. | Under Investigation by firm | DePuy Synthes sent an Urgent Notice to customers affected by the recall to inform them of the potential impacts, and actions to be taken for the return of the identified units. |  | 84 | 84 units | Distributed Nationwide | Terminated | June 01, 2017 | 212 | 84
Z-0647-2017 | 75536 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Multi-Med Single Lumen Catheters sold in Centurion convenience kits. |  |  | October 21, 2016 | 2016 | Centurion Medical Products Corporation | The kits contain Multi-Med Single Lumen Catheters thathave a potential for excess material to remain at the tip of the catheter from the manufacturing process. Itis possible that the excess material may separate from the catheter during use and pose a potential riskof embolism, which could result in serious injury or death. | Process control | Centurion sent an Urgent Recall Notice dated October 21, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of this product immediately, and forward a copy of this notice to other users to whom this product may have been further distributed. In addition, customers were instructed to complete the attached form at their earliest convenience and fax to 517-546-3356.For questions regarding this recall call 517-546-5400. |  | 1000 | 1000 kits | Nationwide Distribution | Terminated | July 20, 2017 | 272 | 1000
Z-0692-2017 | 75594 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force, System x-ray, tomography, computed |  |  | October 20, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is providing software update versionVA50A_SP3 to address the software bugs thatwere identified through normal field monitoringand the Global Complaint Handling Process.Correction for the problems are as follows:1. Correction to volumetricmisrepresentations of high contrastobjects when using ADMIRE.2. Correction to highly sporadic scanaborts due to temporarily tubecurrents at 0mA.3. Correction to missing MicrosoftHotfixes (MS16-001, MS15-135,MS15-088, MS15-048).4. Correction to highly sporadic imageartifacts caused by error signals fromhigh voltage chain at extremely hightube currents.5. Correction to highly sporadicincomplete image reconstruction dueto temporary slow hard disk writespeed.6. Correction to highly sporadic scanaborts due to temporarily high diskload of system disk.7. Correction to highly sporadic imageartifacts generated by scattercorrection in case of tube arcings. | Under Investigation by firm | Siemens sent a Customer Information Letter dated November 2, 2016, to customers affected by the recall to inform them of a release of a software update that will provide bug-fixes to improve system performance. Customers will be contacted to arrange a date for onsite installation. For further questions, please call (610) 448-6471. |  | 48 | 48 units distributed in U.S. | Nationwide Distribution. | Terminated | May 30, 2017 | 222 | 48
Z-0696-2017 | 75756 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core 128 728323 |  |  | November 12, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Software error due to the filament on timer | Software design | Philips sent a Customer Information Letter dated November 14, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to contact their local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377). |  | 509 | 509 | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY and Internationally to Algeria, Argentina, Australia, Austria, Bangladeshm Belgium, Bolivia, Brazil, Cambodia, Canada, Chile, China, Colombia, Croatia, Cuba, Czech Republic, Denmark, Djbouti, Dominican Republic, Egypt, Estonia, Finland, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of Korea, Latvia, Lebanon, Lithuania, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestinian territory, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Russian Federation, Rwanda, Saudi Aarabia, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan, and Vietnam. | Open, Classified |  |  | 1419
Z-0698-2017 | 75632 | No | PAX |  |  | Neurology | HDE - Humanitarian Device Exemption | Elana Surgical KitHUD |  |  | November 09, 2016 | 2016 | Elana, Inc. | Elana has issued this Field Action to voluntarily remove specific serial numbers of our Elana Surgical KitHUD due to potentially compromised integrity of the outer sterile packaging of the Elana Catheter 2.0. | Packaging | Elana sent an Urgent Medical Device Recall letter dated November 7, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to not use the affected device and to return the product for replacement. Customers were also instructed to complete the attached Acknowledgement and Receipt Form to facilitate the product return and exchange. Customers with questions were instructed to call 301-215-7010. |  | 4 | 4 | Worldwide Distribution - US including NY and Internationally to Switzerland and the Netherlands. | Open, Classified |  |  | 4
Z-0702-2017 | 75630 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | MANOS EX (CAT#TY-12-101) |  |  | November 04, 2016 | 2016 | Thayer Intellectual Property, Inc. | The firm failed to notify current customers of IFU update/warning. During a recent FDA inspection, it was discovered that Thayer Intellectual Property, Inc. had revised their IFU to include a warning not to bend, deform or otherwise modify any part of the device. The warning was added as a result of a MAUDE report filed in December 2014 where a piece of the MANOS EX broke during use and was left in the patient's hand. The patient suffered a tendon injury and an additional surgery was required to remove the broken piece of the device and repair the tendon. FDA inspection found that while the firm had updated their IFU, they had not notified customers who had been shipped the device with the older version that did not include the warning. Risk: It is important for MANOS EX users to be aware that bending, deforming, or modifying the device could result in device breakage which could impact patient safety and device performance. | Packaging change control | Thayer sent an Urgent Medical Device Correction letters dated November 4, 2016 by Federal Express. The letter identified the affected product, problem and actions to be taken. Customers were asked to check if any of the affected lots were still available, and if so, to change out the IFU with version 5. For questions call 415-923-9375. |  | 339 | 339 devices | US Nationwide Distribution in the states of CA, TX. MD, MI,VA, DC, UT, CT, ID, SD, IN, and GA. | Open, Classified |  |  | 339
Z-0699-2017 | 75444 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Aequalis Fx2 implant parts and instrument trays. |  |  | September 19, 2016 | 2016 | Tornier, Inc | Tornier is conducting a recall on Aequalis Fx2 ( implant parts and instrument trays) due to reports of dislocations of the poly insert and the stem. | Component design/selection | Wright sent an "Urgent Medical Device Recall" letter dated September 19, 2016. The letter was addressed to Healthcare Professional. The letter identified the affected product, problem and actions to be taken. It advised consignees to stop further distribution or use of product immediately. Requires consignees to notify all accounts that may be end-users, or contact Wright with your request for us to contact them. Advised consignee to complete and e-mail the attached questionnaire to FieldAction@wright.com. For questions contact the same. |  | 30 | 30 | US Distribution TX, CA, and MI | Open, Classified |  |  | 30
Z-0704-2017 | 75796 | No | EBC | Pit and fissure sealant and conditioner. | Dental | Dental | 510(k) | UltraSeal XT Plus Clear Refill. The package contains 4x 1.2 ml UltraSeal XT Plus. |  |  | November 17, 2016 | 2016 | Ultradent Products, Inc. | There was a labeling error with product, UltraSeal XT Plus Clear Refill, Part Number (PN) 565, Lot Number (Lot) BD79T. The package was mislabeled as PQ1, PN 641; while the barcode, IFU and products were labeled for UltraSeal XT Plus Clear Refill, PN 565, Lot BD79T. | Labeling mix-ups | Consignees notified via phone on 17Nov2016 and by letter on 22Nov2016 to identify and return products. Consignees with any questions regarding this field action can be sent to 801-553-4858 or 801-553-4220. |  | 31 | 31 | Worldwide Distribution - US Nationwide | Terminated | June 05, 2017 | 200 | 31
Z-0701-2017 | 75644 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 3M Universal Electrosurgical Pad, REF 9130 (100 pads / case) and |  |  | October 26, 2016 | 2016 | 3M Company - Health Care Business | 3M is recalling Universal Electrosurgical Pads because a report of an unintended material (process liner) in the product. This may prevent or impede the safe return of electrosurgical current following Electrosurgical Unit (ESU) activation. This defect has the potential to increase the risk of a patient burn directly under the pad or at an alternate site of the body. | Equipment maintenance | 3M sent an Urgent Medical Device Recall letter dated October 26, 2016 to affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to complete and return the attached Customer Product Return and Recall Form. Customers with questions can contact 3M Health Care Customer Helpline at 1-800-228-3957, 8:30 am - 7:00 pm EST. |  | 32000 | 32,000 pads | Worldwide Distribution - US Nationwide in the states of AR, AZ, CA, FL, GA, ID, IL, KY, NC, NY, TN, TX, VA and the countries of Canada, Chile, Costa Rica, Germany, Japan, Mexico and South Africa. | Terminated | August 01, 2017 | 279 | 32000
Z-0703-2017 | 75608 | No | NJL |  |  | Orthopedic | PMA | LCS Complete Knee Revision System VVC Insert |  |  | November 08, 2016 | 2016 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is voluntarily recalling one lot of the LCS Complete Knee Revision System VVC Insert Small 22.5mm, due to the insert being manufactured outside of specifications, which prevents the insert from properly seating in the mating tibial tray. | Nonconforming Material/Component | DePuy Orthopaedics, Inc. initiated a voluntary recall on 11/8/2016 via letter mailed to the customer. Only one device was distributed in the U.S., and this was removed from the hospital on 21-Oct-2016. This recall does not affect any other lots of the LCS Complete Knee Revision System VVC Insert Small 22.5mm devices. The recall notice provides instructions for notifying the medical facility that received, used, or purchased the affected lot of the LCS Complete Knee Revision System VVC Insert Small 22.5mm. The purpose of this device recall is to notify medical professionals of the possible effects of using the affected device and to communicate that there are no recalled devices distributed at this time. The customer was asked to do the following: Complete the Reconciliation Form and return to your DePuy Orthopaedics Sales Consultant or fax to 574-371-4939 or email to DPYUS-JointReconFieldActions@its.jnj.com within five (5) business-days of this notice." Retain a copy of the completed Reconciliation Form and this notice in your files." Forward this notice to others in your facility that need to be informed.Effectiveness will be determined by the return of one US Medical Facility Reconciliation Form.For questions about device recall information provided, please contact Kim Earle, Senior Recall Coordinator, at 574-371-4917 or kearle@its.jnj.com or DPYUS-JointReconFieldActions@its.jnj.com. |  | 1 | 1 | US Distribution in the state of TN | Terminated | June 05, 2017 | 209 | 1
Z-0700-2017 | 75458 | No | MHY |  |  | Neurology | PMA | Model 37751 Recharger |  |  | October 03, 2016 | 2016 | Medtronic Neuromodulation | Medtronic has identified an increased number of complaints from customers involving reports of Rechargers (Medtronic Model 37751 Recharger, which is included in Models 37754 and 97754 Charging Systems,Spinal Cord Stimulation)that are in an unresponsive error state, where the Recharger is non-functional with a blank display screen and is beeping every 5 seconds. Medtronic has determined all Rechargers manufactured starting in November 2014 (indicated by serial numbers beginning with NKA4 or NKU4 ) are more susceptible to this error state. | Software change control | Consignees were sent a series of documents which included an introductory Medtronic and "Urgent Medical Device Correction" letters dated September 2016. The urgent letters included one for Deep Brain Stimulation (DBS) and one for Spinal Cord Stimulation(SCS) . Each letter was addressed to Healthcare Professional. Only physicians associated with both SCS and DBS therapy received the cover letter. The letters provided information on "Background, Issues Mitigation, Recommendation, Additional Information and Patient Notification (DBS).Requested consignees to complete and return the Physician /HCP Reply Form.For questions contact Medtronic Technical Services at 1-800-707-0933 weekdays 7 am to 6pm, CT. |  | 42887 | 42,887 | Worldwide Distribution - US Nationwide including DCOUS: Canada and others. | Open, Classified |  |  | 42887
Z-0708-2017 | 75322 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Xcela Power Injectable PICC Maximal Barrier Nursing Kit |  |  | August 29, 2016 | 2016 | Angiodynamics Inc. (Navilyst Medical Inc.) | Medline Industries, Inc. supplies Navilyst with sterile package gloves included in the tandem packing of the Maximal Barrier Nursing Kits. Medline has determined that specific lots of the gloves supplied to Navilyst contain an incorrect expiration date. Navilyst has confirmed that Medlines affected sterile gloves (Medline Part Number MSG1075/MSG3075) have been included in a limited number of kits. | Under Investigation by firm | Navilyst sent an Urgent Voluntary Medical Device Recall notices, reply forms and a copy of Medline's Recall Notice to the their customers. The letter identified the product, the problem, and the action to be taken by the customer. Navilyst Medical, Inc. (an AngioDynamics Company), the manufacturer of the Xcela" PICC w/PASV, the BioFlo PICC w/ Endexo, and the BioFlo PICC w/ Endexo and PASV is conducting a medical device recall based on information received from Medline. Medline has determined that the sterile gloves listed in their Recall Notification dated August 3, 2016 may contain an incorrect expiration date. Actions to be taken: - Customers should immediately remove the affected product from inventory (whether in labs, Central Supply, Shipping and Receiving or ANY other location). - Segregate this product in a secure location for return to Navilyst Medical. - Immediately forward a copy of this recall notification to all sites to which customers may have further distributed affected product. Customers should also promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return) via email to recall@angiodynamics.com. Customers with questions can call 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday  Friday: Eastern Standard Time).For questions regarding this recall 518-792-4112. |  | 75 | 75 packages (boxes of 5) | Nationwide Distribution including AZ, CA, FL, GA, HI, IL&lt; IN&lt; KS, LA, MI, MN, MO, ND, NM, OH, OR, PA, TN, TX, WA &amp; the District of Columbia. | Open, Classified |  |  | 428
Z-0707-2017 | 75270 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge, eFilm Workstation and eFilm Lite |  |  | October 20, 2016 | 2016 | Merge Healthcare, Inc. | A product issues happens for RF projection images. If user measures on RF image, .cal (Calibration) is not displayed and measurement is not correct. The .cal label is not displayed on the projection images after calibration is performed which should be present if the measurement is presented based on Pixel Spacing. | Software design | Merge Healthcare sent an Urgent Medical Device Recall letter dated October 12, 2016, to all affected consignees. The letter described the Product, Issue, Potential Harm, Containment by the Customer / User, Actions by Merge, and Actions by Customers. Advised consignees to discontinue using the product, and identify &amp; notify all customers. Requested consignees for a response to the notification. Customers with questions were instructed to send an email to recall@merge.com. |  | 2163 | 2163 | Worldwide Distribution - US: Nationwide including DC, GU, PR and Internationally to Canada, Argentina, Armenia, Australia, Austria, Belgium, Brazil, Chile, China, Colombia, Congo, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Guatemala, hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Mauritania, Mexico, Netherlands, New Caledonia, New Zealand, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Vietnam, and Yemen. | Open, Classified |  |  | 2163
Z-0712-2017 | 75665 | orthosis, limb brace | IQI | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | TENNIS ELBOW SUPPORT, PneuGel(TM), UNIVERSAL, REF 200A10, QTY: 1 Each, Rx Only, |  |  | November 07, 2016 | 2016 | DeRoyal Industries Inc | DeRoyal received complaints of mold on Tennis Elbow Straps. | Under Investigation by firm | DeRoyal issued their recall on 11/7/16 to the physician / clinician level by sending recall letters to its direct accounts and sales representatives. The foreign customer was notified by email on 11/7/2016. The customers were asked to identify the affected products, destroy them and return the provided notice of destruction form to DeRoyal for credit. |  | 56 | 56 units | US Nationwide Distribution | Open, Classified |  |  | 847
Z-0711-2017 | 75715 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Empowr Knee System |  |  | November 15, 2016 | 2016 | Encore Medical, Lp | A size 4 Empowr baseplate was mislabeled as a size 6. The affected lot was laser marked as M6, marked with the lot number for the size 6 minus lot, and distributed as size 6 minus. | Mixed-up of materials/components | The recalling firm sent an Urgent Field Safety Notice letter dated November 15, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions call 760-734-3551. |  | 7 | 7 units | US Nationwide Distribution in the states of CO, IL, IN, FL, CA, AZ | Open, Classified |  |  | 7
Z-0716-2017 | 75570 | No | LWR |  |  | Cardiovascular | PMA | Perceval Sutureless Aortic Heart Valve size S, Item # ICV1208, REF PVS21; Perceval Sutureless Aortic Heart Valve size M, Item # ICV1209, REF PVS23; Perceval Sutureless Aortic Heart Valve size L, Item # ICV1210, REF PVS25; Perceval Sutureless Aortic Heart Valve size XL, Item # ICV1211, REF PVS27. |  |  | October 31, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field action for the Perceval Sutureless Heart Valve due to updated steps associated with the implantation procedure. | Use error | The firm, LivaNova, sent a "Customer Letter" dated 10/31/2016 to its Consignees with clarifications about the implant of the Perceval Sutureless Aortic Valve. The letter described the product, problem and actions to be taken. The consignees were instructed to reinforce the steps that may influence procedural success and mitigate potential complications and complete and return the Customer Response Form via fax to: LivaNova Canada Corp, Canada at +1 303 895 2335 or email to: HV.FEEDBACK@LIVANOVA.COM.If you have additional questions or request of clarifications, please contact your LivaNova representative or Customer Service.For questions regarding this customer letter, please contact your LivaNova sales representative or email: hv.feedback@livanova.com. |  | 845 | 845 units | Worldwide Distribution-US (Nationwide) including states of: MD, ME, MA, NH, NY, NJ, AL, MS, VA, FL, CA, HI, TN, GA, WI, TX, MI, OH, KS, MN, NC, IL, IN, WA MO, PA, NM, and UT and countries of: Austria, Australia, Belgium, Bulgaria, Brazil, Canada, Chile, Colombia, Croatia, Czech Republic, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Jordan, India, Italy, Iran, Islamic Republic of, Ireland, Israel, Korea, Republic of, Kuwait, Luxembourg, Malaysia, Malta, Mexico, Netherland, New Zealand, Norway, Panama, Poland, Portugal, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South African, Spain, Switzerland, Romania, Russian Federation,Taiwan, Province of China, Thailand, Turkey, United Arab Emirates, United Kingdom, and Uruguay. | Open, Classified |  |  | 845
Z-0715-2017 | 74948 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Peripherally Inserted Central Catheters (PICC), UPN H965750191, Catalog # 75-019 |  |  | June 08, 2016 | 2016 | Angiodynamics Inc. (Navilyst Medical Inc.) | There is a potential for stylet stiffening that is not accurately represented on the labeling. The label states that product includes a stylet with an outside diameter (OD) measuring 0.014, however, the product may contain a stylet with an outside diameter (OD) measuring 0.016. | Mixed-up of materials/components | Angiodynamics, Inc. (Navilyst Medical, Inc.) sent an Urgent Voluntary Medical Device Recall letter dated June 8, 2016, with reply forms to all affected customers. Customers were advised to immediately remove any affected product from inventory, segregate and secure in a location for return. Immediately forward a copy of the recall notification to all sites to which they have distributed affected product. If affected product is located in your institution, please call Navilyst Medical Customer Service at 1-800-772-6446 8am-7pm, EST Monday - Friday to obtain a replacement or credit for your returned product. Promptly complete, sign and return the Reply Verification Tracking Form by email to recall@angiodynamics.com or by Fax to 1-800-782-1357.For questions regarding this recall call 518-792-4112. |  | 50 | 50 cartons (boxes of 5) | Nationwide Distribution to CA, FL, GA, HI, LA, MI, MN, MO, NV, OH &amp; TX | Open, Classified |  |  | 98
Z-0719-2017 | 75620 | No | ELW | Impression material. | Dental | Dental | 510(k) | Virtual XD Test Pack Heavy/Light Fast Set, 2 x 50 ml, REF #/Product Code 646469, Rx ONLY -- |  |  | November 09, 2016 | 2016 | Ivoclar Vivadent, Inc. | The firm received complaints claiming the dental material failed to set up. As the dental material ages, the set time may increase. | Under Investigation by firm | Ivoclar Vivadent sent and URGENT - MEDICAL DEVICE RECALL Letters (dated 11/07/2016) and Recall Response Forms to customers via Certified Mail-Return Receipt Requested. The letter identified the affected product, problem and actions to be taken. Customers were advised to return all affected products in stock. For questions contact Ivoclar Vivadent Customer Service at 800-533-6825. |  | 2886 | US: 2090 units, Canada: 331 units, Australia: 465 units | Worldwide Distribution - US Nationwide and the countries of Canada and Australia | Open, Classified |  |  | 10606
Z-0680-2017 | 75575 | No | RBY |  |  |  |  | Olympus Scientific Solutions Americas Corporation ( OSSA ) Model: VANTA Analytical X-Ray System |  |  | October 27, 2016 | 2016 | Olympus Scientific Solutions Americas | Possible unintentional x-ray emission after users attempt early termination of the sequence. | Radiation Control for Health and Safety Act | The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 41 units located in the United States will consist of the following:1. Shipments were stopped until a software correction was designed and validated. The software correction was designed by October 27, 2016, and was implemented on November 7, 2016 for all units that had not left the place of manufacture and for all new products.2. The software correction will be made available to all distributors, sales personnel and users. Issue 1 is corrected by altering the behavior of the stop button and trigger such that they stop the current test and terminate the test sequence. To correct issue 2 the software will be modified to grey-out the Stop button during this very small period making it consistent with rest of the VANTA Start/Stop test UI paradigm. Both during user Start and Stop actions, users are shown a greyed button to acknowledge user Click and the button is un-greyed when the system completes the user operation. Both user operations typically take 100  400 milliseconds to complete.3. All VANTA units in the United States are currently in the possession of Olympus or your distributors. No units are in the possession of users.4. Olympus or its distributors will apply the software correction prior to delivering a VANTA system to any purchasers in the United States.For further questions, please call (781) 419-3900. |  | 41 | 41 units | US Distribution to the states of : TX, WA, GA, MA, SC, AZ, PA, VA, CA, OH, NJ, MD, CO, IL, LA, and MN. | Open, Classified |  |  | 41
Z-0720-2017 | 75633 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System |  |  | October 11, 2016 | 2016 | RAYSEARCH LABORATORIES AB | An error may occur with the display of dose computed on images other than the planning CT (auxiliary CT) when using multiple patient cases in RayStation 5. If a CT image set with the same Frame of Reference as the displayed auxiliary CT set exists in another case, the dose display may be incorrect. When the error occurs, the display of dose in patient views, including the maximum dose position, the dynamic isodose lines and the dose grid may be incorrect. The dose value normally displayed in the upper left corner when pointing in the 2D view may be incorrect or missing. | Software design | On October 11, 2016 RaySearch Laboratories distributed a Field Safety Notice, Medical Device Correction #13529 to customers via email. Customers were instructed to take caution when a patient has multiple cases. The following may be incorrect if there are image sets that shre the same Frame of Reference residing in different cases: *Display of evaluation dose computed on another CT image set and *Dose tracking fraction dose and deformed dose. Until a corrected version has been installed, all affected users must maintain awareness of this field safety notice. In these cases, the display of the dose values in the 2D view, including the dose value shown when pointing in the view, maximum dose position, dynamic isodose lines and dose grid display may be incorrect. Customers are encouraged to educate staff and all users about this. For regulatory information, please contact David Hedfors, at +46 8 510 530 12 or david.hedfors@raysearchlabs.com. |  | 540 | 540 units | US Nationwide Distribution in the states of AZ, AR, CA, CT, DE, FL, GA, HI, IL, LA, MA, MD, MN, MI, MS, MO, NJ, NY, NC, OH, PA, RI, TX &amp; WA | Open, Classified |  |  | 540
Z-0727-2017 | 74368 | No | NAC | Hydrophilic wound dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | AMD-Ritmed Transparent Dressing - Window with label, 4 " x 4 3/4" (10 cm x 12 cm), Reorder No./Code A60044-W, Sterile |  |  | June 08, 2016 | 2016 | Amd-Ritmed, Inc. | The lot of transparent dressing was found to be contaminated (not sterile) based on FDA sampling and analysis. | Material/Component Contamination | URGENT: MEDICAL DEVICE RECALL Letter (dated 6/8/2016) was e-mailed to the consignees on 6/8/2016. The recall is being conducted to the retail level. For questions. please contact AMD Customer Service by phone at 1-800-445-0340 or by e-mail at contactus@amdritmed.com. |  | 216 | US: 216 cases (200 dressings per case) | NY, MA, VA, TN | Open, Classified |  |  | 216
Z-0726-2017 | 75713 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | Cut lines on the image may present horizontally rather than vertically. | Component design/selection | MERGE sent an Urgent Medical Device Recall dated March 28, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Merge Healthcare is working on correcting the issues and will inform customers once a fix is made available. Customers were asked to reply using the enclosed form and the return addressed envelope. If customers have further distributed the affected product they need to identify their customers and notify them at once of the product recall. Customers with questions were instructed to send an email to recall@merge.com. |  | 537 | 537 sites potentially have the affected versions | Worldwide Distribution - US (nationwide) Canada and foreign distribution | Open, Classified |  |  | 537
Z-0724-2017 | 75683 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Halyard C-Section Drape with Clear Screen and Full Incise (Non-Sterile), Product Code 44965 NS; C-Section Drape with Clear Screen and Full Incise (Sterile), Product Code 44966 00; C-Section Drape with Clear Screen (Non-Sterile), Product Code 44967 NS; C-Section Drape with Clear Screen (Sterile), Product Code 44968 00; C-Section Drape with Pouch and Fenestration (Non-Sterile), Product Code 44977 NS; C-Section Drape with Pouch and Fenestration (Sterile), Product Code 44978 00 |  |  | November 09, 2016 | 2016 | Halyard Health, Inc | Drapes within the affected lots may include a manufacturing variation that prevents convenient removal of the coated release paper, which inhibits easy access to the surgical site. | Nonconforming Material/Component | Accounts were notified of the recall by letter on 11/9/2016. The consignees were instructed to quarantine and discontinue use of the recalled lots and to discard/destroy per the facility's standard disposal procedure for nonwoven materials. A response form was included which was to be sent back. |  | 132 | 132 cases | US Nationwide Distribution in the states of AL, FL, GA, IL, KS, MN, MO, NE, NJ, OH, PH, TN, VA, WI, WY | Open, Classified |  |  | 132
Z-0722-2017 | 75681 | No | DIS | Barbiturate test system. | Toxicology | Toxicology | 510(k) | Barbiturates Intercept Micro-Plate EIA (enzyme immunoassay, barbiturate) |  |  | November 10, 2016 | 2016 | OraSure Technologies, Inc. | Shipping error. Cocaine metabolite Microplate distributed instead of Barbiturate Microplate due to incorrect selection of Microplate during manufacturing. | Mixed-up of materials/components | Due to the small number of customers that were affected by the recall, OraSure notified customers on November 10, 2016, via telephone with a follow-up email. Customers were asked to quarantine the units on hand and they will schedule a pickup by Fed EX to retrieve the affected units.For further questions, please call (610) 882-1820, ext. 1657 |  | 4 | 4 kits | Distributed to KS, and VA | Terminated | April 17, 2017 | 158 | 4
Z-0723-2017 | 75652 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | ADVIA 560 Hematology Systems, Siemens Material Number (SMN) 11170842, IVD. |  |  | October 21, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens is investigating an issue which may cause an incorrect result to be reported. Siemens received two reports of multiple discordant records for the same Sample ID in the ADVIA 560 Hematology System database that occurred during the installation of the systems. The database should only contain one record of a Sample ID number for any given time and date. If there are multiple records for the same Sample ID, it is possible that multiple results may be manually or automatically sent to the Laboratory Information System (LIS), printed or displayed on the results report screen. | Software design | An Urgent Medical Device Correction Letter HI17-02.A.US (dated October 21, 2016) and Response Form were sent to US Customers via Federal Express. An Urgent Field Safety Notice was sent to customers outside the US on October 21, 2016, (distribution is determined at the country level). The letters instruct customers to contact their local Siemens technical support representative if they obtain multiple results for one Sample ID. It also informs the customer what to do if the ADVIA 560 Hematology System generates any of the six error messages listed in the letter. For further questions, please call (312) 275-7795. |  | 164 | US: 23 systems; Foreign: 141 systems | Worldwide Distribution - US Distribution to the states of including Puerto Rico and : CA, CO, FL, IL, MI, NJ, NY, OH, TX, WA, WI, WI and WY., and to the countries of : Angola, Australia, Belgium, Brazil, Canada, Chile, Colombia, Croatia, Czech Republic, Egypt, France, Germany, Greece, India, Italy, Japan, Kenya, Malaysia, Mexico, Myanmar, Netherlands, Norway, Pakistan, Portugal, Saudi Arabia, Slovenia, Spain, Switzerland, Turkey, Uganda, United Kingdom, Vietnam. | Open, Classified |  |  | 164
Z-0721-2017 | 75731 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation 4.0, 4.5, 4.7 and 5.0; Radiation Therapy Treatment Planning System |  |  | November 09, 2016 | 2016 | RAYSEARCH LABORATORIES AB | Software anomaly; an issue was found with the proton Pencil Beam Scanning (PBS) dose calculation in RayStation 4.0, 4.5, 4.7 and 5.0. For treatment plans with a combination of range shifter, large air gap and beams that enter the patient surface at an oblique angle, the dose calculation accuracy may be less than expected. | Software design | Ray Search sent an Urgent Field Safety Notice Medical Device Correction letter (dated 11/8/2016) via e-mailed to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to Inspect their product and identify all installed units with the above software version number, then confirm you have read and understood this notice (contact information included). |  | 66 | US: 66 units | US Nationwide Distribution in the states of California, Florida, Illinois, Louisiana, Michigan, New Jersey, Tennessee, Texas, Washington | Open, Classified |  |  | 66
Z-0725-2017 | 75263 | No | JGJ | Magnesium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Magnesium on RX Imola analyser |  |  | April 14, 2016 | 2016 | Randox Laboratories, Limited | According to the firm, Carry over was observed when the amylase or pancreatic amylase test is run directly before or after the Magnesium assay. An update is required to the RX Imola analyser running order and an acid wash recommended when setting up the magnesium assay. A correction was made to the IFU for the Magnesium assay and all RX Imola customers in the USA were contacted with the updated instruction. | Device Design | Randox sent an Urgent Field Safety Notice to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review and update Rx Imola running order where required. Rework any remaining stock with updated IFU and Important Notice. Complete and return the vigilance response section of the form to technicalservikces@randox.com. |  | 15 | 15 | Nationwide Distribution The product was only distributed to US Consignees. | Open, Classified |  |  | 15
Z-0659-2017 | 75600 | No | LOZ |  |  | Cardiovascular | PMA | 5000 Series Freedom Drivers. Freedom Driver System for Temporary Total Artificial Heart (TAH-t). Part number 595000-001 |  |  | October 21, 2016 | 2016 | SynCardia Systems Inc. | The Main Printed Circuit Board Assembly (PCBA) of the affected Freedom Drivers may fail and cause the Freedom Driver to stop functioning without visual or audible alarms, resulting in the loss of life-sustaining function. | Process control | On 10/21/16 customers were notified by phone. Customers were informed that patients who were being supported by affected Drivers must be switched to unaffected Drivers.The firm will provide unaffected Freedom Drivers to replace the affected Drivers. Affected Drivers must be return to the firm. Instructions for return provided, and Acknowledgment Form must be completed.Contact Information For QuestionsMary Pat SloanSenior Vice PresidentGlobal Certification and LogisticsSynCardia Systems, LLC1992 E Silverlake RdTucson, AZ 85713Direct 520-547-7470Toll-free 866-480-1122 Ext 1308Monday-Friday 7:30 AM to 5:00 PM Pacific/Arizona timeHotline 24/7 866-771-9437www.syncardia.comOn 10/27/16 customer notification letters were sent. The letter provided instructions on how to recognize the device could fail. The patient and caregiver will recognize that an affected Driver is failing because the Freedom Driver will stop pumping without generating any sound or any visual or audible alarms.When functioning, the Freedom Driver exhibits a distinctive pumping sound; when it stops pumping the lack of sound is noticeable.Generally, if the Freedom Driver function is interrupted, a visual and/or audible alarm will occur. However, as a result of this component failure, when affected Freedom Drivers stop functioning, there will be no visual or audible alarms, and no pumping sound.Actions to Be Taken by The Customer/UserSynCardia has provided unaffected Freedom Drivers to the two hospitals to replace the affected 5000 series Drivers. All patients with affected Freedom Drivers have undergone the exchange and received unaffected Drivers.The hospitals have returned all the affected Freedom Drivers to SynCardia.The hospitals are requested to complete the Recall Acknowledgement and Receipt Form below and return to SynCardia. |  | 5 | 5 | US distribution to Virginia and Arizona. | Open, Classified |  |  | 5
Z-0731-2017 | 75853 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | December 02, 2016 | 2016 | Elekta, Inc. | Edits to Particle field definition parameters may not be saved when the field definition window is saved. | Software design | The firm, Elekta, sent an "URGENT IMPORTANT FIELD SAFETY NOTIFICATION" letter dated 11/2016 was sent to affected customers on 12/2/2016. The letter describes the product, problem and actions to be taken. The customers were instructed to save and re-open the field definition form when manual edits have been performed to confirm that the edits were save; post this notice in a place accessible to all users; advise the appropriate personnel working with the product the content of the letter; and complete and return the Acknowledgement Form via fax to: 408 830 8003 immediately upon receipt but no later than 30 days.Elekta will resolve the problem with the MOSAIQ 2.60 SP7F, MOSAIQ 2.60 SP14 and 2.64 SP3. Customers are encouraged to contact Elekta support to get the corrected version.If you have any questions, contact RA Manager Americas at 1-770-300-9725 or email to: Elekta_Global_PMS@elekta.com. |  | 47 | 47 | Worldwide Distribution-US (nationwide) to states of: FL, IL, MO, NJ, OH, OK, TN, TX, VA, and countries of: Italy, Japan, South Korea, Sweden and United Kingdom. | Open, Classified |  |  | 47
Z-0729-2017 | 75825 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | If images are sent without an order in the system, they will be matched with the latest order on the current patient/modality matching potentially resulting in the matching of the report to the incorrect accession number (but still associated to the correct patient). | Software design | MERE sent an Urgent Medical Device Recall letter dated July 11, 2016, to all affected customers via email or via certified mail if they did not have an email address for the customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to respond to the notification using the enclosed form and the return addressed envelope. Customers were also instructed to ensure that all users of the product are provided with the notification and if product was further distributed customers need to identify their customers and notify them at once of the product recall. Customers with questions should send an email to recall@merge.com. For questions regarding this recall call 877-741-5369. |  | 198 | 198 sites potentially have the effected versions | Nationwide Distribution | Open, Classified |  |  | 198
Z-0730-2017 | 75714 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | October 01, 2015 | 2015 | Merge Healthcare, Inc. | Reporting feature times out after inactivity for more than an hour sending the user back to the study list, which causes all reporting data being entered to be lost. | Software design | MERGE sent an Urgent Medical Device Recall letter dated March 28, 2016, to all affected customers. Non-responders were issued a second letter dated July 11, 2016, The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to reply using the enclosed form and the return addressed envelope and also to ensure that all users of the product were provided with the notification. If product was further distributed, customer were asked to identify their customers and notify them at once of the product recall. Customers with questions were advised to email recall@merge.com.For questions regarding this recall call 877-741-5369. |  | 116 | 116 sites potentially have the effected versions | Nationwide Distribution | Open, Classified |  |  | 116
Z-0728-2017 | 74488 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Carry bar which is an accessory to a lift used for patients in hospitals and other facilities. Recall impacts the following model numbers: 360741, 360750, 360751, 360755, 360755 Rev B, 360756, and 360757. |  |  | August 30, 2016 | 2016 | ErgoSafe Products, LLC (DBA) Prism Medical | The black plastic puck on the carry bar is breaking resulting in potential for patient harm. | Device Design | The recalling firm sent field safety notices to their customers by mail on August 30, 2016. The notice instructed customers to inspect their carry bar to see if it used a black plastic puck. If it did, they were instructed to contact the recalling firm to receive a metal insert to install. If the puck is cracked or damaged, they can request a new, replacement carry bar at a discounted price. For further questions, please call (866) 891-6502. |  | 17964 | not greater than 17,964 units | Distributed US nationwide and to Canada, Japan, Australia, Thailand, Great Britain, and South Korea | Open, Classified |  |  | 17964
Z-0746-2017 | 75385 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Harmony Insight Diagnostic Monitor Support System - 100 Spring Arm (EMS) |  |  | October 05, 2016 | 2016 | Steris Corporation | The Drager Pendula spring arm device has a set of screws located on the spring arm that were not properly secured with Loctite and may become loose. | Component design/selection | Steris sent an Urgent Voluntary Field Correction Notice to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remind staff personnel that Harmony Insight Diagnostic Monitor Support Systems, including the affected Drager Pendula spring arms, are not designed or intended to be placed over hospital personnel or patients and should be moved with care. Customers with questions were instructed to call 440-392-7358 or Steris Customer Service at 1-800-548-4873. For questions regarding this recall call 440-392-7231. |  | 63 | 63 | Nationwide Distribution including AK, AZ, CA, CO, FL, IL, KY, MI, MO, NB, NJ, NY, OH, PA, RI, SD, TN, TX, UT, WA &amp; WV. | Terminated | April 14, 2017 | 191 | 63
Z-0732-2017 | 75460 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Lumex, Patriot Full Electric Homecare Bed, &amp; Patriot LX Full Electric Homecare Bed containing the Okin Dumat 3 Head/Foot Motors. |  |  | September 30, 2016 | 2016 | GF Health Products, Inc | GF Health Products, Inc., in conjunction with OKIN America, has initiated a voluntary field correction affecting the OKIN America "DUOMAT 3" Head/Foot motor used in the manufacture of our Lumex Patriot Full Electric Bed Model US0458/US0468 and Patriot LX Full Electric Bed Model US6000.The failure mode is directly related to contaminated material in the plastic housing used to produce the DUOMAT 3 motor. Contaminated motor housings may crack resulting in a motor that will no longer actuate the head section per the intended design.If the head section were to fall it is unlikely but possible an injury could occur. No injuries have been reported. | Process control | Graham Field sent an "Urgent Correction" letter dated September 30, 2016. The letter described the problem and the product involved in the recall. The letter described the " Action Required: Motor Replacement" . Informed consignees to contact GF Health Products at 1-888-974-4347 or fieldcorrection@grahamfield.com within ten (10) business days from receipt of this notification to schedule motor replacement order. Requested consignees to return the completed Certification of Replacement and Return Agreement to: fieldcorrection@grahamfield.com or Fax: 1-920-929-8213. For additional information, please call 1-888-974-4347; or at fieldcorrection@grahamfield.com. |  | 3547 | 3,547 (3063 beds, 484 motors) | Worldwide Distribution - USA ( nationwide) Distribution including Puerto Rico and to the countries of : Bahamas, Canada, Chile, El Salvador and Santo Domingo. | Open, Classified |  |  | 3547
Z-0747-2017 | 75508 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Perkin Elmer RoboRack 25ul, Clear, Non-Conductive, Filter Tips, Pre-sterilized |  |  | June 24, 2016 | 2016 | PerkinElmer, Inc | Filter Tips molded incorrectly and may not seal to the Varispan tip adaptor properly causing inaccurate aspiration and dispensing | Process control | PerkinElmer initiated recall on June 24, 2016 to the Product Manager and was delivered to the customer by email facilitated through PerkinElmer Sales Team.Customers were requested to take the following actions until PerkinElmer's corrective measures are completed:Inspect customer inventory of the 6000689 25 Clear Filter Pipette Tip and identify ifthe affected lots are present. Complete and return the attached Response Form with the required information. Destroy all quantities of tips from the affected lots.Upon receipt of the return Response Form, a shipment of replacement tips will be shipped. |  | 595 | 595 trays (571,200 tips) | Worldwide Distribution - US including CA, MA, NY, IA, OR , PA, WA and Internationally to Belgium, Canada, Finland, France, Germany, Japan, Netherlands, Singapore, and UK | Open, Classified |  |  | 595
Z-0741-2017 | 75638 | No | LYZ | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | McKesson Powder Free Vinyl Exam Gloves Catalog numbers 14-116, 14-118, 14-120, 14-136, 14-138 and 14-140 |  |  | October 12, 2016 | 2016 | CYPRESS MEDICAL PRODUCTS | A container of McKesson vinyl gloves that failed the FDA leak test was inadvertently distributed. | Process control | A letter, dated October 19, 2016, on Cypress Medical Products letterhead, was sent to customers. Customers were requested to stop use and quarantine the gloves. |  | 2005.2 | 2005.2 cases have been distributed | US Distribution. This recall is limited to one distribution center in Auburn, Washington, and the customers serviced from that distribution center. | Terminated | March 27, 2017 | 166 | 2005.2
Z-0733-2017 | 75793 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Cook Spectrum Minocycline/Rifampin Impregnated Double Lumen Central Venous Catheter Tray |  |  | November 28, 2016 | 2016 | Cook Inc. | Cook Medical has received reports of the trays containing the incorrect needle, which in some cases could result in the inability to pass a wire guide through the needle.Potential adverse events that may occur as a result of utilizing needles of incorrect size include delay in treatment, which could be critical in an emergent situation. There have been no reports of illness or injury associated with this issue. | Mixed-up of materials/components | Cook Medical sent an Urgent - Medical Device Recall letter dated November 23, 2016. to all affected customers. The letter identified the product the problem and the action needed to be taken. Action to be taken: 1. Please review the attached list of affected products and lot numbers that were shipped to your account and quarantine any affected products that remain unused. 2. Immediately collect and return all unused affected products to Cook Medical as soon as possible for credit. 3. Please complete the attached Acknowledgement and Receipt Form and return it to Cook Medical. 4. Please report any adverse event to Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235, Monday through Friday between 7:30 am and 5:00 pm Eastern Standard Time or e-mail at fieldactionsna@cookmedical.com. Transmission of this notice: This notice must be passed on to appropriate personnel, including down to the user level, within your organization or to any organization where the potentially affected devices have been transferred. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235. |  | 99 | 99 units | AZCADCGAILLAMNMSNCNVNYOKPATNTXWI | Open, Classified |  |  | 99
Z-0740-2017 | 75802 | No | OFT | Anesthesia conduction catheter. | Anesthesiology | Anesthesiology | Enforcement Discretion | BD Spinal Anesthesia Tray |  |  | November 07, 2016 | 2016 | Becton Dickinson &amp; Company | BD is conducting a voluntary product recall of the BD Spinal Anesthesia Tray, Cat (Ref) 405671, lotB01K089D since a very low number of trays are incorrectly labeled as BD Single Shot Epidural Tray,Cat (Ref) 406069. | Error in labeling | BD sent an Urgent Product Recall letter dated November 7, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately review their inventory for the affected product. Return the affected products with the Business Response Card form following the instruction on the enclosed packing instruction. The completed form should be faxed to 1-866-240-5212 or email the completed form to bd6549@stericycle.com. Customers with questions were instructed to call 1-866-792-5451. For questions regarding this recall call 866-792-5451. |  | 4500 | 4500 | Nationwide Distribution | Open, Classified |  |  | 4500
Z-0739-2017 | 75820 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Performance-LOAD Ambulance Cot Fastener |  |  | November 08, 2016 | 2016 | Stryker Medical Division of Stryker Corporation | Stryker Medical initiated a voluntary recall of Stryker Performance-LOAD Cot Fasteners due to complaints that fastening system may not have been securely fastened or had an inability to fasten into the Performance-LOAD and bounced back during loading which could cause injury to consumers. | Under Investigation by firm | Stryker Medical initiated a voluntary recall of Stryker Performance-LOAD Cot Fasteners MODEL: 6392, due to complaints and service reports that some cots compatible with the Performance-LOAD cot fastening system may not have been securely fastened or had an inability to fasten into the Performance-LOAD (bounced back during loading) which could lead to injury of consumers. Customer notification letters were sent via Fedex on November 8, 2016. Customers were asked to quarantine the impacted devices and submit completed business reply forms by fax (269)488-8691 or email productfieldaction@stryker.com. Upon receipt of the completed business reply form, Stryker will contact you toarrange for the modification/repair of your Performance-LOAD unit to ensure that the unit securely fastens the cot within the ambulance.4. In the interim, until the service can be performed, there are additional actions that users can and should take to ensure that the cot is securely fastened into the Performance-LOAD system.a. Per the PerformanceLOAD Operations Manual (P/N 6392-009-001, pages 26-27. If customers have loaned or sold any of the products listed in this letter, please forward a copy of the recall notice to the new users and advise Stryker of their new location in the space provided on the business reply form. If you have any questions or concerns, please contact Stryker Customer Service (1-800-327-0770). Our normalbusiness hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 749 | 749 units | Worldwide Distribution to AK, AL , AR , AZ , CA , CT , DC , DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA , RI, SC, TN, TX, UT,VA, VT ,WA ,WV ,WY and Internationally to Canada and the Netherlands | Open, Classified |  |  | 749
Z-0742-2017 | 75812 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition AS System, x-ray, tomography, computed |  |  | November 15, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software update that provides software and firmware bug-fixes to improve system performance | Software Design Change | Siemens mailed a Customer Information Letter to affected customers on November 21, 2016 informing them of software updates that would provide bug-fixes to improve system performance and stability for customers with SOMATOM Definition AS, Definition Edge and Definition Flash systems using Software versions VA48A_SP1 and VA48A_SP2. The update will remotely be pushed to their systems or customers will be contacted by one of Siemens service organization to arrange a date for installation of the update on site. Customers with questions were instructed to call 1-800-888-7436. |  | 104 | 104 systems | Nationwide Distribution | Open, Classified |  |  | 104
Z-0734-2017 | 75863 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System using software version 10.6; |  |  | November 22, 2016 | 2016 | Accuray Incorporated | Accuray has become aware of a potential safety issue involving unexpected treatment robot motion when removing an accessory. This issue is limited to certain CyberKnifeÂż Systems with software version 1 0.6. | Software design | Accuray sent an Urgent Safety Advisory Notification letter dated November 22, 2016, to all affected customers. Letters alerted users to the issue and the sequence of events necessary for it to occur. Accuray reps will visit all sites to perform a software upgrade to correct the issue.For more information or any questions about this notification, please contact Accuray Customer Support at 1-877-668-8667 or visit the Accuray Technical Solution Center at http://www.accuray.com/services-support.For further questions, please call (408) 716-4600. |  | 40 | 40 affected devices | Worldwide Distribution - US Distribution to the states of : PA, WA, MT, NY, SC, CA, MN, AK, NJ, FL, KY, WI., and to the countries of : Bulgaria, Canada, France, Germany, India, Japan, Luxembourg, Netherlands, Portugal, Switzerland and Turkey | Terminated | May 22, 2017 | 181 | 40
Z-0738-2017 | 75469 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | BARDIAÂż 60cc Piston Irrigation Syringe, Sterile, Rx Only, The BARDÂż Syringe is available in a single use plastic design which features a 60cc or 70cc capacity graduated barrel. It has a thumb ring plunger for one-handed use. A catheter adapter tip and Luer adapter tip are included. BARDÂż Syringe allows one-hand use with a 60cc or 70cc graduated barrel and thumb-ring plunger and is indicated for catheter irrigation use. BARDÂż Syringe is intended for Cytoscope or Resectoscope irrigation, catheter irrigation and Foley catheter balloon inflation. |  |  | August 03, 2016 | 2016 | C.R. Bard, Inc. | Package Defect: Products may be at risk for having a slit defect on the package and that may affect the product sterility. | Process control | BARD Medical sent an Urgent Medical Device Product Recall letter dated August 3, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to immediately examine their inventory and quarantine product subject to the recall. Complete and return the accompanying Recall &amp; Effectiveness Check Form attached to the letter. If product was further distributed please identify those customers and notify them at once of the product recall by forwarding a copy of the letter. For questions regarding this recall call 770-784-6220. |  |  |  | Worldwide Distribution - US (nationwide) and Internationally to Australia, New Zealand, Canada, Mexico, France and Italy. | Open, Classified
Z-0745-2017 | 75252 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD 23 G x .75 in. BD VacutainerÂż Push Button Blood |  |  | August 08, 2016 | 2016 | Becton Dickinson &amp; Company | BD hasreceived multiple reports of units associated with the 23 G x .75 in. BD VacutainerÂż Push Button BloodCollection Set with 12 in. Tubing whose lots are exhibiting separation offront and rear barrels upon activation of the safety feature, which retracts theneedle. | Device Design | The firm, BD, sent a "URGENT PRODUCT RECALL" letter dated 8/8/2016 to its customer with a return response form via UPS 2nd Day mail on 8/8/2016.The letter described the product, problem and actions to be taken. The customers were instructed to immediately review your inventory for the specific lot numbers, quarantine the product and discontinue the shipment of the affected product; complete and return the Recall Response Form by fax to BD at 1-866-913-0661 or email to bd6835@stericycle.com (even if you do not have any product); return all affected products with completed response form per packing instructions; if you have distributed this product, identify your customers and notify them immediately of this product recall; if you would like BD to conduct the notification to your customers, email your customer list within 3 business days to Matthew.Kelleher@bd.com.if you have any questions or require assistance with the return of the recalled product and/or availability of credit product, contact 1-888-628-0732 between 8Am and 5 PM ET Monday through Friday. |  | 1788150 | 1,788,150 units | Worldwide distribution- US (nationwide) and countries of: Canada, Belgium, Brazil. | Open, Classified |  |  | 1788150
Z-0750-2017 | 75786 | No | POA |  |  | Microbiology | EUA - Emergency Use Authorization | LightMixÂż Zika rRT-PCR Test |  |  | November 16, 2016 | 2016 | Roche Molecular Systems, Inc. | The LightCyclerÂż 480 algorithm used for the LightMixÂż Zika rRT-PCR Test, EUA (catalog number07987897001) occasionally (~1% of the time) calls a positive result for a negative sample if there isan inflection point in the fluorescence signal. | Software design | Roche Molecular Diagnostics Inc. sent an Urgent Medical Device Correction Letter dated November 16, 2016, to all affected customers via UPS. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product until Roche provides updated Instructions for Use to assist in result interpretation. If product was further distributed, ensure the Urgent Medical Device Correction letter is provided to that site. Also, customers were instructed to completed the attached fax form and fax it to 1-888-912-2188. For questions customers should call 1-800-526-1247. |  | 24 | 24 pieces | Nationwide Distribution | Open, Classified |  |  | 24
Z-0748-2017 | 75351 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ellman A CYNOSURE COMPANY Minor Surgery Bipolar Cable, Contents: Five (5) Disposable Bipolar Cables,STERILE |  |  | September 07, 2016 | 2016 | Ellman International, Inc. | Mislabeling of the "expiration date". The expiration date on the outer box and the individual pouches do not match. The expiration date on the box is correct with a date of 2018-02. | Employee error | Cynosure, Inc. dba Ellman distributed Urgent Medical Device Recall Notices dated September 6, 2016, and response forms to their customers via UPS overnight. Cynosure has identified an error in the Expiration Date on the label of the Minor Surgery Bipolar Cable pouch. The Expiration Date on the individual pouches that are provided in the 5 pack boxes do not match the Expiration Date on the box. The Expiration Date on the box is correct with a date of 2018-02, but the label on the individual pouches has a date of 2021-02. Actions To Be Taken - Return any cables subject to recall. You will be contacted by a Cynosure Customer Service Representative who will provide you with a RMA number for any cables that are being returned. Complete and return the response form via email to choy@cynosure.com. Cynosure will replace affected products at no charge upon receipt of the returned cables. Customers with questions should call 978-256-4200 or 800-886-2966.For questions regarding this recall call 516-267-6502. |  | 370 | 370 boxes (US - 159 boxes, OUS - 211 boxes) | Worldwide Distribution - US including CA, FL, MA, MD, MI, MN, MS, OH, PA, VA, TX &amp; WV; **International**Australia, Bangladesh, Greece, Italy, Japan, Lebanon, Russian Federation, Singapore, Spain &amp; Vietman. | Open, Classified |  |  | 370
Z-0752-2017 | 75710 | No | LRW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Metzenbaum Scissors 7 , Curved, Sterile |  |  | October 28, 2016 | 2016 | Integra York PA, Inc | The Chevron seals of 6x10 Tyvek peel pouch used to package the sterile instruments may separate causing the packages to open and compromise sterility of the instrument. | Process control | The firm, Integra, mailed an "URGENT MEDICAL DEVICE RECALL NOTICE" dated October 26, 2016 to customers affected by the recall to notify them of the issue. The notice described the product, problem and actions to be taken. The customers were instructed to immediately review their inventory to determine if they had any of the recalled product on hand and to discontinue distributing the product; review your customer traceability records for shipment of the product and forward a copy of this Recall Notice to any of your customers that purchased the affected lots; complete and return the Acknowledgement Form that accompanied the notice (even if you do not have any product(s)) by email to: FCA3@integralife.com or fax to: 1-609-750-4220.Should you have any questions regarding these instructions, please contact Customer Service at 1-866-854-8300. |  | 50 | 50 units | US Distribution in NC and TX. | Terminated | April 17, 2017 | 171 | 200
Z-0749-2017 | 75682 | No | OUP | Powered surgical instrument for improvement in the appearance of cellulite. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Cellfina Prep Pack, Part No. CP1 |  |  | November 14, 2016 | 2016 | Ulthera Inc | It was discovered that a non-sterile vacuum tube is supplied within the pack, although the Instructions for Use (IFU) describe the use of a sterile vacuum tube as part of the procedure setup. | Other | Ulthera sent an Urgent Medical Device Recall letter dated November 15, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.It was discovered that a non-sterile vacuum tube is supplied within the pack, although the Instructions for Use (IFU) describe the use of a sterile vacuum tube as part of the procedure setup. Customers are informed of the actions to be taken. Customers with any technical or logistical questions are instructed to contact technical support at Ulthera, Inc. at (480) 619-4069. The letter informs the customers with questions related to their Cellfina business are instructed to contact Merz North America at 888-751-7546 or at mymerzsolutions@merz.com. |  | 6562 | 6562 units | Worldwide Distribution - US Distribution and to the countries of :Europe (Austria, France, Germany, Italy, Netherlands, Switzerland, Spain, UK) | Open, Classified |  |  | 6562
Z-0755-2017 | 75584 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Avance, Avance CS2, Amingo. |  |  | November 03, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue where certain Avance CS2, Avance and Amingo anesthesia devices can transition to a System Malfunction state ifthe lower storage drawer containing the optional large tray insert accessory closed with an abnormally high amount of force. | Device Design | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref# 34079 dated November 3, 2016, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 23098 | 23,098 (US = 8,711; OUS = 14,387) | USA ( nationwide) Distribution including Puerto Rico and DC. and to the countries ofAustraliaAUSTRIAAzerbaijanBahrainBangladeshBarbadosBelarusBelgiumBOLIVIABOSNIA-HERZEGOVINABrazilBRUNEIBULGARIACanadaCHILEChinaCOLOMBIACOSTA RICACroatiaCyprusCzech RepublicDenmarkECUADOREgyptEstoniaEthiopiaFinlandFRANCEGeorgiaGERMANYGhanaGibraltarGREECEHONG KONGHUNGARIAIcelandIndiaIndonesiaIRAQIrelandIsraelITALYJAPANJORDANKAZAKHSTANKenyaKOREA, REPUBLIC OF KuwaitLatviaLebanonLibyan Arab JamahiriyaLITHUANIAMACEDONIAMALAYSIAMaltaMexicoMOLDOVAMOROCCONAMIBIANetherlandsNew ZealandNicaraguaNigeriaNORWAYOmanPakistanPanamaPARAGUAYPERUPHILIPPINESPolandPORTUGALQatarROMANIARussian FederationSaudi ArabiaSerbiaSERBIA-MONTENESingaporeSlovakiaSLOVENIASOUTH AFRICASPAINSWEDENSWITZERLANDSYRIATAIWANThailandTunisiaTURKEYUkraineUnited Arab EmiratesUNITED KINGDOMUruguayUZBEKISTANVenezuelaVietNamYemen | Open, Classified |  |  | 23098
Z-0759-2017 | 75693 | No | HCF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Weck Facial Closure Systems have application in laparoscopic procedures for approximation of tissues and percutaneous suturing for closing incision sites. |  |  | November 16, 2016 | 2016 | Teleflex Medical | The devices wings may become partially detached from the EFx Shield during use. | Other | Consignees were notified via letter on 11/16/2016. |  | 2265 | 2265 units | Domestic; US Nationwide; International: Belgium, Canada. | Open, Classified |  |  | 2265
Z-0757-2017 | 75249 | No | EOQ | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Ultrasound Video Bronchoscope |  |  | May 31, 2016 | 2016 | Pentax of America Inc | The EB-1970UK Ultrasound Video Bronchoscope shipped between theclearance date of April 2014 and September 2015 contained the Operation andReprocessing IFUs that were in use prior to the 2014 clearance, and therefore did notinclude the cleared Operation and Reprocessing IFUs. | Process control | Pentax sent a US Urgent Field Correction letter to all affected customers with a field correction/field safety response form via USPS on May 31, 2016. The letter identified the product, the problem, and the action to be taken by the customer. Revised instructions were as follows:The instructions have been revised in the following manner:Âż Operation IFUo Integrated all models of each family into the IFUo Modified the indications for use language to be consistent with similar productso Enhanced the explanation of the operation of the deviceo Revised and added cautions and warningsÂż Reprocessing IFUo Included more specificity regarding reprocessing activities to aid in the proper method for reprocessing theendoscopes, including:Âż number of brush and rinse steps,Âż quantity of fluid volumes,Âż use of additional text, figures, cautions, and warningsIdentification of Affected Endoscope ModelsTable 1 provides an overview of the Affected Endoscope models, and the part number and revision of the historical andrevised IFUs. The accompanying USB flash drive contains all the revised IFUs identified below. Please distribute only thoseIFUs which support the products owned by your facility.Table 1Affected Models Historical Version of IFUs Revised IFUsEB-1970UKUltrasound Video BronchoscopeZ870 Owners IFUS016_R02 Operator IFUS021_R04 Reprocessing IFUNo.388_R01 AddendumNo.392_R00 AddendumEB-70K seriesVideo BronchoscopesZ705 Operation IFU S015_R04 Operator IFUS020_R04 Reprocessing IFUNo.388_R01 AddendumNo.392_R00 AddendumZ706 Reprocessing IFUEB-75K seriesVideo BronchoscopesS003 Owners IFUCustomers were instructed to ensure that user departments discard the prior versions of the Instructions for Use and ensure that all users read and carefully follow the revised Operation and Reprocessing IFUs to ensure the proper use and cleaning of the affected PENTAX endoscopes. Customers were advised to contact their PEN |  | 113 | 113 units | Nationwide Distribution | Open, Classified |  |  | 113
Z-0762-2017 | 75262 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Randox Liquid Cardiac Controls Level 1, Level 2, and Level 3 are liquid controls containing BNP, CK MB Mass, Digoxin, D-Dimer, Homocysteine, hsCRP, Myoglobin, NT-ProBNP, Troponin I, and Troponin T. They have been developed for use in the quality control of BNP, CK MB Mass, Digoxin, D-Dimer, Homocysteine, hsCRP, Myoglobin, NT-ProBNP, Troponin I, and Troponin T assays on various clinical chemistry and immunoassay systems. This in vitro diagnostic device is intended for prescription use only. |  |  | September 16, 2016 | 2016 | Randox Laboratories | An internal real-time stability monitoring program revealed decreased recovery of myoglobin (-20% bias) for Randox Immunoturbidimetric and Siemens Centaur Classic methods and CKMB (-21% bias) for Roche e411, BioMerieux Vidas, and Siemens Centaur Classic methods. | Nonconforming Material/Component | Randox sent an Urgent Field Safety Notice dated September 2, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please note the control value sheets have been updated with the new values. A FSN has been completed to advise customers on the appropriate action to take.For further questions, please call (304) 728-2890. |  | 2817 | 2817 | Worldwide Distribution - US Distribution. | Open, Classified |  |  | 2817
Z-0760-2017 | 75312 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | INOVO Purefill Oxygen Compressor Model PF1100 |  |  | July 21, 2016 | 2016 | Inovo, Inc | Device can fail during operation. | Under Investigation by firm | The firm, Inovo, Inc, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated Sept 2, 2016 to all their consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine and isolate the devices in their inventory and return them to Inovo Inc. Also, if the consignee sold or distributed the device, the consignee needs to retrieve and return the units to Inovo Inc. as soon as possible and complete and return the Medical Device Return Response form in the self-addressed stamped envelope provided or email to: regulatory@inovoin.com or fax to Inovo, Inc. Any questions please contact at 814-443-7602 or via e-mail at regulatory@inovoinc.com. |  | 117 | 117 units | US Nationwide Distribution to states of: AL, AR, CT, FL, IA, KS, KY, MD, MI, MN, MO, NE, NJ, NV, NY, OH, RI, SC, TN, TX, UT, and WA. | Open, Classified |  |  | 117
Z-0756-2017 | 75481 | No | FMH | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | Covidien Precision Specimen Container- 4 oz (118 mL), Plastic Screw Top Cap, OR Packaged, Sterile |  |  | October 14, 2016 | 2016 | Covidien LLC | Sterility of the outer surface of the container compromised due to packaging defect. | Packaging process control | Covidien initiated recall by letter on October 14, 2016. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory of the affected lots to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department.Accounts requested to completed Recalled Product Return From to obtain an RGA.Questions: Contact your Medtronic representative or Customer Service at (800) 882-5878. |  | 5187100 | 5,187,100 Units | NationwideForeign: Japan | Open, Classified |  |  | 5187100
Z-0754-2017 | 75601 | No | FAS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HF Cable WA00014A, Endoscopic electrosurgical unit and accessories |  |  | November 03, 2016 | 2016 | Olympus Corporation of the Americas | Software malfunction that results in incorrect generation or display of error codes. | Under Investigation by firm | Olympus mailed on 11/03/2016 an Urgent Communication Letter to customers affected by the recall. The letter requested that consignees inspect their inventory for the affected lots and call Customer Service to obtain a Returned Material Authorization (RMA) to return the product. |  | 494 | 494 units | Distributed Nationwide | Open, Classified |  |  | 494
Z-0763-2017 | 75770 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Ceramax Skin Barrier Cream: |  |  | November 21, 2016 | 2016 | Oculus Innovative Sciences, Inc. | Firm became aware that professional samples of Ceramax are not meeting stability standards set for the product. | Packaging process control | Oculus issued Medical Device Recall letters dated November 21, 2016, to Sales reps notifying them to stop all further distribution and immediately remove the affected lots of Recall product from sample inventory. Customers were also instructed to sign the Response form verifying receipt of the letter. Customers with questions were instructed to call Oculus Innovative Science-Regulatory Affairs at 707-559-7181. For questions regarding this recall call 559-7191. |  | 68970 | 68,970 units | Nationwide Distribution | Terminated | June 13, 2017 | 204 | 68970
Z-0753-2017 | 75826 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ecolab, Microtek Rolled Probe Cover PC0907NB, Rx Only. |  |  | November 30, 2016 | 2016 | Ecolab Inc | The reason for the recall is that the inner carton box of the Rolled Probe Cover PC0907NB has the incorrect expiration date (2015-10-01) on the label while the individual pouch labels and outer shipper have the correct expiration date (2018-10-01). Themanufacturing date was inadvertently used as the expiration date on the inner carton box labels. | Incorrect or no expiration date | Ecolab sent an "Urgent Medical Device Recall" letter dated November 30, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised to quarantine the product for return. If the product was further distributed notify them of the recall. Requested customers to complete and return the Recall Response Form. For questions contact Customer Service by phone at 1-800-824-3027 or e-mail customerservice@microtekmed.com.For further questions please call (651) 250-4385. |  | 40 | 40 cases (144 eaches per box) | US Distribution to the states of : CA, FL, MI, NJ, TX, UT and VA.. | Open, Classified |  |  | 40
Z-0766-2017 | 75645 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Empowr PS Knee System Box Cut Guide |  |  | September 19, 2016 | 2016 | Encore Medical, Lp | It has been discovered that during Empowr PS Total Knee replacement, while chiseling for bone removal in the box area, femoral chipping has occurred in surgery. | Device Design | DJO Global sent an Urgent Field Safety Notice dated September 19, 2016, to all affected customers. The field safety notice provided a newly revised Surgical Technique to educate users on proper use of the device. Customers were instructed to become familiar with the new surgical techniques and notify all of their applicable users of the revised methods. Customers were also instructed to sign the Acknowledgement and Receipt Form. Customers with questions were advised to call 760-734-3551.For questions regarding this recall call 512-832-6302. |  | 162 | 162 units | Nationwide Distribution | Open, Classified |  |  | 162
Z-0758-2017 | 75663 | No | LZA | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Sterling Nitrile Sterile Powder Free Exam Gloves, S, Sterile. |  |  | November 10, 2016 | 2016 | Halyard Health | Defects: Product may be discolored and brittle resulting in the inability of the user to don the gloves without tearing. | Under Investigation by firm | Consignees were notified by letter via Fed Ex on/about 11/10/2016. |  | 59 | 59 cases | CA, FL, GA, IA, IN, LA, MA, MN, MT, NM, OR, TX and UT. | Terminated | May 04, 2017 | 175 | 59
Z-0765-2017 | 75675 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Reprocessed Xcel Bladeless Trocar (H12LP) |  |  | November 11, 2016 | 2016 | Stryker Sustainability Solutions | An increase in reports indicating that the reprocessed H12LP Trocar adjustable plug may break in the packaging or during use. | No Marketing Application | Stryker sent an Urgent Medical Device Recall letter to all affected customers on November 11, 2016.Customers were instructed to discontinue use of the affected lots and were asked to follow the link below to complete the Recall Effectiveness Check Form and indicate if any devices from the lots referenced in Attachment 1 remain in their inventory. This form must be completed even if no affected product is found.www.novasyte.com/sss142Use the following reference number as you complete the form: [Ship to Account #].If the form indicates that any affected devices remain in your inventory, a prepaid shipping label will be issued for the return of the product. Credit for all affected devices returned will be issued to your facility.Address any questions regarding the recall to Novasyte by calling (760) 279-3180 or emailing sss142@novasyte.com. |  | 40476 | 40,476 devices | Worldwide Distribution - US (nationwide) and Internationally to Israel and CanadaForeign distribution to the following; IL and CA | Terminated | May 05, 2017 | 175 | 40476
Z-0764-2017 | 75735 | No | GDT | Removable skin staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | WECK Visistat 35W, Disposable Skin Stapler, Rx Only, Sterile, Teleflex Medical, NC 27709. |  |  | November 18, 2016 | 2016 | Teleflex Medical | Sterility Deficiency: The packages are not sealed properly and thus the sterility is not assured. | Packaging process control | The firm, Telexflex Medical, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 11/18/2016 to its Consignees/Customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any affected products; complete and return the Recall Acknowledgement Form by fax to 1-855-419-8507, Attn: Customer Service or email to: recalls@teleflex.com (even if you do not have any affected product) and return the product with the completed Recall Acknowledgement Form. A customer service representative will contact you with a RGA # and provide instructions for the return of the product.If you have any other questions, feel free to contact your local salesrepresentative or Customer Service at 1-866-246-6990. |  | 96402 | 96,402 units (9,120 ea US, 87,282 ea OUS) | Worldwide Distribution- U.S. (Nationwide) and countries of: Australia, Belgium, Canada, China, Germany, Hong Kong, India, Japan, Korea, Malaysia, New Zealand, Singapore and Thailand. | Open, Classified |  |  | 96402
Z-0761-2017 | 75743 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 15 monitor/defibrillator is intended for use by trained medical |  |  | March 11, 2010 | 2010 | Physio-Control, Inc. | Customer complaint was received relating to a LifePak15 device that would power on then power down unexpectedly after a short period of use. When the LP15 powered down, power could be immediately restored by activation of the power button. The device was connected to a cellular modem from MultiTech systems for use with LIFE NET services. The corrective action was completed by Velocitor Solutions in December 2010. | Environmental control | Velocitor initiated this field action and notified affected customers on March 11, 2010. This action was reported to Physio-Control as completed on 12/21/2010. |  | 1399 | 1399 units (866 in the US and 533 outside US) | worldwide | Completed |  |  | 1399
Z-0770-2017 | 74587 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE blanket LED Phototherapy system, neonatal phototherapy. Used to provide treatment for neonatal Hyperbilirubinemia. |  |  | October 28, 2016 | 2016 | Natus Medical Incorporated | neoBLUE blanket LED Phototherapy System distributed since release of the product in 2011 is recalled due to customer complaints regarding discoloration/ degradation of fiberoptic bundle at connection of pad and box. | Device Design | Natus began sending the URGENT RECALL NOTIFICATION / FIELD SAFETY NOTICE neoBLUE blanket system performance letter dated October 27, 2016 and the Technical Bulletin to US customers on October 28, 2016.Natus plans to notify their international customers on December 2, 2016. Customers were informed in March 2015 of a potential issue of discoloration/degradation/melting of the fiber optic bundle at the pad connector which is inserted in the neoBLUE blanket light box. Included with that letter was a Technical Bulletin which described the issue and addressed many of your potential questions.The purpose of this current letter (dated October 27, 2016) is to communicate Natus' plan for a follow-on field correction for neoBLUE blanket systems shipped between 2011 through March 21, 2016 (Serial Numbers up to and including xxxx004282), and to remind customers of the instructions contained in the neoBLUE blanket Technical Bulletin.Natus is in the process of confirming an updated neoBLUE blanket configuration which will not be susceptible to the degradation described above. However, the date of project completion is uncertain at this time. Natus will contact customers when the new configuration is available, at which time Natus intend to initiate a field corrective action to replace or upgrade your device(s).Please contact Natus Quality_Programs@natus.com if you have any further questions. Please call (800) 426-0353. |  | 2619 | 2619 units (1925 units in the US and 694 units OUTSIDE the US) | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : including American Samoa and Guam, and in the following countries:Albania, Argentina, Australia, Austria, Bangladesh, Canada, Chile, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Hong Kong, Indonesia, Iraq, Israel, Italy, Korea, Kuwait, Latvia, Malaysia, Mongolia, Namibia, Netherlands, Nigeria, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia (Slovak Republic), Slovenia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Viet Nam, and Virgin Islands (British). | Open, Classified |  |  | 2619
Z-0768-2017 | 75576 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MicroScan LabPro Data Management System. |  |  | October 25, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter has received and confirmed reports of an intermittent and unexpected behavior when loading new panels into a WalkAway instrument usingLabPro Data Management System version 4.42. The issue could cause workflow interruption with a potential of delaying reporting results due to the inability to begin processing new panels. | Device Design | An Urgent Medical Device Recall letter dated 10/25/16 was sent to all their customers to inform them that Beckman Coulter has received and confirmed reports of an intermittent and unexpected behavior when loading new panels into a WalkAway instrument using LabPro Data Management System version 4.42. Customers are informed that the issue could cause WalkAway workflow interruption with a potential of delaying reporting results due to the inability to begin processing new panels. The letter informs the customers of the actions to be taken and the resolution for the product on recall. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding the notice are instructed to contact their Customer Support representative: http://www.beckmancoulter.comor call 1-800-677-7226, option 1 in the United States. |  | 131 | 131 units total (130 units in US) | Worldwide Distribution -- United States and Mexico. | Terminated | May 01, 2017 | 188 | 131
Z-0767-2017 | 75571 | No | FOX | Infusion stand. | General Hospital | General Hospital | 510(K) Exempt | IVEA 500A |  |  | November 01, 2016 | 2016 | Firefly Medical, Inc. | The dowel pin could potentially slip out of its press fit, resulting in the product potentially slumping to one side which could cause the patient to stumble. | Process control | Firm notified customers on **11/18/16** via certified "Urgent: Medical Device Recall (Field Correction) IVEA Model 500A" letter to correct product in the field. Customers were asked to thoroughly read the letter and to immediately follow the actions in the enclosed Product Correction Kit. The instructions include steps for cleaning the correction area and installing the permanent adhesive caps, as well as, completing and returning the product correction form to Firefly Medical. Customers may continue to use their units until the Product Correction actions are complete, as long as pins are not protruding. If the pins are protruding, then customers are to stop using the device immediately and complete the product correction instructions. Customers can call 970-472-5323 or e-mail info@iveamobility.com if they have any questions. |  | 239 | 239 | Worldwide Distribution -- USA and Canada. | Terminated | August 01, 2017 | 273 | 239
Z-0769-2017 | 75228 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Updated Natus neoBLUE blanket LED Phototherapy Systems. It provides phototherapy underneath the infant and can be used in a bassinet, open bed, radiant warmer, incubator, or while holding the infant. |  |  | September 16, 2016 | 2016 | Natus Medical Incorporated | The firm sent the Technical Bulletin to consignees who received the reliability updated neoBLUE blanket LED Phototherapy Systems between March 2016 to August 2016. The reliability update to prevent the early failure of the fiberoptic pads experienced by some customers on older systems. These failures on previous systems involved discoloration and melting of the fiberoptic bundle at the connector that is inserted in the neoBLUE blanket light box. | Device Design | Natus sent the Urgent Recall Notification/Field Safety Notice and Customer Reply Form (DOC-021042A), dated September 16, 2016 and the Technical Bulletin (DOC-021016) to US consignees via email.The letters and technical bulletin for foreign consignees was sent via email to the appropriate Natus Distribution Partner (DP) on November 7, 2016.US Customers can email the customer reply form to Natus_Quality_Progams@natus.com or fax to (847-949-6031). US Customers with questions can call Natus Technical Service at 888-496-2887. International customers should contact their local distributor with any questions. |  | 713 | 713 units total (403 units in US and 310 units outside US) | Worldwide Distribution - USA (nationwide) Distribution to the states of : AL, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, NC, ND, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WI, WV and Puerto Rico. , and to the countries of : Argentina, Austria, Canada, Chile, China, Denmark, Dominican Republic, Finland, France, Germany, Hong Kong, Hungary, Indonesia, Italy, Latvia, Malaysia, Namibia, Paraguay, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Taiwan, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Uruguay, Viet Nam, and Virgin Islands (British). | Open, Classified |  |  | 713
Z-0779-2017 | 75708 | No | CGJ | Follicle-stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DELFIA hFSH kit. In Vitro Diagnostic for the determination of hFSH in serum |  |  | November 16, 2016 | 2016 | PerkinElmer Health Sciences, Inc. | A DELFIA hFSH kit lot failed real time stability testing. The real time stability data demonstrated a decrease in hFSH concentrations between -4.8% to -11.3% for control sera. The -11.3% bias was observed at an hFSH target concentration of 101 IU/L. The change observed in the serum samples within the range 1.76  67.0 IU/L hFSH averaged -2.9%. The reference range of the assay spans 0.60 to 98.0 IU/L. | Other | PerkinElmer initiated recall on November 16, 2016, delivered to the end customer by email facilitated through the PerkinElmer Technical Support group located in Akron Ohio. Technical Support will be asked to make sure the customers receive the Recall Letter and understand the recommended actions.Users asked to inspect inventory, discontinue use destroy the kits and complete the Response Forms and Destruction.Replacement kits will shipped. |  | 79 | 79 kits | CA, CO, NY, WAForeign:Austria, Bolivia, Bosnia , Sweden, Norway, Netherlands, Japan, Great Britain, | Terminated | July 05, 2017 | 231 | 79
Z-0772-2017 | 75420 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS Percutaneous Introducer Set under the following labels: 1) 5 F Stiffened Micro-Introducer Kit, 45CM SS/SS ECHO B PG, Item Number H787065970395, Catalog No.REF 06597039, 2) 5 F Stiffened Micro-Introducer Kit, STF 5F M.I. KIT 45CM SS/SS ECHO S PG, Item Number H787065970415, Catalog No. REF 06597041 |  |  | September 28, 2016 | 2016 | Angiodynamics, Inc. | Greatbatch Medical the manufacturer of the 4F and 5F sheath/dilator components has determined that the products listed in their September 7, 2016 Recall Notification have the potential for the rotating luer to detach from the dilator hub during use. AngioDynamics has confirmed that affected sheath/dilators (Greatbatch Model Numbers 10904-001 and 10904-002) have been included in packaged AngioDynamics Stiffened Micro-Introducer Kits. | Component design/selection | Angiiodynamics sent an Urgent Voluntary Medical Device Recall letter dated September 28, 2016, to all affected customers. Recall notifications were delivered by Federal Express. Customers were instructed to segregate and return all affected devices to AngioDynamics and to complete and return the Reply Verification Tracking Form, provided in the recall notification by fax to 1-800-782-1357. Customers were encouraged to forward a copy of the recall notification to all sites to which have distributed affected product. If affected product is located in your institution, please call AngioDynamics Customer Service at 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday  Friday: Eastern Standard Time) to obtain a replacement or credit for your returned product. Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return); following the directions on this page and the Reply Verification Tracking Form via to recall@angiodynamics.com. Customers with questions should contact customer service at 1-800-772-6446. |  | 162 | 162 boxes | Nationwide Distribution to AL, AR, CA, FL, GA, IL, KY, LA, MI, MN, MO, NC, OH, OK, SC, NY, ME, PA, MD, VA, WA, WI, WV, TN, and TX | Open, Classified |  |  | 2145
Z-0780-2017 | 75588 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PerkinElmer 20ul Clear RoboRack Tip, Pipette (Qty 960) |  |  | October 13, 2016 | 2016 | PerkinElmer, Inc | Defective tips have the potential to pipette/dispense volumes 50-100% lower than is specified for the tips. | Process control | PerkinElmer initiated the recall on Sept 1. 2016 by letter dated . Customers were told of the defect, potential for inaccurate measurements, and requested to inspect for the lots stop using and destroy. Replacement product to be provided. Accounts asked to compelte Response form. Questions call your PerkinElmer representative. |  | 33312 | 33312 | IA, MD, NY, WAForeign: France, Germany, The Netherlands, Hungary, Singapore | Open, Classified |  |  | 33312
Z-0771-2017 | 75781 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RUSCH, Pocket PAC IC Intermittent Cath Kit, 8 FR, 12 FR, 14 FR, 10 FR and 16 FR, Rx Only, Sterile, |  |  | November 22, 2016 | 2016 | Teleflex Medical | Labeling Inconsistency: The products have beenlabeled with the incorrect expiration date which exceeds the actual, validated 35-month shelf life. | Error in labeling | Teleflex sent an Urgent Medical Device Recall Notification letter dated November 22, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers with affected stock, immediately discontinue use and quarantine any affected products. To return product, customers were advised to complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recall@teleflex.com. Customers with no affected stock should also complete the enclosed Recall Acknowledgment Form to confirm receipt of the letter. Customers with questions should contact Customer Service at 1-866-246-6990. |  | 102843 | 102,843 units | Nationwide Distribution to CA, CO, FL, GA, IL, KS, KY, MA, MO, MS, NC, NJ, NY, OH, OK, OR, PA and TX. | Open, Classified |  |  | 102843
Z-0786-2017 | 75794 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AbbVie PEG Kit 20 FR |  |  | November 22, 2016 | 2016 | AbbVie Inc. | AbbVie is issuing an urgent field correction notification due to an increased number of reports of disconnections of the AbbVie 20FR PEG tube from the Y-connectors. | Device Design | The firm sent an URGENT: MEDICAL DEVICE FIELD CORRECTION NOTIFICATION AbbVie PEG Kit 20 FR to affected customers describing the reason for the notification. The letter asks affected customers to forward this letter to all Movement Disorder Specialists and physicians (e.g. GIs, IRs, and surgeons) that place PEG- J Tubes within your facility.The letter describes actions to be taken for patients still using the AbbVie PEG 20FR Tube. The letter requests user facilities to continue to monitor the performance of the Tube. The firm states that there is no need to replace the AbbVie PEG 20FR Tube if it is functioning as expected and maintaining a viable connection with the Y-Connector. The letter provides the following instructions when establishing a connection between Abbvie PEG 20FR Tube with the 20FR Y-connector, always refer to AbbVie J Intestinal Tube 9 FR to PEG 15 and 20 FR Instructions for Use for 62918-001. The letter notes the following:1. Ensure that the set-up process clearly follows the sequence of events as described in the instructions for use and the following clarifications. 2. Ensure that the AbbVie PEG 20FR tube is connected to the corresponding 20 Y- connector (see photograph below which shows the difference between the 15FR and 20FR connection diameter). 3. DO NOT USE ALCOHOL CONTAINING PRODUCTS with the AbbVie PEG 20 FR tube as this may result in damage to the tube.4. Ensure that the Tube Clamp and Tube Clip on Fixation Plate are both closed. 5. Cut off the proximal cone and Fixation Loop of the AbbVie PEG Tube (15FR or 20FR). Note: Ensure that the cut is perpendicular (90 degree angle) to the tube and not on a diagonal as this ensures proper seating of the tube when engaged into the pin of the Y-connector fitting. 6. Slide the colored Fixation Screw (including the outer ring) over the Abbvie PEG Tube:a. For 15 FR, use the Blue/White screwb. For 20 FR, use the Violet/White screw7. Grasp near the |  | 541 | 541 units | U.S. consignees located throughout the country and Puerto Rico. Foreign Consignees: Worldwide. | Open, Classified |  |  | 541
Z-0782-2017 | 75055 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ultroid Hemorrhoid Management Systems 110/220 VAC and procedures kit including the Ultroid Disposable Sterile Probes |  |  | October 19, 2016 | 2016 | Ultroid Technologies, Inc. | Products are not designed, qualified, manufactured, and/or managed under a state of control per internal quality system procedures and external laws, regulations, directives, standards, and/or guidance (such as 21 CFR 820 Quality System Regulation) | Process control | On October 19, 2016 Ultrid sent a recall letter to all their consignees informing them of the recall and to cease further distribution of this product, respond the self-addressed response form and returned for a voucher. For more information contact Melanie Rosario at 727-898-0717 |  | 1129 | 1,129 | AL, AR, AZ, CA, CO, CT, DE, FL, GA, ID, IL, KS, LA, MD, ME, MI, MN, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA and Puerto RicoAsturias, Australia, Chile, Greece, Hong Kong, India, Indonesia, Iraq, Japan, Kuwait, Malaysia, Mexico, Nigeria, Qatar, Russia, Slovenia, Spain, Syria, Trinidad, Turkey, UAE, UK, and Vietnam | Open, Classified |  |  | 61059
Z-0781-2017 | 75459 | No | LGW |  |  | Neurology | PMA | Algovita Spinal Cord Stimulation System, Percutaneous Leads, 8-Electrode Percutaneous Lead Kit; |  |  | October 11, 2016 | 2016 | Nuvectra | Nuvectra is conducting a recall due to two clinical risks that are associated with certain percutaneous leads currently being used in the Algovita Spinal Cord Stimulation System.It has been reported that in some percutaneous leads, the stylet protruded beyond the distal lead tip. | Nonconforming Material/Component | Nuvectra field sales associate or contracted distributor were sent a Nuvectra Urgent Field Safety Notice (Field Action Number: 2016-03) on 10/11/2016 via Fed-Ex. The letter provided background information, described the problem, &amp; the product involved in the recall. The letter listed the potential clinical implications, rate of occurrence, recommendations and mitigations and action needed. Provided consignees with notices for them to deliver to physicians as well as Response forms for them and the physician. For questions contact Kathy Jo Fahery +01.763.951.8219. |  | 1391 | 406 US, 985 OUS | US: Nationwide (CA, TX, KY, FL, IL,PA,OH, GA, NJ, NC, KS,AR, CO, NE, NY, VA, AZ, OK).OUS: Germany | Open, Classified |  |  | 1391
Z-0787-2017 | 75816 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | AMS-651, 120" Microbore extension set with slide clamp; Product Code: AMS-651 |  |  | November 18, 2016 | 2016 | Vygon MFG, Inc., dba/ Churchill Medical Systems, Inc. | Incorrect priming volume is printed on the device package. | Device Design | Vygon sent an Urgent Medical Device Recall letter dated November 18, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to take the following actions: 1. Check your stock for the products included within the scope of the recall. Immediately cease use and distribution, and quarantine all affected product immediately; 2. Count and document your affected inventory on the "Recall Acknowledgement and Inventory Return Form; 3. Fax or email the form to Vygon including customer contact information; 4. Vygon customer service will contact the customer and issue a Return Merchandise Authorization number; 5. Return the product as well as the original signed Recall Acknowledgement and Inventory Return Form using the instructions from Customer Service involving product pick-up; 6. All inventory will be credited or replaced at no cost to you. For questions contact local sales representative or Vygon's Customer Service Department at 1-800-473-5414.or by email at customerservice@vygonus.com. |  | 31500 | 31,500 | US distribution in the state of CA | Open, Classified |  |  | 31500
Z-0788-2017 | 75486 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II Implantable Drug Infusion System. This system includes: Model 8870 Software Application Card, Model 8840 N'Vision Clinician Programmer, and Model 8637 SynchroMed II Implantable Drug Infusion Pump (supplied in 20 ml or 40 ml reservoir size.) |  |  | October 03, 2016 | 2016 | Medtronic Neuromodulation | Medtronic is following up to a May 2013 communication regarding the Priming Bolus function for the SynchroMed Infusion System. Medtronic is updating the Model 8870 software application card (to version AAU01) and the SynchroMed pump labeling to address the priming bolus issue. | Under Investigation by firm | Medtronic Representatives began visiting consignees on 10/3/2016 to replace previous version Model 8870 software application cards, contained within the model 8840 NVision Clinician Programmer, with the updated version (new version is AAU01). Beginning October 20, 2016, Implanting and Following physicians were sent an Urgent Medical Device Correction letter dated September 2016, . This letter to physicians is informing them of the updated Model 8870 software application card, and changes to SynchroMed infusion system labeling. A response form is requested to be returned to Medtronic.Customers with questions were instructed to contact Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 22298 | 22,298 software cards | Worldwide Distribution - (US Nationwide) Internationally to Germany, Greece, Australia, United Kingdom, Switzerland, Italy, Norway, France, Canada, and Malta. | Open, Classified |  |  | 22298
Z-0825-2017 | 75844 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia MR-OR; 781437, 781438 |  |  | December 01, 2016 | 2016 | Philips Electronics North America Corporation | Possibility that a patient or user may be exposed to a very small level of touch current. | Unknown/Undetermined by firm | Philips sent an Urgent Field Safety Notice Medical Device Correction letter dated November 2016 to all affected customers. The letter explains the issue and actions planned by Philips. There is no need for action to be taken by customer or user. Customers with questions should contact their local Philips representative Technical Support Line 1-800-722-9377. |  | 9 | 9 | Worldwide Distribution -US including CA, NY and Internationally to United Kingdom, Australia, Norway, and Thailand | Open, Classified |  |  | 9
Z-0789-2017 | 75742 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | AcuvanceÂż2 Safety Catheter, Product No. 1720-AI |  |  | November 21, 2016 | 2016 | Smiths Medical ASD Inc. | Smiths Medical became aware that three (3) lots of Acuvance2 Safety Catheter, product number 1720-AI, contain devices with missing or incorrectly positioned side port valves. | Process control | Consignees were sent on 11/21/2016 a Smiths Medical "Urgent Medical Device Field Safety Notice: dated November 18, 2016. The letter was addressed to "Clinician who uses the AcuvanceÂż2 Safety Catheter". The letter described the Reason for Recall, Risk to Health, and Instructions to Customers. Requested consignees to complete and return the "Urgent Medical Device Field Safety Notice Response Form" within 10 days of receipt. For questions contact Stericycle via email at the smithsmedicalrecall@stericycle.com. |  | 26750 | 26,750 | Foreign Distribution in the Italy, France, Switzerland, Germany, Sweden | Terminated | August 01, 2017 | 253 | 26750
Z-0821-2017 | 75892 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Tibial Component Metal Cased Small Long; Tibial Component Metal Cased Small Short; Tibial Component Metal Cased Std Long; Tibial Component Metal Cased Std Short |  |  | December 07, 2016 | 2016 | Stanmore Implants Worldwide Ltd. | The recalling firm has received 5 customer inquiries reporting that at least one pouch of the double pouch packaging configuration has been punctured by the tip of the device's stem. | Packaging | Stanmore Implants sent an Urgent Field Safety Notice dated November 25, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately check their internal inventory and quarantine all subject devices, circulate the Field Safety Notice internally to all interested/affected parties, maintain awareness of the notice internally until all required actions have been completed within their facility, inform Stanmore Implants if any of the subject devices have been distributed to other organizations, complete the attached customer response form and list of affected batches form, return the completed form and any affected devices to their S/W Representative. Customers with questions were instructed to contact their local sales representative. |  | 51 | 51 units | Nationwide Distribution to OH, NE, FL, LA, NY, TX, MS, MD, CT, AL, PA | Terminated | July 17, 2017 | 222 | 105
Z-0817-2017 | 75785 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | MEDTRONIC ARRIVE BRAIDED TRANSSEPTAL SHEATH 10F |  |  | October 05, 2016 | 2016 | Oscor, Inc. | Complaint that during insertion of the dilator through the sheath a fragment of inner liner came out of the sheath prior to use. | Nonconforming Material/Component | Oscor sent two(2) Product Recall notifications:-HOSPITALS: Immediately check your inventory to confirm that you do or do not have units from these lot(s) in your possession. Identify and set aside any unit from the identified lot in a manner that ensures the affected product will not be used. Check all storage and usage. Review, complete, sign and return the enclosed Acknowledgement Form attached in the letter, directly to Oscor Inc. at the fax number or e-mail on the form, for immediate replacement (as applicable).Who to call: Please call Oscor Inc. Customer Service Department at 727-937-2511 or email RGA@Oscor.com Hours of operations: Monday to Friday from 8:30AM to 5:30PM Eastern Time.- DISTRIBUTORS: Immediately check your inventory to confirm that you do or do not have units from these lot(s) in your possession. Identify and set aside any unit from the identified lot in a manner that ensures the affected product will not be used. Check all storage and usage. Please pull a list of the hospitals (end user) impacted by the affected lots and retrieve inventory for replacement, if available. Review, complete, sign and return the enclosed Acknowledgement Form (Exhibit A) attached in this letter, directly to Oscor Inc. at the fax number or e-mail on the form, for immediate replacement (as applicable). Who to call: Please call Oscor Inc. Customer Service Department at 727-937-2511 or email RGA@Oscor.com Hours of operation: Monday to Friday from 8:30AM to 5:30PM Eastern Standard Time. |  | 991 | 991 | Worldwide Distribution - US including CA, NJ, MD, TX, FL, NC, MO, and Puerto Rico. Internationally to AUSTRIA, BOLIVIA, CANADA, CHILE, CYPRUS, NETHERLANS, EGYPT, GERMANY, HONG KONG, ISRAEL, ITALY, SWITZERLAND, and VIETNAM | Open, Classified |  |  | 7361
Z-0828-2017 | 75761 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | BIOMET 3i CERTAIN LOW PROFILE 17Âż ABUTMENT 3.4MM(D) X 2MM(H) |  |  | September 14, 2016 | 2016 | Biomet 3i, LLC | A limited number of product pouches may not have been sealed during packaging. | Employee error | BIOMET 3i will conduct this recall with Dentists and Health Care Professionals. Notices will be sent via a carrier with a traceable method of delivery by BIOMET 3i team members." All dentists and health care professionals will be notified via courier: - Dentists and health care professionals will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the product in their control. - Dentists and health care professionals will return on-hand product to BIOMET 3i via a courier agency and ensure all of their products are accounted for using the form provided in the letter. - Dentists and health care professionals will complete and return the Inventory Return Certification Form to the fax number or email address provided.International distributors will be notified via email. Further distribution of the international customer notices will be sent through the local and regional distributors following local language translation. Communications will occur as translations are completed and are anticipated to occur during the week of September 12, 2016.1233. Customers with questions were instructed to call the customer call center at +44 (0) 800 652-1233. |  | 127 | 127 | Worldwide Distribution to Austria, Belgium, Switzerland, Germany, Spain, France, Great Britain, Ireland, Italy, Lithuania, Portugal and Sweden | Open, Classified |  |  | 9363
Z-0826-2017 | 75795 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Affinity Fusion" Oxygenator with Integrated Arterial Filter, |  |  | November 16, 2016 | 2016 | Medtronic Inc | Medtronic is initiating a voluntary product recall for specific lot numbers of the Affinity FusionÂż Oxygenators with BalanceÂż1 Biosurface. These are distributed as stand-alone devices, or as a combination unit with the Affinity Fusion Cardiotomy/Venous Reservoir, or as part of tubing packs. Medtronic has identified an out-of-specification condition exhibiting additional plastic (flash) in the arterial sampling port adjacent to the arterial outlet port of the oxygenator. | Process control | The firm, Medtronic, sent an "Urgent Medical Device Recall" letter dated November 2016 to its Consignees on 11/16/2016. The letter was addressed to Risk Manager or Healthcare Professional. The letter described the problem and the product involved in the recall and actions to be taken. The consignees were instructed to quarantine all unused affected product in your inventory, return the affected product (contact Medtronic Customer Service at 800-848-9300 to initiate a product return) and complete &amp; return the Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality or scan and email to RS.CFQFCA@medtronic.com. For questions contact your Medtronic Representative. If you require assistance in contacting your representative, please contact Medtronics National Answering Service at 800-633-8766. |  | 5609 | 5609 (663 US; 4946 OUS) | Worldwide Distribution-US (nationwide) in states of: AK, AZ, CA, CO, CT, FL, IA, IL, IN, KS, KY, LA, MN, ND, NV, NY, SD, TN, TX, WI, WY and countries of: Canada, Chile, Croatia, Germany, Greece, Hungary, Iran, Italy, Latvia, Lebanon, Poland, Portugal, Romania, Spain, Tanzania, and United Kingdom. | Terminated | August 01, 2017 | 258 | 5609
Z-0827-2017 | 75855 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Zeus-P Lumbar Interbody Device |  |  | November 22, 2016 | 2016 | Amendia, Inc | Labeling error where three boxes in inventory had conflicting patient and primary labels. Evaluation indicated that the primary label and patient label did not match. The primary label matched the label on the inner pouch, but the patient label was incorrect. | Error in labeling | Amendia Inc. sent an Urgent - Medical Device Recall letter dated November 22, 2016, to all affected customers. The letter requested that they examine inventory and quarantine the product. The letter also requested that they contact Amendia to confirm receipt of the letter and to obtain a Return Material Authorization for any product on hand.For further questions please call (404) 610-7215. |  | 25 | 25 units | US Distribution to the states of : GA. MI, OK, SC, TX, VA | Terminated | June 05, 2017 | 195 | 25
Z-0864-2017 | 75692 | No | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | VITROS CRP Slides: 1) VITROSÂż Chemistry Products CRP Slides (250 slides per sales unit), Catalog # 1926740 and 2) , VITROSÂż Chemistry Products CRP Slides (90 slides per sales unit), Catalog # 1926740 |  |  | October 26, 2016 | 2016 | Ortho-Clinical Diagnostics | CRP samples diluted using VITROS Chemistry Products Specialty Diluent, VITROS Chemistry Products FS Diluent Pack 3, or a sample with a low concentration of CRP had positively biased results if the endogenous level of CRP in the diluent is not subtracted from the final result. The IFUs do not provide specific instructions for how to account for endogenous CRP in the diluent when calculating the sample result. Orthos diluents used to dilute samples containing CRP are manufactured using human plasma based fluids that contain endogenous CRP. Without accounting for the CRP in the diluent and removing its contribution from the result, the reported result will include the CRP present in both the sample and diluent, possibly leading to positively biased results. | Process control | Ortho Clinical Diagnostics sent an Important Product Correction Notification dated October 26, 2016., via FedEx overnight courier and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who do NOT use ORTHO PLUS e-Communications and received CRP slides within the last 12 months to inform them of the possibility of positively biased CRP results when using VITROS CRP Slides with diluted samples and to provide them with original instructions for how to perform dilutions when processing CRP samples with concentrations above the measuring range. On 1 November 2016, Ortho Clinical Diagnostics issued customer letter (Ref. CL2016-209) sent via ORTHO PLUS e-Communications to all customers who use ORTHO PLUS e-Communications and received CRP slides within the last 12 months to inform them of the possibility of positively biased CRP results when using VITROS CRP Slides with diluted samples and to provide them with two revised instructions on how to perform dilutions when processing CRP samples with concentrations above the measuring range. Foreign affiliates were informed by email on October 26, 2016, of the issue and instructed to notify their consignees of the issue. This included the original instructions for how to perform dilutions when processing CRP samples with concentrations above the measuring range. On 1 November 2016, two revised Instructions for CRP Dilutions enclosures (Ref. CL2016-209) were sent via FedEx overnight courier and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received CL2016-209 via FedEx overnight courier or US Postal Service Priority Mail that provided additional detailed instructions on how to perform a dilution for samples exceeding the measuring range. If you have further questions, please contact our Ortho Care Technical Solutions Center at 1-800-421-3311. |  | 124063 | 124,063 units | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico, and to the countries of : Bermuda, Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain and Poland. | Open, Classified |  |  | 124063
Z-0865-2017 | 75891 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery Systems, models: G3, G4, VSI. |  |  | December 06, 2016 | 2016 | Accuray Incorporated | Accuray has become aware of a potential safety issue involving the possibility of electric shock duringmaintenance activities from the lon Pump HV Power Supply at the back of the modulator cabinet. This issue is limited to certain G3, G4, and VSI models of the CyberKnifeÂż Robotic Radiosurgery System. | Device Design | Accuray Incorporated sent an Urgent Safety Advisory Notification letter to all affected customers on December 1, 2016 by traceable method.. Letters identified the product, the problem, and the action to be taken by the customer. provided users with affected product and described the issue. The final resolution will result in a modulator cabinet inspection and installation of the Plexiglas cover - as necessary. Customers will be contacted by Accuray to schedule the inspection. Prior to this modulator cabinet inspection, inform all applicable personnel of the potential danger and need to closely follow electrical safety procedures. Customers with questions were instructed to contact Accuray Customer Support at 1-877-668-8667 or visit the Accuray Tehnical Solution Center at http://www.accuray.com/services-support or call:North America +1-877-666-8667, Europe +800-4141-9595, Hong Kong +852-2247-8688, Japan +81-3-6265-1530. Customers were also asked to complete the attached acknowledgement and return to Accuracy.For questions regarding this recall call 608-824-3362. |  | 115 | 115 affected devices. | Worldwide Distribution - US including CA. MN, OK. FL. MO, TN, KS, TX, AZ, MS, IA, MA. NC, NY. IN, NJ, VA. LA, WA, WI, PA. CT, G,. MT, MD, LA., WV, OH, and DC.Foreign distribution to Taiwan, Japan, Korea, Turkey. Malaysia, Vietnam, Greece, China, Spain, France, Italy, Pakistan, and Georgia, | Terminated | June 09, 2017 | 185 | 115
Z-0877-2017 | 75886 | No | MKQ |  |  | Pathology | PMA | BD SurePathTM Preservative Collection Vials, catalog numbers 490505, 490527, 491095, 491096, and 491097. |  |  | February 12, 2016 | 2016 | Becton Dickinson &amp; Co. | BD has confirmed that a portion of BD SurePath" collection vials associated with specific production lots may contain caps that are cracked which can lead to leaking of the preservative fluid. | Nonconforming Material/Component | BD Life Sciences sent the initial recall notification on February 12, 2016. The distributors were provided with a written communication via email and customers were provided with a written communication via UPS.For further questions, please call (410) 316-4000. |  |  | 16,682.396 | Worldwide Distribution - USA (nationwide) Distribution including BB(Barbados) and PR (Puerto Rico) and to the states of : AK, AL, AR, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY., and to the countries of: China, Europe, Hong Kong, India, Korea, Mexico, New Zealand, Panama, Singapore, and Taiwan.480 US customers and 3 US distributors. Product was also shipped to 10 ex-US regions. 15 US government sites received affected product | Completed
Z-0878-2017 | 75909 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge OrthoCase software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | October 08, 2015 | 2015 | Merge Healthcare, Inc. | Measurements in the software are changing after saving a plan with a measurement result, which may result in potential patient injury or delay in diagnosis or treatment. | Software design | Merge issued a letter dated March 28, 2016, via email on April 4, 2016, or via certified mail if they did not have an email address for the customer. The letter notifies the customer of the issue, informs them of a workaround and that a fix is not yet available but they will be notified once one is available. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they are interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 2, 2016.For further questions, please call (877) 741-5369. |  | 30 | 30 sites potentially have the effected versions | Distribution was made to medical facilities in CA, CO, CT, MA, MD, MN, MT, NE, NV, NY NJ, PA, SC, WA, and WI. Foreign distribution was made to Canada, as well as other countries. | Open, Classified |  |  | 30
Z-0867-2017 | 75828 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes VA Implant Rack for the Compact Distal Radius System |  |  | November 29, 2016 | 2016 | Synthes (USA) Products LLC | The screw rack plate was assembled incorrectly which could cause the screws in the screw rack to sit low and touch the surface beneath the screw rack. | Process control | DePuy Synthes mailed an Urgent Notice Medical Device Recall letter to all affected customers informing them of the potential impacts, actions to be taken and instructions for the product return. Customers were instructed to immediately review their inventory to identify and quarantine all affected products, complete the attached Verification Section and send a copy to Depuy Synthes by scan/email: Synthes5748@stericycle.com or fax to 1-888-943-5169. If any of the affected product has been forwarded to another facility, contact the facility to arrange return. Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales Consultant. |  | 31 | 31 units | Nationwide Distribution | Terminated | August 03, 2017 | 247 | 31
Z-0866-2017 | 75716 | No | HIF | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Insufflator, laparoscopic |  |  | November 21, 2016 | 2016 | Northgate Technologies, Inc. | Device could cause a patient overpressure situation without any visual or auditory indication or warning | Software design | Northgate Technologies sent an Urgent Medical Device Recall Notification letter dated November 17, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to examine their records and ascertain the location of affected devices, cease distribution or use and quarantine those products appropriately. Ship all affected products in their possession to Northgate Technologies, Inc., 1591 Scottsdale Court, Elgin, IL 60123. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 847-608-9405. If product was further distributed, those customers should be notified immediately. Customers with questions should call 224-856-2253. For questions regarding this recall call 800-348-0424. |  | 49 | 49 devices | Worldwide Distribution - US including FL and MA. Internationally to Canada | Open, Classified |  |  | 49
Z-0873-2017 | 75819 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | 8450-009-030 3.0MM Precision Round, 13CM, single use, Sterile |  |  | November 30, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling specific lots of sterile MIS Toolsteel Burs because the burs may exhibit signs of corrosion. | Storage | The firm, Stryker Instruments, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated November 2016 to its Customers Overnight service, FedEx on November 30, 2016. The letter described the product, problems and actions to be taken. Stryker Sales Representatives and Stryker International will be notified via email. Customers are instructed to do the following: 1. Immediately review the Recall Notification.2. Check all stock areas and/or operating room storage to determine how many affected Burs from the affected lots are at your facility. Quarantine and discontinue use of the recalled Burs.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of the notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if you dont have anyaffected product on hand. Fax the completed BRF to Stryker Instruments Regulatory Department at 866-521-2762, or scan and email a copy to StrykerInstrumentsRecalls@stryker.com.Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification.4. If you have further distributed this product, please forward this letter and the attached BRF to all affected locations. Please indicate each location on the BRF.5. Fax the completed BRF to Stryker Instruments Regulatory Department at 866-521-2762, or scan and email a copy to StrykerInstrumentsRecalls@stryker.com.6. If the BRF for your facility indicates that recalled product is currently on hand, a FedEx label will be emailed to you. This shipping label should be used to return recalled product. Upon receipt of the recalled product, a credit will beissued to your account.Recalled products will be destroyed, packaging will be defaced, opened and discarded. Product will be discarded.For questions regarding this recall, please contact Stryker Instruments: StrykerInstrumentsRecalls@stryker.com, or |  | 1257 | 1,257 in total | Worldwide Distribution - US (nationwide) including states of: AK, AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, NC, NE, NH, NJ, NM, NY, OH, OK, PA, SC, TN, TX, VA, WA, WI and countries of: Canada, Japan, Venlo (Netherlands), France, and UK (United Kingdom). | Terminated | June 01, 2017 | 183 | 1257
Z-0897-2017 | 75862 | No | PJI |  | Immunology | Immunology | 510(k) | EnLite Neonatal TREC Kit;an in vitro diagnostic device intended for the semi-quantitative determination of TREC (T-cell receptor excision circle) DNA in blood specimens dried on filter paper. The test is for use on the VICTOR" EnLite instrument. The test is indicated for use as an aid in screening newborns for severe combined immunodeficiency disorder (SCID). Product Number: 3401-001U |  |  | December 05, 2016 | 2016 | PerkinElmer Health Sciences, Inc. | Unique Device Identifier (UDI) is missing from outer kit box label. | Employee error | Customers were notified of the recall via email on December 5, 2016. Technical Support will ensure the customers receive the Recall Letter and understand the recommended actions. Customers were asked to retain this Medical Device Correction in your laboratory documentation until the affected lot inventory is either depleted or reaches expiration and to complete the enclosed response form and return it by fax to number 1-330-825-8520 (United States) /+358 2 2678 357 (outside United States) or as scanned by e-mail to TurkuQMresponse@perkinelmer.com as soon as possible, but not later than December 29, 2016. |  | 73 | 73 kits | MA, AK | Open, Classified |  |  | 73
Z-0881-2017 | 75858 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | K-Wire 1.4mm x 12in Non-Threaded, Part Number 9080-12U; K-Wire Threaded 1.4mm x 18, Part Number 9080-18T; K-Wire Non-Threaded 1.4mm x 18, Part Number 9080-18U; 1.4mm x 24 Threaded Guide Wire, Part Number 9080-24T; K-Wire Non-Threaded 24, Part Number 9080-24U; Blunt K-Wire, 18 inches, Threaded, Part Number 9080B-18T; Blunt K-Wire, 18 inches, Unthreaded, Part Number 9080B-18U; Blunt K-Wire, 24 inches, Unthreaded, Part Number 9080B-24U; 1.4mm x 18 Blunt Threaded Nitinol Guide Wire, Part Number 9080B-N-18T; 1.4mm x 18" Blunt Non-Threaded Nitinol Guide Wire, Part Number 9080B-N-18U; Lodestar K-wire Non-Threaded, 18, Part Number 9080L-18U; 1.4mm x 18 Threaded Nitinol Guide Wire, Part Number 9080-N-18T; 1.4mm x 18 Non-Threaded Nitinol Guide Wire, Part Number 9080-N-18U |  |  | November 22, 2016 | 2016 | Amendia, Inc | The rods are packaged in packages with unacceptable seals. In some cases, the integrity of the seal is compromised by channels or voids. | Employee error | Amendia sent an Urgent Medical Device Recall letter dated November 22, 2016, to all affected customers. The letter requested that customers examine their inventory and quarantine the product. The letter also requested that they contact Amendia to confirm receipt of the letter and to obtain a Return Material Authorization for any product on hand. Customers with questions were instructed to call the following numbers:770-575-5224770-575-5236770-575-5369For questions regarding this recall call 404-610-7215. |  | 5833 | 5833 units | Nationwide Distribution to AL, AZ, AR, CA, CO, FL, GA, IA, KS, KY, LA, MA, MI, MN, MO, NV, NJ, NC, SC, TN, TX, VA, WA, WV, WI and Puerto Rico | Open, Classified |  |  | 7151
Z-0891-2017 | 75856 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Optimus Fixed Awl |  |  | November 22, 2016 | 2016 | Amendia, Inc | Complaints that the tip of the awl broke after impaction during surgery. | Device Design | Amendia Inc. sent an Urgent : Medical Device Recall letter dated November 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter requested that they examine inventory and quarantine the product. The letter also requested that they contact Amendia to confirm receipt of the letter and to obtain a Return Material Authorization for any product on hand.For further questions, please call (404) 610-7215. |  | 9 | 9 devices | US Distribution to the states of : FL, GA. IL, MO, TX | Open, Classified |  |  | 9
Z-0879-2017 | 75758 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | OrthoPediatrics PediNail Intramedullary Nailing System |  |  | November 18, 2016 | 2016 | OrthoPediatrics Corp | Labeled as a 48mm Cortical Screw but measuring at 50mm | Labeling mix-ups | OrthoPediatrics sent an URGENT MEDICAL DEVICE RECALL notification on November 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. . It is important that you carefully follow these recall instructions:" Please immediately take possession of any of these screws your customer may have or that you have, and quarantine them. Contact logistics@orthopediatrics.com to receive a Return Authorization Number. It is very important that we contain these devices rapidly and maintain them in quarantine. Replacement devices can be ordered from Customer Service." The FedEx package you receive will contain a pre-addressed FedEx return label for the return of the recalled screw(s)." Included in this FedEx package is a Mandatory Reply Form for you to indicate the total number of affected screw(s) in your sets.We will be providing any additional follow up instructions as needed. If you have any questions call (574) 267-0865.For further questions regarding this recall , please call (574) 268-6379 |  | 22 | 22 units | Worldwide Distribution - US Distribution to the states of : TX, MO, FL, GA, NC, AZ, NY., and to the countries of France and Ireland. | Terminated | July 27, 2017 | 251 | 22
Z-0887-2017 | 75869 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Black Widow 6.6mm x 20mm Screw, Part Number TBS020; Black Widow 6.6mm x 25mm Screw, Part Number TBS025; Black Widow 6.6mm x 30mm Screw, Part Number TBS030 |  |  | November 22, 2016 | 2016 | Amendia, Inc | There is a problem with the mating feature between the plate and screw which could cause the plate opening not to interfere enough with the plate and head of the screw. | Process change control | Amendia sent an Urgent Medical Device Recall letter dated November 22, 2016, to all affected consignees.. The letter requested that they examine inventory and quarantine the product. The letter also requested that they contact Amendia to confirm receipt of the letter and to obtain a Return Material Authorization for any product on hand. Customers with questions should call 1-855-755-3329.For questions regarding this recall call 404-610-7215. |  | 96 | 96 units | Nationwide Distribution to IN, LA, and TX | Terminated | July 31, 2017 | 251 | 144
Z-0899-2017 | 75836 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AQURE System |  |  | November 23, 2016 | 2016 | Radiometer America Inc | Design error when displaying additional information in the patient view window; error may result in misreading a parameter and its value. | Software design | Radiomet4er sent a letter dated November 30, 2016, to all affected customers via email with a request for customers to complete a return form. Customers/Users were advised to not use the "more" button to expand patient results when viewing them in the AQURE patient view. Customers/Users Radiometer Representative will contact them to schedule a visit of a remote session. During the visit or remote session the representative will run a script on the AQURE system, which serve to remove the misalighnment possibility.It is noted in the letter to customers that expanding the results by using the More button will not affect the stored results.If customers are not the end-user of the affected product ensure that the notification letter is distributed to the final end-user.If any questions regarding the notification letter contact Radiometer America Technical Support at 1-800-736-0600 opt.1.For questions regarding this recall call 440-871-8900. |  | 7 | 7 total (2 in US) | Worldwide Distribution - US including MI and NY and Internationally to England, Switzerland, and United Arab Emirates. | Open, Classified |  |  | 7
Z-0895-2017 | 75746 | No | FCH | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BREAST PACK, Surgical Instrument Tray, REF/Catalog No. 5406, STERILE, Rx Only -- busse Hospital Disposables Hauppauge, NY --- Device Listing # E193298. |  |  | November 15, 2016 | 2016 | Robert Busse &amp; Co. Inc. d.b.a. Busse Hospital Disposables | Covidien issued a labeling update for the Covidien Devon Light Gloves. Busse had incorporated the affected gloves in several kits and distributed them prior to receipt of the recall notification. | Other | Busse Hospital Disposables sent a Medical Device Correction letter (dated November 11, 2016) and response form to the three customers via FedEx on 11/15/2016 to inform them of the recall of the kits containing the affected gloves. This was done without notifying NYK-DO. |  | 35 | US: 35 cases | PA, FL, VA | Open, Classified |  |  | 131
Z-0900-2017 | 75740 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Equinoxe Reverse Shoulder Fixed Angle Torque Defining Screw Kit |  |  | October 25, 2016 | 2016 | Exactech, Inc. | Disengaged screw head could potentially not be retained in the Torque Defining Screwdriver. | Component design/selection | Exactech sent an Advisory Notice on October 25, 2016, to Exactech field agents in the U.S. Agents outside of the U.S. were notified on November 2, 2016. Field agents were notified of the recall in the Advisory Notice and asked to notify their accounts that may have the recalled product in stock. Agents were required to complete and return the Field Advisory Notice Response Form within 5 working days. For further questions, please call (352) 377-1140. |  | 2278 | 2,278 units | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and to the states of : AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, MD, ME, MN, MO, NC, NJ, NV, NY, OH, OK, OR, PA, SC, TN, VA, WA and WI., and to the countries of :Austria, Australia, Canada, Germany, France, Spain, India, Switzerland, Italy, South Korea, United Kingdom, Brazil, Luxembourg, Greece, Guatemala, Netherlands, South Africa, and Sweden | Open, Classified |  |  | 2278
Z-0888-2017 | 75649 | sling, arm | ILI | Arm sling. | Physical Medicine | Physical Medicine | 510(K) Exempt | Medi Choice Arm Slings |  |  | October 31, 2016 | 2016 | Owens &amp; Minor Distribution, Inc. | Owens &amp; Minor is the initial importer and private label distributor of Medi Choice Arm Slings (Product). The Product is manufactured in China by Tops Medical Articles Nantong Inc.Owens &amp; Minor first learned of the product issue that prompted the product removal action via customer complaints received in September 2016. Two customer complaints received on September 1"1, 2016 reported surface contamination on the arm slings. Two additional complaints reported on September 19, 2016 and October 7, 2016 reported the same issue.Upon receipt of the complaints on September 1, 2016, Owens &amp; Minor immediately initiated a CAPA investigation for root cause analysis and determination of corrective/preventive actions. A quality inspection of inventory confirmed that the surface contamination was confined to lots manufactured in May, June, and July of 2016. The surface contamination for these lots was found to be frequent. Subsequent August product lots were inspected and no surface contamination was found. Limited inventory from lots prior to May 2016 were available and inspected as well. No surface contamination was found on these prior lots.Owens &amp; Minor has quarantined the affected product lots. The product is on hold at all Owens &amp; Minor Distribution Centers (DC). Samples of contaminated product have been sent to the manufacturer for evaluation and to a certified testing lab for identification of contaminate. These results are pending at this time.Based on the frequency of the product defect as well as evaluation of potential adverse health event risk, Owens &amp; Minor initiated a voluntary recall for the above stated affected product lots. | Under Investigation by firm | The firm, Owens&Minor;, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 11/9/2016 to its consignees/customers including a Recall Response Form. The letter describe the product, problem and actions to be taken. The customers were instructed to examine your inventory and identify all affected product; immediately stop using all affected product; quarantine all affected product per your facility's Standard Operating Procedures; if you have further distributed this product, identify patients/customers who received product and notify them at once; and complete and return Recall Response Form via email: www.owens-minor.com, fax to 804-723-7100, or mail to the O&M; Distribution Center (DC) Attn: Recall Officer (even if you do not have any affected product). Note: upon receipt of the Recall Response Form, the DC will contact you with the appropriate RGA and instructions for product return.If you have any questions concerning credit and/or replacement of affected product or this recall, please call 804-723-7000 or contact your O&M; Sales Representative or Customer Service Representative at the local O&M; Distribution Center. |  | 568 | 568 cases | US Nationwide Distribution. | Open, Classified |  |  | 568
Z-0890-2017 | 75301 | No | FEI | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Best Practices kits containing certain production lots of Clearify Visualization System |  |  | October 04, 2016 | 2016 | Medtronic | Product sterility is compromised due to breach of the sterile barrier | Employee error | Covidien issued an Urgent Medical Device Recall letter on October 4, 2016 and a corrected letter on October 10, 2016 to US Customers via Federal Express. The letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Disposition of product within Medtronic control and customer returned devices will be reworked.Questions or concerns, contact your Medtronic representative or customer service at 800-882-5878. |  | 230 | 230 systems | Worldwide Distribution - US (Nationwide)Foreign:AustriaBelgiumBulgariaBrazilChileCyprusDenmarkEgyptFinlandFranceFranceGermanyGreeceHungaryIranIrelandIsraelItalyJapanKenyaKuwaitLatviaLuxembourgMexicoNamibiaNetherlandsNigeriaNorwayPolandPortugalRomaniaSaudi ArabiaSerbiaSlovakiaSouth AfricaSpainSwedenSwitzerlandTurkeyUnited Arab EmiratesUnited Kingdom | Open, Classified |  |  | 1293399
Z-0885-2017 | 75798 | No | MTF |  |  | Immunology | PMA | IMMULITE 2000/IMMULITE 2000 XPi Systems Prostate-Specific Antigen (PSA) Assay, REF/Catalog Number L2KPTS2, SMN 10706281 and REF/Catalog Number L2KPTS6, SMN 10706282 (US); REF/Catalog Number L2KPS2, SMN 10380986 and REF/Catalog Number L2KPS6, SMN 10380996 (OUS); IVD --- Origin: UK Siemens Healthcare Diagnostics Products Ltd. Llanberis, Gwynedd, LL554EL UK |  |  | November 18, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics confirmed that the IMMULITE 2000/ IMMULITE 2000 XPi Systems PSA assays are not meeting the current High-dose Hook Effect expectation in the Instructions for Use. | Under Investigation by firm | The firm, Siemens, sent an "Urgent Medical Device Correction Letter" (IMC 17-07.A.US, dated 11/2016) was sent via FedEx on 11/18/2016 to the US customers and an "Urgent Field Safety Notice" dated 11/2016 to customers outside the US (distribution is determined at the country level). The letter/notice described the product, problem and actions to be taken. Additionally, Field service personnel were sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. Customers were instructed to review letter with your medical doctor; review products and follow instructions; complete and return the attached Field Correction Effectiveness Check Form within 30 days by fax to Customer Care Center at (312) 275-7795; retain this letter with your laboratory records and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 6697 | 6697 kits (US: 1646 kits; OUS: 5081 kits) | Worldwide Distribution-US (nationwide)and the following countries: Afghanistan, Albania, Algeria, Argentina, Aruba, Austria, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg, Brazil, British Guyana, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Dominican Rep., Ecuador, Egypt, El Salvador, Estonia, Finland, France, Germany, Greece, Great Britain, Grenada, Guatemala, Honduras, Hungary, Iceland, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Jordan, Kazakhstan, Kosovo, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Mexico, Netherlands, Nicaragua, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Spain, South Africa, South Korea, Suriname, Switzerland, Syria, Taiwan, Tadjikistan, Thailand, Trinidad, Tobago, Tunisia, Turkey, Turkmenistan, United Arab Emirates, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 11402
Z-0880-2017 | 75792 | No | NDE | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | GAM Kirschner Wire, Sterile R, T2 K-Wire, Sterile R and T2 K-Wire Recon, Stryker Trauma GmbH, Prof.-Kuntscher-Strasse 1-5, Schonkirchen, Germany 24232; T2 K-Wire, Sterile R |  |  | November 09, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | A review of packaging revealed the seal integrity of the pouch may be compromised. More specifically, there is a potential that the sterile pouch is not sealed at one end due to a manufacturing error. | Packaging process control | Stryker sent an Urgent Medical Device Removal letter dated November 9, 2016, to all affected customers via UPS (with return receipt). Stericycle will be handling the returns. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inform users of the Medical Device Recall and forward the notice to all those individuals who need to be aware within their organization. Return all affected products available at their location to:Stryker C/O Stericycle2670 Executive Dr., Suite AIndianapolis, IN 46241Attn: RA 2016-169 - Event XXXCustomers were also instructed to complete and sign the enclosed Business Reply Form and return to Stericycle.Customers with questions should call +1-201-851-6683.For questions regarding this recall call 201-831-5000. |  | 2084 | 2,084 units | Worldwide Distribution - US (nationwide) Internationally to AU, CA, CL, CH, ES, FR, GB, SE, JP, FR, ZA, and NL. | Open, Classified |  |  | 2084
Z-0892-2017 | 75780 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3, PowerCross .018 OTW PTA Dilatation Catheter, Sterile EO. |  |  | November 10, 2016 | 2016 | Medtronic Inc | Medtronic identified an issue with specific models and lots of the PowerCross .018 Over the Wire (OTW) Percutaneous Transluminal Angioplasty (PTA) Dilatation Catheter. During an unrelated manufacturing inspection of scrapped products, an out of specification condition was observed. The catheter outer shafts were pinched and deformed at the proximal balloon bond. Bench testing of affected devices showed that this resulted in either the inability to inflate the device, a slow deflation, or the inability to deflate the device post expansion. | Process control | Medtronic verbally communicated about the issue on November 20. 2016, to affected customers. Additionally, consignees were hand delivered by Field Sales Representatives a Medtronic "Urgent Medical Device Recall" letter dated November 2016. The letter was addressed to Risk Manger or Health Care Professional. The letter described the problem and the product involved in the recall. Advised consignees to immediately identify, quarantine the affected product, and return it to Medtronic. Requested consignees to complete and email the Customer Confirmation Certificate to RS.CFQFCA@medtronic.com or fax it to Medtronic at 651-367-0612, attention: Customer Focused Quality.For questions, contact your Medtronic representative. If you require assistance in contacting your representative, please contact Medtronics National Answering Service at 800-633-8766. |  | 315 | 315 | NY, NC, WA | Terminated | August 02, 2017 | 265 | 315
Z-0893-2017 | 75737 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gammex Non-Latex Sensitive Neoprene Surgical Gloves |  |  | October 24, 2016 | 2016 | Ansell Healthcare Products LLC | Ansell Healthcare Products is initiating a voluntary recall involving 15 cases of Gammex Non Latex Sensitive Surgical Gloves because the packaged contents incorrectly included a natural rubber Latex surgical glove. | Packaging change control | Ansell Healthcare Products, LLC. sent an Urgent: Voluntary Surgical Glove Recall letter dated October 28, 2016, to all affected customers. The letter identified the problem the product and the action the customer needs to take.Please check your product stock at the warehouse and shelf level in order to advise us directly if any gloves having the recalled lot numbers are found. This information, along with your tracking number, must be communicated to our Customer Service Department by calling 1-800-952-9916, or by returning the attached Response Form to Ansell by fax at1-800-722-8155, or via email customerserviceus@ansell.com . Returned product will either be replaced or credited without delay. Your Ansell representative will be available to follow up with you regarding any concerns you may have. |  | 15 | 15 Cartons | US Distribution to the states of : PA, FL, NJ, IL and MA. | Open, Classified |  |  | 15
Z-0902-2017 | 75817 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Specimen Gate Laboratory 5002-0180, Specimen Gate eReports 5002-0370, Software Version 1.0 |  |  | December 01, 2016 | 2016 | PerkinElmer Health Sciences, Inc. | Issue may cause eReports to consume an excessive amount of disk space causing the software to stop operating. | Software design | PerkinElmer Health Sciences issued a Medical Device Correction letter dated December 1, 2016, to all affected customers email facilitated through our PerkinElmer Software Services group located in Akron Ohio. Software Services will be asked to make sure the customers receive the Recall Letter and implement the recommended corrective measures.Actions to be takenA field correction of the issue is recommended. PerkinElmer will correct the issue with an update installed through remote connection. You will be contacted by your PerkinElmer representative to make the necessary arrangements for implementing the update.In the interim period, please do not attempt to login multiple times in quick succession. Allow the software time to respond to first login command.Please contact your local PerkinElmer representative or Specimen Gate Support for further information.Other InformationPlease inform those affected in your organization accordingly.To comply with regulatory requirements we request that you complete the enclosed response form and return it by fax to number 1-330-825-8520 (United States) /+358 2 2678 357 (outside United States) or as scanned by e-mail to TurkuQMresponse@perkinelmer.com as soon as possible, but not later than December 22, 2016.For questions regarding this recall call 781-663-6900. |  | 21 | 21 | Worldwide Distribution - US including CO, FL, GA, IL, MA, MI, NV, NJ, OH, SC, AZ, KY and Internationally to Italy, Ireland, UK, Spain, and Peru | Terminated | May 01, 2017 | 151 | 21
Z-0904-2017 | 75851 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Innova 2000 |  |  | November 21, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare is conducting a recall due to the potential fall of aged CRT monitors - Prestige 2, Prestilix, Prestige VH, Prestige SI, RFX/SFX, Legacy, Precision 500D, ADVANTX and Innova 2000 systems. | Nonconforming Material/Component | Consignees were sent on 11/21/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# FMI 10913 dated November 21, 2016. The letter was addressed to Director of Radiology, Risk Manager/Hospital Administrator, and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 1598 | Total for all products 1,598 (USA 824, OUS 774) | US: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT,NC, NE NH, NJ, NM, NY, NV, OH, OK, PA, RI, SC, SD, TN, TX,, VA, WA, WO, WY, WV. OUS: AlgeriaArmeniaAustraliaAustriaBrazilCanadaChileChinaDenmarkEgyptFranceGermanyGuatemalaHondurasHong KongHungaryIcelandIndiaIrelandItalyJapanKoreaLibyaMexicoNetherlandsNetherlandsPeruPolandPortugalRussiaSaudi ArabiaSerbiaSierra LeoneSlovakiaSpainTaiwanThailandTurkeyUnited Arab EmiratesUnited KingdomVenezuelaVietnam | Open, Classified |  |  | 1598
Z-0876-2017 | 75383 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 128 Multi-slice CT Scanner System with |  |  | August 31, 2016 | 2016 | Neusoft Medical Systems Co., Ltd. | The Firm discovered during contrast agent tracking scan, when the expected concentration of contrast agent is reached, later planned scan may be interrupted. The system must be powered down and restarted to clear the interruption error before scanning can resume. | Software design | Neusoft Medical Systems Co. Ltd. planned action to bring into compliance.1.Affected Customers will be notified by letter to provide them a description of the affected device, instructions for using the device prior to correction, hazards associated with the defect, and a statement of corrective actions to be taken at no cost to the owner.2.Field Service Engineers will visit owner sites to install a software update which resolves the defect.3.Corrections will be made at no cost to the owner4.Corrections will be completed by May 11, 2017. CDRH approves the CAP subject to the following conditions:.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.Further questions please call (281) 453-1205. |  | 23 | 23 units | US Distribution | Open, Classified |  |  | 23
Z-0901-2017 | 74834 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | NIM EMG Endotracheal Tube, 6.0 mm I.D. x 8.8 mm O.D. (27 FR), Reinforced, Sterile, Rx only. Individually packaged. |  |  | June 28, 2016 | 2016 | Medtronic Xomed, Inc. | Wire in tubing can become exposed, posing potential harm to the intubated patient. | Use error | Medtronic sent an Urgent Product Safety Advisory letter dated June 28, 2016, to all affected customers informing them of the correction and providing instruction to avoid causing injury when using the EMG Endotracheal Tube. Recommended Actions included: inspecting the product thoroughly before use, testing the cuff by inflating it then deflating it before intubation, using a non-malleable stylet that closely matches the natural curve of the tube, not bending and flexing the tube excessively before use, and lubricating the cuff with a non-paralyzing, aqueous lubricant and a lubricated stylet. If customers found an EMG tube that appeared to have an electrode wire protruding from the distal end, the tube is not to be used and practitioners are urged to call Medtronic Product Quality Experience Department at 1-866-849-4003. For further questions , please call (904) 296-9600. |  | 3752 | 3752 | Worldwide Distribution - Nationwide Distribution including US Virgin Island and Puerto Rico , and to the countries of :Japan, Korea, China, Canada, Dominican Republic, Colombia, Chile, Costa Rica, Bahamas, Peru, Panama, Mexico, Argentina, Ecuador, Brazil, Trinidad and Tobago, Hong Kong, Taiwan, United Kingdom, Netherlands, Sri Lanka, Austria, Italy, Germany, United Arab Emirates, Saudi Arabia, Kuwait, France, Belgium, Oman, Ireland, Azerbaijan, India, Spain, Portugal, Luxembourg, Switzerland, Russian Federation, Austria, Greece, Cyprus, Portugal, India, Algeria, Iceland, Iran, Denmark, Hungary, Czech Republic, Ukraine, Turkey, Croatia, Israel, Norway, Oman, Sweden, Libya, Finland, Spain, Portugal, Algeria, Morocco, Poland, Canada, Malaysia, Singapore, Australia, and New Zealand | Open, Classified |  |  | 3752
Z-0918-2017 | 75834 | No | IZO | Diagnostic x-ray high voltage generator. | Radiology | Radiology | 510(K) Exempt | Shimadzu X-ray High Voltage Generator Model No. UD150B-40 on SONIALVISION safire II/17 systems with 0.7/1.2JG326D X-Ray Tube |  |  | November 20, 2016 | 2016 | Shimadzu Medical Systems | Shimadzu Medical Systems is recalling the Shimadzu X-ray High Voltage Generator due to the possibility of oil leaking from the x-ray tube housing under certain conditions. | Under Investigation by firm | The firm, Shimadzu Medical Systems USA, sent an "Urgent Medical Device Recall notice" and Field Safety Notice dated 11/20/16 to customers to inform them that Shimadzu Corporation has identified a potential issue in the Shimadzu X-ray High Voltage Generator UD150B-40. Customers are informed of how to identify the affected products, corrective action plan by Shimadzu, and the actions to be taken by customers. Customers that have not received a call or visit from Shimadzu are instructed to call Shimadzu Medical Systems National Technical Support at (800) 228-1429.Shimadzu sent an updated "URGENT Voluntary Medical Device Recall Notice" letter dated 11/30/16 to its customers on 11/30/16 to inform them that Shimadzu Corp. has identified a potential issue in the Shimadzu X-ray High Voltage Generator UD150B-40 on SONIALVISION safire II/17 systems with 0.7/1.2JG326D X-Ray Tube. The letter informs the customers that when operated for prolonged periods of time at low rate serial exposure, the heat load capacity of the X-ray anode disc (target) may become damaged and leak insulating oil. Customers with questions are instructed to contact Shimadzu National Technical Support at (800) 228-1429. |  | 13 | 13 units | US Distribution to states of: AZ, WA, CA, LA, FL, SC, and NJ. | Open, Classified |  |  | 13
Z-0903-2017 | 75476 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Immunodiagnostic Products Intact PTH Reagent Pack: 1) UPN 10758750006287 / Product Code 6802892 - Reagent Pack, 2) UPN 10758750006250 / Product Code 6802893 - Calibrators |  |  | October 05, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | The device defect is a positive bias of up to 20% relative to the currently marketed Roche device. In addition, the firm states that during the investigation they discovered that the cleared reference interval (RI) was not set optimally at product launch in 2010 and that the upper limit in the IFU of 53.5pg/mL is too low and should be closer to 80 pg/mL. They are in the process of submitting the data to support the revised RI in a new 510K submission and will be changing the IFU once the 510K is cleared. | Under Investigation by firm | Ortho Clinical sent an Urgent Product Correction Notification letter dated October 5, 2016, to all affected customers via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received VITROS Immunodiagnostic Products iPTH to inform them of the issue and the required actions. Foreign affiliates were informed by email on October 5, 2016, of the issue and instructed to notify their customers of the issue and the required actions. Required Actions - Until further notice, be aware of positively biased results when using the device. Discuss any concerns regarding previously reported results with your Laboratory Medical Director to determine the appropriate course of action. Post the notification by your VITROS System or with your user documentation. In accordance with regulatory requirements, complete and return the Confirmation of Receipt form via email to: RA-OCDUS-CONFIRMAD@ORTHOCLINICALDIAGNOSTICS.COM. Customers with questions can call the Ortho Care Technical Solutions Center at 1-800-421-3311.For further questions regarding this recall, please call (585) 453-3452. |  | 7027 | 7,027 Total units | Worldwide Distribution - Nationwide Distribution : and to the countries of : Bermuda, Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain and Poland. | Open, Classified |  |  | 7027
Z-0917-2017 | 75833 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Shimadzu X-ray TV System, Catalog No. SONIALVISION G4 |  |  | November 20, 2016 | 2016 | Shimadzu Medical Systems | Shimadzu Medical Systems is recalling Shimadzu X-ray TV System SONIALVISION due to possibility of oil leaking from the x-ray tube housing under certain conditions. | Under Investigation by firm | An urgent medical device recall notice and a field safety notice dated 11/20/16 was sent to customers to inform them that Shimadzu Corporation has identified a potential issue in the Shimadzu X-ray TV System SONIALVISION G4. Shimadzu informed their customers that they have received a report that oil had leaked from X-ray tube device of SONIALVISION G4 during examination. Customers are informed of the description of the issue, the actions planned by Shimadzu to correct the issue, and the actions that should be taken by the customer. Customers are instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at (800) 228-1429. |  | 39 | 39 units | US: TX, NY, SD, PA, MD, OH, IA, IN, MT, IL, FL, SC, OR, CA, AR, GA, CT, NV, PA, NE, WA | Open, Classified |  |  | 39
Z-0928-2017 | 75824 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Arthex S3 Arm Sleeve Connector; Catalog Number AR-1650-02; manufactured and distributed between March and November 2016. |  |  | November 23, 2016 | 2016 | Allen Medical Systems, Inc. | During use, the unit could separate un-expectantly. This failure would allow the patient's arm to fall out of suspension during a procedure, possibly causing moderate injury to the patient. This potential defect is not apparent to the user. | Under Investigation by firm | The firm's customer was notified of the recall via letter on November 23, 2016. They were asked to take the following actions: 1. Immediately remove from service your Arthrex S3 Arm Sleeve Connector (AR-1650-02). All serial numbers are impacted. 2. Please follow the included instructions for the return of all S3 Arm Sleeve Connectors (AR-1650-02). Allen Medical will inspect, repair or replace all units as necessary at no cost to you. Please take steps to retrieve the units from your customers and return them to Allen Medical via the standard Return Material Authorization (RMA) process. In order to minimize this inconvenience, we will repair the Arthrex S3 Arm Sleeve Connector and return the device to you as soon as possible. For further questions, please call (978) 266-4280. |  | 420 | 420 devices | US Distribution to the state of : FL | Open, Classified |  |  | 420
Z-0930-2017 | 75591 | No | LBR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sterling Diagnostics, Inc., HDL Cholesterol (PEF Method) for IN VITRO DIAGNOSTIC USE in the quantitative determination of analytes. These are CONSUMABLES - a customer buys another when kit is consumed. |  |  | July 20, 2016 | 2016 | Sterling Diagnostics, Inc. | Sterling Diagnostics, Inc. initiated a recall on 7/20/2016 for multiple diagnostic kits due to a recision of the previously extended expiration dates for the kits. | Process change control | Sterling Diagnostics, Inc. initiated their voluntary recall on 07/20/2016 by calling customers by phone to determine what kits remianed at customers' and compare withthe firm's sales records. Sterling Diagnostics followed up by sending a letter to each customer, and instructed them to take note of the corrected expiration dates for the dignostic kits, apply the revised dates to the kits in their possession, and return a customer acknowledgement to the Office Manager Luz Santa Mariaat Sterling Diagnostics, Inc. at 36645 Metro Court, Sterling Heights, MI 48312. Customers with questions can call 586-979-2141. |  | 245 | 245 | Domestic: MIForeign: PhillipinesVA/DOD: None | Terminated | March 22, 2017 | 245 | 1135
Z-0939-2017 | 75876 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DR Systems Unity PACS software, now known as Merge Unity PACS software. |  |  | October 15, 2015 | 2015 | Merge Healthcare, Inc. | The software fails to associate to the correct MG image if there are two images for the same view. | Software design | Merge sent an Urgent : Medical Device Recall letter dated July 11, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Be aware of the following workaround: Customer I User:Ensure both" tome" images are reviewed when more than one is captured for the same view. Merge has a released fix available for this issue.YOUR RESPONSE TO THIS NOTIFICATION IS REQUIREDPlease reply using the enclosed form and the return addressed envelope. Your response is required no later than August 2,2016. For further questions regarding this recall please call (877) 741-5369 |  | 9 | 9 sites potentially have the effected software | Distribution was made to medical facilities located in MT, CA, PA, and TX. | Open, Classified |  |  | 9
Z-0938-2017 | 75813 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific, INGEVITY" MRI Leads, Sterile EO, RxOnly. |  |  | November 30, 2016 | 2016 | Boston Scientific Corporation | Boston Scientific is conducting a recall on Boston Scientific INGEVITY" MRI endocardial pacing leads. There is a possibility that the polyurethane boot at the terminal end of the lead was not securely connected to the lead body. | Process control | Boston Scientific sent an "Important Medical Device Information" letter dated November 30, 2016. The letter was addressed to "Dear Doctor". The letter described the problem and the product involved in the recall. The letter also described the "Clinical Considerations, Recommendations, Affected Population &amp; Further Information". For questions contact your local Boston Scientific representative or Technical Services. For further questions, please call ( 651) 582-4000. |  | 15 | 15 (2 US; 13 OUS) | Worldwide Distribution - US Distribution to the states of : IA and NY., and to the countries of : United Kingdom, Netherlands, Italy, Germany and Belgium. | Open, Classified |  |  | 15
Z-0919-2017 | 75860 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Scope, Computed tomography x-ray system |  |  | November 22, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is providing a new installation of the Residual Current Monitor (RCM) in the Line Connection Box (LCB) to resolve a potential defect that could lead to a power cutout during use of the SOMATOM Scope, Scope Power and Perspective systems. | Nonconforming Material/Component | Siemens mailed on December 01, 2016 a Customer Safety Notice to affected customers to inform them of a corrective action that is being performed to prevent a possible hazard to persons and equipment. The notice explained the underlying issue, the effect the issue has on the system, the associated risks and actions they will take to correct the issue.In addition, customers will receive a Customer Safety Notice via Hand Delivery from the Siemens Customer Service Engineer at the scheduled appointment for the corrective installation. |  | 13 | 13 systems | Distributed to: ILILGASCSDOKNYMIINMDAR | Terminated | May 30, 2017 | 189 | 13