## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities
Z-0490-2018 | 78765 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Hemofiltration Standard Pack including BC 60 Plus-T 8036. Part Number: 701066990. |  |  | May 22, 2017 | 2017 | Maquet | 55 kits were distributed after their expiration date of September 2015. | Incorrect or no expiration date | Medical Device Recall (Removal) Notices were sent to customers via FedEx on 5/22/2017. The letters identified the affected device and the issue. Users are asked to examine their inventory to determine if any of the affected product is on hand. If found, affected devices should be removed and quarantined. Customers should contact Maquet/Getinge Customer Support at 1-888-627-8383 to obtain instructions on returning the affected devices. The Medical Device Recall Removal Response Form provided should be completed and returned. Questions should be directed to your Maquet/Getinge representative or call the Maquet/Getinge Customer Support at 1-888-627-8383. |  | 55 | 55 units | Distributed to the states of NM, CA, and IL. | Terminated | July 12, 2018 | 416 | 55
Z-0307-2018 | 78673 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD-Legacy 1, Model 6400 Ambulatory Infusion Pump, REF 21-6400-51, PC 10008552-001, Smiths Medical ASD, Inc. Rx ONLY |  |  | November 22, 2017 | 2017 | Smiths Medical ASD Inc. | Certain CADD Legacy 1, Ambulatory Infusion Pumps, may contain non-matching serial numbers. The label attached to the back of the pump may list a serial number that does not match the correct serial number programmed into the software of the pump. | Process design | The recall notices were disseminated via FedEx to US consignees on November 22, 2017. The recall notices were disseminated via email to Australian consignees on November 27, 2017. The consignee was directed to turn on the pump and check the serial against the serial number on the label. If they did not match, then they were directed to return the unit. |  | 58 | 58 units | US, Australia | Open, Classified |  |  | 58
Z-0309-2018 | 78613 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay System, Remanufactured, Catalog Number 386220. |  |  | November 06, 2017 | 2017 | Beckman Coulter Inc. | Potential for minor skin burn if the tool contacts the reagent storage interface board (B79917) when performing a pipettor alignment. | Device Design | An Urgent Medical Device Recall letter, dated November 6, 2017, were sent to the affected customers via mail. The letter identifies the affected device, the reason for recall, impact, and actions to be taken. The firm states that customers may continue to operate their system normally. Customers are instructed to contact Technical Support if a device on the Access 2 instrument needs alignment. The alignment tool is only to be used by the Beckman Coulter service personnel and Beckman Coulter trained biomedical engineers. A Beckman Coulter service rep will contact customers to schedule a service visit. Customers are to share this information with their staff and retain a copy of the letter. Customers are to respond via one of the ways identified. Beckman Coulter should be contacted with any questions regarding the notice. |  | 3 | 3 units | Worldwide Distribution -- US, Argentina, Canada, Chile, China, Columbia, Costa Rica, France, Georgia, Germany, Ghana, Greece, Hong Kong, India, Italy, Japan, Kazakhstan, Mexico, Morocco, Myanmar, Nepal, Philippines, Poland, Russia, Rwanda, Saudi Arabia, Slovakia, South Africa, South Korea, Thailand, Turkey, United Kingdom, and Vietnam. | Open, Classified |  |  | 232
Z-0311-2018 | 78833 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Mobile Detector Holder - Stationary X-ray System (RF System). suitable for positioning the wireless portable detector free in the room for X-ray exposures. It supports the wireless portable detector with or without a grid. |  |  | July 14, 2017 | 2017 | Philips Electronics North America Corporation | If the lock mechanism of the movable camp of the mobile detector holder is not adjusted properly or not secured, the portable detector can drop down when the moveable clamp is at the bottom (below of the detector) and additional force is applied. This may happen when the detector is rotated clockwise from landscape to portrait position. | Under Investigation by firm | If the lock mechanism of the movable clamp of the mobile detector holder is not adjusted properly or not secured, the portable detector can drop down when the moveable clamp is at the bottom (below of the detector) and additional force is. applied. This may happen when the detector is rotated clockwise from landscape to portrait position.The Wireless Portable Detector (WPD) is fixed in the Mobile Detector Holder via a lock mechanisrn. The 4 countersunk screws, that are used to attach the lock mechanism, can become loose. If this happens and the mechanism fails, thedetector may fall down when the holder is rotated 90 clockwise from Iandscape. |  | 412 | 412 | 412 units affected. | Open, Classified |  |  | 412
Z-0312-2018 | 78758 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PPS RingLoc+ Acetabular Shell, RNGLC+ LTD HOLE SHELL SZ62 |  |  | June 29, 2017 | 2017 | Zimmer Biomet, Inc. | Device was distributed with the etching missing from the product. | Employee error | The firm, Zimmer Biomet, sent an, "URGENT MEDICAL DEVICE RECALL REMOVAL" letter dated 6/28/2017 via email and FedEx on approximately 06/29/2017 to its consignees. The letter described the product, problem and actions to be taken. The consignees were Instructed to ensure appropriate personnel are notified, immediately locate and quarantine affected product in inventory, complete and return the Certification of Acknowledgement form by email to:corporatequality.postmarket@zimmerbiomet.com or fax to: 574-372-4265, return all affected products, and provide the names of any locations where product was further distributed to.If after reviewing this notice you have further questions or concerns please call the customer call center at 1-574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an on-call representative in the case of an emergency. Alternatively, your questions may be sent by email to corporatequality.postmarket@zimmerbiomet.com. |  | 15 | 15 units | Distributed domestically to MI, MN, ND. | Terminated | March 07, 2018 | 251 | 15
Z-0313-2018 | 77761 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medfusion, Syringe Infusion Pump, Software Version V1.1, Rx Only, Manufactured by Smiths Medical ASD, Inc., 1500 Eureka Park, Lower Pemberton, Ashford, Kent, TN25 4BF, UK. |  |  | August 22, 2017 | 2017 | Smiths Medical ASD Inc. | Potential for the syringe barrel clamp mechanism to loosen, decreasing the pump's ability to accurately detect the diameter of the syringe barrel, resulting in the unit to initiate the "Invalid Syringe Size" alarm. | Component design/selection | The firm,Smiths Medical, sent an, "Urgent Medical Device Recall" Amendment Notice (amending the original August 2017) concerning invalid syringe size alarms on certain MedfusionÂż Syringe Pump Model Series 3500 and 4000. The letter described the product, problem, actions to be to taken, and updated changes, as follows:1. Locate the affected MedfusionÂż Syringe Pump(s) by referring to the list of affected devices on the attached Urgent Medical Device Recall Notice Response Form. This form provides the specific serial number(s) your organization purchased. Each pump has a unique serial number which can be found on the label located on the bottom of the pump.If you have ten (10) or more affected pumps in your possession, please select one of the following repair options:a. Repair at your Facility: Contact Smiths Medical to schedule a technician to visit your facility and complete the necessary repairs onsite.b. Send to Smiths Medical for Repair: Impacted pumps may be sent to a Smiths Medical Service Center, where repairs will be completed. A prepaid shipping label will be provided. Repairs will be completed within ten (10) business days of receipt. Loaner units are available upon request.c. NEW OPTION - Biomedical Department self-repair at your Facility: Contact Smiths Medical to receive the necessary parts and work instruction to have your biomed team complete the necessary repairs onsite.If you have nine (9) or fewer affected pumps in your possession:a. Send to Smiths Medical for Repair: Impacted pumps may be sent to a Smiths Medical Customer Service Center, where repairs will be completed. A prepaid shipping label will be provided. Repairs will be completed within ten (10) business days of receipt. Loaner units are available upon request.b. NEW OPTION - Biomedical Department Self-Repair at Your Facility: Contact Smiths Medical to receive the necessary parts and work instruction to have your biomed team complete the necessary repairs onsite.To initi |  | 9787 | 9,787 units in total | Worldwide Distribution: US (Nationwide) and countries of: Canada, Great Britain, Panama, Germany, India, Philippines, and Oman. | Open, Classified |  |  | 9787
Z-0315-2018 | 78810 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | SafetyGlide Needle, 22GA 1-1/2in, Catalog n. 301701, manufactured by BD, sold in |  |  | July 05, 2017 | 2017 | Centurion Medical Products Corporation | Centurion is recalling specificlots of convenience kits, SafetyGlide Needle, 22GA 1-1/2in, manufactured by BD, sold inCenturion Convenience Kits. distributed between May 2016 and June 2017 due to notice received from BD's recall due to particulate that could be expelled out of the needle if used for injection. | Material/Component Contamination | Centurion Medical Products Corporation sent a notification mailed to the affected customers and distributors via Certified Mail Return Receipt on July 5, 2017 |  |  |  | US. | Terminated | March 09, 2018 | 247
Z-0316-2018 | 77986 | No | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Ceres Self-Drilling Variable Screw 4.0 x 12mm |  |  | July 07, 2017 | 2017 | Amendia, Inc | Incorrect direct marking and color identification for the affected lot of Ceres Self-Drilling Variable Screws. Although the laser mark and color identify the screws as 12 mm in length, some screws in the lot measure 14 mm in length. | Mixed-up of materials/components | The firm, Amendia, sent an "URGENT: MEDICAL DEVICE RECALL" letter to its Customers. The letter described the product, problem and actions to be taken. The customers were Instructed to examine inventory and quarantine affected products, notify customers if affected products were further distributed, coordinate the return of affected products back to Amendia by emailing: jconover@amendia.com or call 770-575-5236 or jwilliamson@amendia or call 770-575-5369, and to complete and return the Response Form. If there are questions, Amendia Customer Service can be contacted at 770-575-5224. |  | 40 | 40 screws | Distributed to states of: AL, FL, GA, NY, TX, and VA. | Open, Classified |  |  | 40
Z-0319-2018 | 78571 | No | LEJ | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sundream 12 V tanning bed |  |  | December 02, 2017 | 2017 | ProSun International, LLC | The 30 minute maximum tanning time has been reduced to 20 minutes. | Other | Customers were initially notified of the recall on 03/24/2016 and then again on 10/17/2017. Instructions included to check current inventory for the affected device and complete and return the response card to ProSun who will then supply a set of 12 lower output lamps that are rated at 30 minutes maximum tanning time, provide a new Owner's Manual, and a new Tanning Schedule Label with the new 30 min lamps listed. |  | 75 | 75 units | Distributed domestically to AK, AL, AZ, CA, CT, FL, GA, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NV, NY, OH, OK, PA, SC, TN, TX, UT, WA, WI | Open, Classified |  |  | 75
Z-0317-2018 | 76569 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Freedom60 Precision Flow Rate Tubing Sets, REF/Part Numbers F900, F2400, F60, F10, F275, F600, F45, F1200 F180. |  |  | March 10, 2016 | 2016 | Repro-Med Systems, Inc. | On 2/24/2016 during post sterilization inspection of part number F180 Precision Flow Rate Tubing, Lot number 1.042/16 the firm found 701 out of 3,195 were rejected due to visual defect, gap in the supplier s manufacturing seal for Part number #317036 5 X 6.5 ST Bag Combo. the firm decided to recall Precision Flow Rate Tubing and RMS HIgH-Flo Subcutaneous Safety Needles Sets due to the sterile barrier packaging being found to have gaps during post-sterilization inspection of one lot of flow tubing, in which 701 out of 3,195 units were rejected for not passing inspection. The firm determined that the issue was large than just the one lot and recalled the product via March 10, 2016 Voluntary Medical Device Corrections and Removal Notification to customers. The recall was conducted without FDA notification. | Nonconforming Material/Component | On March 10, 2016 RMS Medical Products distributed Urgent Voluntary Medical Device Corrections and Removal notices to their customers. On March 12, 2016 &amp; on March 18, 2016, RMS Medical Products distributed a 2nd Urgent Voluntary Medical Device Corrections and Removal notices to their customers advising them that affected products list was reduced after a comprehensive risk assessment was performed. Customers are advised to check stock, quarantine, cease distribution and remove product from use. If customers further distributed product, they are also encouraged to inform the affected parties. Customers are encouraged to complete and return the acknowledgement form via fax to: 845-469-5518. |  |  |  | Nationwide in US | Terminated | February 26, 2018 | 718
Z-0321-2018 | 78799 | No | CAH | Breathing circuit bacterial filter. | Anesthesiology | General Hospital | 510(k) | UNIVERSAL FLEX2 BREATHING CIRCUIT, Item Numbers: D460-6133Z, D390-6121Z, D390-6121Z, D495-61Z, DF4116-6121Z, D466-61Z, DFP170-6121, DF3115-6121Z, DF375-6121Z, DF3115-6121Z, DF3115-6121Z, D360-8031Z, D465-61Z, JD365-6033Z, DF370-6121Z, DF3110-61Z, DF375-17621Z, DF375-6121Z, DF475-17621Z, DF370-6121Z, DF375-6121Z, DF470-6021Z, D366-618Z, D390-60Z, D390-61Z, DP260-61, DF370-6121Z, D3120-6154Z, F95100, JD365-6033Z, DF370-61Z, DF475-61Z, DF3110-6121Z, DF3110-6121Z, DF375-6121Z, DF475-6121Z, F90755, F90798, F95368, PDP260-61, D396R-61Z, F90578, F90622Z, F94980, D360-61Z, D365-6121Z, D366-6121Z, D366-6121Z, SJDF3116-6121Z, DF375-6121Z, DF470-6121Z, DF470-61Z, DF475-6121Z, DF375-6033Z, DF375-618Z, DF470-6127Z, DF475-61Z, DF476-6121Z, F90756, PDB160-6131, PDP160-6121, PDY160-6121, D060-8021, D366-8021Z, D466-6121Z, DF370-6121Z, DF470-6121Z, DF376-6121Z, PDP160-6031, D360-8021Z, D360-80Z, D445-80Z, D465-8021Z, JD365-6033Z, D460-6131Z, D460-6133Z, DF475-6121Z, DF3116S-608Z, F90777, D360-6133Z, D365-80Z, D360-6021Z, D360-6121Z, D360-61Z, DF3110-61Z, DF370-61Z, DF375-17621Z, DF475-6121Z, D365-6121Z, D060-80, PD060-6121, PD160-6121Z, PDP160-6121, DF3115-6121Z, F90767, DF375-17621Z, DF375-6121Z, DF470-6121Z, PD260-61Z, PDY140-6133, D360-16421Z, D360-6121Z, D366-6121Z, DFP2110-6121, DF470-6121Z, D465-8021Z, DF4115-6121Z, DF370-6121Z, DF376-6121Z, PDB140-61, D365-6121Z, DF070-61, PDY160-6121, D365-61Z, D360-6121Z, D360-61Z, D460-6121Z, DF0110-6121, DF3110-61Z, DF370-61Z, DF375-6121Z, DF375-6121Z, DF475-6121Z, DF475-6121Z, DF370-6021Z, DF370-6133Z, DF375-618Z, DF475-17621Z, DFB170-61, PD040-61, PDP160-6121, PDY160-6121, D390-61Z, F94474, D460-61Z, D366-6121Z, D360-6121Z, D360-6121Z,D360-6121Z, D360-6121Z, DF0110-61, DF0110-61, DF3115-6121Z, DF3115-6121Z, DF070-6127, F90752, DF375-17621Z, DF375-61Z, DF376-17621Z, F95100, D360-6021Z, D366-6121Z, D465-8021Z, D360-61Z, DF3110-61Z, DF3110-61Z, DF3110-6121Z, PDP160-6031, D360-6121Z, D360-6121Z, D360-6121Z, D360-6121Z, D360-6121Z, D360-6121Z, D460-61Z, D3120-6121Z, D360-8021Z, D365-80Z, D395-8021Z, D460-8021Z, D465-8021Z, D495-61Z, F90577Z, DF370-6133Z, DF375-6021Z, DF376-6121Z, DF475V-6121Z, DFP170-6121, DF370-61Z, DF470-6121Z, DF470-6121Z, DF470-6121Z, DF470-6121Z, DF470-61Z, DF3115-6121Z, F90767, DF370-6121Z, DF375-6121Z, DF476-6121Z, F95407, F90556, D365V-8021Z, F95383, D460-61Z, F90777, DF370-6121Z, DF370-6121Z, DF375-17621Z, DF475-6121Z, D440-80Z, D360-6021Z, D365-6121Z, DF370-61Z, DF4116-6121Z, DF476-6121Z, F90579, D345-6121Z, D396R-61Z, D460-80Z, D465-6031Z, D465-8021Z, DP2120-61, F90578, F94795, F94980, D360-61Z D365-6121Z, D365-6121Z, D366-6121Z, D366-6121Z, D460-6121Z, D460-6121Z, D460-61Z, D360-6021Z, D360-6121Z, D366-6121Z, D366-6121Z, D060-8021, D3120-6121Z, D360-16421Z, D360-8021Z, D360-8031Z, D366-80Z, D390-61Z, D460-80Z, D465-8021Z, D465V-8021Z, PDP160-6121, DF070-6121, DF370-6021Z, DF375-61Z, DF375-61Z, DF375-61Z, DF470-6021Z, DF475-61Z, DFY170-6121, DF375-17621Z, DF375-6121Z, DF375-6121Z, DF470-6121Z, DF475-6121Z, DF475-6121Z, SJDF3116-6121Z, DF3115-6121Z, DF3116-6121Z, DF4115-6121Z, DF4116-6121Z, DF475-6121Z, DF375-6121Z, DF375-6121Z, DF470-6121Z, DF0110-618, DF370-61Z, DF375-6121Z, DF475-6121Z, DF475-6121Z, DF375-6021Z, F90798, PDP140-6121, D365V-8021Z, D366-8021Z, D445-80Z, D460-8021Z, D460-8021Z, D465-61Z, D365-61Z, DF3116S-608Z, DF370-6121Z, DF375-17621Z, DF375-6121Z, DF470-6127Z, DF475V-6121Z, DF476-6121Z, F95368, F95407, PD060-61, PDB160-6131, PDP160-6121, D460-6131Z, D360-6121Z, D360-61Z, DF370-6121Z, DF375-17621Z, DF475-6121Z, DF376-17621Z, D3120-6121Z, D466-6121Z, DF4110-6121Z, DF370-6121Z, DF376-6121Z, DF475V-6121Z, D440-80Z, DF376R-61Z, DF0110-6121, DF370-6121Z, DF470-6121Z, DF375-17621Z, DF370-6021Z, DF475-61Z, DFP170-6121, DFY270-6121, PDP264C-61, D360-8021Z, D365-8021Z, D365-80Z, D365-80Z, F90578, D360-6121Z, D360-6121Z, D360-61Z, D460-6121Z, DF3115-6121Z, DF3115-6121Z, DF4115-61Z, DFP2110-61, DF375-6121Z, DF470-6121Z, DF475-6121Z, DF476-6121Z, DF470-6121Z, DF475V-6 |  |  | March 15, 2017 | 2017 | King Systems Corp. dba Ambu, Inc. | Potential for leak test failures. | Nonconforming Material/Component | On 3/15/2017, the firm issued a letter to its sales force describing its intent to replace product for consignees that complained about leaking. The firm also issued a copy of the manifold IFU that includes the need for a 1/4 turn in the Universal Flex2 Breathing Circuit (with bacterial filter). |  | 871116 | 871,116 units | Distribution Pattern  Distributors in 44 states: Al, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, WI WY; International Pattern - Canada, China, Chile | Terminated | December 17, 2018 | 642 | 871116
Z-0320-2018 | 78715 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Brainlab Navigation System Spine &amp; Trauma 3D with component Brainlab DrapeLink Registration Kit for C-arm Zeihm Vision RFD: Adapter Left |  |  | November 16, 2017 | 2017 | Brainlab AG | Due to a human error at a supplier, an unknown number of devices within one specific production lot of Brainlab DrapeLink Adapters (Left/Right) were notmanufactured according to defined specifications. | Nonconforming Material/Component | Customers were notified, via email on approximately 11/16/2017. Instructions included to identify all affected devices on hand, cease use of the affected device, and a Brainlab Customer Support Representative will contact the customer to arrange a service visit to replace the item |  | 80 | 80 devices total | Distributed in the US to Kansas.Distributed internationally to Australia, Finland, France, Germany, Hong Kong, Ireland, Japan, Netherlands, Norway, Sweden, Switzerland, Taiwan, Turkey, United Kingdom. | Open, Classified |  |  | 80
Z-0409-2018 | 78840 | No | OSN | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew Taylor Spatial Frame Website software, Catalog Number 71070401. A web-based and mobile application software component of a multilateral external fixation system that is intended for fracture fixation, fixation of long bones and for joint fusions and limb lengthening or deformity corrections which involve cutting of the bone. |  |  | December 13, 2017 | 2017 | Smith &amp; Nephew, Inc. | In rare circumstances, when surgeons use the software to plan the strut adjustment, an inverted image of the frame is generated by the software. The frame will appear inverted in the preview images on the Strut Settings tab. | Under Investigation by firm | The firm initiated their recall by email on 12/13/2017. Users are asked if they notice an inverted frame in the software to repeat the process for the same case and notify a Smith &amp; Nephew representative. In addition, they are asked to double check plannings for strut adjustments and compare with the images printed as described above. |  |  |  | Worldwide | Open, Classified
Z-0322-2018 | 78588 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Syringe Pump, Model 8110. The firm name on the label is CareFusion. |  |  | November 17, 2017 | 2017 | CareFusion 303, Inc. | The syringe plunger grippers may fail to automatically close around the syringe plunger press when the gripper control knob has been closed. | Nonconforming Material/Component | The recalling firm issued "Medical Device Safety Notification, dated November 17, 2017, via UPS. The letter describes the potential issue, affected products, potential risks, and required action for the users. The actions noted that clinicians can continue to use the device and follow the Alaris System User Manual for proper loading of the syringe. Per the manual, clinicians should ensure the plunger grippers lock and hold the syringe plunger in place. The Biomedical Engineering can detect if the device has this issue and if the issue described occurs, the user is to contact BD Support Center or email them to schedule service of the device at the BD Service Depot.Enclosed with the notification letter was a list of affected serial numbers, an excerpt from the user manual to "Prepare Syringe and Administration Set (Syringe Module)", and a customer response card to be returned via mail, email or FAX to confirm the notification has been received.For further questions, please call (858) 617-5977. |  | 73611 | 73,611 | Worldwide Distribution - USA (nationwide) Distribution. There was also government and military distribution. Foreign distribution was made to Canada, United Arab Emirates, Bahrain, Canada, Switzerland, United Kingdom, Kuwait, Mexico, Malaysia, The Netherlands, New Zealand, Qatar, Saudi Arabia, Taiwan, and South Africa. | Open, Classified |  |  | 73611
Z-0259-2018 | 78784 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Mobile Mini C-arm system Part # 1000-0004. |  |  | August 31, 2017 | 2017 | Orthoscan, Inc. | OrthoScan Inc. discovered during investigation of a non-standard work flow a non-conformity of the devices. Specifically, the system software allows a user to activate the Digital Zoom feature in the course of taking a live image exposing a larger x-ray field than can be viewed by the user | Software design | Orthoscan sent an Urgent Recall Notice, dated August 30, 2017, to customers, informing them of the issue and how to address it. Customers are to use the attached list of serial numbers to determine if their device is affected. Customers may continue to use their affected device with caution until the on-site correction has been made. Users should avoid non-conforming workflows by taking the measures outlined in the letter. The firm will contact customers to arrange implementation of the corrective action at no cost to the customer. If there are further questions or concerns, please call National Customer Service Team at (866) 949-4346. |  | 2012 | 2012 total | Worldwide Distribution - USA (nationwide) including Puerto Rico and Hawaii and to the countries of : Australia, Austria, Belgium, Canada, Dominican Republic, Estonia , Finland, France, Germany, Hanover, Hong Kong, Israel, Italy, Japan, Jonesboro, Korea, Malaysia. Netherland, New Zealand, Norway, Poland, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, UAE, and United Kingdom. | Open, Classified |  |  | 2012
Z-0578-2018 | 78899 | bedding, disposable, medical | KME | Medical disposable bedding. | General Hospital | General Hospital | 510(K) Exempt | EMS products from Graham Medical, SnugFit(R) Expandable Mattress Cover, Code 72930, UPC 1 07 47036 72930 3 |  |  | December 27, 2017 | 2017 | LITTLE RAPIDS CORPORATION | The secondary supplier provided material was within specification but did not meet our design needs compared to the primary manufacturer. The product is designed for patient transfer, not lifting. However, we recognize that the product may be utilized outside of its intended design scope, and lifting may occur. Lifting could lead tearing of the equipment cover and the potential for patient falling through. | Device Design | The firm initiated their recall by email on 12/27/2017 to their direct consignees. The consignees were directed to report back with the number of products on hand for return and to notify their customers of the recall. |  | 4590 | 153 cases, 30 units in each case | TX, CT, TN, MO, WI, FL, SD, MS, CA, MS | Terminated | July 30, 2018 | 215 | 4590
Z-0324-2018 | 78674 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cordis S.M.A.R.T. Flex Biliary Stent System |  |  | November 15, 2017 | 2017 | Cordis Corporation | Potential for cracked luer hubs | Nonconforming Material/Component | On November 15, 2017 the firm sent letter to their customers indicating the following:The purpose of this communication is to inform you of a potential leakage scenario with the S.M.A.R.T.Âż Flex Biliary Stent System. The leak may occur due to potentially cracked hubs in some units from within the 29 lots listed in this letter.Inspection details:As stated in the Instructions for use, users are warned that the device should not be used if the outer sheath cannot be flushed. Inspect for leakage during preparation of the device. Leakage may be more likely in the affected product due to cracks in the luer. At time of actual flushing, if there is no leakage observed, the product may be safely used.Available Assistance:We can provide help if you have any questions regarding this Correction.For any health care professional with medical concerns, please contact Cordis to speak to a clinician at 1-800-327-7714 Option 3, Monday through Friday from 9:00 AM to 5:00 PM EST.For questions related to the Correction that are not adequately addressed in this letter, please contact the Cordis Field Action Coordinator Hal Baden at (786) 313-2365. |  | 511 | 511 | Worldwide Distribution - US Nationwide in the states of MI, AL, AR, CA, CO, FL, GA, IL, IA, LA, MS, NJ, NC, OH, OK, OR, PA, TN, TX, VA, WV, WI, MI, NY, MA, NV, RI, IN, WA, AZ, DC, Puerto Rico and the country of Canada. | Open, Classified |  |  | 511
Z-0517-2018 | 78105 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | PC-1000, PC-1000/Laser1000, Model Numbers 800724-2 &amp; 800724-3 |  |  | July 14, 2017 | 2017 | Panoramic Rental Corp. | The lift motor used to raise and lower the overhead chassis will wear if not properly maintained according to the maintenance schedule. | Device Design | On July 14, 2017, Panoramic distributed Urgent Medical Device Correction notices to customers who own and have not taken action as required for the PC-1000 X-ray system machines. This means that your device does not have a critical safety feature that is intended to protect your patients and staff from harm. Customers are encouraged to immediately discontinue use of their machine and follow the instructions in the notice. Customers are encouraged to not resume use of the machine until a certified Panoramic service technician repairs the lift motor. Customers are advised to contact Panoramic to ensure any safety risks associated with the lift motor are identified and addressed. Action Required:(A) If your PC1000 machine is still in service at this location, you must immediately discontinue use and choose one of the following by returning the Acknowledgement Form or contacting the Panoramic Service Team by calling 888-381-2577 Monday through Friday 8 am - 6 pm.1. An upgraded lift motor assembly at no cost to you;2. A digital Panoramic rental unit to replace your PC1000, contingent on a signed one year rental contract (subject to availability); or3. A new digital Panoramic or PanoramicCephalometric unit to replace your PC1000 at a reduced price (startingat $17,995 depending on machine type and subject to availability)(B) If your PC1000 machine is no longer in service or under your control:4. If the machine is no longer in service, Panoramic will provide free removal of your machine. |  | 13340 | 13,340 units | US Nationwide distribution including Puerto Rico | Terminated | July 23, 2018 | 374 | 13340
Z-0325-2018 | 77991 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Mac Pin Non-Cannulated Screw, 6.5 x 60mm |  |  | July 07, 2017 | 2017 | Amendia, Inc | Incorrect labeling for the lot of MAC Pins. Although the pins are laser as 460 mm in length, the pins are 60 mm in length. | Process control | Customers were sent recall notification letters on 07/07/2017. Instructions included to examine inventory and quarantine affected products, notify customers if affected products were further distributed, coordinate the return of affected products back to Amendia, and to complete and return the questionnaire. If there are questions, Amendia Customer Service can be contacted at 770-575-5224. |  | 41 | 41 screws | Distribution to Georgia, USA. | Open, Classified |  |  | 41
Z-0329-2018 | 78727 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EMPOWR Porous Knee System TIBIAL PUNCH, LARGE, REF 801-05-204, NON STERILE, Qty 1 For use in orthopedic surgery. |  |  | December 12, 2017 | 2017 | Encore Medical, Lp | The Locking Punch Guide was not sufficiently guiding the large press fit punch, causing the implant to be placed 2-3mm more posterior than intended. | Device Design | The field safety notice was disseminated by email on 12/12/2017. The notice stated the following: "DJO Sales Agents Action:Please review the bulletin that is provided with this notice. This bulletin contains new instructions on the use of the Locking Punch Guide and Punch. Please ensure your surgeons are familiar with this information.DJO Surgical then requires that you complete the attached acknowledgment and email it to teffany.hutto@djoglobal.com. Please have this form returned by December 15, 2017Once redesigned guides are available, they will be exchanged through a separate field action." |  | 2 | 2 units | One medical device distributor in California. | Completed |  |  | 2
Z-0330-2018 | 78880 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | CombiDiagnost R90 Image-Intensified fluoroscopic x-ray system |  |  | October 19, 2017 | 2017 | Philips Electronics North America Corporation | The tube adapter-plate green/yellow protective earthing (PE) wire was removed. and spacers with different thickness were used to position pre-loc ring in respect to focal spot in specified distance. | Nonconforming Material/Component | Philips Medical System sent an Customer Notification letter dated October 19, 2017, and plans to have a Philips Field Service Engineer go to site to perform the field change order (FCO70900041) including: 1. Check for the presence of the PE cable and, if missing, install it.2 .Perform a modified screw fixation.They expect to complete action by May 2018. For further questions, please call (978) 659-3000. |  | 10 | 10 | Worldwide Distribution - US Distribution and to the countries of : France, Germany, Luxembourg and New Zealand. | Terminated | September 20, 2018 | 336 | 10
Z-0293-2018 | 78738 | Yes | NPT |  |  | Cardiovascular | PMA | Certitude Delivery System and any associated Multipacks Models - 9600CT20A, 9600CT23A, 9600CT26A, 9600CT29A |  |  | July 21, 2017 | 2017 | Edwards Lifesciences, LLC | A component in the Certitude Loader may have amolding overflow defect that could detach during advancement of the delivery system andpotentially embolize from the loader into the patient. | Employee error | Recall notification letters were sent to affected consignees via Fed-Ex on 7/21/2017. |  | 1730 | 1730 units | United States: All states with the exception of: Alaska, Hawaii, Idaho, New Mexico, Rhode Island, South Dakota and Wyoming | Open, Classified |  |  | 1730
Z-0294-2018 | 78574 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK 20e DEFIBRILLATOR/MONITOR, 3201273 |  |  | December 04, 2017 | 2017 | Physio-Control, Inc. | Physio-Control is aware that some devices have had power-related failures as customers prepared their device for initial deployment or during use within the first year of distribution. The symptoms of these failures may include unexpected power on and power off, device lock-up, or a failure to power on or off, any of which has the potential to result in a failure to deliver therapy to the patient and serious injury or death. | Other | The recall was initiated by letter on 12/04/2017. The recall notice stated:"Physio-Controls Planned ActionsPhysio-Control will contact customers with LIFEPAK 20e devices who are potentially affected to arrange for a device correction. This correction will include the replacement of the Power PCBA.Required Customer Actions1. Please forward this letter to all of your sites, trainers, and users that have an affected LIFEPAK 20e device(s) as identified on the attached Confirmation Sheet.2. Promptly return the completed Confirmation Sheet to Physio-Control.3. Follow the recommended daily Operators Checklist steps in accordance with LIFEPAK 20e defibrillator/monitor Operating Instructions  Section 7  Maintaining the Equipment. The checklist can be found in Appendix D of the Operating Instructions.4. You may continue to use your defibrillator/monitor as long as it passes steps 1 and 8 of the Operators Checklist.5. If you experience any of the symptoms described above, or the defibrillator/monitor fails steps 1 and/or 8 of the Operators Checklist, contact Physio-Control immediately to arrange servicing of your device.If your LIFEPAK 20e defibrillator/monitor exhibits any power issues that cannot be resolved, contact Physio-Control immediately. Contact Physio-Control if you have any questions about this matter at 1-800-442-1142, option 7, 6:00 A.M. to 4:00 P.M. (PST), Monday  Friday." |  | 3831 | 3831 units | Worldwide Distribution | Open, Classified |  |  | 3831
Z-0485-2018 | 78042 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Empty IntraVia Container with Non-DEHP Fluid Path, 500mL |  |  | July 27, 2017 | 2017 | Baxter Healthcare Corporation | A leak may allow for microbial contamination of the sterile fluid path. | Process control | US consignees were notified via letter on July 27, 2017 and Canadian consignees were notified on July 31, 2017. Instructions included to locate and remove all affected product from the facility, contact Baxter Healthcare to arrange for return of product, complete and return the Reply Form, and to notify customers if the product was further distributed. For further questions, please call (800) 422-9837. |  | 29088 | 29,088 units | Distribution in the US and Canada. | Open, Classified |  |  | 29088
Z-0398-2018 | 78901 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | OSCILLATING SAW ATTACHMENT |  |  | September 29, 2017 | 2017 | The Anspach Effort, Inc. | This voluntary recall was initiated because a veterinary unit was physically shipped in error to a non-veterinary customer account. | Under Investigation by firm | On November 15, 2017 the firm sent letter to their only consignee.Our records show that your facility has received the product subject to this removal.1. Immediately review your inventory to identify and quarantine the affected product listed above in a manner that ensures the affected product will not be used.2. Review, complete, sign and return the attached reply form on page 3 of this letter to your local DePuy Synthes sales organization in accordance with the directions on the form within 5 business days of receipt of this notification.3. Return the affected product as soon as possible, but within 30 business days.4. Forward this notice to anyone in your facility who needs to be informed.5. If the affected product has been forwarded to another facility, contact that facility to arrange its return.6. Maintain awareness of this notice until the product has been returned.7. Keep a copy of this notice.We apologize for any inconvenience that this product recall (removal) may create and appreciate your cooperation with our request. Should you have any inquiries please do not hesitate to contact your DePuy Synthes sales consultant. Thank you for your attention and cooperation.For further questions, please call (610) 314-2106. |  | 1 | 1 | Internationally to Canada | Open, Classified |  |  | 1
Z-0332-2018 | 78879 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Brilliance iCT SP computed tomography x-ray system |  |  | November 08, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | There is a possibility that screws holding an internal portion of the gantry (Reaction Ring) to the scanner s main bearing may come loose. If this occurs,the part(s) will not be expelled, however, a noise may be heard coming from the gantry. | Device Design | Philips Healthcare is informing consignees of this issue through a Customer Information Letter (CIL_FCO72800690). Philips Key Market representatives will distribute the letter internationally, based on local regulations. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. Customers are notified that if they hear an unusual noise coming from the gantry, or scan halts, the Philips representative should be contacted as soon as possible. Philips will be inspecting the affected systems, tightening and replacing the reaction ring screws as needed at no additional cost. |  | 134 | 134 | 107 domestic accounts15 US Govt accounts682 Foreign Accounts3 Mexico accounts12 Canada accounts | Open, Classified |  |  | 886
Z-0335-2018 | 77617 | Yes | BSK | Inflatable tracheal tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Intellicuff Standalone (PN 951001) |  |  | June 20, 2017 | 2017 | Hamilton Medical, Inc. | Issue related to the performance of the motor in recalled product. During use, motor may cease to function. | Component design/selection | Hamilton Medical sent a Recall Notice dated June 20, 2017, notifying them of the device correction. Consignees were informed that a Hamilton Representative would be in contact with them regarding the return of any devices. For further questions, please call (775) 858-3200. |  | 176 | 176 units | US Distribution to IL and TX. | Open, Classified |  |  | 176
Z-0334-2018 | 78843 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | DeRoyal Sterile Custom kits containing Covidien-Dover(TM) Irrigation Tray with Piston Syringe Containing Medline Aplicare Povidone Iodine Prep Pads, labeled as follows: |  |  | December 19, 2017 | 2017 | DeRoyal Industries Inc | DeRoyal manufactured custom sterile surgical kits containing the Covidien-Dover(TM) Irrigation Tray with Piston Syringe that contains the Medline Aplicare Povidone Iodine Prep Pads. The Prep Pads were subsequently recalled because Medline Aplicare determined the product would not support the 36 month expiration date listed on the label. | Nonconforming Material/Component | The firm, DeRoyal, sent an "URGENT! RECALL NOTICE" letter dated 12/19/2017 to its Consignees overnight. The letter describe the product, problem and actions to be taken. The Consignees were instructed to identify affected product; place in quarantine to prevent further use, and complete and return the Affected Products Listing form to DeRoyal via fax to 865-362-3716 or email to recalls@deroyal.com no later than 1/15/2018, even if you no longer have any inventory. If you have questions or need assistance with the recall, please email jmarsh@deroyal.com or call by phone 865-362-4203 or email us at recalls@deroyal.com. |  | 226 | 226 kits | Distributed to states of: CA and MO. | Terminated | June 21, 2018 | 184 | 226
Z-0337-2018 | 78640 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TRYPTIK(R) 2 C-PLATE SELF-TAPPING POLYAXIAL SCREW 4.0 MM, LENGTH 14 MM, REF TRY-PS-40 14-S, Qty 2 |  |  | November 27, 2017 | 2017 | SPINEART SA | Product labeling for TRYPTIK 2-C-PLATE Self Tapping screw, REF TRY-PS 40, Lot 3-5934 contains the sketch/drawing of a self-drilling screw. | Error in labeling | Spineart sent an Urgent Medical Device letter by email dated November 27, 2017. The letter stated the following:"Out of an abundance of caution", we have decided to recall the concerned products.Please carefully follow instructions detailed below:1. Immediately examine your inventory and quarantine product subject to recall.2. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall by using this document.3. Collect and quarantine all products.4. Send back all products and the enclosed Response Form to Spineart Geneva E-mail: regulatory@spineart.com."For further questions, please call (949) 313-6400.. |  | 120 | 120 units | US Distribution to the state of : CA | Open, Classified |  |  | 282
Z-0347-2018 | 78716 | No | NRY | Percutaneous catheter. | Cardiovascular | Neurology | 510(k) | Trevo XP ProVue System kit containing 1) Trevo XP ProVue Retriever, 4x20mm and 2) Trevo Pro18 Microcatheter |  |  | November 28, 2017 | 2017 | Stryker Neurovascular | Stryker Neurovascular has become aware that at least 1 (one) Trevo XP Provue System was manufactured where the carton sleeve did not match the physical contents of the pack. | Packaging process control | The affected consignees will be notified via letter sent a courier beginning 11/28/17. A response form has been provided which customers have been asked to complete, and correct affected product is being requested. |  |  |  | OUS only - NL, FR, IT, DE, BE | Open, Classified
Z-0352-2018 | 78790 | No | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Medtronic KYPHON(R) Express(TM) Directional Bone Filler Device, SIZE 2, REF FB, QTY 1 EA, Rx only STERILE R |  |  | December 06, 2017 | 2017 | Medtronic Sofamor Danek USA Inc | Medtronic has discovered the directional arrow at the proximal end of the instrument may not correctly align with the cut out opening on the distal end of the instrument. | Process control | Medtronic distributed a voluntary removal notice via FedEx courier service to impacted U.S. consignees on 06 DEC 2017. The recall notice stated the following: "If he/she has not already done so, your Sales Representative will contact you to facilitate the return of any impacted products you may have in your possession. Please disseminate this information to additional personnelwithin your facility as appropriate, and maintain a copy of this notice in your records.Instructions to Customers:1.) Please locate and remove the impacted product from normal storage locations. Do not use this product.2.) Complete the enclosed Customer Confirmation Form and return3.) As applicable, contact your Medtronic Sales Representative to coordinate the return and replacement of affected product." |  | 3319 | 3,319 units | US. Austria, Belgium, Bulgaria, Canada, France, Germany, Italy, Lebanon, Luxembourg, Portugal, South Africa, US | Open, Classified |  |  | 3319
Z-0348-2018 | 78702 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ECG Out Cable - Heartstart MRx |  |  | July 17, 2017 | 2017 | Philips Electronics North America Corporation | Issue with ECG out cables. When a Philips monitor/defibrillator is receiving an ECG signal from an auxiliary bedside monitor via a sync cable or ECG out cable, the following can occur if the monitor/defibrillator experiences interference from electrical fast transients (EFTs):"On the HeartStart MRx and HeartStart XL, EFT noise may be misinterpreted as an R-wave."On the HeartStart XL+, EFT noise can disable ECG monitoring and potentially interrupt demand mode pacing*. *Note: It is contrary to the XL+ Instructions for Use to perform demand mode pacing while using the ECG out cable or obtaining the ECG signal from a bedside monitor. The XL+ Instructions for Use includes the following warning: When pacing in Demand Mode, the ECG cable from the patient must be directly connected to the HeartStart XL+. If the user follows this warning, this problem cannot occur on the XL+. | Device Design | Philips is asking customers to follow the Action to be Taken by Customer/User section of the UrgentMedical Device Correction notification/Field Safety Notice:Customers should remove affected sync cables and ECG out cables from service and destroy them, documenting such action on the attached reply form, which should be returned to Philips. Customers should retain a copy of the reply form.The customer should follow the Instructions for Use that describe the recommended approach for performing cardioversion with the XL+ and MRx monitor/defibrillators. The ECG trunk cable can be disconnected from the Philips bedside monitor and connected to the monitor/defibrillator before delivering therapy; this is the preferred source of an ECG waveform for synchronization. |  | 15811 | 15811 | 1067 US customers163 Canada, 917 International | Terminated | July 12, 2018 | 360 | 15811
Z-0350-2018 | 76847 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Fellows "Mighty 8" Power strip (Part Number 04175-0008) and TrippLite Power strip (PS-415-HGULTRA - Part Number 13-00106-00) are two components sold as part of the Merge Hemo System. |  |  | April 06, 2017 | 2017 | Merge Healthcare, Inc. | Merge is conducting a hardware recall (replacement) of Fellows "Mighty 8" Power Strip and TrippLite Power Strip due to overheat of input resistor causing the power strip to smoke. | Component design/selection | Merge began notify customers via an Urgent Medical Device Recall letter on 4/6/2017 via mail. The letter identified affected product, stated the issue, and informed customers that Merge would replace affected product. A response form is requested to be returned to Merge. Questions can be sent via email to recall@merge.com |  | 1150 | 1,150 power strips | Nationwide. | Open, Classified |  |  | 1150
Z-0349-2018 | 78725 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona Partial Knee System, Size 8, Finishing Guide |  |  | July 05, 2017 | 2017 | Zimmer Biomet, Inc. | Finishing guide does not adequately account for shape differences present in Size 8 of Persona Partial Knee System, potentially resulting in additional surgery time or sensation of tightness during knee flexion. | Device Design | Customers were notified via FedEx mail and email on approximately 07/05/2017. Instructions included to notify appropriate personnel, complete and return the Acknowledgement Form, return affected product, and provide contact information of additional hospitals or surgeons that used the affected product. |  | 158 | 158 devices | Distributed domestically to AZ, CA, CO, ID, IL, IN, MA, MD, MI, NJ, NM, PA, SC, TX, WA, WI. Distributed internationally to Belgium, France, Germany, Italy, Japan, Netherlands, Spain, Sweden, United Kingdom. | Open, Classified |  |  | 158
Z-0353-2018 | 78926 | Yes | MVK |  |  | Cardiovascular | PMA | LifeVest Wearable Defibrillator Model 4000, Product Number 10A0988-A01. |  |  | September 12, 2017 | 2017 | Zoll Manufacturing Corp. | Incorrect service code for properly catching critical defects during self-check. Potential for defibrillation shock failure | Other | On September 14, 2017, ZOLL Manufacturing Corporation issued a patient safety alert notice dated September 12, 2017 to all active patients via courier service. As of November 2, 2017, a copy of this same patient Safety Alert notice is included with all current device shipments. The purpose of the communication is to revise the training patients received about the "Call for service" message. In certain cases, a "Call for service - Message Code 102" could mean that your LifeVest may not be able to deliver therapy if you need it. Users should call ZOLL immediately for a replacement LifeVest if a "Call for service - Message Code 102" appears on the screen. A replacement will be provided within 24 hours. Patients are instructed to continue using the LifeVest as prescribed by their physician. On January 8, 2019, Zoll Manufacturing Corporation issued Urgent Medical Device Correction notices to customers informing all that FDA approved a new software version for the ZOLL LifeVest 4000 that provides enhanced prominence and persistence of patient alerts when a "Call for service---Message Code 102" exists, instructing patients to call ZOLL immediately at 1-800-543-3267 for a replacement LifeVest. All new patients will receive a LifeVest with this new software revision. In addition, over the next 3.5 months, ZOLL will contact all current patients and exchange their monitor for one with this new software. Patients should continue to use the LifeVest as prescribed by their physician during the software rollout and no physician action is required. Customers with questions or concerns, please feel free to contact your ZOLL LifeVest representative or ZOLL Technical Services at 1-800-543-3267. |  | 33670 | 33,670 units | US Nationwide | Open, Classified |  |  | 33670
Z-0351-2018 | 78777 | No | OIJ | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(K) Exempt | Affirm Lateral Arm Upright Biopsy Accessory. Product Number:ASY-09880 |  |  | July 07, 2017 | 2017 | Hologic, Inc. | The firm is implementing a product correction involving its AffirmÂż Lateral Arm Upright Biopsy Accessory, ASY-09880. The correction has been initiated due to a component (the Left-hand Guide or the Blue Needle Guide) issue that could impact needle alignment for all the left-approach lateral biopsy procedures. This may have an impact on the position of the needle within the breast, and thus on the aspiration of the intended tissue. This impact is case-dependent and detectable when imagining is used to confirm needle position and/or tissue acquisition during and after a procedure. There are no known instances of illness or injury. | Component change control | Letter |  | 12 | 12 (U.S.) | US and OUS | Terminated | April 04, 2018 | 271 | 12
Z-0354-2018 | 78787 | Yes | LWP |  |  | Cardiovascular | PMA | ACCOLADE SR Pacemaker |  |  | December 07, 2017 | 2017 | Boston Scientific Corporation | Boston Scientific has received reports of intermittent over-sensing of the Minute Ventilation (MV) sensor signal with certain Boston Scientific pacemaker and cardiac resynchronization therapy pacemaker systems (pacemakers). | Device Design | On December 19, 2017 the firm sent a letter to all their medical consignees stating the following:December 2017Dear Doctor,Boston Scientific has received reports of intermittent oversensing of the Minute Ventilation (MV) sensor signal with certain Boston Scientific pacemaker and cardiac resynchronization therapy pacemaker systems (pacemakers). MV sensor signal oversensing may cause pre-syncope or syncope due to periods of pacing inhibition. This MV behavior may occur with any manufacturers pacing lead system, but Boston Scientific has determined it to be more likely for affected Boston Scientific pacemakers using Medtronic or Abbott/St. Jude(Abbott) leads implanted in either the right atrium (RA) or right ventricle (RV).Boston Scientific is actively developing a software update designed to automatically detect and resolve this MV sensor signal oversensing behavior. We anticipate submitting the software update to Regulatory Agencies inMarch 2018 and pending approval, will release it as soon as possible thereafter. Until this software update is available, Boston Scientific has additional re commendations to mitigate this risk for affected pacemaker systems.Root Cause InvestigationThe MV sensor in Boston Scientific pacemakers can be used for RightRate" (rate adaptive pacing), Respiratory Rate Trend, or AP Scan"1. When the RA/RV pacing leads and lead terminal connections are operating asintended, the MV sensor signal is appropriately filtered and therefore is not detected by the pacemaker or displayed on electrograms (EGMs). However, intermittency related to the lead or pacemaker-lead connection2 hasthe potential to create a transient high impedance condition. A high impedance condition may subsequently alter the MV sensor signal such that it becomes visible on EGMs and potentially subject to oversensing on the RA orRV channels. For a technical description of the Boston Scientifics MV sensor, please refer to Appendix A.Engineering analysis and |  |  |  | AK, AL, AR, AZ, CA , CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, and Puerto RicoWorldwide | Open, Classified
Z-0378-2018 | 78760 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Brand Name: |  |  | June 29, 2017 | 2017 | Roche Diagnostics Corporation | Roche has confirmed that a possible sample mismatch issue may occur on the MODULAR ANALYTICS E 170 module, cobas e 601 analyzer, or cobas e 602 analyzer due to a software limitation. | Software design | Roche sent an Urgent Medical Device Correction (UMDC) sent via UPS Ground (receipt required) dated July 6, 2017, and an amended UMDC on August 31 ,2017. All consignees will be notified by Urgent Medical Device Correction sent via UPS Ground (receipt required).Customers/consignees were instructed the following: If the Tip/Cup pick up error alarm is issued then followed by the Abnormal L2-line Movement alarm on your MODULAR ANALYTICS E 170 module, cobas e 601 analyzer, or cobas e 602 analyzer, after which the system goes into Sampling Stop, follow the steps in this UMDC and the appropriate How to Identify Potentially Mismatched Sample Results within the enclosed document. For further questions, please call (317) 521-4343. |  | 2517 | 2517 | USA (nationwide) Distribution | Open, Classified |  |  | 2517
Z-0496-2018 | 78442 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Hemosphere System with the following three components: HEM1 Advanced Monitor, HEMSGM10 (Swan-Ganz Module), and HEMOXSC100 (Oximetry Smart Cable). |  |  | October 02, 2017 | 2017 | Edwards Lifesciences, LLC | Pre-procedural issues related to software defects | Device Design | Consignees were sent recall notification letters on approximately 10/11/2017. The letter identified the affected product, problem and actions to be taken. instructions included coordination with the respective Edwards representative to schedule the software upgrade. |  | 198 | 198 devices | US Nationwide Distributed in the states of AZ, CA, CO, GA, IL, MO, NM, OK, OR, SC, TX | Open, Classified |  |  | 198
Z-0385-2018 | 78891 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Aspire HD Plus (FDR MS-2500), also known as Aspire HD Plus F and Aspire HD-s (FDR MS-2000), |  |  | July 14, 2017 | 2017 | Fujifilm Medical Systems U.S.A., Inc. | There was a case that the swivel arm went down unintentionally when it was moved up/down. | Employee error | FUJIFILM sent an Urgent: Medical Device Correction Letter dated July 14, 2017, accompanied by a Customer Feedback form to customers.PLEASE READ AND FOLLOW THE INSTRUCTIONSPlease refer to the following page which provides the details of the problem and instructions for actions to be taken. Please follow the instructions in the ACTIONS TO BE TAKEN BY CUSTOMER/USER section.We sincerely regret the inconvenience that this may cause you. FUJIFILM is committed to providing products and services of the highest quality. Your satisfaction with FUJIFILM products and with our response to this issue is very important to us.If you have any questions about this matter, please contact your local FUJIFILM office. For further questions, please call (360) 356-6821. |  | 20 | 20 US | US Distribution and Internationally to Mexico. | Terminated | July 20, 2018 | 371 | 36
Z-0381-2018 | 78483 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Unifine Pentips 5mm x 31G pen needles, packaged in 100 ct pen needles. |  |  | November 06, 2017 | 2017 | Owen Mumford USA, Inc. | Device may potentially be compromised by water in end user packaging. Exposure to liquid may compromise sterility. This may appear as discoloring or yellowing of the device's protective seal. | Environmental control | Owen Mumford sent an Urgent Medical Device Recall Notification letter dated November 6, 2017 via email. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to check inventory for affected product, contact the Customer Service Department to arrange return of affected product, complete and return the response form and to either provide a customer list for Owen and Mumford to notify their customers or confirmation that they will notify their customers if further distributed. |  | 2660 | 2,660 cartons (266,000 pen needles) | US Nationwide Distribution in the states of AL, AZ, CA, CO, FL, GA, HI, IA, IL, MA, MI, MN, MO, MS, NC, NH NJ, NY, OH, OK, PA, TN, NX, UT, VA, VT, WA, WV | Open, Classified |  |  | 3368
Z-0380-2018 | 78900 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation |  |  | July 17, 2017 | 2017 | RAYSEARCH LABORATORIES AB | For the proton QA preparation module in RayStation 6 (including SP1 and SP2), if the snout position or gap is modified in the QA module, the dose in the QA module may be computed for a different setup than what is used for QA measurements. This could result in an incorrect dosage being provided to a patient. | Device Design | Ray Search sent a Field Safety Notice letter dated July 14, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter included a response form with instructions for workarounds until the software is corrected. |  | 117 | 117 units | US Nationwide Distribution - FL, WA, AZ, TN, MD, TX, FL, MI, Washington D.C. | Terminated | August 29, 2018 | 408 | 117
Z-0383-2018 | 78802 | No | KGN |  |  | General &amp; Plastic Surgery | 510(k) | PriMatrixÂż is an acellular dermal tissue matrix derived from fetal bovine dermis. The device is supplied sterile in a variety of sizes to be trimmed by the surgeon to meet the individual patient's needs. |  |  | June 15, 2017 | 2017 | TEI Biosciences | During Final Assembly Operation, it was noticed that some of the PriMatrix Instructions for Use (IFUs) from Lot RM1608011 Part 520-002v10 contained extra pages and/or missing pages:" 22 IFUs contained duplicate pages 15 - 26 &amp; 51 - 62 and were also missing pages 27 - 50 " 3 IFUs were missing pages 27 - 50 " 1 IFU was missing pages 1 - 2 &amp; 75 - 76" 1 IFU was missing pages 15 - 62While pages 27  50, 75  76, and 15  62 contain instructions written in a language other than English, pages 1  2 contain information regarding the indications for use, contraindications, device description, warning and precautions, and potential complications of the device. | Error in labeling | Integra LifeSciences Corporation notified (all U.S.) customers via Letter/Medical Device Correction Field Safety Notice dated June 15, 2017. Letter delivered by courier service, facsimile or email July 15, 2017. |  | 4081 | 4081 units | U.S. | Terminated | March 15, 2018 | 273 | 4081
Z-0384-2018 | 77289 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Small Particle Aerosol Generator Model SPAG-2 6000 Series (SPAG-2). The SPAG-2 medical device is a pneumatic flow system indicated for the administration of Virazole (ribavirin) aerosol. Driven by a pressurized gas (air or blended oxygen) and uses a nebulizer to generate a fine aerosol of hydrated Virazole. |  |  | March 13, 2017 | 2017 | Valeant Pharmacueticals International | SPAG-2 unit nebulizers from Lot J0571670RA were tested at Next Breath for particle sizing under the validated method, NB-TM-261, In November 2016. The particle size specification set in NB-SPEC-039 for Dv10, Dv50, &amp; Dv90 are a mean value of no more than 3.0Âżm, 5.5 Âżm, and 15Âżm respectively. Initial testing of the nebulizers yielded four units that delivered particle sizes that did not meet the specification (NB-SPEC-039). The out of specification (OOS) nebulizers were cleaned per the approved SPAG-2 procedure, SPE-000127228 Rev 2, to remove any product residue and were then re-evaluated. Secondary testing yielded that three of the four nebulizers still did not meet the particle size specification. It was concluded that the 3 nebulizers from SPAG-2 Lot J0571670RA failed to meet (SPE-000127228 Rev 2) did not resolve the high droplet size distribution values observed. | Device Design | On April 4, 2017 Valeant Pharmaceuticals North America distributed Medical Device Voluntary Product Recall and Acknowledgement Forms to their customers via USPS. Valeant Pharmaceuticals are directly contacting those customers who have been shipped affected SPAG-2 devices to arrange pick up and testing of the nebulizer component. Only half of the nebulizers will be picked-up for particle-size testing at each account to prevent shortages of VIRAZOLE, which can only be used with the SPAG-2 device. Nebulizers that pass testing will be cleaned and returned. Nebulizers that do not pass testing will be disposed, and new nebulizers will be distributed to facilities once they become available. Customers are encouraged to complete and return the Acknowledgement Form via email to: valeant3012@stericycle.com. |  | 1228 | 418 SPAG-2 and 808 Nebulizers | US Nationwide &amp; Canada | Open, Classified |  |  | 1228
Z-0603-2018 | 78064 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Exactech GPS Total Shoulder Application 3.2mm Vix Bit Orthopedic surgical tool |  |  | July 11, 2017 | 2017 | Exactech, Inc. | Vix Bit may fracture during use. | Nonconforming Material/Component | Customers were notified on approximately 07/11/2017. Instructions included cease distribution of the affected product, notify customers if further distributed, identify and quarantine any product in inventory and complete and return the Recall Inventory Response Form. For further questions, please call (800) 392-2832. |  | 41 | 41 devices | Worldwide Distribution: US states of Florida and California, Australia, France, Spain, and United Kingdom. | Open, Classified |  |  | 41
Z-0538-2018 | 78936 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Alteon HA Femoral Stem, Press-Fit, Extended Offset, Cementless, HA Coated, Collared, Size 5 |  |  | October 09, 2017 | 2017 | Exactech, Inc. | Wrong device description of the label identifying the implant | Error in labeling | Exactech sent an Urgent Medical Device Recall letter dated October 9, 2017, informing them of the potential mislabel. The notice requested consignees: Immediately cease distribution/use of this product; Extend information to accounts in possession of the product; Identify and quarantine any subject devices in their inventory; Complete and return Recall Acknowledgement and Recall Inventory Response Forms. For further questions, please call (800) 392-2832. |  | 34 | 34 | US Distribution to the states of : AL, CA, CO, FL, MD. and NY | Open, Classified |  |  | 34
Z-0923-2018 | 78838 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT c16000 Cuvette Segment, Part #09D32-05/ 03L77, a component to the ARCHITECT Clinical Chemistry system. There are 15 segments with each cuvette segment holding 11 cuvette pairs (22 cuvettes) for a total of 165 cuvettes pairs or a total of 330 cuvettes (22 cuvettes x 15 cuvette segments) per reaction carousel. The firm name on the label is Abbott Laboratories, Abbott Park, IL. |  |  | December 08, 2017 | 2017 | Abbott Laboratories, Inc | There is a potential to generate falsely depressed patient results in the cuvettes adjacent to the detached segment post for any assay run on the c4000, c8000, and c16000 instrument. | Device Design | The firm, Abbott Laboratories, issued a "Product Correction" letter dated 12/8/2017 via overnight mail on 12/8/2017 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to adhere to the new instructions listed in Appendix A to avoid damaging the cuvette segments; If a cuvette segment with a detached bottom is identified, replace it before performing any additional testing on your ARCHITECT Clinical Chemistry System; If you have forwarded the product listed above to other laboratories, please inform them of this Product Correction and provide to them a copy of this letter and retain this letter for your laboratory records.The procedures will be updated in a future version of the ARCHITECT Operations Manual.If you or any of the health care providers you serve have any questions regarding this information, U.S. Customers please contact Customer Service at 18774ABBOTT (available 24 hours a day, 7 days a week). Customers outside the U.S., please contact your local area Customer Service. |  | 17560 | 1,560 ARCHITECT c16000 Systems | Worldwide Distribution: US (nationwide) including Puerto Rico and countries of: Afghanistan, Albania, Algeria, Angola, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahamas, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Brunei, Bulgaria, Burkina Faso, Cameroon, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Curacao, Cyrus, Czech Republic, Dem Rep of The Congo, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, Gaza &amp; Jericho, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Ivory Coast, Jamaica, Jordan, Kazakhstan, Kenya, Kuwait, Kyrgyzstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malawi, Malaysia, Maldives, Mali, Mauritania, Mexico, Moldova, Montenegro, Montserrat, Morocco, Netherlands, New Zealand, Nigeria, North Yemen, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia (CIS), Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, St. Lucia, St. Vincent, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, UAE, Uganda, United Kingdom, Uruguay, Venezuela, Vietnam, Zambia and Zimbabwe. . | Completed |  |  | 73102
Z-0387-2018 | 78752 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | Phadia 1000 Instrument, Article Number 12-3800-01. |  |  | July 05, 2017 | 2017 | Phadia US Inc | The "Retry" command does not function properly which could cause a shortage of Wash and Rinse solution and affect assay performance and test results. | Device Design | The firm sent letters with a return response form to customers. The firm instructed customers to immediately choose the "stop" command if presented with the 7-101 Liquid Sensor Error. Results from the assay should not be approved. A revision of the current software will have a mandatory installation soon. |  | 70 | 70 units | Nationwide | Terminated | October 04, 2018 | 456 | 70
Z-0415-2018 | 78903 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Equinoxe Cage, Glenoid, Alpha, Medium |  |  | August 31, 2017 | 2017 | Exactech, Inc. | Potentially mislabeled. | Under Investigation by firm | Consignees were sent a Recall Notice dated 8/31/2017 by mail informing them of the potential mislabel. The notice requested consignees: Immediately cease distribution/use of this product; Extend information to accounts in possession of the product; Identify and quarantine any subject devices in their inventory; Complete and return Recall Acknowledgement and Recall Inventory Response Forms. For further questions please call (800) 392-2832. |  | 30 | 30 | Worldwide distribution in the states of OH, WI, AZ, GA, VA, CA, OK, MN, TN, ME, CO, SC, FL, AL, and Puerto Ricoand country of Switzerland | Open, Classified |  |  | 30
Z-0396-2018 | 78826 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | smith&nephew; Mini-Fragment Plating System, VLP(R) Mini-Mod Small Bone Plating System, EVOS(R) Mini Plating system, 2.0MM STRAIGHT PLATE, 8 HOLE VLP(R) MIN-MOD, REF 74442024, QTY: (1), STERILE R |  |  | December 11, 2017 | 2017 | Smith &amp; Nephew, Inc. | Devices, incorrectly labeled as a VLP MINI-MOD 2.0MM STRAIGHT PLATE 8 HOLE, contained a VLP MINI-MOD 2.0MM STOUT PLATE 8 HOLE and vice versa | Mixed-up of materials/components | The firm, Smith &amp; Nephew, Inc., sent an "Urgent Medical Device Recall Notice" letter dated December 11, 2017 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect your inventory, locate unused devices and quarantine immediately; if you are a distributor, you must notify your customers of the field action and ensure that these actions are carried out; return product, and complete and return the response form by email to FieldActions@smith-nephew.com or fax to 901-566-7975 even if you have not product to return. If you have any questions or concerns regarding this recall please contact FieldActions@smith-nephew.com or call 901-566-7084. |  | 2 | 2 units | US Distribution to states of: MI, GA, and OH. | Terminated | June 22, 2018 | 193 | 2
Z-0395-2018 | 78036 | No | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Measurement Specialties Reusable Temperature Probe Autoclavable. Sold as Esophageal/Rectal (Adult), Esophageal/Rectal (Pediatric), and Skin (Adult) |  |  | March 13, 2017 | 2017 | Measurement Specialties Inc | The reusable temperature probes are provided with instructions for sterilization that have not been proven through a validation study. | Other | Customers were notified via mail or email on March 13, 2017. Additional customers were later identified and notified on June 14, 2017. Instructions included to disregard the sterilization instructions listed in the IFU (A08411F)originally provided with probes and instead utilize the updated IFU 08411G which contains validated sterilization instructions. They were also instructed to forward the recall notification to within their firm and to all organizations if further distributed. For further questions, please call (763) 316-4075. |  | 6076 | 6,076 devices total | USA (nationwide) distribution to the states of : AL, CA, KY, IL, KS, MN, NY, OH, TX, VA and WI.Distributed internationally to: Australia, Canada, China, England, Germany, India, Ireland, Israel, Italy, Japan, Malaysia, Netherlands, Singapore, South Africa, South Korea, Taiwan, Vietnam | Open, Classified |  |  | 6076
Z-0400-2018 | 77964 | Yes | NVZ |  |  | Cardiovascular | PMA | MyCareLink Smart Patient Monitors. |  |  | July 14, 2017 | 2017 | Medtronic Inc., Cardiac Rhythm and Heart Failure | Patients monitored on two (2) or more implanted Medtronic heart devices in the Medtronic CareLink Network may have potential impact on the ability to remotely monitor the patient's heart devices. Potential impacts could lead to missed CareAlert notifications or device reports. | Error in labeling | Customers were notified via letter on 07/14/2017. Medtronic communicated that an updated Clinician Instructions for Use has been released with the following updated information: In patients with more than one Medtronic implanted hear device, device data may not be able to transmit from any of the implanted devices to the physician via the Monitor; The patient could potentially transmit device data unintentionally from one of the implanted heart devices when attempting to transmit from another device implanted in the same patient; and a Medtronic Programmer can be used to obtain data from each implanted heart device. Instructions included to ensure all appropriate personnel are notified of this correction. |  | 1519984 | 1,519,984 enrollments total | Domestic: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, Internationally: Austria, Belgium, Brunei, China, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Kazakhstan, Latvia, Lithuania, Malaysia, Malta, Netherlands, Norway, Philippines, Poland, Portugal, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE, United Kingdom | Open, Classified |  |  | 1519984
Z-0393-2018 | 78808 | No | NSU | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BD MAX" UVE Specimen Collection Kit. |  |  | July 18, 2017 | 2017 | Becton Dickinson &amp; Co. | BD recently notified customers in July 2017 regarding an increased rate of unresolved (UNR) and indeterminate (IND) results for the BD MAXTM Vaginal Panel, experienced by some customers. A non-reportable result may lead to repeat of the test and/or re-collection of a specimen. | Use error | Customers were notified via UPS on 07/18/2017. The recall for BD MAX" Vaginal Panel and BD MAX" UVE Specimen Collection Kit was initiated on July 18, 2017 and expanded to include BD MAX" CT/GC/TV on August 11, 2017. The recall expansion letter to customers was dated 08/14/2017. |  |  |  | US and OUS. | Terminated | July 09, 2018 | 356
Z-0388-2018 | 78852 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT |  |  | November 10, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | Philips has identified that certain fasteners cannot be confirmed to have been torqued according to specification. If the fasteners were not torqued as specified and a failure occurs while the system is in use, the parts secured by the fasteners may become loose, which may have a potential for harm. | Under Investigation by firm | An Urgent - Field Safety Notice Medical Device Correction letter, dated November 1, 2017, was sent to customers. The letter identified the affected device, described the issue, and discussed the hazard involved. Users are to cease use immediately and contact their Philips Field Service Engineer, if any unusual noises are heard within the gantry, or if the patient support tabletop is in free float. Philips Field Service Engineers will correct the system by replacing the affected fasteners. If further info or support concerning this issue is needed, customers can contact their local Philips rep or the Customer Care Solutions Center at 1-800-722-9377. |  | 13 | 13 | Worldwide Distribution -- US, Argentina, China, Republic of Korea, Russian Federation, Slovakia, Turkey, and United Kingdom.4 in US, (1 military), 14 foreign accounts | Open, Classified |  |  | 18
Z-0401-2018 | 78594 | Yes | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | CP5 Centrifugal Pump, made up of the drive unit, the emergency system and the pump control panel. It used with an Electrical Remote-Controlled (ERC) tubing clamp and a bubble sensor. |  |  | October 30, 2017 | 2017 | Sorin Group Deutschland GmbH | Possibility that the automatic line closure by the Electronic Remote-Controlled (ERC) clamp after detection of an air bubble could take longer than is intended by design in the CP5 system. | Device Design | Consignees were notified of the recall via USPS trackable mail, beginning on 10/30/2017. The letter informed customers Instructions included to continue following the Instructions for use, including use of an arterial filter on the arterial line, and that a LivaNova representative will contact them to schedule an upgrade to their device. |  | 1318 | 1,318 units | Distributed domestically to Distributed internationally to Algeria, Australia, Austria, Azerbaijan, Bahrain, Belarus, Belgium, Canada, China, Colombia, Croatia, Czech Republic, Ecuador, Estonia, Finland, France, Germany, Hong Kong, Iraq, Italy, Japan, Jordan, Kuwait, Libya, Lithuania, Luxembourg, Malaysia, Morocco, Norway, Palestinian Territory, Panama, Peru, Poland, Qatar, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, The Netherlands, Tunisia, United Kingdom, Vietnam | Open, Classified |  |  | 1318
Z-0405-2018 | 78911 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging version V31 (model # 10014063) is a Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images, including digital mammography images. It supports the physician in diagnosis and treatment planning. |  |  | July 25, 2017 | 2017 | Siemens Medical Solutions USA, Inc. | Siemens is releasing a letter to inform about potential data loss relevant to diagnosis. | Software design | Siemens sent an Urgent Customer Advisory Notice Letter to all customers. All versions before version VB36E_HF01 can be configured to behave like the version VB36E_HF01. Please contact your local CSE to configure this setting and to evaluate an eventual minor performance impact to the system.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the devices new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions, please call (610) 448-6461. |  | 36 | 36 units. | USA (nationwide) Distribution to the states of: MA, CA, NC, WI, NY, NE, PA, AL, NJ, MO, MI, TX, MS, KS, LA, CT, MD, LA, CO, OK, MN and TN. | Terminated | July 23, 2018 | 363 | 36
Z-0404-2018 | 78908 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | daVinci X EndoWrist(R) Stapler 45, REF 470298 |  |  | September 19, 2017 | 2017 | Intuitive Surgical, Inc. | Specific da Vinci Xi EndoWrist Stapler 45 instruments are potentially impacted by a variation in the manufacturing process. In affected instruments, there is a possibility that the Stapler Release Kit (SRK) Wrench interface feature accessed through release hole 2 is manufactured to an incorrect dimension, which may prevent the SRK Wrench from being able to manually unclamp the instrument. | Process control | The firm initiated their recall by email on 09/19/2017 and followed with a letter delivered by FedEx on 09/20/2017. The notice requested the following actions:"Please take the following actions to ensure all affected personnel are fully informed of this Field Safety Notice. Forward this letter to your Risk Manager, OR Director, Purchasing, Biomedical Engineering staff, and other members of your medical staff who are involved with da Vinci procedures.1.Using the lot numbers listed above, please locate and return all affected Staplers at your site via the standard RMA process.a. Credit will be provided for remaining uses on affected instruments returned to Intuitive Surgical.2.Please log into the da Vinci Online Community field action resource to read and/or complete any requested actions related to this issue.a. https://www.davincisurgerycommunity.com/3.If you cannot access the da Vinci Online Community field action resource, please complete the attached Acknowledgement Form and return it via fax or email to Intuitive Surgical per the instructions contained in the Acknowledgement Form.4.Please retain a copy of this letter and a copy of the Acknowledgement Form for your files." |  | 60 | 60 units | US Nationwide Distribution | Open, Classified |  |  | 60
Z-0402-2018 | 78830 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT SP - Model 728311 |  |  | July 11, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | Numerous software issues with Brilliance iCT 4.1.6 version. Issues affect scan start position, data acquisition, exam stop, tube overheat, memory overflow, failure to give error message, length of acquisition, image volume and other functions. | Software design | A Customer Information Letter dated June 29, 2017, was sent to US consignees via certified mail. Philips Key Market representatives will distribute the letter internationally, based on local regulations. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 559 | 559 total | Worldwide distribution, including US nationwide. | Open, Classified |  |  | 559
Z-0410-2018 | 78349 | No | NEU | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TUMARK Q, Length 12 cm, Diameter 18G, GTIN 04250195603081, REF 271500, STERILE |  |  | May 22, 2017 | 2017 | Somatex Medical Technologies GmbH | There is a potential risk of puncture to the packaging material which can compromise the devices sterility, making it non-sterile. | Device Design | On May 22, 2017, Somatex Medical Technologies distributed notices to their customers advising them of the reason for recall. Customers are advised to inspect stock, quarantine and not use affected product. Products will be exchanged free of charge and a replacement delivery will be arranged promptly. All customers are expected to complete and return the response form via fax. |  |  |  | NY - Only one US distributor | Terminated | August 30, 2018 | 465
Z-0408-2018 | 78672 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | smith&nephew;, 2.4 MM, 6 HOLE, EVOS, FLEX PLATE, SST Plate, REF CATALOG NUMBER 72442360N, NON-STERILE, Rx only, (01)00885556456446(10), (11)171116 |  |  | November 29, 2017 | 2017 | Smith &amp; Nephew, Inc. | One lot of EVOS 2.4MM Flex Plate 6 Hole packaging contained templates. | Employee error | Smith &amp; Nephew Inc. initiated its recall by email on November 29, 2017. The consignee was directed to return the recalled product. For further questions, please call (901) 396-2121. |  | 12 | 12 | US Distribution to the state of OR, and Internationally to Colombia. | Open, Classified |  |  | 12
Z-0406-2018 | 78429 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Radial Osteotome Blade, REF 71369314. |  |  | October 23, 2017 | 2017 | Smith &amp; Nephew, Inc. | The expiration date provided on the label is incorrect. | Process change control | The firm initiated their recall by an Urgent Medical Device Recall Notice on 10/23/2017. The letter identified the affected device and the reason for the recall, as well as the potential risk. The firm instructed users to please inspect their inventory and locate any unused devices from the list of product and batch numbers provided in the Notice. Affected devices must be quarantined immediately. Distributors, must notify their customers of the field action and ensure that the actions are carried out. The attached form should be completed and returned. Questions or concerns can be directed to FieldActions@smith-nephew.com. |  | 9 | 9 units | Worldwide Distribution -- US, to the states of TN and NJ; and, Australia. | Terminated | August 24, 2018 | 305 | 9
Z-0411-2018 | 78835 | No | FLD | Sorbent hemoperfusion system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MARS Treatment Kit, Use only with MARS(R) Monitor 1TC, GAMBRO, Product Code 800541 |  |  | December 21, 2017 | 2017 | Baxter Healthcare Corporation | Baxter has received customer complaints regarding leakage in the albumin circuit. The leakage was caused by an inadequate adhesive connection of the tubing to the Hansen connector of the MARS Tube Set which is part of the MARS Treatment Kit. | Nonconforming Material/Component | The firm initiated their recall on 12/21/2017 by letter. The recall notice stated the following actions to be taken by the customers:"1. Locate and remove all affected product from your facility. The product codeand lot number can be found on the individual MARS Treatment Kit box.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.Baxter Healthcare Center for Service can be reached at 888-229-0001between the hours of 7:00 am and 6:00 pm Central Time, Monday throughFriday. Please have your Baxter 8-digit ship-to account number, product code,lot number, and quantity of product to be returned ready when calling.3. If you purchased this product directly from Baxter, complete the enclosedBaxter customer reply form and return it to Baxter by faxing it to 224-270-5457 or scanning and e-mailing it to fca@baxter.com, even if you do not haveany inventory." |  | 186 | 186 units | US, Germany, Sweden, Italy, Belgium, Portugal, Spain, France,Russia, Poland, Turkey, China, Thailand, Taiwan, Hong Kong, Malaysia, Korea, and Mexico. | Open, Classified |  |  | 186
Z-0407-2018 | 78883 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | PURE OXYGEN CONCENTRATOR, CHAD Drive, Models CH5000 and CH5000S |  |  | October 09, 2017 | 2017 | Inovo, Inc | Post-market surveillance for Pure Stationary Oxygen Concentrators model CHSOOO and CHSOOOSdemonstrate an aggregate complaint/service rate greater than 35% for low oxygen purity (outputoxygen below the published specification). | Under Investigation by firm | The firm initiated their recall by letter on 10/09/2017. The letter stated the following:"Actions to be Taken:Based on our distribution records, your firm has received one or more Pure Stationary Oxygen Concentrators. We respectfully request that you take the following actions: 1. Contact all patients and customers that currently have Pure Stationary Oxygen Concentrators in their homes or facilities and inform them of this recall. Retrieve the oxygen concentrators from their homes or facilities.2. Destroy all Pure Stationary Oxygen Concentrator(s) in accordance with the following steps: a. Assure that the unit is not plugged into an electrical outlet.b. Cut the power cord of each oxygen concentrator.c. Take a digital photo of each oxygen concentrator that depicts both the power cord and the serial number of the depicted unit. Note that the serial number is located on a label on the back of the unit.d. Dispose of the oxygen concentrator in accordance with local ordinances.e. Record the serial number of each unit that is destroyed, sign and date the enclosed Recall Return Response Card.3. Email or fax the completed response form and photos documenting destruction to:Email: lnovo4209@stericycle.comFax: 1-877-497-25594. Sales of the Pure Stationary Oxygen Concentrator have been discontinued. Pure Stationary Oxygen Concentrators which have been destroyed will be replaced with a DeVilbiss 525DS Oxygen Concentrator in order to minimize any interruption in patient's supplemental oxygen therapy.5. Pure Concentrators that have been returned to lnovo for repair will not be returned and will be replaced with a DeVilbiss 525DS Oxygen Concentrator.Other Information:Should you require any additional information about this recall action, please call 1-888-679-5132, Monday-Friday 8am-Spm Eastern." |  | 13140 | 13,140 units | worldwide | Open, Classified |  |  | 13140
Z-0346-2018 | 78636 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Adult AirLife dual-limb, dual-heat circuit (Model AH280) with AirLife Humidification Chamber, Model Number AH290. |  |  | November 29, 2014 | 2014 | Vyaire Medical | Manufacturing error in humidification chamber may cause an overfill/overflow condition of the chamber. The may case a water back-up into the patient circuit resulting in an excessive amount of water entering the airway/lungs of the ventilated patient. | Under Investigation by firm | The firm, Vyaire Medical, sent an "URGENT RECALL NOTIFICATION" letter dated November 29th, 2017 to its customers. The recall letter described the product, problem and actions to be taken. The customers were instructed to do the following:"ALL CUSTOMERS:~ Inspect current inventory on-hand. A 100% physical inventory should immediately be performed to identify and remove affected lots of the AirLife Humidification Chamber(s) and Heated Breathing Circuit Kit(s) due to the identified potential patient safety risk.~ Destroy all affected product(s) in-stock in accordance with your facility's destruction protocol.~ If you are a healthcare facility or distributor that has further distributed affected product you must further notify your customers of this URGENT RECALL NOTIFICATION.CUSTOMERS THAT HAVE PURCHASED DIRECT FROM VYAIRE:~ Complete the Customer Response Form (Appendix 2) and return to GMB-GLBALFieldActions@CareFusion.com~ If you wish to receive credit of your in-stock product affected by this recall or to place an order for the replacement Model/Part Numbers (as referenced per #6. Per the FAQ's) please contact Vyaire Medical's Customer Service Department at 800.323.9088 (Option #1, followed by Option #1) Monday-Friday 8:00 am CST - 5:00 pm CST.*Please have your original PO number available when calling to expedite the credit process.CUSTOMERS THAT HAVE PURCHASED THROUGH A DISTRIBUTOR:~ Please work with your distributor to receive credit of your in-stock product affected by this recall.For any additional questions and support concerning this voluntary recall, please contact Clinical Risk Coordinator, at 827-757-0109 or lindy.schenning@vyaire.com. |  | 16670 | 16,670 units in total | US Nationwide Distribution | Open, Classified |  |  | 16670
Z-1288-2018 | 78568 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Southmedic Oxygen Masks with EtCO2 sampling line part number OM-2125-8, 25 masks per case. |  |  | October 17, 2017 | 2017 | Southmedic, Inc. | Potential for loose or missing CO2 diffusers which could result in a non-functioning mask. | Device Design | Customers were notified via letter on October 17, 2017. Instructions included checking inventory for affected devices and quarantining any found, complete and return the Voluntary Recall Notification Confirmation Form, and arrange for the return of any affected products. On February 14, 2018, the recall was expanded and customers were notified. Instructions were the same as the original notification, but updated the Part/Lot information. For further questions, please call 1 (800) 463-7146 |  | 8871 | 8,871 cases (25 masks per case) total | US Distribution to the state of California, and Internationally to Canada and Sweden | Open, Classified |  |  | 8871
Z-0532-2018 | 78989 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare Xhibit Telemetry Receiver, Model 96280, software version 1.1, physiological patient monitor (with arrhythmia detection or alarms) |  |  | January 17, 2018 | 2018 | Spacelabs Healthcare Inc | Spacelabs has received multiple complaints reporting telemetry beds dropping off the Spacelabs Central Station resulting in a loss of monitoring. | Under Investigation by firm | On January , 2018, a customer letter was sent via priority service, return receipt requested, to all U.S. customers. International consignees will be notified on January 24 , 2018, by email. The letter stated: " Please weigh the benefits versus the risks when deciding whether or not to continue to use your Exhibit Telemetry Receiver(s) until updated as set forth below. We recommend that you brief your staff regarding the failure mode listed above." Spacelabs Healthcare will contact you to schedule a convenient time for Spacelabs to update your Exhibit Telemetry Receiver(s) to software version 1.0.2 at no cost." For further questions, please call (800) 522-7025. |  | 12 | 12 systems | US Distribution to the states of : VA, NJ, MI, GA, and Internationally to France | Terminated | November 29, 2018 | 316 | 12
Z-0487-2018 | 78918 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE HEALTHCARE CARESCAPE Central Station (CSCS) software version 2.0.2 |  |  | January 11, 2018 | 2018 | GE Medical Systems Information Technologies, Inc. | CARESCAPE Central Station (CSCS) software version 2.0.2 units may experience unexpected NO COMM (No Communication) and network communication issues after boot-up or system restart. | Software in the Use Environment | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter and email to its consignees on 01/11/2017. The letter described the product, problem and actions to be taken. The consignees were instructed to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions and follow the safety instructions. GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.If you have any questions regarding this notification, please contact Technical Support at 1-800-558-7044 or your local Service Representative. |  | 1803 | 1803 units | Worldwide Distribution: US (nationwide) and countries of: Algeria, Australia, Austria, Canada, Costa Rica, Czech Republic, Denmark, Egypt, Finland, France, Germany, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Italy, Japan, Jordan, Korea, Republic Of, Kuwait, Lebanon, Mexico, Netherlands, New Zealand, Norway, Poland, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uganda, United Arab Emirates, United Kingdom, Uruguay, and Viet Nam. | Open, Classified |  |  | 1803
Z-0412-2018 | 78687 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SST TWIST DRILL, STANDARD, 1.6mm Dia. x 127mm, REF 8054-010, STERILE |  |  | May 22, 2017 | 2017 | MicroAire Surgical Instruments, LLC | One lot was found to contain both the correct twist drill (1.6mm) as well as an incorrect twist drill (2.0mm) | Labeling mix-ups | MicroAire Surgical Instruments, LLC distributed recall letters on May 22, 2017 MicroAire Surgical Instruments, LLC distributed recall letters to the customers who received the mislabeled twist. Letters will be sent by UPS to expedite the delivery. For further questions, please call (434) 975-8000 |  | 42 | 42 | US Distribution. | Terminated | August 29, 2018 | 464 | 42
Z-0413-2018 | 78861 | No | EKR | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Elements Obturation Unit (EOU) used with respective Buchanan Heat Pluggers |  |  | December 27, 2017 | 2017 | Ormco/Sybronendo | Buchanan Heat Pluggers will possibly overheat while used with the Elements Obturation Unit (EOU). | Under Investigation by firm | Sybron Endo sent an Urgent: Medical Device Recall letter dated December 27, 2017.Customers were informed that Sybron Endo was still evaluating the root cause and would provide additional instructions once the root cause is determined. Distributors were asked to cease distributing any Elements Obturation Units (EOU), confirm if any are in inventory, return any affected devices, provide the notification to customers if further distributed, return any affected devices received from customers, and complete and return the acknowledgement form. Customers were instructed to examine current inventory for affected product, check the heat plugger to see if the NiTi icon appears and/or the tip turns glowing red, if it does do not use the heat plugger and return it to Sybron Endo, and complete and return the acknowledgement form. For further questions, please call the Customer Care department at (800) 537-7123. |  | 25765 | 25,765 total | USA (nationwide) distributed to AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. Distributed internationally to Armenia, Australia, Canada, Chile, Ecuador, El Salvador, Guatemala, Hong Kong, India, Japan, Malaysia, Mexico, Netherlands, Pakistan, Panama, Philippines, Singapore, South Korea, Switzerland, Thailand, United Kingdom, Vietnam. | Open, Classified |  |  | 25765
Z-0417-2018 | 78058 | No | OUW | Pediatric position holder. | General Hospital | General Hospital | 510(k) | 9x15" Medium Utility Positioner with Cover |  |  | July 26, 2017 | 2017 | Molnlycke Health Care, Inc | Products may not meet specifications due to possible presence of metal shavings. | Under Investigation by firm | Customers were notified by letter via overnight mail on approximately 07/26/2017. Instructions include identify and quarantine affected product in inventory, contact Molnlycke to arrange return of affected product, and complete and return the Acknowledgement Form. |  | 204 | 204 units total | Distributed domestically to AR, CT, FL, GA, IA, IL, LA, MA, MI, MN, NC, NE, NY, OH, TX, VA, WI. | Open, Classified |  |  | 204
Z-0379-2018 | 78885 | No | RCE |  |  |  |  | Metrotom 800 (130kV CT scanner) |  |  | November 14, 2017 | 2017 | Carl Zeiss Metrology Inc | Due to a software error, it was observed that when the user closes the access door, the system will resume its scan cycle automatically, instead of the user initiating x-ray generation from the control panel. | Radiation Control for Health and Safety Act | Zeiss sent an Safety Notice Letter dated September 28, 2017, to the customers advising them of the recall. The affected device was identified as well as the issue involved. Customers are to complete and return the confirmation- waiver form. During the next scheduled service event, a ZEISS field service technician will update the Metrotom operating system. The corrections are to be made free of charge, and will be completed by March 2018.CDRH approves the CAP subject to the following conditions:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification. For further questions, please call (763) 203-4267. |  | 13 | 13 US | Worldwide Distribution - US, including the states of : WI, NC, GA, IL, AZ, FL, CA, NJ, and territory of PR; and the country of Canada. | Open, Classified |  |  | 13
Z-0418-2018 | 78608 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Triage Drugs of Abuse Plus TCA 25 Test Kit, Model Numbers 92000, 92590, 92591; 25 individually pouched devices, 1 bottle of |  |  | November 14, 2017 | 2017 | Alere San Diego, Inc. | The recalled lots have demonstrated unexpected false positive THC results. | Nonconforming Material/Component | All primary consignees will be contacted in letter provided via fax, email or direct mail. End users will be instructed to discard any remaining inventory they may have of the affected lots. Alere will provide credit to end users for any remaining unused inventory. |  | 2940 | 2940 units | Worldwide distribution - United States Nationwide and the countries of Japan, Canada | Open, Classified |  |  | 2940
Z-0419-2018 | 78886 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | MA3320-PM Premium Adult - SYK Mistral-Air Premium Blankets. MODEL #: MA3302-PM. CATALOG #(s): MA3302-PM |  |  | July 21, 2017 | 2017 | Stryker Medical Division of Stryker Corporation | It was identified that the Mistral-Air Premium blankets may be capable ofconducting electricity during a defibrillation or cardioversion therapy at higherenergy levels (70 Joules or more) when the blanket is applied with the reflectiveside toward the patient (upside down). | Component design/selection | Stryker notified customers of the recall by sending Customer Notification and Business Reply Forms, dated July 21, 2017. |  | 11160 | 11160 | U.S. | Terminated | March 22, 2018 | 244 | 11160
Z-0488-2018 | 78231 | No | EXS | Urine flow or volume measuring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BARD UROS AUTOMATED URINE OUPTUT AND TEMPERATURE MONITOR, MODEL BK10001M |  |  | September 29, 2017 | 2017 | C.R. Bard, Inc. | The possibility exists that use of the Uros Automated Urine Output and Temperature Monitor may lead to the loss of urine output measurement data and/or urine output measurement inaccuracies resulting in an impact on clinical treatment decisions. | Device Design | The firm, Bard Medical Division, sent an "URGENT: MEDICAL DEVICE PRODUCT RECALL" letter dated September 29, 2017 to their consignee. The letter described the product, problem and actions to be taken. The consignee was instructed to immediately examine your inventory and quarantine product; do not use or further distribute any affected product; complete and return the Recall &amp; Effectiveness Check Form by fax to 1-770-784-6460 or email a scanned copy to BMD.FieldAction@crbard.com, Attn: Recall Coordinator, even if you do not have any affected product and return the product. If you have affected product, please contact the BMD Customer Service via phone at 1-800-526-4455 option 5 and again option 5 or email BMD.FieldAction@crbard.com to obtain a Return Goods Authorization number (RGA #).If you have any questions, contact the Senior Director, Quality Systems &amp; Field Assurance at 770-784-6913 or email: Brenda.Sheikey@crbard.com or BMD Customer Service at 1-800-526-4455. |  | 2 | 2 units | US Distribution to the state of: FL | Open, Classified |  |  | 2
Z-0460-2018 | 78763 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Affixus Hip Fracture Nail, RH 125 DEG 9MM X 440MM |  |  | August 07, 2017 | 2017 | Zimmer Biomet, Inc. | During the rework process for the April 17, 2017, field action, it was detected that the set screw and sleeve assembly which are designed to have an interference fit, were no longer in the installed location. This could prevent the product assembly with the mating screw and targeting jig. | Process change control | Customers were notified via FedEx and/or email on approximately 08/07/2017. Instructions include ensure appropriate personnel are notified of the recall, locate and quarantine affected product in inventory, complete and return the Certification of Acknowledgement, return all affected product, and provide the names of hospitals or customers that would have received this product if further distributed. |  | 638 | 638 units | US nationwide distribution. Also Portugal and the Netherlands. | Open, Classified |  |  | 638
Z-0535-2018 | 78977 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Custom Perfusion Systems, indicated for use in the extracorporeal circuit during cardiopulmonary bypass (CPB) surgical procedures. |  |  | January 03, 2018 | 2018 | Medtronic Perfusion Systems | Medtronic custom perfusion tubing packs contain Terumo OPS valves that were recalled due to a lack of flow through the valve. | Component change control | In response to an Urgent Safety Alert issued by Terumo Cardiovascular Systems (CVS) Medtronic initiated their safety alert verbally to their direct consignees on 10/23/2017, providing them with a copy of the Terumo Safety Alert dated 06/27/2017. The firm followed this action with a letter on 1/3/18 in which they stated that the action by Terumo CVS was changed to a recall, and , in response, Medtronic was following up as a recall and requested the products be returned. |  | 2914 | 2914 units | DE, IA, CA and Japan | Open, Classified |  |  | 2914
Z-0495-2018 | 79054 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza Picture Archiving and Communication System (PACS) with software version VB20A; Model numbers: 10863171, 10863172, 10863173 |  |  | August 18, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software update to correct several issues that include (1)Potential data loss, (2) Study mix-up, (3) Incorrect measurements on multi-frame images, (4) Dearchiving issue, and (5) Unauthorized access of data due to inadequate permissions for shared folders. | Software design | Siemens released a software update VB20A_HF04 and a Customer Information Letter was sent with Update Instruction SY010/17/P. Service reports for the affected sites will be monitored to ensure the recall was effective.For further questions, please call (610) 219-6461. |  | 47 | 47 | USA (nationwide) Distribution to the states of : AL, CT, FL, KS, LA, MA, MI, MN, MO, MS, NE, NY, OH, OR, PA, TX and WI. Foreign consignees were not reported as device is manufactured outside the U.S. | Terminated | August 30, 2018 | 377 | 47
Z-0492-2018 | 78776 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TosohÂżSorter 9tray, Product Code 022931and TosohÂżSorter 19tray, Product Code 022932. |  |  | June 11, 2017 | 2017 | Tosoh Bioscience Inc | This recall has been initiated due to a potential malfunction of the photo interrupter (Pl). The potentialmalfunction has been attributed to aging and premature deterioration of the Pl collector current. | Nonconforming Material/Component | Customers were sent a recall notification by email on July 11, 2017 informing them that their local Field Service Engineer would contact them to make arrangements to proactively replace the affected parts. Follow up notifications will be by Federal Express with verification of delivery. |  | 565 | 565 total devices U.S. | Worldwide Distribution - U.S. Nationwide - Latin America | Open, Classified |  |  | 565
Z-0494-2018 | 78692 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | 15 ML REAGENT CUP, B1037307, SAP 1221037307 for use with ABX PENTRA 400 / PENTRA C400 |  |  | November 21, 2017 | 2017 | Horiba Instruments, Inc dba Horiba Medical | The ABX PENTRA Reagent Containers do not always sit properly into the reagent rack. In low reagent level scenarios, this could lead to insufficient reagent volume being pipetted and could lead to incorrect results being reported without an alarm. | Process control | The firm initiated their recall by letter on 11/21/2017. The letter stated the following: "A new production of reagent container compliant for each reference is under process and will be made available very soon. In the meantime and to allow a proper operating of the analyzer ABX PENTRA 400 / PENTRA C400, it Is necessary to push the reagent containers in the reagent racks until they touch the bottom of the rack.Please share this information with your laboratory staff, and retain this notification as part of your Quality System documentation. It is mandatory for you to complete and return the enclosed response form within 10 days so wemay maintain our records." |  | 1000 | 10 bags of 100 units | Nationwide | Open, Classified |  |  | 1000
Z-0498-2018 | 78756 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas e 411 Immunoassay Analyzer-Model/Catalog/Part number: 04775279001/04775279973 - cobas e 411 analyzer (disk system) and 04775201001/04775201973 - cobas e 411 analyzer (rack system), medial device listing number: E116019 cobas e 411 |  |  | July 05, 2017 | 2017 | Roche Diagnostics Corporation | A software malfunction can occur on the cobas e 411 and Elecsys 2010 analyzers in the Sample&amp; Control data file which may lead to a potential data mismatch. | Software design | All consignees were notified by Urgent Medical Device Correction sent via UPS Ground (receipt required) on July 5, 2017.The customer is advised to perform the following actions: cobas e 411: Until the software update is available, use the workaround outlined in the Urgent Medical Device Correction (UMDC). Elecsys 2010: Use the workaround outlined in this Urgent Medical Device Correction (UMDC). Non-responding accounts will be defined as users for which confirmation of receipt of the Urgent Medical Device Correction (UMDC) cannot be obtained. These accounts will be monitored on an ongoing basis and follow-up attempts will be made accordingly. Workaround for customers: he customer is advised to perform the following actions: cobas e 411: Until the software update is available, use the workaround outlined in the Urgent Medical Device Correction (UMDC). Elecsys 2010: Use the workaround outlined in this Urgent Medical Device Correction (UMDC). For further questions, please call (317) 521-7259 |  | 923 | 923 instruments | Nationwide Distribution | Open, Classified |  |  | 2356
Z-0502-2018 | 78789 | No | EMF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | bvi Beaver Xstar Safety Slit Knife: |  |  | November 03, 2017 | 2017 | Beaver-Visitec International Inc. | The protective shield was retracted, and the blade was exposed. | Process control | The firm disseminated their safety advisory notices on 11/03/2017 by email and visit. The letter identified the affected product, problem and actions to be taken. The firm requested the following action:"Prior to opening the Tyvek tray, please visually confirm that the plastic shield is not retracted and the activation slider is fully engaged." |  | 659370 | 659,370 units | Worldwide Distribution - US Nationwide and the countries of Italy, UK, Finland, Germany, Ireland, Turkey | Open, Classified |  |  | 3854000
Z-0497-2018 | 78943 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Graseby 3300 Pump and Graseby 3000 Series Pump Accessories |  |  | August 24, 2017 | 2017 | Smiths Medical ASD Inc. | In certain cases, the lock and key supplied with Graseby 3300 pumps and accessories does not match the standard Graseby 3300 key. | Nonconforming Material/Component | Notification was sent to customers on August 24, 2017. Included with the Product Safety &amp; Recall Notification, Instructions to Customers and the Recall response Form. Customers who find product subject to the recall are to contact Smith Medical Customer Service Department to make arrangements to have the affected products returned to Smiths. For further questions, please call (763) 383-3072. |  | 77 | 77 | Worldwide Distribution to the countries of : Australia, Canada, Germany, France, Great Britain, Hong Kong, and Ireland | Open, Classified |  |  | 77
Z-0183-2018 | 77415 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD Insulin Syringes with the BD Ultra-Fine(TM) needle Âż mL 12.7mm 30G, Catalog/REF 328466 |  |  | May 25, 2017 | 2017 | Becton Dickinson &amp; Company | Product mislabel. Lot 6291768 of the BD Insulin Syringes with the BD Ultra-Fine needle ÂżmL 12.7mm 30G, Cat (Ref) 328466, was found to have some polybags in the lot incorrectly labeled as BD Ultra-Fine needle ÂżmL 8mm 31G, Cat (Ref) 328468. The shelf carton and case carton are correctly labeled as BD Ultra-Fine needle ÂżmL 12.7mm 30G. | Labeling mix-ups | BD sent recall notification letters to all direct consignees (wholesaler, hospital, and retail) on May 25, 2017. BD is requesting wholesalers (distributors) to notify any accounts or additional locations that may have received the recalled product from them. BD gives the option to distributors to provide their customer list to BD if they would like BD to notify their customers. BD is requesting retail pharmacies to notify consumers that may have received the recalled product from them. 8/25/17 Firm notified FDA of recall expansion to include Lot Numbers 6312558 and 6340590. A 2nd notification was sent to consignees in a letter dated Aug 30, 2017 that includes all 3 lots that are in scope of the recall. Most of the direct consignees identified in the first recall action were also identified as affected for the second (expanded) recall action since they received either two or all three recall lots. All affected consignees that received any of the three lots have been combined into one consignee list. These customers will be notified of the expanded recall communication (second recall letter dated Aug 30, 2017). To minimize customer disruption, BD will conduct effectiveness check on the second recall letter only, since it covers the recall action for all three affected lots and combines all affected consignees from both actions (first and second/expanded recall). This will allow us to verify that the affected consignees received the notification and took the necessary actions for all three recall lots without having to contact them multiple times for two different letters that cover the same recall action. Details of the effectiveness are as follows: 1.Letters will be mailed with UPS tracking. The effectiveness check level will be established by the FDA classification letter. 2.A Business Response Card is included with the letter to assist direct consignees in responding to the field action. 3.BD will attempt to contact by phone/email (as applicable) non-respondents |  | 7839340 | 7,839,340 units | Nationwide Distribution including CA, CT, FL, IL, LA, ME, MI, NC, ND, NJ, NY, PA, RI, SC, VA | Terminated | June 27, 2018 | 398 | 7839340
Z-0584-2018 | 78888 | Yes | NIK |  |  | Cardiovascular | PMA | Cardiac Resynchronization Therapy Defibrillators (CRT-Ds), Viva. Labeled as the following: |  |  | December 19, 2017 | 2017 | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | Possible prevention of high and low voltage therapy in Medtronic Implantable Cardioverter Defibrillators (ICDs) and Cardiac Resynchronization Therapy Defibrillator (CRT-D). | Process design | On 12/19/2017, Medtronic verbally notified consignees and retrieved eight devices at sites that had affected inventory on hand. On 01/22/2018, Medtronic provided a recall notification letter to physicians following 48 patients implanted with affected devices and risk managers of those medical facilities, recommending that physicians strongly consider prophylactic device replacement for patients implanted with an affected device. Medtronic Technical Services is available to assist physicians with questions at 800-723-4636. Medtronic Patient Service is available to assist patients at 800-551-5544 (Monday-Friday, 8am-5pm Central Time). The recall was expanded on 03/15/2018 and customers were notified by letter. Affected product, not yet implanted, was retrieved on about 03/08/2018. Instructions in the physicians following patients implanted with the additional affected devices included prophylactic device replacement for those patients at higher risk, including patients whose clinical history indicates prior need for high-voltage therapy and/or for pacemaker-dependent patients. The physicians should carefully weigh the risks and benefits of device replacement. If it is determined that replacement is not warranted, instructions included the following: programming changes to reduce the potential for high-voltage charges associated with arrhythmia detection and therapies, such as enabling ATP before charging for fast ventricular rhythms or programming a separate fast VT via VF zone with ATP; Continue three-month in-clinic or remote follow-ups to verify device functionality because inability to interrogate a device or a failed remote monitoring transmission may be an indication that internal arcing has occurred; and to advice patients to seek medical attention immediately if they experience new or unexpected symptoms suspicious for a ventricular arrhythmia.Medtronic Patient Services is available to assist patients at 800-551-5544 (Monday-Friday, 8am-5pm Centra |  | 951 | 951 devices total | US Nationwide Distribution. | Open, Classified |  |  | 951
Z-0609-2018 | 79082 | No | LYU | Surgical apparel. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | BlockAid Cut Resistant Surgical Glove Liners, combine stainless steel and synthetic fibers woven together to form a cut resistant glove material. Blockaid Cut Resistant Surgical Glove Liners are to be worn between two layers of surgical gloves. Item Number: 650126, 650127, 650128, 650129, 650130, 650131 |  |  | September 11, 2017 | 2017 | Zimmer Biomet, Inc. | Cut resistant surgical glove liners lack sterility assurance. | Process control | 1. Review this notification and ensure affected personnel are aware of the contents. 2. Assist your Zimmer Biomet sales representative quarantine all affected product. 3. Your Zimmer Biomet sales representative will remove the affected product from your facility. 4. Complete Attachment 1  Certificate of Acknowledgement. a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com. b. Retain a copy of the Acknowledgement Form with your field action records in the event of a compliance audit of your facilities documentation. 5. If after reviewing this notice you have further questions or concerns please call the customer call center at 1-574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an on-call representative in the case of an emergency. Alternatively, your questions may be sent by email to corporatequality.postmarket@zimmerbiomet.com. |  | 1863 | 1863 | Distributed in 29 states: AZ, CA, CO, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MO, MT, NC, NE, NJ, NY, OH, OR, PA, SC, TN, TX, VA, WV and Washington, DC. Distributed to 1 Veteran's Administration hospital in MO. Distributed in Netherland, Peru, and Switzerland. | Terminated | August 30, 2018 | 353 | 1863
Z-0507-2018 | 78754 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The ReNew Reusable Grasper and Dissector Tips are to be used with the ReNew Laparoscopic Handpieces and they are indicated for cutting, grasping, dissecting, and coagulation of tissue in endoscopic and laparoscopic surgical procedures. |  |  | July 13, 2017 | 2017 | Microline Surgical, Inc. | The heat-shrink insulation tube may split during autoclave sterilization/re-sterilization. If a split occurs, the heat-shrink could fall off during a surgical procedure and if cautery is needed, there is a potential for burn due to insulation failure. | Device Design | The firm, Microline Surgical, Inc., sent a "URGENT: MEDICAL DEVICE RECALL" letter dated July 13, 2017 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately determine if you are in possession of the product and stop use; quarantine all products both new un-opened Reusable Tips AND Reusable Tips that have been used or are in current use (regardless of whether you have observed the malfunction); If you are a distributor or further distributed this product, we require that you immediately remove affected products from inventory and stop sales and shipping; identify customers immediately and notify them of this product recall; and all customers/distributors complete and return the MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement and Receipt Form (even if you have never received any of these products or have none in stock) by E-MAILING A SCANNED COPY TO: ReusableTipRecall@microlinesurgical.com or Fax to: 978-522-8490, ATTN: Reusable Tip Recall, or MAIL TO: Microline Surgical, Inc, 50 Dunham Road, Suite 1500, Beverly, MA 01915, ATTN: Recall Department.Microline has discontinued manufacture of the affected ReNew Reusable Tips and has initiated a ship-hold and new orders-hold for these products.For more information, please email ReusableTipRecall@microlinesurgical.com or contact Director of RA &amp; QA at 978-922-9810 or email: mleighton@microlinesurgical.com. |  | 61769 | 61,769 units | Worldwide Distribution-US (Nationwide) and countries of: Australia, Brazil, Canada, Chile, China, Colombia, El Salvador, Hong Kong, India, Indonesia, Malaysia, Mexico, Panama, Peru, Argentina, South Korea, Sri Lanka, Taiwan, Thailand, Uruguay &amp; Vietnam. | Open, Classified |  |  | 61769
Z-0508-2018 | 79039 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AArtis Q/ Q.Zen biplane |  |  | August 24, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Within a specific number of Artis zee biplane and Artis Q biplane systems, a connection in the ceiling mount consisting of 12 screws may not have been attached and fastened according to quality standards. A quality inspection will be initiated to ensure all screws in the rotation unit meet quality standards. | Employee error | Siemens sent an Urgent: Medical Device Correction letter dated August 24, 2017 to all affected customers. The letter identified the affected product, problem and actions to be taken. A local management or dispatch will schedule for an inspection to correct the issue. For questions please call 1-800-888-7436 |  | 265 | 265 units | US Nationwide Distribution | Terminated | August 29, 2018 | 370 | 265
Z-0509-2018 | 79009 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | AIRO Mobile CT System, Model Number: MobiCT-32 |  |  | July 28, 2017 | 2017 | Brainlab AG | Risk of unintended motion while the AIRO system is in transport mode. | Component design/selection | AIRO customers were issued Field Safety Notices dated June 29, 2017. The letter identified the affected product, problem and actions to be taken. A warning label was provided to be placed on the system. For questions contact your local Brainlab Customer Support Representative. |  | 102 | 102 | Worldwide Distribution - US Nationwide in the U.S. States: of AK, AZ, CA, CO, FL, GA, IL, KY, LA, MD, MA, MI, MN, MT, MO, NJ, NY, NC, OH, PA, RI, TN, TX, UT, VA, WA, and WV; and the countries of Argentina, Australia, Austria, Belgium, Canada, Egypt, France, Germany, India, Italy, Japan, Kazakhstan, Lithuania, Netherlands, Saudi Arabia, Spain, Switzerland, United Arab Emirates, and Uzbekistan. | Terminated | July 25, 2018 | 362 | 102
Z-0513-2018 | 79088 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Precedence 6 Model # 882351 |  |  | December 15, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | During collimator exchange, the alignment of the collimator cart and detectors may cause a collimator exchange fail, and a collimator may drop to the floor beside the collimator cart. The collimator may be partially supported by the cart (e.g. leaning against the cart), or it may fall/slide all the way to the floor. | Under Investigation by firm | Philips sent an Urgent Field Safety Notice letter dated December 15, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your local Philips representative. |  | 1387 | 1387 in total | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 1387
Z-0515-2018 | 78645 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur Calibrator 30, 2 pack, 6 pack, and Ref 2 pack. |  |  | June 07, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc | Siemens Healthcare Diagnostics has confirmed a negative bias for ADVIA Centaur Enhanced Estradiol (eE2) on the ADVIA Centaur, ADVIA Centaur XP and ADVIA Centaur XPT systems when calibrating with Calibrator 30 kit lots ending in 21 (C3021) as compared to Calibrator 30 kit lots ending in 20 (C3020). | Under Investigation by firm | An Urgent Medical Device Recall letter dated June 7, 2017 was distributed to all customers in the United States and anUrgent Field Safety Notice was distributed to all customers outside the United States who havereceived ADVIA Centaur Systems Calibrator 30 kit lot ending in 21(C3021). These letters instructcustomers who operate the ADVIA Centaur, ADVIA Centaur XP and ADVIA Centaur XPTsystems to discontinue use of C3021 . The letters inform customers that Calibrator 30 kit lotsending in 20 (C3020) and 23 (C3023) are not impacted and can be used. The letters informcustomers who operate ADVIA Centaur CP systems to continue using their inventory of C3021,as they are not impacted by this issue.All affected domestic customers were sent a hard copy of the Urgent Medical Device Recall viaFederal Express. Fax-back forms will be used to confirm that the customers received the UrgentMedical Device Recall letter. |  | 5992 | 5992 units (US and OUS) | US and OUS. | Open, Classified |  |  | 5992
Z-0516-2018 | 77594 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DRILL BIT MINI QUICK CONNECT 1.0MM, REF 2312-20-200 &amp; 2312-20-202, NON STERILE |  |  | June 12, 2017 | 2017 | Zimmer Biomet, Inc. | A raw material anomaly was discovered in four lots of raw material and has the potential to be on or below the surface of the device components which can increase the risk of instrument fracture | Nonconforming Material/Component | On June 12, 2017, Zimmer Biomet distributed Urgent Medical Device Recall Removal and Certificates of Acknowledgement notices to their customers via email and FedEx.Risk Manager Customer Responsibilities are the following:1. Review the notification and ensure affected personnel are aware of the contents.2. Assist your Zimmer Biomet sales representative and quarantine all affected product.3. Your Zimmer Biomet sales representative will remove the affected product from your facility.4. Complete and return the Certificate of Acknowledgement via email to: corporatequality.postmarket@zimmerbiomet.com and retain a copy in your records. 5. If after reviewing the notice and if you have further questions or concerns please call the customer call center at 1-574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Alternatively, your questions may be sent by email to corporatequality.postmarket@zimmerbiomet.com. |  | 1524 | 1,524 units | AK, AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA MD, MI, MN, MO, MT, NC, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA &amp; WV. Internationally to: Australia, Chile, Cost Rica, Ecuador, Netherlands, Panama &amp; Switzerland | Terminated | June 29, 2018 | 382 | 1524
Z-0518-2018 | 78863 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Brilliance 64 computed tomography x-ray system |  |  | October 06, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | Numerous issues related to software Brilliance iCT 4.1.6 software version. | Software design | Philips Medical Systems (Cleveland), Inc. sent the Customer Information Letter via certified letters mail with returned receipts. Also, a Field Service Engineer will perform the software update on the system to resolve the issues. Implementation of the software update will be monitored. |  | 1801 | 1801 | 26 US Govt accounts | Open, Classified |  |  | 1801
Z-0522-2018 | 79004 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT - a integrated diagnostic X-ray Computed Tomography (CT) and Positron Emission Tomography (PET) system |  |  | November 01, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | One of the two wire channels mounted inside the Gantry Separation Unit (GSU) that is also used to secure the top two GSU covers, was missing mounting hardware. | Vendor change control | Philips Healthcare is notifying the affected users of this issue via this Customer Information Letter and correcting via Field Change Order FC088200514. |  | 2 | 2 | One domestic and 1 foreign consignee ( Spain) | Terminated | August 14, 2018 | 286 | 2
Z-0524-2018 | 78677 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza systems with SW VB20A, Model Number - 10863171, 10863172, 10863173 |  |  | August 23, 2017 | 2017 | Siemens Medical Solutions USA, Inc | When a prior study is being replaced in the workflow step, in certain scenarios (based on the Display Protocol configuration) the prior study will only be replaced in the active Workflow Step and not show in all other workflow steps. The other workflow steps will continue to show the initially loaded study. | Software design | Siemens Healthineers distributed Urgent Customer Advisory Notices to their customers via courier service. Customers were advised to always verify date and time of the displayed study during workflow step change to ensure that the desired images are being displayed. Otherwise, please replace the displayed initial study in the workflow step with the expected study. Customers with any questions, call Siemens Customer Care Centre at 888-826-9702. |  | 76 | 76 | US Nationwide Distribution | Terminated | October 01, 2018 | 404 | 76
Z-0523-2018 | 78419 | No | FTA | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Disposable Light Handle Cover, Item Number B1-715-65. Packaged 1 device inside a peel pouch and sealed. Sealed pouches are packaged 10 each inside a poly liner. 5 rows of 10 pouches are placed in a labeled corrugated mailer box. |  |  | October 13, 2017 | 2017 | Ecolab, Inc. | Sterile packaging of the disposable light handle cover may contain gaps in the seal of the pouch. | Under Investigation by firm | Ecolab sent and Urgent Medical Device Recall letter dated October 13, 2017 to customers. The identified the affected product, problem and actions to be taken. Instructions include to immediately quarantine lots in inventory, notify customers since the product has been further distributed, return affected product, and complete and return the Recall Response Form. For questions call 903-597-2568. |  | 5958 | 5,958 cases (50 per case) | Distributed domestically to Michigan. | Open, Classified |  |  | 5958
Z-0531-2018 | 77631 | No | NEY | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Accu2i pMTA Applicator, 19 cm INTERMEDIATE, REF/Catalog # 900-602-US, Model # (UPN) H787900602US0, STERILE --- This is only for use with a Sulis V Generator, REF 806-10. |  |  | June 05, 2017 | 2017 | Angiodynamics, Inc. | The firm is recalling product that has the potential to not function during use. The product affected by this recall will not deliver the desired microwave energy due to coolant ingress in an electrical connection; this condition results in an error code High Reflective Power displayed on the microwave generator. | Under Investigation by firm | URGENT VOLUNTARY MEDICAL DEVICE RECALL Letters and Response Forms (dated June 5, 2107) were sent to consignees via FedEx on June 5, 2017. The letter identified the affected product, problem and actions to be taken. |  | 14353 | 14,353 units in total | Worldwide distribution - US Nationwide | Terminated | September 11, 2018 | 463 | 14353
Z-0525-2018 | 78507 | No | PQE | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | Enforcement Discretion | SUR-FIT Natura Moldable Durahesive Skin Barrier with Hydrocolloid Flexible Collar, 10 pack, 45 mm, REF 411802. |  |  | October 26, 2017 | 2017 | ConvaTec, Inc | A portion of Lot 7H02242 outer boxes (market units) of the SUR-FIT Natura Moldable Durahesive Skin Barrier (10 pack, 45 mm, REF 411802) were incorrectly labeled as Natura Moldable StomahesiveSkin Barrier (10 pack, 45 mm, REF 411803). | Mixed-up of materials/components | Customers were notified of the recall via Urgent: Voluntary Medical Device Recall letters on approximately 10/26/2017. The letter identified the affected device and the reason for the recall. The letter also provided instructions for distributors, retailers, and end users. End Users are to 1.) Inspect all shippers of 411802 (Durahesive), and all loose market units of 411803 (Stomahesive) to confirm whether affected product is under their control; 2.) Perform a count of affected product currently in inventory and complete and return the enclosed response form to the address provided; 3.) If customers have any affected product in inventory they are to contact cic@convatec.com or call the Customer Interaction Center at the toll free number 1-800-422-8811 for further instructions regarding how to return affected product and receive correctly labeled replacement product. |  | 476300 | 47,630 units (4,763 market units, 10 per pack) | Nationwide Distribution - AK, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, and WI. | Open, Classified |  |  | 476300
Z-0537-2018 | 78406 | No | LKK |  |  | General Hospital | PMA | Medtronic Ascenda(R) Intrathecal Catheter, Models: |  |  | September 21, 2017 | 2017 | Medtronic Neuromodulation | There is a possibility that specific lots of catheters that have recently been distributed may be at risk for an increased potential for kinking at the proximal end where the catheter connects to the drug infusionpump. | Other | Medtronic on September 21, 2017, initiated a medical device recall and sent the following via mail (with delivery confirmation) to the consignees.In addition to the mailing, Medtronic field representatives may personally visit the accounts to provide the recall notification, assist the customer as necessary with locating and returning affected product(s), and verify that the customer has or will respond to Medtronic with the Customer Confirmation Form. For further questions, please call (800) 707-0933. |  | 194 | 194 units | Nationwide Distribution | Open, Classified |  |  | 194
Z-0533-2018 | 78221 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | 6.6F Plastic DignityÂż Low Profile CT Port W/Attachable ChronoFlexÂż Polyurethane Catheter, Catalog # MRCTI66041, UDI#- 884908031157 |  |  | May 25, 2017 | 2017 | Medical Components, Inc dba MedComp | The kit was packaged with the incorrect introducer needle size. The label indicates the kit contains an 18G Needle w/Echo Tip. The kits were packaged with a 21G Needle w/Echo Tip. | Mixed-up of materials/components | On May 25, 2017, Medcomp distributed Product Recall notices to their customers via email. Medcomp requested the return of all un-used affected product. Customers were advised to immediately examine inventory and to quarantine product subject to recall. If product was further distributed, notification of the recall to those individuals is required. Contact your customer service representative for a Returned Goods Authorization (RGA) number if necessary. The request can be made by telephone by calling 215-256-4201. |  | 65 | 65 units | Arabia, Australia, Israel, Brunei &amp; Jamaica | Terminated | September 04, 2018 | 467 | 65
Z-0539-2018 | 78917 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Implant Direct, INTERACTIVE HEALING COLLAR, PART NUMBER 6530-15, 5.0mmL: 3.0mmD Platform |  |  | January 23, 2018 | 2018 | Implant Direct Sybron Manufacturing, LLC | InterActive Healing Collar, Lot Number 104203, labeled as sterile with distributed prior to being sterilized. | Process control | The firm initiated their recall on 01/23/2018 by letter. the recall notices requested the following actions by the distributors: "1. Please review your inventory for the affected product.2. Please complete and return the Acknowledgement and Recall Return Form with the affected product, if available, within 48 hours.3. If you are an authorized Implant Direct Sybron Manufacturing distributor, we request that you identify those customers that may have been shipped the affected product lot and contact these customers to inform them of this issue within fortyeight (48) hours of receipt of this notification in order to provide the customers with replacement product."The letter disseminated to the physician requested the following actions:"1. Please review your inventory for the affected product.2. Please complete and return the Acknowledgement and Recall Return Form with the affected product, if available, within 48 hours." |  | 55 | 55 units | CO, SC, IA, NJ, FL, CA, VA, ID, Netherlands | Terminated | August 09, 2018 | 198 | 55
Z-0536-2018 | 79070 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | PowerPICC Provena Catheters, as 3 Fr. Single lumen and 4 Fr. Dual Lumen. Packaged in a poly tray with Tyvek lid within a header bag configuration or a breather bag configuration. |  |  | December 28, 2017 | 2017 | Bard Access Systems Inc. | Possible leaks on the 3 Fr. Single lumen and 4 Fr. Dual lumen PowerPICC catheters at the luer hub extension leg junction. | Labeling design | Customers were notified via letter on approximately 12/28/2017. Instructions include a reminder that if additional extension sets are attached to the luers of the catheters, they should be appropriately secured as per hospital recommended securement protocol, to refrain from having unsupported extension sets that could put stress on the luer extension leg junction of the catheter, and to notify customers if the product was further distributed. Also, the notification contained a statement that Bard Access is in the process of updating their Instructions for Use to include a statement regarding the importance of properly securing additional extension sets per hospital recommended securement protocol. On 04/11/2018, the recall was expanded and customers were notified via letter, which explained that the previous recall has been expanded to involve return of affected product as well as include product distributed from February 2017 through March 2018. Instructions included to examine inventory for affected product, quarantine and immediately discontinue distribution of the product, complete and return the Business Recall Response Form, arrange for the return of affected product, and notify customers if the product has been further distributed. For distributors, Bard Access Systems requested a customer list so that they could notify their customers. |  | 53699 | 53,699 units total | Distribution US nationwide. | Open, Classified |  |  | 53699
Z-0549-2018 | 78219 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Humeral Guide Wire Ball Tip 2.4 mm Diameter 70 cm Length, Item Number 47-2255-008-00 |  |  | June 05, 2017 | 2017 | Zimmer Biomet, Inc. | The design verification for the previous packaging configuration G928 does not cover the 70cm wires. A design verification has been completed to move the 70cm guide wires to a new packaging configuration. As a result, the products packaged in the previous packaging configuration are being removed. | Package design/selection | On June 5, 2017, Zimmer Biomet distributed Urgent Medical Device Recall notices to their customers via FedEx and email. **Hospital Risk Managers are advised to:** 1. Review this notification and ensure affected team members are aware of the contents.2. Complete the Certification of Acknowledgement portion of Attachment 1a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days.3. Assist your Zimmer Biomet sales representative quarantine all affected product.**Distributors are advised to:**1. Review this notification and ensure affected personnel are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Forma. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days.4. Note that any hospitals that received direct shipments of this product from Zimmer Biomet will be sent a copy of the Risk Manager directly. It is important that you review the list of hospitals included with the email notification sent to your facility to identify additional accounts Zimmer Biomet has not notified. Using the Additional Accounts Form provided with the email notice sent to your facility, return contact information for any additional hospitals that may have received or used the affected product within the past 12 months. If there are no additional users to notify, please confirm this by annotating and returning the form provided indicating that no additional users have been identified.5. Return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet within 2 weeks.a. For each return, send a copy of Attachment 1 to corporatequality.postmarket@zimmerbiomet.com.b. Include a hardcopy of Attachment 1 with your shipment for i |  | 7888 | 7,888 units | Nationwide within U.S.A and Puerto Rico; International: Australia, Brazil, Canada, China, Germany, India, Japan, Netherlands &amp; Singapore | Open, Classified |  |  | 27442
Z-0541-2018 | 78950 | No | PWY | Hypodermic single lumen needle. | General Hospital | General Hospital | Enforcement Discretion | Safety-Deluxe Basic Soft Tissue Biopsy Tray without Biopsy Needle, Catalog # 977 |  |  | June 27, 2017 | 2017 | Robert Busse &amp; Co. Inc. d.b.a. Busse Hospital Disposables | BD SafetyGlide Needle 22GA 1-1/2 inch may contain the presence of loose polyprophylene foreign matter above release specification. | Nonconforming Material/Component | On June 27, 2017, Busse Hospital Disposables distributed Urgent Medical Device Recall notices and Response Forms to their customers via courier service. A copy of the supplier's recall notice was also provided to customers. Customers were advised to conduct the following:"Discontinue use of needle immediately."Please check your inventory and cease distribution immediately."If you have distributed the above products, please contact your customers and inform them of this recall immediately. "Contact Busse Hospital Disposables for a Return Good Authorization number. "Complete and Return the Response Form via fax to 631-435-2849 or via email to: ijacobo@busseinc.com. Customers with questions may call 631-435-4711, ext. 254. |  | 28 | 28 cases | US Nationwide Distribution | Open, Classified |  |  | 675
Z-0552-2018 | 78444 | splint, extremity, non-inflatable, external, non-sterile | NOC | Noninflatable extremity splint. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ACE(TM) BRAND, ULTRA LITE ANKLE brace: (a) Size Selection Medium, Cat #901014, UPC 051131019984 2; (b) Size Selection Large, Cat #901015, UPC 0 51131019985 9 |  |  | November 01, 2017 | 2017 | 3M Company - Health Care Business | Mold was found on the non-patient contactsurfaces of the product and was identified asAspergillus candidus. | Packaging process control | The firm initiated their recall on 11/01/2017 by letter. The letter requested the consignee cease distribution of the recalled product, isolate it pending return. Distributors were asked to notify their customers including a copy of the 3M recall notice. |  | 996 | 996 units | AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN,KS, KY, LA, MD, MI, MO, MS, NC, NH, NY, OH,OR, PA, SC, TN, TX, UT, VA, WI, WV | Completed |  |  | 996
Z-0555-2018 | 78914 | No | LYJ |  |  | Neurology | PMA | VNS(R) Therapy Programmer, GMDN: 44077, Model 3000 v1.0.2.2 |  |  | December 02, 2017 | 2017 | Cyberonics, Inc | Two Model 3000 Programmers were distributed in error by prior to FDA approval of version 1.0.2.2 software. | Employee error | The units were collected by visit on December 21, 2017 and December 28, 2017 from the physicians. Both programmers have been successfully removed from the customers and no further actions are being proposed.For further questions, please call (281) 228-7260. |  | 2 | 2 units | US Distribution to the states of : CO and ND | Terminated | August 06, 2018 | 247 | 2
Z-0553-2018 | 78934 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | Rheumatoid Factor (RF) Kit for use on SPAPLUS, Product #LK151.S.A |  |  | January 04, 2018 | 2018 | The Binding Site Group, Ltd. | The low and high controls for the product are returning results above the assigned values. | Under Investigation by firm | Binding Site sent an Urgent Field Safety Notice dated December 22, 2017.Advise on action to be taken by the User:" The new QC values detailed below should be used for the remainder of the shelf life of this Rheumatoid Factor Kit for use on SPAPLUS."Input the revised QC values into your SPAPLUS instrument"Sign and return the TSWS18 E-Back Form accompanying this notification within 5 business days"Refer to Section 9 of the Product Insert SIN282(.A) if QCs are out of range following reassignment. For further questions, please call (858) 453-9177. |  | 49 | 49 kits to the U.S. office (32 kits further distributed) | US Distribution was made to CA. There was no foreign/military/government distribution. | Completed |  |  | 49
Z-0554-2018 | 79005 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | TD60 Transmitter battery charger used with the BeneVision Central Station, Part number 115-030108-00 |  |  | July 14, 2017 | 2017 | Mindray DS USA, Inc. dba Mindray North America | A damaged circuit board in the battery charger can cause the battery to overcharge and the battery case to swell and separate. | Component design/selection | Certified letters were sent to all affected customers on 07/14/2017, to offer a replacement of the battery charger. |  | 203 | 203 | Distributed in 26 states: AL, CA, CO, FL, GA, IA, IN, ME, MI, MN, MO, NC, ND, NE, NH, NJ, NM, NV, NY, OH, PA, SD, TX, UT, WA, and WY; and Puerto Rico. | Terminated | December 17, 2018 | 521 | 203
Z-0489-2018 | 78351 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | MRx Defibrillator Model # M3535A &amp; M3536A, UDI 00884838000018(21)US00601969 |  |  | July 12, 2017 | 2017 | Philips Electronics North America Corporation | Affected units may have a defective component that is intended to protect the internal circuitry from high voltages during shock delivery. The defective component may also prevent delivery of shocks in AED mode. | Nonconforming Material/Component | Philips Healthcare sent an Urgent Medical Device Correction Notification/Field Safety Notice letter dated June, 2017 to their customers. The letter identified the affected product, problem, hazard involved and actions to be taken. Philips is requesting that customers report if their device fails to function as intended. Philips will contact all customers for device replacement once parts are available. If you need further information or support concerning this action, please contact your local Philips representative. |  | 71 | 71 units | No U.S. distribution, Internationally to: India, Ireland, Japan, Republic of Korea, Malaysia, Mexico, South Africa &amp; United Kingdom | Completed |  |  | 71
Z-0534-2018 | 78968 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Covidien Vistec" X-Ray Detectable Sponges, 7317 Vistec" X-Ray Detectable Sponges Sponge. |  |  | July 17, 2017 | 2017 | Covidien LP | Medtronic is voluntarily recalling two production lots of its Covidien Vistec" x-ray detectable sponges de to possible contamination with human blood caused by handling of the product during the manufacturing process and prior to final packaging and Ethylene Oxide sterilization. | Material/Component Contamination | Medtronic sent an Urgent Medical Device Recall letter dated July 17, 2017 to their customers via Federal Express or certified mail. The letter identified the affected product, problem and actions to be taken. Customers are requested to quarantine, discontinue use and return affected product. Customers were instructed to acknowledge receipt of the Urgent Field Corrective Action Notice. For questions contact your Medtronic representative or Customer Service at 800-962-9888, option 2. |  | 33280 | 33,280 | U.S., Gov Accts, and foreign countries: PA, PR, CO. | Terminated | September 14, 2018 | 424 | 33280
Z-0565-2018 | 78622 | No | EXS | Urine flow or volume measuring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BARD(R) CRITICORE(R) MONITOR, Model: 000002N Series |  |  | November 10, 2017 | 2017 | C.R. Bard, Inc. | BMD has identified that some lots of CritiCore(R) monitors may experience failures involving urine output volume measurement fluctuations, core bladder temperature measurement fluctuations and alarm errors. | Software design | Bard Medical sent an Urgent Medical Device Product Recall letter dated November 10, 2017. Action required:"Immediately examine your inventory and quarantine product subject to this recall. Please refer to CritiCore Monitor Affected Serial Numbers (Attachment 2) to help you locate the affected product serial number. Do not use or further distribute any affected product."Please complete and return the accompanying Recall &amp; Effectiveness Check Form (Attachment 1) attached to this letter regardless of whether you have any of the affected product. Additional instructions for product return are contained on the form."If you have further distributed any units with the referenced serial numbers, please identify your customers and notify them at once of this product recall: Your notification should include a copy of this letter and the accompanying enclosures.This recall should be carried out to the user level." For further questions, please call (770) 784-6120 |  | 343 | 343 units | US Distribution and Internationally to Japan | Open, Classified |  |  | 343
Z-0564-2018 | 78824 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Vado Bi-Directional Steerable Sheath 8.8F, REF SS8FMB74, packaged in a lidded tray that is within a pouch which is inserted into a shelf carton, Sterile, Rx. The firm name on the label is Kalila Medical, Campbell, CA. |  |  | December 14, 2017 | 2017 | Kalila Medical | Torn sheath liner material was detected at the proximal end of the device during internal product testing which may be the result of damage incurred during the manufacturing process. | Under Investigation by firm | The recalling firm issued a letter dated 12/14/2017 via FedEx on 12/14/2017. |  | 261 | 261 sheaths | Distribution was made to DC, FL, GA, MD, MI, NJ, NY, OH, PA, SC, and VA. There was also military/government distribution. There was no foreign distribution. | Terminated | December 14, 2018 | 365 | 360
Z-0577-2018 | 79045 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Custom Made Tracheostomy Device |  |  | July 15, 2016 | 2016 | Smiths Medical ASD, Inc. | Smiths Medical determined that a single custom Tracheostomy device had been shipped to the customer without required sterilization. | Under Investigation by firm | The firm, Smith Medical, notified the consignee by phone on July 29, 2017 of the Field Corrective Action Report. The consignee was instructed to return the devices for destruction and replacement.If you have any questions, contact VP, Global Quality and Regulatory at 219-989-9150. |  | 1 | 1 | International Distribution to:Germany | Completed |  |  | 1
Z-0576-2018 | 78878 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Contiplex Nerve Block Tray. Product Catalog Number: 570205. |  |  | June 21, 2017 | 2017 | B. Braun Medical, Inc. | B. Braun Medical Inc. (BBMI) initiated this recall of various BBMI convenience kits containing the 22GA 1-1/2in SafetyGlide Needle due to the presence of loose polypropylene foreign matter was above release specification. This foreign matter has beenidentified as a product hub material which has been tested for biocompatibility per ISO 10993 during the productdevelopment process. | Material/Component Contamination | The firm, BD, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 6/12/2017 to its customers/kit packers. B.Braun Medical Inc. (BBMI), sent a "VOLUNTARY MEDICAL DEVICE RECALL NOTIFICATION" letter dated 6/23/2017 to all BBMI customers in receipt of suspect product. Customers were notified by US Postal Service Certified Mail with registered return receipt mail or FedEx Priority. The letters described the product, problem and actions to be taken. The customers were instructed to review the device notification in its entirety and inform all users of the recall; determine your inventory of the affected product-Do not destroy any affected product; complete and return the "Product Removal Acknowledgement" form to B.Braun Medical Inc. Quality Assurance department via fax to 610-849-1197 or email to PA_QualityAssurance.BBMUS_Service@bbraun.com within two 2 weeks of receipt, even if the total inventory in your possession is zero 0; and if you have any full cases, partial cases or unused individual pieces a BBMI Customer Support Representative will contact you to provide instructions for handling affected product and arrange for return. Should you have any questions or concerns regarding the recall, please contact our Customer Support Department at 800-227-2862. |  | 300 | 300 units | US Nationwide Distribution. | Terminated | April 03, 2018 | 286 | 34690
Z-0562-2018 | 79008 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT (model 882442) running software version 4.0.2 |  |  | December 11, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | A software issue causes PET reconstructions to fail intermittently. It was determined that reconstructions fail due to a negative table position (-1 value is inserted) in the raw data list file, rather than the actual table position.This error has been found to occur in two scenarios:1) When the system operator cancels an acquisitiona. The error will occur every time a scan is cancelled by theoperator.2) Couch position requests within the software sequence were delayeda. The error occurs intermittently, but has been found to occur more frequently when the gantry s network is heavily loaded with multiple retrospective reconstructions running in parallel.b. The error manifests to the technologist by an error message and Failed status on the Reconstruction Monitor and the error message Result {0} failed to reconstruct" on the Acquisition Workflow window during reconstruction after the patient scan has been completed.In both scenarios, the acquisition data will not be able to be reconstructed and will therefore be unusable. | Software design | Philips Healthcare sent an Urgent-Field Safety Notice dated December 11, 2017. U.S. Customers will be receive the letter via USPS Certified mail. Non-U.S.. customers will receive the letter through Philips Global Markets following local regulations. Philips is informing customers of the software issues on the systems, as well as available workarounds to prevent these issues from occurring or to minimize their impact.For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, follow the prompts) |  | 24 | 24 | Worldwide Distribution - US Distribution to the states of : CA, OH, IL, PA and NY. | Open, Classified |  |  | 24
Z-0559-2018 | 79075 | No | NEY | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Emprint Percutaneous Antenna with Thermosphere Technology, 20cm (reinforced), Material CA20L2 |  |  | August 07, 2017 | 2017 | Covidien LLC | The firm received consumer reports of the ceramic trocar tip of the Emprint ablation antenna disengaging from the needle shaft post-ablation. This can result in the tip remaining in the patient. | Device Design | All consignees were notified via Federal Express or certified mail on August 7, 2017, and the letter informs customers of the possible trocar tip disengagement and the actions they should take. Customers are requested to segregate and return affected product and acknowledge receipt and understanding of the Urgent Medical Device Recall Notice. |  |  |  | Nationwide including PR, Canada, China | Open, Classified
Z-0581-2018 | 79135 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | AVID Medical QA Pack AV10654-10 |  |  | July 03, 2017 | 2017 | Avid Medical, Inc. | The swab stick collection device provided in a convenience kit expires prior to the kit expiration date. | Process control | Avid Medical began shipping this product on 05/23/2017. Immediately examine your inventory and quarantine all product subject to this correction. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product correction. Your notification to your customers may be enhanced by including a copy of this Urgent Medical Device Correction notice. Please complete the attached Correction Response Form as soon as possible. Attached is a label template for use in the generation of an appropriate warning label for application on all AV10654-10 inventory of impacted lot 1235229. This label should be applied to each affected convenience kit and product case in a prominently visible location to end-users. The location chosen should not cover any other critical product information found on existing product labeling. If you prefer to have labels printed for your facility, please send quantity and end-user location information request to sharon.moll@owens-minor.com. Replacement component swab stick collection systems are available upon request. This correction will be executed with oversight of the associated account sales rep. Please complete and return the enclosed response form as soon as possible. |  | 330 | 330 | Seventeen sites located in 13 states: FL, IN, KS, KY, MA, OK, NC, NE, NY, OH, PA, VA, and TX. | Terminated | April 17, 2018 | 288 | 330
Z-0579-2018 | 78469 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Carestation 620 A1, Carestation 650 A1, and Carestation 650c A1 |  |  | October 11, 2017 | 2017 | GE Medical Systems China Co., Ltd. | Possible incomplete seal between the CO2 absorber and the breathing circuit CO2 bypass port assembly of the Carestation 600 Series systems. | Other | On approximately 10/11/2017 customers were notified of the recall via letter. Instructions include that if elevated FiCO2 is observed, to increase the flow of fresh gas to reduce the volume of patient gas that is rebreathed. If the FiCO2 cannot be adequately reduced with this action, consider switching to another anesthesia delivery device. GE Healthcare is planning to release revised parts that minimize the likelihood of incomplete gas flow through the CO2 absorbent canister. GE Healthcare will correct all affected devices by installing updated parts at no cost to the customer. The updated parts will be included in the 12-month preventative maintenance schedule for the Carestation 600 Series System. |  | 5148 | 5,148 devices | Distributed domestically to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NV, NY, OH, OK, PA, TN, TX, UT, VA, VT, WA, WI, WV. Distributed internationally to Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Cambodia, Canada, Chile, China, Colombia, Costa Rica, Cote Di'Viore, Croatia, Czech Republic, Democratic Republic of Congo, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kenya, Korea (Republic of), Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Maldives, Malta, Mauritius, Mexico, Moldova, Morocco, Myanmar, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, Zimbabwe. | Open, Classified |  |  | 5148
Z-0566-2018 | 78839 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Medtronic ANGLED DRILL, REF 6630902, QTY: 1 EA Rx only |  |  | December 22, 2017 | 2017 | Medtronic Sofamor Danek USA Inc | Medtronic has discovered the product lot in question was noted as having a longer drill tip extension than what is indicated on the print specification. | Process control | The firm initiated the recall by letter on 12/22/2017, requesting the return of the product. |  | 2 | 2 units | MN, NC | Terminated | July 31, 2018 | 221 | 2
Z-0599-2018 | 79002 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | 3M Bair Hugger(M) Normothermia System, Sterile Full Access Underbody, Model 63700 |  |  | January 23, 2018 | 2018 | 3M Company - Health Care Business | Adesign change to 3M(TM) Bair Hugger(TM) Full Access Underbody Warming Blanket, Models 63500 and 63700, can make these blankets more susceptible to unintended occlusion of airflow resulting in limited or partial inflation during use. | Finished device change control | The firm, 3M, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 1/19/2018 to its customers on 01/23/2018. The end user letter requested the following:"Actions Required:1) Please immediately identify the affected product listed above, remove from your inventory and DO NOT USE.2) Follow the instructions on the attached Customer Product Recall Acknowledgement Form, complete the form and return it to 3M via email at 3Mhealthcare@montage72.com.3) Please dispose of all affected product per your institutional protocol and document disposal in the recall acknowledgment form."The distributor letter requested the following:"Actions Required:1. Immediately discontinue distribution of the 3M(TM) Bair Hugger(TM) Full Access Underbody Warming Blanket, Models 63500 and 63700, with the referenced lot numbers above.2. Determine if any stock of the 3M" Bair Hugger" Full Access Underbody Warming Blanket, Models 63500 and 63700, with the referenced lot numbers above remains in your inventory.3. Complete the attached Distributor Product Recall Acknowledgement Form and return to 3M via email at 3Mhealthcare@montage72.com. It is important to notify 3M whether or not your facility has any inventory of this product.4. Notify any of your customers who may have received the lots. Please use the attached customer notice and form provided in your communications to the customer. Customers should complete and return the enclosed form to 3M per the instructions on the form. Customers should not use the affected product and dispose per institutional protocol.5. Dispose of any affected product in your inventory and document disposal in the recall acknowledgment form."If you have questions regarding this matter, please contact the 3M Health Care Customer Helpline at 1-800-228-3957, option #6, Monday through Friday, 7:30 a.m.  6:00 p.m. Central Daylight Time. |  | 180 | 36 cases of 5 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Canada, Chile, Costa Rica, Deutschland GMBH, Dominican, Ecuador, Hong Kong, Indonesia, Japan, Korea, Paraguay, Philippines, New Zealand, Shanghai, Thailand, and Uruguay. | Open, Classified |  |  | 165720
Z-0594-2018 | 79024 | No | KWL | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CORAIL HIGH OFFSET STEM Collarless SIZE 14 |  |  | July 12, 2017 | 2017 | DePuy Orthopaedics, Inc. | Incorrect device in the package. A package that was labeled as the CORAIL COXA VARA HIGH OFFSET STEM Collared Size 9 (Part No. 3L93709 / Lot No. 5291990), contained the CORAIL HIGH OFFSET STEM Collarless SIZE 14 (Part No. L20314 / Lot No. 5292130). | Mixed-up of materials/components | In August of 2017 customers were notified via an Urgent Medical Device Recall Letter, titled "ONE LOT OF CORAIL COXA VARA HIGH OFFSET STEM Collared Size 9 and ONE LOT OF CORAIL HIGH OFFSET STEM Collarless SIZE 14" with instructions to cease using the affected product immediately.Note: The affected devices may be on consignment at your facility.Return Affected Implants: Determine if any of the recalled devices are still on hand and return the affected device(s) immediately to the DePuy Sales Consultant or return to DePuy Synthes for credit following the normal returns procedures. Reconciliation Form: Complete the Reconciliation Form and return to your DePuy Orthopaedics Sales Consultant or fax to 574-371-4939 or scan and email to DPYUS-JointReconFieldActions@its.jnj.coRecords: Retain a copy of the completed Reconciliation Form in your files along with this notice.Additional Notifications:Notify surgeon users at your facility by providing them with a copy of this notice to ensure surgeon users are aware of this recall.Forward this notice to others in your facility that need to be informed.If any affected product has been forwarded to another facility, contact that facility immediately to communicate this field action with the facility/facilities.Maintain a copy of this notice with the affected devices.Contacts For product-related questions, please contact your local DePuy Orthopaedics, Inc. Sales Consultant. For clinical questions from surgeons, please contact DePuy Orthopaedics, Inc.s Scientific and Medical Affairs mailbox RA-DPYUS-DePSynSc@ITS.JNJ.com. For questions about device recall information provided, please contact Kim (Earle) Long, Senior Recall Coordinator, at 574-371-4917 (M-F; 8 a.m.  5 p.m. EDT) or at DPYUS-JointReconFieldActions@its.jnj.com.Health care professionals may report serious adverse events, side effects or product quality problems with the use of this product to DePuy Orthopaedics, Inc. either by contacting yo |  | 19 | 19 | Nationally; US Consignees for this recall are the affected are US Distributors and US Medical Facilities (Hospitals/User Facilities) that may have purchased, used, or received the affected device. | Open, Classified |  |  | 40
Z-0590-2018 | 77469 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Arctic Sun ArcticGel Pads - Medium, Product Code 317-07 |  |  | June 02, 2017 | 2017 | C.R. Bard, Inc. | Label on product box contains the incorrect reference #317-09 while the shelf box label with barcode and the internal foil pouch label are correctly labeled with product code 317-07 and lot number NDAR0462. | Labeling mix-ups | Customers were notified on approximately 06/02/2017 via FedEx with delivery confirmation and per the individual country's delivery system. Instructions included to examine inventory and identify any product subject to the recall and notify customers if product was further distributed. If customers wished to return product, they were instructed to contact customer service at 1-800-526-4455. |  | 128 | 128 units | Distributed domestically. Distributed internationally to Australia and Belgium. | Open, Classified |  |  | 128
Z-0595-2018 | 79007 | No | NSU | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hologic Panther(R) Fusion(TM), Catalog number ASY-09600 |  |  | January 20, 2018 | 2018 | Hologic, Inc | The Panther or the Panther Fusion glass panel separates from the 19 inch touch screen monitor and could potentially fall and hit an operator. | Device Design | The firm sent an Urgent Medical Device Recall Notification dated January 8, 2018 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed that a Hologic representative will contact your site to arrange a visit in order to provide an interim solution to secure the glass panel until a replacement monitor with the issue corrected can be installed. For questions contact Hologic Technical Support. at +1-888-484-4747 or +1-858-410-8511 or by e-mail at molecularsupport@hologic.com |  | 103 | 103 units | Worldwide distribution | Open, Classified |  |  | 103
Z-0591-2018 | 79042 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | NS357R Vega PS Tibia Trial/Preparation Plateau size T4 (Component of the Vega Knee System Set #ST0468) |  |  | August 10, 2017 | 2017 | Aesculap Implant Systems LLC | The NS357R Vega PS Tibia Trial/Preparation Plateau size T4 may be incorrectly labeled. If the incorrect size trial is used the potential risk is implantation of the incorrect implant size which may lead to injury of ligaments, pain and possible revision surgery. | Mixed-up of materials/components | On August 21, 2017 an Urgent Medical Device Recall notification was issued titled "NS357R Tibia Trial Plateau" asking distributors to inspect Sets #0468 and follow the inspection instructions on the customer notification letter and immediately remove the affected part and return to Aesculap Implant Systems. Questions or concerns can be directed to Customer Service at 314-551-5998 |  | 13 | 13 Units | USA (nationwide) | Terminated | August 08, 2018 | 363 | 13
Z-0592-2018 | 78910 | No | LSL | Neisseria spp. direct serological test reagents. | Microbiology | Microbiology | 510(k) | Xpert CT/NG Catalog GXCT/NG-120 Batch 1000049568, Catalog GXCT/NG-10 Batch 1000053075, Catalog GXCT/NG-10 Batch 1000056156 and Catalog GXCT/NG-10 Batch 1000056157. |  |  | October 09, 2017 | 2017 | Cepheid | Cepheid has received reports that some of our customers have experienced a higher number oftemperature drift errors than would typically be expected when using this test and its intended usespecimen types. Investigation of the issue also revealed an occasional probability of falsenegative results which remains consistent with package insert claims. | Equipment maintenance | An Urgent Medical Device Recall letter (response form included), dated October 9, 2017, was sent to affected consignees. The letter identified the affected product, the issue involved, impact, and actions to be taken. Cepheid will provide replacement product for Xpert CT/NG assay product identified in this recall. Customers are to complete and the return the Response Form provided in order to receive replacement product. Additionally, the letter instructs users to review their laboratory statistics to see if their positivity rates are trending low during use of the affected lots. If so, they are to review the negative patient results according to their laboratory's Quality Management System. Customers are to refer to the applicable contact information provided if there are any questions. |  | 3586 | 3,586 kits (49,400 units) | Worldwide Distribution -- USA, Puerto Rico, and Haiti. | Terminated | February 13, 2018 | 127 | 3586
Z-0604-2018 | 78958 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | 11MM/130 Degree Titanium Trochanteric Fixation Nail 380MM/Right, Sterile |  |  | July 21, 2017 | 2017 | Synthes (USA) Products LLC | The locking mechanism and protective cap were missing from the sterile packed nails of the affected lot. | Employee error | The company issued a recall letter on 7/21/2017 asking customers to quarantine affected product and arrange for it to be returned. |  | 4 | 4 units | TX, GA, PA, UT and Canada | Terminated | July 10, 2018 | 354 | 4
Z-0605-2018 | 78699 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Lyophilized Chemistry Calibrator Level 1 |  |  | August 24, 2017 | 2017 | Beckman Coulter Inc. | Unpaired lots of AU lyophilized chemistry calibrator level 1 and 2 have been distributed. Calibrator level 1 and 2 must be paired when used for calibration and cannot be used in any other combination. If lot numbers are not correctly paired, calibration can be run using incorrect set points which may generate erroneous results. QC failures could occur if incorrect pairs of calibrators are used. | Under Investigation by firm | On August 28, 2017 a letter was sent to all the customers stating the following:Attention Beckman Coulter Customer,Beckman Coulter is initiating a field action for the product listed above. This letter contains important information that needs your immediate attention.ISSUE:Âż Unpaired lots of AU Lyophilized Chemistry Calibrator level 1 and 2 have been distributed.Âż Calibrator level 1 and 2 must be paired when used for calibration and cannot be used in any other combination.Âż Paired calibrator lots are identified by the last two digits of the lot number.Calibrator Level 1  REF DR0070-1, Lot 6101K51Calibrator Level 2  REF DR0070-2, Lot 6102K61Incorrect pairCalibrator Level 1  REF DR0070-1, Lot 6101K61Calibrator Level 2  REF DR0070-2, Lot 6102K61Correct pairIMPACT:Âż If lot numbers are not correctly paired, calibration can be run using incorrect set points which may generate erroneous results.Âż QC failures could occur if incorrect pairs of calibrators are used.Calibrator Level 1  REF DR0070-1, Lot 6101K51Calibrator Level 2  REF DR0070-2, Lot 6102K61Incorrect pairÂż If the above incorrect pair is run with Level 1 lot 6101K51and Level 2 lot 6102K51 set points, erroneous results could be generated for the following assays:o Bicarbonate : maximum 11% bias lowo Calcium Arsenazo : maximum 5% bias lowo Calcium oCPC : maximum 5% bias lowo Inorganic Phosphorous : 6.4% bias higho Magnesium: 3.2% bias lowÂż If the above incorrect pair is run with Level 1 lot 6101K61 and Level 2 lot 6102K61 set points, erroneous results could be generated for the following assays:o Bicarbonate : maximum 15.8% bias higho Direct Bilirubin OSR6x181: 8.7% bias lowACTION: Âż Check the level 1 and level 2 calibrators to make sure you are usingcorrectly paired lots for calibration.Âż Ensure correct set points are being used for the lot. Set points are inthe IFU for each lot and can be found in each kit, or on the BeckmanCoulter websites |  |  | unknown | Worldwide Distribution in the states of AL, AK, AZ, AK, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IO, KS, KY, LA, MA, MD, MA, MI, MN, MO, NE, NV, NH, NJ, NY, NM, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VI, WA, WV, WI, and WYand the countries of Brazil, Canada, Costa Rica, El Salvador, Guam, Paraguay, and Peru | Open, Classified
Z-0601-2018 | 78964 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROW PICC powered by ARROW VPS Stylet, Product Codes: CDC-35052-VPS and CDC-35041-VPS. Peripherally Inserted Central Catheter |  |  | July 25, 2017 | 2017 | Arrow International Inc | Finished good kits may contain incorrect components. Kits that should contain a 5 FR 2 lumen catheter may contain a 4 FR single lumen catheter and kits that should contain a 4 FR single lumen catheter may contain a 5 FR 2 lumen catheter. This may cause a delay in treatment while a replacement catheter is obtained. | Packaging process control | Arrow International notified all U.S. consignees with letter sent on July 25, 2017, via FedEx 2-day. second notification August 22, 2017, third notification November 28, 2017. |  | 2691 | 2691 total | US nationwide distribution. | Terminated | August 29, 2018 | 400 | 2691
Z-0600-2018 | 78902 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Total Knee Replacement , Catalog/Model # TPS-111-1111, UDI: +M572TPS1111111021/$$+31712310398482- and +M572TPS1111111021/$$+31712310399917/ |  |  | July 24, 2017 | 2017 | ConforMIS, Inc. | Implant kit mislabeled | Employee error | On July 10, 2017, ConforMIS notified their Sales Representatives about a possible product mislabeling. On July 11, 2017, ConforMIS secured impacted product. On July 21, 2017, ConforMIS issued a Memorandum to their Sales Representatives. ConforMIS requested that the kits be returned to ConforMIS for investigation and disposition. For further questions, please call (781) 345-9001. |  | 2 | 2 units | US Distribution to the states of :CA and TX. | Terminated | February 01, 2019 | 557 | 2
Z-0486-2018 | 78681 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Radial Artery Catheterization Set, 1) NA-04550-1A &amp; 2) NA-04550-X1A, device permits access to the peripheral arterial circulation or to other small vessels. |  |  | July 10, 2017 | 2017 | Arrow International Inc | Needle supplier indicated the presence of loose polypropylene foreign matter above release specification. | Material/Component Contamination | On July 10, 2017, Arrow International distributed Urgent Medical Device Recallnotice and acknowledgment forms to customers. Customers are advised to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products listed in the notice, so that the affected products can be returned to Arrow.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419- 8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact youwith a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to Arrow International.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document your receipt of this letter. If you have any other questions, feel free to contact your local sales representative or Customer Service at 18662466990. |  |  |  | ... | Terminated | September 14, 2018 | 431
Z-0772-2018 | 79101 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | V-PRO 60 Low Temperature Sterilization System (UDI: 00724995152000) |  |  | September 28, 2017 | 2017 | Steris Corporation | Correction to update sterilizer software to prohibit the use of expired sterilant prior to the start of a sterilization cycle. | Device Design | The V-PRO 60 system software has been updated to ensure the VAPROX Sterilant expiration date is verified prior to the initiation of each sterilization cycle. STERIS Service Technicians will be contacting your facility to update your V-PRO 60 Sterilizer(s) with the revised system software. User Action  Users should continue to review their V-PRO 60 cycle printouts upon completion of each cycle to verify the VAPROX Sterilant was not expired. The VAPROX Sterilant expiration date is identified near the top of each cycle printout following all sterilization cycles as notated in the following illustration: |  | 406 | 406 | Distributed in 44 states: AL, AK, AZ, AR, CA, CO, DE, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY; Puerto Rico, and the following countries: Algeria, Australia, Azerbaijan, Bahrain, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, Colombia, Costa Rica, Cyprus, Denmark, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India,Israel, Italy, Republic of Korea, Kuwait, Latvia, Malaysia, Mexico,Morocco, Myanmar, Panama,Peru, Philippines, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Switzerland, Thailand, Turkey, United Arab Emirates, and Vietnam. | Terminated | September 12, 2018 | 349 | 406
Z-0742-2018 | 79193 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Soft-Vu Omni Flush Angiographic Catheter (UPN H787107322015); SV OF 5F x 65 cm .035 NB PG; Catalog Number: 10732201 |  |  | September 22, 2017 | 2017 | Angiodynamics, Inc. | Product was placed into distribution prior to completion of all required post sterilization release activities. | Process control | Consignees were contacted by phone and by a letter sent via Federal Express. Consignees were instructed to stop using product, segregate affected product, and return the product to the firm along with a reply verification tracking form.For further questions, please call (518) 795-1676. |  | 28 | 28 | US Distribution to the states of : AL, CA, CT, FL, GA, IL, IN, MD, MI, MN, MO, OH, OR, TX. | Terminated | September 19, 2018 | 362 | 28
Z-0629-2018 | 78173 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | HemothermÂż Dual Reservoir Cooler/Heater System Models 400CE and 400MR |  |  | November 27, 2017 | 2017 | Cincinnati Sub-Zero Products LLC | Due to FDA's safety communication relating to nontuberculous mycobacteria (NTM) infections related to heater/cooler devices, an updated Instructions For Use includes a more robust cleaning &amp; disinfection procedure. | Device Design | Between November 27, 2017 - December 8, 2017, Cincinnati Sub-Zero Products LLC, distributed Urgent Medical Device Correction notices to their customers via Email and USPS. This Medical Device Correction has been initiated to provide a labeling update to the cleaning &amp; disinfection procedures. The update addresses the disinfection process, cleaning of reusable accessories, as well as clarifying the water type that should be used with the Hemotherm. Customers are encouraged to do the following: 1) Discard your current Hemotherm Operation &amp; Technical Manual and/or Operation Manual. 2) Replace your current Hemotherm Operation &amp; Technical Manual and/or Operation Manual with the enclosed updated Manual(s). For your convenience, all updates have been highlighted in the manuals. "For Hemotherm 400CE users, utilize the P/N 56418-R Operation &amp; Technical Manual and P/N 56075-G Operation Manuals. "For Hemotherm 400MR users, utilize the P/N 56234-S Operation Manual.All customers are encouraged to complete and return the enclosed response form by fax or email to FA2016-009@genthermcsz.com; Fax: 513-772-9119.Customers with questions or concerns, please contact the Field Action Coordinator: Kathy DeSmidt, E-mail: FA2016-009@genthermcsz.com, Phone: 513-772-8810 or 800-989-7373. |  | 1355 | 1355 units | US Nationwide Distribution | Terminated | July 02, 2018 | 217 | 1355
Z-0606-2018 | 78731 | No | LSI | Blastomyces dermatitidis serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Blastomyces dermatitidis Antiserum 1 mL vial GIBSON BIOSCIENCE. Catalog 200265 |  |  | January 20, 2017 | 2017 | Gibson Laboratories LLC | The firm information that the product was reporting false positive results | Mixed-up of materials/components | On January 20, 2017, the firm sent Urgent Medical Device Recall notification letter and response form to their customers. The letter identified the affected and the reason for the recall. Users are instructed not to use any Blastomyces dermatitidis (catalog # 200265) that they may have in inventory and to quarantine any affected product. Return any affected product per the instructions provided; and, complete and return the attached acknowledgement form. Questions regarding this recall should be directed to Penny Epperson, Quality Manager, at 859-303-5169 or pepperson@gibsonlabs.com. |  | 10 | 10/1 mL vials | Distributed to the states of AL, AZ &amp; MS. | Terminated | February 20, 2018 | 396 | 10
Z-0608-2018 | 79040 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Multi-Unit Abutment Angled, 17Âż, 4mm Cuff 4.1/4.5mm TiLobeÂż Implants, Catalog Number 45324K |  |  | June 20, 2017 | 2017 | Keystone Dental Inc | Abutments are unable to accept the final abutment screw | Other | On June 20, 2017 Keystone Dental telephoned impacted customers and advised them of the problem. On July 3rd, 2017 Keystone Dental issued notices (dated June 29, 2017) and response forms to customers via Fed Ex. Customers were encouraged to complete and return the response form attached form to confirm receipt of the replacement product and that lot 33006 can be returned to Keystone Dental. Customers with questions can call Johanna Drzyzga, Customer Relations Manager at 866-902-9272. |  | 20 | 20 units | US Nationwide Distribution in the states of ND, NJ &amp; MA | Terminated | April 09, 2018 | 293 | 20
Z-0615-2018 | 79131 | No | FAJ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Video Cystoscope models ECY-1570 and ECY-1570K |  |  | July 13, 2017 | 2017 | Pentax of America Inc | The video cytoscopes lack 510(k) premarket notification clearance. | No Marketing Application | Customer Instructions: Customers should inspect their stock of PENTAX Medical ECY-1570 and ECY-1570K video cystoscopes, quarantine them, and discontinue use of these products. If you have sold and distributed these products, please provide a copy of the notice and response form to the new owners. Enclosed with this letter is a customer response form. The form identifies the affected endoscopes (model and serial numbers) which have been sold to your facility. The attached form should be completed and returned to PENTAX Medical. Contact Information: If you have any questions regarding this action, please feel free to contact PENTAX Medical Customer Service. Tel: 800-431-5880 (8:30 AM  5:00 PM, Monday  Friday, EST) Fax: 201-799-4063 (alternate 201-391-4189) Email: customeradvisories@pentaxmedical.com |  | 24 | 24 | Distributed in 8 states: CA, IL, MD, MI, ND, NH, NY, SD, and Guam. | Open, Classified |  |  | 24
Z-0607-2018 | 78792 | No | MOZ |  |  | Orthopedic | PMA | Synvisc-One Âż (hylan G-F 20), 6mL delivered from a 10mL syringe |  |  | December 11, 2017 | 2017 | Sanofi Genzyme | Contamination with Methylbacterium thiocyanatum | Environmental control | On December 11, 2017 Sanofi Genzyme distributed URGENT VOLUNTARY DEVICE RECALL notices and reply cards to their customers via courier service. Customers were advised to inspect inventory, immediately discontinue use and return all affected product. On December 19, 2017 Sanofi Genzyme distributed a 2nd notice titled DEVICE RECALL UPDATE NOTICE (2 of 3) notifying customers that the isolated bacterial organism was Methylbacterium thiocyanatum. On December 22, 2017 Sanofi Genzyme distributed a 3rd notice titled DEVICE RECALL UPDATE NOTICE (3 of 3) which included antibiotic sensitivity analysis guidance for the isolated organism. Customers who wish to return product should contact GENCO Pharmaceutical Services at 1-855-838-5782 for a recalled product return kit. Customers are encouraged to complete and return the Business Reply Card via fax to 1-414-459-8791 or via email to recalls@genco.com. Sanofi Genzyme will issue you a credit or replacement product upon receipt of the recalled product by GENCO. All questions of a medical or clinical nature should be directed to Sanofi Medical Information Services in the U.S. at 1-800-633-1610, option 1. Sanofi US Customer service can be reached at 1-800-633-1610 option 7 then 4. |  | 12380 | 12,380 units | US Nationwide distribution(36 states) | Open, Classified |  |  | 12380
Z-0614-2018 | 78963 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis oneFloor-mounted system for uncompromised imaging; Artis one, Interventional fluoroscopic x-ray system, Artis one is an angiography system developed for diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. X-ray generator POLYDOROS A100 Plus |  |  | August 22, 2017 | 2017 | Siemens Medical Solutions USA, INC | The possibility exists for the monitor display of Artis one systems with a DP Transceiver below revision 03 to experience blinking and/or a blue/black screen | Nonconforming Material/Component | Siemens sent an Urgent- Medical Device Correction letter dated August 22, 2017. Our service organization will contact you shortly to arrange a date to perform this corrective action (A system hardware update will be implemented and all affected DP Transceivers below revision 03 will be updated). Following the system hardware update, the cause is eliminated and a recurrence of the fault is prevented. For further questions, please call (610) 448-6461. |  | 207 | 207 worldwide (13 distributed in US) | US Distribution to the states of : AR, MO, NY, IN, AR, TX , NC, PA, MI, LA, WA, CA and AL. | Terminated | August 21, 2018 | 364 | 207
Z-0597-2018 | 79067 | No | OSH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Response 5.5/6.0 Spine System 7.0mm x 40mm |  |  | August 03, 2017 | 2017 | OrthoPediatrics Corp | The 7mm diameter uniaxial pedicle screws have been color anodized with the same color utilized on 5 mm diameter screws. | Employee error | On August 4, 2017 an urgent medical device recall letter, titled "Response 5.5/6.0 Spine System - 7.0mm x 40mm UNIAXIAL PEDICLE SCREW" was issued to customers with instruction to quarantine affected product, which will be returned to the recalling firm for destruction. Questions or concerns can be directed to logistics@orthopediatrics.com or calling (574) 268-3692 |  | 24 | 24 units | US Nationwide Distribution | Completed |  |  | 24
Z-0596-2018 | 79103 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Persona Partial Knee System |  |  | August 28, 2017 | 2017 | Zimmer Biomet, Inc. | Potential for the Persona Partial Knee Impactor Pad to fracture. | Under Investigation by firm | Consignees were notified of the recall via letter on about 08/28/2017. Instructions for distributors, sales representatives, and distributor operations leaders include to ensure affected team members are aware of the recall, review the list of hospitals included and note any not included that should be notified, and complete and return the Acknowledgement Form. Instructions for risk managers and surgeons include to ensure affected personnel are aware of the recall, ensure that the provided updated surgical technique is used, and complete and return the Certificate of Acknowledgement. Further questions or concerns can be sent to customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Questions can also be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 1275 | 1,275 units total | US Nationwide Distribution in the states of AZ, CA, CO, ID, IL, IN, MA, MD, MI, NJ, NM, PA, SC, TX, WA, WI. Distributed internationally to Belgium, France, Germany, Italy, Japan, Netherlands, Spain, Sweden, Switzerland, UK. | Terminated | August 23, 2018 | 360 | 1275
Z-0613-2018 | 78759 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | A.T.S. 2200TS Tourniquet Systems, Personalized Pressure. Zimmer Surgical, Inc. Dover, OH 44622 |  |  | July 27, 2017 | 2017 | Zimmer Surgical Inc | Updated on-device label and 3 pages of an updated Operator/Service Manual. Also, decrease from 3.8 meters (m) to 1.0 separation distance between the device and RF communication equipment. | Use error | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE CORRECTIONS" letter dated 7/27/2017 to its customers via email or mail. The letter describes the product, problem and actions to be taken. Instructions for distributors include ensure all affected personnel are aware of the recall, confirm receipt of new device label and updated Operator/Service Manual, place the label on the device in the location specified in the notice, replace the 3 pages of the Operator/Service Manual with the updated pages, review the hospital list of customers that accompanied the notice and return contact information for any additional hospitals product was distributed to. Instructions for Risk Managers at hospitals include ensure all affected personnel are aware of the recall, confirm receipt of new device label and updated Operator/Service Manual, place the label on the device in the location specified in the notice, replace the 3 pages of the Operator/Service Manual with the updated pages, and complete and return the Attachment 1 form via email to corporatequality.postmarket@zimmerbiomet.com or ZimmerBiomet8287@stericycle.com or fax to 866-338-2611.If after reviewing this notification you have further questions or concerns please call +1(330) 364-0989 between 8:00 am and 5:00 pm EST, Monday through Friday.Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an on-call representative in the case of an emergency. |  | 8231 | 8,231 devices total | Worldwide Distribution: US (nationwide) to AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, including Guam; and countries to Australia, Brazil, Canada, Chile, China, Denmark, Dominican, El Salvador, England, Finland, France, Hong Kong, India, Italy, Japan, Malaysia, Mexico, Netherlands, Panama, Singapore, Sweden, Taiwan, and Thailand.. | Open, Classified |  |  | 8231
Z-0616-2018 | 78237 | No | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Purely Yours Breast Pump. Model 17070PMW and Model 17070ACA. |  |  | July 10, 2017 | 2017 | AMEDA, INC. | Two specific lots was incorrectly manufactured containing a UK power adapter. | Employee error | Customers were notified initially by phone call on 07/10/2017 and "Device Recall" letters were mailed, via FedEx overnight delivery on 07/10/2017. Instructions included to place any affected product in quarantine, complete and return Correction Response Form, return any product in inventory, notify customers if further distributed, and that replacement product will be provided upon return of their customer's completed Correction Response Form. Any questions can be directed to Ameda at 877-992-6332, press #1. The recall was expanded on 08/18/2017 and customers were notified by letter, via FedEx, on 08/18/2017 through 08/22/2017. Instructions are the same as the initial notification above. |  | 3327 | 3,327 units total | Distributed to CA, DC, DE, FL, GA, IL, MD, NC, NJ, OH, TX, UT, VA, and WA. | Open, Classified |  |  | 3327
Z-0626-2018 | 79104 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | G7 DUAL MOBILITY LINER 44MM F |  |  | August 25, 2017 | 2017 | Zimmer Biomet, Inc. | The affected products are labeled and etched as F-44 mm liners; however, the physical products could potentially be E-42 mm liners. | Mixed-up of materials/components | On August 30, 2017 an Urgent medical device recall removal notice was mailed to customers identified as "G7 Dual Mobility Liner". The recall letter requests the following actions to be taken:1. Review this notification and ensure affected team members are aware of the contents.CF04107 Rev. 2, Eff. Date: 21 Feb. 2017 Ref. CP04102 Recall Activities Page 2 of 4ZFA 2017-3472. Immediately locate and quarantine affected product in your inventory.3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form.a. Return a digital copy to corporate quality.postmarket@zimmerbiomet.com within three (3) days.4. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet.Questions or concerns can be directed to Customer Service at 1-574-371-3071 |  | 8 | 8 | US Nationwide Distribution | Terminated | August 16, 2018 | 356 | 8
Z-0623-2018 | 79092 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | The ELAN 4 Air System is intended for high speed cutting, sawing and drilling of bone in fields of Spine, ENT, Neuro, and Maxillofacial Surgery. |  |  | September 01, 2017 | 2017 | Aesculap Implant Systems LLC | The ELAN 4 Air System was released to the market place prior to an authorized FDA 510K clearance. This issue does not affected sales outside the United States (OUS). | No Marketing Application | On September 1, 2017 an Urgent Medical Device Recall letter was issued to customers for the ELAN 4 Air System which includes the following sub-components:ST0607 ELAN 4 Air Control SetST0608 ELAN 4 Air Neuro Crani SetST0609 ELAN 4 Air Spine SetST06411 ELAN 4 Air HD Spine SetST0614 ELAN 4 Air Micro Saw SetST0615 ELAN 4 Air Hand Control SetST0638 ELAN 4 Air Large Saw SetThe letter directs customers to take the following actions:1.Review the Voluntary Urgent Medical Device Recall Notification Letter in its entirety and ensure that all users in your organization and other concerned persons are informed about this recall and immediately discontinue use. 2.Determine your current inventory of affected products within your facility. Immediately remove the affected products and return to Aesculap Inc. Do not destroy any affected products. 3.Utilizing the attached Product Recall Acknowledgement Response Form on the next page of this letter, record the total number of products in your possession. If you have no inventory remaining, please enter zero (0) on the form and a REASON WHY (example, returned evaluation system to sales representative). 4.Please contact Customer Service and they will provide you with a Returned Goods number and return label.Customer Service: 1-800-282-9000aesculap_emailorders.bbmus_service@bbraun.com |  | 167 | 167 Units | Nationwide Distribution. | Open, Classified |  |  | 167
Z-0624-2018 | 79089 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | JS Series SterilContainer S2 System |  |  | September 08, 2017 | 2017 | Aesculap Implant Systems LLC | The JS Series SterilContainer S2 System was released to the market place prior to an authorized FDA 510K clearance. No other Aesculap SterilContainer System products are impacted by this recall. | No Marketing Application | On August 24, 2017 a Medical Device Recall(Reference 2916714-08/24/2017-010R) was issued requesting that all customers take the following steps: 1.Review the Urgent Medical Device Recall Notification Letter in its entirety and ensure that all users in your organization and other concerned persons are informed about this recall and immediately discontinue use. 2. Determine your current inventory of affected containers within your facility. Immediately quarantine the affected products. A sales representative will come to your facility and remove your affected product and immediately return it to Aesculap Inc. They will assist you with ordering an appropriate replacement container. Do not destroy any affected products.3. Please contact Customer Service and they will provide you a Returned Good number and a return label.Customer Service: Phone: 1-800-282-9000 aesculap_emailorders.bbmus_service@bbraun.com |  | 292 | 292 units | Nationally | Open, Classified |  |  | 292
Z-0620-2018 | 79068 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | ELAN 4 Drill: ELAN 4 1-RING TWIST DRILL D1.5 |  |  | September 05, 2017 | 2017 | Aesculap Implant Systems LLC | ELAN 4 burr/blade accessories were marketed without 510(k) premarket notification clearance. | No Marketing Application | Notifications were sent requesting the customers immediately discontinue use and quarantine affected products. The notification explains that a sales representative will visit the facilities to remove affected product and return it to the firm. |  | 6211 | 6211 | Distributed to 4 hospitals in the following states: NM, OH, TX, WV | Open, Classified |  |  | 17195
Z-0625-2018 | 79001 | No | NDC | Predictive pulmonary-function value calculator. | Anesthesiology | General Hospital | 510(k) | Accu-Chek Connect Diabetes Management App, Catalog number 07562462001 / GTIN number 00365702700000 &amp; Catalog number 07250452001 / GTIN number 00365702700017 |  |  | June 08, 2017 | 2017 | Roche Diabetes Care, Inc. | Certain iOS and Android App versions contain a program error (bug) in the Bolus Advisor feature. Due to a software bug, when the OS region of the phone setting is changed, the unit of measure within the app may unexpectedly change. This creates a risk the app might not transfer the blood glucose result or the user might not correctly input numerical values for carbohydrate used for bolus advice. | Other | On 06/13/2017, Roche Diabetes Care, Inc. made the Urgent Medical Device Correction notice available to their customers through the Accu-Chek.com website and through the Accu-Chek Connect microsite. Patients are instructed to:Update to version 2.1.1 of the Accu-Chek Connect appClear all data from the Accu-Chek blood glucose meter following the steps below, before pairing with the Accu-Chek Connect app for the first time. This clearing of the records in the memory can be accomplished on the PC or Mac by using Accu-Chek Connect Device Link. A micro-USB cable and an Accu-Chek Connect Online account are required. Also ensure that the date and time are set correctly on the blood glucose meter before pairing with the app. Failure to follow these directions might impact first time use of the apps Bolus Advisor feature.Âż Steps to clear data from Accu-Chek meter in Connect Online:1. Create or login to your Connect Online account on your PC or Mac.2. Download Accu-Chek Device Link3. Right-click on Device Link icon in system tray or dock.4. Select Settings &gt; Prompt after every successful upload under the Delete Patient Results from Device heading.5. Connect your meter via micro-USB cable.6. Follow the onscreen instructions to upload your data and clear your meter.If you use the apps Bolus Advisor feature, please carefully review and manage your active insulin and insulin bolus amounts for the first 8 hours after initial installation of the app and pairing your meter. This ensures the software will establish an accurate history for use in future bolus recommendations.They are advised to contact Accu-Chek Customer Care at 1-800-688-4578 if there are additional questions. |  | 67040 | 67,040 unique users of version 2.1.0 | US Nationwide | Terminated | July 27, 2018 | 414 | 67040
Z-0628-2018 | 79095 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AGC 2000 LP Patella, labeled as AGC Porous Patellar Med, Model Number 150804. |  |  | August 25, 2017 | 2017 | Zimmer Biomet, Inc. | The affected AGC 2000 LP Patella products were over-molded leaving the outside diameter nonconforming to the print dimension. | Process change control | Consignees were notified via mail on approximately 08/25/2017. Instructions to distributors, sales representatives, and distributor operations managers include to ensure affected team members are aware of the recall, locate and quarantine affected product in inventory, complete and return the Certification of Acknowledgement, return all affected product, and provide any additional hospitals and/or surgeons that received direct shipments. Instructions to risk managers and surgeons include to ensure affected personnel are informed of the recall, quarantine all affected product, return any affected product in inventory, and complete and return the Certificate of Acknowledgment. Questions or concerns can be sent to the Zimmer Biomet customer call center at 1-574-371-3071 between 8:00 am and 5:00 pm EST, Monday through Friday. Or they can be sent via email to corporatequality.postmarket@zimmerbiomet.com. |  | 11 | 11 devices | US Nationwide Distribution to IN, MO, NY, TX. | Terminated | August 09, 2018 | 349 | 39
Z-0622-2018 | 78997 | No | HBO | Scalp clip. | Neurology | Neurology | 510(k) | Model #: CG8900, Medtronic Clip Gun Scalp Clip Kit (Qty 1); Model #: CG8901, Medtronic Clip Gun Scalp Clip Kit (Qty 10); Model #: CG8901A, Medtronic Clip Gun Applicator (Qty 10) |  |  | December 20, 2017 | 2017 | Medtronic Neurosurgery | Medtronic is initiating a voluntary recall because the products identified above did not adequately demonstrate acceptable packaging performance to current standards. | Package design/selection | A Consignee Notification with response form was sent to affected consignees on 12/20/2017. |  |  | 27,432.8 devices | Algeria, Armenia, Australia, Canada, Cyprus, Egypt, Germany, Ghana, Greece, Guam, Iceland, Indonesia, Iran, Italy, Jordan, New Zealand, Poland, Portugal, Puerto Rica, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Swaziland, Sweden, Switzerland, Syria, Tunisia, Turkey, UAE, UK, USA, Yemen | Open, Classified
Z-0618-2018 | 79043 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Norian Drillable Inject |  |  | August 10, 2017 | 2017 | Synthes USA | The contract supplier notified DePuy Synthes on June 16, 2017 that the kit expiration dates do not align with the earliest expiration date of the component lots within the kit. Some kit expiration dates were later than the earliest expiration date of the component lots. The expiration dates of some components are earlier than that listed on the kit. | Process control | On August 15, 2017 an Urgent Medical Device Recall notice, titled "895204 NORIAN DRILLABLE INJECT" was issued to customers instructing them to review inventory and quarantine all affected product subject to the recall. For questions. concerns or to obtain a return authorization Call 1-800-479-6329. Returned product will be credited. |  | 5 | 5 units | Nationwide Distribution. | Terminated | September 10, 2018 | 396 | 244
Z-0634-2018 | 79050 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Spectrum Central Venous Catheter Tray Minocycline/Rifampin Antibiotic Impregnated, Catalog Identifiers: C-UTLMY-501J-ABRM-HC-FST and C-UTLMY-501J-ABRM-HC-IHI-FST |  |  | July 31, 2017 | 2017 | Cook Inc. | Central venous catheters may not fit over guide wire potentially causing procedure delay. | Process control | Cook sent an Urgent Medical Device Recall dated July 31, 2017 to affected customers. The letter identified the affected product problem and actions to be taken. The letter instructed customers to (1) Examine inventory immediately to determine if you have affected products and quarantine affected products. (2) Return the affected products to Cook Medical with a copy of the Acknowledgement and Receipt Form to receive a product credit. (3) Even if you do not have affected products on hand, you must still complete the Acknowledgement and Receipt Form and send it via fax (812.339.7316) or email (FieldActionsNA@CookMedical.com). (4) Complete and return the Acknowledgement and Receipt Form attached to this letter within 30 days. (5) Report adverse events to Cook Medical Customer Relations by phone at 800.457.4500 or 812.339.2235, Monday through Friday between 7:30 am and 5:00 pm (Eastern time), or by email via CustomerRelationsNA@CookMedical.com. |  | 304 | 304 | Worldwide distribution - US Nationwide in the states of: AL, CA, CO, FL, GA, IA, IL, KY, LA, MD, MI, MN, MO, NC, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA, WI; the District of Columbia and the countries of Belgium, India Italy, Spain, and United Kingdom. | Open, Classified |  |  | 539
Z-0637-2018 | 79052 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom sterile surgical packs labeled as follows: |  |  | January 22, 2018 | 2018 | DeRoyal Industries Inc | DeRoyal manufactured custom sterile surgical kits containing the Terumo Over Pressure Safety Valve, which were subsequently recalled by Terumo CVS due to reports of 'No Flow' through the OPS valve. | Process change control | The firm, DeRoyal, sent an "URGENT! RECALL NOTICE" letter dated 01/22/18 to its Consignees on 1/22/18 by overnight mail. The firm sent a follow-up "URGENT! RECALL NOTICE" letter dated 1/29/18. The letters described the product, problem and actions to be taken. The Consignees were instructed to identify affected products; place in quarantine; complete and return the Notice of Return Form, even if they no longer held any affected product, to DeRoyal via fax to 865-362-3716 or email to recalls@deroyal.com no later than 3/1/18, and return any product they held in stock. If you have further distributed these products, notify your customers of this recall.If you have questions or need assistance with the recall, please contact Regulatory Recall Coordinator and Quality Projects Manager at jmarsh@deroyal.com or by phone 865-362-4203 or email us at recalls@deroyal.com. |  | 181 | 181 units | US Distribution to: NY | Open, Classified |  |  | 181
Z-0639-2018 | 78944 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | ROSA Spine 1.0.2 Stereotaxic instrument Computer-assisted surgical device. |  |  | May 01, 2017 | 2017 | Zimmer Biomet, Inc. | Robot arm being sent to the wrong position | Software design | Urgent Medical Device Recall (Correction) notices were sent to customers.The letter provided to hospital risk managers and surgeons identified the issue and their responsibilities. These responsibilities included:- Stop using the micromovement function in isocentric mode during Brain Surgery with ROSA Brain 3.0 device. -Reviewing the notification to ensure comprehension of the steps required to be taken prior to the permanent correction that will be implemented on site by a service engineer.- Completing and returning Attachment 1 (Certificate of Acknowledgement) to the e-mail address provided.The firm has prepared a software correction that will be deployed and implemented on site by a MEDTECH representative. |  | 1 | 1 | Distributed to CA, DC, FL, IL, KY, MA, MN, NC, NY, OH, and PA. | Open, Classified |  |  | 17
Z-0635-2018 | 78606 | No | MXK | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | OCULUS Pentacam AXL, Model 70100, software versions 1.21r01, 1.21r03, 1.21r11, and 1.21r13. The firm name on the label is OCULUS Optikgerate GmbH, Made in Germany. |  |  | November 21, 2017 | 2017 | Oculus Optikgeraete GMBH | The device software versions have an anomaly which may produce an erroneous marking for the quality specification value | Software design | Oculus issued an email on November 21, 2017 to their U.S. subsidiary which contained a Mandatory Software Update notification. The email informed the U.S. subsidiary that a mandatory software update was planned for the Pentacam AXL in possession of their customers, provided the software versions affected, and the reason for the mandatory update. The email explained the USB thumb drive will be shipped to the U.S. subsidiary for each serial number that may be affected. The U.S. subsidiary was to match the serial numbers to their customers and complete the attached customer letter for issuance to each of the customers. The U.S. subsidiary is then supposed to issue the letter and the USB thumb drive via certified mail to each of their customers who have the affected serial numbers. The customer is to send back the completed, signed form verifying the mandatory software update has been installed on all working places on which the Pentacam AXL software is running or that the procedure was performed by an authorized distributor.For further questions, please call (425) 670-9977 |  | 83 | 83 devices | US Distribution to the states of : CA and FL. There were no foreign/military/government accounts. | Terminated | March 22, 2018 | 121 | 83
Z-0636-2018 | 79169 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Endo GIA Auto Suture Universal Articulating Loading Unit for use in abdominal, gynecologic, pediatric and thoracic surgery for resection, transection and creation of anastomosis. |  |  | September 13, 2017 | 2017 | Covidien LLC | Potential for failure of the device safety interlock. The safety interlock prevents an empty single-use loading unit from being fired a second time. | Under Investigation by firm | All consignees were notified via Federal Express or certified mail on September 13, 2017, and the letter informs customers of the potential for failure of the device safety interlock and the actions they should take. Customers are requested to acknowledge receipt and understanding of the Urgent Field Corrective Action Notice. All in control inventory and product returned from customers will be scrapped. |  | 3816 | 3,816 | Nationally | Open, Classified |  |  | 3816
Z-0641-2018 | 79205 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Allen Medical Bow Frame Model A-70800. Surgical platform used to position patients for lumbar spine procedures. |  |  | October 06, 2017 | 2017 | Allen Medical Systems | The base of the frame has the potential to crack when the patient is initially transferred to the Bow Frame. | Device Design | In October of 2017 an Urgent Medical Device Correction letter was sent to customers notifying them of the recall and requesting customers to immediately remove any frames from use that are in the serial number range subject to the recall and that a service technician will coordinate correction of the device. If the device was sent to another company, to forward the recall notice o that person or firm. If you have any questions concerning this Urgent Medical Device Correction, please contact Allen Medical Customer Service at (800) 433-5774, 8:30  5:00 pm EST. |  | 161 | 161 units total | Worldwide distribution. US Nationwide, Australia, Bangladesh, Brazil, Canada, China, Denmark, France, Germany, Greece, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia, Netherlands, Spain, Switzerland, Taiwan, UAE, and UK. | Open, Classified |  |  | 161
Z-0899-2018 | 79199 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Silicone Fluted Round Drain 24Fr (hubless) &amp; Connector, Sterile. Catalog Number JP-2234. |  |  | October 03, 2017 | 2017 | Degania Silicone, Ltd. | Sterile pouches are missing labels on them. The boxes themselves are properly labeled. The issue is related to the production lot P1607442 only. | Error in labeling | An Urgent Medical Device Correction letter, dated October 03, 2017, was sent to the customer. The letter identified the affected device and the reason for the recall. The letter instructed the consignee to check its stock in order to remove all mislabeled product from the market and eliminate potential risk of misuse. The customer was instructed to return the completed Confirmation Sheet to Degania Silicone, as well as any unlabeled product. Questions regarding the recall should be directed to Katia Kayam, Customer Complaints Officer at Katia@ds-il.com. |  | 4000 | 4,000 units | US Distribution to Puerto Rico. | Open, Classified |  |  | 4000
Z-0643-2018 | 79237 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Pentax Video Duodenoscope Model: ED-3490TK (UDI of design being recalled: 04961333232420) |  |  | February 13, 2018 | 2018 | Pentax of America Inc | The duodenoscopes are being recalled in order to replace the forceps elevator mechanism, the O-rings, and the distal end covering to be consistent with the updated design as well as provide an updated periodic inspection as part of the Operation Manual in order to mitigate the potential risk of infection in flexible endoscopy. | Device Design | Pentax Medical sent an Urgent Medical Device Correction/Removal letter dated February 7, 2018, to the United States Customers. The consignee letter includes a customer response form and a revised operator manual with the added recommended periodic maintenance. The letter requests the return of the form which includes an accounting of the devices (by serial number) owned by the facility. The firm is to contact the consignee to arrange return of the affected units and to provide loaner devices as needed. For further questions, please call 1 (800) 431-5880 ext. 2064. |  | 559 | 559 | USA (nationwide) Distribution to the states of including Puerto Rico : AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NV, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI, WV and WY, and the District of Columbia. | Open, Classified |  |  | 559
Z-0644-2018 | 78946 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Excella II Polyaxial Pedicle Screw used as a pedicle screw system in skeletally mature patients, is intended to provide immobilization and stabilization of spinal segments, as an adjunct to fusion, in treatment of the following acute and chronic deformities of the thoracic, lumbar, and sacral spine: |  |  | January 08, 2018 | 2018 | Innovasis, Inc | Nine lots of 8.5 mm Polyaxial Pedicle Screws are at risk of breakage. | Nonconforming Material/Component | The distributors were notified of the recall on 1/8/2018 via email with an attached pdf. They are also being asked to confirm or deny the quantities showing in inventory via a Product Removal Form sent in the same email. |  | 1010 | 1,010 parts | United States | Terminated | February 27, 2018 | 50 | 1010
Z-0647-2018 | 79197 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Pressure Injectable ARROWg+ard Blue PLUS Multi Lumen Central Venous Catheterization Kit; Product Code: ASK 42703 PUHC; Exp. Dates Feb 2018 - Nov 2018 |  |  | October 17, 2017 | 2017 | Arrow International Inc | These finished good kits may contain the incorrect Springwire Guide (SWG). | Process control | On October 17, 2017 an Urgent Medical Device recall letter from Arrow International was issued to customers requesting that customers review their inventory for product, which should immediately be discontinued and quarantined. The products are then to be returned to the firm. Questions or concerns can be directed to: Customer Service at 18662466990. |  | 5926 | 5,926 total | Distribution US Nationwide | Open, Classified |  |  | 5926
Z-0645-2018 | 79058 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Medtronic AFFINITY NT, Cardiotomy/Venous Reservoir (CVR) with Filter, 540 |  |  | January 16, 2018 | 2018 | Medtronic Perfusion Systems | Possible sterile barrier breach in the pouch of Affinity NT Cardiotomy Venous Reservoir stand alone uncoated and Affinity NT Cardiotomy Venous Reservoir stand alone Trillium coated products. | Process control | Medtronic sent an Urgent Medical Device Recall Letter dated January 16, 2018. Customer Actions:Please review your inventory for product affected by this issue.1.) Immediately identify and quarantine all unused, non-expired, affected product in your inventory.2.) Return all unused, affected product in your inventory to Medtronic by contacting Customer Service at1-800-854-3570 Option 4, or accessing My Orders and referencing this communication to initiate a return and credit of unused product. Your Medtronic sales representative can assist you in the return of affected product as necessary.3.) Complete the enclosed Customer Confirmation Certificate and scan and email to RS.CFQFCA@medtronic.com.This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the potentially affected products have been transferred. For further questions, please call (763) 526-2494. |  | 8644 | 8644 units | Worldwide Distribution - US Distribution and to the countries of : Australia, Canada, Colombia, Denmark, Germany, Guyana, Israel, Italy, Japan, Netherlands, Norway, Saudi Arabia and United Kingdom. | Open, Classified |  |  | 12508
Z-0829-2018 | 79239 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008T, Hemodialysis Delivery System |  |  | October 05, 2017 | 2017 | Fresenius Medical Care Renal Therapies Group, LLC | While reviewing documentation for the next software release of the 2008T, an R&D; technician identified that setting the UF goal to '0' introduces a discrepancy between the UF rate displayed and the actual UF pump rate. The software anomaly is also applicable to the 2008T Machines that contained the 2.63 Bug Fix per DCAF 17-088. 2008T Upgrade kits that contained the 2.63 Bug Fix are also affected. Additionally, SW version 2.64 (in design freeze) is also impacted by this anomaly. | Software Design Change | Fresenius Medical Care sent an Urgent Medical Device Correction letter dated April 24, 2017 to affected customers. the letter identified the affected product, problem and actions to be taken. For questions contact Technical Services at 1-800-227-2572. |  | 15 | 15 machines | US Nationwide Distribution | Open, Classified |  |  | 15
Z-0659-2018 | 79078 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Aspire HD (for FDR-1000AWS) |  |  | August 04, 2017 | 2017 | Fujifilm Medical Systems U.S.A., Inc. | FUJIFILM Medical Systems U.S.A., Inc. (FMSU) identified a potential failure with our Mammography system.The acquisition workstations, FDRÂż1000AWS, FDR-2000AWS, FDR-3000AWS, and CRÂżIR363AWS assign aunique ID number to every image study but very rarely, with the acquisition workstation software versions,VS.O, VS.1, VS.2 , V6.0, V6.1, and V7.0, the system may assign the same ID number to a new set of imagesthat was already assigned to the previous set of images due to the error in ID number generation logic. If anImage with this error (with duplicate ID number) Is transmitted to PACS, it may overwrite the image alreadystored on PACS. | Software design | The firm, FUJIFILM Medical Systems U.S.A., Inc. (FMSU), sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated 7/28/2017 to its customers. The letter described the products, problem and actions to be taken. The customers were instructed to please read and follow the instructions in the "ACTIONS TO BE TAKEN BY CUSTOMER/USER" section. If you have an error, contact your local FUJIFILM office; and complete and return the Customer Feedback Form.FUJIFILM service personnel will contact all of the medical facilities where the applicable products have been installed to arrange for this correction and visit to take corrective measures..If you have any questions about this matter, call Recall contact (203) 276-3445 or email: dharaben.desai@fujifilm.com. |  | 71 | 71 units (68 units Domestic. 3 units Foreign) in total | Worldwide Distribution. | Open, Classified |  |  | 71
Z-0640-2018 | 79154 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | GastriSail" Gastric Positioning Device |  |  | September 21, 2017 | 2017 | Covidien LLC | Reports of gastric or esophageal perforation occurring in procedures where the GastriSail" Gastric Positioning device has been used. | Under Investigation by firm | Medtronic sent an URGENT MEDICAL DEVICE RECALL letter dated September 20, 2017 to their customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to review inventory and notify firm if any product remains at their facility. Arrangements for product return would then be made. For questions or concerns can be directed to customerservice@Covidien.com |  | 17271 | 17,271 units | US Nationwide Distribution | Terminated | September 26, 2018 | 370 | 17271
Z-0664-2018 | 79000 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Sorin Group Smart Perfusion Pack E-Pack, REF 088511400, Rev: L, Single Use, Sterile, Rx. The firm name on the label is Sorin Group USA, Inc., Arvada, CO. |  |  | December 19, 2017 | 2017 | LivaNova USA | The product is labeled with the incorrect expiration date. | Process control | The recalling firm issued a letter dated 12/7/2017 via email on 12/19/2017. |  | 16 | 16 units (6 from lot 1717800062 and 10 units from lot 1720800071 | Distribution was made to one consignee in LA. | Completed |  |  | 16
Z-0670-2018 | 78298 | No | OZO |  |  | Clinical Chemistry | PMA | Medtronic CareLink iPro Therapy Management Software, MMT-7340 |  |  | October 10, 2017 | 2017 | Medtronic Inc. | Due to a time conversion error, data uploaded from the iPro(TM)2 recorder to the CareLink iPro software using the new CareLink iPro Uploader feature may result in incorrect reports. This pertains only to reports generated between September 23, 2017 and September 27, 2017 that used the new CareLink iPro Uploader feature. | Software Manufacturing/Software Deployment | A communication is being sent to all healthcare professionals who had uploaded reports using the universal uploader feature between September 23, 2017 to September 27, 2017. The method is by mail and by email, with a response requested from the healthcare professional to confirm they have received and read the notice. For further questions please call (877) 874-7717 |  | 488 | 488 uploads | Worldwide Distribution - US Distribution | Open, Classified |  |  | 488
Z-0672-2018 | 79246 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Disposable Curved Circular Staplers |  |  | October 23, 2017 | 2017 | Ethicon Endo-Surgery Inc | There is a potential for the devices to create staples that do not meet formed staple height specifications which can result in malformed staples. | Under Investigation by firm | Ethicon Johnson &amp; Johnson Affiliates in Asia Pacific will notify their impacted consignees by providing the customer letter through post, email or sales representative visits. The letter identified the affected product, problems and actions to be taken. |  | 148.913 | 148.913 | China | Open, Classified |  |  | 148.913
Z-0671-2018 | 79182 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Protura Software which utilizes Elekta's iCOM interface: MT6XSM1.4.0, MT6XSM1.4.0-1, MT6XSM1.4.0-2, MT6SXM1.4.0-3, MT6XSM1.5.0-2, MT6XSM1.6.0-1, MT6XSM1.7.2, MT6XSM1.7.2-1, MT6XSM1.7.2-3, MT6XSM1.7.3, MT6XSM1.7.3-1, MT6XSM1.7.3-3. |  |  | November 08, 2017 | 2017 | Med Tec Inc | When an error message remains displayed and is not cleared in the Protura software with the Elekta pedestal coordinates, moving the pedestal could cause the Protura software to not update with the pedestal location and buffers the pedestal movement history. | Other | Customers were notified via email on about November 8, 2017. Instructions include to contact Civco Medical Solutions to arrange an update to the device. Until the update is accomplished, customers were instructed to review and acknowledge the notification, verify the actual pedestal location matches the pedestal location being locked into Protura software, and disable the auto load/zero functionality described in the IFU. For further questions, please call (712) 737-8688. |  | 5 | 5 installations | US Distribution to Iowa and South Dakota. Shipped internationally to the Netherlands. | Open, Classified |  |  | 5
Z-0674-2018 | 79161 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Complete Knee System, Vanguard CR Lipped Tibial Bearing 10x71/75, Item Number: 183540 |  |  | September 29, 2017 | 2017 | Zimmer Biomet, Inc. | Knee implant components (tibial bearings) are labeled with the incorrect size. | Labeling mix-ups | Zimmer Biomet sent an Urgent Medical Device Recall letter dated September 29, 2017 to their customers. The letter identified the affected product, problems and actions to be taken. Customers were instructed to remove the affected product from inventory and provide to their Zimmer sales representative or, if the consignee is a distributor, to ship the affected product back to Zimmer. Consignees were also asked to complete a recall response form and send this back to the firm. Returns are being physically destroyed.If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 24 | 24 | Worldwide Distribution - US Nationwide in the states of CA, GA, KS, LA, MA, MI, MN, OH, TX, VA, WA, WI, and the countries of Netherlands | Open, Classified |  |  | 24
Z-0668-2018 | 79064 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Heliomolar Refill 20x0.25g 110T, Product code 550559AN, Catalog number 550559 |  |  | August 03, 2017 | 2017 | Ivoclar Vivadent, Inc. | Complaints were filed claiming the material consistency was too soft and too sticky in comparison to its usual consistency, causing unexpected handling for the user during packing and contouring. | Nonconforming Material/Component | The firm sent letters to consignees on August 3, 2017 instructing them to segregate the recalled product and arrange for a replacement of stock. |  | 2 | 2 units | US and Canada | Terminated | August 28, 2018 | 390 | 6654
Z-0676-2018 | 78832 | No | LKV |  |  | Clinical Chemistry | PMA | Rapid for the TLiIQ(R) System, HOLOGIC(R), Catalog number: 01200 |  |  | November 29, 2017 | 2017 | Hologic, Inc. | Rapid fFN cassettes were manufactured with material (A137 ascites) from a new supplier (SDIX). The products were distributed to customers before FDA approved the change to use the material from SDIX. | Mixed-up of materials/components | Hologic initiated the recall on Wednesday, November 29, 2017. The method of notification was via overnight FedEx. |  | 165 | 165 kits | United States (US), Guam (GU), Japan (JP) | Terminated | August 22, 2018 | 266 | 165
Z-0678-2018 | 78951 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Nutrisafe 2 35 mL Syringe for Safety Connection. Product Code: 1015.352M |  |  | July 24, 2017 | 2017 | Vygon MFG, Inc., dba/ Churchill Medical Systems, Inc. | Vygon is conducting a voluntary recall of the products and lots listed in this recall, beinginitiated because our Tyvek supplier has notified Vygon of issues with their product, specifically theTyvek may not be coated completely causing gaps in the adhesive used to seal the product packaging.This potential lack of adhesive may affect the sterility of the product. | Under Investigation by firm | Vygon sent recall letters dated July 24, 2017, to customers via email and fax with read receipts and delivery confirmations. The recall letter informed customers of the reason for recall, product description, code and lot number, and instructions to customers, with a Recall Acknowledgement and Inventory Return Form provided/accompanying the recall letter. |  |  |  | U.S. | Open, Classified
Z-0663-2018 | 78960 | No | NJL |  |  | Orthopedic | PMA | NexGen LPS Femoral Component, Left and Right, Size G |  |  | July 24, 2017 | 2017 | ZIMMER ORTHOPEDIC MFG LTD | There is a possibility that the protective foam insert used during the packaging of the affected products is undersized, which may cause a breach in the inner cavity during transportation. | Reprocessing Controls | On approximately July 24, 2017, customers were notified via letter of the recall. Instructions for distributors included to ensure affected personnel are aware of the recall, locate and quarantine any affected product in inventory, complete and return the Inventory Return Certification Form, make arrangements to return the product Zimmer Biomet, and provide customers if product was further distributed. Instructions for Risk Managers at medical facilities include to ensure affected personnel are aware of the recall, to assist the Zimmer Biomet sales representative in identifying and quarantining affected devices so they can be removed by the sales representative, and to complete and return the Certification of Acknowledgement form. Further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. or 1 (800) 613-6131 |  | 7 | 7 devices | Worldwide Distribution - US Distribution to the states of : CA, MA, ME, MI, MN and WI., and to the countries of Germany, France and Italy. | Terminated | April 16, 2018 | 266 | 7
Z-0675-2018 | 79046 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas 8000 Modular Series system; Software Version 05-02 and 06-03; UDI: 05641446001 |  |  | July 28, 2017 | 2017 | Roche Diagnostics Corporation | A software failure may incorrectly set the system settings to "default" settings, creating a risk of incorrect results. | Device Design | Consignees were notified by letter (receipt required) and requested to follow instructions to determine if the software issue has occurred. If the software issue is detected the consignee contacts customer support to have system reconfigured. |  | 399 | 399 | Distributed in 36 states: AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA, WI and Puerto Rico. | Open, Classified |  |  | 399
Z-0669-2018 | 79021 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LCP Drill Sleeve 1.5 Drill Bit 1.1mm Part Number 03.114.001 |  |  | August 07, 2017 | 2017 | Synthes, Inc. | May have difficulty attaching the drill sleeve to the plate. If the drill sleeve will not connect to the plate intraoperatively, surgical delay could occur. If the pilot hole is drilled without the drill sleeve, the pilot hole could be off-center which could result in screw and/or plate loosening. | Process design | In August of 2017 customers were mailed an URGENT FIELD SAFETY NOTICE titled "884115 - Locking Compression Plate (LCP) Drill Sleeve" requesting that customers who received product subject to recall to please contact: RA-DPYCH-flactionint@ITS.JNJ.com or call 610-719-5450. In the event that you have product requiring use:Pre-assemble the drill sleeve with the plate preoperatively. If you experience difficulty connecting the drill sleeve to the plate, use the additional 1.1 mm locking drill sleeve provided in the graphic case, or:o Align the 1.1 mm drill bit to the plate as perpendicular as possible and then drill freehand and the locking screw can be engaged in this manner.o If the drill hole must be used and no option to lock remains, a non-locking screw can be used. |  | 567 | 567 units | Distribution to US, Canada, and Germany. | Terminated | July 25, 2018 | 352 | 567
Z-0685-2018 | 78587 | Yes | NIP |  |  | Cardiovascular | PMA | Innova Self-Expanding Stent System, 200 mm. Labeled with Material Number and description as follows: |  |  | November 06, 2017 | 2017 | Boston Scientific Corporation | Stent possibly unable to be fully released form the delivery system. | Nonconforming Material/Component | Boston Scientific sent an Urgent Medical Device letter dated November 6, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers\ were instructed to immediately discontinue use of the product and remove all affected products from inventory, complete and return the Verification Form, return affected products and provide the notification to any healthcare professionals affected. |  | 22916 | 22,916 units total | Worldwide Distribution - US Nationwide Distributed domestically to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV. Distributed internationally to Canada | Open, Classified |  |  | 22916
Z-0689-2018 | 79158 | No | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Philips Avent Smart Ear Thermometer (SCH740), Philips Ear Thermometer (DL8740) |  |  | February 02, 2018 | 2018 | Philips Consumer Lifestyle | Device may provide inaccurate measurements. | Use error | " Philips advises users of the Device to only use the Device as directed in the instructions for use including the additional advice as per this Field Safety Notice" Philips advises users against relying on the Device as the sole basis for a decision seeking treatment and/or professional help" Users who wish to return the Devices based on the information provided in this Field Safety Notice can contact their local Philips product support representative |  | 45141 | 45141 | Domestic (2463 units) and Foreign (42678 units) | Open, Classified |  |  | 45141
Z-0690-2018 | 79200 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROWÂż PICC |  |  | October 17, 2017 | 2017 | Arrow International Inc | The kit may contain an incorrect catheter component. Kits that shouldcontain a 40cm catheter may contain a 55cm catheter. | Mixed-up of materials/components | On October 17, 2017 an URGENT Medical Device Recall letter titled "ARROW PICC powered by ARROW VPS Stylet" was mailed to customers instructing them to review inventory and quarantine all affected product which will be returned to the firm. Questions or concerns can be directed to Customer Service at 18662466990. |  | 601 | 601 eaches | AZ, CA, GA, MA, MO, NY, VA | Open, Classified |  |  | 601
Z-0687-2018 | 78704 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | Nitinol TC Reusable Electrodes (TCN), Model Numbers: 1)TCN-10, 2) TCNK-10-R*, 3)TCN-10-3M, 4) TCNK-15, 5) TCN-15, 6) TCNK-15-C, 7) TCN-15-3M, 8) TCNK-15-R*, 9) TCN-20, 10) TCNK-20, 11) TCN-20-3M*, 12) TCNK-20-C*, 13) TCN-5, 14) TCNK-20-R*, 15) TCN-5-3M, 16) TCNK-5, 17) TCNK-10, 18) TCNK-5-C, 19) TCNK-10-C &amp; 20) TCNK-5-R*. |  |  | July 05, 2017 | 2017 | Cosman Medical, LLC | After multiple reprocessing cycles, the epoxy resin which holds the TCN Electrode in the hub can exhibit signs of damage. In some cases, this damage may result in the inability to fully remove blood and/or tissue residuals prior to cleaning and re-sterilizing the device. | Device Design | The firm, Cosman Medical, sent an "Urgent Medical Device Correction" notice and Field Correction Reply Verification Tracking Form dated July 5, 2017 to their customers. The notice described the product, problem and actions to be taken. The Customers were instructed to do the following:1. Immediately discontinue use of devices that have already been used.Used TCN Electrodes should be disposed of in accordance with your institutions infectious material/biohazard waste control procedures.2. Unused/un-opened TCN Electrodes should be identified as single use only (for one patient). These devices are supplied as non-sterile which require cleaning and sterilization prior to use. These TCN Electrodes may be used once the cleaning and sterilization steps are completed. After single use, the used TCN Electrode should be disposed of in accordance with your institutions infectious material/biohazard waste control procedures.3. Complete and return the Field Correction Reply Verification Tracking Form.via Email: Cosmanorders@bsci.com or via Fax to: 781-272-6563 - Identify your email with TCN Electrodes Action in the subject line.Customers with questions may call 1-888-8COSMAN or email Cosmanorders@bsci.com. |  | 11250 | 11250 units (10,460 USA; 790 OUS) | Worldwide Distribution: US (nationwide) and country of Canada. | Terminated | September 14, 2018 | 436 | 11250
Z-0688-2018 | 78961 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | PreludeSNAP Splittable Hydrophilic and |  |  | December 14, 2017 | 2017 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is voluntarily conducting a recall of specific lots of PreludeSNAP Splittable Sheath Introducers due to an intermittent failure of the splittable hub. | Under Investigation by firm | Communication to consignees began on December 14, 2017. Merit requested that consignees stop using the PreludeSNAP devices for human use, discontinue distribution, and return remaining product to Merit Medical. |  | 60945 | 60,945 devices | US, Australia, EU, Serbia, Israel, Malaysia, South Africa, Bahrain | Open, Classified |  |  | 60945
Z-0683-2018 | 79016 | No | LIX | Cardiopulmonary resuscitation (CPR) aid. | Cardiovascular | Cardiovascular | 510(k) | Philips QCPR Meter - used with FR3 |  |  | January 29, 2018 | 2018 | Philips Electronics North America Corporation | Inform users of correct placement of the disposable adhesive pad on the QCPR meter and inform users of injuries associated with proper performance of CPR. As required by Consent Decree of Permanent Injunction No. 1:17-cv-11955-DJC, entered October 31, 2017 between United States of America and Philips North America | Device Design | Philips sent an Medical Device Recall letter dated January 19, 2018 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately return the attached reply card. For question contact your local Philips representative. |  | 28094 | 28094 in total | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 28094
Z-0697-2018 | 78573 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | MOOG Curlin Infusion Administration Set, Ref 340-4114, Non-DEHP Microbore Tubing with Non-Vented Bag Spike, packaged 20/box, for use with the MOOG Curlin Ambulatory Volumetric Infusion Pump, RX, Sterile. |  |  | November 10, 2017 | 2017 | Zevex Incorporated (dba Moog Medical Medical Devices Group) | Particulate found on the Curlin spike. | Nonconforming Material/Component | The recalling firm began issuing letters via USPS certified mail with return receipt on 11/10/2017 and concluded on 11/13/2017. Updated letters dated 11/20/2017 were issued via certified mail to distributors and to medical facilities. The distributor letter emphasized their instructions to notify downstream accounts. The medical facility letter updated lot numbers.The letter described the issue, risk to health, and actions the customer was to take. A response form was enclosed to report whether the customer has inventory they are not returning, that they no longer have inventory of the product, or the amount of product being returned for replacement. The form was to be returned via FAX or email. |  | 1124880 | 56,244/20 administration set cases | Distribution was nationwide. There was government and military distribution. Foreign distribution was made to Canada, Saudi Arabia, Turkey, Croatia, United Arab Emirates, and Australia. | Terminated |  |  | 1139314
Z-0704-2018 | 79139 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | digital x-ray detector |  |  | December 11, 2017 | 2017 | Philips Electronics North America Corporation | It the WiFi connection between the SkyPlate detector and HP transfer point is weak, an image may fail to transfer from the SkyPlate detector to the system. The image remains in the memory of the detector, but cannot be transferred wirelessly or by use of the backup cable. To continue, the operator can reset the SkyPlate detector by removing its batteries, but the acquired image is lost and are-take is necessary. | Under Investigation by firm | Philips sent an Urgent- Field Safety Notice date December 11, 2017.Philips plans to release a new HP firmware to the installed base. In addition the existing Access Point Hardware in the field needs to be patched. A Philips Service Engineer will contact the affected Customers when the Field Action Kit is available to be implemented. For further questions, please call (978) 659-7275. |  | 157 | 157 | Worldwide Distribution - USA (nationwide) Distribution, Internationally to Canada. | Open, Classified |  |  | 157
Z-0695-2018 | 79117 | No | JCY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 1113 Gravity Tube Set Pk/24 -SH |  |  | August 31, 2017 | 2017 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Leakage of FMS gravity tub sets (281113) at the bulb and tube connection | Nonconforming Material/Component | On September 6, 2017 an URGENT MEDICAL DEVICE RECALL letter was issued to customers titled "FMS Gravity Tubing" requesting that customers cease further distribution or use, quarantine all remaining products and contact sales consultant for the coordination of return and replacement. Questions or concerns can be directed to the recall coordinator at 508-828-6609 |  | 15678 | 15,678 | US Nationwide Distribution | Open, Classified |  |  | 15678
Z-0692-2018 | 79116 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Berchtold Chromophare Dual Monitor Carrier Arm or Dual Flat Panel, a component of Berchtold Chromophare Ceiling Mounted Surgical Light System. |  |  | January 26, 2018 | 2018 | Stryker Communications | The mounts on the monitor may have an insufficient (missing or incomplete) weld. | Process control | The recalling firm issued letters dated 1/26/2018 via FedEx on 1/26/2018 explaining the issue and the potential adverse consequences. |  | 850 | 850 | Distribution was made to medical facilities nationwide, including government and military distribution. Foreign distribution was made to Canada, Belgium, China, United Kingdom, UAE, France, Germany, Spain, Japan, Ecuador, and Russia. | Open, Classified |  |  | 850
Z-0711-2018 | 79214 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis model Interventional fluoroscopic x-ray foot switch accessory; Model Numbers: 4787797, 4787805, and 4787813. |  |  | September 28, 2017 | 2017 | Siemens Medical Solutions USA, Inc | The possibility exists for the wireless foot switch to fail due to impacts by external factors, such as electrostatic discharge that exceed a certain intensity. If the foot switch fails, it will no longer be able to be used to release radiation. The hand switch provided can be used to release exposures; however, fluoroscopic examinations will not be possible. | Device Design | Letters were sent to all consignees with instructions to contact their service provider to arrange for an appointment to update the device hardware. |  | 142 | 142 | Distributed in XX states: Ak, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, the District of Columbia, Guam, and Puerto Rico. | Open, Classified |  |  | 2209
Z-0696-2018 | 79017 | No | GCP | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | A20975A "Working insert, with ramp, one way" (Lots 13ZW to 172W; A20976A "Bridge, one way" (Lots 146W to 172W) and A20977A "Bridge, two way" (Lots 146W to 174W) |  |  | August 18, 2017 | 2017 | Olympus Corporation of the Americas | Olympus has received complaints about fragments of adhesive which detached from inside the working channel of the referenced products. Cracking, chipping, missing pieces, and delamination of the adhesive have also been reported. In one case, a fragment of the adhesive was detected inside a patient's urethra during a procedure and was retrieved without injury to the patient.The manufacturer, Olympus Winter &amp; lbe GmbH ("OWi") has determined that adhesive can be dislodged during the intended use of the cystoscopy bridge or working insert, e.g., when inserting an instrument through the working channel. As a result, a fragment of the adhesive may fall inside the patient's bladder, urethra or uterus and would need to be retrieved, extending the length of the procedure or requiring additional surgical treatment.The adhesive is used to seal gaps inside the working channel of the cystoscopy bridges and working inserts. In 2013, the adhesive used until then was discontinued and had to be replaced. Current investigations revealed that the durability of the new adhesive is insufficient | Component change control | On August 18, 2017 an URGENT Medical Device Removal and Corrective Action letter was sent to customers regarding the "Cystoscopy Bridges and Working Inserts" instructing customers to review their inventory, quarantine affected product and call customer service at 800-848-9024 to arrange return and replacement. |  | 8741 | 8741 | US Nationwide Distribution | Open, Classified |  |  | 8741
Z-0691-2018 | 79073 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | Biofinity XR Toric contact lenses 8.7/14.5 - 6PK |  |  | September 01, 2017 | 2017 | CooperVision Inc. | The firm identified a labeling error affecting Made-To-Order(MTO) Biofinity XR Toric blister labels. The primary blister reads Biofinity Multifocal Toric in error, and should read Biofinity XR Toric. The outer carton correctly reads Biofinity XR Toric. | Employee error | Notifications were sent with instructions to examine inventory, discard any affected inventory, and contact the firm for credit or replacement. For further questions, please call (585) 756-9273. |  | 120 | 120 | USA Distribution | Terminated | July 30, 2018 | 332 | 120
Z-0693-2018 | 79086 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coated VICRYL RAPIDETM (Polyglactin 910) Suture |  |  | August 17, 2017 | 2017 | Ethicon, Inc. | On August 17, 2017, Ethicon had decided to initiate a medical device recall(removal) of the twenty-five (25) lots of VICRYL RAPIDETM (Polyglactin 910) Suturethat were distributed to the field. During finished goods testing, a suture damage wasidentified on a lot of VICRYL RAPIDETM Suture. During the course of the subsequentinvestigation, the suture also failed knot tensile strength testing. | Equipment maintenance | On August 31, 2017 a Medical Device recall letter (Reference 2210968-08.31.17-003-R) was issued requesting customers to do the following:1. Examine your inventory immediately to determine if you have product subject to this recall (removal) on hand and quarantine such product(s).2. Remove the product subject to this recall (removal) and communicate the issue to relevant operating room or materials management personnel, or anyone else in your facility who needs to be informed.3. If any product subject to this recall (removal) has been forwarded to another facility, contact that facility to arrange return.If you have additional questions regarding this recall (removal) or to report any customer complaints, please contact Ethicon Customer Support Center at 1-877-ETHICON (1-877-384-4266).As with any medical device, adverse reactions or quality problems experienced with the use of this product should be reported to the FDA's MedWatch Adverse Event Reporting program online, by regular mail or by fax:" Online: www.fda.gov/medwatch/report.htm" Regular Mail:Use postage-paid FDA form 3500 available at: www.fda.gov/MedWatch/getforms.htmMail to: MedWatch 5600 Fishers Lane, Rockville, MD 20852-9787" Fax: 1-800-FDA-0178 |  | 129372 | 129,372 Eaches (Internationally) | Nationally | Open, Classified |  |  | 129372
Z-0707-2018 | 79097 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis Zee and Zeego, Model Numbers: 10094135, 10094137, 10094141, 10094142, 10094143, 10280959, 10848283 &amp; 10848354 |  |  | June 06, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Potential for liquid entry | Device Design | On June 6, 2017, Siemens Healthineers distributed URGENT MEDICAL DEVICE CORRECTION notices to their customers via USPS Certified Mail.Siemens conducted a corrective field action in June 2015 to protect the Artis table against liquid entry at critical locations. However, even with this corrective action, if there are larger amounts of liquid on the table surface, they can still enter the table. Siemens Medical Solutions USA, Inc. will contact customers to arrange a date to perform the corrective action. Please feel free to contact the service organization for an earlier appointment by calling 610-448-6461. |  |  |  | U.S. Nationwide | Open, Classified
Z-0712-2018 | 79293 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ProLite Mesh - mesh, surgical, polymeric |  |  | November 22, 2017 | 2017 | Atrium Medical Corporation | An incorrect ProLite Mesh configuration was packaged as product code 1010306-06, with lot number 412298. ProLite Mesh manufactured as product code 1010306-04 was inadvertently placed inside the pouch of packages labeled with product code 1010306-06 and lot number 412298. | Under Investigation by firm | Please examine inventory to determine if you have any of the affected products listed on page 1. If so, please remove affected products, quarantine them, and place in a secure location. If you have affected product, please contact Getinge Customer Service at internationalcs@atriummed.com for a Return Authorization (RA) and shipping instructions to return any affected product. Pack the product to be returned with the appropriate return documents, and using the shipping instructions provided, arrange for pick up with the designated delivery service provider. Please complete and sign the enclosed Urgent Medical Device Field Safety Notice to acknowledge receipt. Return the completed form to the following e-mail address: MESH2017@getinge.comFor further questions, please call (973) 709-7654. |  | 60 | 60 units | Internationally to Spain only. | Terminated | July 30, 2018 | 250 | 60
Z-0694-2018 | 79171 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Footswitch accessory to the CUSA Clarity Ultrasonic Tissue Ablation System; Catalogue No: C7000 |  |  | September 26, 2017 | 2017 | Integra LifeSciences Corp. | A faulty footswitch may result in unintended ultrasonic fragmentation activation. There are two scenarios when the failure mode may occur: 1) Ultrasonic fragmentation may be inadvertently activated at start-up or at the initial use once the footswitch is connected to the console and after the user primes the system, performs a handpiece test, and enters the Main Screen. 2) Ultrasonic fragmentation may remain activated during surgical use once the foot pedal is disengaged slowly. | Process control | Integra I sent an Urgent Voluntary Medical Device Removal Field Safety Notice letter dated September 26, 2017 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check the footswitch you received with your purchase of the CUSA Clarity Console, Catalog # C7000for the lot number on the back of the footswitch. Reference pictures are provided on the next page. Compare it to the affected footswitch lot numbers on the next page. If you do have an affected lot number, replace this footswitch with the one that has been provided to you by your sales representative. If you have not yet been provided with the replacement footswitch, please contact your sales representative as soon as possible for a replacement footswitch.2.Complete the information on the form and either provide it to your sales representative or email or fax as indicated on the form. Keep a copy of the form for your records.For questions regarding the instructions contact SSS E&L; at sssel@integralife.com or 888-419-0625 option 2. |  | 155 | 155 | Worldwide Distribution - US Nationwide in 33 states: AL, AR, CA, CO, CT, FL, GA, IA, ID, IL, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NM, NY, OH, OK, OR, PA, TN, TX, VA, WA, WI in the countries of Greece, South Africa, France, Germany, Italy, United Kingdom, Sweden, Netherlands, and Dominican Republic. | Terminated | August 14, 2018 | 322 | 155
Z-0703-2018 | 79107 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Foot Controls impacted for the following devices: |  |  | July 31, 2017 | 2017 | Oakworks Inc | The foot control used on the OAKWORKSÂż Inc., CFPM300, CFPM301, CFPM400, CFPM401, CFPMB301, CFUR301, CFUR401 and CFLU401 Imaging Tables can become lodged under the base or under the column of the table. | Device Design | OAKWORKS, Inc. issued a MEDICAL DEVICE ADVISORY NOTICE dated 8/18/2017 to customers providing more information in regards to the Medical Device Advisory Notice sent on May 12, 2017. The advisory notice was initiated in response to reported issues with the foot control. OakWorks, Inc. customer service will place the order for the appropriate Foot Control Kit listed below based on what table is impacted and the date of manufacture. PN2015-204 - Foot Control Kit - Standard Riser2015-209 - Foot Control Kit - Thinner Riser2015-210 - Foot Control Kit - Holder Only2015-211 - Foot Control Kit - Riser RemovalCustomers with questions may contact 1-800-558-8850 or 1-717-235-6807. |  | 1104 | 1,104 | USA (nationwide) Distribution | Open, Classified |  |  | 1104
Z-0705-2018 | 78929 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Rusch(R) Greenlite MAC 1, Catalog Number 004551001 |  |  | January 10, 2017 | 2017 | Teleflex Medical | The pouch may contain the incorrect RuschGreen Lite product size. Pouches that are labeled as containing RUSCH GREEN LITE MAC 1 mayinstead contain RUSCH GREEN LITE MAC 2. | Package design/selection | The firm initiated their recall by letter on 01/10/2018 via FedEx 2-day. Affected consignees were instructed to take the following actions:"1. If you have affected stock, immediately discontinue use and quarantine any products with the product code and lot number listed above, so that the affected products can be returned to Teleflex.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document your receipt of this letter." |  | 2140 | 2140 units | US only | Open, Classified |  |  | 2140
Z-0717-2018 | 79091 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD and and Xper FD ORT - Philips X-ray Systems |  |  | December 20, 2017 | 2017 | Philips Electronics North America Corporation | Extra monitors, so called 7th or 8th monitor, mounted on the Monitor Ceiling Suspension (MCS), might disengage from the MCS and fall down. The two small set screws securing the two gliding rods into an aluminum block ofthe support arm can come loose because of vibrations and forces exerted on them. If the set screws have come loose gravity can keep the rods in position. However, if a force is exerted on the monitor pushing it forward, the monitor cantopple forward. Once the monitor has toppled over it is only attached with its wires. If the monitor breaks loose of its wires it will drop on the floor. | Device Design | Philips sent an Urgent Field Safety Notice letter December 20, 2017 dated to all affected customers. The letter identified the affected product, problem and the actions to be taken. For question contact your local Philips representative. |  | 1486 | 1486 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 1486
Z-0716-2018 | 79099 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper systems with Certeray X-ray Generator located in regions with high mains (480V) voltage. |  |  | January 04, 2018 | 2018 | Philips Electronics North America Corporation | There is an increase in the failure rate of certain Anode Drive Units (ADU5) used in these products. Early life failure of the ADUs only occurs when the hospital mains provide 480V to the system. This current voltage may lead to saturation and overheating of the coils that protect the IGBTs (insulated-gate bipolar transistor) of the ADU. The overheating may also generate a peculiar burning smell that may be noticed. In some instances this burning smell may be noticed before the ADU fails. When the ADU fails the X-ray performance of the system is reduced toEmergency Fluoroscopy. Exposure will not be possible and image quality is reduced. | Device Design | Customer Notification letter dated August 2017 were sent to affected customers. The letter identified affected product, problem and actions to be taken. The customer is requested to place this Field Safety Notice with the documentation of the system until this correction has been implemented. For questions contact your local Philips representative. |  | 705 | 705 | Worldwide Distribution - US Nationwide; foreign distribution to Brazil, Germany, Iceland, Japan, Mexico, and (updated 6/25/18) Saudi Arabia. | Open, Classified |  |  | 705
Z-0715-2018 | 79159 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | AIRO Mobile CT System, Model Number: MobiCT-32 |  |  | September 26, 2017 | 2017 | Mobius Imaging, LLC | If the tilt-drive motor has an intermittent connection or broken encoder or sensor wire there is a risk of unintended or unpredictable device motion while the CT system is homing the axis at startup or preparing for a scan. | Under Investigation by firm | Mobius sent an Urgent Field Safety Notice letter dated September 26, 2017 to their customers. The letter identified the affected product, problem and actions to be taken. Mobius Imaging is advising all customers to discontinue using the AIRO until update is completed on their system. Please provide a copy of this Field Safety Notification to all AIRO operators that use the AIRO system, and quarantine the system per hospital policies and procedures. |  | 105 | 105 | Worldwide Distribution - US Nationwide in the states of: AZ, CA, CO, FL, GA, IL, KY, LA, MA, MD, MI, MN, MO, MT, NC, NY, OH, OR, PA, RI, TN, TX, UT, VA, WA, WV and the countries of Argentina, Australia, Austria, Belgium, Canada, Egypt, France, Germany, India, Italy, Japan, Kazakhstan, Lithuania, Netherlands, Pakistan, Russia, Saudi Arabia, Spain, Switzerland, UAE, Uzebekistan | Terminated | August 16, 2018 | 324 | 105
Z-0714-2018 | 79207 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Go.Now computed tomography (CT) x-ray system table, Model Number: 11061618 |  |  | October 02, 2017 | 2017 | Siemens Medical Solutions USA, Inc | The table may become unstable if the nuts on the bolts were not adequately tightened and thus have the potential to loosen and several bolts may gradually loosen at the same time. | Unknown/Undetermined by firm | Siemens sent a Customer Safety Advisory Notice letter date October 2, 2017 to their customers. Customers were instructed not to use the table if there is any unusual noise or vibration and to contact Siemens service organization immediately at 1-800-888-7436. Customers were informed that a service technician will schedule an on site repair. |  | 1 | 1 | US Nationwide Distribution in the states of FL, MO, NY, TN | Terminated | September 27, 2018 | 360 | 7
Z-0718-2018 | 79066 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Max Implant System - Bone Tap for use with 7.0mm diameter Max Implant, Catalog Number D-TAP-MAX7-7 |  |  | June 16, 2017 | 2017 | Keystone Dental Inc | An 8mm diameter bone tap was laser marked as 7mm diameter and was packed with a 7mm label. | Unknown/Undetermined by firm | Consignees were initially called between June 13, 2017 and June 20, 2017. Customer Relations Fax Back forms were completed to document the phone calls. A notification letter was also sent via FedEx or hand-delivered on 6/19/17. All returned product will be scrapped. For further questions, please call (781) 328-3313. |  | 72 | 72 | USA (nationwide) Distribution to the states of : CA, CO, FL, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, NH, NY, OH, PA, TN, TX, UT, VA, VT, WA, WI and WV; and Internationally to Canada | Open, Classified |  |  | 72
Z-0724-2018 | 78663 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integrity R1.2 |  |  | November 20, 2017 | 2017 | Elekta Limited | Potential for positional errors following automatic table movement (ATM). | Software Design Change | Customers were notified via letter on about [DATE]. Instructions include to post the notice in a place accessible to all users until the action is closed and to advise appropriate personnel working with the affected product of the recall action. Additionally Elekta plans to release Integrity R4.0.0, which should identify positional errors over 5 mm, in the first half of 2018. |  | 539 | 539 units total | Distributed domestically to AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. Distributed internationally to ALBANIA, ALGERIA, ANTIGUA AND BARBUDA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOSNIA AND HERZEGOVINIA, BOTSWANA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, LAOS, LITHUANIA, MALAYSIA, MALTA, MEXICO, MOROCCO, MYANMAR, NAMIBIA, NEPAL, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKASUDAN, SURINAME, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UK, UKRAINE, VENEZUELA, VIETNAM, ZIMBABWE | Open, Classified |  |  | 539
Z-0720-2018 | 79170 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA CentaurÂż T3 |  |  | July 28, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc | Siemens has identified an issue when scanning the 20 barcode for the Master Curve Card(MCC) with ADVIA Centaur systems T3 kit lots ending in 198 on the ADVIA Centaur XPTsystem. | Under Investigation by firm | Siemens sent an Urgent Medical Device Correction letter dated July 28, 2017, that was distributed to all customers in the United States and an Urgent Field Safety Notice was distributed to all customers outside the United States. On August 15, 2017 Follow-up communication was sent to all customers.For further questions, please call (508) 668-5000. |  | 33298 | 33,298 | Worldwide Distribution - US Distribution. | Open, Classified |  |  | 33298
Z-0721-2018 | 78762 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Segmental System, ZSS Distal Femur. Sizes as follows: |  |  | August 01, 2017 | 2017 | Zimmer Biomet, Inc. | The products being recalled exceeded the weight previously tested leading to a possibility of damage to the package and/or a compromised sterile barrier. | Package design/selection | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL REMOVAL"-Lot Specific letter dated 8/1/2017 to its Customers via email and/or FedEx. The letter describes the product, problem and actions to be taken. The Instructions to Distributors, Sales Representatives, and Distributor Operation Managers included to ensure affected team member are aware of the contents of the notification letter, immediately locate and quarantine affected product in inventory, complete and return the Certification of Acknowledgement form by email to: corporatequality.postmarket@zimmerbiomet.com., return all affected product, and note if there are any hospitals and/or surgeons for which need to be notified. Instructions for Risk Managers and surgeons are to ensure affected personnel are aware of the contents of the notification letter, quarantine any affected product on hand, complete and return the Certificate of Acknowledgement form by email to: corporatequality.postmarket@zimmerbiomet.com., and return any affected product on hand. For questions or concerns, please call the 411 call center at 1-574-371-3071 between 8:00 am and 5:00 pm EST, Monday through Friday or email corporatequality.postmarket@zimmerbiomet.com |  | 3350 | 3,350 units total | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and Puerto Rico; and to countries of: Austria, Australia, Bahamas, Brazil, Canada, China, Dominican Republic, England, France, Germany, Israel, Japan, South Korea, Lebanon, Malaysia, Mexico, Netherlands, Panama, Singapore, Spain, Switzerland, and Thailand. | Open, Classified |  |  | 3350
Z-0719-2018 | 78973 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | MEDISTIM VERIQ SYSTEM, MODEL VQ4122 |  |  | April 25, 2017 | 2017 | Medistim Asa | A system component overheated causing melting of the connector and emission of smoke through the systems loudspeaker opening. | Nonconforming Material/Component | A letter was sent describing the device, reason for the recall, and risk involved. The firm stated that it would ship a VeriQ replacement system without the affected connector being used. |  | 2 | 2 | Distributed to New York. | Open, Classified |  |  | 2
Z-0725-2018 | 79167 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Destination Guiding Sheath catheter introducer, Catalog Number: RSR06 |  |  | September 12, 2017 | 2017 | Terumo Medical Corporation | The tucking depth feature of this catheter introducer is marginally outside the upper specification. | Process control | Terumo sent an Urgent Medical Device Recall dated September 12, 2017 to affected customers. The letter identified the affected product problem and actions to be taken. The informed customers that a territory managers will visit each institution to examine the inventory, identify the affected product and lot, and prepare the product for return to Terumo Medical Corporation. There is an inventory confirmation form that is returned by the sales representative to the recalling site. |  | 33 | 33 | To be added. | Terminated | August 31, 2018 | 353 | 33
Z-0743-2018 | 79100 | No | FRG | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Vis-U-All High Temp 7.5"x13" Self Seal Pouch 200 pouches per box; 4 boxes per case |  |  | June 15, 2017 | 2017 | Steris Corporation | On June 15, 2017, STERIS identified that the purple ink located on the outside label of Vis-U-All Pouches may migrate following steam processing causing cosmetic defects on the exterior of the pouch and in some instances, transference of the ink onto the device(s) contained within the pouch. | Component change control | Steris sent an Urgent Voluntary Recall Notice letter dated July 2017 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately inspect on-hand inventory for recalled product. If you have product in your possession, please quarantine it and contact firm to coordinate shipment of replacement product. If you have questions regarding this matter, please contact STERIS Customer Service at 1-800-548-4873. |  | 127 | 127 | US Nationwide Distribution | Terminated | July 17, 2018 | 397 | 1870
Z-0741-2018 | 78479 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Cosmedent's Etching Gel, 3g, 37% Phosphoric Acid, Rx only. Packaged in 3g Luer Lock syringes. |  |  | September 28, 2017 | 2017 | Cosmedent, Inc. | Product packaged in a syringe that is not a lure lock syringe and the tip cold come off when expressing the material. | Under Investigation by firm | Customers were sent recall notification letters on approximately 09/28/2017. The letter identified the affected device and discussed the reason for the recall. Instructions included discontinuing use of the product, and returning it to Cosmedent. The firm will issue a credit or refund for the recalled device. |  | 84 | 84 units | Worldwide Distribution -- USA, to the states of IL, LA, MI, PA, and WI; and, to the countries of Australia, Hungary, and Spain. | Open, Classified |  |  | 84
Z-1012-2018 | 79257 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The GLOBAL UNITE Platform Shoulder System is intended for cemented or uncemented total shoulder or hemi-shoulder replacement procedures, Product labeling: LEÂżO-1100-30-000 REV D.pdf.pdf, LEÂżO-1100-40-000 REV D.pdf.pdf, LEÂżO-110030100_REV D.pdf.pdf, LEÂżO-110030110_REV D.pdf.pdf, LEÂżO-110030120_REV D.pdf.pdf |  |  | November 28, 2017 | 2017 | DePuy Orthopaedics, Inc. | The screw in specific lots of the GLOBAL UNITE Anatomic Body and GLOBAL UNITE Fracture Body was inverted during assembly to the body, which will cause the humeral stem to sit proud and may cause surgical delays. | Employee error | DePuy Orthopaedics, Inc. is voluntarily recalling 11 lots of the GLOBALÂż UNITEÂż Platform Shoulder System. The device is being recalled because the screw in specific lots of the GLOBAL UNITE Anatomic Body and GLOBAL UNITE Fracture Body was inverted during assembly to the body, which will cause the humeral stem to sit proud and may cause surgical delays. Once the screw is properly assembled, there is no additional risk to the patient. Further distribution or use of the affected lots is to cease immediately.DePuy Orthopaedics, Inc. is not recommending prophylactic revision. The company recommends that surgeons discuss potential clinical implications and risks with symptomatic patients that received the affected lots and may have experienced a surgical delay. Please take the following actions:Please cease using the affected components immediately.Return affected devices: Medical facilities are to determine if any of the recalled devices are still on hand, and return affected devices immediately to their DePuy Synthes Sales Consultant or return them to DePuy Synthes Joint Reconstruction for credit following normal procedures. Reconciliation forms should be faxed to 574-371-4939. Note: The affected lots may be on consignment at your facility.Reconciliation Form: Complete the Reconciliation Form and return to your DePuy Synthes Sales Consultant, email to DPYUS-JointReconFieldActions@its.jnj.com, or fax to 574-371-4939 within five (5) days of this notice. Records: Retain a copy of the completed Reconciliation Form in your files along with this notice. Additional Notifications:Notify surgeons at your facility by providing them with a copy of this notice to ensure surgeons are aware of this recall. Forward this notice to others in your facility that need to be informed.If any affected product has been forwarded to another facility, contact that facility immediately to communicate this field action with the facility/facilities.Maintain a copy of this notice wi |  | 85 | 85 | Distribution Level: The device removal is extended to the companies US Distributors/Sales Offices and Medical Facilities that may have used, purchased, or received the affected devices. Approximately 53 units distributed to firm's international affiliates. | Open, Classified |  |  | 85
Z-0729-2018 | 78942 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | OptiView DAB IHC Detection Kit, Ventana Part Number 760-700, Roche GMMI 06396500001 in vitro diagnostic |  |  | December 19, 2017 | 2017 | Ventana Medical Systems Inc | Increased reports of leaking and sticking reagent dispensers for horseradish peroxidase (HRP) dispensers (part of the iView, ultraView and OptiView detection kits, as well as CINtec PLUS Cytology Kit) and with Hematoxylin II. | Process control | On Dec 19th all Roche affiliates (the direct consignees) worldwide were notified via Safety Board Notification (SBN) and subsequent Field Safety Notification (EX-US) and UMDC TD00090 (US) SBN and FSN revised Dec 20th due to 3 affected part numbers missed on original letter.The instructions are as follows:"Actions taken by Roche Diagnostics All affected product has been placed on hold. Ventana has reworked all product in its inventory and is in the process of manufacturing new lots for distribution and replacement of customer affected kits. Customers will be notified when corrected product is available. Actions to be taken by the customer/userAffected kits may continue to be used by customers until corrected product is available, however Ventana is mandating that the affected IHC detection kits (iView, ultraView, OptiView) must only be used in conjunction with same-slide controls. These controls must be appropriate for each assay and capable of detecting false negative results due to a complete or partial reagent dispense failure. CINtec PLUS Cytology does not have the capacity for same slide controls, so system-level controls should be maintained. ultraView SISH Detection is used for HER2 analysis, and employs internal positive controls; no external control is required. For assays that directly relate to clinical therapy decision making (e.g. ER/PR, HER2, ALK, etc.), it is additionally important to select a same slide positive control tissue with sufficient sensitivity to detect small decreases in intensity that may cause borderline positive cases to appear as negative (e.g. HER2 2+ vs. 1+). Although the use of same slide controls is considered optimal laboratory practice and strongly recommended by Ventana, customers may revert to standard run controls once non-impacted product is received.In order to reduce the risk of this issue impacting patient care, customers not using same slide controls as a standard practice should follow their |  | 12028 | 12028 units | Worldwide distribution, including US nationwide, Argentina, Australia, Hong Kong, Indonesia, Philippines, Brazil, Chile, Thailand, Korea, Malaysia, New Zealand, Ecuador, Mexico, Pakistan, Vietnam, Germany, Japan, India, Colombia, Singapore, Uruguay, Peru, Canada, Taiwan, China, | Open, Classified |  |  | 336416
Z-0728-2018 | 79015 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | JANUS &amp; JANUS G3 Automated Workstation is an automated programmable liquid handling instrument intended for use in clinical settings to facilitate the pipetting and dilution of solutions in sample preparation work flows. |  |  | August 02, 2017 | 2017 | Perkinelmer Life Sciences, Inc. | We have become aware through internal testing that JANUS and Janus G3 Automated Workstations were failing production final acceptance testing due to a Horizontal Motion Error message displayed on the system. Further investigation concluded that two printed circuit boards (PCB) used in MDT models of Janus and Janus G3 were manufactured with the incorrect resistor values. This anomaly affects component part number 7101640 (PCB-Quad Motor Controller MDT) and part number 7101647 (PCB-Quad Motor Controller Gripper). This anomaly could result in the Horizontal Motion Error being displayed and result in instrument down time.Until PerkinElmer s corrective measures are completed, we request you;Âż Be aware that there is potential for the Horizontal Motion Error message to display which could result in instrument down time.Âż Follow guidance provided in the Janus Automated Workstation User Manual when the Horizontal Motion Error message displays.Âż If you have additional questions, contact PerkinElmer Technical Support for further assistance (800) 762-4000. | Nonconforming Material/Component | On August 4 2017 an URGENT Medical Device Correction letter titled "Horizontal Motion Error" was issued to customers informing them of the error notice and to follow guidance provided in the Janus Automated Workstation User Manual when the Horizontal Motion Error message displays. Question or concerns can be directed to Technical Support at 800-762-4000 |  | 7 | 7 | Nationwide | Terminated | July 31, 2018 | 363 | 7
Z-0727-2018 | 79110 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Stage-1Âż RDS Indexed Abutment for RDS Stage-1Âż Single Stage |  |  | August 25, 2017 | 2017 | Keystone Dental Inc | The recalled units are mislabeled with an incorrect expiration date. The expiration date printed onthe label is 2202-06. The correct expiration date should be 2022-06. | Error in labeling | On August 18, 2017 a letter titled "Product Action  S2400K, Stage-1 RDS Indexed Abutment, Lot 34276" was issued to customers requesting persons to review their inventory and respond if product has been used or not. A response form is requested from these customers with instruction to return any unused product. Questions or concerns can be directed to:Customer Relations:Toll Free 866-902-9272 |  | 20 | 20 | Nationally. | Open, Classified |  |  | 20
Z-0775-2018 | 79146 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 979195 Loaner Kit for femoral intramedullary nail. Orthopedic instrument. |  |  | September 13, 2017 | 2017 | Zimmer Biomet, Inc. | Fracture of the bolt connecting the insertion guide to the intramedullary nail during insertion of the femoral nail. | Device Design | On September 13, 2017 an URGENT MEDICAL DEVICE RECALL REMOVAL was issued to customers requesting that they review inventory and quarantine affected product, which will be immediately returned to recalling firm. Also to notify any customers who may have received product. Questions or concerns can be directed to Customer Service at 574-371-3071 |  | 337 | 337 total | Distribution includes US Nationwide. | Open, Classified |  |  | 337
Z-0776-2018 | 79132 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Intellivue iX Information Center Software - All PIIC iX Surveillance stations including: |  |  | January 03, 2018 | 2018 | Philips Electronics North America Corporation | Once a surveillance station is restarted on January 1, 2018 or later, the station will be unable to perform patient discharge and transfer operations. Any subsequent attempt to perform these operations will cause the station to restart, resulting in a short period of loss of monitoring at the Surveillance station during such restart. Until this issue can be corrected, users should avoid intentionally restarting their Surveillance stations in 2018. | Software design | To avoid this issue, do not intentionally restart any surveillance stations that have been running normally since the new year. You may use the workflows described in attachment Alternative Workflows for PIIC iX to avoid operations that may cause the surveillance station to restart. When a patient is discharged, confirm that the patient is discharged.Philips has initiated a correction to address this issue and will provide this software update to customers with affected devices at no charge. Software updates for PIIC iX A.0x and B.0x are expected to be available by January 8th.A Philips Healthcare representative will contact customers with affected devices to arrange for correction of the issue. Contact your local Philips representative if you have a question about any device affected by this issue. |  | 20357 | 20357 | Domestic and Foreign | Open, Classified |  |  | 20357
Z-0803-2018 | 79141 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The 2nd Assist Knee Positioner |  |  | September 07, 2017 | 2017 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a medical device recall for the 2nd Assistant Knee Positioner due to the potential for the sterile packaging barrier to be breached during distribution. | Package design/selection | On September 12, 2017 an URGENT MEDICAL DEVICE RECALL notice was issued to distributors and customers titled "The Second Assistant Disposable Knee Positioner". The notice requests customers to review their inventory and to quarantine any remaining product, which will be returned to the firm. Also, to notify any additional accounts which may have received product from you. Questions or concerns can be directed to Customer Service at 574-371-3071 |  | 38879 | 38,879 units | Nationally | Terminated | August 23, 2018 | 350 | 38879
Z-0778-2018 | 79226 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Alcon 25+Âż TOT ALPLUSÂż Vitrectomv Pak |  |  | January 12, 2018 | 2018 | Alcon Research, LTD. | Alcon is conducting this Voluntary Medical Device Recall for a specific lot (2060953H) of its 25+ TOTALPLUS Vitrectomy Pak due a potential that a 23GA ULTRAVIT probe may be included within the 25+ TOTALPLUS Vitrectomy Pak, instead of a 25+ ULTRAVIT probe. | Employee error | Alcon sent an Urgent Field Safety Notice Recall notification letters to affected foreign medical users/distributors beginning January 12, 2018.For further questions please call (713) 668-9100. |  | 300 | 300 units | Worldwide Distribution :ArgentinaAustraliaFranceItalyMexicoPolandSpainSwedenRussiaSouth KoreaUkraine | Open, Classified |  |  | 352
Z-0804-2018 | 79129 | No | LDR | Infusion pump. | General Hospital | General Hospital | 510(k) | STAT 2Âż I.V. Gravity Flow Controller and STAT 2Âż Primary, Secondary Administration Sets and Extension Sets (Catalog Numbers: (1) S2, STAT 2Âż I.V. Controller, (2) S2-12, STAT 2Âż Extension Set Gravity Flow Controller, 12 length, (3) S2-20 MD, STAT 2Âż Secondary Set Gravity Flow Controller, 20 Drops/ml, 13 length, (4) S2V-20, STAT 2Âż Primary Administration Set Gravity Flow Controller, 20 Drops/ml, 84 length, and (5) S2V-60, STAT 2Âż Primary Administration Set Gravity Flow Controller, 60 Drops/ml, 84 length) |  |  | June 19, 2017 | 2017 | ConMed Corporation | Intravascular (I.V.) fluid administration sets are exhibiting inaccurate or inconsistent flow rates. | Nonconforming Material/Component | The firm, ConMed, sent an "URGENT: DEVICE RECALL" letter dated 6/21/2017 to its domestic and foreign customers. The letter described the product, the problem and the actions to be taken. The customers were instructed NOT to use any STAT 2Âż I.V. Gravity Flow Controllers manufactured between April 27, 2012 and April 12, 2017; review your inventory for any of the devices listed on Attachment I; contact all of those departments within your facility and any other facilities within your organization that may have received affected products. It is imperative that all end users of these devices receive this notice and respond immediately; if you HAVE inventory of any of the devices listed, complete the Business Reply Form and return it with devices to: CONMED Corporation, 525 French Road, Utica, NY 13502 via UPS Account # W5Y243; if you DO NOT HAVE any affected devices complete and return the form by email: stat2@conmed.com; fax: Field Action Support Team at 315-624-3225 or US mail to CONMED Corporation, 525 French Road, Utica, NY 13502;Attn: Field Action Support Team.PLEASE DO NOT RETURN USED DEVICES.If you have any questions or requests, please dont hesitate to contact the Field Action Support Team at 1-800-448-6506, fax to 315-624-3225, or email stat2@conmed.com. |  | 1612440 | 1,612,440 in total | Worldwide Distribution: US (nationwide) to states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, Washington, DC, Puerto Rico, and countries of: Lebanon, Panama, Saudi Arabia, Germany, India, and Canada. | Open, Classified |  |  | 1612440
Z-0780-2018 | 79162 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | AVID TruCustom DAVINCI PACK convenience kits, Item Code: ADVU037-09 |  |  | June 22, 2017 | 2017 | Avid Medical, Inc. | Certain kits packed by the firm contain BD 22GA 1-1/2 inch SafetyGlide Non Sterile Bulk needles that may contain loose polypropylene foreign material above the release specification. | Nonconforming Material/Component | An email notification was sent to all sales representatives with the BD recall notification and the following instructions: " Examine inventory for the affected AVID catalog numbers and lots noted in the Where Used Chart. If affected product is found: (1) discontinue use and quarantine impacted product immediately; and (2) promptly apply a recall label to each affected lot found. Once labeled with the recall label, product may be taken out of quarantine and made available for use. " All other components in the MAI/ AVID tray are unaffected by this recall and may safely be used. Labels may be printed using the label template attached. Or, upon receiving the following information I will be happy to overnight the required labels. Any revisions or alterations to the labels are prohibited." Label Quantity- 1 for each affected tray, plus 1 for each outside carton." Contact Name" Address" Phone NumberSales representatives were asked to confirm to QualityAssurance@owens-minor.com that your customers were notified. |  | 50 | 50 | Distributed in 10 states: IL, KY, LA, MA, MO, NY, OH, TX, VA, WI. | Terminated | October 12, 2018 | 477 | 3722
Z-0691-2017 | 75291 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Sterile IV Start Kits |  |  | July 29, 2016 | 2016 | B. Braun Medical, Inc. | Positive results from non-routine sterility testing commissioned by B. Braun Medical Inc. (BBMI) for the finished product of this lot. | Unknown/Undetermined by firm | BBMI (B. Braun) provided written notification on August 05, 2016, to all BBMI customers in receipt of the suspected product. These customers included both BBMI direct end customers and BBMI direct distributors. Customers were asked to determine if they had the affected lots in their possession. If they did, customers were asked not to destroy the product and to immediately discontinue use of and quarantine if found. Customers were also asked to complete the Product Removal Acknowledgement form and return to B. Braun. B. Braun will contact each customer if they have any full cases, partial cases or unused pieces of the affected products to provide instructions for handling and return of the affected products. Customers with questions were instructed to call 1-800-227-2862. |  | 15950 | 15,950 units | Nationwide Distribution to GA, IL, and FL | Terminated | March 06, 2018 | 585 | 15950
Z-0777-2018 | 79289 | No | LWK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 3M Surgical Clipper Professional 9681 |  |  | February 02, 2018 | 2018 | 3M Company - Health Care Business | Failure to follow proper charging practices can result in lithium-ion battery degradation, characterized by excessive heat during operation or failure to properly charge. In rare circumstances, battery degradation has been associated with battery venting, which is a rapid release of battery energy. | Labeling design | The firm initiated their field safety action on 02/08/2018 by certified mail to customers and distributors. Larger distributor were notified by email on 02/12/2018. The notices to the distributors stated:"3M is notifying all users of the 3M" Surgical Clipper Professional 9681 of an added product warning, stressing the importance of following proper clipper charging practices as documented in the Model 9681 Clipper instructions for use. Current charging instructions recommend leaving the clipper body in the drop-in charger in between uses.Failure to follow proper charging practices can result in lithium-ion battery degradation, characterized by excessive heat during operation or failure to properly charge. In rare circumstances, battery degradation has been associated with battery venting, which is a rapid release of battery energy. 3M has received two reports of battery venting resulting in minor burn injury.3M is adding the following product warning to emphasize the importance of proper charging practices:Warning: To avoid potential battery degradation, charge the clipper after each use and store in the drop-in charger stand when not in use.Distributors are being asked to take the following actions:" Identify those customers who have ordered the 3M" Surgical Clipper Professional 9681 or the 3M" Surgical Professional Clipper Blades 9680 beginning January 1, 2016 to date." Communicate this Field Safety Notice to those customers."The notice to the customer stated the following:"3M is notifying all users of the 3M" Surgical Clipper Professional 9681 of an added product warning, stressing the importance of following proper clipper charging practices as documented in the Model 9681 Clipper instructions for use. Current charging instructions recommend leaving the clipper body in the drop-in charger in between uses.Failure to follow proper charging practices can result in lithium-ion battery degradation, characterized by excessive heat durin |  | 120000 | 120000 units | Worldwide distribution including US nationwide, including Puerto Rico. Singapore, Australia, Canada, Chile, Colombia, Costa Rica, Brazil, Dominica, Ecuador, Guatemala, UAE, Israel, Japan, Korea, Mexico, new Zealand, Paraguay, Peru, Taiwan, Uruguay, India, Bolivia, El Salvador, Saudi Arabia, and Thailand. | Open, Classified |  |  | 120000
Z-0828-2018 | 79251 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Berchtold Chromophare Ceiling Mounted Surgical Light System, Catalog #CH13000000. |  |  | January 29, 2018 | 2018 | Stryker Communications | The device could have a missing spanner nut and snap ring which could potentially result in the equipment falling. | Employee error | Stryker sent an Urgent Medical Device Correction letters to affected customers. The letter identified the affected product, problem and actions to be taken. The customer is informed a firm representative will be conducting a visual inspection of the equipment. No response form was enclosed. For questions email CommunicationsRA1@stryker.com |  | 1162 | 1,162 devices | Worldwide Distribution - US Nationwide Distribution and in the countries of Canada. | Open, Classified |  |  | 1162
Z-0812-2018 | 79248 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | 731 Series Ventilators running software version 05.20.00 |  |  | June 30, 2017 | 2017 | ZOLL Medical Corporation | A software anomaly in the 731 software version 05/20/00, was identified, which can lead to a user inadvertently changing device settings. | Under Investigation by firm | On June 30, 2017 an urgent recall letter was mailed to customers with the following instructions:Customers who have affected devices should immediately take the following steps: (1)Alert all users of 731 Ventilator models to this problem.(2)Direct users to always verify device parameters after selecting the device Confirm button.(3)Contact ZOLL's Technical Service Department or your local ZOLL Service Provider to schedule a software update for affected devicesOur 24/7 technical support numbers 1(800) 348-9011 or +l (978) 421- 9460 are available to assist users with any aspect of this notice. |  | 46 | 46 | Outside US | Terminated | October 09, 2018 | 466 | 46
Z-0818-2018 | 79105 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Ziptight Ankle Syndesmosis Fixation Device, Titanium Implant Titanium Implant for surgical repairs of soft tissue. |  |  | August 22, 2017 | 2017 | Zimmer Biomet, Inc. | Ziptight Ankle Syndesmosis Fixation was packaged without the tip protectors. | Packaging process control | Customers were notified via letter on about 08/22/2017. Instructions for distributors, sales representatives, and distributor operation managers include to ensure affected personnel are aware of the recall, locate and quarantine affected product in inventory, complete and return the Certification of Acknowledgement, return affected product, and identify any additional hospitals or surgeons that should be notified. Instructions for surgeons and risk managers include to ensure affected personnel are aware of the recall, identify and quarantine any affected product on hand, complete and return the Certificate of Acknowledgement, return any affected product. Questions or concerns can be addressed to the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST, Monday through Friday. Additionally, questions may be sent by email to corporatequality.postmarket@zimmerbiomet.com. |  | 885 | 885 units total | Distributed domestically to AK, AL, AR, AZ, CA, CO, FL, ID, IN, KY, MA, MD, ME, MI, MN, MO, NC, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI. Distributed internationally to ARGENTINA, AUSTRALIA, CANADA, COSTA RICA, JAPAN, NETHERLANDS. | Open, Classified |  |  | 885
Z-0832-2018 | 79192 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Wills-Oglesby Percutaneous Gastrostomy Set |  |  | October 19, 2017 | 2017 | Cook Inc. | Cook has identified that the affected lots may contain wire guides that were loaded backwards into the wire guide holder. | Under Investigation by firm | On October 19, 2017 an URGENT MEDICAL DEVICE RECALL letter was mailed to customers requesting that they examine inventory for affected product, immediately quarantine and return product to recalling firm for credit. Questions or concerns can be directed to 1-800-457-4500 or 812-330-2235. |  | 1 | Quantities co-mingled into line 1 | US Nationwide Distribution | Open, Classified |  |  | 1233
Z-0821-2018 | 79262 | No | OLP | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ultra Renew Plus |  |  | December 04, 2017 | 2017 | Accord Media, LLC | Distribution of medical devices with unapproved green LED light used to treat hyperpigmentation | Device Design | On December 4, 2017, Accord Media distributed Urgent Medical Device Field Action notices and Response Forms to customers via courier service. Customers were instructed to inspect stock and identify affected product whereby the would apply a label to the product which reads "DO NOT USE GREEN LIGHT FUNCTIONALITY - NOT FDA CLEARED FOR USE". Customers are encouraged to complete and return the response form via email to: customerservice@truthinaging.com. |  | 1040 | 1040 units | US Nationwide | Open, Classified |  |  | 3322
Z-0809-2018 | 79093 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | proSA Tool Set used to adjust and verify the proSA Adjustable Shunt System |  |  | September 21, 2017 | 2017 | Aesculap Implant Systems LLC | A proSA adjusting pin used to adjust the proSA Adjustable Shunt System does not set the correct valve opening pressure. | Process control | Aesculap sent an Urgent Medical Device Recall Notification letter dated September 21, 2017.1. Review the Urgent Medical Device Recall Notification Letter in its entirety and ensure that all users in your organization and other concerned persons are informed about this recall and immediately quarantine the unit. 2. Determine the current inventory of the affected lot(s) within your facility. Please quarantine the product and do use. Do not destroy any affected product. A Sales Representative will come to your facility and remove the affected sets and provide you replacements. 3. Please complete the attached Product Recall Acknowledgement Response Form on the next page of this letter and record the total number of units in your possession. If you have no inventory remaining, please enter zero (0) on the form and reason WHY. The Sales Representative will return the form along with the affected product to:4.The Sales Representative will contact Customer Service to order replacements sets and get a Recall RGR number to return the affected sets.For further questions, please call (610) 984-9414 Ext. 5414 |  | 2 | 2 | US Distribution to PA. | Terminated | August 08, 2018 | 321 | 2
Z-0808-2018 | 79178 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via systems with SW VB20A |  |  | August 31, 2017 | 2017 | Siemens Medical Solutions USA, Inc | A software update was released on August 22, 2017. Performing this software installation and upgrading syngo.via system to the software version VB20A_HF02 will resolve various bugs and inconsistent diagnosis with the old system. | Software Design Change | On August 25, 2017 an URGENT letter was issued to customers notify of a system update for Syngo.via systems with SW VB20A. The software was released on August 22, 2017. Several bugs and inconsistent diagnoses were detected and this software is intended to correct these. If you have any questions, call Siemens Customer Care Centre at 888-826-9702 (24/7). |  | 30 | 30 | Nationwide Distribution. | Terminated | October 01, 2018 | 396 | 30
Z-0826-2018 | 79176 | No | JLS | Progesterone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ST AIA-PACKÂż PROG III, Progesterone III Assay, Part Number: 025240 |  |  | February 13, 2018 | 2018 | Tosoh Bioscience Inc | Falsely elevated progesterone assay values will occur when using certain progesterone assay part numbers due to cross reactivity of DHEA-S, a metabolite of the DHEA (dehydroepiandrosterone) supplement. | Device Design | The assays identified in the above table must not be used to report results for patients who are taking DHEA supplements." For patients that are taking DHEA supplements, an alternative testing method which is not expected to show cross reactivity to DHEA-S should be used such as Liquid Chromatography-Mass Spectroscopy (LCMS)." Please review the information in this Urgent Medical Device Recall with your Medical Director and/or Lab Director, as soon as possible." Continue to use the Progesterone Assays for patients who are not receiving DHEA supplements or therapy." Complete and return the attached Urgent Medical Device Recall Acknowledgement Form within 15-days." Maintain this notification with your laboratory records and forward this information to others who may have received this product." If you have obtained unexpected test results or received any complaints of illness or adverse events associated with the use of the products identified in the table above, contact Tosoh Technical Support 24 hours a day, seven days a week at 1(800) 248-6764 or email Bernadette.OConnell@tosoh.com." Should you have any questions regarding this medical device recall, please feel free to contact Bernadette OConnell at 1(800) 248-6764, or by email at Bernadette.OConnell@tosoh.com. We will be available to answer any questions Monday - Friday, from 8:00 AM to 5:00 PM (PST). |  | 3322 | 3322 | Distributed in 39 states: AK, AL, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI. | Open, Classified |  |  | 17802
Z-0823-2018 | 79136 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | REVOLVE ADVANCED ADIPOSE SYSTEM, for aspiration, harvesting, filtering, and transferring of autologous adipose tissue for aesthetic body contouring. |  |  | August 08, 2017 | 2017 | LifeCell Corporation, A Wholly Owned Subsidiary of Allergan PLC | Presence of bacterial endotoxins levels above the acceptable limit | Material/Component Contamination | The firm's initial notice to customers (dated May 25, 2017) advising customers that Allergan has placed a hold on shipping the device until verification is complete. Within this notice, customers were advised to continue use of the product. On June 14, 2017, a second notice was issued to customers informing customers that product shipments will resume by August 1, 2017 and that manufacturing process improvements will be implemented. On August 8, 2017, Allergen distributed URGENT MEDICAL DEVICE RECALL notices and Business Reply Forms (dated August 2, 2017) to US customers advising them to take the following Actions: Actions to be taken1. Immediately examine your inventory and quarantine all product subject to this recall. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall, and that all product should be quarantined. Your notification to your customers may be enhanced by including a copy of this recall notificationletter.2. Carry out a physical count of the affected product in your possession and record the count on the enclosed postage paid Business Reply Form (BRF) and Packing slip.3. Mail the postage paid BRF immediately. To assure that we can account for all recalled product, it is imperative that you return the BRF.4. When returning the recalled product, attach the prepaid FedEx Authorized Return shipping label to the outside of the return carton. Return the recalled product and completed PackingSlip to:GENCO Pharmaceutical Services (GPS), a subsidiary of FedEx Supply Chain6101 North 64th Street, Milwaukee, WI 53218For Product Returns: Contact Genco Pharmaceutical Services, a subsidiary of FedEx Supply Chain at: 1-877-319-8961 during the hours of 7 AM - 5 PM CST.For Adverse Events/Product Complaints: Contact LifeCell at 1-800-367-5737, during the hours of 9 AM - 5 PM EST or via Email: complaintfeedbackregistration@lifecell.comFor Credit or |  | 25798 | 25,798 units | US Nationwide, Canada, Australia, Chile, Brazil, Israel, Japan New Zealand &amp; Singapore | Terminated | September 07, 2018 | 395 | 25798
Z-0827-2018 | 79053 | No | GJA | Thrombin time test. | Hematology | Hematology | 510(k) | BC Thrombin Reagent Kit Lot Number 46751; UDI - 008427680131024675120180908 and |  |  | October 17, 2017 | 2017 | Siemens Healthcare Diagnostics | The BC Thrombin Reagent kit lot 46751 (contains Thrombin Reagent lot 517468) and kit lot 47184 (contains Thrombin Reagent lot 517469) produce unexpected prolonged Thrombin Time (TT) results for expected normal samples, and may recover above the upper limit of normal (&lt; 21 sec.) as stated within the IFU (Instruction For Use). | Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notice letter dated October 2017 to affected customer.The letter identified the affected product, problem and actions to be taken by the customer. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter and the Urgent Field Corrective Action letter within 30 days.For questions contact your Siemens Customers Care Center or local Siemens technical support representative. |  | 7631 | 7631 units | Worldwide distribution. US nationwide, Austria, Belgium, Bosnia-Herzegovina, Croatia, Finland, France (incl. Guadeloupe, Martinique, Reunion, French Guinea overseas departments), Germany, Greece, Hungary, Italy, Kazakhstan, Macedonia, Latvia, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Serbia, Slovakia, Slovenia, Spain, Sweden, Turkey, Argentina, Brazil, El Salvador, Mexico, Peru, Uruguay, Canada, and Australia. | Open, Classified |  |  | 7631
Z-0807-2018 | 78724 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products Performance Verifier is used as an assayed control to monitor the performance on VITROS 250/350/5, 1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System. |  |  | July 10, 2017 | 2017 | Ortho Clinical Diagnostics Inc | VITROS Performance Verifier II (PVII) Lot K4852 does not meet long term stability requirements for the AST Analyte (only) when the product is stored refrigerated for up to 6 months. | Unknown/Undetermined by firm | On July 10, 2017 the firm sent notification letters and an e-communications to their customers. Foreign customers were notified by email. Customers were instructed to complete a response form with the amount of product on hand and to discontinue use of the recalled product. For further questions, please call (585) 453-3452. |  | 10021 | 10,021 units | Worldwide | Open, Classified |  |  | 10021
Z-0813-2018 | 79274 | No | LRY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical punch; sealed pouches packed in punch shelf box |  |  | October 23, 2017 | 2017 | Alto Development Corp | Surgical punches are being recalled due to potential packaging damage which may have occurred during transportation resulting in loss of sterility. | Package design/selection | A&E; Medical sent an Recall Notice Action Required letter dated October 20, 2017. Surgical Punches are being recalled due to potential packaging damage which may have occurred during transportation resulting in a loss of sterility. Recent testing identified certain shipping configurations which may cause damage to the sterile barrier pouch if the Surgical Punches experience excessive movement from multiple drops during transportA&E; is recalling all product currently on the market. Customer and Distributor Letters were sent with Acknowledgement Form. Product will be quarantined pending rework. For further questions, please call (800) 323-4035. |  | 411486 | 411,486 units | Worldwide Distribution - USA Distribution | Open, Classified |  |  | 411486
Z-0810-2018 | 79147 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | AccuLIF TL Insertion Handle |  |  | September 08, 2017 | 2017 | Howmedica Osteonics Corp. | When the lever of the Insertion Handle is in the "locked" position, the implant may be able to pivot rather than maintaining a rigid connection. | Under Investigation by firm | A recall letter, dated August 25, 2017, titled "Product Recall Re: AccuLIFTL Insertion Handle" instructing customers to meet with the Stryker Representative issuing the letter, complete the requested form and return any affected product to this representative. Questions or concerns can be directed to Customer Service at 201-749-8699 |  | 130 | 130 | Nationally | Terminated | July 31, 2018 | 326 | 130
Z-0834-2018 | 79115 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | DELTA XTEND" Reverse Shoulder System Modular Centered and Eccentric Epiphysis Implants intended for use in total shoulder or hemi-shoulder replacement procedures in patients with non-functional rotator cuffs. HA components are for cementless use only. |  |  | August 31, 2017 | 2017 | DePuy Orthopaedics, Inc. | The affected lots are being recalled because the epiphysis may not assemble to the stem, which may cause a surgical delay. | Nonconforming Material/Component | On September 11, 2017 an URGENT MEDICAL DEVICE RECALL notice was issued to customers for specific lots of the DELTA XTEND Reverse Shoulder System Modular Centered and Eccentric Epiphysis Implants. The letter instructs customers to inspect inventory and to remove and return the affected lots to the recalling firm. Questions or concerns can be directed to Kim (Earle) Long, Senior Recall Coordinator, 574-371-4917. |  |  | Quantities co-mingled under initial product | Nationally | Open, Classified
Z-0806-2018 | 79090 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | Sterile EXACTAMED Oral Dispenser, Clear, Individually Packaged, 5 mL, Product Code H93876105 |  |  | February 06, 2018 | 2018 | Baxter Healthcare Corporation | Affected lot of 5 mL oral dispensers was packaged in individual pouches that are mislabeled with the incorrect product code (H93876101) and volume (1mL). | Employee error | Customers were notified via letter on about 02/06/2018. Instructions include to locate and remove affected product in inventory, confirm mislabeled 5 mL oral dispensers (Lot number 1203395) in the individual pouch are not mixed together with 1 mL oral dispensers (Product code H93876101) or other stored product, contact Baxter Healthcare Center to arrange return of affected product, and complete and return the customer reply form if purchased directly from Baxter. Also, customer are instructed to notify customers and conduct a consumer-level recall if the affected product was further distributed. Questions and concerns can be directed to the Public Contact, Center for One Baxter with phone number (Toll Free) 800-422-9837 and 847-948-4770, Monday through Friday 8:00 AM to 5:00 PM CST. |  | 2800 | 2,800 units | Distributed domestically to IL, MA, MD, OH, SD, TX, VA, WI. Distributed internationally to Canada. | Open, Classified |  |  | 2800
Z-0815-2018 | 79019 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | OTTO BOCK Kenevo knee joints Model 3C60=ST |  |  | September 12, 2017 | 2017 | Otto Bock Healthcare Product | Otto Bock Healthcare Products GmbH has identified a design issue through extended durabilitytesting where the pylon clamping mechanism has the potential to cause damage to thepylon. When used longer than 1.45 million cycles (which represents approximately 2.3years of use of the knee joint under heavy usage conditions), AND the pylon clamp is overtightenedaccording to the labeled torque specification, AND the patient weighs at most125kg it may be remote that the pylon breaks, and causing the patient to fall.There have been no failures or complaints reported directly attributed to the failure mode collapsed pylon .In the US Market and in Canada, labeling requires a condition based service at three years.In the rest of the world, a mandatory two year service interval is specified. Only the devicesin distribution in the United States and Canada are affected by this action. | Device Design | On September 14, 2017 a letter was sent to all their consignees via 2nd day signature receipt required to customers instructing them to contact their patients and send in their devices for service within the first 2.3 years of use. The pylon clamp will be replaced at no charge. The pylon will also be inspected and replaced if necessary. The customer will be requested to complete a reply form indicating they have received the notice and intend to contact their patients to arrange for the requested service.Customer Service: (800) 328-3458  Option 1Loaner Service: (800) 328-3458  Option 3, then Option 1 |  | 4 | 4 | AK, AR, AZ, CA, CT, DC, GA, IL, IN, KY, LA, MD, MI, MN, NC, ND, NJ, MN, NY, OH, OK, OR, PA, PR, TX, VA, and WAAU, BE, DE, FR, GB, IT, KW, LU, NL, NO, AT, SA, SE, CH, ES, ZA, and HU | Open, Classified |  |  | 158
Z-0811-2018 | 79185 | Yes | FKR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage Dialysate Sack: 1) REF SAK-301 UDI-M535SAK3010, 2) REF SAK-302 UDI-M535SAK3020, 3) REF SAK-303 UDI-M535SAK3030, 4) REF SAK-304 UDI-M535SAK3040, 5) REF SAK-306 UDI-M535SAK3060, 6) REF SAK-307 UDI- M535SAK3070, 6) REF SAK-402 UDI-M535SAK4020, 7) REF SAK-405 UDI- M535SAK4050, 8) REF SAK-406 UDI-M535SAK4060 &amp; REF SAK-407 UDI- M535SAK4070 |  |  | January 29, 2018 | 2018 | NxStage Medical, Inc. | May contain endotoxin levels which have been confirmed to exceed the ANSI/AAMI quality standard dialysis fluid of 0.5 EU/ml when used to prepare product. | Material/Component Contamination | On January 29, 2018 Product Recall Notices &amp; Recall Reply Forms included were emailed to those patients for whom the firm had an email address. Hard copy letters were sent via UPS second day air to patients for whom NxStage did not have an email address. All Dialysis centers were contacted via email. In addition to the active patients and dialysis center consignees of affected product, recall letters were also emailed to the Dialysis Center Home Training Nurse Manager for all active or dropped patients who were shipped affected product, and corporate contacts for national dialysis accounts. Customers are advised to inspect inventory, quarantine, discontinue use and return product. Customers are also encouraged to return the completed Recall Reply Form. Contact NxStage Customer Service to arrange for return of all affected product and for replacement product to be sent. Customers with questions may call Customer Service at 1-866-697-8243. |  | 32078 | 16,039 cases/2 units | Worldwide Distribution - USA (nationwide) and to the countries of : United Kingdom &amp; Spain | Open, Classified |  |  | 32078
Z-0816-2018 | 79160 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Ultrasound Transducer, Model Number: UST-5550-R, is used in conjunction with a diagnostic ultrasound system evaluation during robotic and non-robotic intra-operative and laparoscopic procedures. |  |  | October 20, 2017 | 2017 | Hitachi Medical Systems America Inc | The ultrasound probe may not have adequate protection against electrical shock hazards. | Mixed-up of materials/components | Hitachi must recall the lot of UST-5550-R Transducers from S/N 1 to S/N 155. If you have a serial number in this range, please contact Hitachi Service at (800) 800-4925 to schedule the recall. Hitachi Sales will contact you regarding the options for a replacement transducer. |  | 55 | 55 | Distributed in 21 states: AR, AZ, CA, CO, FL, GA, IA, IL, IN, MA, MO, NC, NY, OK, OR, PA, SC, TX, VA, WA, WV. | Open, Classified |  |  | 55
Z-0824-2018 | 79336 | No | FWB | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Camera System, Television, Surgical without audio |  |  | December 15, 2017 | 2017 | Trumpf Medical Systems, Inc. | The radiated emission from the TruVidia" Wireless Receiver is above the normative accepted limits, in which the radiated emission may interact with other devices in the operating room and cause them not to function properly. | Other | Trumpf Medical Systems sent an Urgent Medical Device Correction letter dated December 15, 2017. Instructions include to complete and return the Confirmation of Receipt, arrange for the Trumpf Medical Engineer or an authorized Service Engineer to upgrade the affected device, and ensure that all users of the device are aware of the recall. For further questions, please call (843) 329-0543 |  | 212 | 212 devices | Worldwide Distribution - US Distribution to the states of : GA, IL, IN, PA, SC, TX, UT, VA. and to the countries of : Austria, Australia, Bangladesh, Belgium, France, Switzerland, Czech Republic, Germany, Denmark, Spain, United Arab Emirates, United Kingdom, Ireland, Italy, Kuwait, Lithuania, Poland, Portugal, Romania, Sweden, Slovenia, Thailand, Turkey. | Open, Classified |  |  | 212
Z-0505-2018 | 78897 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Peripheral Diamondback 1.25 Solid OAD, a percutaneous orbital atherectomy system indicated for use as therapy in patients with occlusive atherosclerotic disease in peripheral arteries and who are acceptable candidates for percutaneous transluminal atherectomy. |  |  | September 14, 2017 | 2017 | Cardiovascular Systems Inc | Cardiovascular Systems, Inc. (CSI) is removing the products because it was identified that the labeling for two lots of product s were potentially inadvertently switched during production. | Labeling mix-ups | On 09/14/2017 the firm sent letters to their consignees explaining the recall and given the following instructions:Please immediately remove from use and return the affected Product Name devices to CSI. To do so:1) Complete the attached Customer Acknowledgement Forma. Per lot, document the quantity of all devices that are to be returned.b. Per lot, document the quantity of all devices that have been consumed.2) Place the Acknowledgement Form with the device3) Return the device and Acknowledgement Form to CSI using the pre-paid FedEx labela. If you would like, a CSI Sales Representative can return the device and form. Please contact the CSI Recall Coordinator to organize pick up and return of the product. If all products specified in this recall notice have previously been consumed, used or are no longer available for return to CSI, please complete the Customer Acknowledgment Form per the above, and return the form to CSI in the provided FedEx envelope.Contact Information: csifieldactionhotline@csi360.com651-259-2800 - Tel |  | 8 | 8 | FL, IL, MI, MN, MO, NY, SC, TX, and WI | Completed |  |  | 8
Z-1171-2018 | 79196 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | **Recall Expanded to add these models: Bigger Better-Bladder cardiopulmonary bypass blood reservoir; Models: BBB38 (individual, sterile, pouched) and BBB38NS (sold bulk to kit manufacturers) |  |  | September 22, 2017 | 2017 | Circulatory Technology Inc | The device is used as part of extracorporeal membrane oxygenation (ECMO) cardiopulmonary bypass circuits and may collapse during use which can increase resistance to flow in the venous line and cause a drop in circuit blood flow. | Process control | The firm is requesting that consignees immediately inspect inventory, and quarantine all units having the aforementioned lots. These units must be returned to CTI. CTI will provide UPS shipping labels for returning the units and will provide replacements or credit for the returned units. Should the inventory consist of only the affected lots then, until the consignee receives replacement BB14s from an unaffected lot, the firm recommends using the BB14 from the affected lots only if, prior to its use, they pull on the pigtail of the Better-Bladder to test its bond integrity prior to its use. If the pigtail dislodges then return the unit for credit or replacement. The consignees should complete and return the enclosed response form via email to ytamari@cirtec.com as soon as possible. |  | 410 | 410 | **Requested update** Distributed in 19 states: AL, AR, CA, FL, GA, IL, IN, KY, MD, MO, NJ, NY, OH, OR, PA, TN, TX, VA, WA, and the District of Columbia. | Open, Classified |  |  | 1250
Z-0726-2018 | 78679 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | The GC80 Digital X-ray Imaging System |  |  | June 22, 2017 | 2017 | NeuroLogica Corporation | An image was obtained with over exposure during a Thorax examination using the AEC function. | Other | On June 22, 2017, Samsung distributed Urgent Field Safety Notices &amp; acknowledgment forms to their customers via email. INSTRUCTIONS FOR END USERS:1) Please read the Field Safety Notice documentation in detail.2) Follow the instructions provided in this FSN in relation to any actions which are required of you.3) Complete the provided acknowledgement form and return to your supplier as soon as possible. Your organizations reply is the evidence we need to monitor the progress of the corrective actions.4) If required and instructed by this FSN documentation, please undertake all relevant service actions as soon as is practicably possible.INSTRUCTIONS FOR DISTRIBUTORS:1) Please read the Field Safety Notice documentation in detail.2) Please review and complete the acknowledgement form provided in this FSN and return to Samsung as soon as possible.Please pass this entire package of FSN documentation onto your customers and ask them to return the completed acknowledgement form. Please forward copies of acknowledgements from end users onto to the email addresses stated above. Samsung require this evidence to monitor the progress of the corrective actions.4) If required and instructed by this FSN documentation, please undertake all relevant service actions as soon as is practicably possible.Customers with questions were advised to contact their local representative. |  | 45 | 45 units | Nationwide US | Terminated | October 05, 2018 | 470 | 45
Z-0846-2018 | 79133 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | A-Series A3/A5 Anesthesia Delivery System; Model Numbers: 0633F-01000-0X (A3) and 0631F-01000-0X (A5) |  |  | July 14, 2017 | 2017 | Mindray DS USA, Inc. dba Mindray North America | A software issue may result in the previous settings being applied instead of the default settings or the unit may skip the startup leak test. | Device Design | Mindray sent an Urgent Medical Device Correction letter dated June 9, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. Corrective Action: In order to eliminate any potential for the EEPROM issue to manifest, Mindray is offering two options for upgrading affected systems in your facility: Option 1: You may perform the software upgrade if your facility has Biomedical Engineering or other staff that have been trained by Mindray to perform maintenance operations on the ASeries. Please contact Mindray's Technical Support team who will assist with providing the software and instructions for performing the software upgrade. The Technical Support can be reached at (877) 913-9663 (Option 1). Option 2: A Mindray Service Representative will upgrade the affected systems in your facility at no charge. Please contact your Mindray Service Representative to arrange for this upgrade at (877) 913-9663 (Option 8). Both the Technical Support and Service teams may be reached Monday through Friday, 8:30am - 5:30pm, ET. |  | 2875 | 2875 in total | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 2875
Z-0838-2018 | 78668 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ACL DISPOSABLE PACK BONE-TENDON-BONE, Model 0234020280, labeled sterile. |  |  | November 17, 2017 | 2017 | Stryker Corporation | Routine bioburden testing of certain lots were found to have levels higher than internal acceptable rates. It is indicated to be a single surgery use, disposable set of instrumentation for anterior cruciate ligament (ACL) reconstruction surgeries. | Under Investigation by firm | On about 11/17/2017 letters were sent by certified mail with tracking information to all domestic accounts and International Stryker sites were notified by email. Instructions include to inform individuals who need to be aware of the device recall, examine all stock areas and/or operating room storage for affected product quarantine and discontinue use of the recalled devices, complete the acknowledgement form, and if affected product is found contact Stryker customer service at 1-800-624-4422 (Option 3) or email endocustomersupport@stryker.com to arrange for product return. |  | 60753 | 60,753 units total | Distributed domestically to . Distributed internationally to Australia and Mexico. | Open, Classified |  |  | 60753
Z-0840-2018 | 79134 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance Big Bore Oncology CTwith 4.2.0 software version. Model #' 728243 |  |  | January 03, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | Software issues including:Software errors that may result in Extended Field of View check message not displayed when Field of View is set &gt;600 mm for a paused scan;Request for change to automatically turn on the center x/y 0 setting and the disable result rotation setting when turning on the CT simulation Exam Card;Software errors that may result in Tumor LOC allowing POI (Point of Interest) locking when a non-isocenter slice is displayed;Software errors that may result in partial set of images or inability to generate CT images. | Under Investigation by firm | Philips sent an Urgent Field Safety Notice Medical Device Correction letter dated January 2, 2018 to their customers. The letter identified the affected product problem and actions to be taken. For questions contact your local Philips representative or local Philips Healthcare office. |  | 11 | 11 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 130
Z-0845-2018 | 79153 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | INFX-8000H Fluoroscopic X-Ray Systems |  |  | December 29, 2017 | 2017 | Toshiba American Medical Systems Inc | During an examination a device error occurred and the ceiling drive operation was disabled. It was found that the fastening bolts of the driving portion of the ceiling drive loosened cause the error. | Process control | A recall notification letter was sent to affected consignees on 12/29/2017. The following information was included in the notification letter: Should this problem occur prior to the above corrective action and in the middle of a clinical case, youwill be able continue the case tentatively. The x-ray emissions will be still available and the inhibitedceiling support device can be overridden by pressing the override switch on the left side of theoperation console.If any abnormalities are found during use, please stop using your system and contact your Toshibaservice representative.Please share this information with all users and reviewing radiologist as well as clinical engineering andthe biomedical group at your facility. |  | 2 | 2 units | United States | Open, Classified |  |  | 48
Z-0841-2018 | 79148 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IQon Spectral CT with 4.7.2 software version.- Model 728332, a Philips Computed Tomography X-Ray System |  |  | January 11, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | 23 Software related issues for the IQon Spectral CT with 4.7.2 software version. These issues include:"Software errors that may result in CT spectral images misrepresentation or incorrect treatment;"Software errors that may result in partial set of images or inability to generate CT images. | Under Investigation by firm | Philips Healthcare is notifying the affected users of these issues via a Field Safety Notice and will be releasing a software fix via Field Change Order to correct these issues. They also provide the following instructions.:Issue 1, Uric acid application for gout provides inconsistent results - When reviewing images using the Uric Acid Spectral Result type inconsistent results were provided on gout studies. The tendons would occasionally be displayed as containing Uric Acid even with no or normal uric acid blood levels.Issue 2, Multi-phase Cardiac CTA Scans should include a planned conventional result for every planned spectral result - When planning a Cardiac multi-phase study, it is possible to create a phase that consists of only a Spectral Result.Issue 3, 0-MAR information is missing in DICOM when saving as original or Secondary Capture - When an SBl that was created with 0-MAR enabled is used to create and save any Spectral result:" DICOM information about 0-MAR is missing but the 0-MAR label appears on a result saved as Original" DICOM information and the 0-MAR label are both missing on a result saved as Secondary Capture Issue 4, Attenuation plot presents incorrect HU values for MonoE - When comparing HU values of an ROl to the values in the attenuation plotgraph when using the Show spectral plots for this ROI option, there is adifference of 10-20 HU values between the ROI HU value and the valueshown on the graph Issue 5, Windows Operating System Time Zone Updates - Microsoft Windows Daylight Savings Time Zone updates have not been applied since December 2014. This may cause the system time to be incorrect at the start and end of daylight savings time. Issue 6, Sending Calcium Score findings to report using Mass Scoreprotocol provides confusing information - The Heartbeat CS (Calcium Scoring) option provides confusing information in the Impressions section of the Calcium Scoring report. The report provides m |  | 63 | 63 | Domestic and Foreign | Open, Classified |  |  | 63
Z-0849-2018 | 79072 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Concorde Lift Torque Limiting Handle. |  |  | July 26, 2017 | 2017 | DePuy Orthopaedics, Inc. | Potential for Intra-operative breakage of driver tips | Under Investigation by firm | On August 1, 2017, an Urgent Product Recall notice titled "CONCORDE LIFT DRIVER Driver Shaft and Torque Handle" was mailed to customers that received the affected instruments. The letter described the issue, potential hazard, and actions to be taken. The notice instructs customers to cease further distribution or use and to contact a DePuy Synthes Spine sales consultant to return the products subject to recall. Customers are to review, complete, sign, and return the business reply form provided to the firm within 5 business days of receipt of the notification. A copy of the notice should be forwarded to all staff that need to be informed, as well as, any facility that the affected device was further distributed to. Customers should direct any questions regarding this recall to their DePuy Synthes Spine Sales Consultant or the Clarke Madigan, DePuy Synthes Spine Recall Coordinator, at 508-828-609 or DPYUS-SpineFieldActions@its.jnj.com. |  | 130 | 130 units total | Nationwide Distribution | Open, Classified |  |  | 130
Z-0850-2018 | 78849 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Network Firewall (Cisco ASA 5506), Model # 867098 |  |  | October 05, 2017 | 2017 | Philips Electronics North America Corporation | Firewall installed with Philips IntelliVue Information Center iX or Information Center Classic may have a defective component, which may result in loss of connection to the Information Center iX. The defect involves the clock signal component within the firewall. This component has a high probabilityof failing in appliances that have been running for greater than 18 months. If the clock signal component were to fail, the firewall will stop functioning, will not boot, and is not recoverable.This failure will result in loss of communication between devices that are separated by the firewall, which may cause the Information Center to reboot. | Nonconforming Material/Component | Philips sent an Urgent Recall Notice letter dated October 2017 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that a Philips representatives will contact affected customers to arrange for replacement of all affected units supplied by Philips. For questions contact your local Philips representative: Philips Customer Care Solution Center 800-722-9377. |  | 332 | 332 units | US Nationwide | Open, Classified |  |  | 332
Z-0896-2018 | 78609 | No | LKK |  |  | General Hospital | PMA | Medtronic Synchromed II Model 8637 supplied in 20mL and 40 mL reservoir size and a catheter. |  |  | August 09, 2017 | 2017 | Medtronic Neuromodulation | Medtronic implemented a final design change intended to address the motor stall due to corrosion and is recalling devices with the previous design that are not implanted. | Device Design | Medtronic sent an Urgent Medical Device Recall letter dated September 2017, to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the enclosed Customers Confirmation form. For questions contact Technical Services at 1-888-854-0978. |  | 40572 | 27,547 US and 13,025 OUS | US Nationwide Distribution in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY | Open, Classified |  |  | 40572
Z-0895-2018 | 79165 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Hemosll ReadiPlasTin (20 ml Size), Part Number 0020301400 |  |  | August 02, 2017 | 2017 | Instrumentation Laboratory Co. | Instrumentation Laboratory Co. received customer reports of performance issues with some vials of Lot Nos. N1166235 and N0177760, including increased imprecision, out of range quality controls and prolonged sample results. | Under Investigation by firm | On August 2, 2017, Instrumentation Laboratory sent an Urgent Medical Device Removal letters dated August 2, 2017 to their customers. On October 5, 2017 Instrumentation Laboratory expanded their recall adding 3 additional lots to a second Urgent Medical Device Removal letter. On February 26, 2018, Instrumentation Laboratory expanded their recall a 3rd time to include 1 additional lot to a third Urgent Medical Device Removal letter. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete a Mandatory Tracking Response form. For questions contact your Customer Service Representative at 1-800-955-9525, option #1. |  | 14894 | 14,894 kits | Worldwide Distribution - US Nationwide | Completed |  |  | 14894
Z-0897-2018 | 79096 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Auriga 30 Laser System, Universal Part Number: M0068S30G0. |  |  | September 22, 2017 | 2017 | Boston Scientific Corporation | Certain Auriga consoles have incorrectly assembled trigger wire components. The trigger wire functions in the start-up of the console. An incorrectly assembled trigger wire may result in error codes or failure to start the console. | Process design | Urgent Medical Device Field Correction letters, dated September 2017, were sent to distributors and medical facilities with instructions for acknowledging the notice and making corrections. The letter identified the affected device and the reason for the correction. Customers were to check their inventory, segregate, and cease distribution for the affected devices. Customers were to complete and return the acknowledgment form.**NO PRODUCT IS BEING RECALLED AND YOU ARE NOT REQUIRED TO RETURN PRODUCT TO BOSTON SCIENTIFIC** |  | 2 | 2 | International Distribution Only -- Australia, Vietnam, France, Germany, Italy, Spain, and Romania. | Terminated | August 14, 2018 | 326 | 11
Z-1189-2018 | 78667 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Tapered Screw-Vent MTX, Dental Implant, sterile. Model TSVTB10. |  |  | November 29, 2017 | 2017 | Zimmer Dental Inc | Two lots of the Tapered Screw-Vent Implants, with different outer thread diameters, were inadvertently commingled during the manufacturing process and may be mislabeled. | Packaging process control | On 11/29/2017, customers were notified via letter. Instructions for distributors include to ensure that affected team members are aware of the recall, immediately locate and quarantine affected devices in inventory, notify consignees if product was distributed, make arrangements to immediately return affected devices, and complete and return the Attachment 1 - Inventory Return Certification Form. Instructions for clinicians include to review the recall notice, examine inventory for affected devices, complete and return the Attachment 2 - Certificate of Acknowledgement, and return all affected devices. |  | 995 | 995 distribution events total | Worldwide distribution US nationwide, Canada, Chile, Costa Rica, Denmark, France, Hungary, Italy, Japan, Lebanon, Portugal, Serbia, Spain, and Taiwan. | Terminated | February 08, 2019 | 436 | 995
Z-0891-2018 | 76874 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | LASEREDGE Knives, 2.0MMX2.2MM TRAPEZOID ANGLED (6/BOX), REF/PRODUCT CODE E7601, STERILE, Rx Only |  |  | March 03, 2017 | 2017 | Bausch &amp; Lomb Inc Irb | The complaints database was reviewed and there have been an atypical volume of dull blades complaints for various lots. out of an abundance of precaution, Bausch + Lomb is voluntarily recalling all product codes SKU manufacturing without the electro-polish control and penetration force testing. | Process design | On March 3, 2017 Bausch + Lomb , distributed Urgent Medical Device Recall notices and Acknowledgement Forms to customers via FedEx. Customers were advised to: 1.Please review your inventory and quarantine all unused (full and partial) boxes of LaserEdge" knives (6/Box or individual packaged knives) affected lots. 2.Complete the enclosed Medical Device Voluntary Recall Acknowledgement Form and contact Bausch + Lomb to obtain a Return Material Authorization Number (RMA) and arrange for a pickup of the identified product. You can contact the Bausch + Lomb Surgical Customer Service team by calling 1-800-338-2020, Customer Service (option 2), Product Returns and Adverse Event Reporting (option 2).Please contact the Bausch + Lomb Surgical Customer Service team with any questions or concerns regarding this process: 1-800-338-2020. |  | 1236749 | 1,236,749 knives in total | US Nationwide Distribution | Terminated | June 25, 2018 | 479 | 1236749
Z-1927-2018 | 79340 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ALM XTEN Surgical Light intended to provide visible illumination for the surgical field or for the examination of the patient. Part Numbers: ARD567801093, ARD567801094 |  |  | November 16, 2017 | 2017 | GETINGE US SALES LLC | Getinge has received complaints concerning the X TEN Surgical Lights whose spring arms have broken, causing light heads to separate from the assembly. The issue relates to cracking of the spring arms at the weld seams that has the potential to develop into breaks with heavy use and aging of the equipment. | Unknown/Undetermined by firm | Maquet/Getinge notified customers via an Urgent Medical Device Recall Notice dated November 16, 2017. |  | 667 | 667 | U.S. and Gov. accts. | Open, Classified |  |  | 667
Z-0936-2018 | 79352 | No | GAZ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Argyle Suction Tubing (Product # 8888301515) |  |  | October 24, 2017 | 2017 | Medtronic Minimally Invasive Therapies Group | Incorrect packaging. In one lot of Argyle Suction Tubing, the pouch may be incorrectly labeled as the Salem Sump Dual Lumen Stomach Tube. This product may have been distributed within the case of Argyle Suction Tubing. | Packaging process control | An URGENT MEDICAL DEVICE RECALL letter was issued to customers on October 24, 2017. This letter notified customers of the products subject to the recall and stated the following: The primary package labeling (Label closest to the product) may be incorrect. Customers can identify the affected product by consulting Attachment B. Please quarantine and return any unused products of the item codes and lot detailed above. Unused products from the affected item code and lots should be returned as described in the Required Actions section below. If you have distributed the Argyle" Suction tubing listed above, please promptly forward the information from this letter to those recipients. All unused products from the affected item codes and lots must be returned. If you have any questions regarding this communication, please contact your Medtronic representative or Customer Service at 800-962-9888, option 2. |  | 58954 | 58,954 | USA (nationwide) Distribution | Open, Classified |  |  | 58954
Z-0906-2018 | 79157 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 2.5 mm Titanium Calibrated Reaming Rod, 850 mm, Part Number: 359.083S |  |  | July 26, 2017 | 2017 | Synthes (USA) Products LLC | Possible lack of product sterility due to potential gaps/channeling in the seal of the package. | Process control | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated 7/26/2017 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to do the following: If you DO have any of the identified devices and intend to re-sterilize the product (TRAUMA Products only) per the package insert, please take the following steps:Âż Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located and will be re-sterilized. Also, please indicate the number of devices found. Please include your name, title, telephone number and signature in the spaces provided.Âż Send a copy of the completed Verification Section to Synthes by:o Fax: 844-721-3045 oro Scan/email: Synthes5735@stericycle.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 4 of the notification.If you DO have any of the identified devices, and are returning the product for replacement, please take the following steps:Âż Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number.Âż Complete the Verification Section (page 4 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the ReturnÂż Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided.Âż Return the Verification Section (page 4 of this letter) with the product to:o Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132.Âż Send a copy of the completed Verification Section to Synthes by:o Fax: 844-721-3045 oro Scan/email: Synthes5735@stericycle.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clea |  | 63 | 63 | Worldwide Distribution: US (Nationwide) to states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, the District of Columbia, including Puerto Rico, and country of: Canada. | Terminated | September 10, 2018 | 411 | 8667
Z-1241-2018 | 79353 | No | PHX | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AltiVate Reverse INSTRUMENTATION, HUMERAL SOCKET SHELL TRIAL, STANDARD, REF 804-06-152, Qty: 01, Encore Medical, LP Orthopedic surgical instrument. |  |  | January 30, 2018 | 2018 | Encore Medical, Lp | Broken or missing screws and retaining clips. This may be due to screw tolerance and material strength of the trial. | Device Design | The firm disseminated their field safety notices on 01/30/2018 by email. The letter identified the affected product, problem and actions to be taken. The notices requested the trails be returned for replacement. |  | 387 | 387 units | US Nationwide Distribution | Open, Classified |  |  | 387
Z-1134-2018 | 78922 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medfusion Syringe Pump 4000 Series |  |  | November 13, 2017 | 2017 | Smiths Medical ASD Inc. | Certain Medfusion Syringe Pump Models, Series 3100, 3500, and 4000, may not recognize or may misidentify loaded medication syringes. The inability of a pump to recognize a syringe (i.e. the size of the syringe is unknown to the pump) results in an inability to complete pump programming. Misidentification of a syringe is where the pump misinterprets the syringe size. | Process change control | Consignees were notified via mail on 11/13/2017, 01/09/2018, and 02/12/2018, as a result of expansions of the recall and the need to provide additional instructions and information to consignees. Instructions in the first letter date 11/13/2017 instructed customers to always confirm the accuracy of infusion values to the original order, including verifying syringe size, to utilize the enclosed Reference Tool when verifying syringe size, and to complete the return the Response Form. Instructions in the letter dated 01/09/2018 included to always confirm the accuracy of infusion values to the original order, including verifying syringe size, to advise all clinicians overseeing use of the affected devices to utilize the enclosed Reference Tool, to remove the pump if the syringe is not recognized or is read as an incorrect size during programming or infusion, to notify impacted customers if the devices were further distributed, and to complete the return the Response Form. Instructions in the letters dated 02/12/2018 included to identify any affected devices on hand, complete and return the Urgent Medical Device Recall Response Form, and arrange for repair of the affected as per the location of the consignee in Canada or the USA. If the consignee is located in the USA, the three options include to return to Smiths Medical for repair, arrange for the shipment of the required parts and have the consignee's onsite Biomed Team complete the repair, or arrange for onsite repair by Smiths Medical if there are more than 10 pumps onsite. If the consignee is located in Canada, the two options include to return to Smiths Medical for repair or arrange for the shipment of the required parts and have the consignee's onsite Biomed Team complete the repair. Consignees not located in the USA or Canada were instructed to contact their Smiths Medical Representative to discuss repair options. For further questions, please send inquiries to SmithsMedical3033@stericycle.com. |  | 16600 | 16,600 pumps total | Worldwide distribution- US nationwide Distribution and to the countries of : Australia, Barbados, Canada, Hong Kong, Indonesia, Jamaica, New Zealand, Philippines, Saudi Arabia, Singapore, South Africa, Trinidad and Tobago, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 16600
Z-1166-2018 | 79342 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROWg+ard Blue PLUS Pressure Injectable Multi-Lumen CVC Kit. |  |  | November 14, 2017 | 2017 | Arrow International Inc | Products do not contain the Central Venous Catheter Placement Procedures Note or the Central Line Insertion Checklist, which is a tool intended to guide clinicians in the use of best practices during central line insertions. | Labeling mix-ups | On November 16, 2017, an Urgent Medical Device Notification titled "ARROW PSI and Pressure Injectable Multi-Lumen CVC Kits" was issued to customers informing them of the issue, and urging all to take the following actions:1. Place a copy of this notification and the enclosed Central Line Insertion Checklist with each unit of affected product currently in your inventory.2. Complete the enclosed Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or e-mail to recalls@teleflex.com.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 18124 | 18,124 total | Nationwide Distribution | Open, Classified |  |  | 18124
Z-0908-2018 | 79166 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Endo GIA" Radial Reload with Tri-Staple" Technology |  |  | September 15, 2017 | 2017 | Covidien LLC | The device cartridge disengaged during use due to manufacturing error. | Process control | Medtronic sent an "URGENT MEDICAL RECALL LETTER" dated September 13, 2017, was issued to customers titled "Covidien Endo GIA Black Radial Reload with Tri-Staple Technology" urging customers to quarantine and return unused product to recalling firm. Questions or concerns can be directed to: feedback.customerservice@Covidien.com. For further questions, please call (203) 492-5000. |  | 163 | 163 | Internationally, including Japan. No USA Customers | Terminated | December 19, 2018 | 460 | 163
Z-0920-2018 | 79219 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Signature Ceramic Femoral Head. Model Number: 111-152-621. |  |  | February 02, 2018 | 2018 | Signature Orthopedics Pty Limited | This recall has been identified due to the batch being mislabeled, with the packagedcomponent's size not corresponding to the label. Therefore, the product intended to be usedfor the procedure may not be included in the packaging resulting in surgical delay asreplacement product is retrieved. | Labeling mix-ups | Signature Orthopaedics notified customer(s) of the recall by sending an Urgent Medical Device Recall Letter dated February 2, 2018 via email. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the Customer Acknowledgement Form and firm's Effectiveness Check Questionnaire. |  | 10 | 10 | Worldwide Distribution - U.S. Nationwide and the country of Australia. | Terminated | August 08, 2018 | 187 | 10
Z-0919-2018 | 79206 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Sirus Drill 13mm diameter trauma instrument used for preparation of bone in Trauma, Hip and Shoulder surgeries.; Model Number: 02.00020.040 |  |  | October 03, 2017 | 2017 | Zimmer Biomet, Inc. | The Sirus Drill and Flexible Shaft instruments may not be adequately cleaned when utilizing the standard cleaning instructions potentially resulting in infection and subsequent complications. | Labeling design | Review the notification and ensure affected personnel are aware of contents. Assist Zimmer Biomet sales representative in locating and quarantining affected product. Sales representative will remove product from facility. Complete certificate of acknowledgement and return to Zimmer Biomet. Distributors are asked to locate and quarantine all affected product in inventory and from affected hospitals in their territory. |  | 127 | 127 | Distributed in 31 states: AL, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NJ, NM, NY, OH, OR, PA, TN, TX, VA, VT, WA, WI and the District of Columbia. | Terminated | August 20, 2018 | 321 | 584
Z-0916-2018 | 79087 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Stryker Sustainability Solutions Reprocessed BW Lasso 2515 NAV eco Variable Diagnostic EP Catheters, Ref D134301, Sterile, Rx. |  |  | January 02, 2018 | 2018 | Stryker Sustainability Solutions | An EEPROM chip error code may occur when the catheters are used with CARTO EP Navigation Systems. | Device Design | Stryker sent an Urgent Medical Device Recall letter dated January 2, 2018, explains the reason for recall, risk to health, and instructs them to discontinue use of the affected products. The customer is to check their inventory for the affected serial numbers provided in the attachment and complete the recall effectiveness check form to indicate the amount of product in inventory. Return instructions are provided. For further questions, please call (888) 763-8803. |  | 170 | 170 devices | USA (nationwide) Distribution was made to medical facilities in AZ, CA, CT, FL, GA, IL, KS, LA, MN, MO, OH, PA, TX, VT, WA, and Wi. Foreign distribution was made to Canada. | Terminated | May 02, 2018 | 120 | 178
Z-0910-2018 | 79130 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas 8000 Modular Analyzer Series; Part Numbers: 1. 05964067001 cobas 8000 c 502 module |  |  | July 18, 2017 | 2017 | Roche Diagnostics Corporation | A warning is being added to the cobasÂż 6000 analyzer and the cobasÂż 8000 modular analyzer series Operators Manual to prevent operator injury during maintenance. | Device Design | Actions Taken by Roche Diagnostics: A warning will be added to the cobasÂż 6000 analyzer series Operators Manual and the cobasÂż 8000 modular analyzer series Operators Manual. Actions to be Taken by the Operator: Operators are advised not to overtighten the thumb screws when replacing the USM Cover 3 after maintenance. When removing the screws for maintenance, avoid touching the edge of the USM Cover 3. Enclosure: Fax back form 7727-00-0717Actions Required: " Do not overtighten the thumb screws when replacing the USM Cover 3 after maintenance. When removing the screws for maintenance, avoid touching the edge of the USM Cover 3. " Complete the attached fax form and fax it to 1-888-276-5901. " File or post this Urgent Medical Device Correction (UMDC) in an appropriate location within your laboratory, for future reference. |  | 245 | 245 | Distributed in all 50 states: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MN, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, the District of Columbia, and Puerto Rico. | Terminated | July 13, 2018 | 360 | 3678
Z-0900-2018 | 79235 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | BIO CAL(R) Temperature Controller, Models: (a) 370, (b) 370I |  |  | February 08, 2018 | 2018 | Medtronic Perfusion Systems | Medtronic has been unable to develop a cleaning protocol to satisfy current industry concerns and expectations. As a result, an updated cleaning protocol will not be developed by Medtronic and it has been determined that the best course of action is to request BIO CAL users to discontinue use and dispose of BIO CAL devices. | Device Design | Medtronic sent an Urgent Medical Device Recall dated February 2018 via UPS 2-Day.to their customers. The letter identified the affected product, problem and actions to be taken. Medtronic recommended that the affected device no longer be used for clinical procedures. Customers are asked to dispose of any devices in their possession per their normal equipment obsolescence procedures. Customer are also instructed to not resell any devices for clinical use. For questions contact your Medtronic Sales Representative. |  | 1980 | 1980 units | Worldwide Distribution | Open, Classified |  |  | 1980
Z-0912-2018 | 79155 | No | DIF | Clinical toxicology control material. | Toxicology | Toxicology | 510(K) Exempt | VITROS Chemistry Products DAT Performance Verifier II in vitro diagnostic |  |  | September 22, 2017 | 2017 | Ortho-Clinical Diagnostics | This product is not meeting the current storage labeling for unopened VITROS DAT Performance Verifiers causing customers to obtain lower than expected VITROS BENZ results when using this control. | Under Investigation by firm | On September 18, 2017 a letter was sent to customers who received the BENZ reagent within the past 18 months of affected lots. The letter stated that customers may continue use of the product if QC results using VITROS DAT Performance Verifiers are within established ranges for each assay. If QC BENZ results are below established ranges, the customer should discontinue use of the product. The letter requests customers post the notification by each system that processes the VITROS BENZ Reagent. If the product was transferred outside of the facility, to notify those customers as well. Affected lots will be replaced in 6-8 weeks. Questions or concerns can be directed to 585-453-3452 |  | 323 | 323 | Worldwide distribution: US Nationwide (including Puerto Rico), Bermuda, Australia, Brazil, Canada, Chile, Colombia, India, Japan, Mexico, Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain, Poland, and Russia. | Open, Classified |  |  | 1150
Z-0924-2018 | 79181 | No | MIU | Trypanosoma spp. serological reagents. | Microbiology | Microbiology | 510(k) | Wiener Lab CHAGATEST ELISA RECOMBINANTE V. 3.0 |  |  | November 10, 2017 | 2017 | Wiener Laboratories | Internal program testing shows that product may not reach the indicated expiration date. | Nonconforming Material/Component | On about November 10, 2017, customers were notified via email of the recall. Instructions include to examine inventory and cease distribution of the device, make arrangements to return affected devices, acknowledge receipt of the recall notification, and notify customers if further distributed. |  | 1563 | 1,563 units | Worldwide Distribution - US Distribution to the states of : Florida and Washington DC. , and to the countries of : Argentina, Bolivia, Brasil, and Costa Rica. | Open, Classified |  |  | 1563
Z-0932-2018 | 79341 | [6 empty] | 241-McKesson Premium Skin Staple Remover Kit, Item #100123, McKesson I.V. Start Kit, 100124-McKesson Suture Removal Kit, and 82-15 Cypress Medical Products Skin Staple Remover Kit. McKesson Item # 241, 100123, and 100124, Cypress Item # 82-15 |  |  | November 17, 2017 | 2017 | Cypress Medical Products LLC | Product not meeting the iodine assay level requirements to support 36 month expiration dating. | Nonconforming Material/Component | The firm, Cypress Medical Products, sent an "URGENT VOLUNTARY MEDICAL DEVICE RECALL" letter dated 11/16/2017 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to immediately discontinue use of any product matching the affect items; If you have product affected by this notice, a recall label must be applied to each affected convenience kit. Please contact Cypress Medical Products, LLC customer service via email at cypress.customerservice@cypressmed.com to request the number of labels required. Upon opening the kit during procedure time, please dispose of the affected Aplicare Povidone Iodine Prep Pad in accordance with your institutions policies and procedures. If you have further distributed the product, please notify your customers and provide the customer with a copy of this notification. Complete and return the Reply form via email to: Cypress Medical Products LLC, Attn: Cypress Customer Service at: cypress.customerservice@cypressmed.com, even if you have no affected product.To ensure product availability, new orders for the affected convenience kits will be shipped with the affixed recall labels until new product is made available from the manufacturer.If you have any questions concerning this recall, please contact Cypress Customer Service at the email: cypress.customerservice@cypressmed.com. |  | 12206 | 12,206 cases | US Distribution to states of: CO, FL, GA, KY, MI, MO, NC, NH, OH, TN, TX, WA, AND WI. | Open, Classified |  |  | 12206
Z-0925-2018 | 79407 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | QuikClot TraumaPad, Part# 460 a topical dressing for local management of bleeding wounds such as cuts, lacerations, and abrasions. It may also be used for temporary treatment of severely bleeding wounds such as surgical wounds (operative, postoperative, dermatological, etc.) and traumatic injuries |  |  | November 17, 2017 | 2017 | Z-Medica, LLC | A customer complained that one of the pouches in a box of 10 had "QuikClot Roll" pouch instead of "QuikClot TraumaPad" pouch. The carton as well as the remaining 9 pouches were properly identified as QuikClot Trauma Pad. The customer alleged that when the pouch was opened, it contained a QuikClot TraumaPad. All 10 pouches in the box had the correct part number, lot number, and expiration date for the prod uct (QuikClot TraumaPad). | Mixed-up of materials/components | On November 22, 2017 an URGENT ADVISORY NOTICE was issued to customers stating the following:The QuikClot TraumaPad is packaged with 10 pouches in a carton. This lot may contain some pouches that incorrectly identify the product as QuikClot Roll. All pouches (identified as TraumaPad or Roll) are identified with the correct Part N umber (REF 460), Lot Number, and Expiration Date for QuikClot TraumaPad, and contain the QuikClot TraumaPad product. Please note that the labeling on the cartons is also correct and we believe that this only impacts a small num ber of pouches.In addition, QuikClot TraumaPad and QuikCot Roll have the identical ingredients, chemical composition, indications for use, packaging material and surface area. The only difference is that the QuikClot TraumaPad is a 12 inch x 12 inch, 3-ply pad and the QuikClot Roll is a 3 inch x 4 yard roll. Both products are equally effective in being used to stop bleeding.This was the result of an isolated operator error. Please note that actions have been taken to prevent this issue in the future and further corrective actions are under review.Ifyou have any inventory of this lot number, you may choose to use the prod uct as the correct prod uct is packaged into each pouch (only pouch labeling is incorrect). Ifyou chose to destroy the prod uct or have any questions regardi ng this matter, please contact us at the email and telephone number below. |  | 4890 | 4890 | Nationally | Terminated | June 04, 2018 | 199 | 4890
Z-0930-2018 | 79227 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, graphite block 8 |  |  | October 17, 2017 | 2017 | Ion Beam Applications S.A. | IBA identified incorrect screw holes depth in graphite block 8 with respect to the degrader | Component design/selection | On July 17, 2017 An Urgent Field Safety Notice was issued to customers with the following user action: In case a patient treatment plan requires the use of the block 8, it is recommended to perform a full range verification during the patient treatment plan QA check. After the replacement of the graphite block 8, it is recommended to verify there is no effect on the Bragg peak. |  | 3 | 3 | WA | Terminated |  |  | 3
Z-0927-2018 | 79272 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | O-armÂż O2 interventional fluoroscopic x-ray system; Catalog Number: B1-700-02000 |  |  | September 30, 2017 | 2017 | Medtronic Navigation, Inc.-Littleton | During servicing of certain internal components of the O-Arm O2 Surgical Imaging System, it is possible that the high-power cable connector may come in contact with surrounding metal during the disconnecting/reconnecting process which may result in an electrical short and a shock to the field service technician. | Employee error | Medtronic sent an Medical Device Safety Notice dated September 29, 2017, reminding them to follow servicing procedures. In addition, Medtronic is developing a retrofit service kit that will be installed during the next service visit for each O-arm O2 system, no later than 15 months from service kit availability. The user can continue to use the system prior to installation of the service kit.Medtronic will also implement the same changes in the manufacturing process for newly manufactured O-arm O2 systems. For further questions, please call (978) 698-6000. |  | 449 | 449 | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and to the states of : AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, SD, TN, TX, UT, VA, WA, WI, WV., and to the countries of : Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Luxembourg, Martinique, Mexico, Netherlands, New Zealand, Poland, Portugal, Qatar, Romania, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 449
Z-0926-2018 | 79150 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | i-STAT DE handheld data processing module for clinical use, Software Version 2.8, List Number:08K46-01115200 (UDI: 00054749001255) |  |  | September 12, 2017 | 2017 | Abbott Point Of Care Inc. | Issues resulting from upgrade to software version 2.8: (1) Location, operator, stored patient lists will not update, and (2) Customized Reference Ranges, Action Ranges, and Custom Reportable Ranges are reset to factory default values. No erroneous results are generated as a result of this issue. | Device Design | The following actions are recommended if you have already upgraded to version 2.8:1. If you are using Location, Operator, or Patient lists.a. Delete the following files:i. C:\Istat32\LocInstData\LocInstData.txtii. C:\Istat32\OpData\OpData.txtiii. C:\Istat32\PatData\PatData.txtiv. C:\Istat32\Log\iSTATDMI Errors.txtv. All files named C:\Istat32\Log\DeCustLogXX.txt; where XX is any two digits.b. Stop and restart the DECOMSRV service, which in turn restarts the DEDPSRV service.2. If you are using customized Reference Ranges, Action Ranges, and Custom Reportable Ranges:a. Open DE Customization Workspace.b. For each profile, select preferences by clicking the button with the profile name.c. Navigate to the Results tab.d. Verify the Action and Custom Reportable Ranges. Click Edit to make changes.e. If changes are made, click the OK button to save.If you have not upgraded to i-STAT/DE version 2.8, make arrangements with your data management vendor to update to the newer version of i-STAT/DE at your earliest opportunity. DO NOT UPGRADE without coordinating with your data management vendor so that continuity of your i-STAT testing is not impacted. How should you verify what version of i-STAT/DE is installed on a computer?Using the Program and Features app in Windows is the best way to determine the version of i-STAT/DE installed on your system. To do so, follow these steps:1. Open Control Panel and the Program and Features applet.2. Locate the program named i-STAT/DE Version 2.8 in the list.3. The version column will show 2.8 for the old version and 2.8.0.1 for the new version.If you determine the version 2.8.0.1 is not installed, make arrangements with your data management vendor to update to the latest validated version compatible with your data management system.Enclosed is an updated i-STAT/DE version 2.8, version qualifier, 2.8.0.1 CD-ROM and a Business Reply Card (BRC). Please complete and return the Business Reply c |  | 894 | 894 | Distributed in 48 States: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, the District of Columbia, Australia, Italy, Sweden, and Canada. | Terminated | August 23, 2018 | 345 | 894
Z-0931-2018 | 79228 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Pencil Beam Scanning version PTS-6.4.11.X |  |  | October 19, 2017 | 2017 | Ion Beam Applications S.A. | Range for the same energy could be different depending if the layer isthe first layer of the irradiation plan or not. | Device Design | IBA sent an Urgent Field Safety Notice dated October 19, 2017 , issued to customers stating the following: IBA recommends performing a check to detect the consequence of this potential power supply regulation issue on a regular basis (IBA recommends Monthly), see Annex 1 for details.IBA is developing a corrective action in order to reduce the hysteresiseffect and detect magnet current overshoot.This solution will be deployed to all systems using JEMA analog powersupply. This solution will be validated and deployed by April 2018. For further questions, please call (215) 972-7985 |  | 3 | 3 | US Distribution to the state of :FL | Open, Classified |  |  | 3
Z-0928-2018 | 79275 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | O-arm 1000 mobile image-intensified fluoroscopic x-ray system; (1) Model: BASE OARM Bl70000028100 SYS 100V, Product Number: Bl70000028100 (UDI: 00643169354418); (2) Model: BASE OARM Bl70000028120 SYS 120V, Product Number: Bl70000028120 (UDI: 00643169353411); (3) Model: BASE OARM Bl70000028120R SYS 120V RWK, Product Number: Bl70000028120R (UDI: 00643169353459); (4) Model: BASE OARM Bl70000028230 SYS 230V, Product Number: Bl70000028230 (UDI: 00643169353992); (5) BASE OARM Bl70000028230R SYS 230V RWK, Product Number: Bl70000028230R (UDI: 00643169354081) |  |  | January 25, 2018 | 2018 | Medtronic Navigation, Inc.-Littleton | Software update ("SW Update 3.2.1") is being implemented to address the following issues: Inability to power the system after shutdown, System stays in standalone mode, Image reconstruction, System/Pendant bootup, Dose display/report, Gantry motion, Network communication, System shutdown, and Early termination of 3D spin. | Software design | Medtronic sent an Urgent Medical Device Correction letter to their customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that a local service representative will contact you to install the update. For questions contact Medtronic Technical Support at 1-800-595-9709. |  | 253 | 253 | Worldwide Distribution - US Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, LA, MA, MI, MN, MO, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WY, and the countries of Puerto Rico. Distributed in Algeria, Australia, Chile, China, Czech Republic, Finland, France, Germany, Greece, India, Israel, Jordan, Lebanon, Sri Lanka, New Zealand, Nicaragua, Poland, Portugal, Russia, Saudi Arabia, Singapore, Spain, Switzerland, Thailand, United Arab Emirates, United Kingdom. | Open, Classified |  |  | 253
Z-0929-2018 | 79183 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | DIASTAT anti-Nuclear antibody (ANA) / DIASTAT ANA ELISA. Catalog Number: FANA200. |  |  | October 15, 2015 | 2015 | Euro Diagnostica AB | Complaint investigation concluded that although product qualityrequirements stated in the instructions for use were met at lot release, the high background in the ELISA plate, thestrong IgM conjugate, and the low reference control led to an increased risk of equivocal and/or false positive sampletest results for the two FANA200 kit lots SS 1500 and SS 2009. | Nonconforming Material/Component | Euro Diagnostica sent a Customer notifications letter to affected customers via email on 10/15/2015. The letter identified the affected product problem and actions to be taken. For question contact Euro Diagnostica . |  | 1090 | 1090 | Worldwide Distribution and US. Nationwide | Open, Classified |  |  | 1090
Z-0944-2018 | 79186 | No | PHL | Powered lower extremity exoskeleton. | Physical Medicine | Neurology | 510(k) | ReWalk Personal 6.0. Catalog number: 50-20-0004. |  |  | August 14, 2017 | 2017 | Argo Medical Technologies Ltd | Firm received complaints for ReWalk Personal 6.0 of an injury to tibia and fibula. | Other | Customers were notified via letter on 08/14/2017. |  | 127 | 127 | US and foreign. | Open, Classified |  |  | 127
Z-0939-2018 | 79397 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trabecular MetalÂż Primary Hip Prosthesis; Part Numbers: 1. 00-7864-013-00 (UDI: (01) 00889024139251 (17) 250731 (10) 63703649), and 2. 00-7864-013-20 (UDI: (01) 00889024139268 (17) 250731 (10) 63703681) |  |  | January 03, 2018 | 2018 | Zimmer Biomet, Inc. | Two lots of hip implants are mislabeled. | Labeling mix-ups | DISTRIBUTORS: 1. Review this notification and notify affected team members of the contents. 2. Immediately locate and quarantine affected product. 3. Immediately return all affected product from your distributorship and from affected hospitals within your territory. a. Complete Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days. Mark RECALL on the outside of the returned cartons. 4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com RISK MANAGERS: 1. Review this notification and notify affected personnel of the contents. 2. Quarantine all affected product. Your sales representative will remove the affected product from your facility. 3. Complete Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 24 | 24 | Distributed in 2 US states: LA and NY. Distributed in Canada, China, Japan, and Switzerland. | Open, Classified |  |  | 24
Z-0945-2018 | 79051 | orthosis, limb brace | IQI | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | FUTURO (TM) SPORT, QUICK STRAP ANKLE SUPPORT, REF 47736GEN, UPC 0 51131 19692 6 |  |  | January 31, 2018 | 2018 | 3M Company - Health Care Business | 3M discovered that the Futuro" Quick Strap Ankle Brace (Cat. #47736EN) and ACE" Brand Deluxe Ankle brace (Cat. #207736) do not have the correct labeling. The products contain natural latex rubber, but do not include the appropriate caution statement. | Error in labeling | The firm initiated their recall by letter on 01/30/2018. The letter stated the following:"Actions required:1.Immediately discontinue sale of Futuro" Quick Strap Ankle Brace.2.Isolate or quarantine any stock of the Futuro" Quick Strap Ankle Brace.3.Call 1-866-746-2524, option #2, Monday to Friday, 8:00 am  4:30 pm Central time, to arrange for the return of the Futuro" Quick Strap Ankle Brace, Cat# 47736EN.4.If you have further distributed the Futuro" Quick Strap Ankle Brace, please notify any of your customers at once. Please include a copy of this letter in your communications.We apologize for any inconvenience this may cause you. Should you have any questions regarding the product and return process, please feel free to contact 3M Customer Service at 1-866-746-2524, option 2, Monday to Friday, 8:00 am 4:30 pm Central Time." |  | 158179 | 158,179 units | US and Australia, Jordan, Republic Of Korea, Lebanon, New Zealand, Oman, Qatar, Taiwan, United Arab Emirates | Open, Classified |  |  | 320527
Z-0937-2018 | 79035 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid Cardiac Control CQ5053 |  |  | August 08, 2017 | 2017 | Randox Laboratories, Limited | Randox has confirmed that Troponin T and Troponin I in Liquid Cardiac quality control CQ5053 lot 4069CK does not meet the shelf-life claim in the product labelling. As the control may now fall outside the provided control ranges the product is to be recalled from the field. | Unknown/Undetermined by firm | A recall notice was issued to customers on August 1, 2017 instructing them to discontinue use of product, discard remaining product for reimbursement and provide documentation of destruction. |  | 450 | 450 kits | Nationally | Terminated | September 05, 2018 | 393 | 450
Z-0950-2018 | 79195 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Amplatz Ultra Stiff Wire Guide |  |  | October 13, 2017 | 2017 | Cook Inc. | Label does not state that the product is heparin-coated. | Labeling Change Control | On October 31, 2017 an URGENT MEDICAL DEVICE CORRECTION was issued to customers requesting the following, including a corrective action:Step 1:Determine if your product isincluded in the affected lotranges by comparing your lotnumber to the listingprovided." If your product is notaffected, complete theAcknowledgement andReceipt Form." If your product is affected,proceed to Step 2.Step 2:Remove a Heparin Coated label from the label sheetprovided and apply to the topright corner of your affectedunit.o If you need additionallabels, please contactStericycle at855.215.4967.If you would like assistance,representatives are available to supportyou. To request assistance, pleasecontact Stericycle at 855.215.4967 |  | 12462 | 12,462 | Nationally | Open, Classified |  |  | 1424570
Z-0941-2018 | 79038 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Accolade and Restoration hip product; Catalog Number: 6021-0030, 6276-1-127, 6276-5-216, 6020-2530, 6020-4535, 6021-0740, 6021-0230, and 6021-4535 |  |  | July 21, 2017 | 2017 | Howmedica Osteonics Corp. | Inner and outer sterile barriers not fully sealed. | Process control | Recall notices were sent to customers with a request to return all affected product. |  | 255 | 255 | Distributed in ten (10) states: AR, DE, FL, IL, IN, MI, NJ, OH, TN, and TX; and Canada, Sweden, United Kingdom, France, Spain, Italy, Japan, Romania, Colombia, Netherlands, and Australia. | Open, Classified |  |  | 255
Z-0935-2018 | 79217 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Difco" Agar Noble, 500g bottle |  |  | October 03, 2017 | 2017 | Becton Dickinson &amp; Co. | HighPuragar was mislabeled and inadvertently sold as Noble Agar. | Mixed-up of materials/components | On October 3, 2017 an URGENT MEDICAL DEVICE PRODUCT RECA:LL letter was mailed to customers informing them of the recall and requesting that their inventory be checked for product, which should be quarantined and then thrown away. Also, to share notification of recall to any other customers that may have received product. Firm will replace any discarded inventory. Questions or concerns can be directed to: 1-800-638-8663. |  | 216 | 216 bottles | USA (nationwide) Distribution to the states of and including Puerto Rico : AL, AZ, CA, FL, GA, IA, IL, IN, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NY, OH, OR, PA,TN, TX, UT, VA, WA | Terminated | October 09, 2018 | 371 | 216
Z-0938-2018 | 79230 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Proton Therapy System a medical device designed to produce and delivers a proton beam for the treatment of patients with localized tumors and other conditions susceptible to treatment by radiation. |  |  | November 17, 2017 | 2017 | Ion Beam Applications S.A. | There is a risk of misalignment of the patient when using treatmentplans with multiple isocenters. | Device Design | On November 17, 2017 an Urgent Field Safety Notice was issued to customers informing users that when creating or importing a treatment plan with multiple isocenters, the user has to enter Patient Positioning System positions manually. |  | 37 | 37 | Nationally | Open, Classified |  |  | 37
Z-0933-2018 | 79318 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Vygon IV Start Convenience Kits; Product Codes AMS-623T and AMS-626TGSLF. |  |  | October 03, 2017 | 2017 | Vygon MFG, Inc., dba/ Churchill Medical Systems, Inc. | The Iodine Prep Pads within the Vygon IV Start Convenience Kits, had stability samples that were not meeting iodine assay level requirements to support the 36 month expiration date. All kits containing the affected pads were recalled. | Under Investigation by firm | Vygon sent an Urgent Medical Device Recall letter dated October 3, 2017. Customers requested to check stock and quarantine affected devices and to complete Recall Acknowledgement and Inventory Return Form. Customers instructed to fax or email the form to Vygon including customer contact info. Vygon will contact the customer and issue a Return Merchandise Authorization Number. Distributors instructed to communicate recall to customers who received affected product and provide them copy of the recall notification. Returned products were received in Lansdale, PA warehouse location and placed into quarantine. For further questions, please call Customer Service at (800) 473-5414. |  | 21350 | 21350 | US Distribution. | Open, Classified |  |  | 21350
Z-0942-2018 | 79209 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | INFINITY DUAL HEMO MCable Pod |  |  | October 04, 2017 | 2017 | Draegar Medical Systems, Inc. | Dual Hemo MCable Pods with RevisionIndex (RI) 15 and 16, which are used with the DrÂżger Infinity Acute Care System, maypermit liquid ingress inconsistent with their IPX4 rating, which could potentially result in anincorrect measured value of invasive blood pressure. | Nonconforming Material/Component | In September of 2017 an Urgent medical Device Recall letter was issued to customers informing them of the recall and to request that inventory be checked to see if the product subject in the recall is on the premises. To do this, the customer will check the label for the revision index number which is located on the upper middle portion of label. RI number 15 and 16 are included in the recall. No other numbers are subject to recall. Affected product will receive a sleeve from the company to correct the issue. Questions or concerns can be directed to 1-800-437-2437 (Press 1, then press 32349) or contact Customer Care at 1-800-437-2437 (Press 4, then press 4 again) |  | 2760 | 2,760 | Nationally | Terminated | June 15, 2018 | 254 | 2760
Z-0967-2018 | 79218 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | SYSTEM 26. Part Numbers: |  |  | October 10, 2017 | 2017 | ORTHO SOLUTIONS, LTD. | Following the identification of corrosion residue found inside the cannulated devices in Forefoot andMidfoot trays Orhto Solutions Group has taken the decision to recall ALL x106 System 26 Forefoot and MidfootImplants and Instruments tray from the US market. | Process control | Ortho Solutions notified customers of the recall via letter 10/10/2017. The recall notification letter was accompanied by instructions to Distributors and a Product Removal Form. |  | 120 | 120 | U.S. | Terminated | August 31, 2018 | 325 | 120
Z-0934-2018 | 79231 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Uniform and Pencil Beam Scanning |  |  | December 05, 2017 | 2017 | Ion Beam Applications S.A. | An accessory (range shifter, ridge filter&amp;) can be improperly inserted insnout 300x400 with one rail only and the proton therapy system can stilloperate. | Device Design | Ion Beam Applications (IBA) sent an Urgent Field Safety Notice dated December 5, 2017, to customers informing them of the issue and request that all users verify that the accessory is correctly inserted in snout 300x400 with both rails in the guide. A solution to this issue including a hardware modification will be deployed on or before March 31, 2018. For further questions, please call (215) 972-7985. |  | 14 | 14 | US Distribution to the states of : MA, FL, VA, PA, OK, TN | Open, Classified |  |  | 14
Z-0943-2018 | 79229 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Pencil Beam Scanning, PTS-9.4.X, PTS-10.0.1.1 |  |  | November 01, 2017 | 2017 | Ion Beam Applications S.A. | The internal configuration of the electrometers and the Real-Timecontrol boards of the Dekimo Scanning Controller is not checked beforeeach patient treatment. | Device Design | The firm, IBA, sent an "Urgent Field Safety Notice" dated 10/31/2017 to customers on 10/31/2017. The notice described the product, problem and actions to be taken. stating the following: There is no specific recommendation for the user. IBA will provide an Internal User notice to the IBA site team, requesting the operator to select the Beam Delivery Point CGTR, by clicking on the BDP button, before each patient treatment preparation (see Annex 1 fordetails). This will begin the check of the internal configuration of allthe elements of the Dekimo Scanning Controller. An updated system will be deployed on February 28, 2017.If you have any questions, contact Customer Complaints &amp; Vigilance Director at Vigilance@iba-group.com or +32 10 203 787 or helpdesk at +32 2 507 20 81 (available 24/7). |  | 3 | 3 | Worldwide Distribution: US (nationwide) to: MI only and country of: France. | Open, Classified |  |  | 3
Z-0970-2018 | 79076 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Immunoassay Premium Plus Controls (3x4x5mL and 12x5mL packages) |  |  | August 11, 2017 | 2017 | Randox Laboratories | The ACTH in Immunoassay Premium and Premium Plus quality control does not meet the quoted reconstituted stability claim in the product IFU. | Device Design | On July 24, 2017 an Urgent notice was issued instructing customers to place a copy a notice in all kits remaining in stock. This notifies customers that they should amend the claims for open stability in the instructions for use, to state use immediately after the 30 minute reconstitution period. Questions or concerns can be emailed to technical.services@randox.com |  | 199 | 199 Kits | Distribution US nationwide. | Completed |  |  | 316
Z-0973-2018 | 79190 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LCSU 4, Laerdal Compact Suction Unit |  |  | August 31, 2017 | 2017 | Laerdal Medical Corporation | A shipment of LCSU 4 suction units and canisters attempted importation into the United States but was refused by FDA for lack of necessary approvals with CDRH by the foreign firms. | No Marketing Application | On September 5, 2017, a letter was mailed to customers notifying them that they received one of the units subject to recall. The letter requested that the customer return the affected item for a replacement. Questions or concerns can be directed to 1-877-523-7325 or email: customerservice@laerdal.com , referencing the following number: 2017-R-05. |  | 204 | 204 | Nationally | Terminated | June 13, 2018 | 286 | 204
Z-0974-2018 | 79030 | No | OLC | Surgical mesh. | General &amp; Plastic Surgery | Orthopedic | 510(k) | CMI (Collagen Meniscus Implant) device, Ivy Sports Medicine s collagen-based meniscus implant |  |  | January 25, 2018 | 2018 | Stryker Corporation | The recalled products were shipped without the required temperature control packaging, therefore could potentially have been exposed to elevated temperatures during transit. | Packaging | The international Stryker site was notified by email on 1/25/2018. Customers are instructed to: 1. Inform individuals within your organization who need to be aware of this device removal. 2. Review Part Numbers (4600, 4601, 4607, and 4612) for affected lot numbers. Please determine if you have theaffected product in stock. Response is required. 3. If no product is found, notify your local Stryker office.4. If you do have product, segregate the product and call your local Stryker office to arrange for product return andissuance of credit. |  | 155 | 155 devices | Netherlands | Open, Classified |  |  | 155
Z-0971-2018 | 79177 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter SIGMA Spectrum Infusion Pump with Master Drug Library (Version 6) (GTIN 00085412091570); Model Numbers: 35700BAX and 35700ABB. |  |  | October 06, 2017 | 2017 | Baxter Healthcare Corporation | Certain infusion pumps may exhibit excessive wear of the pumping mechanism caused by a lack of lubrication, this could affect the occlusion of the tubing and result in over- or under-infusion. | Nonconforming Material/Component | An Urgent Medical Device Correction, dated October 6, 2017, was sent to customers. The letter described the problem, identified the affected device, discussed the hazard involved, and action to be taken. Customers were asked to locate the affected products at their facility. Customers should continue to follow the instructions in document 41394 Rev. A, PREVENTIVE MAINTENANCE, for their entire fleet of pumps. Baxter will inspect affected pumps. To arrange for return of devices that have failed the flow rate accuracy test, contact Baxter Healthcare Medina. If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please notify your customers of this Urgent Medical Device Correction. If there are additional questions, customers can contact their Baxter Sales Rep or Baxter Healthcare Medina at 800-356-3454. |  | 173375 | 173,375 | Worldwide Distribution - US, including the District of Columbia, Puerto Rico, and Canada. | Open, Classified |  |  | 173375
Z-0972-2018 | 79243 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Vial2BagDC, Product Number 6070112. |  |  | October 06, 2017 | 2017 | Medimop Medical Projects Ltd. | Medimop is recalling certain lots of this product Vial2Bag DC due to the potential of the IV Port separating from the Via12BagDC body during the drug reconstitution I IV Set attachment process. | Process control | The recalling firm's sent an Urgent Medical Device Recall Letter dated October 13, 2017. was supplied to the consignee which contains the instructions for the recall process to the consignee. The notification letter will be sent to the Distributor via email.The notification letter instructs the distributor to notify the Customer/Users via the Medimop consignee recall letter. For further questions, please call (610) 594-4353. |  | 453600 | 453,600 units | US Distribution | Open, Classified |  |  | 453600
Z-0968-2018 | 79041 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD VacutainerÂż K2 EDTA (K2E) 3.6mg Blood Collection Tube (13 x 75 mm x 2.0 mL). Catalog Number: 367841. |  |  | August 29, 2017 | 2017 | Becton Dickinson &amp; Company | A limited portion of the lot was manufactured with less than the required amount of K2EDTA additive, an anticoagulant, which may cause erroneous results that could lead to specimen recollection and potential delay of treatment or misdiagnosis/mismanagement of treatment and serious complications, such as failure to detect thrombocytosis or postponing surgery. | Under Investigation by firm | In August of 2017 All US direct consignees were contacted via e-mail and/or UPS mail. An Urgent Product Removal Recall letter was issued, identifying the affected device and the reason for the recall. It also requested customers to review their inventory and quarantine any affected devices, which will be returned to the firm for replacement. The Customer Recall Response Form should be completed and returned regardless of whether there is any affected devices on hand. For questions or concerns please contact Customer Service at 1-855-215-4992. |  | 8494000 | 8,494,000 | Nationwide Distribution | Open, Classified |  |  | 8494000
Z-0977-2018 | 79204 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | OneStep CPR Complete, Adult Multi Function Electrodes, Part Numbers: 8900 0214 01 (case of 8 single electrodes, UDI: 10847946016279) and 8900 0224 01 (single electrode, UDI: 00847946016272) |  |  | August 04, 2017 | 2017 | ZOLL Medical Corporation | A portion of one lot of ZOLL OneStep Complete Electrodes for adults will cause the defibrillator to deliver pediatric energy levels instead of adult energy levels when used with the ZOLL R Series Defibrillator. Defective electrodes will display PEDIATRIC PADS IN USE on the R Series defibrillator. Using the affected product with the M Series Defibrillator poses no risk. | Process control | Zoll sent an Urgent Device Correction letter dated August 4, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. The following actions should be taken until the affected electrodes are removed from use: 1. Alert staff to the possibility of this occurrence. 2. Instruct staff to inspect the lot code on the electrode package and remove affected lot 2517A from use. 3. Inform users that if they encounter this situation in the process of defibrillating an adult patient, they should set the defibrillator energy level manually. 4. Complete the attached form and return to ZOLL for replacement electrodes. ZOLL will be replacing the effected lot of the electrodes with properly configured electrodes. For questions contact technical support at 1-800-348-9011 or 1-978-421-9460. |  | 3400 | 3400 | Worldwide Distribution - US Nationwide in the states of AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and in the countries of Canada. | Terminated | August 31, 2018 | 392 | 3400
Z-0975-2018 | 79360 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | IMRIS ORT100 Table, Part Number 110470-000 |  |  | February 20, 2018 | 2018 | Deerfield Imaging, Inc. | The rotational locking mechanisms on the ORT 100/ T2X tables can become disengaged or unlocked while in use which could potentially result in unintended movement of the table. | Device Design | The firm disseminated the notices on 02/23/2018 and followed with emails on 02/26/2018. The notices stated:"To minimize the risk of injury, IMRIS is taking the following corrective actions:" ORT100 Table: The rotational locking mechanism shall be replaced in all ORT100 tables that do not contain the most current design." T2X Table: Tables will have their rotational lock inspected. If components are worn or performance is other than intended, the rotational lock assembly will be serviced or replaced entirely.IMRIS has identified that your facility has one or more of these OR Tables which require evaluation. Please confirm details on page 2 and return your response to IMRIS.IMRIS Customer Service will contact you to implement this corrective action.IMRIS recommends these preventive actions:" Per the ORT Operator Manuals the table should be tested daily in all standard positions and inspected for anomalies including the fit of the covers, back section, headrest, pendant connection and swivel release pedal free engagement" During usage, verify the table is locked by manually pushing the table to make sure it does not move rotationally." If the rotational lock is observed to be not functioning properly, please contact IMRIS Customer Service.The appropriate competent authorities have been notified of this corrective action. IMRIS is committed to patient and user safety, and apologizes for any inconvenience that this may cause. Please contact customer support via our service line at 1-866-475-0525 or Email: customersupport@i mris.com if you have any questions or concerns." |  | 15 | 15 tables | US, Canada, Australia, China, Qatar | Open, Classified |  |  | 31
Z-0978-2018 | 79191 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart XL+ Defibrillator/Monitor |  |  | February 07, 2018 | 2018 | Philips Electronics North America Corporation | Update XLÂż device software to version A.03. This includes enhancements to the Operational Checks, event logs, and troubleshooting messages that provide as complete of information as needed to users ondevice readiness. | Labeling mix-ups | Philips is initiating a correction to affected devices. A software upgrade is available free of charge for all units affected by the issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrades. Customers may continue to use XL+ prior to receiving the software upgrade. For further questions, please call (978) 659-3000. |  | 4315 | 4315 | Foreign Only - China | Open, Classified |  |  | 4315
Z-0983-2018 | 79377 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Alden Classic 38 Sphere Daily Wear Soft contact; Catalog Number: AOCLS38 |  |  | November 03, 2017 | 2017 | Alden Optical | Contact lenses lack sterility assurance. | Process design | Consignees were initially called and told over the phone not to further distribute the contact lenses. For two consignees whose patients could not be reached by phone, a letter was sent. The consignee letter states the specific product and lot should not be further distributed due to the lack of sterility assurance. A business reply card was provided with the letter and contacts should be returned to the manufacturer using a label provided by the firm. |  | 4 | 4 | Worldwide Distribution - US Nationwide in the states:of CA, FL, GA, ID, IL, IN, KS, KY, MD, MI, MO, MS, NC, NJ, NM, NV, NY, OH, OR, PA, TX, VA, WI, WY and the country of Canada. | Terminated | June 19, 2018 | 228 | 148
Z-0980-2018 | 79385 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Humeral Stem 55mm(Size 13) |  |  | December 08, 2017 | 2017 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a medical device recall for two lots of Comprehensive Humeral stems. The lots were commingled during the manufacturing process and were etched with incorrect lot numbers. The size difference between the two stem sizes is 1.1 inches. It is visually recognizable by the user that the box label does not match the contents of the box. | Under Investigation by firm | On December 13, 2017 an URGENT MEDICAL DEVICE RECALL letter was issued to customers notifying them of the mix-up and requesting that they review the notification, discuss with team members, see if any product is in inventory and if so, quarantine and discontinue use. Affected product should be returned to the recalling firm. Questions or concerns can be directed to customer service at 574-371-3071. |  | 12 | 12 | US Nationwide in the states of MN, NJ, CA, TX, ND | Terminated | September 04, 2018 | 270 | 12
Z-0992-2018 | 79242 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | DELTA CER FEM HD 32/0MM T1 |  |  | June 30, 2017 | 2017 | Zimmer Biomet, Inc. | Product may contain the presence of melted sterile packaging upon opening. | Packaging process control | On June 2, 2017 an URGENT Medical Recall Letter was issued to customers regarding affected product(s). The letter requests customers to review the content and notify affected personnel of the recall. Assist the sales representative in quarantining affected product, which will be removed and returned to recalling firm and then replaced. All customers who may have received product from distributors are to be notified. Questions or concerns can be directed to 1-574-372-4847. |  | 30 | 30 | US in the state of Texas | Terminated | April 13, 2018 | 287 | 80
Z-1239-2018 | 79400 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Coonrad/Morrey Elbow Implant System humeral assembly 4IN XSML, Item Number: 32-8105-027-04 (UDI: (01)0088902427 4006(17)220731(10)63694912) |  |  | January 24, 2018 | 2018 | Zimmer Biomet, Inc. | One lot of humeral stem implants do not include the humeral bushing components. | Under Investigation by firm | 1. Immediately locate and quarantine affected product in your inventory.2. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. Request a Return Authorization Number. Mark RECALL on the outside of the returned cartons. |  | 9 | 9 | Distributed in China and Taiwan. | Open, Classified |  |  | 9
Z-1007-2018 | 79319 | Yes | MPB | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MAHURKAR"* Elite 12 FR x 16cm Acute Dual Lumen Dialysis |  |  | October 14, 2017 | 2017 | Medtronic Minimally Invasive Therapies Group | Medtronic is voluntarily recalling one lot of itsMAHURKAR"* Elite 12 FR x 16cm Acute Dual Lumen Dialysis Catheter with Curved Extensions, IC Trays, Lot 1717900160, following notification from our supplier that one production lot of MAHURKAR"* Elite 12 FR x 16cm Acute Dual Lumen Dialysis Catheter with Curved Extensions, IC Trays was inadvertently shipped to Medtronic and distributed without being sterilized. | Process control | Medtronic began notifying customers via telephone on October14, 2017, and all consignees were notified via Federal Express orcertified mail on October 17, 2017. The letter informs customersthat they were shipped devices that were not sterile. Customersare requested to acknowledge receipt and understanding of theUrgent Field Corrective Action Notice. |  |  |  | U.S. | Terminated | June 15, 2018 | 244
Z-1041-2018 | 78706 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CR PRECOAT FEM COMP SIZE BMR |  |  | November 29, 2017 | 2017 | Zimmer Biomet, Inc. | The LDPE bag packaging for various highly polished hip and knee implants may adhere to the highly polished surface, leaving residue or material from the LDPE bag on the implant after it is removed from the bag. | Packaging | A similar recall was initiated in January 2016 to remove remaining inventory packaged in the old bag. Further evaluation identified additional lots packaged in the old bag that were not included in the January 2016 recall; therefore, resulting in this new recall for the additional lots packaged in the old bag.A firm, Zimmer Biomet, sent an "Urgent Medical Device Recall" notices dated November 29, 2017 and response forms to customers. The notice described the product, problem and actions to be taken. The Customers should take the following action:Risk Manager Responsibilities:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.Distributor Responsibilities:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.c. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.d. Mark RECALL on the outside of the returned cartons.4. Return the Additi |  | 13227 | 13,227 in total | Worldwide Distribution: US (nationwide) and countries of: Argentina, Australia, Brazil, Bolivia, Canada, Chile, China, Colombia, Ecuador, El Salvador, India, Japan, Malaysia, Mexico, Netherlands, New Zealand, Nicaragua, Panama, Seoul-Korea, Singapore, Taiwan, Thailand, and UAE. | Open, Classified |  |  | 13227
Z-1191-2018 | 79406 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer M/L Taper Hip Prosthesis, 7711 Series, Femoral Stem 12/14 Neck Taper Plasma Sprayed Press-Fit Cementless Size 7.5 Extended Offset Reduced Neck Length; Part Number: 00-7711-007-40 |  |  | March 01, 2018 | 2018 | Zimmer Biomet, Inc. | Certain packages labeled as size 7.5 Extended Offset Reduced Neck length hip implant, actually contained a Size 4 Standard Reduced Neck length taper hip implant. | Labeling mix-ups | RISK MANAGER: 1. Review this notification and notify affected personnel of the contents. 2. Quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility. 3. Complete Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. DISTRIBUTOR: 1. Review this notification and notify affected team members of the contents. 2. Quarantine affected product in your inventory. 3. Return affected product from your distributorship and from affected hospitals. a. Complete Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com Mark RECALL on the outside of the returned cartons. 4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 12 | 12 | Distributed in the following U.S. states: AZ, CA, CT, IN, MD, MI, MN, NJ, NY, OR, WA, and Puerto Rico. Distributed in China and India. | Open, Classified |  |  | 36
Z-1009-2018 | 79273 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | fastener, fixation, nondegradable, soft tissue |  |  | October 25, 2017 | 2017 | Zimmer Biomet, Inc. | It was identified that the product was not assembled correctly causing the implant shaft to shift and fall out during the opening of the device, which potentially could break the device. | Process control | In October of 2017 an URGENT MEDICAL DEVICE RECALL LETTER was issued to customers requesting that any remaining product be quarantined and returned to recalling firm. Questions or concerns can be directed to 574-371-3071. |  | 429 | 429 units | TX, NC, FL, PA, IN, WI, VA, NY, NJ, AL, CT | Open, Classified |  |  | 429
Z-1008-2018 | 79254 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 36MM COCR MODULAR HD -6MM |  |  | September 15, 2017 | 2017 | Zimmer Biomet, Inc. | One piece from a different lot mixed with lot 236840 and is incorrectly etched as -6MM offset instead of +3MM offset. As the incorrect etching and related size difference affects the functional use of this product, the affected lot is being removed from the field. | Error in labeling | Foreign risk managers were emailed notification of recall with the following instructions:Risk Manager Responsibilities:1.Review this notification and ensure that affected personnel are aware of the contents.2.If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3.Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4.Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5.If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 23 | 23 | Outside USA | Terminated | August 20, 2018 | 339 | 23
Z-1225-2018 | 79408 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | ACE Trochanteric Nail System |  |  | December 20, 2017 | 2017 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a medical device recall for the ACE Trochanteric Nail (ATN) system after a Device Performance Review (DPR) failed clinical analysis for survivorship pursuant to the acceptance criteria for the DPR protocol for the ATN system. | Device Design | On December 22, 2017 an URGENT MEDICAL DEVICE RECALL letter was issued to customers with a complete list of products subject to recall and requesting that the customer review the notification and notify all affected personnel of the recall. If there is affected product at the facility it is requested that these units be quarantined and returned to a sales representative for return to the firm. Questions or concerns can be directed to 574-371-3071. |  | 3454 | 3454 | USA (nationwide) | Open, Classified |  |  | 3454
Z-1010-2018 | 79236 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NaturaLyteÂż Liquid Acid Concentrate 2.0 mEq/L, Model # 08-2251-0 |  |  | September 29, 2017 | 2017 | Fresenius Medical Care Renal Therapies Group, LLC | The label specifies the potassium concentration is 2.0 mEq/L. Some of the product in the affected lot wasfound to have a potassium concentration of 2.25 mEq/L. | Release of Material/Component prior to receiving test results | The firm, Fresenius Medical Care, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated 10/4/2017 to its Consignees. The letter described the product, problem and actions to be taken. The Consignees are instructed to:" Immediately examine your stock to determine whether you have any NaturalyteÂż Liquid Acid Concentrate from the lot listed." If any product of this lot is found, discontinue use immediately." Place all units in a secure, segregated area." Contact your FMCRTG, LLC Customer Service Team at 1-800-323-5188 for instructions on how to return the recalled product." Promptly fill out and return the attached Reply Form via fax to: 781-699-9796 ATTN: Quality Department or SCAN and email to: NOTIFYRA@fmc-na.com.' and"Route the letter to others in your organization that may have any of the affected product that is subject to this recall.Please note that the NaturalyteÂż Liquid Acid Concentrate matching the Lot Number listed in the notification is the only lot subject to this recall.If you have any additional medical concerns or questions, please contact Medical Information and Communication: 855-616-2309 or Website: www.fresenius-medinfo.com. |  | 598 | 598 cases | US Distribution to states of: CA, FL, KS, LA, and TX. | Terminated | August 09, 2018 | 314 | 598
Z-1197-2018 | 79409 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | AIM Tibial Nails intramedullary fixation rod - NON-STERILE; Item numbers: 810508255NAIL TIB DYN 08 X 25.5 |  |  | February 06, 2018 | 2018 | Zimmer Biomet, Inc. | The AIM Tibial Nails failed a clinical analysis where the surgical revision rate exceeded the acceptance criteria in the Device Performance Review protocol for the implant system. | Under Investigation by firm | Zimmer Biomet sent an Urgent Medical Device Recall letter dated February 06, 2018 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to: 1. Review this notification and notify affected personnel of the contents. 2. Quarantine all affected product. Your Zimmer Biomet sales representative will remove it from your facility. 3. Complete the Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. DISTRIBUTOR: 1. Review this notification and notify affected team members of the contents.2. Quarantine affected product in your inventory. 3.Return affected product from your distributorship and from affected hospitals within your territory. a. Complete Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days. b. International returns, request an International Return Authorization (IRA) Number by emailing zimmerbiometintlirarequests@zimmerbiomet.com 4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com. |  |  | Not known | Worldwide and US of Indiana and country of the Republic of Korea. | Open, Classified
Z-1013-2018 | 79302 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | T/Pump temperature therapy pump, Models TP700 and TP700C |  |  | November 07, 2017 | 2017 | Stryker Medical Division of Stryker Corporation | The T/Pump Operations Manual and Maintenance Manual were incorrectly revised to add the sentence, Temperature therapy can also be for heat conservation for normothermia maintenance to the Indications for Use statement. | Error in labeling | The following actions are included in the consignee letter: 1. Locate the operations (101515 rev D.0) and maintenance (101176 rev F.0)manuals that accompanied the units listed on the attached business reply form.2. Find the corrected pages included in this letter and insert one into each affectedmanual at your location.3. Complete the enclosed business reply form to acknowledge that you havereceived and understand this notification.4. Return your completed business reply form by fax 269 488 8691 or emailproductfieldaction@stryker.com to Strykers regulatory department.5. If you have loaned or sold any of the products listed in this letter, please forwarda copy of this notice to the new users and advise us of their new location in thespace provided on the business reply form.6. If you have disposed of any of these units and they are no longer in use, please,advise us of their obsolescence by providing us with their serial number in thespace provided on the business reply form. |  | 9902 | 9902 | Distributed in all 50 states, the District of Columbia, and Guam. | Open, Classified |  |  | 9902
Z-1011-2018 | 79250 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard CR Porous Femoral 62.5mm (Left Femur) |  |  | October 03, 2017 | 2017 | Zimmer Biomet, Inc. | The affected products are labeled as a 62.5mm LEFT Femur; however, the contents are a Vanguard 70mm RIGHT Femur. | Error in labeling | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL REMOVAL" letter dated October 3, 2017 to distributors and customers. The letter described the product, problem and actions to be taken. The customers were instructed to:1. Review notification and ensure affected team member are aware of the contents2. Immediately locate and quarantine affected product in your inventory.3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form via digital copy to: corporatequallity.postmarket@zimmerbiomet.com within three (3) days.4. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet via SMS, via email to: rgarequest@zimmerbiomet.com-specify "RECALL" as RGA type.5. Identify additional accounts Zimmer Biomet has not notified.6. Retain a copy of your recall acknowledgement and product return forms for your records.Questions or concerns can be directed to the customer call center at 1-574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. |  | 8 | 8 | US Distribution to states of: FL, WI, MD, MO, and MI. | Terminated | August 23, 2018 | 324 | 8
Z-1293-2018 | 79464 | No | OSP | External uterine contraction monitor and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Uterine Electromyographic Monitor -Monica IF24 Interface System |  |  | December 15, 2017 | 2017 | GE Healthcare, LLC | if the cables of the Monica IF24 System are removed by a user, and then incorrectly re-connected the Fetal Heart Rate (FHR) and Maternal Heart Rate (MHR) signals would be displayed incorrectly (swapped). | Device Design | GE Healthcare sent an Urgent Medical Device correction Letter dated December 15, 2017 to their customers. The letter identified the affected product, problem and action to be taken. Customers are instructed to:1. Follow the attached safety instructions to replace the screws, see attached page 2.2. Follow the attached safety instructions to perform the Monica Novii function TEST to confirm correct cable connection, see attached page 3.3. If the Monica Novii System fails the Novii function TEST, contact a GE Healthcare Representative to schedule repair/replacement.4. If the Monica Novii System passes the Novii function TEST, you may continue to use Monica Novii System without any further action.5. Complete the attached fax back form (page 4&amp;5) for all Monica Novii System as soon as you have completed the screw replacement.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative |  | 355 | 355 | Worldwide - US Nationwide Distribution. | Open, Classified |  |  | 4194
Z-0997-2018 | 79339 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Pressure Injectable Multi- Lumen Central Venous Catheterization Kit with Blue FlexTip(R), ARROWg+ard Blue PLUS(R) |  |  | November 27, 2017 | 2017 | Arrow International Inc | Some of the products listed in the recall may be missing a sticker that contains the lot number, expiration date, and Unique Device Identification (UDI) information. | Error in labeling | On November 29, 2017 an Urgent Medical Device Correction notice was issued to customers stating the following:Our records indicate that you have received products that are subject to this action. We are now notifying our distributors to take the following actions:1.Immediately discontinue distribution and quarantine any products with the product codes and lot numbers listed above.2.Inspect affected products within your control to identify whether they have the sticker with the lot number, expiration date, and UDI information. Products missing this sticker should be returned to Arrow International so they are not distributed to health care providers in error. Products which are properly labeled should not be returned.3.Using the provided Customer Letter and Acknowledgement Form templates, communicate this correction to any of your customers who have received products included within the scope of the correction. Customers should return mislabeled products. Products that are properly labeled should not be returned.4.If any of your customers have mislabeled products to return, have them return the affected products to you, together with a completed Acknowledgement Form, for consolidation and return to Arrow International. In the event that an alternative approach is needed, contact Customer Service for more information at 1-866-246-6990.5.To return products to Arrow International, complete the enclosed Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email it to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Arrow International.6.If you and your customers have no affected stock or if the products in your stock are properly labeled, please complete the enclosed Acknowledgment Form and fax it to 1-855-419 |  | 23904 | 23,904 in total | US in the states of OH, FL, SC, TX, KY, NJ, MI | Open, Classified |  |  | 23904
Z-1582-2018 | 79411 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AQURE, Software version 2.3.0 and 2.3.1 |  |  | February 27, 2018 | 2018 | Radiometer Medical ApS | There is a potential problem relating to the AQURE System, versions 2.3.0 and 2.3.1, that may result in patient mix-up. | Device Design | Radiometer Medical sent a Notification letter dated February 27, 2018 to customers. The letter identified the affected product, problem and actions to be taken For questions contact your Radiometer representative. |  | 2 | 2 | US Nationwide Distribution in the states to Georgia and Wisconsin. | Completed |  |  | 2
Z-1194-2018 | 79474 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROW ENDURANCE" Extended Dwell Peripheral Catheter System, Catalog Numbers: EDC-00820 &amp; EDC-00820-B |  |  | December 15, 2017 | 2017 | Arrow International Inc | Certain lots of peripheral catheter systems may be missing the safety latch, intended to minimize the risk of sharps injuries. | Process control | Notification letters were sent to consignees with the following instructions: 1. If you have affected stock, quarantine so that the affected products can bereturned to Arrow.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to Arrow.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to Arrow. |  | 305 | 305 | Distributed in the following U.S. states: CO, FL, IL, NC, PA, SC, and TX. | Open, Classified |  |  | 305
Z-1006-2018 | 79266 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac versions 6.5 through 6.5 intended to be used to place patients at an accurately defined point within the treatment beam of a medical accelerator for stereotactic radiosurgery or radiotherapy procedures, in order to treat lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. ExacTrac may also be used to monitor the patient position during the treatment. |  |  | February 14, 2018 | 2018 | Brainlab AG | The usage of workflows that deviate from the recommended specifications in the User manual for ExacTrac Patient Positioning System (versions 6.0, 6.1, 6.2, 6.5) with Auxiliary Device Interface (ADI) and Varian Clinac or Varian TrueBeam, which may result in misinterpretation of beam authorization via ADI. | Other | The firm disseminated their notices by email on 02/14/2018. The notice requests the following actions:"User Corrective Action:Always follow the instructions below during use of ExacTrac Patient Positioning System (versions 6.0, 6.1, 6.2, 6.5) with Auxiliary Device Interface (ADI) and Varian Clinac or Varian TrueBeam:Do not use beam authorization by ExacTrac via ADI as a safety feature to ensure correct patient positioning before treatment.For any treatments performed with ExacTrac always verify that ExacTrac confirms the target position by displaying the OK icon. Do not treat the patient if the OK icon is not displayed.This indicates that the isocenter is in the planned treatment position. The treatment beam should not be activated unless the OK icon is shown.If using ExacTrac on a Varian Clinac or Varian TrueBeam system equipped with the Auxiliary Device Interface (ADI) always perform treatment in the following sequence:1.Start the ExacTrac Software and login using your user name and password2.Load the Plan / Patient on the Varian Treatment Application / Console3. Open the Plan / Patient on ExacTrac by clicking Open in the message box (refer to Figure 2)4. Prepare the beam on the Varian Treatment Application / ConsoleIf you intend to perform treatments for which ExacTrac is not used for patient positioning, please click Cancel in the message box to confirm not opening the patient and to continue the treatment without ExacTrac.ExacTrac can be configured to directly load the patient automatically in ExacTrac if it is opened on the Varian Treatment Application / Console - meaning that the message box (refer to Figure 2) is deactivated. However, this may not be the preferred solution for each user, as automatic loading causes delays in treatment preparation if no positioning via ExacTrac is to be used, since in these cases the patient data has to be closed manually. Please contact your local Brainlab Customer Support Representat |  | 587 | 587 systems | worldwide | Open, Classified |  |  | 587
Z-1070-2018 | 79245 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | RaPETÂż RF Latex Reagent used as a component in the following kits: RaPET RF (1155-100, Control# 173201 expiry 2019-04-30) |  |  | February 13, 2018 | 2018 | Stanbio Laboratory, LP | All patient serum samples yield a weakly positive test result. | Component design/selection | Recall notification letters were sent to affected consignees on 2/14/2018. The letter identified the affected product, problem and actions to be taken. The letter instructed the consignee to stop using affected products, quarantine products, and exchange the kit for an unaffected one. |  | 686 | 686 units | Worldwide Distribution - United States nationwide , Canada, Korea, Mexico, Trinidad | Terminated | September 27, 2018 | 226 | 686
Z-1065-2018 | 79328 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound System |  |  | January 26, 2018 | 2018 | Siemens Medical Solutions USA, Inc. | The application may underestimate the EROA (Effective Regurgitant Orifice Area) in comparison to the same patient results obtained with the 4Z1c volume transthoracic echocardiography transducer. | Under Investigation by firm | Siemens sent a Customer Notification letter to their customers. The letter identified the affected product, problem and actions to be taken. For questions contact your Customer Care Center. |  | 564 | 564 systems | Worldwide Distribution - U.S. nationwide and the countries of Algeria, Argentina, Australia, Austria, Azerbaijan, Brazil, Brunei, Canada, Chile, China, Columbia, Ecuador, Egypt, El Salvador, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Israel, Italy, Japan, Malaysia, Malta, Mexico, Netherlands, Norway, Pakistan, Philippines, Poland, Portugal, Republic of Korea, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Taiwan, Thailand, Turkey, U.A.E, Ukraine, United Kingdom, United States, Uzbekistan, and Vietnam. | Open, Classified |  |  | 564
Z-1073-2018 | 79429 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | PORT-A-CATH II PS/Titan, Tray PAC II. Labeled as: |  |  | December 04, 2017 | 2017 | Smiths Medical ASD Inc. | Certain models and lots of PORT A CATH implantableport kits may have been misassembled. The kits may incorrectly contain a 6Fr or 6.5Fr size introducer instead of the correct 8.5Fr size introducer. The Power PORT A CATH II kit may also contain the incorrect locking collars in addition to the incorrect sized introducer. | Process design | Customers were notified via letter on about 12/04/2017. Instructions include to examine current inventory for affected product, collect and isolate any potentially affected product, determine the total number of potentially affected product in possession and complete and return the Urgent Medical Device Recall Response Form, return any affected product on hand, and notify customers if the affected products were further distributed. |  | 1077 | 1,077 units total | Worldwide Distribution. US nationwide including Puerto Rico, Austria, Bahamas, Belgium, Dominican Republic, Netherlands, Spain, and Turks and Caicos. | Open, Classified |  |  | 1077
Z-1063-2018 | 79269 | No | LCO | Platelet factor 4 radioimmunoassay. | Hematology | Hematology | 510(k) | ZYMUTEST HIA MonoStrip IgGAM. |  |  | November 09, 2011 | 2011 | Hyphen BioMed | A voluntary recall is initiated for ZYMUTEST HIA lgGAM - Ref RK040D - Lot F1700415 and ZYMUTEST HIA MonoStrip lgGAM - Ref RK041D - Lot F1700416 because they may produce false positive results. | Other | The manufacturer, HYPHEN BioMed, initiated a voluntary recall of the two lots via email. Distributors and end users are being notified and instructed to: 1) acknowledge receipt and understanding of notification, and 2) destroy affected products remaining in stock. For further questions, please call (513) 770-1993. |  | 30 | 30 | Worldwide Distribution - US Distribution and to the countries of : Austria and France. | Terminated | June 21, 2018 | 2416 | 124
Z-1064-2018 | 79387 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 6.5mm x 125mm x 16mm Cannulated Screw, Item Number: 110009948 (UDI: (01)00887868111327 (17)270421 (10)120910) |  |  | December 07, 2017 | 2017 | Zimmer Biomet, Inc. | One lot of 3.5mm x 24mm Cortical Locking screws is labeled as 6.5mm x 125mm x 16mm Cannulated screws. | Error in labeling | Zimmer Biomet sent an Urgent Medical Device Recall letter dated December 2017 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to: 1. Review this notification n and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory. 3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. Request an International Return Authorization Number by emailing zimmerbiometintlirarequests@zimmerbiomet.comc. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.d. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.e. Mark "RECALL" on the outside of the returned cartons.4. Return a copy of the your Inventory Return Certification and product return form.5. For questions call customer service at 574-374-3071. |  | 6 | 6 | Canada. | Open, Classified |  |  | 6
Z-1069-2018 | 79292 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | STA - NeoplastineÂż Cl Plus Âż (ref. 00606) |  |  | October 25, 2017 | 2017 | Diagnostica Stago, Inc. | Confirmed defect of homogeneity where some reagent vials will give prolonged Prothrombin Time (decreased PT%). | Device Design | Stago sent an Urgent Field Safety Notice letter dated October 25, 2017 to their customers. The letter identified the affected product, problem and actions to be taken: Customers were instructed to:If it is not already done, run a Quality Control test at every change of vial (if you have affected lots of product).Return to company, by fax or email, an acknowledgement form confirming that the recall letter was read by the receiving company.For questions contact the Stago Hotline (1-800-725-0607). |  |  |  | Nationally | Open, Classified
Z-1072-2018 | 79263 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | TunneLoc Tibial Fixation Device (Within an ACL Kit) |  |  | October 11, 2017 | 2017 | Zimmer Biomet, Inc. | Gamma sterilization validation testing indicates the sterilization dose for the product may not be adequate. | Vendor change control | Zimmer Biomet sent an URGENT MEDICAL DEVICE RECALL letter dated October 16, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review inventory and quarantine affected product, which will be returned to Biomet through a sales representative. Questions or concerns can be directed to 574-371-3071 |  |  |  | US Nationwide Distribution | Open, Classified
Z-1076-2018 | 78655 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | GENICON EZEE Retrieval, sterile, Rx only. |  |  | July 19, 2017 | 2017 | Genicon, Inc. | Plastic handle is breaking off of shaft inhibiting proper removal per Instructions for Use. | Device Design | Consignees were notified via email on 07/19/2017. Instructions included |  | 2967 | 2,967 units | Distributed domestically to FL, IL, MI, TN, UT. Distributed internationally to Dominican Republic, Ireland, Denmark, Spain, Switzerland, United Kingdom. | Open, Classified |  |  | 2967
Z-1075-2018 | 79386 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Iridex TruFocus LIO Premiere" Laser Indirect Ophthalmoscope ( LIO ) |  |  | February 27, 2018 | 2018 | Iridex Corporation | It was discovered that the laser energy was not going where intended to put it, resulting in focal cataracts and iris burns. | Device Design | Action Taken by IRIDEX: 1. A investigation was started immediately upon receipt of the returned devices from the Sloan Kettering complaint on November 3, 2017. CAPA 0125 was initiated to document the events, root cause, correction, corrective and preventive action and effectiveness of these actions.2. A ship hold was initiated on November 15, 2017.3.As described below, IRIDEX plans to notify customers and distributors that it is voluntarily recalling all models of the TruFocus LIO Premiere.4. Public WarningA press release was prepared and sent out on February 23, 2018.For further questions please call (650) 962-8848 Ext. 3672 |  | 104 | US - 58 OUS - 46 | Worldwide Distribution - USA Distribution and to the countries of : China, Germany, Australia, Canada, Ireland, Lebanon, Japan, Israel and Qatar | Open, Classified |  |  | 104
Z-1077-2018 | 79364 | No | PHP |  |  | Pathology | PMA | Epi proColon Plasma Quick Kit; Catalog Number: M5-02-001: Consists of 1x Lysis Binding Buffer, 1x Wash B Concentrate, 1x Dilution Buffer, 4x Bisulfate Solution, 1x Protection Buffer; UDI: (01)04260400670036(17)180228(10)1603908; |  |  | November 02, 2017 | 2017 | Epigenomics Ag | The diagnostic test kit may produce invalid test runs which may delay test results. | Unknown/Undetermined by firm | Users were instructed to discontinue use of the kits, complete the acknowledgement form, and distribute the notification to relevant persons within organization. If kits were further distributed the consignee was asked to forward the safety notice to these entities. |  | 200 | 200 | Shipped to one U.S. distributor in NY. Further distributed in 9 states: CA, MI, NC, NE, NJ, PA, TN, TX, UT. | Terminated | June 19, 2018 | 229 | 200
Z-1078-2018 | 79329 | No | MCZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | SKIN STAPLE REMOVER |  |  | October 23, 2017 | 2017 | Centurion Medical Products Corporation | Convenience Kit contains a PVP Prep Pad, received by a supplier, that does not meet the iodine assay level required to support the 36 month shelf life. | Nonconforming Material/Component | On October 16, 2017 a PRODUCT CORRECTION letter was issued to customers notifying them that if they had any kits in their inventory, Centurion will supply labels to affix to the front of each kit instructing the user to discard the prep pad upon opening the kit. Other components of the kit may continue to be used and are NOT within the scope of this recall. |  |  |  | Nationally | Terminated | June 21, 2018 | 241
Z-1099-2018 | 79424 | No | GBW | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Malecot Nephrostomy Catheter System, Product Number: M001224110, GTIN: 08714729030980 |  |  | December 12, 2017 | 2017 | Boston Scientific Corporation | Certain Malecot Nephrostomy Catheters are breaking at the mid-shaft bond during use. | Under Investigation by firm | The firm, Boston Scientific, sent an "Urgent Medical Device Product Removal - Immediate Action Required" letter dated 12/12/2017 to its consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to: identify and remove all affected batches from your inventory; immediately discontinue use of and segregate the affected product, and return it to Boston Scientific. Distributors were notified that this removal is to the customer level and asked to notify any customers who received affected product. Consignees were asked to complete and return the Replay Verification Tracking Form (even if you do not have any product to return) via email: BSCFieldActionCenter@bsci.com or Fax to: Field Action Center at 1-866-213-1806.If you have any questions, your local Sales Representative can answer any questions that you may have regarding this removal or contact Vice President , Regulatory Affairs at 508-683-4398 or email: Donna.Gardner@bsci.com or BSCFieldActionCenter@bsci.com. |  | 422 | 422 | Worldwide Distribution: US (nationwide) in the following: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KA, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, including the District of Columbia and Puerto Rico; and countries of: Argentina, Bolivia, Brazil, Canada, Chile, Costa Rica, Dominican Republic, Ecuador, El Salvador, Guatemala, Mexico, Panama, Peru, Venezuela, Australia, Bangladesh, Botswana, China, Kuwait, Lebanon, Macau, Malaysia, Philippines, Saudi Arabia, Singapore, South Africa, South Korea, Turkey, United Arab Emirates, and Zimbabwe. | Open, Classified |  |  | 12218
Z-1105-2018 | 79247 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | smith&nephew;, LEGION HK FEMORAL WEDGE (WEDGE AND SCREW), 5 MM, SIZE 5, REF 71422159, QTY: (1), STERILE |  |  | February 01, 2018 | 2018 | Smith &amp; Nephew, Inc. | The screws package with devices within lot 14AAP6352A were 5 MM instead of 10MM and are too short to be used with the Legion HK Distal Wedge Size 5. | Process control | The firm initiated their recall by email on February 1, 2018. The letter stated the following:"Required Actions:1. Please inspect your inventory and locate any unused devices from the listed product and batch numbers on the first page of this Field Action Notification, and quarantine them immediately. a. If you are a distributor, you must notify your customers of the field action and ensure that these actions are carried out.2. If you have no product to return, please put an X in the appropriate location below.3. If you have product to return, please list the batch numbers and quantities of each batch that you are returning in the appropriate boxes below.4. Complete the remainders of the form sign and send to FieldActions@smith-nephew.com or fax to 901-566- 7975.Please Note  even if you have no product to return, this form must be completed, signed and returned.5. Once the form is received by Smith &amp; Nephew, you will be sent a Return Authorization (RA) number.If you have any questions or concerns regarding this recall please contact FieldActions@smith-nephew.com." For further questions, please call (901) 396-2121. |  | 28 | 28 units | Worldwide Distribution - US Distribution to the states of: NY, IL, NJ, TX, CO, and to the countries of : Belgium, Great Britain, Australia, Germany, Netherlands, Norway. | Terminated | October 22, 2018 | 263 | 28
Z-1127-2018 | 79436 | Yes | NFK | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | REPAIR KIT, REF RPK-01, UDI 884908039856 |  |  | February 23, 2018 | 2018 | Medical Components, Inc dba MedComp | There is the potential for connection failures related to the female two-part luer to hemodialysis blood tubing lines, etc. which may contribute to blood loss or air emboli. | Nonconforming Material/Component | On February 23, 2018, Medcomp distributed Product Recall and Distribution notices to international distributors via email and to U.S. facilities via courier service on February 26, 2018. Medcomp is requesting the return of all un-used affected product. Customers should immediately examine inventory and quarantine product subject to recall. If for any reason the product was further distributed, please identify the recipients and notify them at once of this product recall. Your notification may be enhanced by including a copy of this recall notification letter. Contact your customer service representative for a Returned Goods Authorization number if necessary by calling (215) 256-4201. Complete and return the Distribution Notice via fax (215-256-9191) or email (complaints@medcompnet.com. |  | 1575 | 1575 kits | Worldwide Distribution - US Nationwide in the states of AL, AR, CA, CO, CT, DE, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MS, NC, NH, NJ, NM, NV, NY, OH, OK, PA, RI, TN, TX, UT, VA, VT, WA &amp; WI. International: Australia, Austria, Belgium, Canada, Ecuador, France, Germany, Great BritainGreece, Ireland, Israel, Italy, Netherlands, Netherlands, South Africa, SpainSweden, Turkey &amp; United Arab Emirates. | Open, Classified |  |  | 3215
Z-1133-2018 | 79529 | No | FBN | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SpyScope DS Access and Delivery diagnostic biliary catheter, M00546600 |  |  | December 20, 2017 | 2017 | Boston Scientific Corporation | Certain lots of devices may have the working channel sleeve protruding from the camera cap which may result in injuries from minor tissue damage to significant bleeding. | Process design | The notification letter requests the consignee to identify and return the affected product along with the response form. Distributors are asked to share the recall information to the customer level. 12/20/2017 by mail. |  | 5026 | 5026 | Worldwide Distribution US nationwide, Austria, Belarus, Belguim, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Hungary, Ireland, Israel, Italy, Netherlands, Norway, Poland, Portugal, Romania, Slovakia, Spain, Sweden, Switzerland, Bahrain, China, Hong Kong, India, Malaysia, Pakistan, Qatar, Saudi Arabia, Singapore, South Africa, South Korea, Thailand, Tunisia, Turkey, Vietnam, Brazil, Canada, Cost Rica, Ecuador, Guatemala, Mexico, Peru | Open, Classified |  |  | 5026
Z-1117-2018 | 79268 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Apex Knee System, Tibial Baseplate Size 6 Left, non-porous - KC-2206L |  |  | November 20, 2017 | 2017 | OMNIlife science Inc. | The Apex Knee System Tibial Baseplate identified has an improper expiration date marked on the external packaging. The improper expiration date (2202-03) ismarked on the external packaging only. The correct lot number (2022-03) is present in the internal packaging and sterile batch documentation. | Incorrect or no expiration date | OMNIlife science, Inc. sent an Urgent Medical Device Advisory Notice dated November 20, 2017.Actions to be taken by the Customer/User"Immediately quarantine lot 26519 for KC-2206L"Return Product Using RMA# XXXX to OMNI."This advisory notice only includes LOT 26519. It does not affect other lots.You must acknowledge this notice by returning the attached Acknowledgement and Receipt Form and accompanying by FAX to 508-819-3390 or my email to CS@omnils.com. For further questions, please call (508) 824-2444. |  | 32 | 32 | Worldwide Distribution - US Distribution and to the country Spain. | Completed |  |  | 32
Z-1130-2018 | 79296 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Situate Gauze RF and X-Ray detectable. |  |  | October 12, 2017 | 2017 | Covidien LLC | This voluntary recall is being conducted due to the potentialfor failure to detect the affected Situate" gauze - RF and Xraydetectable products by the Situate" detection console dueto a manufacturing error. | Nonconforming Material/Component | All consignees were notified via FedEx or certified mail on October 12, 2017 and the letter 'Urgent Field Corrective action Notice' informs customers of the potential for failure to detect the affected Situate" gauze RF and X-Ray detectable products by the Situate" detection console and the actions they should take. Customers are requested to acknowledge receipt and understanding of the Urgent Field Correction Action Notice. For further questions, please call (203) 492-5000. |  | 20800 | 20,800 | U.S. | Open, Classified |  |  | 20800
Z-1128-2018 | 79321 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | U2 Total Knee System Tibial Insert, Posterior Stabilized, #3, |  |  | November 24, 2017 | 2017 | United Orthopedic Corporation | UOC received a customer compliant report from our distributor on October 31, 2017which stated that contents of the package do not match the product labeling. | Labeling mix-ups | Urgent Field safety notice was distributed to customers on November 24, 2017 via e-mail. The letter identified the affected product, problem and actions to be taken. For further questions, please call (949) 328-3366. |  | 69 | 69 pieces | Worldwide Distribution - US Distribution and to the countries of: UK, Iran, Greece, Pakistan, Switzerland, France, China, Taiwan, Italy, Korea and South Africa. | Terminated | October 09, 2018 | 319 | 83
Z-1132-2018 | 79276 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow B Solution - Premixed Dialysate for Hemodialysis with NxStage System One |  |  | November 10, 2017 | 2017 | NxStage Medical, Inc. | Using different fluid formulations of NxStage PureFlow B Solution at the same time during treatment may cause patient health risk. The NxStage PureFlow B Solution is for use with systems that use one premixed dialysate during hemodialysis. The original labeling did not warn against mixing dialysate formulations and did not instruct staff to verify that the fluid is flowing during use. | Labeling design | Customers are advised to :1. Always make sure when using PureFlow B Solution that fluid is flowing freely from all bags.2. Recognize that any practice of hanging more than one fluid type at the same time is not an intended use of the NxStage System One" (as it does not monitor fluid composition) and is not a labeled indication of PureFlow B solutions.3. Please make sure that all users of PureFlow B Solution in your organization and other concerned persons are informed regarding this Field Safety Notice.4. Complete the enclosed customer reply form and return it to NxStage. Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices. |  | 1243400 | 1243400 units | US nationwide distribution | Open, Classified |  |  | 1243400
Z-1131-2018 | 79213 | No | KNG | Fetal ultrasonic monitor and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical 2 MHZ Waterproof OB Probe |  |  | October 06, 2017 | 2017 | CooperSurgical, Inc. | The 2MHZ prove was incorrectlhttp://cts.fda.gov/division-tracking/images/trash.pngy labeled as a 3MHZ probe | Error in labeling | CooperSurgical sent an Urgent Medical Device Recall letter dated October 6, 2017 to affected customers via UPS with confirmed delivery receipt. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to discontinue use of affected products and complete the attached Acknowledgement and Receipt Form for replacement. If you have any further questions please feel free to contact me at 203.601.5200. |  | 13 | 13 units | Worldwide Distribution - US Nationwide and the countries of ( Canada and England) | Completed |  |  | 13
Z-0194-2018 | 78315 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MONTERIS MEDICAL NEUROBLATE SYSTEM, used for MRI-guided neurosurgical ablation. |  |  | October 05, 2017 | 2017 | Monteris Medical Corp | The firm received complaints that involved an unanticipated interaction between a NeuroBlate system and 1.5 Tesla (1.ST) MRI systems. | Device Design | The firm issued advisory notice 10/05/2017. The notice discussed complaints received that involved an unanticipated interaction between a NeuroBlate system and 1.5 Tesla (1.5T) MRI systems. The notice provided added recommendations for all NeuroBlate users to mitigate the anomaly. |  | 52 | 52 systems | Nationwide and Canada | Open, Classified |  |  | 52
Z-1317-2018 | 79515 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | D-RAD Self-Tapping Locking Screw intended for the fixation of fractures involving the distal radius. smith&nephew; D-RAD 2.4MM X 10MM S-T LOCKING SCREW, REF 74692410, QTY (1), STERILE R |  |  | March 07, 2018 | 2018 | Smith &amp; Nephew, Inc. | One lot of D-RAD Self-Taping Locking Screws used with the Distal Radius Fracture Kit were machined out of specification. Screws measured over tolerance within the head thread form by 0.002-0.003 inch. | Under Investigation by firm | The firm initiated their recall by letter and email on 03/07/2018. The firm requested return of the product. |  | 59 | 59 units | US, Puerto Rico, Malaysia | Terminated | October 31, 2018 | 238 | 59
Z-1148-2018 | 79538 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis one fluoroscopic x-ray system and patient table |  |  | January 12, 2018 | 2018 | Siemens Medical Solutions USA, Inc | Certain patient table systems may have a missing or incorrectly installed snap ring that could cause the bolt to slide out of its base and the tabletop carriage to slide to the floor. | Process control | Siemens sent an Urgent - Medical Device Correction letter dated January 5, 2018. Siemens will conduct an inspection of all affected Artis tables and verify the correct positioning of the bolt as well as the snap rings. The letter recommends the consignees follow Please follow the standard emergency procedures they have in place in the event of system failures. It is recommended to consider reinforcing these emergency procedures until the update has been completed. The Siemens service organization will contact the consignees to arrange a date to perform the corrective action. For further questions, please call (610) 448-6461. |  | 16 | 16 | US Distribution to the states of : AL, AR, CA, FL, IN, LA, MI, MO, NC, NY, PA, TX, and WA. | Terminated | August 28, 2018 | 228 | 16
Z-1172-2018 | 79536 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | HEMASHIELD Knitted Double Velour Cardiovascular Fabric |  |  | January 03, 2018 | 2018 | GETINGE US SALES LLC | A probable mislabeling of the product: a different-sized product is referenced on the product packaging label (i.e. 2.5 cm (width) x 7.6 cm (length)) while the actual product inside the box is 2.5 cm (width) x 15 cm (length). | Labeling mix-ups | On January 3, 2018 an URGENT MEDICAL DEVICE RECALL NOTICE was issued to the (1) affected customer. A return authorization was issued with the recall letter, which is intended to be returned to the recalling firm. Questions or concerns can be directed to 1-800-528-7486 |  | 1 | 1 | FL | Terminated | July 03, 2018 | 181 | 1
Z-1303-2018 | 79532 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via. Medical Device Software. Picture archiving and communications system. |  |  | January 08, 2018 | 2018 | Siemens Medical Solutions USA, Inc | A software functionality in the report sections Findings Information and Summary of Measured Findings is not performing properly. The system will use the original values instead of displaying and saving corrected/modified values. Outdated information could be sent to the referring physician when creating the formal report in the reporting system. | Software design | Customer Advisory Notice SY065/17/P was sent to direct accounts. Firm is continuing to work on software solution. A workaround is available and is described in the Customer Advisory Notice. |  | 236 | 236 units | Device is software only. No products are distributed to wholesale dealers, distributers or retailers | Open, Classified |  |  | 236
Z-1219-2018 | 79475 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | (1) Sigma HP Cemented Trochlea Size 1 Narrow Left, Catalog Number: 102403100, GTIN: 10603295001614; (2) Sigma HP Cemented Trochlea Size 2 Narrow Left, Catalog Number: 102403200, GTIN: 10603295001621; (3) Sigma HP Cemented Trochlea Size 3 Narrow Left, Catalog Number: 102403300, GTIN: 10603295001638; (4) Sigma HP Cemented Trochlea Size 4 Narrow Left, Catalog Number: 102403400, GTIN: 10603295001645; (5) Sigma HP Cemented Trochlea Size 5 Narrow Left, Catalog Number: 102403500, GTIN: 10603295001652; (6) Sigma HP Cemented Trochlea Size 1 Narrow Right, Catalog Number: 102404100, GTIN: 10603295001669; (7) Sigma HP Cemented Trochlea Size 2 Narrow Right, Catalog Number: 102404200, GTIN: 10603295001676; (8) Sigma HP Cemented Trochlea Size 3 Narrow Right, Catalog Number: 102404300, GTIN: 10603295001683; (9) Sigma HP Cemented Trochlea Size 4 Narrow Right, Catalog Number: 102404400, GTIN: 10603295001690; (10) Sigma HP Cemented Trochlea Size 5 Narrow Right, Catalog Number: 102404500, GTIN: 10603295001706 |  |  | December 21, 2017 | 2017 | DePuy Orthopaedics, Inc. | The SIGMAÂż HP PFJ Cemented Trochlear Implants, a standalone component of the partial knee system is being recalled based on elevated revision rates observed as part of the company s post market surveillance process. | Under Investigation by firm | Notifications were sent to the firm's consignees with the following instructions. Immediately Inspect Inventory: DISTRIBUTORS &amp; SALES OFFICES: Inspect all inventory to identify any affected devices. 2. Return Affected U.S. Distributor/Sales Office Inventory: If any affected devices are found in a U.S. Distributors/Sales Offices inventory, return to: DePuy Synthes Joint Reconstruction 3. To Expedite the Return Process: Be sure to reference H17-19 on all return paperwork and/or online return forms and on the outside of the box when returning recalled devices. 4. Upon receipt of the affected devices, a credit will be issued using the normal returns process. 5. Identify Consignment Hospitals: Within 5-business days, complete the Distributor / Sales Office Response Card: 5.1. Provide the Medical Facility contact information and indicate the US Distributor/Sales Office has returned all affected devices. MEDICAL PROFESSIONALS: Steps to Take: " Cease using the affected components. " Return Affected Implants: Complete the reconciliation form and return to the sales consultant or fax to 574-371-4939 within five (5) days of this notice. Additional Notifications: o Notify surgeon users at your facility by providing them with a copy of this notice to ensure surgeon users are aware of this device recall. Forward this notice to others in your facility that need to be informed. o If any affected product has been forwarded to another facility, contact that facility immediately to communicate this field action with the facility/facilities. For further questions, please call (574) 372-7136. |  | 2110 | 2110 | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and to the states of : AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, and the District of Columbia., and to the countries of : Australia, Austria, Belgium, Canada, Croatia, Denmark, Finland, France, Germany, India, Ireland, Italy, Korea, Luxembourg, Malaysia, Netherlands, Portugal, Singapore, Slovakia. Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, and United Kingdom. | Open, Classified |  |  | 2110
Z-1337-2018 | 79548 | No | EQJ | Ear, nose, and throat bur. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Gyrus ACMI Round Cutting Burrs - to be used with the Diego Elite system |  |  | August 03, 2017 | 2017 | Gyrus ACMI, Inc. | "Chattering" when cutting burrs are used in bone. |  | As a result, OSTA will issue a Field Safety Notice (FSN) to all customers that received variable exposure burrs of 5mm or greater diameter. The FSN willinform them of the concern and will advise to exercise due caution if using the affected device in proximity to sensitive or vital anatomical structures.On September 6, 2017 the firm sent a URGENT: MEDICAL DEVICE FIELD SAFETY NOTICE to the attention of the Surgical Risk Management Departmentfor the following products: MBUR5060FRCV - Burr, 5mm Round Cutting, Variable Exposure, Standard Length, MBUR6060FRCV - Burr, 6mm Round Cutting, Variable Exposure, Standard Length and MBUR7060FRCV - Burr, 7mm Round Cutting, Variable Exposure, Standard Lengthfor All product with an expiration date prior to August 3, 2017Description of the problem:Olympus has become aware of an issue that requires your attention. This letter pertains to the Olympus Round Variable Cutting Burrs ( Burrs ) referenced above. Our records indicate that you may have purchased affected burrs. We are in receipt of a number of non-injury-related complaints that has made Olympus aware of a possibility that the burr may chatter when drilling bone. Although this chatter phenomenon can occur with any type or brand of cutting burr, depending on technique and bone density, Olympus has become aware that the larger diameter variable exposure burrs listed above can exhibit more chatter than other similar devices in some situations, particularly when they are extended to longer lengths. After extensive testing, Olympus has determined that our variable exposure burrs of less than 5mm diameter, and all fixed exposure burr offerings, are not affected by this phenomenon.Because of this, Olympus has elected to issue this field safety notice (FSN). Please refer to the enclosure with this letter that cautions of the possibility of the chatter phenomenon in the larger variable offerings. The FSN also recommends that surgeons exercise due caut |  | 324 | 324 | NM, MI, and TXEU, Australia, Japan, and Hong Kong | Open, Classified |  |  | 324
Z-1150-2018 | 79540 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis angiographic x-ray system, Model Numbers: 7412807, 7413078, 7728392 |  |  | January 05, 2018 | 2018 | Siemens Medical Solutions USA, Inc | Certain systems may have leaking of the plastic coolant couplings, the liquid can enter into the flat panel detector leading to a failure of the detector. It may be necessary to cancel or restart the treatment or transfer the patient to an alternate or another system. | Process control | The firm, Siemens Healthineers, sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated 1/5/2018 to its consignees. The letter describes the product, problem and actions to be taken. The consignees were instructed to follow their standard emergency procedures in place in the event of system failures; to consider reinforcing these emergency procedures until the update has been completed. Siemens will replace all affected plastic couplings with metal couplings. The Siemens service organization will contact the consignees to arrange a date to perform this corrective action.If the device has been sold and is therefore no longer in your possession, please forward this field correction notice to the new owner.Our service organization will contact you shortly to arrange a date to perform this corrective action. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436. |  | 22 | 22 | US Distribution in the following U.S. states: AZ, CA, FL, IA, IL, LA, MI, NC, NJ, OH, OK, and TX. | Terminated | August 23, 2018 | 230 | 22
Z-1137-2018 | 79547 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | MyLab Ultrasound System; Models 6100, 6150, 6200, 6250; PA023E Probe 7.5/10 |  |  | January 03, 2018 | 2018 | ESAOTE S.P.A. | The ultrasound system used in certain settings may result in overheating of the ultrasound probe head. | Software design | The January 3, 2018, notification letter requests the consignee complete and return a recall response form. The letter states that Esaote customer service will contact the consignee about the software update. In the meantime, it is recommended that the consignee does not use the PA023E probe in Pulsed Wave (PW) Doppler mode and that they refrain from using the system in Elaxto mode where the patient is unconscious or sedated. The consignee is asked to check the temperature of the probe before use. |  | 1 | 1 | US distribution in the states: AZ, FL, OK, PA, and TX. | Open, Classified |  |  | 6
Z-1278-2018 | 79478 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008 K2 Hemodialysis Machine with software version 5.40, Models: (1) Hemodialysis SYS OLC/Diasafe PLS 190610, UDI 00840861100859 (+Serial Number), (2) Machine Short Cap, OLC.DP HP 190630, UDI 00840861100866 (+Serial Number) |  |  | December 14, 2017 | 2017 | Fresenius Medical Care Renal Therapies Group, LLC | When the recirculation ultrafiltration (UF) Goal is set to a value greater than 200ml in service mode and the user starts a treatment using the SLED (Sustained Low Efficiency Dialysis) program the display will show an invalid message and the UF pump will run at the recirculation UF rate which may be up to 4000ml/hour. The SLED program is required to limit the UF rate to a maximum of 1000 ml/hour. | Device Design | The recall notification instructs the user to the UF Goal with values less than 200ml when in Service Mode and to monitor and set the UF Rate to the desired setting after starting the SLED treatment. The consignee is advised that a software update is forthcoming to correct the issue. For further questions, please call (800) 227-2572. |  | 127 | 127 | US Distribution to the states of: AR, CA, CO, FL, GA, IL, MA, MD, MN, NC, NY, OH, PA, SC, TN, and TX. | Open, Classified |  |  | 127
Z-1145-2018 | 79426 | No | OJV | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom Dialysis Kit CF-4230 Without Gel, which includes Flex-NeckÂż Classic Peritoneal Dialysis Catheter 2 Cuffs Infant Coiled, Catalog # K14-00078 |  |  | September 20, 2017 | 2017 | Merit Medical Systems, Inc. | 1 Cuff catheters were incorrectly configured and labeled as 2 Cuff product and placed into these dialysis kits. | Mixed-up of materials/components | On 09/20/17, an URGENT PRODUCT RECALL NOTICE was sent to all direct consignees via mail, informing them of a voluntary recall of specific lots of peritoneal dialysis catheters and catheter containing kits. Consignees were advised to immediately do the following: identify and quarantine effected devices; inform personnel to whom the devices were distributed of the field action; fill out and email the customer response form to response@merit.com within 5 days, and return effected devices within ten days using Federal Express account 112952756. Consignee questions were directed to phone: 801-208-4381 or fax: 801-316-4880. Sales representatives were sent an email communication informing them of the recall and instructing them to contact effected customers to help them return Customer Response Forms and product. |  | 8 | 8 | United States, Austria, Finland, Germany, Ireland, Switzerland, and United Kingdom | Open, Classified |  |  | 82
Z-1146-2018 | 79324 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | ClariVein IC infusion catheter, Model Number: 65-018-E4S |  |  | November 09, 2017 | 2017 | Vascular Insights, LLC | Certain catheters could have a compromised catheter shaft due to a manufacturing defect which could result in some delamination of the catheter shaft including during introduction of the catheter into a vascular access introducer. | Process control | Vascular Insights sent an Urgent Voluntary Medical Device Recall letter to their customers. The letter identified the affected product, problem and the actions to be taken. Customers were instructed to 1. Do not use or distribute any recalled product. 2. Immediately remove all recalled product from your inventory.3. Segregate recalled product in a secure location and follow the instruction herein to return the recalled product and receive replacement product. 4. Immediately forward a copy of this recall notification to any/all sites to which you may have distributed the affected product and provide directions to follow the instructions contained herein. Customers were asked to complete and return the enclosed Reply Verification Tracking Form. |  | 928 | 928 | U.S Nationwide in the states: AZ, CA, CT, FL, GA, IA, IL, IN, KS, KY, MA, MD, MO, NE, NJ, NY, OH, PA, SC, TN, TX, VA, WA | Terminated | December 18, 2018 | 404 | 928
Z-1149-2018 | 79264 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Wash Kit (REF 10680) Lots 16279, 17045 and 17215 |  |  | November 21, 2017 | 2017 | Medica Corporation | The firm has initiated a field recall of the Wash1 Kit REF 10680 due to reported inconsistent Lipase values when analyzing patient samples. The firm's internal investigation identified the underlying cause to be an ineffective wash solution resulting in an interference carryover from the triglyceride (TRIG) assay if they are on the same worklist. | Under Investigation by firm | Firm emailed all consignees and instructed them to remove/discard affected wedges (bottles) and contact Medica for immediate replacement1 Immediately discontinue use of the Wash1 wedges (REF 10680) (Lot# listed above) and discard any remaining kits/wedges. 2. Please discuss possible concerns about lipase results achieved using these lots of Wash wedges with your laboratory manager and/or clinician in order to decide on appropriate measures. Report any concerns to Medica for evaluation and review.3.Submit claims for replacement of affected product to your authorized Medica EasyRA dealer. For further questions, please call (781) 541-7560. |  | 392 | 392 | Worldwide and US Distribution. | Open, Classified |  |  | 392
Z-1153-2018 | 79398 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | J-Plasma Precise(R) OPEN, Catalog Numbers: BVX-044-BPP, BVX-044-BPS, BVX-150-BPP, BVX-150-BPS |  |  | February 08, 2018 | 2018 | Bovie Medical Corporation | Bovie Medical is issuing a correction to address incomplete insertion of the J-Plasma(R) handpiece cable plug into the generator receptacle resulting in an incomplete helium seal. The potential exists for patient fluid backflow into the handpiece and the generator receptacle leading to risk of cross-contamination if the generator is reused. | Nonconforming Material/Component | The firm initiated their recall on 02/08/2108 by email to distributors, and letter to the remaining consignees.Distributors were directed as follows:"We ask that you immediately:1. Quarantine your inventory of the J-Plasma Handpieceso Provide us with a list of catalog numbers and quatities that you have on your possession using the attached Inventory List template.o Email the information to me at Rubiela.Maldonado@boviemed.comNote: as soon as we receive the information from you we will ship the box labels and insert stuffer notices for the stock that you have on your possession.o After receiving these labels, please follow the attached instructions on how to add each of the insert stuffer notice and the outer box label to each of the boxes that you have on your possession.o Document the number of labels/stuffers added to each of the products on the Inventory List provided, sign and date to confirm this was performed.o Return the signed document to me via email at Rubiela.maldonado@boviemed.com.2. Inform your customers about the correction:o Identify and contact your customers that have received the product affected by this correction.o Provide them with a copy of the Urgent Medical Device Correction Template attached to this letter.o Track and document the customer responses to ensure 100% effectiveness.o Provide us a summary of the effectiveness of this correction (number of customers impacted and number of responses received)."The remaining consignees (users of the device) were instructed as follows:"ACTIONS:At the end of each procedure, inspect the handpiece cable plug and generator receptacle for the presence of fluids. If fluids are present, immediately discontinue use of the generator.In addition to following the J-PlasmaÂż handpiece Instructions for Use, immediately begin to follow the supplemental instructions provided below to ensure that the J-PlasmaÂż handpiece cable plug is properly and co |  | 1005 | 1005 units | worldwide | Open, Classified |  |  | 13934
Z-1147-2018 | 79355 | [6 empty] | DERMABOND" PRINEO" Skin Closure System |  |  | November 27, 2017 | 2017 | Ethicon, Inc. | Ethicon discovered that specific lots of DERMABOND"PRINEO"System may not dry within the specified time after proper application, and thus may fall off. | Component change control | Ethicon sent an URGENT MEDICAL DEVICE RECALL REMOVAL letter dated November 2017 to customers titled: "DERMABONDTM PRINEOTM SKIN CLOSURE SYSTEM (22 CM)". The letter identified the affected product, problem and action to be taken. The letter instructed customers to do not use or distribute any product which is subject to recall. Customers are asked to examine inventory, quarantine products that are subject to recall and contact the firm for return of product. Also, post the recall letter in a visible location at facility. If you have additional questions regarding this recall (removal) or to report any customer complaints, pleasecontact Ethicon Customer Support Center at 1-877-ETHICON (1-877-384-4266). The Customer SupportCenter is open Monday through Friday, 7:30 AM to 6:30 PM ET. |  | 20090 | 20,090 eaches (10,545 units) | US Nationwide Distribution | Open, Classified |  |  | 20090
Z-1152-2018 | 79528 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROWg+ard Blue PLUS catheter, Catalog Number: ASK-45854-NG |  |  | December 21, 2017 | 2017 | Arrow International Inc | Certain lots may be missing information at the bottom of the lid including the lot number, expiration date, and Unique Device Identification (UDI) information. | Process control | Arrow International sent Notification letters to DISTRIBUTORS that include the following instructions: 1. Discontinue distribution and quarantine any affected products. 2. Inspect affected products to identify whether they have the lot number, expiration date, and UDI information. Products missing this information should be returned. 3. Using the provided Customer Letter and Acknowledgement Form templates, communicate this recall to any customers who have received affected products. Customers should return mislabeled products. Products that are properly labeled should not be returned. 4. Any customers with mislabeled products should return them with a completed Acknowledgement Form. 6. If you and customers have no affected stock please complete and return the Acknowledgment Form. Notification letters to RISK MANAGERS include the following instructions: 1. Discontinue use and quarantine any affected products. 2. Inspect affected products and identify whether they have the lot number, expiration date, and UDI information. Products missing this information should be returned to the Distributor. 3. Complete the Acknowledgement Form whether or not there is affected product that must be returned.For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 420 | 420 | United States in the state of FL. | Open, Classified |  |  | 420
Z-1151-2018 | 78999 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate 2550 "High Speed" Catalog Number ODL25255; AutoMate 1250 "Standard" Catalog Number ODL25125. Chemistry analyzer sample processing and sorting system. |  |  | January 18, 2018 | 2018 | Beckman Coulter Inc. | A misaligned pipette tip sensor poses the remote risk of cross-contamination when failing to cause a no-tip error, resulting in pipetting patient samples without pipette tips. | Device Design | A recall notification letter was sent to affected US consignees on 1/19/2018 notifying them of the issue. In addition, the recalling firm is instructing field service engineers to inspect the tip present sensor for misalignment at the next service visit. |  | 73 | 73 units in US | Worldwide distribution. US nationwide, Algeria, Australia, Austria,Belarus, Belgium, Bosnia and Herzegovina,Brazil, Canada, China,Croatia, Cyprus, Czech Republic,Denmark, Finland, France,Germany, Gibraltar, Greece,Hong Kong, Hungary, India,Ireland, Israel, Italy,Korea, Kuwait, Latvia,Lithuania, Macao, Mayotte,Netherlands, New Zealand, Norway,Panama, Poland, Portugal,Russia, Saudi Arabia, Slovakia,Slovenia, South Africa, Spain,Sweden, Switzerland, Taiwan,Turkey, and United Kingdom. | Open, Classified |  |  | 73
Z-1167-2018 | 79396 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | SwishTapered Implant |  |  | January 19, 2018 | 2018 | Implant Direct Sybron Manufacturing, LLC | The extender that is packaged with the SwishTapered" and SwishPlus" Implant may have been packaged in the incorrect position (upside down in the plastic retainer). | Process control | Customers were sent Urgent: Safety Alert/Advisory Notice letters around 01/19/2018. The Notice identified the affected device and the reason for the correction. Instructions for both distributors and doctors include to review inventory for affected product, complete and return the Acknowledgement form within 48 hours of receipt of the letter, follow the included instructions to correct the affected product, and distributors are to notify those customers that may have been shipped the affected product lot and inform them of the issue. Questions should be directed to Implant Direct Sybron Manufacturing LLC Customer Care at 1-888-649-6425. |  | 1117 | 1,117 units total | Worldwide Distribution -- US, Europe, Canada, Australia, and Japan. | Open, Classified |  |  | 1117
Z-1140-2018 | 79223 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD VacutainerÂż EDTA Tan Top Tube; Catalog Number: 367855 (UDI: 50382903678551) |  |  | March 22, 2018 | 2018 | Becton Dickinson &amp; Company | BD Vacutainer EDTA Lavender, Tan, and Pink Top Tubes and BD Vacutainer Lithium Heparin Green Top Tubes can cause an underestimation of lead in blood samples when used with Magellan Diagnostics LeadCare assays, employing the Anodic Stripping Voltammetry (ASV) methodology, or any other assay employing ASV methodology. | Device Design | On March 22, 2018, BD issued a press release. BD intends to distribute URGENT MEDICAL DEVICE CORRECTION notices to their customers in the upcoming days via courier service. The URGENT MEDICAL DEVICE CORRECTION notification has been posted on BD's website along with a link to their press release. BD has updated the IFU for BD Vacutainer EDTA and BD Vacutainer Lithium Heparin Tubes to include the noted precaution for use of Lavender, Tan, Pink and Green Top tubes on Magellan Diagnostics Leadcare assays using ASV methodology. **Distributors are advised to Take the Following Actions:1) This is a Medical Device Correction and there is no need to return or discard product. The product can continue to be used with other non-ASV blood lead level test technologies such as GFAAS and ICP-MS and all assays which do not employ ASV methodology.2) If you have distributed this product, please identify your customers and notify them of this recall by sharing a copy of the correction notice.3) Complete the attached Customer Response Form and return to the BD contact noted on the form whether or not you have any of the covered product.4) If you would like BD to conduct the notification to your customers, please email your customer list within 3 business days to BDPASRC@bd.com. Please include the contact name, address, phone number, email address, and/or fax numbers for each customer.Distributors who require further assistance, please contact BD at 1-888-237-2762 (select Option #3 and then Option # 4) between 8AM and 5 PM, CT Monday through Friday.**Medical Directors, Risk Managers, Medical Device Safety Officers &amp; Lab Managers are advised to Take the Following Actions:1) This is a Medical Device Correction and there is no need to return or discard product. The product can continue to be used with other non-ASV blood lead level test technologies such as GFAAS and ICP-MS and all assays which do not employ ASV methodology.2) Continue to follow FDA's safety |  | 3956900 | 3,956,900 | Distributed in 46 states: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NM, NY, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and Puerto Rico. | Open, Classified |  |  | 1302570453
Z-1169-2018 | 79006 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate 1200 Catalog Number ODL25120; |  |  | January 19, 2018 | 2018 | Beckman Coulter Inc. | Belts on the x-axis (and on the y-axis, to a lesser extent) manufactured from 2015 are showing premature failure, posing the remote risk of cross-contamination. | Device Design | Beckman Coulter plans to replace the belt via Service Modification during the next planned service maintenance visit. At that time, the customer will be informed of the belt replacement by the service engineer. |  | 98 | 98 units in the US | Australia, Austria, Belarus,Belgium, Bosnia and Herzegovina, Brazil,Canada, China, Croatia,Cyprus, Czech Republic, Denmark,Finland, France, Germany,Gibraltar, Greece, Hong Kong,Hungary, India, Ireland,Israel, Italy, Korea,Kuwait, Latvia, Lithuania,Macao, Mayotte, Mexico,Netherlands, New Zealand, Norway,Panama, Poland, Portugal,Qatar, Russia, Saudi Arabia,Singapore, Slovakia, Slovenia,South Africa, Spain, Sweden,Switzerland, Taiwan, Turkey,United Kingdom, United States | Open, Classified |  |  | 98
Z-1170-2018 | 79560 | No | PMP |  |  | Neurology | PMA | Proclaim DRG Implantable Pulse Generator, Model Number 3664 |  |  | March 09, 2018 | 2018 | St. Jude Medical, Inc. | The firm received complaints of error messages that occurred during routine impedance checks on Proclaim DRG IPGs, model 3664. Some complaints were also associated with transient over stimulation which created discomfort for the patients. | Software change control | Notification material in the form of a Customer Letter was hand-delivered to twoconsignees in the US beginning 09 Mar 2018. The Physician Letter was hand-delivered to impacted implanting physicians, beginning 14 Mar 2018. A Patient Letter was delivered via certified mailing to all active implanted patients in the US, beginning 14 Mar 2018. |  | 17 | 17 units | US | Open, Classified |  |  | 17
Z-1173-2018 | 79399 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Comprehensive Segmental Revision System (SRS) |  |  | February 28, 2018 | 2018 | Zimmer Biomet, Inc. | Zimmer Biomet received reports that during assembly, an SRS Humeral Distal Body implant failed to engage correctly at the point of reduction with a Discovery Elbow System Humeral Condyle. A mismatch occurred between the screw holes in the Distal Humeral Body &amp; the Discovery Elbow Condyle implant during vendor machine operation. The condition would be noticed during assembly and therefore, if the product has already been implanted, no further action is required. The mismatch does not always prevent engagement of the screw, based on complaints received it is estimated to prevent engagement of the screw in less than 1% of cases. | Under Investigation by firm | An URGENT MEDICAL DEVICE RECALL letter will be issued to customers titled "Affected Product: Comprehensive Segmental Revision System (SRS)" instructing customers to do the following:Risk Manager Responsibilities:1.Review this notification and ensure that affected personnel are aware of the contents.2.If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3.Complete Attachment 2  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4.Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5.If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  |  |  | Nationally | Open, Classified
Z-1179-2018 | 79297 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | STAGE-1Âż RDS COC Abutment Plastic Sleeve, Non-Lock, Multi-Unit; Catalog Number: S2437-01-1K |  |  | October 16, 2017 | 2017 | Keystone Dental Inc | A locking plastic abutment sleeve, used by dental technicians in dental laboratories to aid in the creation of the final restorative crown, was packed as a non-locking plastic abutment sleeve. | Labeling mix-ups | The firm, Keystone Dental, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 10/24/217 to its consignee. The letter described the product, problem and actions to be taken. The consignee was instructed to do the following: Return any product remaining in inventory to Keystone Dental and replacement product will be sent to you. Note: These lots were the only affected product and all other inventory you may have can be used. Complete and return the Customer Relations Fax Back Form to Customer Relations (1-866-903-9272-US only/ 1-781-328-3400).If you have any questions, please contact Customer Relations Manager at:Toll Free 866-902-9272; Fax 866-903-9272 and CustomerSupport@KeystoneDental.com. |  | 55 | 55 | International Distribution to: Germany only. | Open, Classified |  |  | 55
Z-1174-2018 | 79337 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | MINI INFANT &lt; 3KG, RADIOTRANSLUCENT ELECTRODE, Universal/Anderson Connector, REF/Catalog Number 2602 |  |  | February 23, 2018 | 2018 | ConMed Corporation | Wire sets of certain devices were assembled with adhesive on the contact wires, which may impede conductivity and could cause failure | Process design | The firm, ConMed,issued an "URGENT: DEVICE RECALL" notice dated 2/28/2018 and business reply forms to customers on 2/28/2018 via courier service. The notice described the product, problem and actions to be taken. The Customers were instructed to do the following:Step 1: Review your inventory for any of the devices with the lot codes listed on the notice. Customers are encouraged to contact all of those departments within your facility and any other facilities within your organization that may have received affected products. It is imperative that all end users of these devices receive this notice and respond immediately.Step 2a: If you HAVE inventory of any of the devices with the lot codes listed on the notice, complete the business reply form and return it with the devices to: CONMED Corporation 525 French Road Utica, NY 13502 USA Attn: Ed KovacStep 2b: If you DO NOT HAVE any affected devices to return, please complete the business reply form, indicating you have no devices and return by one of the means listed below:1. Email to: padpro2602@conmed.com2. Fax to: Field Action Support Team at 315-624-3225.3. US mail, using enclosed postage-paid envelope, to: CONMED, 525 French Road, Utica, NY 13502; Attn: Field Action Support TeamIf you have any questions or requests, please don't hesitate to contact the Field Action Support Team at 1-800-448-6506, fax to 315-624-3225, or email padpro2602@conmed.com. |  | 18720 | 18,720 units in total | Worldwide Distribution: US (Nationwide) and countries of: New Zealand, Saudi Arabia, Switzerland, United Arab Emirates Uruguay &amp; Republic of Macedonia (The former Yugoslav Republic of Macedonia). | Open, Classified |  |  | 18720
Z-1180-2018 | 78909 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD VacutainerÂż Barricor" Lithium Heparin Plasma Blood Collection Tubes for Invitro Diagnostic Use, Catalog Numbers: 365043, 365044, 365045 &amp; 365047 |  |  | September 18, 2017 | 2017 | Becton Dickinson &amp; Company | Increased amount of residual blood present on the top of stopper well after venipuncture | Nonconforming Material/Component | On September 22, 2017, BD distributed URGENT PRODUCT REMOVAL RECALL notices to distributors via email. Distributors and customers were mailed via USPS. Customers are advised to take to take the following actions:1. Discontinue use of the tubes.2. Complete and return the Customer Recall Response Form via fax to 1-855-544-4851 or via email to BD7021@stericycle.com. 3. Return all affected products using the enclosed packaging instructions and include the completed Recall Response Form with the product. Upon receipt of the returned product, BD will issue alternate product. Customers with questions may call 1-855-215-5168, Monday through Friday between 8AM and 5 PM, EST.For further questions, please call (201) 847-6800. |  | 1976000 | 1,976,000 tubes | USA (nationwide) | Terminated | February 04, 2019 | 504 | 1976000
Z-1186-2018 | 79471 | No | MQF | Assisted reproduction catheters. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Echosight Jansen-Anderson Intratubal Transfer Set, Catalog Number: K-J-JITS-572900, Global Product Number: G17946 |  |  | December 19, 2017 | 2017 | Cook Inc. | Certain lots of in vitro fertilization (IVF) products were distributed without completed finished device quality testing with potential adverse events that may include nonviable gametes and/or nonviable embryos. | Process control | Cook Medical sent an Urgent Medical Device Recall dated December 19, 2017 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to:1. Examine inventory immediately and quarantine affected product(s).2. Return products to Cook with the Acknowledgement and Receipt Form to receive a product credit.3. Complete and return the Acknowledgement and Receipt Form within 5 business days of receiving this letter. 4.Report adverse events to Cook.For questions contact Cook Medical Customer Relations at 800-457-4500 or 812-339-2235. |  | 57 | 57 | Worldwide Distribution - U.S. Nationwide in the states of: CA, CO, FL, IN, MA, NE, NJ, OH, TN, TX, and WI. and countries of Belarus, Belgium, Brazil, Canada, Chile, China, Egypt, Italy, Japan, Poland, Saudi Arabia, Spain, and United Kingdom. | Open, Classified |  |  | 2451
Z-1184-2018 | 79412 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista System LOCI 8 CALIBRATOR |  |  | November 09, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc. | An increase in the number of failed calibration events or negative bias with Quality Control (QC) and patient samples when using Dimension Vista LOCI 8 Calibrator lot 7CD076 to calibrate Estradiol (E2) Assay. | Under Investigation by firm | The firm, Siemens, issued an "Urgent Medical Device Correction (VC18-01.A.US)" letter dated 11/3/2017, to US customers on November 3, 2017 and an "Urgent Field Safety Notice (VC18-01.A.OUS)" to customers outside the US in November 2017. Customers were instructed to:"Discontinue use of and discard impacted LOCI 8 CAL lot 7CD076"Calibrate the E2 assay using an alternate lot of LOCI 8 CAL. "Review the letter with their Medical Director. "Review their inventory of this product to determine theirs laboratory's replacement needs and to provide information to Siemens for reporting to the regulatory authorities."Complete and return the Field Correction Effectiveness Check Form attached to the UMDR/UFSN letter within 30 days via Fax to Customer Care Center at (312) 275-7795 or send scanned copy via email to uscctsfcaefax.healthcare@siemens.com. " Retain this letter with your laboratory records, and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center, your local Siemens Technical Support Representative or call recall contact at 800-441-9250; email: heydi.calderon@siemens-healthineers.com. |  | 304612 | 989 (681 USA/308 OUS) | Worldwide Distribution: US (nationwide) and countries of: Austria, Belgium, Canada, Denmark, France, Germany, Italy, Netherlands, Slovakia, Spain and Switzerland. | Open, Classified |  |  | 304612
Z-1182-2018 | 79402 | No | MNL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | First Step Flexible Endoscope Bedside Pre-Clean Kit, EP-4. Intended for pre-clean of outer surface of flexible scopes and surgical instruments to remove gross contamination prior to manual cleaning and disinfection. Product Labeling: "The multi-tiered enzymatic detergent starts working instantly to break down all forms of bioburden.". |  |  | November 03, 2017 | 2017 | Madison Polymeric Engineering | Affected lots may have exceeded its microbial limits. | Process design | Cygnus sent an Recall Letter dated November 3, 2017.Actions Required:1. Immediately discontinue distribution of First Steps EP-4 and EP-4D with the referenced lot numbers above (as applicable).2. Determine if any stock of the First Steps EP-4 and EP-4D with the referenced lot numbers above (as applicable) remains in your inventory.3. Please complete the attached form and return to Cygnus Medical as indicated. It is important to notify Cygnus Medical whether or not your facility has any inventory of this product.4. Please notify any of your customers who may have received the lots in the tables above between November 28, 2016 and January 11, 2017 with the letter provided.Once Cygnus Medical has received your completed form, if you have indicated you have any remaining product from these lots, we will issue a Return Goods Authorization, call tag, and/or FedEx/UPS account number.If you have questions regarding this matter, please contact Customer Service at 800-990-7489 or Quality Control at 203-488-4554 at x105 during business hours 8:30am-5:00pm Eastern Standard Time. |  | 285 | 285 cases | US Distribution. . | Terminated | July 17, 2018 | 256 | 285
Z-1195-2018 | 79375 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Nucleic Acid test intended for screening for MRSA/SA. Each kit has 20 tubes. in vitro diagnostic. |  |  | October 27, 2017 | 2017 | Roche Molecular Systems, Inc. | The cobas MRSA/SA tests may have decreased performance compared to the analytical sensitivity for MSRA detection. A screening failure can lead to potential infection in the colonized patient being tested as well as the spread of MSRA to others. The United States is not impacted, as the cobas MRSA/SA nucleic acid test for use on the cobas liat system is not available in the US. | Software design | The cobas MRSA/SA tests has the potential for decreased performance compared to the analytical sensitivity for MSRA detection. A screening failure can lead to potential infection in the colonized patient being tested as well as the spread of MSRA to others. |  | 136 | 136 kits | Worldwide distribution. Austria, Denmark, Germany, Singapore, Netherlands, Hong Kong, and Switzerland. US product is not affected. | Open, Classified |  |  | 136
Z-1183-2018 | 79271 | No | PSZ | Influenza virus antigen detection test system. | Microbiology | Microbiology | 510(k) | Sofia 2 Analyzer Catalog Number 20299 |  |  | December 15, 2017 | 2017 | Quidel Corporation | The calibration for the affected batch of analyzers was incorrectly set during the manufacturing process which may lead to a failure when calibration is first performed. There may also be e remote possibility that incorrect qualitative test results for patient samples near the assay cutoff may be reported. | Software Manufacturing/Software Deployment | Quidel sent and Urgent Product Recall letter dated December 15, 2017.The letter instructs the user to discontinue use of the device until new firmware can be installed. For further questions, please call (800) 874-1517. |  | 482 | 482 units | US Distribution | Open, Classified |  |  | 482
Z-1181-2018 | 79452 | No | FOX | Infusion stand. | General Hospital | General Hospital | 510(K) Exempt | Rotating IV Pole |  |  | December 01, 2017 | 2017 | Ohmeda Medical | There is the potential for the Rotating IV Pole to fall. A fall could result in an injury to a bystander holding the patient or to the operator. | Device Design | On December 1, 2017 an URGENT MEDICAL DEVICE CORRECTION letter was issued to customers notifying of the recall. The letter requested customers to discontinue use of the affected Rotating IV poles and remove from Service. Place affected Rotating IV poles in a quarantine marked area. A GE Service Representative will contact you to arrange for the removal &amp; replacement of a new IV pole. If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. For further questions, please call (410) 888-5376. |  | 1179 | 1,179 (1,004 US; 175 OUS) | USA (nationwide) | Open, Classified |  |  | 1179
Z-1199-2018 | 79493 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | DATEX-OHMEDA AISYS CS2 ALADIN2 CASSETTE |  |  | October 04, 2017 | 2017 | GE Healthcare, LLC | Non-Conforming devices were shipped from the service center without proper evaluation and therefore may be nonconforming. | Process control | The firm, GE healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 4th, 2017 to all their customers. The letter described the product, problem and actions to be taken. The customers were instructed to contact GE Service to obtain an RMA number. GE Healthcare has arranged with Piramal Inc. to replace the affected products at no cost. If you have any questions or concern regarding this notification please contact GE healthcare Service at 1-800-437-1171 or your local representative. |  | 2 | 2 | US Distribution to states of: AZ, IL NC, NH and PA. | Open, Classified |  |  | 30
Z-1193-2018 | 79301 | No | GDT | Removable skin staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Teleflex MEDICAL WECK Visistat 35. Used for the external skin closure of lacerations. |  |  | February 15, 2018 | 2018 | Teleflex Medical | Some boxes are labeled as Visistat Wide devices, when in fact the contain Visistat Regular devices. The 6 individually packaged units within the boxes are correctly labeled. | Packaging process control | The firm initiated the recall on 02/14/2018 by letter and followed up with another letter on 03/12/2018 to clarify the labeling issue. The letter to the distributor stated:"1. Immediately discontinue distribution and quarantine any products with the catalog and lot number listed above.2. Inspect affected products within your control to identify if the box label correctly identifies the products inside the box. Products with an incorrect label on the box should be returned to Teleflex Medical so that they are not distributed to health care providers in error.3. Using the provided Customer Letter and Acknowledgement Form templates, communicate this recall to any of your customers who have received product included within the scope of the recall. Customers may return mislabeled products if they choose to do so. Note the attached Customer Letter template was updated since this letter was first sent to you.4. If any of your customers wish to return affected products, have them return the affected products to you, together with a completed Acknowledgement Form, for consolidation and return to Teleflex Medical. In the event that an alternative approach is needed, contact Teleflex Customer Service for more information at 1-866-246-6990."The customer letter requested discontinuation of use and quarantining of products and return of affected stock. |  | 12 | 12 units | KS | Terminated | October 31, 2018 | 258 | 12
Z-1200-2018 | 79568 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound System. The firm name on the label is Siemens Medical Solutions USA, Inc., Buffalo Grove, IL. |  |  | March 05, 2018 | 2018 | Siemens Medical Solutions USA, Inc. | The ECG signal may flatline due to electromagnetic interference during the use of electrosurgical equipment. | Device Design | Undated recall letters were issued via email (if they had an email address for the customer) or via certified mail (if no email address was available) on 3/5/2018. |  | 229 | 229 systems | Distribution was made to medical facilities in AZ, CA, CT, FL, IL, IN, KS, KY, MI, MN, MO, NJ, NY, NC, OH, PA, TX, VA, and WA. There was no military/government distribution.Foreign distribution was made to Argentina, Australia, Austria, Belgium, Bosnia Herzeg, Colombia, Croatia, Denmark, France, Germany, India, Ireland, Israel, Italy, Japan, Mexico, New Zealand, Philippines, Poland, Republic Korea, Romania, Singapore, Slovakia, Spain, Sweden, Taiwan, Thailand, Turkey, UAE, and Vietnam. | Open, Classified |  |  | 229
Z-1224-2018 | 79394 | No | JHW | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PERSONAÂż THE PERSONALIZED KNEE SYSTEM "PSN TIB STM 5 DEG SZ D R" indicated for patients with severe knee pain and disability due to: - Rheumatoid arthritis, osteoarthritis, traumatic arthritis, polyarthritis. |  |  | December 12, 2017 | 2017 | Zimmer Biomet, Inc. | A complaint was received for Item 42-5320-067-02, lot 63236047. It was report that there was no set screw or taper plug included in the box. Complaint investigation confirmed that the entire lot of 63236047 was missing the set screw and taper plug. | Under Investigation by firm | On December 12, 2017 an URGENT MEDICAL DEVICE RECALL LETTER was issued to customers will the following guidance:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.Questions or concerns can be directed to 1-574-371-3071. |  | 3 | 3 units | Non-USA Customer | Open, Classified |  |  | 3
Z-1216-2018 | 79335 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | The Stryker Thoracic Pedicle Feeler 6002-350-000 is a smart instrument for remote control of |  |  | October 16, 2017 | 2017 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments and Stryker Leibinger GmbH &amp; Co. KG arevoluntarily recalling certain serial numbers Thoracic Pedicle Feelers that are used in conjunction with the StrykerÂż NavigationSystem because the settings onthe machine were higher than intended. This change led to increased energy directed onto the tip of the Thoracic PedicleFeeler which caused material degradation where the laser marks exist. There is the potential for the tip of the Thoracic Pedicle Feeler to break. | Software Manufacturing/Software Deployment | US customers will be notified via FedEx overnight. US Sales Representatives will be notified via email. International entities will be notified via an internal notification system. International customers will be notified per local regulations/procedures.Customer Instructions:Actions to be taken by the Customer/User:1. Immediately review this Recall Notification.2. Immediately check all stock areas and/or operating room storage for affected equipment. Quarantine anddiscontinue use of any affected Thoracic Pedicle Feelers.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this Notification and identify how manyaffected items are currently in your inventory. Please complete and return the BRF even if you don't have any affected product on hand. Note: Your signature on the BRF indicates that you received and understand this Notification and have followed theinstructions in the Notification.4. If you have further distributed this product, please forward this Notification and the attached BRF to all affected locations. Please indicate each location on the BRF.5. If the BRF for your facility indicates that recalled product is currently on hand, we will contact you to arrange for the return of the product for repair/replacement.If discontinuing use of the specific Thoracic Pedicle Feeler(s) listed will cause disruption at your facility. For further questions, please call (269) 323-7700. |  | 28 | 28 | International , US and US Govt. | Open, Classified |  |  | 28
Z-1223-2018 | 79425 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Vital-Port Vascular Access System, Subcutaneous, implanted, intravascular infusion port and catheter, IP, IP-S, |  |  | November 28, 2017 | 2017 | Cook Vascular Inc. | Non-coring needle provided with the Cook Vital-PortÂż Vascular Assess System (Vital-Port) may cut or dislodge a core or sliver of material from the Vital-Port septum when the non-coring needle is inserted into the Vital-Port. This needle is used on the initial implant of the Vital-Port. Vital-Port products that have been successfully placed in patients are not impacted by this recall. | Nonconforming Material/Component | Removal for all non-expired lots of the Cook Vital-Port Vascular Access System. Written recall communication letters will be mailed via courier to US Consignees.plan to STORE recalled product(s)? (if returned)Recalled products will be returned to global regional distribution warehouses to be quarantined in a marked receiving area of returned goods.8.9 Provide plan for final DISPOSITION for recalled product(s)?All recalled products will be destroyed at the regional distribution warehouses. A certificate of destruction will be provided upon completion of destruction prior to the completion of the recall. Products remaining in the control of Cook will be assessed for rework to replace the non-coring needle. |  | 31006 | 31,006 Units | The devices have been distributed within the United States and the following countries: United Arab Emirates, Argentina, Austria, Australia, Barbados, Belgium, Bulgaria, Bermuda, Brazil, Canada, Switzerland, Colombia, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Spain, Finland, Faroe Islands, France, United Kingdom, Guatemala, Hong Kong, Hungary, Indonesia, Ireland, Iceland, Italy, Jamaica, Jordan, Japan, Kenya, South Korea, Kuwait, Lebanon, Latvia, Morocco, Macedonia, Mexico, Netherlands, New Zealand, Oman, Panama, Poland, Portugal, Serbia, Saudi Arabia, Sweden, San Marino, El Salvador, Turkey, Trinidad &amp; Tobago, Ukraine, Uruguay, Venezuela, Virgin Islands, Kosovo, China, and South Africa. | Terminated | October 03, 2018 | 309 | 31006
Z-1208-2018 | 79371 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Rechargeable Li-ion battery used with Puritan Bennett TM 980 Series Ventilator System including Pediatric-Adult Ventilator, Neonatal Ventilator, Universal Ventilator; Battery Product Number: 10086042; UDI: 20884521700861 |  |  | November 21, 2017 | 2017 | Covidien LLC | Rechargeable lithium-ion batteries with incorrect firmware that are used in certain Puritan Bennett 980 ventilators may not fully charge after installation. | Process control | Consignees are instructed to immediately notify all care environments in which the PB980 ventilators are used about the notification. If the facility has distributed PB980 ventilators to other persons or facilities, they are asked to please promptly forward a copy of the letter to those recipients. The consignee is asked to complete an attached form and return it as directed to confirm receipt and understanding of the notification.The firm developed a tool to inspect and identify the firmware on the batteries. Medtronic service engineers will inspect batteries at consignee facilities during scheduled or preventive maintenance. Batteries found to have the incorrect firmware will be replaced.For further questions, please call (203) 492-5415. |  | 15885 | 15,885 | Worldwide Distribution - USA (nationwide) Distribution to the states of : AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, WY, and the District of Columbia. and to the countries of : Argentina, Chile, Mexico, Panama, Canada, Australia, New Zealand, Austria, Bahrain, Belgium, Bulgaria, Croatia, Czech Republic, Egypt, Estonia, France, Germany, Greece, Hungary, Ireland, Italy, Jordan, Kuwait, Latvia, Lebanon, Lithuania, Morocco, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom, Israel, Netherlands, Tunisia, Brunei, Indonesia, Malaysia, Myanmar, Philippines, Singapore, Thailand, and Vietnam | Open, Classified |  |  | 15885
Z-1213-2018 | 79455 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Sonimage HS1 Ultrasound Kit AC adapter, Product Number: A8WA, UDI: 04560141944534 |  |  | December 04, 2017 | 2017 | Konica Minolta Medical Imaging USA, Inc. | There is a risk of circuit board breakage within the main housing of the AC Adapter unit. Smoke may be emitted from the AC Adapter as a direct result. | Device Design | Konica Minolta is preparing for the final AC Adapter replacement. Upon receipt of the new AC Adapter, owners of SONIMAGE HSl System will be asked to replace the existing AC Adapter. Additionally, instructions for the return of the current AC Adapter will be provided. |  | 85 | 85 | Distributed in the following U.S. states: AL, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, WA, WI and the District of Columbia. Distributed in Canada and Brazil. | Open, Classified |  |  | 220
Z-1236-2018 | 79487 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Phoenix Nail System, intended for use with an intramedullary nail system for alignment, stabilization and fixation of fractures. 3.5mm Inserter Connector, Short; Model Number: 14-441045; Kit Model Numbers: 14-400082, 14-400082S, 14-440065S, 14-442000-00, 14-442000S, 14-444000-00, 979105, 14-441045-00, and 14-444000S |  |  | March 12, 2018 | 2018 | Zimmer Biomet, Inc. | Certain lots of the Phoenix Tibia Nail 3.5mm Inserter instruments are being recalled due to reports of fracturing and remaining in the device implant. | Device Design | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL (REMOVAL)" notice dated 3/12/2018 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to do the following:Risk Manager:1.Review this notification and ensure affected personnel are aware of the contents.2.Assist Zimmer Biomet sales representative and quarantine all affected product.3.Zimmer Biomet sales representative will remove the affected product from the facility.4.Complete Certificate of Acknowledgement.a.Return a digital copy to corporatequality.postmarket@zimmerbiomet.com. Distributors: 1.Review this notification and ensure that affected team members are aware of the contents.2.Locate and quarantine affected product in your inventory.3.Complete Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days. 4.Immediately return all affected product from your distributorship and from affected hospitals within your territory. Mark RECALL on the outside of the returned cartons. 5.Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.comEXPANSION: The firm expanded the recall to include additional lots and the new notification sent 3/23/18, also explains the additional issue that devices were etched with a supplier lot number and the packages were etched with the firm lot number. Both numbers are included in an Appendix to the notification letter. |  | 2470 | 2470 | Worldwide Distribution: US (nationwide) AK, AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and Puerto Rico; and countries of: Australia, Brazil, Canada, Chile, Colombia, El Salvador, Netherlands, Japan, Costa Rica, Spain, and Thailand. | Open, Classified |  |  | 2470
Z-1210-2018 | 79152 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ACUVUE ADVANCE Brand Contact Lenses with HYDRACLEAR |  |  | October 19, 2017 | 2017 | Johnson &amp; Johnson Vision Care, Inc. | Some contact lenses were associated with complaints of a foreign matter caught between the blister package and foil. | Packaging process control | On October 19, 2017, Johnson and Johnson sent letters to their customers stating the recall of ACUVUE and given the instructions to review their inventory, stop using the affected lots, use the enclosed label to return any affected products, contact customer service to arrange the return, even if they do not have any recall product inventory. A contact was provided to each individual letter. |  | 165750 | 165,750 | Worldwide distribution. US nationwide including Puerto Rico, Korea, Czech Republic, China, Japan, Great Britain, Austria, Malaysia, Hong Kong, Russia, Taiwan, Singapore, India, and Brazil | Open, Classified |  |  | 347583
Z-1222-2018 | 79380 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | QUANTA-Lyser 240 EIA, Part #GS0241. |  |  | January 26, 2018 | 2018 | Inova Diagnostics Incorporated | The device power supply was not manufactured according to specification which can result in the loss of the protective grounding of the instrument. | Nonconforming Material/Component | The letters were issued via email on January 26, 2018, informing the consignee of the issue and the actions to be taken. The response forms were requested to be completed and returned by February 23, 2018, indicating receipt of the notification and confirmation that the power supply has been replaced.The firm reported on February 23, 2018 that responses have been received from both consignees. For further questions, please call (858) 586-9900 x 1381. |  | 1 | 1 instrument | There was no U.S. distribution. Distribution was made to Panama and United Kingdom. | Terminated |  |  | 1
Z-1237-2018 | 79483 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Signia" Power Handle (SIGPHANDLE, SIGNVAL1 and SIGVAL20 - Kits) |  |  | December 13, 2017 | 2017 | Covidien LLC | Product reportedly shutting down during use. | Component design/selection | The firm, Medtronic, sent a "Field Corrective Action Notice" dated 12/13/2017 to all consignees via Federal Express or certified mail on 12/13/17. The notice described the product, problem and actions to be taken. Customers were instructed to: immediately notify all users of Signia power handles of this important notification; follow the User manual and Instructions for Use; If your facility has distributed Signia" power handles to other persons or facilities, please promptly forward a copy of this letter to those recipients, and complete and return the Field Corrective Action Acknowledgement Form via Fax to (203) 492-7719 or email to: FCAMITG@Medtronic.com. All in-control inventory and product at customers locations will have the software update installed.If you have any questions regarding this communication, please contact your Medtronic representative or Customer Service at 800-962-9888, option 2. |  | 2121 | 2121 | Worldwide Distribution: US (Nationwide). | Open, Classified |  |  | 2121
Z-1206-2018 | 79454 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | Johnson &amp; Johnson BAND-AIDÂż Brand First Aid Products HURT-FREEÂż Wrap (size: 2in) |  |  | December 18, 2017 | 2017 | Johnson &amp; Johnson Consumer, Inc. | Johnson &amp; Johnson Consumer Inc. is initiating a voluntary, Class II recall at the retail level in the U.S. of select COACHÂż Self-Adhering Sports Wraps, BAND-AIDÂż Brand First Aid Products SECURE-FLEXÂż Wraps and BAND-AIDÂż Brand First Aid Products HURT-FREEÂż Wraps products. The current labeling claim states not made with natural rubber latex which needs to be updated based upon recent awareness that natural latex was used as a base ingredient in the early-stage manufacturing process, which reduces the allergic protein found in natural latex. | Labeling design | On December 18, 2017 an URGENT MEDICAL DEVICE RECALL letter was issued to retailers and warehouses with intentions to notify all affected individuals, including customers of the recall. The recall letter requests that the following actions be taken:1. Immediately examine your inventory and quarantine the above-mentioned product subject to recall. In addition, if you may have further distributed this product to retailers, warehouses or distributors, please identify and send a list of these locations to Stericycle at BandAid8687@stericycle.com and they will be notified.2. Please complete the enclosed Business Reply Form and return immediately by emailing to BandAid8687@stericycle.com or faxing to (1-888-965-5802).3. If you have any questions regarding the form, please call Stericycle at (1-855-215-5023).Please return ALL inventory of the identified product above. No product other than that specified is to be returned.Questions or concerns can be directed to 1-855-215-5023 |  |  |  | Nationally | Open, Classified
Z-1217-2018 | 79366 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | IMRIS ORT200 Removable Operating Room Table, 50-60 Hz Version is intended for use during diagnostic examinations or surgical procedures to support and position a patient. |  |  | February 23, 2018 | 2018 | Deerfield Imaging, Inc. | The rotational locking mechanisms on the ORT 200/300 tables fail to operate such that it cannot change states between locked and unlocked. | Under Investigation by firm | The firm disseminated their notices on 02/23/2018 by mail and followed with an email on 02/26/2018. The notice stated:"It has recently come to our attention that the ORT200/300 rotational lock fails to operate such that it cannot change states between locked and unlocked. This may pose a risk of injury to the patient.To minimize the risk of injury, IMRJS is taking the following corrective actions:" ORT200 and 300 tables: all customer sites with these tables shall have the rotational lock operation verified and adjusted, if necessary.IMRIS has identified that your facility has one or more of these OR Tables which require evaluation. Please confirm details on page 2 and return your response to IMRIS. IMRIS Customer Service will contact you to implement this corrective action.IMRIS recommends these preventive actions:" Per the OAT Operator Manuals the table should be tested daily in all standard positions and inspected for anomalies including the fit of the covers, tabletops, headrest, pendant connection and rotation." During usage, verify the table is locked by manually pushing the table to make sure it does not move rotationally." If the rotational lock is observed to be not functioning properly, please contact IMRIS Customer Service.The appropriate competent authorities have been notified of this corrective action. IMRIS is committed to patient and user safety, and apologizes for any inconvenience that this may cause. Please contact customer support via our service line at 1-866-475-0525 or Email: customersupport@imris. com if you have any questions or concerns." |  | 17 | 17 units | US, Canada, China, Sweden, Germany, Japan, Qatar, France, Audtralia | Open, Classified |  |  | 30
Z-1242-2018 | 79423 | No | BZK | Monitoring spirometer. | Anesthesiology | Anesthesiology | 510(k) | Gas Module 3 monitoring spirometer, Part Number: 115-030108-00 |  |  | December 01, 2017 | 2017 | Mindray DS USA, Inc. dba Mindray North America | The firm made multiple changes to the product requiring the submission of a new 510(k) premarket notification. | Device Design | Mindray sent an Safety Notification letter dated December 1, 2017 to all customers. The letter identified the affected product, problem and actions to be taken. For question call 201-995-8407. |  | 1719 | 1719 | Worldwide Distribution - U.S. Nationwide in the states: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WY, including Puerto Rico and distributed to 30 foreign consignees. | Terminated | August 29, 2018 | 271 | 1719
Z-1267-2018 | 77160 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyCAL HbA1c Calibrator; REF 10658 |  |  | April 11, 2017 | 2017 | Medica Corporation | Patient results determined using this calibrator lot might deviate up to 20%. Controls might not indicate false values. | Process design | Consignees were notified of the recall via email bulletin dated 4/11/17. The bulletin informs customers of the issues and explains the actions to be taken. |  |  |  | USA: MIFLCATXIndia, Russia, Puerto Rico, N. Mariana Island/Palau. | Open, Classified
Z-1238-2018 | 79383 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona Partial Knee spacer block instruments; (1) 8mm block, Model Number: 42-5399-035-08; (2) 9mm block, Model Number:42-5399-035-09; (3) 10mm block, Model Number: 42-5399-035-10; (4) 12mm block, Model Number: 42-5399-035-12; (5) 14mm block, Model Number: 42-5399-035-14 |  |  | November 29, 2017 | 2017 | Zimmer Biomet, Inc. | Certain partial knee implant spacer blocks do not fit in the alignment tower instrument or handle. | Nonconforming Material/Component | Zimmer Biomet sent an Urgent Medical Device Correction Recall letter dated November 29, 2017. RISK MANAGER PHASE I: 1. Review this notification and notify affected personnel of the contents. 2. Product does not need to be returned at this time. You will receive a follow-up notice when the removal phase is initiated (Phase 2). If you identify a spacer block that will not assemble with these instruments, and you are unwilling to use the spacer block without this functionality, the spacer block may be returned to your sales representative. 3. Complete Certificate of Acknowledgement (Phase 1) and send to CorporateQuality.PostMarket@zimmerbiomet.com. SURGEON PHASE I: 1. Review this notification for awareness of the contents. 2. The spacer block may be used without the alignment tower and inserted and removed by hand from the joint, without the Persona Partial Knee handle. If you identify a spacer block that will not assemble with these instruments, and you are unwilling to use the spacer block without this functionality, the spacer block may be returned to your sales representative. 3. There are no specific patient monitoring instructions related to this recall that are recommended beyond your existing follow-up schedule. 4. Complete Certificate of Acknowledgement (Phase 1) and send to CorporateQuality.PostMarket@zimmerbiomet.com. DISTRIBUTOR PHASE I: 1. Review this notification and notify affected team of the contents. 2. Product does not need to be returned at this time. You will receive a follow-up notice when the removal phase is initiated. 3. Should you or your account prefer to perform a functional check, this may be done by using the instructions in the surgical technique for assembling the alignment tower and handle to the spacer block. If you or your account identify a spacer block that will not assemble with these instruments, and the user is unwilling to use the spacer block without this functionality, the spacer block may be returned as a complaint to Pr |  | 1023 | 1023 | USA (nationwide) Distribution to the states of : AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI., and Internationally to Japan and Netherlands. | Open, Classified |  |  | 1023
Z-1260-2018 | 79241 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Altivate Anatomic(TM) INSTRUMENTATION, PEGGED GLENOID TRIAL, SIZE 54mm, REF 804-07-540 |  |  | January 29, 2018 | 2018 | Encore Medical, Lp | It was discovered during a surgery that the design of the radius at the base of the pegs of the Anatomic Glenoid Trial (Part Numbers: 804-07-380, 804-07-420, 804-07-460, 804-07-500, and 804-07-540) is much larger on the trails than on the implants | Device Design | DJO Global sent an Urgent Field Safety Notice letter dated January 29, 2018, to affected customers. The letter identified the affected product, problem and actions to be taken. For questions call (512) 834-6255 |  | 93 | 93 units | US Nationwide Distribution | Open, Classified |  |  | 461
Z-1248-2018 | 79418 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Advanced Cement Mixing (ACM) Kit w/ Femoral Breakaway Nozzle &amp; Prox. Med. Press., Product Number: 0306-573-000, UDI: (01)4546540055415 |  |  | November 09, 2017 | 2017 | Stryker Instruments Div. of Stryker Corporation | Lack of sterility assurance as bioburden levels were found to be higher than internal acceptable rates. | Environmental control | Letters dated November 16, 2017, were sent to customers. Actions to be taken by the Customer/User:1. Immediately review this Recall Notification.2. Check all stock areas and/or operating room storage to determine if any devices from the affected lots are at your facility.3. Quarantine and discontinue use of the recalled devices.4. Complete and sign the enclosed Business Reply Form (BRF).5. If you have further distributed these products, please forward this letter and the attached BRF to all affected locations. Please indicate each location on the BRF.6. Return the completed BRF to Stryker Instruments Post Market Regulatory Compliance via fax (866-521- 2762) or email (erin.bissonnette@stryker.com).7. If the BRF for your facility indicates that recalled product is currently on hand, a FedEx label will be emailed to you. This shipping label should be used to return recalled product. Upon receipt of the recalled product, a credit will be issued to your account.8. Report any adverse events associated with the use of these devices to Stryker Instruments 1-800-253-3210, including any infections potentially associated with the use of affected product. |  | 1542 | 1542 | Distributed in all 50 U.S. States and the District of Columbia. Distributed in Australia, Argentina, Japan, Canada, Chile, Korea, Spain, France, Italy, Netherlands, and United Kingdom. | Open, Classified |  |  | 12859
Z-1261-2018 | 79513 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via software is intended to be used for viewing, manipulation, communication, and storage of medical images. It can be used as a standalone device or together with a variety of cleared and unmodified syngo based software options. |  |  | January 10, 2018 | 2018 | Siemens Medical Solutions USA, Inc | When the archiving configuration is changed, data received/created after the upgrade may be flagged as "Not to be archived". This is caused by the automatic function for cleaning up temporary data being disabled by the software upgrade. Due to the disabled cleanup function, disc capacity for free space decreases faster than usual. Unless the archiving configuration and the cleanup automatic function are reverted back to the original state (prior to the software upgrade), data that was incorrectly flagged "Not to be archived" must be manually prevented from being deleted by either clinical administrator or service engineer. | Software Manufacturing/Software Deployment | Customers received letters with information related to the issue and regarding the upgrade to the software version VB20A_HF04. The software was released December 22, 2017 and is available for install. |  | 165 | 165 units | Medical device software which needs to be installed. | Terminated | October 02, 2018 | 265 | 165
Z-1266-2018 | 79488 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | IMRIS ORT300, Removable Operating Room Table, Part Numbers: |  |  | February 23, 2018 | 2018 | Deerfield Imaging, Inc. | The Operating Room Table may drift while in use, which could potentially result in unintended movement of the table. | Other | The firm initiated their field notification on 02/23/2018 by letter. The notice stated the following:"It has recently come to our attention that the ORTT2X/100/200/300 tables may drift while in use, which could potentially result in unintended movement of the table. This may pose a risk of injury to the patient.To minimize these risks, IMRIS is taking the following corrective actions:" ORT T2X/100/200 tables: The tables will be tested for drift symptoms. No further action will occur for tables that pass the drift test. For tables that do not pass the drift test, a hydraulc fluid flush and filter change will be performed." ORT300 tables: all customer sites with these tables shall have the hydraulic fluid flushed and exchanged for new and the filter element will be changed.A drift test will also be performed.IMRIS has identified that your facility has one or more of these OR Tables which require evaluation. Please confirm details on page 2 and return your response to IMRIS. IMRIS Customer Service will contact you to implement this corrective action.IMRIS recommends these preventive actions:" Per the ORT Operator Manuals, the table should be tested daily in all standard positions and inspected for anomalies. Before each surgery, inspect the table for any visibly loose components or hydraulic fluid leaks and check that the headrest and tabletop safety latches are secured..Please contact customer support via our service line at 1-866-475-0525 or Email: customersupport@imris.com if you have any questions or concerns." |  | 13 | 13 units | US, Canada, China, Australia, Sweden, Germany | Completed |  |  | 70
Z-1262-2018 | 79613 | [6 empty] | F-T3SYSTEM, T3 (Trumpf Table Top) System |  |  | February 21, 2018 | 2018 | Allen Medical Systems | The recalling firm has become aware that the rail section clamps used to attach the leg support of the device may fail during use, which could result in critical harm to the patient. | Error in labeling | On February 2, 2018, the firm issued an Urgent Medical Device Correction notification to their distributors to make them aware of the device correction. The firm requested that distributors further disseminate an attached recall letter to consignees.The communication addressed to consignees advised of the issue affecting the device. Consignees were asked to inform users that the device is for armboard use only, and not to be used with the product system's leg support. The notice informed consignees that the firm would provide new units upon completion of development along with revised Instructions for Use.Customers with questions concerning this Urgent Medical Device Correction were referred to Allen Medical Customer Service: (800) 433-5774, 8:30AM-5PM (EST) |  | 83 | 83 | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Australia, France, Saudi Arabia, Germany, and France. | Open, Classified |  |  | 83
Z-1268-2018 | 79074 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | c6000, Chemistry (Photometric, Discrete), for clinical use. - cobas e 601 module: Immunoassay Analyzer, Catalog Numbers: 04745922001, 05036348001, 05860652001, 04745922692, 05036348001 &amp; 05036348692 |  |  | December 18, 2017 | 2017 | Roche Diagnostics Corporation | Complaints have been received concerning questionable results on analyzers due to customers utilizing 13 sample tubes without tube adapters. | Labeling False and Misleading | Roche Diagnostics Corporation issued Urgent Medical Device Correction (UMDC) notices and Fax Forms to their customers dated 12/18/2017 via UPS Ground (receipt required). Customers are encouraged to do the following Mandatory Steps to improve the vertical alignment and handling of 13 mm sample tubes and false bottom tubes: FOR cobas e 411 &amp; cobas e 601 analyzers and MODULAR ANALYTICS E 170 module:1) Always use Roche rack cup adapters with 13 mm tubes2) Ensure sample tubes are straightFOR cobas e 411 disk system:1) Always use 13 mm Sample Disk Tube Adapters (SDTAs) with 13 mm sample tubes.2) Ensure sample tubes are straight.3) Ensure 13 mm SDTAs are not exposed to direct sunlight.4) If a larger diameter tube is left inserted in a 13 mm SDTA for more than 3 days, the adapter may become distorted. Do not insert smaller diameter tubes into distorted adapter as they may not aligncorrectly.**All customers are encouraged to DISCARD Analyzer Bulletin 10-009U3.Follow the care and handling instructions for using adapters as outlined in this Urgent Medical Device Correction (UMDC) and complete &amp; return the UMDC fax form via fax to 1-877-271-1472 |  |  |  | US Nationwide | Open, Classified
Z-1240-2018 | 79395 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Persona Partial Knee Spacer Blocks |  |  | October 24, 2017 | 2017 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a medical device recall for the Persona Partial Knee Spacer Blocks in sizes 8mm, 9mm, 10mm, 12mm, and 14mm. Complaints received indicated that flexion has the potential to feel tight during range-of-motion trialing and that the final trial has the potential to appear looser in extension. | Component design/selection | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 10/31/2017 to customers titled "Affected Product: Persona Partial Knee Spacer Blocks Size 8mm, 9mm, 10mm, 12mm, 14mm". The letter described the product, problem, and actions to be taken. The letter instructed customers to review the notification and ensure that affected personnel are aware of the contents. Review inventory for affected product which would then be quarantined and ultimately removed by a firm sales representative. The customers were also instructed to complete and return the Certificate of Acknowledgement Form (even if you do not have affected products at your facility) via email to: CorporateQuality.PostMarket@zimmerbiomet.com or Fax to: 574-372-4265.If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 430 | 430 | US Nationwide Distribution. | Open, Classified |  |  | 430
Z-1272-2018 | 79379 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | The FEMTO LDV(TM)Z8 Femtosecond Surgical Laser |  |  | January 21, 2018 | 2018 | SIE AG, Surgical Instrument Engineering | This Correction is being initiated based on a customer complaint received on November 16, 2017 which indicated a complication with a cataract procedure. An investigation completed by the manufacturer determined that although the complication was due to user error improvements could and should be made on our part to improve the usability and performance of our device. | Device Design | The firm, Ziemer Ophthalmology, initiated their field correction by telephone on 01/21/2018 and followed with a letter on 01/28/2018. The letter described the product, problem and actions taken. The letter stated:"...we will be releasing new software which contains an added safety feature which will prevent unintended shift of the planned ARC Incisions and/or alert you in case of unforeseen eye movement between configuration and execution of the ARC incisions. We will be contacting your practice the week of March 12th to set up a time to install this new software and provide in-depth training.Prior to the installation of the new software, you may continue to use the Z8 for cataract surgery; however, we would advise that you not perform both arcuate and clear corneal incisions on the same patient until the new software is installed.If you have any questions, do not hesitate to contact Director of Operations at 618-462-9301 for any further information you require. |  | 125 | 125 devices | Worldwide Distribution: US (nationwide) including states of: AL, CA, FL, GA, IL, IN, MN, OR, TN, TX, VA, and WA; and countries of: Argentina, Austria, Belgium, Brazil, Canada, France, Germany, India, Iraq, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, Namibia, Pakistan, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Spain, South Africa, Switzerland, Taiwan, and United Arab Emirates. | Terminated | November 02, 2018 | 285 | 125
Z-1271-2018 | 79445 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | 8F Plastic DignityÂż Low Profile CT Port W/Pre-Attached ChronoFlexÂż Polyurethane Catheter |  |  | November 22, 2017 | 2017 | Medical Components, Inc dba MedComp | The kit was packaged with the incorrect guidewire. The label indicates the kit contains a marked guidewire. The kits were packaged with an unmarked guidewire. | Mixed-up of materials/components | Recall notification was issued, by mail to the customers on November 22, 2017. This notification requested that customers review inventory and return all un-used affected product after quarantine. Questions or concerns can be directed to 215-256-4201/ |  | 24 | 24 Eaches | CA, AR. AL | Terminated | July 11, 2018 | 231 | 24
Z-1207-2018 | 79449 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | 6F Taiga Guiding Catheter |  |  | September 26, 2017 | 2017 | Medtronic Inc | Complaints regarding 6F Taiga guide catheter tip detachment and/or tip splitting. | Nonconforming Material/Component | A recall notification was sent to all costumers in September 2017. Customers actions to be done: Identify and quarantine all unused affected products as listed in your inventory. Share notification with others who need to be aware where affected product may have potentially been transferred. |  | 15923 | 15,923 | Japan | Open, Classified |  |  | 15923
Z-1275-2018 | 79582 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Ray Station 4.9 Ray Station 5, Ray Station 6 and Ray Station 7 Software build numbers: 4.9.0.42, 5.0.0.37, 5.0.1.11, 5.0.2.35, 6.0.0.24, 6.1.0.26, 6.1.1.2, 6.2.0.7 or 7.0.0.19 |  |  | January 26, 2018 | 2018 | RAYSEARCH LABORATORIES AB | Software issue with Center Beam in Field functionality. Issue can result in incorrect treatment volume delivered to patient. | Device Design | On January 26, 2018 RaySearch Laboratories AB sent Field Safety Notice, Medical Device Correction #25955 to each customer via e-mail. Customers were advised of the following: be aware that Center Beam in Field affects all beams that share the same isocenter and that apertures may be modified in a way not intended; do not use Center Beam in Field after having created an opposed beam. If an opposed beam has been created, delete the opposed beam and use Center Beam in Field on the remaining beam. Then recreate the opposed beam; please educate planning staff and all users about this workaround; inspect your product and identify all installed units with the above software version number(s), then confirm you have read and understood this notice by replying to the notification via email, fax or by calling local support. Customers with questions were instructed to contact David Hedfors via e-mail at david.hedfors@raysearchlabs.com |  | 2484 | 2484 | Nationwide | Open, Classified |  |  | 2484
Z-1273-2018 | 79374 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | ProSpore Biological Indicator Part Numbers: PS-6-50, PS-5-50, PS-4-10, PS-4-50, 3471 |  |  | February 12, 2018 | 2018 | Mesa Laboratories Bozeman Manufacturing Facility | The concentration of bromocresol purple in the recalled lots does not conform to specification. | Other | Per the Recall Strategy, consignees have been requested to immediatelydiscontinue use and discard according to site procedures. |  | 28930 | 28,930 units | United States, Republic of Korea, Israel, Poland, France, China, Canada, Japan, Singapore, Hong Kong, Jordan | Terminated | July 03, 2018 | 141 | 68930
Z-1279-2018 | 79437 | No | PQP |  |  | Pathology | PMA | Oncomine Dx Target Test User Guides and Assay Definition File, Model: A32461; UDI: (01)10190302006071(11)170911(10)1.8(240)A32461 |  |  | December 04, 2017 | 2017 | Life Technologies Corporation | The Oncomine Dx Target Test may report erroneous results when used off-label for samples other than non-small cell lung cancer (NSCLC) tumor specimens. | Device Design | The following actions are described in the consignee later: Review your testing results from the Oncomine Dx Target Test to ensure that none of the reports containing BRAF V600E positive samples are impacted. Utilize Figure 1 and 2 to aid in identifying this discrepancy. If you have any questions, please contact 408.802.2341. |  | 9 | 9 | US Distribution to the states of: AZ, CA, NC, TX, and VA. | Open, Classified |  |  | 9
Z-1280-2018 | 76464 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Prismaflex Control Unit. Dialyzer, high permeability with or without sealed dialysate system. |  |  | February 15, 2017 | 2017 | Baxter Healthcare Corporation | Firm has received reports of device operators failing to adhere to instructions for use pertaining to the safe unloading of disposable sets from the Prismaflex Control Unit. Additionally, for software versions 5.10 and 6.10, the programmed syringe size for the syringe pump may revert to safe default values unintentionally. | Use error | On February 15, 2017, the firm sent an Urgent Device Correction notification and response form to affected customers via USPS first class mail in order to update software versions of the Prismaflex Control Units. The update was intended to prevent user errors related to the safe unloading of disposable sets. On May 30, 2017, the firm issued an updated notification letter notifying the same customer base of the issue with the programmed syringe sizes. The software update deployed to correct the issue with the safe unloading of disposable sets was also deployed to correct the issue with the programmed syringe sizes. All non-responding customers will be contacted via telephone. Prismaflex control units will be updated to a new software version. For further questions, please call (800) 422-9837. |  | 3440 | 3,440 units | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Algeria, Armenia, Australia, Bahrain, Belarus, Belgium, Canada, China, Cyprus, Denmark, Egypt, France, Germany, Great Britain, Greece, Ireland, Italy, Albania, Bulgaria, Croatia, Czech Republic, Estonia, Hong Kong, Hungary, Iceland, India, Iran, Israel, Ivory Coast, Jordan, Kazakhstan, Kenya, Kosovo, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Malta, Montenegro, Morocco, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sudan, Sweden, Switzerland, Syria, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Ukraine, Vietnam, Yemen, Argentina, Brazil, Chile, and Dominican Republic | Completed |  |  | 3440
Z-2472-2018 | 79665 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Emotion 6 (Model 10165888) |  |  | February 01, 2018 | 2018 | Siemens Medical Solutions USA, Inc | A potential risk of unnecessary radiation exposure due to a software issue | Software design | Customers were sent notification letter via USPS with Certified Mail Receipts and Electronic Receipts. |  |  |  | Nationwide | Open, Classified
Z-1277-2018 | 79479 | No | FPB | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CORTRAK* 2 Nasogastric/Nasointestinal (NG/NI) Feeding Tubes with Electromagnetic Transmitting Stylet and CORFLO* Enteral Feeding tube, ANTI-IV* Connector, 10 Fr, 140 cm |  |  | December 13, 2017 | 2017 | Halyard Health, Inc | The transmitting stylet within CORTRAK* 2 nasogastric/Nasointestinal (NG/NI) Feeding Tubes with ANTI-IV* connector (Product Code 20-9551Trak2A) may be visible within the window of the exit port of the feeding tube. Based on specifications, the transmitting stylet in the NG tube should not be visible within the port window. | Under Investigation by firm | The firm, Halyard, sent an "URGENT: Product Advisory Notice" letter dated 12/13/2017 to its customers on about 12/13/2017. The letter described the product, problem and actions to be taken. The Customers were instructed to do the following:Instructions for distributors included to examine unsold inventory for impacted lots, complete and return the Distributor Product Advisory Notice Acknowledgement Form, and send the attached End-User/Customer Product Advisory Notice and Acknowledgement Form to all end user customers who would have potentially received impacted product by email to: monica.king@hyh.com or by FAX to: 678-254-0347. A Halyard Customer Service Representative will contact the distributors if affected product was on hand. Instructions for End-User/Customers included to evaluate the clinical facility's unused inventory for impacted lots, inform all clinicians in the facility who may use potentially impacted devices to verify that the stylet tip is not visible within the window of the exit port after engaging the stylet with the port connector of the NG/NI Feeding Tube, and if it is visible, do not use the device. Additionally they were asked to complete and return the Product Advisory Notice Acknowledgement Form by email to: monica.king@hyh.com or by FAX to: 678-254-0347.If you require further assistance, please call Halyard Regulatory Affairs at 678-477-4165. You may also contact Halyard Customer Service at 1-844-HALYARD (1-844-425-9273, OPTION 1, OPTION 3). |  | 185 | 185 cases (10 units per case) | Worldwide Distribution: US (nationwide) to states of: AR, CA, CO, GA, IA, LA, MI, MN, NM, TX; and country of Germany. Distributed internationally to Germany. | Open, Classified |  |  | 185
Z-1276-2018 | 79438 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | Phadia Prime software, article number 12-4101-00, as used in combination with the Phadia 250 Instrument, article number 12-3900-01 and EliA Assays. This recall is for any Phadia Prime software version up to and including the current version, 2.1.4. |  |  | November 20, 2017 | 2017 | Phadia Ab | We want to inform all Phadia 250 system operators performing EliA Assays not to use the function OK to All in any version of Phadia Prime, up to and including 2.1.4, when rejecting and retesting samples with any EliA assay (the OK function may be used for rejecting single tests and dilution of samples in accordance with product DfU). | Software design | An urgent recall notification was issued to customers in November of 2017, and informs all Phadia 250 system operators performing EliA Assays not to use the function OK to All in any version of Phadia Prime, up to and including 2.1.4, when rejecting and retesting samples with any EliA assay (the OK function may be used for rejecting single tests and dilution of samples in accordance with product DfU). For further questions, please call (269) 492-1940. |  | 2307 | 2307 | US Distribution to the states of :TX, VA, NJ, MA, IN, GA, UT, TN, CA, MS, NC, ME and OR | Open, Classified |  |  | 2307
Z-1283-2018 | 79314 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Radial Artery Catheterization Kit; Model: AK-04020-C |  |  | October 26, 2017 | 2017 | Arrow International Inc | The povidone iodine prep pads included in certain kits and trays do not meet the requirements for a 36-month shelf life and should be labeled with a 24-month shelf life. | Nonconforming Material/Component | Teleflex/Arrow sent a letter dated October 26th, 2017. Customers are instructed to discontinue use and quarantine any products with the affected product codes and lot numbers. There is a Recall Acknowledgment Form that is requested to be completed that will allow the return and replacement of the affected product. Consignees without affected product on hand are also asked to return the form. |  | 24192 | 24,192 | US nationwide distribution, including Puerto Rico. | Open, Classified |  |  | 85994
Z-1294-2018 | 79362 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Alpha Air 6 ENT Microscope |  |  | January 01, 2017 | 2017 | Seiler Instr. &amp; Mfg. Co. Inc. | The microscope may tip over if attachments are fully extended in one line and there is an attempt to move the device not by the handles mounted on the post. | Use error | The customer was notified via email on about January 4, 2017. Instructions included that the scope only tips if the user has it extended all the way out and is pulling on the handles of the scope. The end user must not pull on the handles of the scope when it is extended. They were also instructed to educate personnel on this matter. For further questions, please call (314) 968-2282. |  | 44 | 43 units, but 1 affected | Worldwide Distribution - US Distribution to the states of : CA, CO, FL, IL, MN, NV, OK, TX., and to the countries of : Canada, China, Ecuador, England, Greece, Guatemala, Jordan, Lebanon, Mexico, Romania, Russia, South Korea, Taiwan, Japan and United Arab Emirates. | Open, Classified |  |  | 44
Z-1289-2018 | 79590 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DynaCad software consist of an MR Analysis Server software and the viewer workstation software. The MR Analysis software consists of DynaCad Breast, DynaCad Prostate, and DynaCad Advanced PK for other MR analyses modules. |  |  | January 24, 2018 | 2018 | Invivo Corporation | Following update from 3.3 to 3.5 of the DynaCAD software it was noted that the Ktrans map was not rendered correctly on the DynaCad Client. The defect causes the pharmacokinetic (PK) color maps to display incorrectly when viewed from remote DynaCAD client computers and could result in visually underestimating calculated Ktrans, Kep, and iAUGC values. The defect also impacts DynaCAD s on-the-fly calculation of Apparent Diffusion Coefficient (ADC) maps. If the ADC values are computed on-the-fly by DynaCAD, the ADC values and colors will also display incorrectly if viewed on a remote DynaCAD client. | Other | Invivo sent an Urgent Medical Device Correction - Field Safety Notice letter dated January 24, 2018 to affected customers. The letter identified the affected product, problem and actions to be taken. This document contains important information for the continued safe and proper use of your equipment. Customers were instructed to review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. Customers were asked to retain a copy with the equipment Instruction for Use. For questions contact Invivo Customer Support:1-877-468-4861 Opt. 3 |  | 109040 | 470 distribution events, 232 actual units (updated 6/26/18) | Device is software. Customer notification letters recommended that users discontinue use of the Ktrans, Kep, and iAUGC colormaps when assessing studies from a remotely connected DynaCAD client computer. It is further advised that users refrain from using the colormaps and values derived from ADC maps calculated by DynaCAD. ADC maps originating natively from the MRI system should be used as an alternative as these values are unaffected. Firm will provide a v4.0 software update for the affected software versions (v3.4, v3.5) to correct the defect at no charge to the user. | Open, Classified |  |  | 109040
Z-1291-2018 | 79530 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DS Titanium Ligation Clip Appliers, used in laparoscopic surgery |  |  | January 22, 2018 | 2018 | Aesculap Implant Systems LLC | Added a contraindication to IFU: Do not use the DS clips for living donor nephrectomies. Added other contraindications. Updated cleaning and sterilization instructions. | Labeling design | A letter and copy of the updated IFU was sent to affected customers the week of January 15, 2018. Product does not need to be returned. IFU is revised and put on the Aesculap Website for usage. https://www.aesculapusa.com/products/instructions-for-use. |  | 270 | 270 | US and Canada distribution | Open, Classified |  |  | 9891
Z-1300-2018 | 79494 | No | NDN | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | Confidence Plus Kit Spinal Cement System, Product Codes: 283910000 (GTIN: (01)10705034209623) and 283910000 (GTIN: (01)10705034209630) |  |  | December 15, 2017 | 2017 | DePuy Spine, Inc. | Certain lots of the spinal cement may contain incorrect amounts of contrast medium so that the product is difficult or impossible to visualize on x-ray. | Nonconforming Material/Component | The firm, DePuy Synthes, sent an "URGENT FIELD SAFETY NOTIFICATION-RECALL CONFIDENCE KIT SPINAL CEMENT SYSTEM" dated 12/2017 to its consignees. The notification described the product, problem and actions to be taken. Notifications were sent to consignees with the following instructions: 1. Immediately identify and quarantine all unused products listed below in a manner that ensures the affected products will not be used. 2. Review, complete, sign and return the attached Business Reply Form to your local DePuy Synthes sales organization in accordance with the directions on the form within 5 business days of receipt of this notification. 3. Return any affected product within 30 business days. A credit note will be issued for the returned items. 4. Forward this notice to anyone in your facility that needs to be informed. 5. If any impacted products have been forwarded to another facility, contact that facility and provide them with this letter. 6. Maintain awareness of this notice until all products listed below have been returned to DePuy Spine. 7. Keep a copy of this notice.If you have questions about this Field Safety Notification, please contact your DePuy Synthes Spine Sales Consultant or the Recall Coordinator via telephone: 1-508-828-6609 or via email: DPYUS-SpineFieldActions@its.jnj.com). |  | 1972 | 1972 | International Distribution (outside the U.S.) only: Mexico, Canada, Ireland, Great Britain, Malaysia, Germany, Austria, Colombia, Spain, Austria, Switzerland, Australia, Hungary, Croatia, and Sweden. | Open, Classified |  |  | 1972
Z-1302-2018 | 79537 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare Centricity PACS-IW, Model Numbers: |  |  | February 23, 2018 | 2018 | GE Healthcare | A database handling error could occur during the image acquisition process affecting the completeness of acquired images with Centricity PACS-IW. There is a potential that one or more image series (i.e. all images within an image set) may be missing from an exam without indication to the user. While this is rare, this can occur with imaging studies that consist of a very small number of images per series. E.g. CR Thorax exam with 1 image per series. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated February 23, 2018 to their customers. The letter provided the affected product, problem and provided the following safety instructions:"Until a product modification is available to correct this issue, the following actions are recommended: 1. Utilize the image count within the QC process to alert the user of a discrepancy in the number of transmitted images from the modality to the number of images available in the Universal Viewer. a. If a discrepancy is identified, attempt to retransmit the exam to PACS. b. If retransmission is unsuccessful, a GE Healthcare Service representative will need to be contacted to assist with resolution of the impacted exam. Urgent cases impacted by this issue should be interpreted at the modality.2. If DICOM storage commit is configured and in use, no commit notification is sent to the modality for images that were impacted by this issue.3. Attention should be given to prior cases, acquired before the application of the safety instructions given in this communication as they may be affected by the safety issue."If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 569 | 569 units | Worldwide Distribution-US Nationwide | Open, Classified |  |  | 569
Z-1305-2018 | 79433 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Dynavox T15 Speech Generating Device ( including Batteries) |  |  | December 14, 2017 | 2017 | Dynavox Systems Llc | The battery is integral to the device. If the device sits unused and is not charged for a lengthy period of time (a few months) the battery will go into a deep discharge state and the battery will expire. The battery may out-gas which may be described as swelled or puffed, as gas is trapped inside the pack surrounding the cells. This battery may cause the device case to separate a small amount. | Component design/selection | On February 26, 2018 an URGENT MEDICAL RECALL LETTER was issued to customers with the subject line listing the serial number specific to that customer along with the date of purchase. The letter informs customers of the issue and instructs customers to return product to one of their facilities to have the product repaired until an RA. Questions or concerns can be directed to 1-800-344-1778, Press 1 and then Press 2. |  | 307 | 307 | Nationally | Open, Classified |  |  | 4922
Z-1306-2018 | 79393 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | A.L.P.S. MIS Calcaneus Locking Plate LG RT F.A.S.T. Guide inserts, Item Number: 816209003,UDI: 00887868044670989403 |  |  | October 25, 2017 | 2017 | Zimmer Biomet, Inc. | The drill guides were manufactured with the incorrect color anodization. | Unknown/Undetermined by firm | DISTRIBUTORS: Review this notification and ensure that affected team members are aware of the contents. 1. Review this notification and ensure affected team members are aware of the contents. 2. Immediately locate and quarantine affected product in your inventory. 3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form.a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days. 4. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet.a. Request a Return Authorization Number 5. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com RISK MANAGERS: 1. Review this notification and notify affected personnel of the contents. 2. Assist your Zimmer Biomet sales representative and quarantine all affected product. 3. Complete Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 22 | 22 | Distributed in the following US States: AZ, MO, MT, NJ, NY, OH, TX, and WA. Distributed in Chile and Netherlands. | Open, Classified |  |  | 22
Z-1301-2018 | 79499 | No | JZG | Beta-2-microglobulin immunological test system. | Immunology | Immunology | 510(K) Exempt | Binding Site Optilite Beta-2-Microglobulin Urine Kit, Ref LK043.L.OPT.A, IVD. The firm name on the label is The Binding Site Group Ltd., Birmingham, U.K. |  |  | March 09, 2018 | 2018 | The Binding Site Group, Ltd. | The product is not meeting the antigen excess protection claim of up to 760mg/L, as stated in the Instructions for Use. | Under Investigation by firm | The recalling firm notified their U.S. location by letter dated 3/6/2018 via email on 3/9/2018. |  | 20 | 20 kits distributed to the U.S. | US Distribution was made to CA. There was no foreign/military/government distribution. | Completed |  |  | 20
Z-1298-2018 | 79294 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ArrowClarke" PleuraSealÂż Thoracentesis Kit |  |  | October 24, 2017 | 2017 | Arrow International Inc | Product packaging may not be completely sealed. | Component design/selection | On October 24, 2017 and Urgent Medical Device Recall letter was issued to customers to review inventory and if there's any affected stock, to immediately discontinue use, quarantine and prepare to return to recalling firm. Questions or concerns can be directed to 1-855-419-8507 or 1-866-246-6990 |  |  |  | Nationally | Open, Classified
Z-1296-2018 | 79279 | No | DOF | Lead test system. | Toxicology | Toxicology | 510(k) | LeadCare Plus Blood Lead Analyzer |  |  | January 09, 2018 | 2018 | Magellan Diagnostics, Inc. | The incorrect placement of AA batteries into certain LeadCare Plus Blood Lead Analyzer may cause the batteries to overheat, and the chance that a user could burn their fingers if the batteries are removed while hot. | Device Design | An Urgent Medical Device Correction letter, dated January 9, 2018, was sent to customers. The letter identified the affected device and provided a description of the issue. If two of the four AA batteries are inserted incorrectly, the batteries may overheat, causing a minor burn if the user opens the battery compartment and tries to remove the hot batteries. Customers are to follow the steps outlined in the letter, including placing the enclosed label on the upper back panel of the analyzer to remind users who are changing batteries of the importance of correct placement. Proper battery placement should be provided in annual training. Customers are to complete and return the acknowledgement card. Please direct any questions to 800-275-0102 or LeadCareSupport@magellandx.com. |  | 79 | 79 | Worldwide Distribution | Open, Classified |  |  | 6222
Z-1311-2018 | 79462 | No | PMT | Device to detect and measure non-microbial analyte(s) in human clinical specimens to aid in assessment of patients with suspected sepsis. | Microbiology | Microbiology | 510(k) | Elecsys BRAHMS PCT |  |  | December 22, 2017 | 2017 | Roche Diagnostics Corporation | Product exhibits a decreased recovery of patient samples on the cobas e 801 module compared to the MODULAR ANALYTICS E 170, cobas e 411, 601, and 602 analyzers. | Software design | The firm, Roche, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated December 22, 2017 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:" Immediately discontinue use of Elecsys BRAHMS PCT (catalog number 07301715190) reagent lotnumber 204084 on the cobas e 801 module." Discard affected product according to local guidelines." Complete the attached fax form and fax it to 1-800-722-7222." File this Urgent Medical Device Removal (UMDR) for future reference.Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 if you have questions about the information contained in this UMDR. |  | 47 | 47 | US Distribution in states of: CA, FL and IN. | Completed |  |  | 47
Z-1309-2018 | 78642 | No | EZD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Personal Intermittent Catheters- 16" and 10" lengths - 30 catheters per box, each catheter is packaged in an individual pouch. |  |  | September 14, 2017 | 2017 | C.R. Bard, Inc. | Several instances of a catheter caught in the sterile barrier seal and in some instances cut catheters were observed. | Process control |  |  | 10949013 | 10,949,013 units in total | US Nationwide in the states of AK, AL, AR, AZ, CA, FL, GA, IA, ID, KS, KY, MD, ME, MI, MN, MO, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, including Hawaii and Puerto Rico | Open, Classified |  |  | 10993653
Z-1313-2018 | 79444 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | ImmunoCAP Allergen c6, Amoxicilloyl ( Article number 14-4451-016). |  |  | December 18, 2017 | 2017 | Phadia US Inc | The stability specifications of ImmunoCAP Allergen c6, Amoxicilloyl were not met. | Device Design | Customers are being asked to please stop using all current product, and scrap any unused products. Âż Please disregard any reported test result below the assay Limit of Quantification (LoQ 0.1 kUA/l) from the lots, as these results may not be valid.Please note that assay results above the Limit of Quantification, will only be affected by giving a lower result than the real value.Âż Please fill in the Recall Letter Acknowledgement Form provided, including the table in the response form to obtain credit, and return the signed response by e-mail or fax. |  | 950 | 950 total | Nationwide Distribution | Open, Classified |  |  | 950
Z-1315-2018 | 78732 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IntelliVue X3 Patient Monitor. |  |  | November 17, 2017 | 2017 | Philips Electronics North America Corporation | The NBP measurement of Intellivue X3 Patient Monitor shows intermittently only mean values instead of the diastolic and systolic blood pressure values. This is caused by a falsely detected NBP cuff, which is leading to wrong internal NBP setting. Furthermore, occasionally the Monitor shows the Check Touch Input message and the monitor is inoperable with the touch interface. | Under Investigation by firm | Philips will contact customers with affected devices to arrange a correction of the issues. Firm plans a Software update of the affected devices and Replacement of the display assembly. |  | 975 | 975 (updated 8-14-18) | 48 Foreign Accounts | Open, Classified |  |  | 975
Z-1316-2018 | 79010 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | HomeSafe AutoAlert Pendant works in conjunction with a compatible Lifeline communicator. |  |  | November 13, 2017 | 2017 | Lifeline Systems Company | A programing error in some Model FD100 HomeSafe AutoAlert Pendants will render the fall detection featureinoperable. | Software Manufacturing/Software Deployment | Firm has contacted all recipients and informed them of an issue that affects Philips Lifeline AutoAlert help button from automatically signaling the Lifeline Response Center for help in the event of a fall. Further instructed them that they would need to press the Lifeline Help Button in order to connect with the emergency response center. They then instruct them to return their inventory for replacement. New AAHB Replacements will be provided when available. Philips will work with the Programs to ensure buttons swaps are completed. |  | 20201 | 20201 | US and Canada distribution | Open, Classified |  |  | 20201
Z-1314-2018 | 79524 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Terumo Cardiovascular Procedure Kits containing Pall LG6NS LeukoGuardÂż Leukocyte Reduction Arterial Blood Filters. |  |  | December 18, 2017 | 2017 | Terumo Cardiovascular Systems Corp | Possible blood leaks through the hydrophobic portion of the Pall LG6NS LeukoGuard Leukocyte Reduction Arterial Blood Filters. | Unknown/Undetermined by firm | On December 18, 2017 a MEDICAL DEVICE RECALL letter was issued to customers with the specific lot codes and distribution dates listed on the customer response form. This letter requests customers to do the following: Review this Medical Device Recall Notice. Assure that all users receive notice of this issue. Refer to the Customer Response Form to identify your product that is subject to this action. Confirm receipt of this notification by completing and returning the attached Customer Response Form to the email address or fax number indicated on the form. Terumo CVS will issue a Returned Goods Authorization upon receipt of the Customer Response Form. Return all affected products to Terumo CVS.Questions or concerns can be directed to 1-800-521-2818 |  | 2677 | 2677 | US Distribution to the states of : AL, OH, VA, NC, GA, NY, FL | Open, Classified |  |  | 2677
Z-1318-2018 | 79601 | refractor, manual, non-powered, including phoropter | HKN | Manual refractor. | Ophthalmic | Ophthalmic | 510(K) Exempt | Refractor Suspension Arm (P/N: 525004, 05 and 525006, 07) |  |  | December 06, 2017 | 2017 | Reliance Medical Products Inc | Product has been found with a missing Washer, Retaining Ring, and/or Stop Screw. Recalling firm was notified by distributor that Refractor Suspension Arm fell on a patient s lower torso along with the arm housing | Employee error | Recalling firm sent letter to customers via certified mail for US customers and via regular mail to foreign customers. Verify product has the presence of the one or more of the three parts: Refractor Suspension Arm missing Screw, Washer and/or Retaining Ring (Snap Ring) that may be missing, and to assure it is functioning properly. For further questions, please call (513) 398-3937 |  | 90501 | 933 ( of which 97 has been quarantined) | Worldwide Distribution - USA(nationwide) Distribution to the states of : NJ, CA, OH, VA, MA, MN, IL , UT, TN, AR, LA, AZ, PA, TX, IA and HI., and to the countries of : Costa Rica, Jamaica, South Africa, New Zealand, South Korea, Dubai, Saudi Arabia and Jordan. | Open, Classified |  |  | 90501
Z-1319-2018 | 79516 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Progressa Bed - AC-powered adjustable hospital bed |  |  | January 12, 2018 | 2018 | Hill-Rom, Inc. | A component used within the braking system may break resulting in one or more brake pedals on the bed staying in brake or not being engaged which could result in unwanted bed movement and possible patient or user injury. | Component change control | Firm has developed a correction for the braking system and plan to have a Hill-Rom technician to update affected Progressa Beds. Hill-Rom sent a letter, dated January 12, 2018.Until beds are updated, follow the safety information as outlined in the Progressa User manual each time when setting the brakes. Warning: Always set the brakes when the bed is occupied, except during patient transport. To help make sure the bed will not move, push and pull on the bed to check it after the brakes are engaged. Brakes should always be set when the bed is occupied and especially when moving a patient from one surface to another. Patients often use the bed for support when getting out of bed and could be injured if the bed unexpectedly moves. After setting the brakes, push and pull the bed to make sure of stability.If the brakes do not hold after following the instructions above, take the bed out of service until the bed is updated. |  | 2117 | 2117 | Worldwide distribution. US nationwide, Canada, Chile, Colombia, Costa Rica, France, Hong Kong, Israel, Jamaica, Japan, Jordan, Malaysia, Mexico, Oman, Panama, Russian Federation, Saudi Arabia, Singapore, Spain, Taiwan, UAE, and UK. | Open, Classified |  |  | 2117
Z-1320-2018 | 79556 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | Hematology Diluent M-30D used with BC-3200 and BC-3600 Hematology Analyzer; Part Numbers: (1) 105-005405-00 (2 x 5.5L container), (2) 105-005406-00 (20L container) |  |  | January 24, 2018 | 2018 | Mindray DS USA, Inc. dba Mindray North America | Certain lots of M-30D diluent used with BC-3200 and BC-3600 Hematology Analyzers and M-53D diluent used with BC-5390 Hematology Analyzers may cause the system to produce an elevated platelet background count. | Material/Component Contamination | Letters dated January 24, 2018, were sent to consignees with the following instructions: The issue manifests intermittently, therefore, you may continue to use affected lots if the background check is successfully completed. Discard affected lots which cause the system to display the error message background abnormal . Mindray will replace the potentially affected lots at no cost. Please contact Technical Support to order replacement product. For further questions, please call (201) 995-8407. |  | 4838 | 4838 | US Distribution including Puerto Rico and to the states of : CA, MI, NC, TX, and WA. . | Open, Classified |  |  | 5334
Z-1341-2018 | 79381 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | AQUIOS CL Flow Cytometer System, Catalog #B30166. The AQUIOS CL Flow Cytometer system is an automated analyzer that use a no-wash sample preparation process.The firm name on the label is Beckman Coulter Ireland, Inc., Co. Clare, Ireland. |  |  | January 29, 2018 | 2018 | Beckman Coulter Inc. | The device may process the same sample with two different sample IDs and sample information while using the single tube loader, which has the potential for erroneous results due to the mis-identification. | Software design | The recalling firm issued letters dated 1/25/2018 via email on 1/31/2018 or to some customers via regular mail during the week of 1/29/2018. |  | 288 | 288 devices | Distribution was made nationwide in the U.S. and in Puerto Rico. There was also government distribution, but no military distribution. Foreign distribution was made to Angola, Australia, Belgium, Botswana, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Finland, France, Germany, Ghana, Hong Kong, India, Israel, Italy, Japan, Kazakhstan, Kenya,Lesotho, Malawi, Malaysia, Mexico, Namibia, Netherlands, New Zealand, Norway, Pakistan, Panama, Poland, Portugal, Romania, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Uganda, Ukraine, United Kingdom, and Zambia. | Open, Classified |  |  | 288
Z-1349-2018 | 79583 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | VOLISTA StandOP Surgical Light; Model numbers VLT600 SF AIM STP, VLT600 DF AIM STP, Ref. codes ARD568811901, ARD568811911 |  |  | February 07, 2018 | 2018 | GETINGE US SALES LLC | The manufacturer received complaints indicating that the device's central handle holder detaches easily from the handle holder ring on the light head during clinical use. In the event of such detachment, particles from the handle holder ring could fall into the sterile field during a procedure. | Component design/selection | The firm, Getinge, sent an "URGENT MEDICAL DEVICE RECALL" Notice dated 2/7/2018 to all affected consignees via mail. The notice described the product, problem and actions to be taken. The consignees were instructed to do the following:1) Ensure that all users of the device are made aware of this field correction.2) Check the device serial number against the list of affected serial numbers.3) Contact the recalling firm to arrange for servicing of all affected units. Getinge Customer Service: (800) 950-9912, 8AM - 5PM (EST) Monday to Friday.4) Complete and sign the enclosed customer response form, and return it by fax to: (973) 860-1675 or by email to: VolistaLights.FA@getinge.com.If you have any questions, please contact your Getinge representative or call Getinge Customer Service at (800)950-9912 (Select option 1) between 8:00 AM and 5:00 PM (EST). |  | 5214 | 5214 | Worldwide Distribution: US (nationwide) to states of: AR, AZ, CA, CO, FL, ID, IA, KS, KY, ME, MO, NE, NY, OR, TN, TX, UT, and WA; and countries of: Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cameroon, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Ghana, Hong Kong, Hungary, India, Iran, Iraq, Italy, Japan, Jordan, Kenya, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Namibia, Netherlands, Nigeria, Oman, Pakistan, Panama, Paraguay, Poland, Portugal, South Korea, Romania, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Trinidad-Tobago, Turkey, UAE, UK, Yemen and Zimbabwe. | Open, Classified |  |  | 5214
Z-1340-2018 | 79541 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PDS Plus Antibacterial (Polydioxanone) Suture |  |  | January 10, 2018 | 2018 | Ethicon, Inc. | Three (3) lots were distributed with the incorrect expiry date on the label. The lotswere labeled with an expiry date of five (5) years from date of manufacture insteadof the correct expiry date of two (2) years from date of manufacture. | Process control | On January 10, 2018 an URGENT MEDICAL DEVICE LETTER for removal was issued to customers requesting that all affected personnel to discontinue use or distribution of the recalled product. The following actions being required:1. Examine inventory and quarantine all affected product2. Communicate this recall notice to all affected personnel at facility, or other affected facilities.3. Complete the business reply form attached to the recall letter.4. Keep the recall notice posted visibly at facility and maintain a copy with your records5. Return unused product to recalling firm by April 13, 2018. For further assistance contact 1-855-215-5039. |  | 180 | 180 | Distribution US Nationwide, Canada, India, and Colombia. | Open, Classified |  |  | 7272
Z-1350-2018 | 79559 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Spinal Jaxx Interbody Fusion Device, Model: 100001-511, Size: 9 H x 10 W x 25 L |  |  | January 18, 2018 | 2018 | Neuropro Spinal Jaxx | The implant size is incorrectly etched on the implant. The size listed on the box is correct. | Labeling Change Control | The firm called the consignee using a phone script. Additional information requested regarding whether any instructions were provided to the consignee regarding what to do with the recalled devices. |  | 6 | 6 | The devices were distributed in California. | Terminated | July 13, 2018 | 176 | 6
Z-1348-2018 | 79553 | No | HKO | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | AUTO REF/KERATOMETER ARK-1s |  |  | December 20, 2017 | 2017 | Nidek Inc. | There is possibly a difference between the measured dioptric value of the subjective measurement mode and the objective measurement mode on ARK-1s. | Equipment maintenance | Customers were notified via letter on about 12/20/2017. The letter instructed consignees that a representative from Marco Ophthalmic, Inc. will make arrangements for the affected unit to be picked up and replaced with an ARK-S unit that meets factory specifications. |  | 63 | 63 untis | Distributed domestically to AK, AL, AR, CA, CO, DE, FL, GA, ID, IL, IN, MD, MI, MO, MT, NM, NY, OH, OK, PA, SC, TN, TX, VA, WA, WV. | Completed |  |  | 63
Z-1322-2018 | 79508 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Immunodiagnostic Products TSH Reagent Pack |  |  | January 11, 2018 | 2018 | Clinical Diagnostic Systems | VITROS TSH reagent lots have been found to generate higher than expected Calibrator level 2 signals on some customer systems. These Calibrator Level 2 responses cause a higher than expected frequency of customer generated calibration curves to fall outside of the calibration quality parameter measuring the gradient of the Customer Calibration Curve between Calibrator Level 2 and Calibrator Level 3. | Nonconforming Material/Component | Distributors instructed to disgard remaining inventory, complete Confirmation of Receipt form, and firm promised to replace product. Distributors instructed to send letter to customers. |  | 16078 | Lot #5470 (US=8570/OUS=537); Lot #5430 (US=0/OUS=9135); Lot #5178 (US=5178/OUS=4396) | Distributors were notified via letter.Lot 5430 was not distributed in the US.Lot 5178 was also affected by this issue, but this lot expired on 31-Mar-2017 so Ortho took no action on this lot. | Open, Classified |  |  | 16078
Z-1335-2018 | 79512 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | AGB+ Multi Lumen CVC Kit (1) Part Number (PN): CDC-42703-B1A, (2) PN: CDC-42802-1A, (3) CDC-45703-1A, (4) CDC-45703-A1A, (5) PN: CDC-45703-B1A; AGB+ Quad-Lumen CVC Kit (6) PN: CDC-42854-1A, (7) PN: CDC-45854-1A; AGB+ Two-Lumen CVC Kit (8) PN: CDC-45802-1A, (9) PN: CDC-45802-B1A, (10) PN: CDC-46702-1A, (11) CDC-47702-1A; (12) Multi-Lumen Central Venous Catheterization Kit with Blue FlexTip(R), ARROWg+ard Blue PLUS(R) Catheter and Sharps Safety Features, PN: ASK-45703-AH |  |  | December 22, 2017 | 2017 | Arrow International Inc | Certain lots of cardiovascular catheters may not be completely sealed and so sterility of the product cannot be guaranteed. | Process control | The notification letter includes the following instructions for CUSTOMERS: 1. Discontinue use and quarantine any affected products. 2. Complete and return the Recall Acknowledgement Form to start return process. 3. Complete the Recall Acknowledgment Form if there is no product on hand. Instructions for DISTRIBUTORS: 1.Discontinue distribution and quarantine any affected products. 2. Communicate this recall to any customers with affected product. 3. Have the customers return any affected product to you with a completed RecallAcknowledgement Form. 4. Return affected products from inventory using the Recall Acknowledgement Form. 5. Once you have returned recalled products from your own inventory, and collected the recalled products from your customers, please check the box on the enclosed Recall Acknowledgment Form that indicates that you have completed the recall and return the form. |  |  |  | Distributed in all 50 U.S. states and the District of Columbia. Distributed in Bahamas, Canada, Mexico, Panama, and Singapore. | Open, Classified
Z-1342-2018 | 79077 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 6.5" (17 cm) Smallbore Ext Set w/Remv MicroClaveÂż Clear, NanoClaveÂż T-Connector, Clamp, Rotating Luer, Bulk Non-Sterile, 500 per case, Item A1009-NS, Rx. The firm name on the label is icumedical, San Clemente, CA. |  |  | February 05, 2018 | 2018 | ICU Medical, Inc. | The infusion sets were manufactured with an incorrect spin collar. | Process control | The firm, ICU Medical Inc., sent an "URGENT: Medical Device Recall Notification" letter dated 2/5/2018 to its customers via UPS on 2/5/2018 in the U.S. The letter described the product, problem and actions to be taken. The customers were instructed to: inspect your inventory for specific product and lot numbers; quarantine all affected devices; complete and return the Recall Response Form as applicable (even if you have no affected devices) via Fax to: 1-866-899-7474 or email to: ICUMedical2805@stericycle.com; if you have distributed the product further, notify your accounts that received the product and ask them to contact Stericycle at 1-866-629-6182 (M-F, 8am-5pm ET) to receive a Response Form. Return affected product to Stericycle using labels provided, if you have not received a return label or reply form or require additional assistance, please contact Stericycle at 1-866-629-6182 (M-F, 8am-5pm) the return labels are for single use only. For additional labels visit http://expertezlabel.com.For further inquiries, please contact ICU Medical Inc. for Global Complaint Management at ProductComplaintsPP@icumed.com or ICU Customer Service at 1-866-829-9025, option 2 (M-F, 8:30am-4pm PT). |  | 11500 | 11,500 pouches | Worldwide Distribution: US (nationwide) to state in AZ, CA, CO, FL, SC, and TX; and countries of: Canada and Japan. | Completed |  |  | 22950
Z-1360-2018 | 79495 | No | KDG | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Normed Charcot Osteotome Chisel (Various sizes) |  |  | March 15, 2018 | 2018 | Zimmer GmbH | Product may contain a gap between the handle and metal chisel which can lead to failure to properly disinfect or sterilize between use. | Under Investigation by firm | On March 15, 2018 an URGENT MEDICAL DEVICE RECALL LETTER was issued to affected hospitals with the following instructions:1. Review this notification for awareness of the contents.2. Assist your Zimmer Biomet sales representative to quarantine immediately all affected instruments.3. Your Zimmer Biomet sales representative will remove the affected instruments from your facilityQuestions or concerns can be directed to a zimmer biomet representative. |  |  |  | CO, IN, TX, IL | Open, Classified
Z-1356-2018 | 79507 | No | FPD | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard(R) Adult/Pediatric Feeding Tube, ENFit(TM), REF EN0036420 |  |  | February 27, 2018 | 2018 | C.R. Bard, Inc. | BMD has identified that the product code/lot number combination identified above may have ineffectively fused tubing and ENFit(TM) connector. | Finished device change control | The firm initiate their recall by letter on 02/27/2018 by letter. The recall letter stated:"Action required:" Immediately examine your inventory and quarantine product subject to this recall. Please refer to the sample product labeling on Attachment 1 to help you locate the product. Do not use or further distribute any affected product." Please complete and return the accompanying Recall &amp; Effectiveness Check Form attached to this letter regardless of whether or not you have any of the affected product. Additional instructions for product return are contained on the form. Upon receipt of the form indicating there is product to return, the BMD Recall Coordinator will issue you a return authorization number for return of the affected product." If you have further distributed any units from the affected lots, please identify your customers and notify them at once of this product recall by forwarding a copy of this letter. Your notification should include a copy of this letter and the accompanying enclosures.This recall should be carried out to the user level." |  | 3200 | 3200 units | US Nationwide Distribution | Open, Classified |  |  | 73600
Z-1359-2018 | 79656 | No | FHA | Urological clamp for males. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Strauss Penis Clamp 130MM/General Instruments |  |  | February 16, 2018 | 2018 | Aesculap Implant Systems LLC | The plastic sleeves of the Penis Clamp and Mouth Gag may stick together during or after reprocessing rendering the instruments unusable. | Mixed-up of materials/components | An Urgent Medical Device Recall notification will be sent via Fed-ex overnight to our customers on February 16, 2018. Customers will be required to review the recall notification letter in its entirety and ensure all users in the organization are informed about this recall and the affected products |  |  |  | nationwide | Open, Classified
Z-1353-2018 | 79674 | No | LRR |  |  | General &amp; Plastic Surgery | Enforcement Discretion | First aid/emergency kits containing Honeywell eyewash: |  |  | March 20, 2018 | 2018 | Certified Safety Mfg Inc | Firm is recalling first aid/emergency kits and cabinets which contain Honeywell Eyewash (3 lots) which is being recalled due to leaking bottles. | Other | The firm initiated the recall by letter on 03/20/2018. The firm requested return of the recalled Honeywell eyewash component. |  | 748 | 748 units | US Nationwide Distribution in the states of NE, KY, CA, OH, TX, NC, AL, NJ, TN, MN, IA, MO, MA, MI, ME, BY, GA, WA, LA, IL, VA, NY | Open, Classified |  |  | 748
Z-1363-2018 | 79498 | No | NKG |  |  | Orthopedic | 510(k) | ORTHOFIX, Connector System, SMALL SET SCREW, REF 79-2003 |  |  | February 23, 2018 | 2018 | Orthofix, Inc | Orthofix is conducting a voluntary recall of the Connector System Large Set Screw (PN: 79-2002) with affected Lot Numbers 001 and 002, and Small Set Screw (PN: 79-2003) with affected Lot Numbers 002, 003 and 004. Through the company's manufacturing controls, size discrepancies were noted with the drive feature for both the large and small sets screws. As a result, it is possible that (1) the set screw drive feature will not accept the Set Screw Driver (PN: 79-1006) or(2)the driver may get wedged in the corners of the drive feature (false bottom) which may lead to a stripped set screw. | Nonconforming Material/Component | The firm disseminated the recall notices on 02/23/2018 by mail. The firm requested the return of the screws. |  | 397 | 397 units | US | Open, Classified |  |  | 585
Z-1352-2018 | 79577 | No | OBP | Repetitive transcranial magnetic stimulation system. | Neurology | Neurology | 510(k) | Magstim Rapid Therapy System consisting of: Rapid Mainframe, Rapid Single Power Supply Unit , Rapid User Interface, Air Film Coil and Main Filters |  |  | September 28, 2017 | 2017 | The Magstim Company Limited | Users manuals not supplied with 4800-00T US Rapid Therapy System | Employee error | On September 28, 2017 the recalling firm sent a Field Safety Notice that included the Operating Manual to all their customers involved in this event. An Acknowledgement card was attached to the letter to be sent back to the recalling firm informing that they received the Operating Manual |  | 43 | 43 | US Nationwide Distribution in the states of AR, CA, GA, KS, MA, MD, MO, NC, NJ, NW, NY, TX and Hawaii | Open, Classified |  |  | 43
Z-1362-2018 | 79566 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BECKMAN COULTER iQ200 Series Urine Microscopy Analyzer, Catalog Numbers: 700-3345, C10684, 700-3375, C10683, 700-3325, 700-3347, 700-3320, 700-3322, 700-3370 |  |  | March 02, 2018 | 2018 | Beckman Coulter Inc. | Beckman Coulter has determined that there is a potential for under-reporting casts. This can occur if per high-power field (/HPF) units of measurement for casts are selected in the iQ200 software, but the abnormal threshold and/or grading format is set up based on reporting per low-power field (/LPF)* or *per microliter. This may occur during initial method validation or if settings are altered after the initial validation. | Use error | The customers were contacted initially by the Urgent Medical Device Recall letter on 03/028/2018 delivered by email and US Mail. Beckman Coulter will schedule a sites visit to verify the reporting units of measurement for casts for the laboratory within 90 days of this recall initiation. The use of Casts/HPF is not a common setting worldwide. |  | 5247 | 5247 units | Worldwide | Open, Classified |  |  | 5247
Z-1351-2018 | 79533 | No | EBW | Dental handpiece and accessories. | Dental | Dental | 510(k) | Foot Controls |  |  | December 21, 2017 | 2017 | W &amp; H DentalWerk Burmoos GMBH - Site 1 | During continuous postmarket surveillance activities, we have noticed a potential malfunction of individual foot controls of the type S-N1/S-N2.In case the error occurs, the foot control might not switch back to zero position when the pedal is released. Consequently, the motor keeps rotating. If this happens during automatic thread cutting or automatic screwing-in of implants, it may, in the worst case, lead to a serious injury of the patient. | Process change control | On December 22, 2017 a field safety notice was issued to customers notifying them of the potential problem with the foot controls. On February 15, 2018 a replacement foot control was sent to customers with instructions to unplug the original foot control, cut the cable once disconnected to an electrical source and properly dispose of. The replacement foot control is to be used instead. Questions or concerns can be directed to footcontrol@wh.com |  |  |  | US Nationwide Distribution | Open, Classified
Z-1371-2018 | 78599 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-C2 Ventilator Unit |  |  | March 12, 2018 | 2018 | Hamilton Medical AG | After two and a half years, the battery fuel gauge may indicate a higher battery charge than what is available. | Other | Customers were notified via letter on about 03/12/2018. The letter identified the affected device and described the issue. Instructions for medical facilities are to inspect the batteries for the battery manufacturing date and replace any batteries that are found to be in use for two years or more. Instructions for determining the battery manufacturing date are included in the Medical Device Field Safety Corrective Action Update attached to the letter. Questions or concerns can be directed to Bret Everett at 817-909-0308 or Hamilton Technical Support at 1-800-426-6331. |  | 739 | 739 units total | Worldwide Distribution -- USA, including the states of AZ, CA, CO, FL, GA, HI, IL, IN, KY, MA, MI, MN, MO, MT, NC, ND, NE, NV, OH, OK, PA, SC, SD, TN, TX, UT, VA, VT, and WA; and, the countries of Algeria, Angola, Argentina, Armenia, Australia, Austria, Austria, Azerbaijan, Bangladesh, Belarus, Belgian, Bolivia, Brazil, Bulgaria, Bulgaria, Cambodia, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czechia, Denmark, Dominican. Rep., Ecuador, Egypt, El Salvador, Estland, Finland, France, Gabon, Germany, Greece, Guatemala, Holland, Honduras, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Malaysia, Maldives, Mauritius, Macedonia, Mexico, Mongolia, Morocco, Myanmar, Nepal, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Rwanda, Saudi Arabia, Scotland, Serbia, Seychelles, Slovakia, Slovenia, South Africa, Spain, SRI Lanka, Sudan, Suriname, Sweden, Sweden, Switzerland, Syria, Taiwan, Tajikistan, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, UAE, Uganda, UK , Ukraine, Uruguay, Venezuela, Vietnam, and Yemen. | Open, Classified |  |  | 739
Z-1373-2018 | 79224 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Hex Driver Part Number HXDRIVER-2.0 AO; Hex Driver Part Number HXDRIVR-2.5 AO; Hex Driver Part Number HXDRIVR-2.5CL AO; Hex Driver Part Number DVHX-2.5/130C AO; Hex Driver Part Number HXDRIVR-3.0 AO; Hex Driver Part Number HXDRIVER-3.5AO |  |  | January 11, 2018 | 2018 | TriMed Inc. | Following the completion of sterilization cycle(s) of the impacted Hex Driver, the epoxy color band on the impacted Hex Driver can migrate down the shaft and/or potentially onto the sterilization tray or caddy. | Process control | Recall notification letters were sent to affected consignees on January 16, 2018. The consignee letter requests that the consignee return an initial Acknowledgement &amp;Receipt Form to TriMed by January 23, 2018. |  | 1525 | 1525 devices | US, China, Latin America | Open, Classified |  |  | 1525
Z-1367-2018 | 79579 | No | BTC | Reservoir bag. | Anesthesiology | Anesthesiology | 510(K) Exempt | Latex Rebreathing Bags: |  |  | March 15, 2018 | 2018 | Teleflex Medical Europe Ltd | These products contain latex, but the following statement is not included on any level of product packaging: Caution: This Product Contains Natural Rubber Latex Which May Cause Allergic Reactions. | Employee error | The firm initiated their field correction on 03/15/2018 by letter. The notices stated the following:Affected distributors are instructed to take the following actions:1. Place a copy of this notification with each unit of affected product currently in your inventory.2. Using the provided customer letter template and acknowledgment form, communicate this notification to any of your customers who have received product included within the scope of this notification.3. Have each of your customers who received the affected product return a completed acknowledgment form to you.4. Once you have finished collecting and consolidating all of the acknowledgment forms, from your customers and placing a copy of this notification with each unit of affected product in your inventory, please complete the enclosed Distributor Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document completion of this recall. Affected direct customers are instructed to take the following actions:1. Place a copy of this notification with each unit of affected product currently in your inventory. Users should note This Product Contains Natural Rubber Latex Which May Cause Allergic Reactions. 2. After step (1) is complete, fill out the enclosed Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com.Upon finding additional lots, the firm issued additional letter on 06/27/2018 with the same message as the letter disseminated in March. |  | 107962 | 107962 units | US Nationwide Distribution | Open, Classified |  |  | 237444
Z-1365-2018 | 79625 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Liberty Select Cycler (SW v.2.8.7), Material Number RTLR108343 |  |  | January 24, 2018 | 2018 | Fresenius Medical Care Renal Therapies Group, LLC | The recalling firm identified a software issue related to the Patient Line Check (PLC) which may result in an increased risk of Overfill (also known as Increased Intraperitoneal Volume, IIPV). Overfill/IIPV may result in serious injury or death. | Software design | On 01/25/2018, the recalling firm sent Urgent Medical Device Correction letters to affected Home Therapy Nurse Managers (HTNs). On 01/31/2018, the firm sent letters to affected patients. The letters informed customers of the recall and stated that patients who are considered "slow drainers" and experience M65 Scale warnings during one or more drain cycles are at increased risk.On 03/29/2018, the firm sent out a second Urgent Medical Correction Letter to HTNs, followed by a patient communication sent on 04/02/2018. In the second communication, the firm amended that slow drainers can develop Overfill/IIPV even without receiving M65 scale warning.Health care providers and patients using the device were advised to monitor drain rates and address the reasons for slow draining, such as constipation, fibrin deposition, peritonitis, and catheter malposition. Health care providers were advised to continue to review risk factors and symptoms of Overfill/IIPV with PD nursing staff and patients with extended drain times.The firm is not requesting the return of the recalled device at this time. For additional information and alternative treatment options, please contact the FMCRTG Medical Information Line at 1-855-616-2309. Hours of operation are 8:30AM-5:00PM (EST), Monday - Friday. Online requests may be submitted at any time at www.fmcna-medinfo.com. |  | 9293 | 9293 | Nationwide distribution. | Open, Classified |  |  | 9293
Z-1366-2018 | 79368 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | iQ200 Series Urine Microscopy Analyzer with Barcode Reader Model NFT-2100, Catalog #s: 800-3046, 800-3047, 800-3052, 800-3053, 800-3802, 800-3803, 800-3900, 800-3920, 800-3925, 800-3931, 800-3933, 800-3934, 800-3935, 800-3937, 800-3950, 800-3951, 800-3042, 800-3043, 800-3044, 800-3049, |  |  | November 10, 2017 | 2017 | Beckman Coulter Inc. | iQ200 Series Urine Microscopy Analyzer may intermittently fail to read urine sample dilution barcode labels causing erroneous results that could delay treatment for health conditions such as hematuria. | Device Design | An urgent medical device recall letter, dated 11/10/2017, was sent to customers informing them that Beckman Coulter was initiating a field action for the iQ200 Series Urine Microscopy Analyzer with Barcode Reader, which may fail to read urine sample dilution barcode labels instead defaulting to no dilution, dilution factor of 1:1. In addition, this product may also fail to read urine sample dilution barcode labels, when the rack is rejected, leading to no results. Customers were informed of the issue, impact, action, and that the resolution was being investigated. Customers were instructed to respond within 10 days: electronically via email or manually by completing and returning the response form accompanying the device recall letter. Customers with questions regarding the notice were instructed to contact the Customer Support Center: electronically at http://www.beckmancoulter.com/customersupport/support or via phone at 1-800-854-3633, in the United States and Canada. Customers with questions located outside the United States and Canada, were instructed to contact local Beckman Coulter representatives. |  | 4350 | 4350 total units (1713 in U.S.) | United States: nationwide, including Puerto Rico; and Foreign: Canada, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Bermuda, Bolivia, Brazil, Brunei Darussalam, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, France, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Israel, Italy, Japan, Kenya, Korea, Kuwait, Lebanon, Lithuania, Macao, Macedonia, Malaysia, Maldives, Martinique, Mexico, Moldova, Republic Of, Montenegro, Morocco, Netherlands, New Zealand, Nigeria, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, U.A.E, Uganda, United Kingdom, United States, Uruguay, Venezuela, Vietnam | Open, Classified |  |  | 4350
Z-1375-2018 | 79581 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | WRIGHT, Wright Medical Technology, VALOR(R) Nail, Right, Medium, OD: 10mm, L: 200mm, REF415101020 |  |  | March 14, 2018 | 2018 | Wright Medical Technology, Inc. | Tow lot of VALOR(TM) Nail could potentially be missing the assembled compression screw. | Under Investigation by firm | Firm initiated the recall on 03/14/2018 by email communication to distribution centers and distributor consignees, along with a field note to all sales representatives. The firm requested return of the product. |  | 47 | 47 units | US and Argentina, Belgium, Brazil, Chile, Colombia, Germany, Italy, Mexico, Peru, Thailand, and United Kingdom | Open, Classified |  |  | 97
Z-1374-2018 | 79373 | No | NWJ | Absorbable poly(hydroxybutyrate) surgical suture produced by recombinant DNA technology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Fornier Phantom Fiber(TM) Sutures: |  |  | February 08, 2018 | 2018 | Tornier, Inc | Data has shown that degradation of a raw material used in Phantom Fiber may occur more quickly than expected if improperly stored. This may impact the shelf-life of the product which could cause the suture to lose strength more quickly than expected. | Nonconforming Material/Component | The firm, Wright, sent an "Urgent Field Safety Notice" letter by email on 02/08/2018. The firm followed up with a letter on April 02, 2018 disseminated by FedEx Priority mail to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following actions:" Cease using the affected devices immediately." Medical facilities are to determine if any of the recalled implants are on hand, and return affected implants immediately to Wright Medical."Review the notice and complete the Acknowledgement Section Attachment C to signify that your facility has been informed of this Field Safety Notice (Recall)."Return the completed Acknowledgement to Wright Medical via email to: FieldAction@wright.com within five (5) working days of this notice."Contact customer service at 1-888-494-7950 to facilitate any product returns.For any inquiries regarding the Phantom Fiber Field Safety Notice (Recall) contact: Wright Medical Quality Systems Manager by Email: FieldAction@wright.com or Phone: 952.683.7482. |  | 32253 | 32253 units | US Nationwide Distribution. | Open, Classified |  |  | 32253
Z-1378-2018 | 79711 | No | MQG | Assisted reproduction accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | MINC Benchtop Incubator (Mini Incubator) |  |  | January 12, 2018 | 2018 | Cook Inc. | The temperature of the incubation lids of the MINC Bench top Incubators may become too high or too low for proper gamete/embryo incubation due to fatigue failure of the flat flex lid heater. | Device Design | Cook plans to begin notifying consignees of the recall on January 12, 2018Written recall communication letters will be mailed via courier to US consignees |  | 4451 | 4,451 | Nationwide (K-IMNC-100 -US G46022) . (K-IMNC-100 G20079 manufactured in Australia and distributed worldwide) | Open, Classified |  |  | 4934
Z-1387-2018 | 79378 | No | KDN | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cold Storage Solution, 2L, labeled as Waters IGL Cold Storage Solution, Catalog No. CSS/2000/US; and, Belzer UW Cold Storage Solution, Catalog No. BUW_002. |  |  | February 26, 2018 | 2018 | Preservation Solutions, Inc. | Bags of solution may have a small leak at the base of the injection port where the port is welded into the body of the bag.. | Nonconforming Material/Component | An Urgent: Product Recall Notification letters, dated February 26, 2018, were sent to consignees. The instructions include, halting distribution of any 2 Liter bags in inventory; and, arranging for the return of affected units. Preservation Solutions, Inc. provided a notification letter for their consignees to provide to their customers. The letter instructed customers to return any 2 Liter bags to their supplier, and to complete and return the acknowledgement form. |  | 31104 | 31,104 units | Worldwide Distribution -- USA and Europe. | Open, Classified |  |  | 31104
Z-1384-2018 | 79646 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Intra-Aortic Balloon Catheter Mega 7.5FR 30cc |  |  | February 21, 2018 | 2018 | Datascope Corporation | The U.S. Instructions For Use for the Mega Intra-Aortic Balloon (IAB) Catheter was missing the Contraindications section. | Error in labeling | Maquet/Getinge issued a recall letter on February 21, 2018 to affected customers, via FEDEX with confirmation of delivery, to notify them of the fact that the IFU is missing the Contraindications section and to provide the missing information. Affected customers are instructed to examine their inventory immediately to determine if they have affected product. An Addendum to the Mega IAB Catheter IFU, containing the Contraindications, is provided as an attachment in the recall letter. It is suggested that affected customers post copies of the Addendum near all affected lot numbers of Mega IAB Catheters. Customers are provided with Maquet/Getinge contact information should their facility require a full copy of the correct Mega IAB Catheter IFU. Lastly, affected customers are instructed to complete and sign a response form, acknowledging that they have received the notification, and return the form to Maquet/Getinge. |  | 655 | 655 | US Nationwide Distribution | Open, Classified |  |  | 16113
Z-1382-2018 | 79648 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity Acute Care System (IACS) Monitoring Solution, Catalog Numbers: MS18841, MS20401, MS20407, MS20724, MS22956, MS25510, MS25520, MS25643, MS26196, MS26372, MS31818; UDI Information: 4049098054454, 4049098054447, 4049098054409, 4049098009799, 4049098009751 |  |  | February 21, 2018 | 2018 | Draegar Medical Systems, Inc. | The invasive pressure visual and audible alarms may not activate. | Software Manufacturing/Software Deployment | On February 21, 2018 , Drager Medical Systems Inc. mailed an Urgent Medical Device Recall Notification to affected customers, affected customers that informs the end users of the recommended connecting sequence when using the hemodynamic pod with the M540 patient monitor. A Customer Reply Card will be included as a receipt from the customers that they received and acknowledged the recommended sequence when using the hemodynamic pod with the M540 patient monitor.To avoid encountering the identified issue, the user should follow the below identified steps whenever the hemodynamic pod is plugged into the Infinity M540 patient monitor and the unit is placed in Standby. When the user is ready to monitor the patient:1. Exit Standby mode2. Unplug, pause for a minimum of 3 seconds and re-plug the hemodynamic pod3. Continue per your normal procedureIf there are any questions regarding the notification please call 1-800-437-2437 (press 1 at the prompt, then 32349). For questions regarding the operation and/or servicing of your Drager Infinity Acute Care System (IACS) please contact Drager Service Technical Support at 1-800-437-2437 (press 4 at the prompt and then 4 again). For further questions, please call (978) 379-8077. |  | 28317 | 28317 total products | Worldwide Distribution - USA (nationwide) and the following countries:Angola, Argentina, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bolivia, Botswana, Brazil, Brunei Dar-es-S, Bulgaria, Canada, Chile, China, Columbia, Costa Rica, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Japan, Jordan, Kenya, Kuwait, Latvia, Liechtenstein, Lithuania, Macedonia, Malaysia, Maldives, Mexico, Monaco, Mozambique, Namibia, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Poland, Portugal, Qatar, Reunion, Russian Fed., Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Trinidad and Tobago, Turkey, Turkmenistan, United Kingdom, Utd.Arab.Emir. and Vietnam. | Open, Classified |  |  | 28317
Z-1386-2018 | 79653 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | SOUTHERN IMPLANTS REF IBR12d-13, Implant, Ext. Hex, tapered 12o Co-Axis(TM) intended for surgical placement in the upper or lower jaw to provide a means for prosthetic attachment of crowns, bridges or overdentures utilizing delayed or immediate loading. |  |  | March 08, 2018 | 2018 | Southern Implants, (Pty.) Ltd. | The UDI label correctly reads IBR12d-13, however, the inner label incorrectly read IBNT12d-13. | Process control | The firm initiated their recall by email on 03/08/2018. The firm requested return of the product. |  | 49 | 49 units | US, South Africa, France | Completed |  |  | 49
Z-1381-2018 | 79684 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Segmental System Proximal Femoral Component 38mm Offset, Model Number 00-5850-030-38 |  |  | March 29, 2018 | 2018 | Zimmer Biomet, Inc. | There is a potential for debris in the hole on the superior lateral aspect of the device as a result of machining which may not have been adequately removed during the subsequent cleaning process. | Under Investigation by firm | Urgent Medical Device Recall letters and emails were distributed on March 29, 2018. Distributors Responsibilities:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. International distributors can request a International Return Authorization by emailing zimmerbiometintlirarequests@zimmerbiomet.com for returnsc. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.d. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.e. Mark RECALL on the outside of the returned cartons.4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com.a. Review the list of hospitals and/or surgeons included with the email notification sent to your facility, which includes a list of hospitals and/or surgeons that have already been notified of this recall.b. Identify whether there are any additional hospitals and/or surgeons that Zimmer Biomet has not notified and list these accounts on the Additional Accounts form. Please provide the form in Excel format.c. If there are no additional accounts or surgeons to notify, please indicate that there are no additional accounts, or indicate None or NA on the form.5. Retain a copy of your Inventory Return Certification and product return forms for your records in the event of a compliance audit of your facility.6. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating ho |  | 1145 | 1,145 distribution events. Note that this product is lot controlled and may be distributed multiple times if it was returned and returned to the shelf | Worldwide Distribution - USA (nationwide) to the states of : AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY., and to the countries of : Australia, South Korea, Malaysia, Netherlands, and Singapore. | Open, Classified |  |  | 1145
Z-1379-2018 | 79641 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge AGS, FSLC, FSUC, RC accessory to Washer- Disinfector 86 Series UDI Info: 86 Series Washer Disinfector 07340153700246, 07340153700253, 07340153700260, 07340153700277 |  |  | March 02, 2018 | 2018 | Getinge Disinfection Ab | There is potential for a cart with disinfected goods to unload without the trolley being docked. Disinfected goods could end up falling to the floor causing risk to nearby users and damage to goods and equipment. | Software design | A letter (Urgent Medical Device Recall Notice) was issued March 2, 2018 via 2 Day Fedex with customer signature required to allow for tracking of delivery. The letter includes a customer response form requesting feedback including if device is still in use at the facility. The letter instructs the customer to contact Getinge Service to schedule a visit and service by a trained Getinge technician at no charge to the customer. In addition Getinge Service will proactively contact affected customers and perform the correction when affected installed devices are identified.Recalled product will be corrected by a trained Getinge Service Representative. The correction will involve a software update to remove any risk to operators. Additionally a preventive hardware update that provides a longer wearing end stop pin will be performed during the field correction. For further questions please call (585) 214-6049. |  | 170 | 170 total products | Worldwide Distribution - USA (nationwide) and the following countries: AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, CZECH REPUBLIC, GERMANY, DENMARK, NETHERLANDS, NORWAY, SPAIN, FINLAND, FRANCE, GREAT BRITAIN, HONG KONG, ITALY, JAPAN, KOREA, KUWAIT, MALAYSIA, NORWAY, NEW ZELAND, SAUDI ARABIA, SWEDEN, SINGAPORE, SLOVENIA, THAILAND | Open, Classified |  |  | 170
Z-1390-2018 | 79558 | No | OEQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | Enforcement Discretion | Fluid Management Convenience Kits (Angioplasty Kits) intended to be used in Fluid Management |  |  | January 18, 2018 | 2018 | Angiodynamics Inc. (Navilyst Medical Inc.) | NAMIC Fluid Management Convenience Kits contains a Fluid Delivery Set that was not assembled in accordance with specification, and prevents the device from being used, resulting in a minor delay in procedure. | Under Investigation by firm | Recall Notifications were delivered by Fed Express. Consignees instructed to segregate and return all affected devices to Angiodynamics and complete and return the Reply Verification Tracking Form. |  | 38 | 38 units | Recall conducted to end user level. Recall Notifications were delivered by Fed Express. | Terminated | September 21, 2018 | 246 | 38
Z-1398-2018 | 79766 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR Avant-garde, Digital Linear Accelerator, Model No. 05863472 |  |  | February 14, 2018 | 2018 | Siemens Medical Solutions USA, Inc | Control Console software has been updated to reduce the risk for collision when delivering automatically sequenced treatments with automatic movements of the gantry and/or the treatment table. | Software design | On February 16, 2018, the recalling firm issued a Customer Advisory Notice via USPS Certified Mail to inform customers about the update of the Control Console software version 9.2.400 of the Digital Linear Accelerator to version 9.2.502. Customers were informed of the issues that would be fixed through this software update. Consignees were asked to instruct affected personnel accordingly, place the safety advisory notice in the device's instructions for use, and have personnel maintain awareness until the fix is implemented. |  | 1 | 1 | Worldwide Distribution- US Nationwide in the states: AL, CA, FL, GA, IL, IN, KY, MA, MI, MO, MS, NJ, NY, OH, OK, PA, RI, SC, TN, TX, WI, WV, and WY; and PR. International distribution to Nassau. | Open, Classified |  |  | 61
Z-1392-2018 | 79696 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT, Model No. 882442 |  |  | February 15, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | Following a period of inactivity, the mass storage device may cause the acquisition console to become unresponsive. This may prevent an acquisition from proceeding. | Under Investigation by firm | On February 15, 2018, the recalling firm notified customers of the recall via letter. The notice stated that to avoid the hazard of performing a rescan or reinjection, all users must be informed not to view studies or access data on theAddonics MSD while a patient is being scanned. If a user is attempting to view information stored on the Addonics MSD and theconsole appears to be unresponsive, reset power on the Addonics MSD or pullout and reconnect its USB cable.It is recommended that the Addonics USB Mass Storage Device be turned offwhen not being accessed, and powered on only when data transfer is needed.Following data transfer, the Addonics USB Mass Storage Device should beturned off.Customers were instructed to inform those who need to be aware within their organization or any organization wehre potentially affected devices have been transferred; transfer the notice to other affected organizations; and maintain the recall notice with their system Instructions For Use until the correction is made to the system. Customers were advised that Philips will be providing an updated Mass Storage Device that does not experience the same issue that will be provided at no additional cost to the customer. The Field Correction will be released in Q1 2018.Customers who need further information or support regarding the recall were directed to contact their local Philips representative. Customer Care Solutions Center: 1-800-722-9377. |  | 12 | 12 | Worldwide Distribution - US Nationwide in the states of: OH and OR, and Foreign distribution to Australia, China, Cuba, France, and South Africa. | Open, Classified |  |  | 12
Z-1396-2018 | 79759 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 with 4.1.7 XX026 software version model number 728231 |  |  | February 21, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | Surview scan with tube current lower than 30 mA is unable to be initialized. | Under Investigation by firm | On February 21, 2018 Philips Medical Systems sent a Customer Information Letter regarding the Philips Brilliance 64 Software Model 4.1.7. Customers were informed of how to identify affected products and instructed to use the Instructions For Use recommended 30mA for surview scans. Customers with questions or concerns should contact their local Philips representative or local Philips Healthcare office Customer Care Solutions Center at 1-800- 722-9377. |  | 291 | 291 | Worldwide Distribution - US Nationwide in the states of GA, VA, MI, AZ, and the countries of Canada, Switzerland | Open, Classified |  |  | 291
Z-1391-2018 | 79628 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Solid state x-ray imager (flat panel/digital imager) |  |  | January 12, 2018 | 2018 | NeuroLogica Corporation | While transporting the portable machine with the retrofit equipment in a backwards motion, there is potential for foot injury from the underside of the detector bin of the retrofit unit. | Device Design | A field safety notice was sent stating that the manufacturer, without charge, will remedy the defect. The notice stated that field upgrades would be performed by December 31, 2017. The notice provided instructions to end users and to distributors. For further questions, please call (978) 564-8503. |  | 16 | 16 units | US Distribution . Only one consignee was shipped 16 units of the affected device. | Terminated | August 14, 2018 | 214 | 16
Z-1412-2018 | 79666 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 16 Slice (Air) - Model no. 728246, Computed Tomography X-Ray system |  |  | February 08, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | Closing the Emergency Stop may disable the horizontal motor, allowing the couch top to move during patient load or unload activities. | Software design | The recalling firm notified customers of the recall on February 8, 2018, via Urgent Field Safety Notice. The notice advised customers of the issue and affected products and lot numbers. Customers were advised that once the E-stop has been closed, the operator/technologist must initate a brief horizontal motion (in or out) using any of the operator control panels, CT-Box or tape/foot switches. This action will ensure the couch carbon top (upper pallet) is in a controlled state. Customers were informed that a Field Service Engineer would contact them to schedule site updates to correct the issue.Customers who need further information or support were advised to contact their local Philips representative. Additionally, they may contact the Customer Care Solutions Center (1-800-722-9377. |  | 2 | 2 | Worldwide Distribution - U.S. Nationwide. International distribution to Canada, Mexico, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Cambodia, Chile, China, Colombia, Croatia, Cuba, Czech Republic, Denmark, Djibouti, Dominican Republic, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Greece, Guyana, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Republic of Korea, Laos, Lebanon, Lithuania, Macedonia, Malaysia, Maldives, Mali, Martinique, Monaco, Mongolia, Myanmar, Nepal, Netherlands, New Calcedonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Rwanda, Saint Pierre and Miquelon, Samoa, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 1790
Z-1388-2018 | 79603 | No | GHQ | Fetal hemoglobin assay. | Hematology | Hematology | 510(K) Exempt | Human Fetal Hemoglobin (HBF-1) FITC Conjugate |  |  | February 07, 2018 | 2018 | Life Technologies Corporation | Three lots of Human Fetal Hemoglobin FITC Conjugate are displaying high background and unusual bright staining in the controls and patient samples which could potentially lead to erroneous results (i.e. false positives) | Use error | All consignees which have purchased the Human Fetal Hemoglobin FITC antibody including facilities/customers that have shipped the lots outside of their facility will be notified via certified mail to check inventory to see if affected product is on hand. Consignees instructed to discontinue use and discard per internal procedure. The effectiveness check will begin 10 days after the distribution date of the Medical Device Recall Notification. The field removal is considered effective when 10% of the affected customers acknowledge receipt and understanding of the notification instructions. |  |  |  | MI, CA, IL, TX, SC, NY, MN, DE, NJ, WA , OR, IN, MA, DE, FL, PA | Open, Classified
Z-1403-2018 | 79670 | No | BZA | Airway connector. | Anesthesiology | Anesthesiology | 510(K) Exempt | Teleflex MEDICAL HUDSON RCI One Way Valve w. Capped Monitoring Port, REF 1644, QTY 50 |  |  | March 27, 2018 | 2018 | Teleflex Medical Europe Ltd | These one-way valves may disconnect at the joint between the two components that make up the device, which would cause an interruption of ventilation to the patient. | Device Design | Teleflex initiated the recall by letter on 03/27/2018. The firm directed the consignee as follows:"1. If you have affected stock, immediately discontinue use and quarantine any products with the product code and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document your receipt of this letter." |  | 103910 | 103910 units | US Nationwide in the states of AU, CA, NZ | Open, Classified |  |  | 103910
Z-1395-2018 | 79710 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | GE Healthcare: |  |  | March 16, 2018 | 2018 | GE Medical Systems Ultrasound &amp; Primary Care Diagnostics, LL | Under certain conditions, when using DICOM Worklist along with DICOM MPPS, a report for a bone density exam may be sent to PACS with the incorrect patient information in the DICOM header. The correct patient information will be listed on the DICOM report image; however, the report may appear under a different patients name in the PACS. | Software design | The firm initiated their recall by letter on March 16, 2018. The notice stated:"Under certain conditions, when using DICOM Worklist along with DICOM MPPS, a report for a bone density exam may be sent to PACS with the incorrect patient information in the DICOM header. The correct patient information will be listed on the DICOM report image; however, the report may appear under a different patients name in the PACS. There have been no injuries reported as a result of this issue.You can continue using your system by disabling the DICOM MPPS feature as follows:1. In enCORE, select Tools ¦ Connectivity Options.2. Select DICOM tab.3. Press Configure Connections in DICOM Worklist section.4. In the MPPS SCP Settings, clear the AE Title and the TCP/IP Address fields5. Press OK to all.A GE Healthcare representative will contact you to arrange for the correction."For further questions, please call (608) 826-7492 |  | 303 | 303 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and to the countries of : Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Egypt, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Lebanon, Malaysia, Malta, Mexico, Morocco,Norway, Poland, Portugal , Romania, Russia, Saudi Arabia, Spain, Switzerland, Taiwan, Tunisia, Turkey, United Arab, United Kingdom and Venezuela. | Open, Classified |  |  | 372
Z-1424-2018 | 76620 | No | FSH | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | JAMSHIDI (TJ) NEEDLE BONE MARROW [a.) 13G x 3.5, b.) 8G x 4, c.) 11G x 4, d.) 8G x 6, e.) 11G x 6]; STERILE; |  |  | February 23, 2017 | 2017 | Carefusion 2200 Inc | The integrity of the sterile packaging is potentially compromised. | Package design/selection | On 2/23/2017, firm notified customers of recall via mailed letter. The letter included a recall response form. Customers were directed to either return the affected product along with the response form or destroy it and return the response form, indicating the product had been destroyed. All affected product will be destroyed. |  |  |  | Nationwide, Australia, Brazil, Canada, China, Japan, Kuwait, Malaysia, Mexico, Netherlands, New Zealand, Philippines, Singapore, South Africa, Taiwan, Thailand, United Arab Emirates, Vietnam | Open, Classified
Z-1427-2018 | 79531 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | Halyard Closed Suction System for Adults, Double Swivel Elbow (DSE). Used to aspirate liquids or semisolids from a patient's upper airway. Labeled as the following: |  |  | February 07, 2018 | 2018 | Halyard Health, Inc | Certain Flex Connectors supplied with Halyard Closed Suction Kits with Flex Connector may become loose or disconnect before use or during use. | Labeling design | Customers were notified via mail and/or email on about 02/07/2018. Instructions for distributors included to notify all affected customers who may have received potentially impacted products and to complete and return the Distributor Acknowledgement Form. Instructions for customers included to ensure adequate connection of the Flex Connector to Swivel Connector before use of the product by following. This accomplished by ensuring a tight fit is achieved on the Swivel with the Flex Connector advanced approximately three-quarters over the Swivel, so a gap remains visible. If and adequate connection is observed, the Flex Connector may be replaced with a new, sterile Flex Connector and the system connection should be re-evaluated. Lastly, if use of a Flex Connector with the ventilator circuit is optional, the closed suction Swivel Connector can be connected directly to the ventilator circuit. During use, if a disconnection between the Flex Connector and Swivel Connector is observed during use, replace the Flex Connector with a new, sterile Flex Connector (e.g. HYH Product # 1115) and ensure adequate connection as described in this notice. If an inadequate connection occurs, Halyard Health instructs customers to safe the Flex Connector and report it via the Product Complaint Process (PIQ) and the local Halyard Health Sales Representative will assist. Additionally customers were instructed to complete and return the Acknowledgement Form and distribute the recall notice to all clinicians in the facility that may be impacted. |  | 20722634420 | 32,189 cases (20 units per case - 643,780 units) total | Worldwide distribution. US nationwide including Puerto Rico, ARGENTINA, AUSTRALIA, BAHAMAS, BELGIUM, BOLIVIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DOMINICAN REP., ECUADOR, EL SALVADOR, GUATEMALA, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, LEBANON, MALAYSIA, MEXICO, NEW ZEALAND, PERU, PHILIPPINES, QATAR, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, THAILAND, UNITED KINGDOM, and UNITED ARAB EMIRATES. | Open, Classified |  |  | 20722634420
Z-1419-2018 | 79534 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | Cocaine Metabolite MICRO-PLATE EIA, Catalog Number: 1122SC, UDI: (0100608337000429(17)180228(10)000 |  |  | December 20, 2017 | 2017 | OraSure Technologies, Inc. | The controls and calibrators contained within the cocaine test kit may cause false positive results with the negative control. | Under Investigation by firm | The email to the firm's consignees states the controls will be replaced. |  | 10 | 10 | Distributed to 2 consignees in KS. | Open, Classified |  |  | 10
Z-1417-2018 | 79652 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Sarns TCM, Catalog No. 15747 |  |  | March 16, 2018 | 2018 | Terumo Cardiovascular Systems Corporation | Update to cooler-heater cleaning instructions. | Labeling design | The firm,Terumo Cardiovascular Systems Corporation, mailed customers an "Urgent Medical Device Correction" letter dated 2/16/18. The letter described the product, problem and actions to be taken. The Customers were instructed to do the following: 1) discard the current Sarns TCM, TCM II Cooling and Heating, or tHX2 Temperature Management Systems cleaning guides; 2) Review the notice and updated cleaning guides, and ensure all users receive notice of the updated cleaning guides and; 3) Return the Customer Response Form acknowledging receipt of the notice via Terumo Recall Email: tcvs.recall@terumomedical.com or Terumo Recall Fax: 734-741-6149. The letter also includes a list of the updates to the cleaning guides.Customers with questions may contact Terumo CVS Customer Service (800) 521-2818, Monday-Friday, 8am-6pm ET or Fax (734)741-6149. |  | 2393 | 2393- total for all products | Worldwide Distribution: US (Nationwide), Australia, Belgium, Canada, Chile, China. Colombia, Dominican Republic, England, Finland, France, Germany, Greece, HK, India, Indonesia, Iran. Israel, Italy, Japan, Korea, Mexico, Netherlands, New Zealand, Norway, Philippines, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates,and Vietnam. | Open, Classified |  |  | 2393
Z-1420-2018 | 79627 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Product: Milford Smartbase |  |  | March 07, 2018 | 2018 | Autochair | Recalling firm was informed that as a result of combination of a weak spring and incorrect alignment of the locking pin for latching the leg positions, product legs folded inwards while lifting a person off a bed and the individual was dropped. | Component design/selection | Autochair will issue a Retrofit Failsafe Kit. In the meantime, please quarantine current stock of SmartBase products until the failsafe device has been fitted and tested. |  |  |  | PA, and worldwide | Open, Classified
Z-1430-2018 | 79451 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Multiva systems, 1) Multiva 1.5T 8 R5 Model 781072, 2) Multiva 1.5T 16 R5 Model 781073, 3) Multiva 1.5T 8ch R5 based Model 781074, 4) Multiva 1.5T Model 781076, &amp; 5) Multiva 1.5T 16ch R5 based Model 781078 |  |  | March 16, 2018 | 2018 | Philips Electronics North America Corporation | Potential risk for helium gas inside the MR examination room during a magnet quench | Nonconforming Material/Component | On March 16, 2018, Philips Healthcare distributed URGENT MEDICAL DEVICE CORRECTION notices to their customers via certified mailing. On May 24, 2018, the recall letter was distributed to additional affected customers identified by the firm. **Action to be Taken by Customer**:During a quench, a large amount of helium evaporates and is vented outside the building through a venting system. A quench causes immediate removal of the static magnetic field. A quench can occur spontaneously or can be induced if there is an emergency. In case helium gas escapes into the examination room during a magnet quench ensure strict adherence to the Instructions for Use: Release of helium gas in the examination room"If helium gas is not vented properly after the Magnet Emergency Off button is used or during a quench (for example if the helium vent pipe is blocked) a high concentration of helium gas may build up in the examination room. This gas forms clouds of cold mist. Helium gas dilutes the oxygen in the air. High concentrations of helium gas can lead to suffocation.If helium enters the examination room:" Immediately remove all persons from the examination room." Do not switch off air circulation and ventilation in the examination room." Do not reenter the examination room until it is confirmed that the air oxygen content is at a safe level."Emergency procedures"The User is required to establish emergency procedures for the following situations:" A medical emergency" A fire" An emergency that requires immediate removal of the magnetic field" The release of helium gas into the examination roomPhilips MRI systems have an Emergency Table Stop button in case there is an emergency during tabletop movement. (Instructions for Use R5.3)Philips will schedule an inspection of all MRI systems that may be affected by this notice. If applicable the spare metal burst disk on site will also be replaced during this inspection. Metal Burst disks on all aff |  | 8205 | 8,205 units in total | Worldwide Distribution - US Nationwide; International - Albania, Algeria, Andorra, Argentina, Australia, Austria, Bahamas, Bahrain, Bangladesh, Barbados, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Botswana, Brazil, Bulgaria, Burkina Faso, Chile,China, Colombia, Congo (Democratic Republic of the), Costa Rica, CÂżte D'Ivoire, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic,Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Gabon, Georgia, Germany, Ghana, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary,Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jersey, Jordan,Kazakhstan, Kenya, Republic of Korea - South Korea, Kuwait, Kyrgyzstan, Latvia, Lebanon, Lesotho, Libya, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Mauritius, Mexico, Monaco, Mongolia,Montenegro, Morocco, Myanmar, Nepal, Netherlands, New Caledonia, New Zealand, Nicaragua, Nigeria, Norway, Oman, Pakistan, Palestine, State ofPanama, Papua New Guinea, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Rwanda, Saudi Arabia, Serbia, Singapore,Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sudan, Sweden, Switzerland,Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Venezuela,Viet Nam, Yemen &amp; Zimbabwe. | Open, Classified |  |  | 8205
Z-1461-2018 | 79713 | No | MQG | Assisted reproduction accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cook Vacuum Pump ( Cook Aspiration Unit) K-MAR-5200-US, GPN G49275. |  |  | March 22, 2018 | 2018 | Cook Inc. | Pump was found to be constructed from crimp connectors that do not meet the fixation wiring requirements and can cause failure of the device to operate or electric shock or burn to the user. | Nonconforming Material/Component | An Urgent: Medical Device Correction letter, dated March 20, 2018, was sent to consignees. The letter identified the affected device and provided a description of the problem. The firm stated that it would be replacing the mains wiring for each unit. An authorized service agent will contact customers to arrange for the affected devices to be corrected. Customers are to follow the instructions regarding the residual current device (RCD) provided in the letter, while they are waiting for their devices to be corrected. Additionally, customers should complete the Acknowledgement and Receipt Form within 5 business days of receiving the letter. Questions or concerns should be directed to Cook Medical Capital Equipment Service at 855-207-7214. |  | 2267 | 2,267 | Worldwide Distribution -- USA, Argentina, Bolivia, Brazil, Canada, Colombia, Guatemala, Mexico, Panama, Peru, Uruguay, and Venezuela. | Open, Classified |  |  | 2399
Z-1444-2018 | 79614 | No | LZD | Total temporomandibular joint prosthesis. | Dental | Dental | PMA | Patient-Fitted Temporomandibular (TMJ) Right Mandibular Component part of TMJ Reconstruction Prosthesis System. Model /catalog # TYY-NNNNRM |  |  | March 12, 2018 | 2018 | TMJ Solutions Inc | This device was mislabeled and is not the correct component for this patient. | Error in labeling | TMJ Concepts sent an Medical Device Recall Communication letter dated March 12, 2018.INSTRUCTIONS:The implants for this case should not be used and should be immediately quarantined. We request that you return the entire case (the anatomical bone model together with both the fossa and mandibular components) to us by FedEx Priority Overnight Delivery. Any bone screws and drills that were provided with the shipment can be retained at your hospital and do not need to be returned. For further questions, please call (805) 650-3391. |  | 1 | 1 | US Distribution to GA. | Open, Classified |  |  | 1
Z-1443-2018 | 79615 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Stage-1 Tissue Punch, Disposable, 4.5mm; catalog number: ISS141K. |  |  | January 24, 2018 | 2018 | Keystone Dental Inc | The cutting edge of tissue punches may be potentially damaged with small dents and nicks. | Under Investigation by firm | An Urgent - Medical Device Recall letter, dated January 31, 2018, was sent to the impacted customers on 02/01/2018. The letter identified the affected device and the reason for the recall. Customers were instructed to immediately cease using the recalled product and return unused products to Keystone Dental. Customer Relations Reply Forms are being utilized to document the communication. Completed Relations Reply Forms should be faxed/ emailed to Customer Support; 866-903-9272 or customersupport@keystonedental.com. |  | 101 | 101 | Worldwide Distribution -- US, to the state of California; and, the countries of Canada and Germany | Terminated | July 30, 2018 | 187 | 101
Z-1478-2018 | 79624 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Flowmeter Module (accessory to Terumo Advanced |  |  | January 19, 2018 | 2018 | Terumo Cardiovascular Systems Corporation | Inaccurrate flow readings. Depending on the degree of inaccuracy, this issue may not be easy for the user to detect during setup or use (for example, following a Flow Probe relocation or manipulation). | Employee error | The firm notified customers via phone alerting them of this affected device, issue, potential hazard, correction, and instructions. The communication also included scheduling an expedited service call for a Field Service Technician to replace the affected Flowmeter Module with a corrected Flowmeter Module. When necessary to avoid delaying or cancelling life-sustaining surgery, users can continue to use the Flowmeter Module while awaiting replacement. Once the issue is recognized by the user, if a replacement Flowmeter Module is available, replacement and reassignment of safety connections of the Flowmeter Module can be accomplished in less than 15 seconds. In the event that a replacement Flowmeter Module is not available, a less common mitigation is the use of a back-up stand-alone centrifugal pump or a stand-alone ultrasonic flowmeter system to provide flow data. Customers should receive a copy of the Urgent Medical Device Recall phone script by e-mail along with a Customer Response Form. The e-mail should be reviewed, and the form completed and returned as indicated. Questions or concerns can be directed Terumo CVS Customer Service at 1-800-521-2818. |  | 7 | 7 | Worldwide Distribution -- USA, including the states of IL, KY, and MI; and, the country of Singapore. | Terminated | July 18, 2018 | 180 | 7
Z-1479-2018 | 79591 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic DxTerity Diagnostic Catheter, JL 4.0, REF DXT5JL40, marked with 6F color code on box. |  |  | March 13, 2018 | 2018 | Medtronic Vascular | The color coded French size indicator on the outer box may incorrectly reflect a 6 French size, rather than the correct 5 French size. The catheters inside the carton are the correct DXT5JL40 item, the inner pouch has correct 5F labeling, information and configuration details. | Packaging | Medtronic initiated consignee communications on 13-Mar- 2018 to Risk Managers and Health Care Professionals. US notifications were sent via 2-Day UPS delivery. OUS communications were sent per locally approved methods.Consignees are asked to quarantine and return any unused potentially affected product to Medtronic. |  | 855 | 855 units | US and AUSTRIA, Canada, Germany, ROMANIA, SLOVENIA, | Open, Classified |  |  | 855
Z-1463-2018 | 79597 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | The humeral stems trials are packaged in a kit for Titan Modular Shoulder System under the following labeling: |  |  | December 20, 2017 | 2017 | Integra LifeSciences Corp. | Incidents of stem trial breakage were reported to the firm suggesting that these fractures all occurred during insertion/impaction or extraction of the humeral stem trial whiling preparing the humeral canal and/or trialing. All cases resulted in a delay in surgery, with a variance of medical intervention required. | Device Design | Firm sent letters to consignees on December 20, 2017. Firm asked consignees to examine inventory and determine if consignee had affected product. Firm promised to send new stem trials along with a return shipping label. Firm asked to complete the Acknowledgment and Return Form and email or fax back to firm. For further questions, please call (609) 275-0500. |  | 1619 | 1619 units | Worldwide Distribution - USA (nationwide) Distribution. | Open, Classified |  |  | 1619
Z-1465-2018 | 79681 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | M/DN Intramedullary Fixation Humeral Guide Wire - Smooth 2.4 mm Diameter 70 cm Length Single Use Only; Item Number: 00-2255-025-00 |  |  | March 26, 2018 | 2018 | Zimmer Biomet, Inc. | Potential failure of sterile packaging seal. | Under Investigation by firm | Zimmer Biomet will conduct this recall to distributors and hospital risk managers that have consigned inventory or inventory within their territory, as well as distributors that have accounts in their territory that received a direct sale of affected products. Affected products that remain in consignment will be referred to as fielded inventory throughout the remainder of this recall strategy.All distributors will be notified via email. Hospital risk managers, as well as distributors, with direct sales and/or fielded inventory, will be notified via courier. Notification letters were sent to customers on March 26, 2018." Distributors will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the product in their territory, as well as identifying hospitals who have fielded inventory." Distributors will return on-hand product to Zimmer Biomet and ensure all affected fielded inventory is accounted for using the form provided in the letter." Hospital risk managers will be provided with a letter identifying the issue and their responsibilities. These responsibilities include:o Assisting the Zimmer Biomet sales representative with the quarantine of fielded inventory and completion and return of the Attachment 1 form.o Complete and return Attachment 1 (Acknowledgement of Responsibility) to email address provided. |  | 19483 | 19483 total products | US Nationwide and the following countries: Brazil, Singapore, Canada, Australia, Japan, Malaysia, Chile, Netherlands, Trinidad | Open, Classified |  |  | 19483
Z-1480-2018 | 79649 | No | PQP |  |  | Pathology | PMA | Oncomine Dx Target Test, (Model No. A32451), Oncomine Dx Target RNA/DNA Panel (Model No. A32441). |  |  | February 08, 2018 | 2018 | Life Technologies Corporation | The lower concentration of the RNA panel could cause a NO CALL or FALSE NEGATIVE for the reporting of ROS1 fusion. If a physician relies on a false negative result, crizotinib could be improperly excluded from the patient's treatment regimen. | Process control | The firm issued an Urgent: Medical Device Recall Notification, dated February 16, 2018, via e-mail and phone call to affected consignees. The letter identified the affected device and advised customers of the potential "No Call" or false positive outcomes. Users were advised that they may need an alternate method to determine the presence of ROS1 variant in previously tested patient samples, and that any future tests using the affected lot cannot be considered valid if "No Call" or "Negative" is the result. Requested action included reviewing previous results which used the affected lot. Customers are to complete the Customer Response Sheet provided and e-mail a scanned copy along with the letter to vineet.bansal1@thermofisher.com.Questions and comments should be directed to Vineet Bansal, 408-802-2341, vineet.bansal1@thermofisher.com. |  | 30 | 30 | Distributed to states: AZ, CA, NC, TX, and VA. | Open, Classified |  |  | 30
Z-1462-2018 | 79600 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Revision Knee Catalog #5545-A-301 and Catalog #5546-A-601 |  |  | January 30, 2018 | 2018 | Howmedica Osteonics Corp. | Customers may receive the incorrect size implant from what is labeled on the box. | Under Investigation by firm | On January 31, 2018 branches/agencies received notification of the recall via e-mail and were asked to quarantine affected devices. Notification Letters and Product Accountability Forms were sent to branches, agencies &amp; hospital risk management via UPS (with return receipt) on January 31, 2018 with an expected arrival date of February 02, 2018. Customers were asked to complete a Notification Business Reply Form and return via e-mail or fax. Affected product is to be returned to the manufacturer. Customers with questions may contact (201)831-6693. |  | 77 | 77 total products | USA (nationwide) Distribution to the states of : AL, AZ, CA, FL, GA, KS, MA, MD, MI, NC, NJ, NY, OH, PA, SC, TN, TX, VA, and WA | Open, Classified |  |  | 77
Z-1464-2018 | 79798 | No | EMO | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Double Blade Shoe F/Unigraft Knife 1.5 Thick X 9MM SKU 6500009 |  |  | March 28, 2018 | 2018 | Ace Surgical Supply Co., Inc. | The product inside the package is incorrect. A single blade shoe was mistakenly packaged instead of a double blade. | Error in labeling | ACE is contacting the 5 affected customers via certified mail for US customers and via email for non-US Customers. Follow-up notifications will be sent 12 weeks after initial notification. Firm will rework incorrectly packaged product and will use completed reply forms as effectiveness check. For further questions, please call (508) 588-3100. |  | 6 | 6 | US Distribution to the states of : MI, CA, OK, and Internationally to China, Canada, | Terminated | September 11, 2018 | 167 | 6
Z-1361-2018 | 79669 | No | REU |  |  |  |  | Crimson Trace Corporation laser sights. |  |  | October 10, 2017 | 2017 | Crimson Trace Corporation | A previously "in specification" Finished Product with a Model 25-0657 diode had power output higher than 5 mW. | Radiation Control for Health and Safety Act | Crimson Trace Corporation (CTC) will bring defect into compliance: 1.CTC will send notification letters or emails to all relevant dealers, distributors, and purchasers,2.CTC will provide return merchandise authorization numbers and shipping labels to all relevant dealers, distributors, and purchasers,3.CTC will replace all returned product units with compliant diodes,4.The corrections will be conducted at no cost to the purchaser, and5.The corrective action will be completed by January 31, 2019.For further questions, please call (800) 442-2406. |  | 184390 | 184,390 | USA (nationwide) Distribution | Open, Classified |  |  | 184390
Z-1487-2018 | 79636 | bandage, cast | ITG | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | TCC-EZ, Total Contact Cast System, Model No. TCC2-4051 |  |  | August 18, 2017 | 2017 | Integra LifeSciences Corp. | A quarantined lot of the devices was released for sale. Devices from this lot may experience increased exothermic reaction causing higher levels of heat output which exceed the product specifications. | Process control | On 8/18/17, the recalling firm's sales reps notified affected consignees of the recall and distributed customer letters. Customers were asked to examine their inventory for product from the affected lot, warned not to use the product, complete the included "Acknowledgement and Return Form", and return the form by email or fax. Customers were provided with a Customer Service number: 1-800-654-2873. |  | 250 | 250 | US nationwide distribution. | Terminated | July 19, 2018 | 335 | 250
Z-1496-2018 | 79629 | No | DAK | Complement components immunological test system. | Immunology | Immunology | 510(k) | MicroVue CIC-C1q EIA, Model A001. An enzyme immunoassay for the detection of circulating immune complexes in human serum or plasma. |  |  | April 04, 2017 | 2017 | Diagnostic Hybrids, Inc. | There is a potential for the slope of the standard curve to fail to meet the assay validation requirements for the assay. | Other | Customers were notified via letter on 4/4/17. Customers were asked to review their current inventory for the specified lot and complete and return the response form. If the product was further distributed, customers were instructed to provide them with the notification letter. |  | 101 | 101 kits | The products were distributed to the following US states: NY, TX, CO, UT. The products were distributed to the following foreign countries: Switzerland. | Open, Classified |  |  | 101
Z-1502-2018 | 79779 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General Hospital | 510(k) | T4 Pullover Toga, Regular ; Catalog Number: 0400-750-000 |  |  | February 23, 2018 | 2018 | Stryker Instruments Div. of Stryker Corporation | Separation of material layers may occur, causing a potential risk of exposure to contaminants. | Nonconforming Material/Component | On March 7, 2018 Stryker Instruments mailed Urgent Medical Device Recall Notifications to affected customers. Distributors and Sales Representative were notified via e-mail. Customers were instructed to:1) Immediately review this Recall Notification; 2) Immediately check all stock areas and/or operating room storage for affected products. Quarantine and discontinue use of any affected T4 and T5 Togas; 3) Complete the enclosed Business Reply Form (BRF) to confirm receipt of this Notification and identify how many affected items are currently in your inventory. Please complete and return the BRF even if you dont have any affected product on hand.Fax the completed BRF to Stryker Instruments at 866-521-2762, or email to kara.spath@stryker.com. Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification; 4) If you have further distributed this product, please forward this Notification and the attached BRF to all affected locations. Please indicate each location on the BRF; 5) If the BRF for your facility indicates that recalled product is currently on hand, a FedEx label will be emailed to you and should be used to return recalled product. Upon receipt of the recalled product, a credit will be issued to your account.Customer with questions or concerns may call (800)253-3210. |  | 34570 | 34570 total products | US Nationwide and Ireland, Japan, South Korea, UK | Open, Classified |  |  | 34570
Z-1511-2018 | 79763 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Flex Arm |  |  | March 05, 2018 | 2018 | Synthes, Inc. | Quick connect feature of instruments may be unable to connect to the male features of Flex Arm Adaptors, Insight Retractors, or other mating parts. | Device Design | The firm sent an Urgent Product Recall letter dated March 21, 2018, to affected customers. The letter identified the affected product, problem and actions to be taken. For questions call 1-610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 84 | 84 | Worldwide distribution - US Nationwide and the countries of Canada, Switzerland and Japan | Open, Classified |  |  | 84
Z-1512-2018 | 79661 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Knee Prosthesis, Sterile, Single-Use devices |  |  | January 12, 2018 | 2018 | Howmedica Osteonics Corp. | Firm became aware that the specific product/lot combinations referenced above may contain the incorrect size implant from what is labeled on the box. | Error in labeling | Branches/Agencies notified by email on 01/16/2018 and were asked to quarantine affected devices. Notification letters and product accountability forms were sent via UPS (with return receipt) on 01/15/2018. Customer notifications were updated and sent to consignees on 03/16/2018 as the technical and medical assessments associated with the issue were completed and additional information related to potential hazards and harms were identified. |  | 6 | 6 units of catalog 5521-B-400 | Affected products were sent to three medical facilities. Affected product will be returned to the manufacturer for disposition. | Open, Classified |  |  | 6
Z-1497-2018 | 79642 | No | DAK | Complement components immunological test system. | Immunology | Immunology | 510(k) | MicroVue CIC-C1q EIA, Model A001. The MicroVue CIC-C1q EIA is for detection of circulating immune complexes (CIC) in human serum or plasma. |  |  | November 03, 2015 | 2015 | Diagnostic Hybrids, Inc. | A component of the kit was found to contain bacterial contamination. | Material/Component Contamination | Customers were notified via letter dated November 3, 2015, to review their inventory for the specified lots, discard the kits, and contact the recalling firm for replacement. Certificates of destruction were also provided to customers. For further questions, please call (740) 589-3300. |  | 19 | 19 kits | Worldwide Distribution - USA (nationwide) to the states of : CO, CT, IA, IL, MN, MO, NC, NY, OH TX, and UT., and to the countries of : Switzerland, India, Hong Kong, and Japan. | Open, Classified |  |  | 198
Z-1485-2018 | 79749 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Genesis Surgical Cassette Tapered |  |  | March 13, 2018 | 2018 | Keystone Dental Inc | The surgical cassette contained extra-large white grommets making the 3.8 mm and 6.5mm dental burrs wobble within the grommets in the surgical cassette. | Error in labeling | Notification was mailed to all impacted customers on March 13, 2018, via overnight Federal Ex-press. Customer Relations will contact the customers via phone to ensure the customers received the notification, instructions, and replacement grommets. Customer Relations Reply Forms are being used to document the communication. For further questions, please call (866) 902-9272. |  | 19 | 19 | USA (nationwide) Distribution and Internationally to Turkey. | Terminated | July 23, 2018 | 132 | 19
Z-1486-2018 | 79678 | No | LTK | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur CA 125II Assay; Cat No. 09427226, SMN 10310443 (1-pack/100 tests); Cat no. 01678114, SMN 10315696 (5-pack/500 tests) |  |  | February 21, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc | Upon dilution, some patient samples may exhibit over-recovery outside the representative data provided in the assay instructions for use. | Other | A Customer Notification was sent to customers on 2/21/18. The letter instructs customers to refer to the information provided until the appropriate updates are available in the assay instructions for use. The letter is to be retained and forwarded to those who may have received the product. |  | 1031044300 | SMN 10310443 (1-pack/100 tests) - 13257 kits; SMN 10315696 (5-pack/500 tests) - 9930 kits | US Nationwide, and the following countries: Albania, Angola, Argentina, Armenia, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia Herzegovina, Brazil, Bulgaria, Cambodia, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Latvia, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Mexico, Morocco, Myamar, Nepal, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan,Thailand, Turkey, United Arab Emirates, Ukraine, United Kingdom, Uruguay, Uzbekistan,Vatican City, and Vietnam. | Open, Classified |  |  | 1031044300
Z-1513-2018 | 79514 | No | PEN | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | FilmArray BCID Panel (Blood Culture Identification Panel), Catalog numbers RFIT-ASY-0126 (30 test kit) and RFIT-ASY-0127 (6 test kit), when used with all lots of BD BACTEC Blood Culture Bottles, Catalog numbers 442020, 442021, 442023, 442192, 442265, 442194, 442022, 442193, 442260, and 442191. |  |  | March 14, 2018 | 2018 | BioFire Diagnostics, LLC | There is an increased risk of false positive Proteus results when the product is used with specific types of BD BACTEC blood culture bottles. | Material/Component Contamination | The recalling firm initiated the recall on several different dates and methods: Via email on 3/14/2018 with the letter dated 3/3/2018 attached, via USPS on 3/23/2018, and via FedEx between 3/24/2018 and 3/28/2018.The recalling firm conducted an expansion and issued letters dated 6/11/2018 via FedEx or USPS on 6/13/2018. |  |  |  | Distribution was nationwide. There was military/government distribution. Foreign distribution was made to Canada, Algeria, Argentina, Australia, Austria, Bahrain, Botswana, Brazil, Chile, Colombia, Costa Rica, Czech Republic, Dominican Republic, Ecuador, Egypt, France, French Polynesia, Georgia, Germany, Greece, Guatemala, Hong Kong, India, Iraq, Israel, Italy, Japan, Jordan, Republic of Korea, Kuwait, Lebanon, Malaysia, Mexico, Myanmar, The Netherlands, New Caledonia, Nicaragua, Panama, Peru, Republic of the Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, Uganda, United Arab Emirates, United Kingdom, Uruguay. | Terminated | October 24, 2018 | 224
Z-1521-2018 | 79861 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | AdvanceÂż 35LP Low-Profile PTA Balloon Dilatation Catheter; Catalog numbers: PTA5-35-135-6-2 .0 and PTA5-35-135-4-4 .0; UDI: 10827002522647 and 10827002522524 |  |  | April 16, 2018 | 2018 | Cook Inc. | Incorrect product labeling. Product labeled as 6mm x 2cm balloon are packaged with a 4mm x 4 cm balloon label and vice-versa. | Labeling mix-ups | On April 16, 2018 Cook Medical Inc. mailed an Urgent Medical Device recall notification to affected customers. Customers were instructed to: 1) Examine your inventory immediately to determine if you have affected product(s), and quarantine affected product(s). Immediately cease all distribution and use of these products; 2) Return the affected product(s) to Cook Medical with a copy of the Acknowledgement and Receipt Form to receive a product credit.NOTE: Unaffected products that are returned will not be credited; 3) Please complete the Acknowledgement and Receipt Form within 5 business days of receiving this letter. Even if you do not have affected product(s) on hand, you must still complete the Acknowledgement and Receipt Form and return it via fax (812.339.7316) or email (FieldActionsNA@CookMedical.com); 4) Immediately report adverse events to Cook Medical Customer Relations by phone at 800.457.4500 or812.339.2235, Monday through Friday between 7:30 am and 5:00 pm (Eastern Time) or by email to CustomerRelationsNA@CookMedical.com Cook Incorporated |  | 40 | 40 total products | Worldwide Distribution - US Nationwide in the states of AZ, DE, FL, GA, ID, IL, KY, MI, MO, NH, NJ, PA, RI, TX, UT and the country of Peru | Open, Classified |  |  | 40
Z-1522-2018 | 79692 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Orthopedic Salvage System (OSS) Modular Arthrodesis 0 Degree Locking Collar, Model Number CP260600 |  |  | April 03, 2018 | 2018 | Zimmer Biomet, Inc. | There is a potential dry blast irregularity on the locking collar screws packaged with the locking collar. | Nonconforming Material/Component | Zimmer Biomet sent an Urgent Medical Device Recall letters to their affected customers. The letters identified the affected product, problem and the actions to be taken. Distributor Responsibilities:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. For each return, complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.c. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.d. Mark RECALL on the outside of the returned cartons.4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com.a. Review the list of hospitals and/or surgeons included with the email notification sent to your facility, which includes a list of hospitals and/or surgeons that have already been notified of this recall.b. Identify whether there are any additional hospitals and/or surgeons that Zimmer Biomet has not notified and list these accounts on the Additional Accounts form. Please provide the form in Excel format.c. If there are no additional accounts or surgeons to notify, please indicate that there are no additional accounts, or indicate None or NA on the form.5. Retain a copy of your Inventory Return Certification and product return forms for your records in the event of a compliance audit of your facility.6. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transf |  | 16 | 16 units (21 distribution events. The product is lot controlled and may be distributed multiple times if it was returned and returned to the shelf) | The product was distributed to the following states: CT, FL, GA, IN, MA, MI, NY, and TX. The product was distributed to the following countries: Malaysia and Netherlands. | Terminated | July 02, 2018 | 90 | 16
Z-1520-2018 | 79595 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Quidel MicroVue Intact PTH EIA, Model 8044. An enzyme immunoassay for the quantification of Intact Parathyroid Hormone (PTH) in serum. |  |  | October 26, 2016 | 2016 | Diagnostic Hybrids, Inc. | There may be a loss of separation between Calibrators A and B which may impact the assay calibration at the lower end of the assay. | Nonconforming Material/Component | Customers were notified via letter and email on 10/26/16. Customers were asked to complete and return the attached response form to the recalling firm. Customers were asked to review inventory, quarantine the products with the specified lot numbers, and, if further distributed, identify and notify any additional customers. |  | 26 | 26 | The product was distributed to the following countries: Singapore, Switzerland | Open, Classified |  |  | 26
Z-1519-2018 | 79526 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Roche Acetaminophen assay |  |  | January 02, 2018 | 2018 | Roche Diagnostics Corporation | The previous version of method sheets for the Acetaminophen assay on the COBAS INTEGRA 400plus/800/800 CTS analyzers and the MODULAR ANALYTICS P module are missing the followingstatement in the Limitations-interference section:"In very rare cases, gammopathy, in particular type IgM (WaldenstrÂżm s macroglobulinemia), may causeunreliable results." | Error in labeling | On January 2, 2018 an URGENT MEDICAL DEVICE CORRECTION letter was issued to customers with notification to the issue and information regarding the updated method sheets and that they were available for use. The COBAS INTEGRA 400 plus/800/800 CTS was updated from affected method sheet version 3.0 to 4.0. The MODULAR ANALYTICS P was updated from affected method sheet version 4.0 to 5.0 No product return is required. Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a weekat 1-800-428-2336 if you have questions about the information contained in this UMDC. |  | 35111 | 35,111 | Nationwide Distribution | Open, Classified |  |  | 35111
Z-1514-2018 | 79698 | No | HWB | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stem Extractor f/Guide Bar, Part number 03.401.072. Instrumentation for extraction and revision of Epoca shoulder implants. |  |  | March 19, 2018 | 2018 | Synthes, Inc. | There is a potential for the connection screw of the stem extractor to break. | Device Design | The firm, DePuy Synthes, sent an "URGENT FIELD SAFETY NOTICE" dated 3/19/2018 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to the following: Âż Review the instructions listed in this notification.Âż Forward this notice to anyone in your facility that needs to be informed.Âż If any of the affected product has been forwarded to another facility, contact that facility.Âż Keep a copy of this notice.Âż Return the Verification Section (page 3 of this letter) to Synthes by:o Fax: 1-888-731-7954 oro Scan/email: Synthes3091@stericycle.com. Follow the Instructions if connecting screw for the stem extractor breaksIf you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 31 | 31 | Worldwide Distribution: US (nationwide) and to countries of: Canada and Switzerland. | Open, Classified |  |  | 31
Z-1515-2018 | 79755 | No | MVN | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | clariti 1 day/Fresh Day 90 pack (-1.00), UV Blocking, 90 silicon hydrogel, daily disposable contact lenses |  |  | February 21, 2018 | 2018 | CooperVision Inc. | The secondary packaging is labeled with an incorrect expiration date. | Incorrect or no expiration date | The firm, CooperVision, sent an "Firm Initiated Voluntary Removal - ACTION REQUIRED" letter to its customers on 2/21/18. The letter describes the product, problem and actions to be taken. The letters instructed customers to: Examine your inventory for any affected lot numbers; Stop any further distribution of the specified lot(s) of clariti 1 day and Fresh Day contact lenses; discard product and contact CooperVision, Inc. Customer Service department at 1-800-341-2020 to receive credit and/or complimentary replacement product.The removal only affects clariti 1 day lenses with the lot numbers specified on the enclosed letter. No other clariti 1 day and Fresh Day products  or any other CooperVision products  are affected.If you have any further questions, please feel free to contact your Account Representative or Customer Service at 1-800-341-2020, 9:00 AM  7:00 PM ET, Monday - Friday. |  | 43560 | 22 packs/1980 lenses | US Nationwide Distribution to states of: AL, AR, CA, CO, FL, GA, HI, IA, IL, IN, KS, KY, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI, and WY. | Open, Classified |  |  | 1321810
Z-1530-2018 | 79872 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CELLSEARCHÂż CIRCULATING Tumor Cell Kits (IVD) (Product Code 7900001) |  |  | April 04, 2018 | 2018 | Menarini Silicon Biosystems | Unusually high number of total images/unassigned events may require additional time and effort for the operator to review all the images and may cause a delay in reporting the results. | Process design | Menarini silicon biosystems sent an Urgent Medical Device Correction Notification dated April 3, 2018.Required Actions." Please dispose of any affected product in your current inventory ." Report any previous occurrence of unusually high number of total images Iunassigned events in patient samples to Customer Technical Support." Complete and return the enclosed Confirmation of Receipt form no later than April 20,2018. For further question, please call (215) 830-0777. |  |  |  | Worldwide Distribution - USA (nationwide) and to the countries of : Canada and Italy | Open, Classified
Z-1533-2018 | 79687 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD VacutainerÂż Ultra TouchTM Push Button Blood Collection Set 0.6 x 19 mm x 305 mm 23G x Âż x 12. Catalog no 367364 |  |  | March 20, 2018 | 2018 | Becton Dickinson &amp; Company | Products do not meet the labeled sterility claim of a Sterility Assurance Level (SAL) 10-6. | Material/Component Contamination | Immediately review your inventory for the specific Catalog (Ref) and lot numbers, and quarantine product subject to the recall. Immediately discontinue the distribution of the affected product. If you have inventory of the recalled product, please complete the attached Business Response Form and return the recalled product following the enclosed packing instruction. This is required so that BD may process your product replacement.If you do not have the recalled product, please complete the attached Business Response Form as well so that BD receives an acknowledgement of your receipt of this recall notification |  |  |  | US (Nationwide) Canada, Kuwait, Guyana, Guatemala, Taiwan, UK, Belgium, Sweden, Ireland, Israel, Spain, Switzerland and Lithuania | Open, Classified
Z-1529-2018 | 79631 | [6 empty] | Accelerate PhenoTest BC kit, Ref #10101018. The firm name on the label is Accelerate Diagnostics, Tucson, AZ. |  |  | March 14, 2018 | 2018 | Accelerate Diagnostics Inc | The kits may produce an elevated rate of false positive calls for the Staphylococcus aureus identification probe (SAU). | Under Investigation by firm | The recalling firm contacted all customers on 3/14/2018 via telephone, and a follow-up e-mail was also sent containing a technical bulletin #TB000079, entitled Product Correction Notice, dated March 14, 2018. The notice identified the affected devices and the reason for the recall. Customers were asked to please discontinue use of and dispose of all remaining kits from the affected lots, and document the disposal by completing the 'Certificate of Kit Disposal.' Replacement kits will be provided. |  | 95 | 95 kits | Distributed to IL, TX, CA, SC, and DE. | Completed |  |  | 95
Z-1534-2018 | 79668 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Evicel Accessory Device - 45 cm Flexible Tip (SET OF 3), Product Code EVT45. Accessory Tips for use only with EVICEL Fibrin Sealant (Human) Application Device. |  |  | March 07, 2018 | 2018 | Ethicon, Inc. | Products labeled for distribution outside the United States may have been distributed to customers in the United States. | Other | Letters distributed on 3/7/18 to hospitals instructed them to: 1. Examine your inventory immediately to determine if you have the product lot subject to this recall (removal) on hand and quarantine such product(s).2. Remove the product lot subject to this recall (removal) and communicate the issue to all relevant operating room or materials management personnel, or anyone else in your facility who needs to be informed. If any of the devices subject to this recall (removal) have been forwarded to another facility, please contact that facility to arrange return.3. Complete the Business Reply Form (BRF) (Attachment 2) confirming receipt of this notice and fax or email it to Stericycle at 1-877-473-7068 or Ethicon4528@stericycle.com within three (3) businessdays. Please return the BRF even if you do not have the product lot subject to this recall(removal).4. Customers are required to return all unused EVICELÂż 45 CM Flexible (Accessory) Tip devices for the lot subject to this recall (Lot 17-0102 Code EVT45) that are in their inventory immediately. Only unused EVICELÂż 45 CM Flexible (Accessory) Tip devices for the lot subject to this recall (removal) returned by June 30, 2018 will be eligible for replacement. Any unused EVICELÂż 45 CM Flexible (Accessory) Tip devices for the lot subject to this recall (removal) returned after June 30, 2018 will not be eligible for replacement.5. To return unused EVICELÂż 45 CM Flexible (Accessory) Tip devices for the lot subject to this recall (removal), photocopy the completed BRF, place it in the box with the subject product(s), and affix the pre-paid authorized shipping label included with this recall (removal) notification letter. Ethicon will pay for the shipping charges only if the authorized label is used. Extra shipping labels may be obtained by calling Stericycle at 1-877-597-9582 and referencing Event 4528. Your JJHCS number and mailing address have been pre-populated on the BRF.Distributors were instructed to:1. E |  | 738 | 738 devices (246 boxes of 3) | The device was distributed to the following states: AK, AL, AR, AZ, CA, CO, FL, GA, IA, IL MD, MI, MN, MO, NC, ND, NJ, NV, NY, OH, PA, TN, TX, WA, WI, and WV. The device was also distributed to Puerto Rico. | Open, Classified |  |  | 738
Z-1536-2018 | 79700 | No | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew Whipknot Soft Tissue Cinch #5 Sutures. PN: 7211015 |  |  | December 11, 2017 | 2017 | Smith &amp; Nephew, Inc. | Packaging design is insufficient to prevent potential sterile barrier breach (pin hole) in Tyvek film of pouch. A limited number of previously recalled devices were inadvertently redistributed. | Employee error | On 12/11/2017, recall notifications were sent to Hospital Representatives, Affiliates/Distributors, and Sales Personnel. Notified consignees were instructed to inspect inventory and quarantine any affected devices, return the completed Inventory Return Certification Form, and return the affected product. In addition to the above, Affiliates and Distributors were instructed to forwarding the recall letter to customers and asking them to return any recalled product.Consignees were instructed to contact the firm via email at FieldActions@smith-nephew.com |  | 22 | 22 | Distributed to accounts in 7 states: AZ, ME, NY, OH, PA, TN, and VT; and PR. | Open, Classified |  |  | 22
Z-1525-2018 | 79677 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Avenir Muller Stem; Item Number: 01.06010.104, Lot Number: 2905496, UDI: (01)00889024479500 (17)220430 (10)2905496 |  |  | February 28, 2018 | 2018 | Zimmer GmbH | A potential commingle that could result in the product in the package not matching the product on the outer label. | Error in labeling | On March 7, 2018 Field Safety Corrective Action Notices were hand delivered by Zimmer Biomet team members or their representatives. Customers were instructed to complete the following: 1.Review this notification for awareness of the contents.2.Assist your Zimmer Biomet sales representative to quarantine immediately all affected implants.3.Your Zimmer Biomet sales representative will remove the affected implants from your facility.4.Complete Attachment 1  Certificate of Acknowledgement.a. Return a digital copy to fieldaction.emea@zimmerbiomet.com. b. Retain a copy of the Certificate of Acknowledgement with your field action records in the event of a compliance audit of your documentation.5.If after reviewing the notice you have further questions or concerns please contact your Zimmer Biomet representative. |  |  |  | There was no distribution to the US. The distribution occurred to the following countries: Belgium, Switzerland, Germany, France, Canary Islands, Italy, and Thailand. | Open, Classified
Z-1535-2018 | 79453 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Widefield OCT software feature for the Spectralis HRA+OCT and variants |  |  | December 08, 2017 | 2017 | Heidelberg Engineering | FDA notified Heidelberg Engineering that a 510(k) clearance is needed for the SPECTRALIS Widefield OCT feature, and, on November 27, 2017, Heidelberg Engineering submitted a new 510(k) for this feature. FDA has asked Heidelberg Engineering to instruct customers not to use the Widefield OCT software feature on the SPECTRALIS device until 510(k) clearance has been obtained for this software feature. | No Marketing Application | On December 14, 2017 an URGENT Medical device software field correction letter was issued to customers via email or ground mail to notify all customers of the recall action. The letter states the following:FDA has asked us to instruct you not to use the Widefield OCT feature on your SPECTRALIS device until a decision is made on this pending 510(k) submission. Heidelberg Engineering will notify you immediately upon completion of the FDAs review. In the meantime, you may continue to use your Widefield Objective for cSLO fundus imaging (including Fluorescein Angiography, ICG Angiography, MultiColor, Infrared and Red-Free Fundus Imaging). Questions or concerns can be directed to (508) 530-5904 |  | 382 | 382 | Nationally | Open, Classified |  |  | 382
Z-1528-2018 | 79905 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | EEA" Hemorrhoid and Prolapse Stapler Set with DST Series" Technology 4.8MM, Product Number HEM3348, GTIN Number 1 - 10884521047563, GTIN Number 2 - 20884521047560 |  |  | April 09, 2018 | 2018 | Covidien LLC | There is a potential for improper welding of the yellow staple guide to the instrument. Use of a device with an improperly welded staple guide may result in improper staple formation potentially leading to bleeding or anastomotic leak. | Nonconforming Material/Component | Notification letters were distributed on 4/9/18. The letters instructed customers to:Medtronic requests that you quarantine and return any unused products of the item codes and lots detailed above. Unused products from the affected item codes and lots should be returned as described in the Required Actions section below. If you have distributed Covidien EEA" Hemorrhoid and Prolapse Stapler Set with DST Series" Technology listed above, please promptly forward the information from this letter to those recipients. All unused products from the affected item codes and lots must be returned.Required Actions:1. Please quarantine and discontinue use of the affected item code and lots listed on page one.2. Please return affected product as indicated below.3. Complete the Recalled Product Return Form even if you do not have inventory.Purchased directly from Medtronic:Complete the Recalled Product Return Form and fax to: 800-895-6140Ship affected product with RGA# provided by Customer Service to:MedtronicAttn: Field Returns195 McDermott RoadNorth Haven, CT 06473 USAComplete form and check the box indicating no inventory Fax to 800-895-6140 or email to feedback.customerservice@Covidien.comException: Customers with zero inventory, fax to 203-492-7719 or email to FCAMITG@Covidien.comPurchased from a distributor:Complete all fields on the form and contact your distributor directly to arrange for return of productComplete form and check the box indicating no inventory Fax to 203-492-7719 or emailFCAMITG@Covidien.com |  | 475 | 475 | Worldwide Distribution - US Nationwide in the US states of : AL, CA, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OR, SC, TN, TX, UT, and VA. The devices were distributed to the following foreign countries: UAE, Australia, Austria, Belgium, Bulgaria, Canada, Czech Republic, Denmark, Estonia, Finland, France, Germany, Indonesia, Italy, Korea, Malaysia, Netherlands, Poland, Portugal, Russia, Saudi Arabia, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Uganda, UK, and Vietnam. | Open, Classified |  |  | 6022
Z-1526-2018 | 79751 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Hinge Coupling Hoffmann LRF Catalog # 4933-0-800 |  |  | March 12, 2018 | 2018 | Stryker GmbH | The wrong parts were picked for packaging. The connecting Nuts were packed in the packaging labeled with Hinge coupling | Labeling mix-ups | Stryker sent an Urgent Product Recall letter dated March 12, 2018, to all affected customers. .Product will be returned to the manufacturer for disposition.Non-responding consignees will be attempted to be contacted a minimum of three times in total. Returned product will be quarantined and returned to the manufacturer.Product will be returned to the manufacturer for disposition.For further questions, please call (201) 831-5272. |  | 18 | 18 | USA (nationwide) Distribution | Open, Classified |  |  | 18
Z-1523-2018 | 79662 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Versatex Monofilament Mesh 50 x 50cm, Product No. VTX5050M |  |  | February 23, 2018 | 2018 | Sofradim Production | There have been patient reports of abdominal hernia recurrence following hernia repair. The recalling firm has added a statement to the device's Instructions for Use to make surgeons aware of risks in using the device in Transversus Abdominis Muscle Release (TAR) procedures. | Other | All consignees were notified via Federal Express or certifiedmail on February 23, 2018. Customers were advised of the added warning on the product labeling. Customers were requested toacknowledge receipt and understanding of the Urgent FieldCorrective Action Notice. Customers with questions were directed to contact their Medtronic Representative, or Customer Service at 800-962-9888, option 2. |  | 703 | 703 | Worldwide Distribution - USA (nationwide) Distribution to the states of: CA, CT, FL, GA, IL, IN, KY, MA, MD, MI, MN, MO, NC, NH, NJ, NV, NY, OH, PA, TN, TX, VA, WA and WI, and to the countries of : Belgium, Finland, France, Germany, Ireland, Israel, Italy, Netherlands, Russian Federation, Spain, Switzerland, and UK. | Open, Classified |  |  | 703
Z-1894-2018 | 79895 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | iChem VELOCITY Automated Urine Chemistry System, Catalog #700-7176-001 and iChem VELOCITY Automated Urine Chemistry System Computerless, Catalog #700-7177-001. |  |  | March 29, 2018 | 2018 | Beckman Coulter Inc. | The Strip Provider Modules (SPM) on some of the instruments had not been inspected during manufacture for correct seal placement and integrity. | Process control | The recalling firm, Beckman Coulter/IRIS, sent an "URGENT MEDICAL DEVICE RECALL" letter dated dated 3/26/2018 via email on 3/29/2018 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:-Follow the on-board strip stability claims per Table 1 to determine the frequency at which the strips need to be changed.-Whenever the Strip Provider Module is accessed (i.e., opened or when loading strips), remove and discard the remaining on-board strips and desiccant before replenishing.-It is suggested that for the tests performed since the installation of the iChemVELOCITY/SPM housing assembly, consult with your Laboratory Director to determine whether a retrospective review of results is clinically warranted.-Complete and return the CUSTOMER RESPONSE FORM within 10 days by email to: Regulatory.Notifications@beckman.com, Please include FA-33476 iChemVELOCITY Automated Urine Chemistry System in e-mail subject field or by mail to:Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 31-B0611800 S.W. 147th Avenue, P O BOX 169015, Miami, FL 33116-9015.-Beckman Coulter will contact you by phone to schedule a site visit for replacing the SPM on your instrument.-Beckman Coulter will be implementing process improvements to prevent this issue from recurring.If you have any questions regarding this notice, please contact:- From our website: http://www.beckmancoulter.com- By phone: call 800-854-3633 in the United States and Canada. Âż -Outside the United States and Canada, contact your local Beckman Coulter representative. |  | 6 | 6 devices | US Distribution to states of: AL, IA, GA, and FL. | Completed |  |  | 6
Z-1539-2018 | 79908 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | WET-FIELD OSHER THERMODOT MARKER, 25GA FINE TIP, CURVED; Part number 221270. |  |  | February 26, 2018 | 2018 | Beaver-Visitec International Inc. | Incorrect tip was used during manufacturing of the Osher Thermodot Marker. The markers have a needle tip instead of the proper flat tip. | Process control | On February 26, 2018, Beaver Visitec International Inc. emailed Product Recall Notices to affected customers. The letter identified the affected device and the reason for the recall. Customers were instructed to destroy affected product, provide evidence of the destruction, and complete a customer response form. Customers will receive a credit. Customers with questions regarding the recall may contact Customer Service at (866)906-8080. |  | 110 | 110 total products | Distributed to the states of OH, NY, FL, MN, and KS. | Terminated | January 28, 2019 | 336 | 110
Z-1702-2018 | 79972 | No | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | BARDÂż BrachySourceÂż Iodine125 radioactive seeds consist of a welded titanium capsule containing the I125 absorbed onto a nickel / copper-coated, gold-cored aluminum wire (reference Figure 1 below). The I125 has a half-life of 59.6 days and decays by electron capture with the emission of characteristic photons and Auger electrons. The principal photon emissions are 27.4 and 31 keV x-rays and 35.5 keV gamma. The titanium wall of the BrachySourceÂż Seed Implants absorbs the electrons. |  |  | April 10, 2018 | 2018 | Bard Brachytherapy, Inc. | Bard Medical Division is recalling the BARDÂż BrachySourceÂż Iodine125 Radioactive Seeds in MickÂż Applicators and BARDÂż BrachySourceÂż Iodine125 Radioactive Seeds Loose in Vial with Sterile Spacers. Because the incorrect dosage level radioactive seeds were mixed up and sent to two hospitals. One Hospital ordered BrachySourceÂż Iodine125 0.36 mCi radioactive seeds but instead received BrachySourceÂż Iodine125 0.28 mCi radioactive seeds and vice versa for the other hospital affected. | Labeling mix-ups | The notification to the customer was completed by Bard Medical Customer Service via a telephone call to alert them of the incorrect product shipped on April 10, 2018.Bard Medical Division is recalling the BARDÂż BrachySourceÂż Iodine125 Radioactive Seeds in MickÂż Applicators and BARDÂż BrachySourceÂż Iodine125 Radioactive Seeds Loose in Vial with Sterile Spacers. The incorrect dosage level of radioactive seeds was sent to 2 hospitals. One hospital ordered 0.36 mCi radioactive seeds but instead received 0.28 mCi radioactive seeds intended for the other hospital. For further questions, please call (770) 385-0562. |  | 170.64 | 70 Seeds at 0.28mCi and 100 seeds at 0.36mCi | US Distribution to the states of : California and Georgia | Open, Classified |  |  | 170.64
Z-1538-2018 | 79685 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Modular Head Component, Biomet Hip System; Item No. 11-363661, Hip Joint, metal/polymer/metal, semi-constrained, porous-coated, uncemented prosthesis |  |  | April 02, 2018 | 2018 | Zimmer Biomet, Inc. | Two lots of different sized modular heads potentially commingled. Risks include delay in surgery and dislocation due to use of an incorrect sized head. | Nonconforming Material/Component | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 4/2/2018 to its customers. The letter described the product, problem and actions to be taken. Distributors will be notified of the recall via email. Hospital risk managers and distributors with product will be notified via courier. The distributor letter includes instructions to locate and remove the product in their territory, identify hospitals who have previously used the product; return on-hand product to the recalling firm and ensure all products are accounted for. The letter for hospital risk managers instruct the risk managers to assist the recalling firm's sales reps with quarantine of the product; to review notification and ensure affected personnel are aware of the contents; and complete and return the Certificate of Acknowledgement form to CorporateQuality.PostMarket@zimmerbiomet.com, even if you do not have product, and retain a copy.Customers with questions are directed to call customer service at 574-371-3071 between 8AM-5PM EST, Monday through Friday. Alternatively, questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 18 | 18 | Worldwide Distribution: US (nationwide) in states of: AL, AR, FL, GA, IN, KY, LA, MO, MS, OH, OK, TX, and VA; and countries of: Canada, Costa Rica and Chile. | Open, Classified |  |  | 48
Z-1540-2018 | 79593 | No | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension VistaÂż Calcium FlexÂż reagent cartridge, Dimension VistaÂż CA, K1023, SMN # 10445160 |  |  | January 30, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | May produce erroneously low results from specific well sets. | Under Investigation by firm | An Urgent Medical Device Recall and an Urgent Field Safety Notice were issued on January 30, 2018 to all Dimension Vista customers who purchased the impacted lot to notify them of the issue, the potential risk to health, and actions to be taken by the customer. Customers were instructed to: "Discontinue use of and discard the Dimension Vista Calcium Flex reagent lot 17171BD."Please review this letter with your Medical Director."Review your inventory of these products to determine your laboratorys replacement needs and to provide information to Siemens for reporting to the authorities."Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days.If you have received any complaints of illness or adverse events associated with the product listed in Table 1, immediately contact your local Siemens Customer Care Center or your local Siemens Technical Support Representative. For further questions, please call (312) 275-7795. |  | 6000 | 6000 | Worldwide Distribution - USA (nationwide) and to the following countries:Australia,Austria,Belgium,Canada,Czech Republic,Denmark,France,Germany,Italy,Japan,Norway,Portugal,Slovakia,Slovenia,South Korea,Spain,Switzerland | Open, Classified |  |  | 6000
Z-1541-2018 | 79421 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Specimen Gate Screening Center, Part Number 5002-0500, All software versions from 1.0 to 1.8 (current software version) |  |  | November 17, 2017 | 2017 | PerkinElmer Life and Analytical Sciences, Wallac, OY | Potential errors in patient results generated by the Screening Center product that include both false negative and false positive results. | Software design | PerkinElmer sent an Urgent Medical Device Correction letter dated November 14,2017.Instructions:1. Implement additional control measures concerning demographic information, by taking one of the following actionsa. Fully enter demographics before loading specimens on instruments (i.e. specimens can be punched, but the assay plates cannot be loaded to the measurement instruments).b. If the above mentioned is not possible or if previously entered demographic information changes after loading the assay plates on the measurement instruments, recalculate all your results in Specimen Gate laboratory Result Viewer prior to assay acceptance.2. Follow unreported specimens with available tools to avoid any delay in reporting screening results. For further questions, please call (317) 418-1735 |  | 8 | 8 | Worldwide Distribution - US Distribution to the state of Georgia., and to the countries of : Belgium, Canada, Denmark, Italy, and United Kingdom. | Open, Classified |  |  | 8
Z-1544-2018 | 79769 | general purpose reagent | PPM | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | BD MAX DNA MMK (SPC) For Laboratory Use, catalog no. 442829 |  |  | March 02, 2018 | 2018 | Becton Dickinson &amp; Co. | The package insert incorrectly lists the volume of the primer and probe diluent to use for the reconstitution of the master mix tube. Adding the incorrect amount of diluent could have an effect on the assay. | Error in labeling | On March 2, 2018, the firm sent Urgent Medical Device Product Advisory notices to customers via UPS. The notice identified the errors and the correct information, as well as potential health risks. Customers were asked to take the following actions: 1) Share the information with all users of the product within their facility to ensure awareness; 2) Complete the customer response form and return to the recalling firm; and 3) Report any associated adverse health consequences to the recalling firm.The recalling firm is not asking customer to send back product or discard. The recalling firm will provide users with an updated insert when available.BD Customer/Technical Support: 800-638-8663, Monday-Friday, 8:30AM-5PM (EST) for US customers. |  | 13898 | 13,898 | Distributed to accounts in 8 states: CA, FL, NJ, NY, OR, PA, VA, and WI. International distribution to Argentina, Australia, Brazil, Canada, Chile, China, Colombia, Europe, Hong Kong, Japan, Singapore, Taiwan, and Thailand. | Terminated | July 02, 2018 | 122 | 13898
Z-1542-2018 | 79549 | No | LFH | Tricyclic antidepressant drugs test system. | Toxicology | Toxicology | 510(k) | Alere iScreen Dx Urine Drug Screen Card |  |  | December 15, 2017 | 2017 | Ameditech Inc | The kit box label incorrectly identifies Phencyclidine (PCP), where it should list Tricyclic Antidepressants (TCA) as one of the 10 drug analytes included in the kit box. | Labeling Change Control | Customers were notified on about 12/15/2017. Instructions included to become aware of the recall so as to relay the information to customers with questions. |  | 828 | 828 kits (20,700 devices) | Distributed domestically to AZ, FL, IL, ND, NV, PA, TN, TX, VA. | Terminated | February 11, 2019 | 423 | 828
Z-1543-2018 | 79778 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS PanelsÂż CHOL+GLU Test Strips (REF 1765) |  |  | March 14, 2018 | 2018 | Polymer Technology Systems, Inc. | Potential for loss of activity of the glucose analyte (under-recovery) prior to the full stated shelf life (August 2018) | Under Investigation by firm | Letters will be sent to each ship-to address, as well as each bill-to address via UPS with signature requirement. Customers will be instructed to immediately quarantine and destroy any materials remaining from these two lots. |  |  |  | OH, PR, VA, CAL, IL, FL, WA, TN, KY, OH, MN, NY and Belgium, Hong Kong, Sweden, Australia, Russia, Sweden, Brazil, England | Terminated | September 13, 2018 | 183
Z-1704-2018 | 79818 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ATTUNE Revision CRS Femoral RT SZ 5 Cemented, Catalog No. 150440205, Femoral Revision Implant, Knee |  |  | April 16, 2018 | 2018 | DePuy Orthopaedics, Inc. | There may be burrs on the extraction hole threads. | Process control | Notification letters distributed on 4/17/18 instructed customers to perform the following:Steps to Take Please take the following actions:" Immediately return unused affected units: Note: On March 19, 2018, one unit was removed from consignment. Two units were implanted prior to the medical device recall. No additional units are in the market." Reconciliation Form: Complete the reconciliation form and return to your sales consultant or fax to 574-371-4939 or email to klong16@its.jnj.com within five (5) days of this notice." Records: Retain a copy of the completed reconciliation form in your files along with this notice." Additional Notifications:o Notify surgeons at your facility by providing them with a copy of this notice. Forward this notice to others in your facility that need to be informed.Clinical Implications and Patient ImpactDePuy (Ireland) is not recommending prophylactic revision in the absence of symptoms. The company recommends that surgeons discuss potential clinical implications and risks with symptomatic patients that received the affected units. |  | 1 | 1 unit | The products were distributed to the following US states: FL, MO, and PA. | Open, Classified |  |  | 3
Z-1546-2018 | 79702 | No | GCX | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Medi-Vac(TM) Flex Advantage(TM) Suction Canister Liner, Material Numbers: |  |  | March 28, 2018 | 2018 | Cardinal Health 200, LLC | products may be missing a one-way valve on the fluid patient port. | Process control | Cardinal Health sent an Urgent Product Recall letter dated March 28, 2018,Please take the following actions:1. INSPECT your inventory for the affected product code and lot number (listed on package)2. SEGREGATE and QUARANTINE all on-hand product3. RETURN the enclosed acknowledgment form via fax to 847-689-9101 or 614-652- 9648, whether or not you have affected product, indicating the product code, lot and quantity of product you've quarantined4. CONTACT Customer Service at 800-964-5227 to arrange for return and credit of any affected product . For further questions, please call (800) 292- 9332. |  | 3600 | 3600 units | US Distribution to the state of MA | Terminated | November 29, 2018 | 246 | 3600
Z-1448-2018 | 79428 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Perspective 16 |  |  | February 15, 2018 | 2018 | Siemens Medical Solutions USA, Inc | to inform customers of possible incorrect tube current calculations by the CARE Dose4D algorithm for head scans based on p.a. (posterior-anterior) or a.p. (anterior-posterior) topograms. Depending on the geometrical shape of the skull bone, it may happen in rare cases that the calculated dose distribution is not appropriate and could lead to unnecessary radiation exposure | Software design | The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 1301 units will consist of the following:1. You will contact customers and initiate the actions to correct the failure to comply issue.2. You will provide a software update to correct tube current calculations by the CARE Dose4D algorithm for head scans.3. You will follow-up with consignees to verify the effectiveness of the software update.CDRH approves the CAP subject to the following conditions:1.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Any future customer notification letters to be sent must comply with the requirements of 21 CFR 1003.21. This includes, but is not limited to, the prescribed packaging and the statement that you will remedy the defect or bring the product into compliance at no charge.You may proceed or continue with implementation of the CAP (subject to the conditions noted above). For further questions, please call (610) 219-4834. |  | 1301 | 1301 | USA (nationally) Distribution. | Open, Classified |  |  | 1301
Z-1587-2018 | 79683 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Atellica Solution and Atellica Sample Handler Prime for in vitro diagnostic testing of clinical specimens. |  |  | February 26, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | When more than one immunoassay analyzer is connected to a single solution system, reporting and conversion factors may not be applied to all analyzers | Software design | Recall initiated 02/23/2018Firm released an update software (SW) version (V), 1.12.1, which corrects the behavior. A Mandatory Update (MU) campaign was launched to install the new SW V on all systems, whether or not they are impacted by this behavior. The MU campaign will be tracked to 100% completion globally. |  |  |  | FL andFranceSpainSwedenUnited KingdomGermany | Open, Classified
Z-1584-2018 | 79825 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | 5 mm Apple-Hunt Secondary Cannula/Pyramidal Tip Trocar - extended length (5 count); Model Number 900-840 |  |  | March 12, 2018 | 2018 | CooperSurgical, Inc. | Products have been packaged with the wrong Trocar length type. | Under Investigation by firm | On March 26, 2018 CopperSurgical Inc. mailed a Medical Device Recall Notification to affected customers. Customers were instructed to discontinue the use of the affected lot number 223333 and complete an Acknowledgement and Receipt Form to schedule refund or exchanges. If customers supplied products to any other facility, they have been asked to provide the recipients a copy of the notification letter. Customers with questions may contact CopperSurgical at recall@coopersurgical.com. |  | 520 | 520 products | Worldwide Distribution - US Nationwide in the states of FL, GA, IN, NJ, and countries of Canada, Austria, Chile, Italy, Singapore | Open, Classified |  |  | 520
Z-1583-2018 | 79939 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DeRoyal(R) SURGI MATE(R) Electrosurgical Pencil, Button Switch, with 15 ft (4.6 m) Cord and Holster, REF 88-001502 |  |  | April 23, 2018 | 2018 | DeRoyal Industries Inc | Non-sterile product was packaged inside a case that was labeled as sterile. | Process control | DeRoyal sent an Urgent Recall Notice letter dated April 23, 2018 to affected customers. The letter identified the affected products, problems and actions to be taken. The recall notices requested destruction of the product at the consignee level. Distributors were asked to notify their consignees or to provide the list to DeRoyal for notification. For questions contact your DeRoyal Representative, customer service at 888-938-7828 or by email at recalls@deroyal.com |  | 1800 | 1800 pencils | Honduras, Belgium, and The Netherlands | Terminated | November 09, 2018 | 200 | 1800
Z-1589-2018 | 79770 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper R8.x.25.5 and UNIZ systems (only with a FlexVision large screen monitor), Interventional fluoroscopic x-ray system |  |  | March 14, 2018 | 2018 | Philips Electronics North America Corporation | After continuous operation for more than one and a half days, the image on the large screen monitor may freeze for approximately 15 seconds after which the system will restore itself. | Software design | Field Safety Notices were distributed on 3/14/18. The notices instructed customers:To prevent occurrence of the issue customers shall restart the system at least once a day. Customer shall ensure that all staff with access to the affected systems are informed of the contents of this Field Safety Notice.A copy of this Field Safety Notice shall be placed together with the documentation of the system until the system has been corrected by Philips.The problems will be resolved by a software update, which is expected to be available by the second half of 2018.You will be notified by your local Philips representative when the software update is available for installation.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722-9377, 8am to 8pm (EST) Monday thru Friday. |  | 682 | 682 | Worldwide Distribution in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, and WY. The product was distributed to the following countries: Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Croatia, Cuba, Czech Republic, Denmark, Estonia, France, Germany, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Republic of Korea, Malaysia, Mexico, Morocco, Myanmar, Netherlands, New Caledonia, Norway, Pakistan, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, and United Kingdom. | Open, Classified |  |  | 682
Z-1389-2018 | 79630 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Power-LOAD Cot Fastening System, Model #6390, Catalog #6390-000-000, UDI: 07613327261523 |  |  | March 12, 2018 | 2018 | Stryker Medical Division of Stryker Corporation | The power load floor plates which hold the ambulance cot anchors in place during patient transport, may be broken. | Nonconforming Material/Component | On February 27, 2018, affected customers received an Urgent Medical Device Correction Notification via FedEx notifying them of the issue and potential hazard. For distributed units, the four Floor Plate Attachment Brackets that hold the anchor in place will be replaced in the field by Strykers field service personnel. Customers will be asked to return a business reply form to acknowledge that they have received and understand the notification. For further questions, please call (269) 389-8306. |  | 4724 | 4724 | Worldwide Distribution - USA (nationwide) and to the countries of : Australia, Canada, France, Germany, Hong Kong, Italy, India, Netherlands, Portugal, Saint Martin, Switzerland, United Kingdom | Open, Classified |  |  | 4724
Z-1585-2018 | 79439 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System |  |  | December 08, 2017 | 2017 | Hitachi Ltd., Medical System Operations Group | Due to a system controller software anomaly, the patient positioning deviation correction may not be applied and may result in the wrong part of the patient being irradiated. | Device Design | Hitachi sent a Field Safety Notice letter dated November 28, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. The notification letter includes instructions for the user on how to verify patient position even with the software issue. The letter also explains that correction software is being developed with a tentative release date of, "End of January, 2018.". A customer service representative will contact affected customers about installing the software. For questions contact your our Customer Service Representative. |  | 1 | 1 | US Distribution in the state of NY. | Open, Classified |  |  | 1
Z-1477-2018 | 79605 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, ProteusPLUS or Proteus ONE produces and delivers a proton beam for the treatment of patients with localized tumors and other conditions susceptible to treatment by radiation. |  |  | January 26, 2018 | 2018 | Ion Beam Applications S.A. | When switching from Pencil Beam Scanning treatment mode to Uniform Scanning treatment mode, the tuning setpoints of the scanning magnets are not always taken into account when the proton beam is requested. The mismatch between the setpoint and feedback leads to an error message. Depending on user action, two scenarios are possible:1. The user resumes the treatment field without analyzing the cause for the error message. This may lead to under-irradiation.2. The user reboots the Scanning Magnets Power Supply Electronic Unit. This may lead to delay in patient treatment and the patient may need to be re-aligned, which leads to an additional X-ray dose. | Under Investigation by firm | On January 26, 2018, the firm sent Field Safety Notices to all its consignees by email or delivered personally, alerting them of the issue. The notice recommended that users not resume pauses in Uniform Scanning more than twice without informing on-site firm personnel. The firm also recommended rebooting the Scanning Magnets Power Supply Electronic Unit if the issue occurs. The recalling firm stated it would implement an upgrade of the Scanning Magnet Power Supply Electronic Unit at all affected sites by June 30, 2018.Consignees were instructed to sign and return a copy of a notice read confirmation within 10 working days.Recalling firm Helpdesk Contact: +32 2 507 20 81 (available 24/7) |  | 5 | 5 | Consignees in four states: VA, IL, NJ, and WA. One consignee located in Germany. | Completed |  |  | 5
t | o | o |  | b | i | g
Z-2185-2018 | 80018 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Accu-Chek Aviva Plus Test Strip, Blood glucose monitoring system, Catalog nos. 06908217001, 06908268001, 06908349001,06908217001 |  |  | May 07, 2018 | 2018 | Roche Diabetes Care, Inc. | The recalling firm identified four US lots of test strips out of specification. Affected strips show increased potential for strip errors prior to applying blood, strips not detected once inserted, or biased results, which may not be detectable. In the case of a biased result, incorrect action could be taken by the user in response to the displayed blood glucose result. | Equipment maintenance | "On May 7, 2018, the firm sent Urgent Medical Device Correction letters to affected Retailers, Suppliers, and Patients. Patients were instructed to check their test strip supply lot numbers against the recalled lots. Any test strips from the affected lots must be discarded. To get replacement test strips, patients should use one of the following options:o For fastest replacement of your product, please submit your request athttps://notices.accu-chek.como Email Accu-Chek Customer Care at accu-chek.care@roche.com and includethe following information: your name, full mailing address, phone number,test strip lot number and number of affected test strip vials.o For questions regarding this recall, contact Accu-Chek Customer Careline at 1-800-358-4866 Monday through Friday between the hours of 8:00 amand 8:00 pm Eastern Standard Time.Suppliers were instructed to discontinue distribution and dispose of any remaining stock of the affected lots immediately. Suppliers were also asked to forward recall notifications to their customers, including both supplier- and consumer-level customers. Suppliers were asked to complete a faxback form indicating the total number of vials disposed of by the supplier and supplier customers. If you are a supplier and have questions about the information in this notification, please contact your Roche Account Manager.Retailers were instructed to discontinue sale or distribution and dispose of remaining affected lots in stock. Retailers were also asked to provide recall notification to their consumer customers. Retailers may note the quantity disposed and contact their Suppliers to arrange for credit. For questions please contact your wholesaler or the Accu-Chek Customer Care line at 1-800-358-4866 Monday through Friday between the hours of 8:00 am and 8:00 pm Eastern Standard Time." |  | 295248 | 295, 248 | Worldwide Distribution - US Nationwide distribution. Foreign distribution to ArgentinaArmeniaAustraliaAustriaAzerbaijanBahamasBahrainBarbadosBelgiumBrazilBulgariaCanadaChileChinaColombiaCosta RicaDenmarkDominican RepublicEl SalvadorFinlandGeorgiaGermanyGhanaGreeceGuatemalaHondurasHong Kong (SAR China)HungaryIndiaIndonesiaIsraelItalyJamaicaJapanJordanIrelandKenya Korea Kuwait Lebanon Lithuania Luxembourg Malaysia Malta Mexico Myanmar Netherlands New Zealand Norway Oman Pakistan Philippines Poland Russian Federation New Zealand Saudi Arabia Serbia Singapore Slovenia South Africa Spain Sweden Switzerland Taiwan (Republic of China), Thailand, United Arab Emirates, United Kingdom, Vietnam, Portugal" | Open, Classified |  |  | 295248
Z-2267-2018 | 80126 | No | NXT |  |  | Orthopedic | PMA | Birmingham Hip Resurfacing Femoral Head: |  |  | May 11, 2018 | 2018 | Smith &amp; Nephew, Inc. | The affected devices are manufactured to the same design specification as products intended for sale and use in the US but do not contain the FDA approved labeling (i.e., IFU and Patient Information Brochure). The affected devices were distributed in the US with the IFU intended for use outside of the US | Process control | The firm initiated the recall by email and letter on 05/11/2018. Consignees are asked to inspect inventory, locate and quarantine affected, unused devices. Affected product will be returned to Smith &amp; Nephew. |  | 2 | 2 units | OH, TX, CO, TN, MI, PA, NJ | Open, Classified |  |  | 8
Z-2107-2018 | 80095 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | MicroAire SmartRelease Endoscopic System, Endoscopic Carpal Tunnel Release, SmartRelease Handpiece, containing IFU rev. D |  |  | April 19, 2018 | 2018 | MicroAire Surgical Instruments, LLC | The Instructions for Use (IFU) is incorrect with regards to sterilization parameters and must be replaced. | Other | Urgent Medical Device Recall letters were distributed to customers. The letters instructed customers to do the following:If you have some quantity of IFUs remaining, we're asking you to destroy your current copies, IM-SMARTRELEASE Rev. A-D, and replace them by going on-line at MicroAire at www.microaire.com/resources. Click Carpal Tunnel, then SmartRelease Instruction Manual to obtain the latest version, IM-SMARTRELEASE Rev. E. Please complete and send back page two of this letter to notification@microaire.com or via fax at 800-648-4309.If you have no quantity of IFUs remaining, there are no further activities needed by you except for the mandatory returning of page two of this letter to notification@microaire.com or via fax at 800-648-4309. For further questions, please call (800) 722-0822. |  | 2009 | 2009 | Worldwide Distribution - USA (nationwide) including Puerto Rico and to the countries of: Australia, Belgium, Brazil, Canada, Chile, Czech Republic, Denmark, France, Germany, Hong Kong, India, Iran, Ireland, Israel, Italy, Jordan, Korea, Lebanon, Mexico, Nicaragua, Poland, Russia, Saudi Arabia, Serbia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, UAE, and UK. | Open, Classified |  |  | 2009
Z-2055-2018 | 79709 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD VACUTAINER PLUS SST Tube with HEMOGARD closure 13x75mm 3.5ml PLBL Gold, Catalog No. 367983 |  |  | February 13, 2018 | 2018 | Becton Dickinson &amp; Company | Tubes may exhibit stopper creepout/pullout, where the stopper dissociates from the tube | Nonconforming Material/Component | On March 15, 2018 BD Life Sciences sent Urgent Medical Device Recall Notifications to affected customers. Customers were instructed to : 1. Immediately review your inventory for the subject lots and quarantine the product subject to the recall.2. Share this recall notification with all users of the product to ensure they are also aware of this recall.3. Complete the attached Customer Response Form and return to the BD contact noted on the form whether or not you have any of the impacted material so that BD may acknowledge your receipt of this notification and process your product replacement.Return all affected products using the enclosed packaging instructions and include the completed Recall Response Form with the product.4. Report any adverse health consequences experienced with the use of this product to BD. Events may also be reported to the FDA's MedWatch Adverse Event Reporting program. Web: Medwatch website at www.fda.gov/medwatch Phone: 1-800-FDA-1088 (1-800-332-1088); Mail: MedWatch, HF-2, FDA 5600 Fisher's Lane, Rockville, MD 20853-9787. Customer who require further assistance should contact BD Preanalytical Systems at (877)870-4486, Mon-Fri 8:00am-5:00pm |  | 9568900 | 9,568,900 total products | US Nationwide, Argentina, Belize , Brazil, Canada, China, Guam, India, Japan, Mexico, Nicaragua, Pakistan, Philippines, Singapore, Suriname, Taiwan | Open, Classified |  |  | 9568900
Z-2177-2018 | 80185 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Modular Primary Tibial Tray Implants; Modular Interlok Primary Tray (size 75 mm); Item No. 141215; UDI: (01) 00880304005310 (17) 280306 (10) 385340 |  |  | May 18, 2018 | 2018 | Zimmer Biomet, Inc. | One lot of 75mm tibial tray is potentially etched and labeled as a 79 mm tibial tray. This may lead to delay in surgery, or immediate revision surgery to remove an incorrectly sized tibial component after it has been cemented. | Under Investigation by firm | On May 18, 2018, the firm notified affected distributors and user facilities of the recall through Urgent Medical Device Recall letters sent via email and FedEx. The letters informed affected consignees of the issue and stated that the affected units were distributed in March 2018. Distributor Responsibilities:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.c. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.d. Mark RECALL on the outside of the returned cartons.4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com.a. Review the list of hospitals and/or surgeons included with the email notification sent to your facility, which includes a list of hospitals and/or surgeons that have already been notified of this recall.b. Identify whether there are any additional hospitals and/or surgeons that Zimmer Biomet has not notified and list these accounts on the Additional Accounts form. Please provide the form in Excel format.c. If there are no additional accounts or surgeons to notify, please indicate that there are no additional accounts, or indicate None or NA on the form.Risk Manager Responsibilities:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected |  | 16 | 16 | US Distribution to the states of : FL, GA, IN, MA, NM, UT, and WI. Foreign distribution to Canada and Mexico | Open, Classified |  |  | 16
Z-2050-2018 | 78466 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | nordicBrainEx 2.0 provides analysis and visualization capabilities of dynamic MRI data of the brain, presenting the derived properties and parameters in a clinically useful context. |  |  | September 17, 2014 | 2014 | NordicNeuroLab AS | An error was discovered in the interpretation of certain DICOM header tags that may lead to incorrect orientation labeling, and thus and indirect left-right, up-down or anterior-posterior flipping of images. | Software design | The firm notified their consignees by email on 09/17/2014. |  | 22 | 22 licenses | USA, Sweden, Spain, Russia, Hungary, India, Norway, Iceland, Italy, South Korea, United Kingdom | Open, Classified |  |  | 22
Z-2049-2018 | 80066 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | HARMONIC ACE Shears + Adaptive Tissue Technology, 5mm Diameter / 36cm Length / Curved, Product Code HAR36, GTIN 10705036014447 |  |  | May 03, 2018 | 2018 | Ethicon Endo-Surgery Inc | The devices may have been assembled with an internal component that may cause continuous or inadvertentactivation of the device. | Other | Urgent Voluntary Medical Device Recall (Removal) letters were distributed on 5/3/18. The letters included the following instructions for customers:1. Examine your inventory immediately to determine if you have product lots subject to this recall on hand and quarantine such product(s).2. Remove the product lots subject to this recall and communicate the issue to relevant operating room or materials management personnel, or anyone else in your facility who needs to be informed. If any product lots subject to this recall have been forwarded to another facility, contact that facility to arrange return.3. Complete the Business Reply Form (BRF) (Attachment 2) confirming receipt of this notice and fax or email it to Stericycle at 1-866-913-0665 or ethicon5872@stericycle.com within three (3) business days. Please return the BRF even if you do not have the product lots subject to this recall.4. Keep this notice visibly posted for awareness until all product lots subject to this recall have been returned to Stericycle. While processing your returns, please maintain a copy of this notice with the product lots subject to this recall and keep a copy for your records.5. Customers are required to return all unused HARMONIC ACEÂż+ Shears with Adaptive Tissue Technology product lots subject to this recall that are in their inventory immediately. Only unused HARMONIC ACEÂż+ Shears with Adaptive Tissue Technology product lots subject to this recall returned by August 31, 2018 will be eligible for replacement. Any unused HARMONIC ACEÂż+Shears with Adaptive Tissue Technology product lots subject to this recall returned after August 31, 2018 will not be eligible for replacement.6. To return unused HARMONIC ACEÂż+ Shears with Adaptive Tissue Technology product lots subject to this recall, photocopy the completed BRF, place it in the box with the subject products, and affix the pre-paid authorized shipping label included with this recall notification letter. Ethicon wi |  | 146326 | 146,326 | Worldwide distribution - US Nationwide.The devices were distributed to the following foreign countries: Algeria, Angola, Argentina, Armenia, Australia, Austria, Belgium, Bolivia, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Columbia, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, French Guiana, French Polynesia, Georgia, Germany, UK, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Martinique, Mayotte, Mexico, Mongolia, Montenegro, Morocco, Netherlands, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, Ukraine, UAE, Uruguay, Vietnam. | Open, Classified |  |  | 171983
Z-2051-2018 | 80016 | No | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoris MCK Onlay Insert Extractor. |  |  | January 19, 2018 | 2018 | Mako Surgical Corporation | Possible hinge pin disassociation from the Mako Onlay Insert Extractor. | Under Investigation by firm | Customers were initially notified on approximately 01/19/2018, and again notified when the recall was expanded on 04/25/2018. Instructions included to inform all applicable individuals within the organization that need to be notified, complete and return the Recall Notification Business Reply Form, and make arrangements for return all affected products on hand. On about 12/12/2018, Stryker expanded the scope of the recall to include two additional lot codes and notified customers via "URGENT MEDICAL DEVICE REMOVAL" letter. Instructions were the same as the original recall notification. |  | 575 | 575 units total (400 US, 175 OUS) | Distributed domestically to AL, AR, AZ, CA, CO, CT, FL, IL, IN, KS, LA, MD, MI, MN, MO, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, TN, TX, UT, WI. Distributed internationally to Australia, China, Italy, Netherlands, Singapore, South Africa, Taiwan, Turkey, United Kingdom. | Open, Classified |  |  | 575
Z-2052-2018 | 79903 | No | FDX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Hobbs Bronchial Cytology Brushes, Model Number 4202; 4204; 4206 |  |  | May 01, 2018 | 2018 | Hobbs Medical, Inc. | Device was marketed for uses outside the 510(k) clearance. | Labeling False and Misleading | The firm, Hobbs Medical, Inc., sent an "Voluntary Medical Device Recall Advisory Notice" letter to customers on 5/1/18. The letter described the product, problem and actions to be taken.Our records indicate that you have received the following Bronchial Cytology Brushes since the beginning of 2018:Please complete the quantity used and quantity returned section of the attached form and you may voluntarily return any unused device(s) referencing Return Authorization Number. Hobbs Medical will credit your account in a timely fashion.f you have any questions, please contact customer service at 800-344-6227. |  | 1820 | 1820 devices | US Nationwide Distribution to states of: AZ, CA, GA, IL, MD, NY, TN, TX, and VA. | Open, Classified |  |  | 1820
Z-2057-2018 | 79978 | No | PPR | Arthroscope. | Orthopedic | Orthopedic | Enforcement Discretion | 3910-090-800 ASSY, PKG, 8.0MM X 90MM DRI-LOK THREADED CANNULA, 5 units per box. |  |  | December 21, 2017 | 2017 | Stryker Corporation | Inadvertent shipment of expired units. | Incorrect or no expiration date | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 12/27/2017 to its customers via mail on about 12/28/2017. Customers were initially notified via phone on 12/21/2017 and 12/22/2017. The letter described the product, problem and actions to be taken. The customers were instructed to inform those individuals in the organization who need to be aware of the recall; examine stock for affected devices; quarantine and discontinue use of recalled device; complete and return the Business Reply Form to Stryker Endoscopy via fax (408-855-6314) or email (endorecall@stryker.com), and if affected product is on hand segregate it and contact Stryker customer service at 1-800-624-4422 (Option 3) or email endocustomersupport@stryker.com to arrange for product return andissuance of credit or replacement (upon availability).If you have any questions, contact Director of Regulatory Affairs and Clinical Science at 408-754-2664 or email: to endorecall@stryker.com. |  | 2 | 2 boxes (5 units per box) | US distribution to Massachusetts only.. | Completed |  |  | 2
Z-1884-2018 | 78975 | No | RDZ |  |  |  |  | Santajoy Ornament &amp; Giftware Co., Ltd E-Z Illuminations Northern Lights E-Z Illuminations Galaxy |  |  | November 21, 2017 | 2017 | SANTA JOY ORNAMENT &amp; GIFTWARE CO | It was discovered that the product units it imported may, at peak levels, emit radiation that exceeds the federal standard for Class IIIa products. | Nonconforming Material/Component | The Firm will bring the device into compliance by : 1. Santa Joy Ornament (SJO) will provide a full refund in exchange for the return of the models listed above. Returned products will be destroyed.2. SJO has purchased optical measurement equipment and intends to measure the output of each laser product and ensure all laser products comply with applicable performance standards before shipment.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification. For further questions, please call (410) 540-9840. |  | 433128 | 433,128 | USA (nationwide) distribution | Open, Classified |  |  | 433128
Z-2242-2018 | 80193 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.2MM X 560MM COCR THD TIP; 14-441053 |  |  | May 16, 2018 | 2018 | Zimmer Biomet, Inc. | The firm is recalling various trauma guide wires due to insufficient packaging design verification and exceeding the expected occurrence rate of complaints for sterile barrier failure. Guide wires packaged in the affected packaging configuration have the potential for sterile barrier breach, leading to possible risk for infection. | Under Investigation by firm | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 16, 2018 on May 16, 2018, to affected consignees via email and FedEx. The letter describes the product, problem and actions to be taken. Consignees were informed of the recall and instructed to take the following actions:Risk Manager Responsibilities:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.Distributor Responsibilities:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.comb. International returns, please request an IRA by emailing zimmerbiometintlirarequests@zimmerbiomet.comc. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.d. Mark RECALL on the outside of the returned cartons.4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com.a. Review the list of hospitals included with the email notification sent to your facility, which includes a list of hospitals that have already been notified of this recall.b. Identify whether there are any additional hospitals that Zimmer Biomet has not notified and list these accounts on the Additional Accounts form. Please provide the form in Excel format |  | 2396 | 2396 | Worldwide Distribution: US (Nationwide) and countries of: Argentina, South Korea, Chile, Colombia, Trinidad and Tobago, El Salvador, Costa Rica, Canada, Australia, Japan, and Netherlands. | Open, Classified |  |  | 12863
Z-2053-2018 | 79496 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur BNP (B-type Natriuretic Peptide), Catalog Numbers: 02816138 (100 tests), 02816634 (500 tests) |  |  | December 18, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc. | In certain scenarios the ADVIA Centaur CP, XP and XPT immunoassay analyzers may exhibit reduced dilution recovery or become inoperable when used with Multi-Diluent 15 and certain lots of BNP and TSH3RUL Assays. | Device Design | Consignees are asked to return the response form and follow the instructions pertaining to their system and assays: Actions for ALL ADVIA Centaur Systems: Do not perform onboard dilutions on any ADVIA Centaur system if any of the affected kits are present on the system. Onboard dilutions for the ADVIA Centaur BNP and/or TSH3UL assay(s) are only supported when using ADVIA Centaur BNP kit lots ending in 210 and above and ADVIA Centaur TSH3UL kit lots ending in 310 and above. Actions for ADVIA Centaur XPT system: If you intend to perform onboard dilutions for the ADVIA Centaur BNP and/or TSH3UL assays, you must restart the workstation after scanning the master curve card for the BNP and/or TSH3UL assay kit lot, before ordering an onboard dilution for these assays. Please perform the attached steps to restart the workstation. Actions for ADVIA Centaur CP system: If any of the affected kit lots have been in use on the system, you must manually track the time that the Multi-Diluent 15 has been in use on board the system. Replace the ancillary pack of Multi-Diluent 15 every 7 days when used, until a follow-up communication is received from Siemens stating this issue has been resolved. |  | 21687 | 21,687 | Distributed in the following U.S. states: AL, CA, DE, GA, IA, IL, IN, KY, MA, MI, MS, MT, NC, ND, NJ, NM, NV, NY, OH, PA, RI, TN, TX, VA, WA, WV, WY and Puerto Rico. Distributed in the following OUS countries: AustraliaBelgiumÂżÂżÂżÂżÂżÂżÂżÂżCanadaChileFranceÂżÂżÂżÂżÂżÂżÂżÂżÂżGermanyÂżÂżÂżÂżÂżÂżÂżÂżGreeceÂżÂżÂżÂżÂżÂżÂżÂżÂżIsraelItalyÂżÂżÂżÂżÂżÂżÂżÂżÂżÂżJapanLatviaÂżÂżÂżÂżÂżÂżÂżÂżÂżLuxembourgÂżÂżÂżÂżÂżNetherlandsÂżÂżÂżÂżNew ZealandNorwayÂżÂżÂżÂżÂżÂżÂżÂżÂżPolandÂżÂżÂżÂżÂżÂżÂżÂżÂżPortugalÂżÂżÂżÂżÂżÂżÂżRomaniaÂżÂżÂżÂżÂżÂżÂżÂżSingaporeSlovakiaÂżÂżÂżÂżÂżÂżÂżSouth AfricaSpainÂżÂżÂżÂżÂżÂżÂżÂżÂżÂżSwitzerlandÂżÂżÂżÂżU.A.E.ÂżÂżÂżÂżÂżÂżÂżÂżÂżUnitedÂżKingdomÂż | Open, Classified |  |  | 95902
Z-2045-2018 | 80162 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | nordicBrainEX |  |  | December 21, 2012 | 2012 | NordicNeuroLab AS | There is an error related to the relative geometry between fiber tracts in a fiber group and images.The error occurs in the following situations: A group of fibers has been selected using the VOI functionality and put into a fiber group. This is done with images series A visualized in the 3D viewer. Then a new images series B is visualized in the 3D viewer. If image series A and B have the same geometrical resolution (pixel sizes and slice distances), the coregistration matrix taking A to B will not be applied to the fiber group, and therefore the fiber group will in general not be positioned correctly on B in the 3D viewer. Furthermore, if the fiber group is exported as a new image series the fibers could be misplaced, and similarly if the fibers are visualized in the MPR on B, the positioning will not be correct. The misalignment will be equal to the rotations/translations necessary to align A with B. | Software design | The firm distributed the recall notification letter by email. The letter identified the affected product, problem and the actions to be taken. For questions contact NordicNeurolab Customer Service or if you would like to arrange a return customerservice@nordicneurolab.com. |  | 36 | 36 licenses | Worldwide Distribution - US Nationwide - Saudi Arabia, Brazil, Sweden, Spain, Russia, Belgium, Norway, South Korea, Japan | Open, Classified |  |  | 36
Z-2183-2018 | 80172 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Knee System PS Tibial Bearing, Part Number 183621 |  |  | May 11, 2018 | 2018 | Zimmer Biomet, Inc. | The label master file was errantly set up as a 63/37mm instead of a 63/67mm. There is no 63/37 size offered, and the product is laser marked with the correct size. | Under Investigation by firm | The firm, Zimmer Biomet, sent an "Urgent Medical Device Recall (Removal)" letters dated 5/11/2018 to customers on 5/11/18. The letters described the product, problem and actions to be taken. The letters instruct customers to perform the following actions:1. Review this notification and ensure affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form.a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days.4. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet.a. For each return, send a copy of Attachment 1 to corporatequality.postmarket@zimmerbiomet.com.b. Include a hardcopy of Attachment 1 with your shipment for immediate processing.c. Mark the outside of the returning product box(es) clearly with RECALL. 5. Distributors will notify their accounts to which the product was further distributed and to locate and return products from these accounts in accordance with local regulations.6. Retain a copy of your field action acknowledgement and product return forms for your records in the event of a compliance audit of your facility.7. If after reviewing this notice you have further questions or concerns please call the customer call center at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voice mail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 169 | 169 | Worldwide distribution: US (nationwide) to states of: AR, CA, FL, GA, IL, IN, KS, MI, MN, MS, NC, OH, PA, TN, TX, and WV and country of: Canada. | Open, Classified |  |  | 169
Z-2266-2018 | 80173 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CARESCAPE R860 ventilators with software version 10SP05 |  |  | December 08, 2017 | 2017 | Datex-Ohmeda, Inc. | Potential for the display processor to experience anunexpected failure. | Software change control | On December 12, 2017 GE Healthcare sent letters to their consignees stating that There is a potential for the display processor to experience an unexpected failure. If this failure occurs, the device screen will temporarily go blank and will then display an Error Code EC01 XX . The blank display results in loss of monitored data being displayed to the user. There have been no injuries reported as a result of this issue. You may continue to use the system. Device performance is not affected on devices where no such display failure occurs.In the event this issue occurs, an audible alarm will sound. Mechanical ventilation will continue at previous settings. A device restart is required to clear the alarm condition. Consider using appropriate alternative means of ventilation such as a self-inflating bag or another ventilator to prevent patient injury while restarting the device.GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 349 | 349 (US) + 2,051 (OUS) | Worldwide Distribution - USA (nationwide) Distribution to the states of : AL, AZ, CA, GA, ID, IL, IN, MS, NC, NJ, NK, NV, OH, OK, RI, TX, and WI., and to the countries of : Albania, Argentina, Australia, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Ecuador, El Salvador, Estonia, Ethiopia, Egypt, Finland, France, Germany, Ghana, Hungary, India, Indonesia, Italy, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Malaysia, Mexico, Morocco, Myanmar, Nepal, New Zealand, Nigeria, Norway, Pakistan, Palestine, Peru, Philippines, Poland, Russia, Saudi Arabia, Senegal, Singapore, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Vietnam | Open, Classified |  |  | 349
Z-1614-2018 | 80020 | [6 empty] | 11-VI Incorporated Laser diode component |  |  | December 13, 2017 | 2017 | Ii - Vi Incorporated | Affected diodes incorporated in the laser sights emitted greater than the 5 mW limit. | Device Design | Please note that II-VI Inc. is responsible for the correction of all noncompliant products should the reported epoxy processing defect adversely affect any other laser diodes manufactured by II-VI, Inc.The single customer, Crimson Trace Corporation CTC was formally notified by you in writing. The defective laser diodes, models 25-0655 &amp; 25-0657 will be or have already been replaced at no cost to CTC. The replacement will be done in batches and is expected to take approximately six to nine months, or no later than January 31, 2019. Defective laser diodes will be either returned to II-VI or destroyed at CTC. For further questions, please call (724) 352-4455. |  | 323972 | 323,972 | US Distribution to Oregon. | Open, Classified |  |  | 323972
Z-2062-2018 | 79932 | No | CDQ | Urea nitrogen test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Chemistry Urea Nitrogen Reagent (material number 10309051) for use on ADVIA 1800, 2400 and XPT Chemistry Systems |  |  | March 27, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | Urea Nitrogen reagent kit demonstrated an increased incidence of Calibration Failures. Results cannot be generated by the system when calibration fails. | Material/Component Contamination | Date recall was initiated: 03-27-18Customer communication was sent on 04-03-18Urgent Medical Device Recall (CHC18-04.A.US) was distributed to all customers in the United States and an Urgent Field Safety Notice (CHC18-04.A.OUS) was distributed to all customers outside the United States who have received ADVIA Chemistry Urea Nitrogen reagent kit lots 408143 and 418290 (Lot 418290 was not sold in the US). These letters instruct the customers to discontinue use of reagent kit lots 408143 and 418290 and to contact their Customer Care Center for replacements kits.In the US the method of distribution is Fed Ex and outside the US the distribution is determined at the country level. |  |  |  | Nationwide and worldwide | Open, Classified
Z-2061-2018 | 78467 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | nordicICE v 2.3.14 |  |  | July 03, 2015 | 2015 | NordicNeuroLab AS | The bug causes BOLD activation maps to be visualized as overlays without taking the coregistration into account. Any output created from these activation maps as overlays will not be adjusted according to the coregistration. | Software design | The firm notified their customers by email on July 3, 2015.Actions to be taken by the customer/User:We advise users to immediately stop using the BOLD module in nordiclCE and contact NordicNeuroLab for further instructions on how to upgrade the nordiclCE BOLD module, as described more fully below.For further questions, please call (262) 337-2909. |  | 121 | 121 licenses | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Australia , Belgium, Brazil, China, Czech Republic, Denmark, France, Germany, Greece, Hungary, India, Italy, Japan, Kuwait, Malaysia, Netherlands, Norway and Poland. | Open, Classified |  |  | 121
Z-2289-2018 | 80186 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Complete Knee System; E1 Vanguard PS Tibial Bearing; Model Numbers EP-183664, EP-183642 |  |  | May 24, 2018 | 2018 | Zimmer Biomet, Inc. | Two lots of tibial bearings were commingled. There is a possibility that a 14mm tibial bearing is packaged in a box labeled as 12 mm tibial bearing and vice versa. This may lead to extension of surgery to find another implant. | Under Investigation by firm | On May 24, 2018, Zimmer Biomet issued an Urgent Medical Device Recall letter to its affected consignees via email and FedEx. The letter informed customers of the product commingling issue. Customers were instructed to take the following actions:Distributors:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.c. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.d. Mark RECALL on the outside of the returned cartons.4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com.a. Review the list of hospitals included with the email notification sent to your facility, which includes a list of hospitals that have already been notified of this recall.b. Identify whether there are any additional hospitals that Zimmer Biomet has not notified and list these accounts on the Additional Accounts form. Please provide the form in Excel format.c. If there are no additional accounts to notify, please indicate that there are no additional accounts, or indicate None or NA on the form.Risk Managers:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.Po |  | 48 | 48 | Distributed to accounts in CA, CT, GA, IA, IN, LA, MA, MD, MN, MS, NJ, OH, PA, TN, TX, UT, VA, WI. Foreign distribution to Taiwan, Japan, and the Netherlands. | Open, Classified |  |  | 48
Z-2059-2018 | 79968 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | RX Imola (RX4900) |  |  | May 04, 2018 | 2018 | Randox Laboratories | Abnormal premature termination of the software could affect the use of the analyser when running patient samples. The impact being a delay in reporting test results. | Under Investigation by firm | All customers will be contacted directly by email and will be followed up with once every 2 weeks over a period of 6 weeks. Customers are advised to update the software in all analyzers. For further questions, please call 1 (866)472-6369. |  | 41 | 41 analyzers | US Distribution including Puerto Rico and the state of : West Virginia. | Open, Classified |  |  | 41
Z-2100-2018 | 80012 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | VGI MEDICAL VerteLP LATERAL INTERBODY FUSION instrument kit |  |  | April 09, 2018 | 2018 | VGI Medical, LLC | IFU contained an automated cleaning process for the VerteLP Instrument Tray that was not validated | Process design | The firm notified their direct consignees by letter on 04/09/2018 and followed with an e-mail on 04/11/2018. The letter directed the direct consignees (distributors) to instruct their accounts (user) to discontinue the automated cleaning process and to provide users with the new IFUs. |  | 6 | 6 units | Distributed to MI, FL, and NY. | Open, Classified |  |  | 6
Z-2256-2018 | 80180 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | C2 CryoBalloon Focal Catheter, REF FG-1009. C2 CryoBalloon Ablation System is used as a cryosurgical tool in the field of general surgery, specifically for endoscopic applications. |  |  | March 30, 2018 | 2018 | C2 Therapeutics, Inc. | The catheter alignment arrow printed onto the strain relief is not aligned with the detachment feature of the catheter connector, resulting in the potential for the catheter to be attached to the controller in the detach orientation. In this orientation, the catheter is not securely attached to the controller and may be partially (~1cm) ejected from controller. | Process control | On 3/30/2018, C2 Therapeutics issued URGENT FIELD CORRECTION notices to customers via certified mail, return receipt requested.Customers are advised to take the following Action:- Forward a copy of the notice and the enclosures to the department in which is referenced in the notice. - Inspect stock and quarantine products identified within the notice. - The end user of the affected products should complete and return the response form via email to RMA@c2therapeutics.com. The response form contains an RMA number to allow for the return of devices affected by this field action, and enclosed in the package is a Federal Express shipping label for the returned merchandise.Customers with questions can call 650-318-5899, Monday through Friday 8:30 AM - 5 PM, PST. |  | 57 | 57 units | Distribution US nationwide. | Open, Classified |  |  | 57
Z-2063-2018 | 80114 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TL-HEX TruLok Hexapod System, Double Row Footplates, TL-HEX Sterile, All Sizes: 120 MM, 140 MM, 160 MM, 180 MM, 200 MM. |  |  | January 30, 2018 | 2018 | Orthofix Srl | The firm received complaints regarding the incompatibility and dimensional interference between the double row footplates manufactured in a limited period of 2014 and struts manufactured after September 2016. The firm determined that certain batches of the double row footplates are not compatible with the current version of the TL-HEX struts. | Device Design | Customers were notified on approximately January 30, 2018, via letter. Instructions included to examine inventory for affected devices, remove any affected devices and arrange for the return of them, notify customers if the devices have been further distributed, and complete and return the response form. The notification also provided alternative procedures to complete the surgical intervention and treatment: 1) the surgeon can use parallel external support (either footplate or ring) that does not have dimensional issues; 2) the surgeon can complete fixation temporarily using non-hexapodal TrueLok connection elements, but the solution requires post-operative non-surgical intervention such as additional parallel external support that does not have dimensional issues or the surgeon completes the frame for treatment using TruLok non-hexapodal elements; 3) in case the issue arises in a strut changed during treatment, the surgeon can use an Emergency Tab Kit (see attached diagram in letter). For further questions, please call (214) 937-2000. |  | 408 | 408 units | Worldwide Distribution - US Distribution within Texas., and to the countries of : Australia, Belgium, Chile, Denmark, France, Germany, Israel, Italy, Japan, South Africa, Spain, Sweden, Switzerland and United Kingdom. | Open, Classified |  |  | 408
Z-2068-2018 | 79813 | No | JFH | Acid phosphatase (total or prostatic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ST-AIA PACK PAP; Part Number: 025262 |  |  | March 05, 2018 | 2018 | Tosoh Bioscience Inc | Asfotase Alfa (Strensiq) interferes with certain Tosoh assays which potentially causes falsely increased or falsely decreased immunoassay test results, depending upon the assay. | Device Design | The firm TOSOH Bioscience Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 23, 2018 to affected customers. The letter described the product, problem and action to be taken. The customers were instructed to do the following: *Based on our revised labeling, the assays identified in Table 1 should not be used for patients who are being treated with Asfotase Alfa." For patients that are being treated with Asfotase Alfa, use an alternative test method that does not utilize alkaline phosphatase technology." Inform all medical professionals that assays utilizing alkaline phosphatase-based technologies must not be used for patients receiving Asfotase Alfa treatment. Continue to use the assays listed in Table 1 for any patients who are not receiving Asfotase Alfa treatment." Complete and return the attached Acknowledgement Form to Tosoh Bioscience, Inc., within 15-days of receiving this notification via Fax: 1-650-636-8651, Email: biorecallresponse@tosoh.com or Regular mail: 3600 Gantz Rd., Grove City, OH 43123." Maintain this notification with your laboratory records and forward this information to others who may have received this product." If you have obtained unexpected test results or received any complaints of illness or adverse events associated with the use of the products identified in the table above, contact Tosoh Technical Support 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to call (800) 248-6764 or email: bernadette.oconnell@tosoh.com. Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 1083 | 1083 | Worldwide Distribution: US (Nationwide) and countries of; Canada, Mexico, Columbia, Chile, Uruguay, Nicaragua, Argentina, Peru, Chile, Honduras. Equador, Caymen Islands, Panama, and Costa Rica. | Open, Classified |  |  | 300971
Z-2064-2018 | 80150 | No | EZC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bardex Lubricath 3-Way 75cc Continuous Irrigation Foley Catheter (Creevy), 22Fr. Rx (01267522) |  |  | December 13, 2017 | 2017 | C.R. Bard, Inc. | Foley Catheter with Incorrect Retainer Cap Print | Process control | On December 13, 2017 a letter was sent to all their customers informing them of the following: Bard Medical Division (BMD), C.R Bard, Inc., would like to inform you of a potential labeling issue involving the BARD product code and lot numbers listed below. Manufacturing of these affected products began May 2017.Due to a manufacturing issue, there is a potential for catheters with the two lot numbers listed above to have the incorrect information printed on the purple retainer cap. The incorrect print specifies that the balloon should be inflated with 10mL of water. However, the correct volume to use to inflate the catheter balloon is 75cc . If the user underinflates the balloon with only 10mL of water, the expected effect of helping stop the bleeding as an adjunct to conventional hemostasis post-TURP may not happen as quickly as desired, likely leading to a marginal severity of harm. Additionally, the catheter may become dislodged from the patient.Instructions to Customer:"Examine your inventory and identify any product subject to this communication. During use, adhereto the correct inflation volume as stated in the packaging."Product return is not required. If a unit does not meet your needs or you have any questions,please contact our customer service at 1-800-526-4455."If you have further distributed or transferred this product, please identify the respective organizationsand notify them at once of this product communication. Your notification to these organizations maybe enhanced by including a copy of this medical device communication letter.We appreciate your cooperation and assistance in dealing with this matter and sincerely apologize for any inconvenience that may result from this action. If you have any questions please do not hesitate to contact your Bard representative. |  |  |  | FL, NY, MN and TXCanada, Japan and Taiwan | Open, Classified
Z-2105-2018 | 80086 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet Cardiopulmonary Bypass Custom Tubing kit BEQ-TOP 15801 Neonatal Special, Part Number 701055490 |  |  | March 11, 2016 | 2016 | Datascope Corporation | There is a potential problem with the cuvette where failure of the cuvette would cause an error message on the monitor display of H/S DISCONNECT AT CUVETTE. | Under Investigation by firm | Urgent Safety Alert notifications were sent to customers on 3/11/16. The notices included the following instructions for customers:Please examine your inventory immediately to determine if you have the lot/batch numbers listed. If so, please following the recommended actions outlined in the attached Terumo Cardiovascular Systems Urgent Safety Alert.URGENT: Please complete attached Customer Response Form on page 3, to acknowledge that you have received the Terumo Cardiovascular Systems Urgent Safety Alert. Please fax the completed Customer Response Form to 1-973-396-3607 or send via email to fieldactions@maquet.com. |  | 194 | 194 total kits | The products were distributed to the following US states: AL, OR, and TX. | Open, Classified |  |  | 194
Z-2103-2018 | 78848 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | fMRI Hardware System - AudioSystem, Version 1.0 |  |  | October 28, 2017 | 2017 | NordicNeuroLab AS | Report was of an interrupted MR scan after patient complained about acoustic distortion, and heat felt form the headphones. | Device Design | The firm notified the consignees by email on September 28, 2017. The units were to be returned. For further questions, please call 1 (855) 665-3674 option 1. |  | 64 | 64 units | Worldwide Distribution - USA (nationwide) and to the countries of : Russia, Israel. Norway, South Korea, China, France, South Africa, Japan, Poland, Chile, Switzerland. Australia, Belgium, Lithuania, United Kingdom and Bulgaria | Open, Classified |  |  | 64
Z-2101-2018 | 80159 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS-IW with Universal Viewer |  |  | January 25, 2018 | 2018 | GE Healthcare | Potential that one or more image series may be missing from an exam without a user warning displayed in the Viewer. | Software design | On January 28, 2019 GE healthcare notified their customers and gave the following instructions for the Centricity Universal Viewer with PACS-IW foundation 6.0.x; GTIN 00840682103800 and Centricity PACS-IW with Universal Viewer version 5.0.x with PACS-IW foundation.; Until a product modification is available to correct this issue, the following actions are recommended.1. Utilize the image count within the QC process to alert the user of a discrepancy in the number of transmitted images from the modality to the number of images available in the Universal Viewer.a. If a discrepancy is identified, attempt to retransmit the exam to PACS.b. If retransmission is unsuccessful, a GE Healthcare Service representative will need to be contacted to assist with resolution of the impacted exam. Urgent cases impacted by this issue should be interpreted at the modality.2. If DICOM storage commit is configured and in use, no commit notification is sent to the modality for images that were impacted by this issue.3. Attention should be given to prior cases, acquired before the application of the safety instructions given in this communication as they may be affected by the safety issue.For historical data inquiries, please contact GE service for assistance in identifying any affected images.GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction. For questions please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 44 | 44 | Worldwide Distribution - US Nationwide in the states of AK, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, LA, MA, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OK, OR. PA. SD, TN, TX, UT, WA, WI, WV, and WY and the countries of Austria, Belgium, Brazil, Canada, Colombia, France, Germany, Ghana, India, Ireland, Israel, Italy, Kuwait, Malaysia, Mexico, Netherlands, Panama, Poland, Russia, Saudi Arabia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, United Kingdomand Uruguay | Open, Classified |  |  | 296
Z-2046-2018 | 79963 | No | MQM | Bilirubin (total and unbound) in the neonate test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Draeger Jaundice Meter JM-103 |  |  | May 15, 2018 | 2018 | Draeger Medical Systems, Inc. | Users have misinterpreted the display for out of range measurement indicated by the blinking" ---" to mean a zero measurement. | Device Design | The firm, Drager, sent an "Urgent Medical Device Safety Notices" were distributed to customers on 5/15/18. The letter described the product, problem and actions to be taken. The letters instruct customers to perform the following:Enclosed with this letter are label/s and instructions on how and hwere to apply the label onto your device/s. Please make sure that all JM-105/JM-103 devices in your hospital are labeled according to the instructions in a timely manner.Please make sure that in your clinical environment the users are trained in the use of the equipment and the clinical workflow is set up in a way that the device is not used as a diagnostic device. In case of any uncertainty (e.g. risk factors cannot be evaluated), a blood test needs to be performed. The risk factors according to the Guideline of American Academy of Pediatrics are summarized in the attached Sample Usage Protocol Template as well as the recommended method to use the device as intended in a clinical environment.Please completed and return the attached Customer Reply card.If you have any questions regarding, please call 1-800-437-2437(press 1 at the prompt than 32349) between the hours of 8:00am-4:30pm EST. |  | 10363 | 10363 | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY; and countries of: Algeria, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahamas, Bahrain, Bangladesh, Barbados, Belgium, Benin, Bolivia, Bosnia-Herz., Botswana, Brazil, British Virgin Islands, Brunei Darussalam, Bulgaria, Burkina-Faso, Cambodia, Canada, Chile, Columbia, Congo, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, French Guiana, Macedonia, Georgia, Germany, Ghana, Greece, Greenland, Guadeloupe, Guam, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jordan, Kazakhstan, Kosovo, Kuwait, Laos, Latvia, Lebanon, Libya, Lithuania, Malaysia, Mayotte, Mexico, Mongolia, Morocco, Mozambique, Myanmar, Namibia, Nicaragua, Nigeria, Norway, Nepal, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, St Vincent, St Pier Miquel, Sweden, Switzerland, Tadzhikistan, Tanzania, Thailand, Trinidad Tobago, Tunisia, Turkey, Turkmenistan, United Kingdom, Uruguay, UAE, Venezuela, Vietnam, and Zimbabwe. | Open, Classified |  |  | 17074
Z-2110-2018 | 80101 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NaturaLyteÂż Liquid Acid Concentrate for Bicarbonate Dialysis, 3.43 Liters, Part Number: 08-3231-1 |  |  | May 02, 2018 | 2018 | Fresenius Medical Care Renal Therapies Group, LLC | The bottles are mislabeled with an incorrect part number. | Employee error | Phone calls and Urgent Medical Device Recall notification letters were issued on 5/2/18. Customers were instructed to do the following:Immediately examine your stock to determine whether you have any NaturaLyte Liquid Acid from the lot listed above.If any products of this lot is found, discontinue use immediately.Place all units in a secure, segregated area.Contact your FMCRTG, LLC Customer Service Team at 1-800-323-5188 for instructions on how to return the recalled product and arrange for replacement product.Promptly fill out and return the attached reply form. |  | 1060 | 1060 cases (4240 bottles) | Worldwide Distribution - US Nationwide in the states: CA, IA, IL, IN, KS, MD, MI, MN, MS, NE, NJ, NV, OK, PA, SD, TX, and WI. | Open, Classified |  |  | 1060
Z-2114-2018 | 80092 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | BivonaÂż Tracheostomy Tube Tracheostomy Tubes |  |  | April 02, 2018 | 2018 | Smiths Medical ASD Inc. | Carton labeling is printed with "Sterile" indicated labeling, however the device is not sold as sterile. | Error in labeling | Customers were notified via letter on approximately 04/02/2018. Instructions included to determine if there are any affected Customized BivonaÂż Tracheostomy Tubes in inventory, notify customers if the devices have been further distributed, arrange for the return of affected devices, and complete and return the acknowledgement response form. |  | 16 | 16 units | Domestic distribution to CA, FL, GA, ID, NJ, NM, OH, VA. International distribution to France. | Terminated | November 14, 2018 | 226 | 16
Z-2108-2018 | 80061 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | ROSA(TM) Robotized Stereotactic Assistant, Version 2.5 |  |  | May 08, 2018 | 2018 | MEDTECH SAS | Replacement of units lacking an updated device approval. | No Marketing Application | Medtech would like to inform you of the planned removal of all ROSA 2.5 devices from the US market by 31st October 2018.The ROSA 2.5 design has been superseded by a newer generation product, ROSA Brain 3.0, and a decision has been made by Zimmer Biomet to exchange all ROSA 2.5 devices in the US market with ROSA Brain 3.0 devices. This decision has been taken in response to a regulatory clearance compliance issue, and there is no safety concern associated with this action.No specific action is required and the ROSA 2.5 can continue to be used at this time.A Medtech representative will be in contact to schedule the exchange of the RO 15 061 device with a ROSA Brain 3.0 device. For further questions, please call (574) 371-3071. |  | 29 | 29 units | US Distribution | Open, Classified |  |  | 29
Z-2121-2018 | 79819 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Non-Locking Screw, D 2.7MM X 8MM, Sterile, Rx Only, ORTHOFIX S.r.r. Italy on label, Size (Part Numbers): a) D 2.7MM X 8MM (002-D-27008); |  |  | April 06, 2018 | 2018 | Orthofix, Inc | Firm is is conducting a voluntary recall of the Rival Reduce Fracture Plating System and Rival View Plating System. As the result of a complaint investigation, the locking screws were confirmed to not properly engage the plates during surgery. | Device Design | On 04/06/2018, Orthofix Inc. (U.S. Agent), sent written notification of the recall to distributors and hospitals via e-mail and/or FedEx overnight courier. |  | 935 | 935 | U.S Consignees: NJ, FL, PA, SC, LA, VA, MN, TX, GA, DC, NV, IN, CA, MI, IL, MD, NC, OH, OK, | Open, Classified |  |  | 7213
Z-2109-2018 | 80219 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 12C (AdaPTinsight) |  |  | October 28, 2016 | 2016 | Ion Beam Applications S.A. | IBA is initiating this recall to address an issue identified with AdaPTinsight software and to reduce risk related to this problem. | Software change control | IBA mailed an Urgent Field Safety Notice to affected customers on 10/28/2016 to inform them of the issue. The issue will be solved in AdaPTinsights versions starting from Al 2.0. The solution was to be deployed on impacted IBA Proton Therapy sites by September 2017. |  | 5 | 5 units (12C) | Worldwide Distribution - US Nationwide in the states of LA, TX, Italy, Sweden and France. Each country's National Competent Authorities were notified of the firm's Field Safety Notice. | Open, Classified |  |  | 5
Z-2112-2018 | 80014 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter SIGMA Spectrum Infusion Pump with Master Drug Library (Version 6), infusion pump, 35700BAX and 35700ABB |  |  | May 08, 2018 | 2018 | Baxter Healthcare Corporation | The firm will be updating the Instructions for Use for the product. Current IFUs instruct users to lower the primary bag 20 inches below the secondary bag using a hanger when performing a secondary infusion. However, the firm supplies 12.5-inch hangers with the Baxter disposable secondary sets. | Other | On May 8, 2018, Baxter issued a Medical Device Correction letter to its customers. Customers were advised of the issue and instructed to do the following:1. The SIGMA SPECTRUM Infusion Pumps can continue to be safely used withthe Baxter-provided 12.5-inch hanger.2. The updated Instructions for Use will be made available by May 31, 2018. Atthat time, you may access the updated IFU online at the Baxter GlobalTechnical E-Service Center at URL: https://service.baxter.com/tsportal/#/login.For additional questions and concerns about this issue please contact BaxterHealthcare Medina at 800-356-3454, Option 1, between the hours of 7:00 amand 7:00 pm Eastern Time.3. If you purchased this product directly from Baxter, complete theenclosed Baxter customer reply form and return it to Baxter by faxing it to224-270-5457 or scanning and e-mailing it to fca@baxter.com, even if you donot have any inventory. Returning the Baxter customer reply form promptlywill prevent you from receiving repeat notices.4. If you purchased this product form a distributor, please note that the Baxter customer reply form is not applicable. If a reply form is provided by your distributor or wholesaler, please return it to the supplier according to their instructions.5. If you distribute this product to other facilities or departments within your institution, please forward a copy of this communication to them.6. If you are a dealer, wholesaler, distributor/reseller, or original equipment manufacturer (OEM) that distributed any affected product to other facilities, please notify your customers of this Medical Device Correction in accordance with your customary procedures. |  | 333961 | 333961 | Worldwide Distribution - US Nationwide and the country of Canada. | Open, Classified |  |  | 456850
Z-2115-2018 | 80057 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Xper Flex Cardio Physiomonitoring System |  |  | November 20, 2017 | 2017 | Invivo Corporation | While in Full Disclosure playback, a user may inadvertently close the Full Disclosure Control Window using the ESC key function, instead of pressing the X icon in the upper right corner of this window. | Software design | The firm, Philips Healthcare, sent an "URGENT - Medical Device Correction" letter dated November 2017 and Field Safety Notice to each known affected customer. The letter describes the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. The customers were instructed to do the following:Follow the Action to be Taken by Customer/User section of the Field Safety Notice as follows:-You may continue to use the device provided that each monitored patient in the Pre or Post-Op Holding Area is closely observed by a qualified health care professional and is not left unattended, as specified in the devices instructions for use. The most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment. Also, users should always immediately respond to all audible alarms given by the device.Until your device can be corrected, please inform users of the content of this notice and instruct them that to properly exit Full Disclosure playback, use the X icon in the upper right corner of the Full Disclosure Playback control window and do not use the ESC key function.A Philips representative will contact you regarding any affected PCC modules. Each affected PCC will require a software update by Philips. Additionally, all Software and Documentation Media Kits will be replaced. This corrective action will be implemented free of charge by Philips.If you need any further information or support concerning this problem, please contact your local Philips representative: 800-669-1328 option 2, then option 3. |  | 1040 | 1,040 units (USA) and 7 units (Foreign) | Worldwide Distribution: US (nationwide) to states of: AK, AL, AR, AZ, CA, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, NC, NH, NJ, NM, NY, NV, OH, PA, RI, SC, TN, TX, VA, WA, and WI ; and countries of: Canada, China, Great Britain and Korea. | Open, Classified |  |  | 1040
Z-2111-2018 | 80093 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Alaris Extension Sets with microbore tubing Model 20022 and 30914 |  |  | November 06, 2017 | 2017 | CareFusion 303, Inc. | Reports of incorrect priming volume due to manufacturing with incorrect tubing. | Under Investigation by firm | The firm, BD, sent an "URGENT: Medical Device Recall Notification" letter dated November 2017 to all their customers. The letter described the product, problem and actions to be taken. The customers were instructed to: immediately Discontinue use of the Alaris Extension Sets model 20022 lot numbers 16087775 and 16127995, and model number 30914 lot number 16127518. Contact BD for a replacement lot.; return the product for replacement using the following instructions; and immediately complete and return to BD the enclosed, pre-addressed and postage paid, Recall Response Card to: BD Support Center,10020 Pacific Mesa Blvd, San Diego CA 92121, by FAX: 1-858-617-4851 or by EMAIL: SupportCenter@bd.com. Note on the card your distributors name and quantities that will be returned." Once BD receives the Recall Response Card, the BD Support Center will provide instructions for return of the affected lot numbers and replacement of the affected lots. " " All recalled product should be returned directly to the distributor from whom it was purchased. Customers will receive their return goods processed through their distributor if the products were purchased through a distributor. "" Your distributor has already been notified of this recall. If you have any questions, please contact your distributor directly, or call the BD Support Center at the number listed below."A second letter was sent on May 2018 adding an additional lot # for Alaris Extension Set with microbore tubing Model number 30914 lot number 17097167Recall Related Questions BD Support CenterPhone: 1-888-562-6018Hours: 7am to 4pm PST, Monday -Friday Email: SupportCenter@carefusion.com |  | 25400 | 25,400 units | US Distribution to states of: CA, IL, IN, KS, NE, NJ, TX, UT, VA and WA | Open, Classified |  |  | 25400
Z-2140-2018 | 80096 | No | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | lmmunalysis EDDP Urine Control Set 1 |  |  | November 28, 2017 | 2017 | Alere San Diego Inc. DBA Immunalysis Corporation | The High (125 ng/ml) Control does not meet specification - it has been confirmed to produce values below 100 ng/mL when run on validated clinical chemistry analyzers, which is below the 100 ng/mL cutoff of the EDDP Specific Urine Enzyme Immunoassay with which this control is utilized. | Under Investigation by firm | On November 28, 2017, the firm sent their customer notifications of the recall stating the following:CUSTOMER REQUIRED ACTION1. If you have product in your possession:" Immediately discontinue use and/or sale of affected product lot and place in quarantine." Complete the enclose Customer Verification Form, within 10 days, indicating the quantity of product you have discarded or will return to lmmunalysis." FAX completed Customer Verification Form to 909-482-0850 or scan and email to rtimbang@immunalysis.com" Upon receipt of your Customer Verification Form, we will initiate shipment of replacement product in the quantity indicated on the form and, if applicable, will contact you to arrange product return." Return product as directed.2. If you have no remaining product in your possession:" Complete the enclosed Customer Verification Form, within 10 days, indicating "We do not have any affected product - it has been consumed"." FAX completed Customer Verification Form to 909-482-0850 or scan and email to rtimbang@immunalysis.com3. Share this information with your staff and retain this notification as part of your records.Please complete this form and return within 10 days even if you do not have anyremaining product. Send form to lmmunalysis at Fax 909-482-0850 or by e-mail to rtimbang@immunalysis.comFor further questions, please call (207) 730-5738. |  | 37 | 37 kits | US Distribution to the states of : CA, CO, CT, FL, LA, MD, NC, NJ, OH, OK, TN, TX, WI, and WV | Open, Classified |  |  | 37
Z-2137-2018 | 80068 | No | JTZ | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | remel SUPERPACK CS/ 500 SEG/FLASK, REF R01626, Qty:50, plate label: REMEL MH W/OXACILLIN |  |  | May 04, 2018 | 2018 | Thermo Fisher | Product may fail performance testing for S aureus ATCC 43300. | Under Investigation by firm | The firm initiated the recall by letter on 05/04/2018. The letter stated the following: "Please notify any personnel who need to be aware of the potential for failure of performance for this product. Accordingly, in keeping with our Quality Policy, we request that you inspect your stock and destroy any remaining inventory of the lots listed above and contact Customer Services or your local distributor regarding replacement. You should review results and laboratory reporting associated with use of the listed lots and consider retesting and/or seek appropriate expert advice for further action."For questions contact Technical Services Department at 800-255-6730 or 913-888-0939. |  | 67 | 67 units | US Nationwide Distribution | Terminated | December 19, 2018 | 229 | 67
Z-2139-2018 | 80221 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus ONE and Proteus Plus |  |  | September 16, 2016 | 2016 | Ion Beam Applications S.A. | IBA is initiating this recall to notify its customers that the use of non-IBA accessories and/or third party equipment or unauthorized modifications to an IBA medical device may pose a safety risk. | Finished device change control | IBA mailed an Urgent Field Safety Notice to affected customers on 09/16/2016 to notify them of the risks associated with the use of this device to prevent serious injury/death. Customers were asked not to make any unauthorized modification to the IBA System or IBA accessories unless approved in writing or in advance by IBA. IBA will be updating their manuals to include these warnings. |  |  |  | Units were delivered in the USA, Korea, Germany, France, Czech Republic, Italy, Poland and Sweden. | Terminated | February 08, 2019 | 875
Z-2248-2018 | 80200 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Biograph mCT Systems: |  |  | May 29, 2018 | 2018 | Siemens Medical Solutions USA, Inc. | Siemens Medical Solutions USA, Inc., Molecular Imaging business line became aware, through our supplier that some laser cover windows may fall off a limited population of the Biograph mCT and Biograph mCT Flow systems. | Device Design | The firm notified the consignees of the issue by email on May 29, 2018. The letter identified the affected product, problem and actions to be taken. They are tracking completion of the letter delivery through the hand-delivered approach. For further questions, please call (847) 304-7700.. |  | 405 | 405 units | Worldwide Distribution | Open, Classified |  |  | 405
Z-2058-2018 | 80115 | No | REM |  |  |  |  | Cell Observer SD, DirectFRAP, Laser TIRF 1, Laser TIRF 2, Laser TIRF3, LightSheet, Elyra, and multiple |  |  | May 08, 2018 | 2018 | Zeiss, Carl Inc | In certain eyepiece configurations the laser shutter may not close completely, permitting reflected beams that may be greater than the Class I limit. | Radiation Control for Health and Safety Act | 1.Your firm will send an initial email to all known users advising them of the issue, the risk involved and associated mitigations until the units can be evaluated by Carl Zeiss service representatives,2.Your firm will send a second written communication to each first line purchaser and to each known user site with the notice of action and warning to discontinue use of the component until it has been verified as defective or properly functioning, and3.Your firm will dispatch factory trained service representatives to perform verification tests and to replace defective units as necessary, free of charge.CDRH approves the CAP subject to the following conditions:For further questions, please call (914) 747-1800.. 1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Corrective actions will be provided at no cost to the purchasers and completed by March 1, 2019. |  | 1070 | 1070 units to US | USA (nationwide) | Open, Classified |  |  | 1070
Z-2158-2018 | 80201 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | MEDLINE ADULT MANUAL RESUSCITATOR WITH MEDIUM ADULT MASK, TUBE RESERVOIR, PEEP VALVE AND 7 FT. OXYGEN TUBING, REF CPRM1126P |  |  | May 24, 2018 | 2018 | Sun Med, LLC | There is a potential for the patient port retaining ring to not fully seat allowing the patient port to detach during use. This may impact air delivery to the patient. | Mixed-up of materials/components | On May 25, 2018 Sun Med, LLC issued URGENT MEDICAL DEVICE PRODUCT REMOVAL notices to two of their customers. On June 1, 2018 Sun Med, LLC issued a press release. Customers were encouraged to take the following actions:1) Examine inventory, quarantine and cease distribution of impacted product.2) If customers further distributed product, please identify your customers/consignees and notify them of this product removal. Your notification may be enhanced by including a copy of the removal notice. If you need any assistance contacting your customers, please inform SunMed.3) Complete and return the Removal Return Response Form via fax to 1-800-400-8820 or via email to info@sun-med.com as soon as possible. 4) Once you return the affected product(s), new corrected replacement resuscitation bags will be provided to you. If you need replacement products to be sent to you urgently, please call SunMed directly.Customers with questions can call 1-800-433-2797 or email info@sun-med.com |  | 30 | 30 cases | IL &amp; OH distributors | Open, Classified |  |  | 3084
Z-2136-2018 | 79805 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas e 801 immunoassay analyzer, cobas 8000 Modular Series |  |  | April 03, 2018 | 2018 | Roche Diagnostics Corporation | The syringe plunger on the device may be mounted in a tilted position, which applies mechanical stress to the plunger. In the worst case scenario, the plunger may crack leading to impaired syringe function. | Process design | On April 3, 2018, The Firm sent an Urgent Medical Device Correction letter via UPS Ground. Consignees were notified of the possible failure. Consignees were informed that a Roche representative would contact them to install a syringe plunger guide for the affected syringes to resolve the issue. In the meantime, consignees were instructed to only measure assays that do not require prewash steps and assays that are not impacted by missing prewash step until the syringe plunger guide is installed on the affected module. If the customer identifies broken syringe plungers on their device, the notice instructed them to contact the Roche Support Network Customer Support Center at 1-800-428-2336, available 24 hours a day, 7 days a week. Customers were asked to complete a customer response form and return it to the firm. |  | 35 | 35 | US Distribution to the states of : CA, CT, IA, IN, KY, ME, MI, NC, NY, OH, SC, TX, and VA. | Open, Classified |  |  | 35
Z-2161-2018 | 80076 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Aortic Arterial Cannula 24 Fr (O.D.); 23 cm (L); with 3/8-inch connector with LL, Part Number 701002267, Article Number A24-7107. For use in cardiopulmonary bypass |  |  | May 14, 2018 | 2018 | Maquet Cardiopulmonary Ag | This lot of arterial cannula may have been supplied with 3/8-width de-airing plugs that could break during the process of removal from the arterial cannula. | Component design/selection | Urgent Medical Device Recall letters were distributed beginning 5/14/18. The letters instructed customers to do the following:1. Read and understand this notification, and disseminate the information provided to all users, or prospective users, of the affected arterial cannulae.2. Check your arterial cannula inventory and compare with the Getinge part number, article number and lot number provided on Page 1 of this notification.3. Remove any affected product that are still in your inventory and prepare it for shipment back to Getinge.4. Call Getinge Customer Support at 1-888-627-8383 (Press Option 2 and then Option 2) for a Return Authorization (RGA) and for shipping information and label. Customer Support will issue a credit for the product return.5. Complete, sign and promptly return the filled out response form (Page 4 of this notification) either by email to FieldActions@getinge.com or by fax to 1-973-860-1675. |  | 150 | 150 US | Worldwide distribution including US state of VA and Puerto Rico, Australia, Belgium, Brazil, Canada, Chile, Colombia, Cuba, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Iran, Ireland, Israel, Italy, Japan, Lebanon, Malaysia, Martinique, Mauritius, Mexico, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Reunion, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Syria, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 660
Z-2142-2018 | 80064 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IntelliVue MX40 Patient Monitor, Software Versions B.05, B.06, B.06.5X, Model 865351 |  |  | April 30, 2018 | 2018 | Philips Electronics North America Corporation | Five warning statements are missing from the instructions for use. | Error in labeling | Urgent Medical Device Correction notification letters were distributed on 4/30/18. The notification included the following instructions:Customers were asked to please review the information provided in the letter with all members of their staff who need to be aware of its contents. The letter provides information on the missing warnings and instructions for actions to be taken. The enclosed MX40 IFU Errata Sheet must be attached to the first page of Chapter 6 of the Instructions for Use for ready reference. Complete and return the attached Customer Reply Form. |  | 7178 | 7,178 | Worldwide Distribution -- US, to the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, and to the territories of Guam and Puerto Rico; and, countries of Argentina, Australia, Austria, Belgium, Canada, Chile, China, Croatia, Czech Republic, Denmark, Estonia, Finland, France, German, Ghana, Greece, Guadeloupe, Hong Kong, Iceland, India, Ireland, Israel, Italy, Japan, Latvia, Lebanon, Malaysia, Martinique, Mexico, Myanmar, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UK, UAE, and Vietnam. | Open, Classified |  |  | 37259
Z-2163-2018 | 80125 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | smith&nephew; GENESIS(R) II TIBIAL Punch, 13 MM, REF 71440408, QTY: (1) |  |  | May 04, 2018 | 2018 | Smith &amp; Nephew, Inc. | One lot of GENESIS II 13 mm tibial punches are 2 mm smaller than specification. | Error in labeling | The firm initiated the recall by email on 05/04/2018. Consignees are asked to inspect inventory, locate and quarantine affected, unused devices. Affected product will be returned to Smith &amp; Nephew. All consignees are asked to complete the Response Form, even if they do not have product to return. |  | 5 | 5 units | International only: United Arab Emirates, Spain, Mexico | Terminated | February 22, 2019 | 294 | 5
Z-2162-2018 | 80112 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Prima Plus Âż4.1 x 11.5mm Implant, Catalog Number 15737K |  |  | May 01, 2018 | 2018 | Keystone Dental Inc | The implants are mislabeled. | Error in labeling | The Firm sent an Urgent Medical Device Correction letter dated April 30, 2018. In this discrepancy, the 5 is missing from the implant length description. Please note this discrepancy is limited to the implant package label only. The patient chart labels are all correct. Keystone Dental has performed a thorough review and determined there are no health hazards associated with this issue. There is no issue with the product within the package and the product can be safely used without issue in accordance with the defined surgical protocols. There is no risk of confusion of product size. The diameter is correct. Prima Plus implants are only available in 11.5mm lengths. The catalogs, surgical protocols and manuals only define procedures and provide instrumentation for placing 11.5mm length implants. Additionally, the patient chart labels contained within the implant package are correct so that patient procedure history will be properly maintained. The issue is limited to the label on the outside of the product package only. The lot identified above is the only affected lot. No other lots are affected with this issue. The product referenced above can be used safely without issue and no further action is required. For further questions, please call (866) 902-9272. |  | 72 | 72 | US Distribution to the states of : CA, CT, FL, IN, MA, MD, MI, MN, MT, NJ, NV, NY, TX, and VA. | Terminated | September 14, 2018 | 136 | 72
Z-0029-2018 | 78093 | Yes | NKE |  |  | Cardiovascular | PMA | Anthem Model Numbers: PM3210, PM3212; Allure/Relieve Model Numbers: PM3222, PM3242, PM3262, PM3542, PM3224, PM3244 |  |  | August 28, 2017 | 2017 | St Jude Medical Inc. | New pacemaker firmware was developed to further mitigate the risk of unauthorized access to our pacemakers that utilize radio frequency (RF) communications. | Unknown/Undetermined by firm | Abbott sent an Important Cybersecurity Advisory dated August 28, 2017, to all affected customers to notify customers of the availability of the programmer software update and associated pacemaker firmware update. The notification material is in the form of a Physician Letter and Hospital Letter and will be delivered to physicians and hospitals with inventory in the US by overnight service. Customers with questions were instructed to contact their Abbott representative or the customers technical support hotline at 1-800-722-3774. |  | 102684 | 102684 units | Worldwide Distribution | Open, Classified |  |  | 1318873
Z-2166-2018 | 80111 | No | JSP | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | The VITEKÂż 2 Anaerobic and Corynebacteria identification card (ANC) is intended for use with the VITEKÂż 2 Systems for the automated identification of most clinically significant anaerobic organisms and Corynebatcerium species. The VITEKÂż 2 ANC identification card is a single-use disposable. |  |  | March 23, 2018 | 2018 | BioMerieux SA | A shipment to Guadeloupe suffered a temperature excursions up to 47.8ÂżC, exceeding the acceptable limits. After detection of the issue, Global Supply Chain QA asked the bioMÂżrieux distributor in Guadeloupe to put products in quarantine in July 2017; After several reminders, the distributor informed in November 2017 that products were delivered to customers. | Environmental control | The firm, Biomerieux, sent an "Urgent product removal notice" letter on April 23 to affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1. Immediately acknowledge receipt (AR) of this FSCA.2. Identify all countries and customers for which you are responsible that are impacted by this FSCA.3. Define additional actions required (e.g. Translate customer letter, send to impacted countries/customers, etc.)4. Distribute the customer letter to all customers that have received the products 5. Discard any impacted product remaining in local inventory6. After all actions above are complete, please return the acknowledgement of completion (AC) for this FSCA. The due date for completion of the required actions and submitting the AC is 23-APR-2018SubsidiariesÂż Subsidiaries using SAP, please manage both the Acknowledgment of Receipt and the Acknowledgment of Completion of activities in SAP.Âż Subsidiaries not using SAP, please immediately upload the Acknowledgement of Receipt in LiveLink after signing and completing sections #1 and #2. Section #3 must be signed, completed and uploaded in LiveLink before the due date.DistributorsÂż Please immediately Acknowledge Receipt of this FSCA by completing sections #1 and #2 of the acknowledgement form and sending it by email to fieldactions@biomerieux.com or by fax to +33 4 78 87 21 79. Please remember to sign and complete all of sections #1 and #2.Âż Sections #1 and #3 must be signed, completed, faxed or emailed before the due date.If you have any questions, contact the Specialist, Regulatory &amp; Quality Compliance at +33 4 78 87 56 16 or email to: charlotte.thollet@ext.biomerieux.com. |  | 135 | 135 (78 boxes of various product) in total | International distribution. | Open, Classified |  |  | 1257
Z-2176-2018 | 80094 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 16 P with DEPMED HARDENING KIT 12NC: 453567400741, Model Number 728246. Computed Tomography X-ray system |  |  | March 30, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | The patient support head holder could contact the back ISO shelter wall during manual or motorized motion through the gantry bore. This may result in a serious injury that would require medical intervention to preclude permanent impairment in certain scenarios. | Device Design | Urgent Field Safety Notices dated 3/28/18 were distributed to customers. The letters instructed customers to perform the following actions:Please review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication.Please retain a copy with the equipment Instruction for Use.As outlined in the Instructions for Use, always monitor the patient during allmovements (manual or motorized) of the patient support. If you recognize an impending collision with the back shelter wall, halt the motion or activate the system E-stop.Mandatory Field Change Order (FCO) 72800686 which will add additionalwarning labels specific to this issue to the system. The FCO will be implemented free of charge. |  | 20 | 20 | Worldwide distribution including US states of CA, MD, NV, OH, TX, and WI, Afghanistan, Canada, Germany, Iraq, Italy, Korea, Kuwait, and Lebanon. | Open, Classified |  |  | 25
Z-2173-2018 | 80222 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235 |  |  | August 07, 2016 | 2016 | Ion Beam Applications S.A. | IBA is recalling to provide information through a Safety Notice to reduce risk for maintenance technicians during ion source replacement. | Employee error | IBA mailed out a Field Safety Notice on August 2, 2016, to affected customers where this maintenance is not performed by IBA technicians to inform them of the issue. For further questions please call (215) 972-7985. |  | 19 | 19 units | Units were distributed in the USA, Korea, Germany, France, Czech Republic, Italy, Poland, and Sweden. Each country's national Competent Authorities have been informed. | Open, Classified |  |  | 19
Z-2182-2018 | 80009 | No | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | Dimension Vista Gentamicin Flex reagent cartridge, DV GENT, gentamicin enzyme immunoassay. Affected lots are 17135BC, 17150BF, 17159BD, 17178AC, 17269BD, and |  |  | April 19, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | The affected lots may exhibit inaccuracy for Quality Control and patient samples at the low end of the Analytical Measurement Range. The bias is due to the calibration curve not showing separation between the Level 1 and Level 2 calibrators. The negative bias observed for patient, QC, and calibrator samples, when using any of the listed lots, has the potential to impact the interpretation of gentamicin results during therapeutic monitoring at concentrations up to approximately 2.5 Âżg/mL [5.4 Âżmol/L]. | Under Investigation by firm | On April 19, 2018, the firm sent an Urgent Medical Device Recall letter to customers. The letter identified the affected lots and advised customers to do the following: Inspect stock, discontinue use of and discard the Dimension and Dimension Vista GENT lotslisted in Table 1.Âż Please review this letter with your Medical Director.Âż Review your inventory of these products to determine your laboratorys replacement needsand to provide information to Siemens for reporting to the regulatory authorities.Âż Complete and return the Field Correction Effectiveness Check Form attached to this letterwithin 30 days.Âż If you have received any complaints of illness or adverse events associated with the productlisted in Table 1, immediately contact your local Siemens Customer Care Center or yourlocal Siemens Technical Support Representative. |  | 20373 | 20373 | Nationwide distribution. Foreign distribution to Argentina, Australia, Brazil, Bahamas, Canada, Germany, India, Japan, Myanmar, New Zealand, Saudi Arabia, Taiwan, and South Africa. | Open, Classified |  |  | 27916
Z-2184-2018 | 80137 | No | JOY | Automated cell-locating device. | Hematology | Hematology | 510(k) | CellaVision DM Software versions 6.0.1 or 6.0.2 installed on the following products: CellaVision DM96, DM1200, DM9600 and DI-60 |  |  | December 04, 2017 | 2017 | Cellavision AB | A software malfunction was found where WBC, RBC and PLT comments added after a slide is signed, are not sent to the LIS. This can only occur where customers process multiple slides per blood sample. | Software design | CellaVision sent an Urgent Medical Device Recall letter dated December 4, 2017. The letter identified affected product, problem and the actions to be taken. Customers are requested to determine if they are affected or not. If affected they are requested to contact their local distributor for installing a patch.Customers are asked to confirm receipt of the notice by signing and dating theattached Acknowledgement Form and emailing a copy to their local distributor.For technical questions contact CelleVision (service@cellavision.se) |  | 22176 | Instruments: 224 Software: 99 (US) | US Nationwide in the states of FL, IL, NY, and NC | Open, Classified |  |  | 22176
Z-2279-2018 | 80211 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Randox Liquid Assayed Specific Protein Control - Level 2 (SP Control 2), Cat. no. PS2683 |  |  | May 22, 2018 | 2018 | Randox Laboratories Ltd. | The recalling firm has confirmed that Free Kappa light chains in Liquid Assayed Specific Protein Controls increase over the shelf life of the product. This may lead to a delay in reporting Free Kappa Light Chains results, which are used as an aid in diagnosis of monoclonal gammopathy of undetermined significance and for Multiple Myeloma through review of the ratio with free lambda chains. | Under Investigation by firm | On May 22, 2018, Randox issued an Urgent Field Safety Notice to affected distributors along with an Important Notice and updated IFU. The letter advised customers of the product issue. It further instructed distributors to do the following:1. Discuss the contents of the notice with your Medical Director.2. Place a copy of the important notice and updated IFU into any remaining stock.3. Complete and return the vigilance response section of this form to technical.services@randox.com within five working days.If you have any questions after reviewing this information, you may contact the firm at technical.services@randox.com. |  |  |  | Distributed to accounts in CA, IN, KS, ME, and WV, and PR. Foreign distribution to Australia, Bangladesh, Belarus, Belgium, Brazil, Cambodia, Canada, Chile, China, Colombia, Cyprus, Denmark, Dutch Caribbean, Ecuador, Egypt, El Salvador, Ethiopia, Falkland Islands, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, India, Iraq, Ireland, Italy, Jordan, Kenya, Saudi Arabia, Korea, Kuwait, Latvia, Lebanon, Peru, Lithuania, Macedonia, Malaysia, Malta, Mauritius, Mexico, Morocco, Myanmar, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Qatar, Romania, Russia, Serbia, Slovakia, South Africa, Spain, Sri Lanka, Oman, Thailand, Netherlands, Turkey, United Arab Emirates, United Kingdom, Ukraine, Vietnam, and Zimbabwe. | Open, Classified
Z-2212-2018 | 80169 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | RS TANDEM 5F DL XCELA PICC PASV, Catalog Number 60M181582 |  |  | May 10, 2018 | 2018 | Angiodynamics Inc. (Navilyst Medical Inc.) | A component of the kits might contain unsafe levels of bacterial endotoxins (pyrogens). | Material/Component Contamination | Urgent Medical Device Recall letters were sent to customers on 5/10/18. The letters instructed customers to do the following: IMMEDIATELYStop using the product subject to recall.Remove any affected (recalled) product from your inventory (whether in labs, Central Supply, Shipping and Receiving or ANY other location).Segregate this product in a secure location for return to AngioDynamics, Inc.Forward a copy of this recall notification to all sites to which you have distributed affected product.Complete and return the Reply Verification Tracking Form.If affected product is located in your institution, please call AngioDynamics Customer Service at 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday  Friday: Eastern Standard Time) to obtain a replacement or credit for your returned product. |  | 13 | 13 boxes | The products were distributed to the following US states: AZ, CA, CO, CT, DC, FL, GA, IL, IN, KS, KY, LA, MD, MI, MN, MO, MS, MT, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, WA, WI, and WV. The products were distributed to the following foreign countries: Canada, China, Ireland, Spain. | Open, Classified |  |  | 1911
Z-2287-2018 | 80279 | No | LRW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Teleflex MEDICAL Pilling(R) UNIVERSAL SHEARS 7 1/2", REF 142542 |  |  | June 12, 2018 | 2018 | Teleflex Medical Europe Ltd | The incorrect plastic was used to manufacture the handles of these shears, which causes the handles to melt if sterilized. | Under Investigation by firm | The firm initiated the recall by letter on 06/12/2018. All affected consignees are instructed to take the following actions:"1. If you have affected stock, immediately discontinue use and quarantine any products with the product code and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document your receipt of this letter." |  | 573 | 573 units | Worldwide Distribution - US Nationwide and the countries of Canada, India | Open, Classified |  |  | 573
Z-2228-2018 | 80140 | No | GAX | Nonpneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Part #: SAM XT-C, Model: Hi-Viz Blue |  |  | May 02, 2018 | 2018 | The Seaberg Company Inc | Based on internal testing, results indicated a possible failure of the stitches securing the buckle to the nylon belt could occur, posing a potential risk when used on a human patient to stop arterial blood flow. There is a remote chance that a serious injury and/or death could occur as a result of the failure of the device. | Employee error | On 05/02/2018, the firm posted a press release notifying customers of a recall of all unused SAM XT Extremity Tourniquets (SAM XT) due to a possible failure of the stitches securing the buckle to the nylon belt that could pose a potential risk when used on a human patient to stop arterial blood flow. On 05/03/2018, recall notification forms were mailed to customers via overnight mail service. |  |  |  | USA: CA, MD, OH, FL, ID, CO, CT, OR, AZ, IL, ME, MO, NY, SC, MI, AR, MT, CT, WA, TX, KY, NC, NJ, PAOUS: Australia, Canada, China, Columbia, Costa Rica, Europe (EU), Hong Kong, Iceland, India, Iraq, Israel, Japan, Kenya, Korea, Lebanon, Malaysia, Russia, Saudi Arabia, Singapore, South Africa, Thailand, UAE | Open, Classified
Z-2060-2018 | 80067 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Sonialvision Safire II, Model #: DAR-8000f |  |  | January 12, 2018 | 2018 | Shimadzu Medical Systems | When selecting serial radiography with a pulse rate of 7.5fps (including selecting a preset or changing thepulse rate using a temporary edit function) it was observed the indicated "irradiation time" on the X-rayGenerator Console, as well as the "integral dose" on the external console of fluoroscopy, were reset to 0. There has been one report of this event. | Software design | Shimadf0u sent an Urgent Voluntary Medical Device Recall Notice letter dated January 2018. The letter identified the affected product, problem and actions to be taken. For questions contact Shimadzu representative or contact Shimadzu National Technical Support at 800-228-1429. |  | 17 | 17 | US Nationwide in the states of IL, WA, MI, TX, AZ, CA, LA, FL, SC, and NJ. | Open, Classified |  |  | 17
Z-2244-2018 | 80198 | No | HXD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Farabeuf Clamp for 3.5mm Screws PRO, Catalog # 705463 |  |  | May 04, 2018 | 2018 | Stryker GmbH | Laser etching which indicates whether the device is in compression or distraction mode appears to be on the wrong side of the lever arm | Other | On May 8, 2018, URGENT PRODUCT RECALL and BUSINESS REPLY notices were issued to U.S.A customers via UPS. Customers are encouraged to take the following actions:1. Inform individuals within your organization who need to be aware of this device recall.2. Immediately check all stock areas and/or operating room storage to determine if any devices from the affected product list are at your facility. Response is required by submitted the completed response form via email to: StrykerOrtho6487@stericycle.com, even though you may not have any physical inventory on site anymore.3. Quarantine and discontinue use of the recalled devices.4. Maintain awareness of this notice internally until all required actions have been completed within your facility5. Inform Stryker if any of the subject devices have been distributed to other organizations.a) Please provide contact details so that Stryker can inform the recipients appropriately.b) If you are a Distributor, note that you are responsible for notifying your affected customers.Customers with questions may call 201-831-5000. |  | 18 | 18 parts (10 parts quarantined) | AR, CA, ID, IL, MD, MI, NJ, NY, OK, SD &amp; TX | Open, Classified |  |  | 18
Z-2229-2018 | 80234 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | DermaPro Waterproof Silicone Tape, SNS57232 |  |  | December 14, 2017 | 2017 | Safe N Simple, LLC | The recalling firm received a customer complaint stating the tape was difficult to peel off the roll and had poor adhesion. | Process control | The recalling firm notified its distributors via phone call and email on January 3, 2018. The firm informed distributors of the issue with the affected product, and offered to replace affected product, and suggested replacing product distributed to customers. |  | 318 | 318 rolls | Distributors in two states: AL and OH. | Open, Classified |  |  | 318
Z-2243-2018 | 80085 | No | EXS | Urine flow or volume measuring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | CritiCore Automated Urine Output and Temperature Monitor |  |  | November 10, 2017 | 2017 | C.R. Bard, Inc. | Issues identified with the monitor including urine output measurement errors, temperature measurement errors and undesired alarms. | Other | On Nov 10, 2017 Bard Medical sent notification letter to all their consignees informing them of the following: Action required:"Immediately examine your inventory and quarantine product subject to this recall. Please refer to CritiCore MonitorAffected Serial Numbers (Attachment 2) to help you locate the affected product serial number. Do not use or furtherdistribute any affected product."Please complete and return the accompanying Recall &amp; Effectiveness Check Form (Attachment 1) attached to thisletter regardless of whether you have any of the affected product. Additional instructions for product return arecontained on the form."If you have further distributed any units with the referenced serial numbers, please identify your customers and notifythem at once of this product recall: Your notification should include a copy of this letter and the accompanying enclosures.Please complete this form by 11/24/2017 and fax to 1-770-784-6469 or email a scanned copy to BMD.FieldAction@crbard.com . If you have affected product, please contact the BMD Customer Service via phone at 1-800-526-4455 option 5 and again option 5 or email BMD.FieldAction@crbard.com to obtain a Return Goods Authorization number (RGA #). Fax this form to 1-770-784-6469 or email a scanned copy to BMD.FieldAction@crbard.com. Attn: Recall CoordinatorThis recall should be carried out to the user level. |  |  |  | AR, AZ, CA, FL, GA, HI, IN, JP, KS, LA, MD, MN, MO, MS, NC, NJ, OH, OK, TN, TX, VA, WA, WI, and WY | Open, Classified
Z-2246-2018 | 79754 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Part no. 690.379, Graphic Case F/6.5mm Cannulated Screw Set |  |  | March 05, 2018 | 2018 | Synthes (USA) Products LLC | Affected screw gauges incorrectly measure screws 5.41mm shorter than the actual overall screw length. | Process control | On March 8, 2018, the recalling firm sent Urgent Product Recall letters to its customers to inform them of the recall. Customers were advised to take the following actions: 1. Immediately review their inventory to identify and quarantine all affected products listed above in a manner that ensures the affected products will not be used.2. Complete the Verification Section and send a copy to the recalling firm. 3. Forward this notice to anyone in the facility that needs to be informed.4. If any of the affected products have been forwarded to another facility, contact that facility to arrange return.Customers with questions were advised to call 610-719-5450, 8AM-5PM (EST), Monday through Friday, or contact their DePuy Synthes sales consultant. |  | 4 | 4 | Distributed to accounts in CA, IL, IN, KS, LA, MD, NC, NE, NJ, OR, TX, and VA. Foreign distribution to Canada. | Open, Classified |  |  | 7
Z-2249-2018 | 80135 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Drainer(R) Centesis Catheters: |  |  | May 04, 2018 | 2018 | Vascular Solutions, Inc. | The supplier of Drainer(R) centesis catheters notified Vascular Solutions, Inc. that there is a potential for certain production lots to contain unsafe levels of bacterial endotoxins. | Other | Galt Medical Corp. sent an Urgent Medical Device Recall dated May 2, 2018 to affected customers. The letter identified the affected product, problem and actions to be taken. The firm requested the following actions be taken:Those with affected stock, were directed to immediately discontinue use and quarantine any product. If the affected product was further distributed, the distributor was directed to immediately contact all parties who may have received the affected stock and recover any unused product. For questions contact your local sales representative or Recall Coordinator at 214-778-1306 or email quality@galtneedletech.com |  | 276 | 276 units | US Nationwide Distribution | Open, Classified |  |  | 276
Z-2255-2018 | 80062 | No | MUZ |  |  | Neurology | PMA | VNS Therapy Programmer, Model 3000, v1.0 System |  |  | November 13, 2017 | 2017 | Cyberonics, Inc | Certain Model 3000 programming events can result in miscalculation of parameters stored in the Models 103, 104, 105, and 106 generators. During these programming events, the miscalculations can lead to: " Delivery of more stimulation than intended, resulting in painful stimulation or other common side effects (Model 106 only); " No stimulation in the case of device disablement (Burst Watchdog Timeout), resulting in no therapy to the patient (Model 106 only);" Delivery of less stimulation than intended, resulting in therapeutic settings not being achieved within device specification (Models 103, 104, 105, or 106); and/or " Delays or absence of the 75% and 50% battery life indicators displayed by the programming software (Models 103, 104, 105, or 106). | Device Design | On November 13, 2017 LivaNva sent letters to all the doctor and surgeons that use this device. "Dear Doctor:This notification describes an issue identified with the VNS Therapy Âż Programmer, Model 3000 version 1.0 which may impact patients implanted with Models 103, 104, 105 or 106 VNS Therapy generators. This issue does not affect Model 1000 SenTiva or Model 102/102R generators.What is the Issue?Certain programming events performed with a Model 3000 Programmer on a Model 103, 104, 105 or 106 VNS Therapy generator can result in a miscalculation of some parameters stored in the generator, which could lead to:" Output current delivered being different than that programmed;" Device disablement; and/or" Errors in the 75% and 50% battery life indicators.The actual battery life and battery function are not affected by this issue.Risk to HealthThis issue presents the following risks:" Delivery of more stimulation than intended, resulting in painful stimulation or other common side effects 1 (Model 106 only);" No stimulation in the case of device disablement, resulting in no therapy to the patient (Model 106 only);" Delivery of less stimulation than intended, resulting in therapeutic settings not being achieved (Models 103, 104, 105 or 106) ; or" Sudden decrease from 100% to 25% battery life indicator, though the actual battery life and battery function are not affected (Models 103, 104, 105).Which Patients are Potentially Impacted?Any patient with a Model 103, 104, 105 or 106 generator that has been programmed using a Model 3000 programmer may be affected by this issue. Your LivaNova representative will collect data from your Model 3000 programmer to identify which patients are impacted by this notice.Actions to be taken by the Physician1. Stop using the Model 3000 programmer to program Model 103, 104, 105 or 106 generators. You shall continue using the Model 3000 programmer with Model 1000 SenTiva generators and may continue to use it wi |  | 252 | 252 | AL, CA, CO, DE, FL, GA, ID, IL, IN, ME, MO, MS, NC, NJ, NY, PA, TN, TN, TX, UT, WA and WI | Open, Classified |  |  | 252
Z-2251-2018 | 80209 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Fabius GS Tiro Anesthesia Machine; Cat. no. 8606000 |  |  | May 24, 2018 | 2018 | Draeger Medical, Inc. | The Draeger anesthesia device may be able to dose 100% N2O. In the event of a fault, the S-ORC module would not prevent setting an N2O flow that would resultin a hypoxic mixture from being dosed to the patient. Potential adverse outcomes include death of the patient. | Process change control | On May 24, 2018, the firm issued Urgent Medical Device Recall letters to affected end users (medical facilities) via certified mail. The letter identified the affected product, problem and actions to be taken. The recalling firm stated that they are currently manufacturing replacement parts for all potentially affected devices, which will be installed as part of the next service. The firm's Service dept. will contact end users to arrange the replacement Free of ChargeIn addition, end users were asked to complete and return a Customer Reply Card to acknowledge the receipt of this information.If you have any questions regarding this recall please contact Michael Kelhart between the hours of 8:00 AM  4:30 PM EST at 1-800-437-2437 (press 1 at the prompt, then 32349). For questions regarding the operation and/or servicing of your Draeger anesthesia machine please contact DraegerService Technical Support between the hours of 8:00 AM  8:00 PM EST at 1-800-437-2437 (press 4 at the prompt then 4 again). |  | 31 | 31 | US Nationwide Distribution - Healthcare facilities in AK, AL, AZ, CA, CO, FL, GA, IA, IL, IN, MA, MD, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, TN, TX, UT, WA, WI, and WV. | Open, Classified |  |  | 202
Z-2424-2018 | 80298 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Version PTS-11.0.1.2 and PTS-11.0.2 |  |  | May 16, 2018 | 2018 | Ion Beam Applications S.A. | Correction vector confirmation message is lost if access point is changed after sending the correction vector. As a result, the patient will be treated in the setup position or treatment position instead of the corrected position. | Software design | On May 18, 2018 IBA issued URGENT FIELD SAFETY NOTICES to their customers. Customers were advised to take the following actions:It is recommended to ensure that the correction vector confirmation message appears on the expected access point. When the correction vector is applied (if deemed necessary1), ensure that the Patient Positioning System GoTo the corrected position for treatment before starting irradiation:- verify on adaPTdeliver screen that the corrected target position is selected before engaging the movement. - verify on adaPTdeliver screen that the corrected target position appears on the screen after the movement.Moreover, as advised in the Proton Therapy System Clinical User Guide, it is recommended to verify the patient position after each correction vector application.A software update will be deployed on all affected sites by July 2018.Customers with questions may email: vigiliance@iba-group.com |  | 3 | 3 units | Worldwide Distribution - US Nationwide in the states of FL and the countries of Sweden, The Netherlands | Terminated | October 16, 2018 | 153 | 3
Z-2264-2018 | 80099 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Nitinol Staple, Elastic Staple Small-Asymmetrical, Sterile R, Model #: a) ES-10x13x15; b) ES-10x16x19; c) ES-10x17x19; d) ES-12x13x15; e) ES-12x15x17; f) ES-12x16x19; g) ES-15x13x16 |  |  | April 13, 2018 | 2018 | TriMed Inc. | During a retrospective review of sterilization records, TriMed has determined that evidence of quarterlydose audits/verification required to comply with the regulations could not be found. Sterilizationrecords indicate that all lots of staples were irradiated to at least the defined dose. There is a risk that this issue could result in patient infection if not properly sterilized. | Process design | On 04/13/18, customer recall communication forms were sent via mail informing customers to quarantine effected devices, call the manufacturer to arrange device return at 800-633-7221, and report adverse events to the FDA's MedWatch. |  | 706 | 706 | U.S.: IL, UT, PA, NJ, WA, TX, NV, CA, NC, MS, KS, MO, MN, FL;OUS: Australia and Sweden | Open, Classified |  |  | 2755
Z-2261-2018 | 80181 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | P28 JAWS Nitinol Staple System kits are designed for use in an osteotomy, arthrodesis and fragment fixation of bones and joints of the foot including fixation of small bone fragments. |  |  | December 19, 2017 | 2017 | Paragon 28, Inc. | Missing drill guide as stated in the package. | Error in labeling | The firm, Paragon 28 Inc. sent an "MEDICAL DEVICE RECALL" letter dated Dec 19, 2017 to their consignees. The letter describes the product, problem and actions to be taken. The consignees were instructed to immediately review their inventory to determine if they have the devices; if you have possession of any of the devices, you must ship entire kit back to Paragon 28 Inc., 96 Inverness Drive, E. ste 1; Englewood CO 80112. If you have any questions, contact Senior Quality Manager at 720-643-1301 or email to: epohl@paragon28.com. |  | 1626 | 1,626 | Worldwide Distribution: US (nationwide) to states of: AL, AZ, CA, CO, DC, FL, GA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, NC, NJ, NV, NY, OH, OR, PA, SC, TN, TX, VA, WA, and WI; to countries of: Australia and South Africa. | Open, Classified |  |  | 1626
Z-2257-2018 | 80179 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dropper Plus Point-of-Care Urinalysis Dipstick Control kit |  |  | December 21, 2017 | 2017 | Quantimetrix Corporation | The kits contained an incorrect lot of Dropper Plus Level 1 vials | Mixed-up of materials/components | Communication with their only customer for this recall was via e-mail and phone. Upon notification by the customer the following activities were initiated:All lots of the Dropper Plus product in inventory were 100% inspected. The customer was ask to return all of the kits to Quantimetrix. For further questions please call (310) 536-0006. |  | 20 | 20 kits | US Distribution to the state of : Kansas | Open, Classified |  |  | 20
Z-2260-2018 | 79901 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | The BD Vacutainer Plus Plastic Citrate Blood Collection Tube is an evacuated blood collection tube that provides a means of collecting, transporting, and processing blood in a closed tube. The buffered sodium citrate additive provides an anticoagulated specimen that may be used for clinical laboratory coagulation assays. |  |  | April 11, 2018 | 2018 | Becton Dickinson &amp; Company | Approximately 3.4% of BD Vacutainer Plus Plastic Citrate Blood Collection Tubes (2.7mL) (light blue top tube) will exhibit over fill draw volumes from ~11-14%. | Nonconforming Material/Component | BD began notifying distributors on April 11, 2018. Customer communications were mailed on April 12-13th, 2018.US customers will be sent a letter via UPS beginning April 12, 2018.Distributors were e-mailed on April 11, 2018 and requested that they provide a list of their consignees to facilitate mailing to end usersBD Regions will be contacted via email regarding distribution outside of the USDistributors were instructed to return customer lists to BD. Letters will be mailed with UPS tracking, which provides for a 100% effectiveness check. Customer response form will be tabulated.100% of non-responders will be contacted by phone. |  |  |  | US/PR-All 50 states and PRGlobal-AR, BO, BR, BS, BZ, CA, CO, CR, CW, DO, EC, GT, GY, IN, JM, JP, KR, MX, NI, PA, PE, SG, SR, SV, TT, TW, UY, VIDates of distribution: June 2, 2017 to the present. | Open, Classified
Z-2269-2018 | 79998 | No | FET | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus JF-140F duodenoscopes (Model NumberJF-140F) |  |  | May 11, 2018 | 2018 | Olympus Corporation of the Americas | Issuance of validated, new reprocessing procedures. This corrective action is being undertaken as a result of ongoing manufacturer and FDA activities relating to reported patient infections associated with duodenoscopes. The new cleaning procedure requires additional recess flushing and forceps elevator raising/lowering steps during precleaning, manual cleaning and manual disinfection. | Under Investigation by firm | Firm is extending recall to the user levelLetters will be sent on May, 11, 2018 via overnight mail to customersFollow-up phone calls to consignees to offer training on new reprocessing instructionsNo action can be taken with hospitals that are no longer in business |  | 994 | 994 | Nationwide | Open, Classified |  |  | 1062
Z-0256-2018 | 78546 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | TERUMO(R) Overpressure Safety Valve, REF: (a) LH130 (b) LH130J |  |  | June 27, 2017 | 2017 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received reports of no flow through the over-pressure safety (OPS) Valve. | Process change control | The firm sent a safety alert to its consignees on 07/27/2017 by mail. The notice requested that the consignee test the product prior to using. If the product did not operate, the consignee was asked to return it.Beginning on 12/11/2017 the firm conducted a removal by mail of any remaining product. The customer instructions are as follows:CUSTOMER INSTRUCTIONSReview this Medical Device Recall Notice.Assure that all users receive notice of this issue.Refer to the Customer Response Form to identify your product that is subject to this action.Confirm receipt of this notification by completing and returning the attached Customer Response Form tothe email address or fax number indicated on the form .Terumo CVS will issue a Returned Goods Authorization upon receipt of the Customer Response Form.Return all affected products to Terumo CVS.QUESTIONS?We encourage you to contact Terumo CVS with any questions or concerns: Terumo CVS Customer Service: 1.800.521.2818Monday- Friday, 8 a.m . - 6 p.m. ETTerumo Safety Alert Fax: 1.410.392.7183 Terumo Safety Alert Email: TCVSCustomerResponseElkton@terumomedical.com |  | 297222 | 297,222 units | Worldwide distribution: US, Canada, Colombia, UK, Belgium, and Japan | Open, Classified |  |  | 511368
Z-2271-2018 | 80199 | No | OHA |  |  | Cardiovascular | 510(k) | Cryovalve(R) SG Pulmonary Human Heart Valve, Catalog Number SGPVVOO |  |  | May 14, 2018 | 2018 | CryoLife, Inc. | Devices, which met FDA and AATB eligibility criteria, were distributed. Final autopsy and toxicology reports indicated the presence of morphine in the urine. | Under Investigation by firm | The firm notified the consignees by letter on 05/14/2018. |  | 1 | 1 unit | AR, OH | Open, Classified |  |  | 1
Z-2282-2018 | 80220 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | ImagePilot, Model No. D9MA |  |  | May 09, 2018 | 2018 | Konica Minolta Medical Imaging USA, Inc. | Panasonic has issued a Mandatory Product Safety Notice, indicating a possibility of battery ignition (fire) in CF-C2 "Toughbook" laptops. Since Konica Minolta uses this laptop in some of its product configurations, it notified its customers to utilize a BIOS Utility for the laptops to minimize the risk of an incident while Panasonic develops a final solution. | Unknown/Undetermined by firm | The firm issued an Important Product Recall Information letter to their affected customers dated May 3, 2018, delivered via UPS Overnight Delivery. The letter notified customers of the Panasonic recall and how it affected the Konica Minolta product. The letter instructed customers to utilize a BIOS Utility for CF-C2 laptops to minimize the risk of an incident while Panasonic develops a final solution. Customers were provided with Technical Bulletin which detailed the steps.Correction Option 1: Update BIOS Environmental Setting using Utility Tool1. Download the BIOS Utility Tool from the Panasonic Website:ftp://ftp.panasonic.com/ computer/ software/ Panasonic-CF-C2-Charge-Tool.zip2. Load the Utility Tool onto the CF-C2 Desktop3. Double-click to run batch file: ChgCtrlUty.bat4. Optional tool: If Integrated Portable, an external keyboard may be required.Visit KM University at https://km-university.com/ for video instructions:https://km-university.com/course/notification-of-mandatory-safety-bulletin-of-panasonic-toughbook-cf-c2/Correction Option 2: Manual change to the BIOS Environmental Setting (for Integrated Portable, an external keyboard will be required)1. If power is on, turn the power off and restart prior to following instructions.2. Upon boot-up, at Panasonic splash screen, continuously hit the F2 key on the keyboard to access the BIOS.a. Go to the "MAIN" tab associated with the BIOS setting.b. Select "ENVIRONMENT Current Status" setting.c. Ensure "HIGH TEMPERATURE" is selected. Otherwise change this setting to high and click ENTER.d. Go to the "EXlr' tab associated with the BIOS setting.e. Select "SAVE VALUES AND REBOOr' and hit ENTER key on the keyboard.f. A prompt will appear to state, "SAVE VALUES". Select YES and hit ENTER.If you have any questions, need a keyboard, require assistance with these instructions, and/or a safety verification please contact our Customer Care Team at (800) 945-045 |  |  |  | Nationwide distribution. Foreign distribution to Argentina, Brazil, Canada, Chile, Colombia, Ecuador, Japan, Mexico, Netherlands, and Peru. | Open, Classified
Z-2277-2018 | 80176 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core128/Elite China 728324 |  |  | March 28, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | A problem has been detected in the X-Ray Tube. The affected tubes have a production defect that allows oil leakage from the tube housing, if it were to re-occur, could pose a risk for patients or users. | Other | Urgent Field Safety Notice letters were sent to customers on 3/28/18. The letters did not request any actions to be conducted by the customers.For Brilliance Big Bore/ Ingenuity CT / Ingenuity Core / IngenuityCorel 28 Customers:A Philips Field Service Engineer will contact you to schedule the FCOupdates at your site.Field Change Order (FCO 72800691 and 72800692) will be released to correct the issue.For Ingenuity TF PET/CT Customers:Field Change Order (FCO 88200520 and 88200511) will be released to correct this issue.A Philips Field Service Engineer will contact you to schedule the FCOupdates at your site. |  |  |  | The systems were distributed to the following US states: HI, KS, and MO.The systems were distributed to the following foreign countries: Algeria, Austria, Brazil, China, Indonesia, Japan, Kenya, Korea, Mexico, Poland, Turkey, and Uganda. | Open, Classified
Z-2283-2018 | 80272 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Pheno Interventional Fluoroscopic X-Ray System, Model 10849000 |  |  | May 01, 2018 | 2018 | Siemens Medical Solutions USA, Inc | If the C-arm is manually moved in the longitudinal direction to a fully extended position, the system's safety stop mechanism may be activated. As a result, the C-arm stops the movement and can drop up to 20 mm (0.8 ). Under certain circumstances, the cover of the detector or collimator can touch and possibly injure the patient due to the C-arm drop. This could lead to a situation in which it is necessary to stop the clinical treatment or continue the treatment on an alternative system. | Device Design | On May 1, 2018, the firm distributed Urgent Medical Device Correction letters to its affected customers.The letter instructed customers to follow standard emergency procedures they have in place in the event of system failures. The firm recommended considering reinforcing the emergency procedures until the update has been completed. As an immediate measure, operators were instructed to avoid driving the C-arm into a fully extended longitudinal position.The firm is currently developing a solution to eliminate the root cause of the problem. The solution will be distributed to all affected customers upon completion.Customers were asked to forward the safety information to any other organizations that could be affected. |  | 39 | 39 | Healthcare facilities in AL, CA, CO, FL, IL, KY, MA, MI, MN, MS, NE, NJ, NY, OH, TN, TX, UT, and VA. Foreign distribution to Australia, Austria, Belgium, Denmark, France, Germany, Italy, Japan, Kuwait, Netherlands, Norway, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 39
Z-2284-2018 | 80217 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The Proton Therapy System - Proteus 235 (brand names: Proteus Plus and Proteus ONE) |  |  | April 11, 2016 | 2016 | Ion Beam Applications S.A. | IBA is conducting a voluntary recall to address a PTS (Proton Therapy System) software issue and to reduce the risk related to this issue. | Software Manufacturing/Software Deployment | IBA mailed an Urgent Field Safety Notice to affected customers to inform them and prevent risk associated with this issue. IBA also recommends users always check for any discrepancy between the prescribed and delivered MUs on their clinical screen after an irradiation in which a pause occurred. IBA also recommends users always check for any discrepancy between the prescribed and delivered MUs on their clinical screen after an irradiation in which a pause occurred. A software upgrade will be deployed on impacted units to correct this issue. |  | 2 | 2 | Units were distributed to Korea and Jacksonville, FL. | Open, Classified |  |  | 2
Z-2285-2018 | 80142 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Space Pump IV Set, 2 Caresites, ASV, 0.2 Micron Air Eliminating Filter, Product Code 363420 |  |  | May 09, 2018 | 2018 | B. Braun Medical, Inc. | There is a potential for a different set to be mixed within the lot, where the incorrect set does not contain an air eliminating filter. | Process control | B. Braun Medical Inc. sent an Urgent Voluntary Medical Device Recall Notification letter dated May 9, 2018. The letter included the following instructions for customers:1. Review the Medical Device Recall Notification in its entirety and ensure that all users in your organization of the above mentioned product, and other concerned persons, are informed about this voluntary product recall and the affected product. Further use of this product should be discontinued immediately and affected product quarantined.2. If you are a distributor, please immediately forward this recall notification to your customers.3. Determine your current inventory of the affected lots within your facility. Do not destroy any affected product.4. Utilizing the attached "Product Removal Acknowledgement" form, record the total number of individual units (within partial cases) and the number of full-unopened cases. If you have no inventory remaining, please enter zero (0) on the form.5. Return the completed "Product Removal Acknowledgement" form to B. Braun Medical Inc. Quality Assurance department by faxing the form to (610) 849-1197 or e-mail to PA_QualityAssurance.BBMUS_Service@bbraunusa.com within two (2) weeks of receipt, even if the total inventory in your possession is zero (0). It is important this form is returned, so BBMI can meet its United States Food and Drug Administration regulatory requirement.6. If you have any full cases, partial cases or unused individual pieces of these affected products as identified in the " Product Removal Acknowledgement" form that was submitted to BBMI Quality Assurance Department, a BBMI Customer Support Representative will contact you to provide instructions for handling the affected product and arrange for return to BBMI. This may include credit or replacement product. For further questions, please call (610) 596-0500. |  | 1872 | 1,872 | US Distribution to the states of : CA, IL, and TX., and Internationally to Canada. | Open, Classified |  |  | 1872
Z-2286-2018 | 80022 | No | PFT | Reagents for molecular diagnostic instrument test systems. | Clinical Chemistry | Toxicology | 510(K) Exempt | ABX Pentra Micro ALBUMIN 2 CP, Serum/Plasma, REF 1300032563 |  |  | April 26, 2018 | 2018 | Horiba Instruments, Inc dba Horiba Medical | When the ABX Pentra Micro ALBUMIN 2 CP (ALBT2) Reagent 2 is sampled and dispensed immediately after an ABX Pentra Creatinine 120 CP (CREA3) Reagent 2 dispense, the final result of the microalbumin may be elevated more than 10%. This increase may affect serum, plasma and urine sample types. | Device Design | The firm, Horiba Medical, sent a "FIELD SAFETY NOTICE" dated April 26, 2018 to the customers on the customer list on April 26, 2018. The notice described the product, problem and actions to be taken. The notice directed the consignee to manually program the system to address the issue and the share the information with the laboratory staff and to complete and return the enclosed Acknowledgment Form within 10 days via Fax: (949)753-0533, Email: azita.hedayati@horiba.com or Mail: HORIBA Medical, 9755 Research Dr., Irvine, CA 92618.If you have any questions regarding this Product Corrective Action, please contact your local Horiba Medical representative or our Technical Service Department at 888-903-5001, option 3, option 2. |  | 686 | 686 devices (233 devices US, 453 devices OUS) | Worldwide Distribution: US (nationwide) and to countries of: Bulgaria, Burma, Cameroon, Djibouti, France, Gabon, Germany, Italy, Lebanon, Mali, Mauritania, Poland, Portugal, Romania, Saudi Arabia, South Africa and Thailand. | Open, Classified |  |  | 686
Z-2288-2018 | 80295 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | HLD Systems 600 Series Washer/Pasteurizer, UDI (01) 0085249007009 (11) 160929 (21) 61160; Model 610HT, 115V &amp; Model 610HT, 230V |  |  | April 09, 2018 | 2018 | Cenorin, LLC | There is a supplier quality issue with the terminal blocks which may cause the terminal block to overheat, char and/or burn. | Nonconforming Material/Component | On April 9, 2018, Cenorin telephoned customers and then followed-up with a formal letter titled URGENT MEDICAL DEVICE RECALL via email.To mitigate this quality issue, Cenorin has a Corrective Kit available to replace the existing terminal blocks in the Washer/Pasteurizer as the long-term solution. In the meantime, here is how you can help with the recall process:1. Please immediately examine your inventory for the recalled Washer/Pasteurizers listed and quarantine them to prevent further distribution. In addition, if you have further distributed this product, please notify your customers immediately using the enclosed notice.2. If the Washer/Pasteurizers does not exhibit any of the symptoms of failure as noted, continue to use the Washer/Pasteurizers as normal until the service can be completed. If the Washer/Pasteurizers exhibit any of the symptoms of failure, immediately cease its use and contact Cenorin immediately at recall@cenorin.com or call Technical Support at 800-426-1042 ext. 224.3. Complete the Recall Acknowledgement &amp; Response Form and return to Cenorin as noted on form via email to recall@cenorin.com within 5 business days. Response is required.Customers with questions, please email Cenorin at recall@cenorin.com or call 253-395-2400 or toll-free 800-426-1042, extension 236 during the hours of business Monday through Friday, 7:30AM to 4:30PM, Pacific Standard Time. |  | 14 | 14 units | Worldwide Distribution - US Nationwide in the states of CA, FL, IN, NC, NE, UT &amp; WV; and countries of Malaysia and Taiwan | Open, Classified |  |  | 14
Z-2291-2018 | 80300 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Amalgatome SD, COAXIAL PNEUMATIC HOSE, SCHRADER, Model X101004 |  |  | April 17, 2018 | 2018 | Exsurco Medical | Improper installation of hand piece may result in a weakening of the internal hose connection which can cause loss of pressure to the device or a hose burst. | Use error | On April 17, 2018 Exsurco Medical hand delivered notices to customers. Customers were advised to take the following actions:-Review, complete and return the Customer Response Card. -Upon verifying the information, provide the information to all Burn Account Managers.-An Exsurco Field Service Technician will perform the correction.Starting July 1, 2018, Exsurco Medical will issue URGENT Medical Device Recall notices and response cards to Burn Account Managers via certified mail or by hand delivery. Exsurso Medical will also provide replacement hoses simultaneously.Burn Account Managers are advised to take the following actions:-Review, complete and return the Customer Response Card. -Replace all impacted X101004 Pneumatic Hoses with replacement hoses.-Discard impacted X101004 Pneumatic Hoses as outlined in the recall notification.Should you have any questions regarding this matter please contact (440) 204-3239/ email: saigal@exsurco.com. |  | 70 | 70 coaxial hoses | CA, GA, MO, OH &amp; TX | Open, Classified |  |  | 70
Z-2290-2018 | 80252 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Radiation Therapy Treatment Planning System; 6.0, 6.1, 6.2, 7.0 |  |  | March 28, 2018 | 2018 | RAYSEARCH LABORATORIES AB | The dose calculation accuracy may in some situations be less than expected. The user must be aware in order to avoid incorrect dose calculations during treatment planning. | Software design | On March 28, 2018, the firm issued a Field Safety Notice, Medical Device Correction letter to its affected customers.The firm instructed users to take the following actions:-Do not use MLC only collimated to set output factor corrections (OFCs) in beam modeling of Elekta Synergy machines with MLCi/MLCi2 or any other machine with the MLC closer to the source than both the x- and y-jaws. Always use ÂżJaws and MLC collimatedÂż.-Be aware that there is an error in the field measure calculation for machines with the MLC closer to the source than both the x- and y-jaws, and for machines with fixed jaws. Make sure to perform patient-specific QA before treatment delivery for treatment plans with fields where: 1) closed MLCleaves are positioned inside the jaw field and 2) the minimum static/dynamic tip gap is larger than 0.001 cm.Please educate planning staff and all users about this workaround. Inspect your product and identify all installed units with the above software version number(s), then confirm you have read and understood this information by replying to the notification email.This issue will be resolved in the next version of RayStation/RayPlan, scheduled for market release in June 2018 (subject to market clearance in some markets). If customers wish to continue using versions of RayStation/RayPlan affected by this Field Safety Notice, all users must maintain awareness of this FieldSafety Notice. Alternatively, customers can choose to upgrade to the new version once it becomes available for clinical use. |  | 728 | 728 | US Nationwide Distribution in the states of AL, AR, AZ, CA, CT, DC, DE, FL, GA, HI, ID, IL, MA, MD, ME, MI, MO, MS, MT, NC, NJ, NY, OH, OR, PA, TN, TX, UT, VA, VT, WA, WI. | Open, Classified |  |  | 728
Z-2314-2018 | 80286 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis zeego, Material no. 10280959 |  |  | May 01, 2018 | 2018 | Siemens Medical Solutions USA, Inc | After the Large Display returns from power save mode, it may not show an image, and stay black without showing an error message although X-ray is still possible. The problem does not occur during an ongoing procedure. If the problem occurs, the system cannot be operated normally. It may be necessary to cancel or restart the treatment or transfer the patient to an alternate or another system. | Device Design | On May 1, 2018, the firm notified customers via Urgent Medical Device Correction letters. Users were warned that the Large Display may not show an image after it returns from power save mode. The Large Display stays black without showing an error message although X-Ray is still possible. The problem occurs sporadically, and does not occur during an ongoing procedure.The firm stated that there should be no influence on the previous treatment of patients.Users were instructed to follow standard emergency procedures in the event of system failures. It is recommended to consider reinforcing emergency procedures until the update has been completed. The firm's service organization will contact customers to arrange a date to correct the error by updating the firmware of the affected devices. Users were asked to promptly notify and instruct accordingly all staff within their organizations who need to be aware of the issue. |  | 56 | 56 | Worldwide distribution. US nationwide including Puerto Rico, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Canada, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Luxembourg, Mexico, Myanmar, Netherlands, New Zealand, Norway, China, Paraguay, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 232
Z-2300-2018 | 80271 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OmniDiagnost-Eleva X-ray system Angiographic and Diagnostic, Product Nos. 708026, 708027, 708028 |  |  | May 01, 2018 | 2018 | Philips Medical Systems Nederlands | The fixation of the upper and lower tilt actuator, of the Omnidiagnost system might break off and the table will start to rotate from 0 to + 90 /- 90 degree withhigh speed. This rotating movement can not be stopped by the user. This could potentially lead to the patient falling off the table. There is a risk of injury for the patient, user, or bystander. | Device Design | On May 17, 2018, the firm issued Urgent Medical Device Correction letters to affected consignees. Customers were advised of the product issue. Customers were instructed to do the following:Stop using the system and call your Philips representative if any of the following situations occurs:. the table height or tilt movement is not working.. you notice a cracking sound or snapping sound different than the normalsound during height or tilt movement.. you experience a blockade in the table height or tilt movement duringsystem movements.. the system is having an unexpected collision or the system had anunexpected collision in the past 2 months.. you have used a CPR stand that obstructed the system, or. you notice any (other) unusual system behavior beyond the normal useof the system.Avoid collisions with the table and base stand of the system by ensuring that noobstacles are placed around the system.Customer shall ensure that all staff with access to the affected systems areinformed of the content of this Field Safety Notice.A copy of this Field Safety Notice shall be placed together with thedocumentation of the system until the system has been corrected by Philips.The firm plans to correct all potentially affected products in the field by means of a field change order free of charge. Customers will be contacted by a local Philips representative to schedule the corrective action, which will start effective July 2018.Customer with further questions may contact Philips Technical Support Line: 1-800-722-9377. |  | 505 | 505 | Healthcare facilities in AZ, CA, CO, FL, IA, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, NJ, NY, OK, OR, PA, TN, TX, VA, WA, and WI. | Open, Classified |  |  | 980
Z-2302-2018 | 80250 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | Flexible Smoothbore Breathing Circuit with Exhalation Port 22mm, Product No. 80062000 |  |  | April 25, 2018 | 2018 | Intersurgical Inc | When assembling an HME or Filter on this breathing circuit the tubing may be pushedup and inadvertently cover or block the exhalation port. This could potentially pose aserious health risk to the patient. | Device Design | On April 25, 2018, the firm issued an Urgent Medical Device Recall letter to its distributors. Customers were advised of the product issue and recall.Recall Instructions:1.Immediately discontinue use and quarantine any stock with the product number B0062000 described above in this recall notification.2.If you have distributed any of this product, please immediately contact your customers, advise them of the recall situation and have them return their outstanding recalled stock to you.3.Contact Intersurgical Customer Support at 800-828-9633 for a Return Goods Authorization Number (RGA). Hours of operation are Monday through Friday, 8:00 a.m. to 4:30 p.m. EST.4.Once you have received the RGA Number, please enter it in the space provided on the Customer Product Recall Acknowledgement Form: a.Return the completed form via email to Intersurgical Inc. at Support@intersurgicalinc.com. b.The subject line should contain your Company name and RGA number. i.This will allow us to document your receipt of this letter and the amount of product you have on-hand for return. 5.Intersurgical will credit your account when the product is received.If you have any other questions, feel free to contact Intersurgical Customer Support at 1-800-828-9633. |  | 1340 | 1340 | Distributors in IL, NJ, NV, TX, UT. | Open, Classified |  |  | 1340
Z-2336-2018 | 80227 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Multi-Lumen Central Venous Catheterization Set with Blue FlexTip Catheter, CS-15853-E |  |  | May 11, 2018 | 2018 | Arrow International Inc | The recalling firm has received complaints related to increased resistance encountered when introducing the CVC over the spring wire guide. This may lead to procedural delay, and there may be need of medical intervention by exchanging the catheter for another CVC catheter. | Employee error | On May 11, 2018, the firm sent an Urgent Medical Device Recall letter to affected customers via FedEx.Distributors were instructed to take the following actions:1. Immediately discontinue distribution and quarantine any products with the product codeand lot number listed above.2. Using the provided customer letter and Recall Acknowledgement Form templates,communicate this recall to any of your customers who have received product includedwithin the scope of the recall.3. Have the customers return any affected product to you, together with a completed RecallAcknowledgement Form, for consolidation and return to Arrow International. In the eventthat an alternative approach is needed, contact Arrow International Customer Service formore information at 1-866-396-2111.4. To return affected products from your inventory, complete the enclosed RecallAcknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or emailto recalls@teleflex.com. This will allow us to document the amount of product you haveon hand for return. A customer service representative will contact you with a ReturnGoods Authorization (RGA) Number and will provide instructions for the return of productto Arrow International.5. Once you have completed returning all of the recalled products from your own inventory,and collecting and consolidating all of the recalled products from your customers, pleasecheck the box on the enclosed Recall Acknowledgment Form that indicates that youhave completed the recall and fax it to 1-855-419-8507, Attn: Customer Service, or emailit to recalls@teleflex.com. This will allow us to document completion of the recall.6. If you and your customers have no affected stock, please complete the enclosed RecallAcknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email torecalls@teleflex.com. This will allow us to document your receipt of this letter. |  | 1200 | 1200 | Distributor in FL. International distribution to Czech Republic, Dominican Republic, France, Greece, and Italy. | Open, Classified |  |  | 1200
Z-2317-2018 | 80267 | No | DTE | External pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | Single Chamber Temporary External Pacemaker, Model 53401 |  |  | June 06, 2018 | 2018 | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | A subset of Medtronic Model 53401 Single Chamber External Pulse Generators (EPGs) may revert from programmed settings to nominal settings during use with a patient due to an unexpected Power on Reset (POR). | Device Design | Medtronic began notifying customers, via letter, on about 06/06/2018. Customers were instructed to verify if they possessed affected 53401 EPGs, if they did have affected EPGs to determine if the firmware was updated, to notify all appropriate personnel, to notify customers if the affected devices have been transferred, and to complete and return the Customer Confirmation Certificate. Instructions for determining if the EPG firmware has been updated include: 1) powering the unit on and viewing firmware version; 2) if the firmware version is 01.03.00 or greater, the EPG firmware has been updated with the correction; and 3) if firmware version is 01.02.00, the EPG firmware has not been updated. Medtronic requests that customers arrange to have their EPGs returned for the firmware update. |  | 1108 | 1,108 devices | Domestic distribution to AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV. International distribution to Armenia, Austria, Bangladesh, Belgium, China, Dominican Republic, Finland, France, Germany, India, Ireland, Israel, Italy, Korea, Luxembourg, Malaysia, Nepal, Netherlands, Norway, Poland, Portugal, Romania, Saudi Arabia, Sierra Leone, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Tanzania, Turkey, United Arab Emirates, United Kingdom. | Open, Classified |  |  | 1108
Z-2334-2018 | 80244 | No | MXC | Medical magnetic tape recorder. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical CONFIRM Rx Insertable Cardiac Monitor, Sterile EO, Model #: DM3500 |  |  | April 16, 2018 | 2018 | St Jude Medical Inc. | The firm is advising physicians that exposure to sub-freezing temperatures during the supply chain process caused a transient battery voltage drop for a small number of Confirm Rx Model DM3500 Insertable Cardiac Monitoring (ICM) devices. This drop is an expected and normal behavior for this battery chemistry, but causes an incorrect display of a low battery indicator even after the battery voltage returns to normal. | Nonconforming Material/Component | The firm, Abbott, sent an "URGENT MEDICAL DEVICE CORRECTION" letters dated 05/18/2018, to its consignees notifying them that updated programmer software version 24.2.x will be made available over the next several weeks and will allow physicians to detect the presence of incorrect battery indicator prior to implant. Additionally, the programmer software will provide a mechanism to resolve the incorrect display for already implanted devices. If a low battery indicator is observed, customers were advised to contact firm technical services at 1-800-722-3774 to assist in confirmation and correction of the battery indicator display. |  | 986 | 986 | U.S. Nationwide Distribution: PR, MI, NJ, IL, FL, AR, PA, KY, CA, TN, IN, AL, TX, AZ, SC, OK, VA, NH, NM, WA, CO, IA, WI, NV, UT, OH, NC, MO, DE, WV, MD, OR, NE, VT, KS, MA, IA, WY, CT, GA, MS, LA, DC, LA, | Open, Classified |  |  | 986
Z-2337-2018 | 80242 | No | MGB |  |  | Cardiovascular | PMA | Angio-Seal Evolution Vascular Closure Device, Cat. No. C610135 |  |  | April 09, 2018 | 2018 | Terumo Medical Corp | Two lots of product were released for distribution prior to the completion of all required validation and lot acceptance activities. Although sterile, products from these lots may not function as intended and therefore should not be used. | Release of Material/Component prior to receiving test results | On April 9, 2018, the firm sent Urgent Voluntary Medical Device Recall letters and Inventory Confirmation Forms to customers via FedEx. The letter described the product issue.Customers were asked to:1. Review your Angio-Seal Evolution inventory immediately to identify and isolate affected inventory in order to prevent future use. 2. Your Terumo Territory Manager will visit your facility to confirm product inventory and assist you in completing the enclosed Recall Acknowledgement/Inventory Confirmation Form. This required documentation, along with the enclosed Fed-Ex product shipment labels will facilitate your return of affected Angio-Seal Evolution product to Terumo Medical Corporation. |  | 240 | 240 | Distributed to accounts in AL, CA, FL, IA, IN, MI, MO, NC, NY, SD, TN, and TX. | Open, Classified |  |  | 240
Z-2330-2018 | 80232 | No | NEC | Liquid bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Safe N Simple No-Sting Skin Barrier Film, Ostomy Skin Barrier Wipe, SNS00807 |  |  | June 18, 2012 | 2012 | Safe N Simple, LLC | The firm received a customer complaint of observed spots on the wipes. The manufacturer confirmed that the product had bacterial contamination. This leads to a potential risk of infection. | Process design | On June 18, 2012, the firm disseminated "Important-Action Required" recall information to affected distributors via email or phone call. The notification requested all distributors to immediately stop selling the product. The following instructions were provided:1. Quantity on hand should be emailed to safen'simple customer service manager, MichelleBliszack, mbliszack@snsmedicalwipes.com. We would like to know the current quantity youhave in inventory of the subject product by Tuesday, June 19th, 2012.2. We request the product with lot number 20120115 be either discarded or returned back tosafen'simple. We will arrange and pay for shipment back to our warehouse. safen'simple willreplace all existing stock of lot number 20120115 at no charge with a new lot number.3. Until we ship replacement product, appropriate substitutes in performance are productsSNS80725 and SNS80744. They are both similar product, however the size of the wipe issmaller. Packaging is 25 wipes/box and 100 wipes/box respectively. |  | 17592 | 17592 | Distributors in CA, FL, OH, MI, NJ, NY, PA, TX, and VA. | Open, Classified |  |  | 17592
Z-2315-2018 | 80019 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Diagnostic Ultrasound System, Aplio I-series. Labeled as i700, i800, i900 |  |  | January 04, 2018 | 2018 | Toshiba American Medical Systems Inc | The power supply unit may be damaged and may fail to start up when the power plug is inserted into a power outlet while the main power switch on the rear of the system isset to ON. | Software design | Canon sent an Urgent - Medical Device Correction letter dated January 4, 2018. Customers were notified of the issue and informed that a modified firmware will be installed to correct the issue. When that modified firmware becomes available, the Toshiba service representative will contact customers to schedule an appointment to install the firmware. Until the modified firmware is installed, instructions were provided to confirm that the main power switch is sent to OFF, insert the power plug into the power outlet, and to set the main power switch to ON. The system should be operated according to the procedure described in section 9.1, Connection the Power Cable and Protective Grounding, in the operation manual. Additionally, customers were asked to complete and return the attached Customer Reply Form. For further questions, please call (714) 730-5000. |  | 162 | 162 units total | USA (nationwide) Distribution to the states of : CA, CO, DE, GA, HI, IA, ID, IL, IN, MD, MI, MN, MO, NC, ND, NV, OH, OR, PA, SC, SD, TN, TX, VA, WA, WV and WY. | Open, Classified |  |  | 162
Z-2319-2018 | 80212 | Yes | LWS |  |  | Cardiovascular | PMA | Ellipse, Sterile EO, Model #/ Part #: |  |  | April 16, 2018 | 2018 | St Jude Medical Inc. | The firm is releasing new firmware for high voltage implantable cardiac devices (ICD) and cardiac resynchronization therapy defibrillators (CRT-D) intended to strengthen security by providing an additional layer of protection against unauthorized device access. For older generation devices not capable of accepting the firmware, the firm is advising customers of the option to permanently disable the RF communication capability in these devices. | Software design | On 04/16/2018, Important Medical Device Advisory notifications were mailed via overnight service to U.S. physicians and hospitals notifying customers of the availability of the programmer software update and associated ICD/CRT-D firmware update. Advisory notifications were mailed to customers outside the U.S. on different dates depending on geography and local regulatory requirements. Prophylactic replacement of affected devices is not recommended. The firm recommends firmware upgrades for all eligible patients at the next regularly scheduled visit or when appropriate depending on the preferences of the patient and physician. For devices not eligible for the cybersecurity firmware update the firm recommends a discussion of the risks of cybersecurity vulnerabilities and proven benefits of remote monitoring with patients at their next regularly scheduled visit. RF communication may be permanently disabled in devices not eligible for firmware updates during an in-clinic device interrogation with a programmer that has received the software update. Customers with questions about the firmware update can call the customer technical support hotline at 1-800-436-5056. |  | 139351 | 139351 | U.S. Nationwide distribution | Open, Classified |  |  | 811702
Z-2331-2018 | 79620 | No | CGJ | Follicle-stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VIDAS FSH, Ref 30407-01, IVD, Rx. The firm name on the label is bioMerieux SA, France. |  |  | April 09, 2018 | 2018 | BioMerieux SA | Invalid calibration with low calibrator S1 while using the product. | Under Investigation by firm | The recalling firm issued letters to the end users dated April 2018, via FedEx 2 Day Express explaining the issue, potential risk to health, and requested the impacted lots be destroyed. For further questions, please call (314) 731-8805. |  | 59 | 59 units distributed to the U.S. | Distribution was made to the recalling firm's distribution center in KY. There was no military/government distribution. | Open, Classified |  |  | 59
Z-2333-2018 | 80160 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ECOLAB MICROTEK Rolled Probe Cover, 2.6 cm x 20 cm (1 in x 8 in), Ref PC0906, packaged 1 probe cover/pouch, 25 pouches/carton, 2 cartons/shipping case, Sterile, RX. |  |  | November 22, 2017 | 2017 | ECOLAB INC | The incorrect IFU was shipped with the product. | Labeling False and Misleading | The recalling firm issued letters dated November 21, 2017, informing the customer the incorrect IFU was shipped with the product and a correct version of the IFU is being provided with the letter. The consignee was requested to notify downstream customers to the user level and that a copy of the letter could be forwarded to enhance communication. A Recall Response Form was enclosed for completion to be returned via FAX or email as soon as possible. The letter also informs the consignee that the recalling firm will send additional replacement IFU's if they have product in inventory. For further questions, please call (651) 250-4385. |  | 65 | 65 cases | US Distribution to the states of : CA, IL, LA, SC, TX, and WI. There was no foreign/government/military distribution. | Completed |  |  | 65
Z-2178-2018 | 80081 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet Cardiopulmonary Bypass Custom Tubing kit BO-TOP 38600 ADULT CPB, Part Number 701056035 |  |  | July 27, 2017 | 2017 | Datascope Corporation | There is a potential for the over-pressure safety (OPS) valve in the tubing kits to not allow flow through. | Nonconforming Material/Component | On 7/27/17, an Urgent Safety Alert was issued to customers of the Cardiovascular Procedure Kit with LN130B Over-pressure Safety Valves. The safety alert instructed customers to perform the following:Please examine your inventory immediately to determine if you have unused kits of the batch numbers listed. If so, please follow the recommended actions outlined in the attached Terumo Cardiovascular Systems Urgent Safety Alert. It is possible that custom tubing kits with the affected valves may have already been used without issue. Customers were instructed to test the valve prior to use.The notice was sent to customers again on 10/3/17 and 12/12/17.On 1/11/18, Urgent Medical Device Recall letters were sent to customers. The letters included the following instructions:Terumo CVS now recommends removal of any remaining product from your inventory.A review of our records indicates that you have received the following Maquet/Getinge Custom Tubing Kit batch numbers affected by the Terumo Cardiovascular Systems Medical Device Recall (refer to information provided in the table below).Please examine your inventory immediately to determine if you have unused custom tubing kits of the batch numbers listed. If so, please remove the affected products, quarantine them and follow the instructions for returning the product to Maquet/Getinge for full credit. It is possible that custom tubing kits with the affected valves may have already been used without issue.The Terumo Cardiovascular Systems Medical Device Recall affects the Custom Tubing Kit with LN130B Over-pressure Safety Valve distributed by Maquet/Getinge having lot/batch numbers listed above; no other product lot/batch numbers that were delivered to you are affected.Instructions for returning affected product:1. Review this notice and Terumos Medical Device Recall Notice (attached).2. Ensure that all users (i.e. perfusionist, operating room staff and biomedical staff) receive notice |  | 10 | 10 | The products were distributed to the following US states: IA and FL. | Open, Classified |  |  | 45
Z-2428-2018 | 79993 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BioMerieux mini VIDAS System in combination with Quality Control VIDAS* (QCV*) |  |  | January 04, 2018 | 2018 | bioMerieux, Inc. | Increase of events affecting performance regarding pump clogging and displacement of the color-coded SPR label (DOT). These performance issues could contribute to erroneous results for any assay performed on the VIDASÂż and mini VIDASÂż system due to a decrease of pipetted volume. | Device Design | Subsidiaries and distributors were initially notified of the recall on approximately 01/04/2018. U.S. customers were notified via letter on approximately 01/22/2018. On 03/28/2018 a follow-up notification was issued world-wide to customers to expand the scope. Instructions included to increase the frequency of the Quality Control VIDASÂż (QCVÂż) testing to weekly rather than monthly to reduce the period of potential retrospective analyses needed, continue to conduct a visual inspection of the SPRÂż after each run (including patient sample test and QCVÂż test) to ensure that the dot on the SPRÂż is still in place, continue to perform External QC testing in accordance with laboratory policy and in compliance with Governmental Regulations &amp; Accreditation Requirements, which includes CLIA, ensure the notification letter is distributed to and reviewed by all appropriate personnel, store the notification letter with the bioMÂżrieux VIDASÂż and /or miniVIDASÂż documentation, and complete and return the Acknowledgement Form in Attachment A by Fax to confirm receipt of the notice. For further questions, please call (314) 731-8694 |  | 1866 | 1,866 units | Worldwide Distribution - US Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, WV, WY. International distribution to Algeria, Angola, Anguilla, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Benin, Bolivia, Bosnia - Herzegovina, Brazil, Brunei, Bulgaria, Burkina Faso, Cambodia, Cameroon, Canada, Cape Verde, Chad, Chile, China / Hong Kong, Colombia, Congo, Costa Rica, Croatia, Czech Republic, Denmark, Djibouti, Dominican Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, Gabon, Georgia, Germany, Ghana, Greece, Guam, Guatemala, Guinea, Honduras, Hungary, India, Indonesia, Iran, Iraq, Israel, Italy, Ivory Coast, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Mali, Malta, Mauretania, Mexico, Mongolia, Morocco, Mozambique, Myanmar, Nepal, Netherlands, Nicaragua, Niger, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Fed., Saudi Arabia, Senegal, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syria, Tanzania, Thailand, Togo, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, Vietnam, Zimbabwe. | Open, Classified |  |  | 1866
Z-1817-2018 | 79974 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Endo GIA" Auto Suture" Universal Articulating Loading Unit 45mm - 2.5mm( Item Code 030454) |  |  | April 26, 2018 | 2018 | COVIDIEN MEDTRONIC | Device may be missing a sled component. The sled component is responsible for staple deployment. This may result in failure to form a staple line when tissue is divided, leading to bleeding or leakage of luminal contents. | Process control | On April 26, 2108, all consignees were issued were issued Urgent Medical Device Recall notices and Return Forms via Federal Express or certified mail. On May 22, 2018, all consignees received an UPDATED URGENT MEDICAL DEVICE RECALL notice. The updated notice included NEW product which has been processed under FDA Ref # 80423. The notice informs customers of the potential for use of a device with a missing sled and the actions they should take to discontinue use of the product and return all inventory of the affected lots to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. All in-control inventory and devices returned from customers will be reworked. **FOR INTERNATIONAL CUSTOMERS**Please be aware that the recalled units are packed within various convenience kits as identified on the Attachment B provided with your customer letter.Customers are also encouraged to EMAIL OR FAX THE ACKNOWLEDGEMENT FORM TO:- "Product purchased directly from Medtronic": feedback.customerservice@Covidien.com or fax to (800) 895-6140.- "Product purchased through distributor": FCAMITG@Covidien.com or fax it to (203) 492-7719. |  | 171271 | 171,271 units in total | Worldwide and US Nationwide | Open, Classified |  |  | 171271
Z-0706-2017 | 75755 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Spectral CT Applications supports viewing and analysis of images at energies selected from the available spectrum in order to provide information about the chemical composition of the body materials and/or contrast agents. |  |  | November 14, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | A software defect is causing issues with the IQon Spectral CT and Philips Spectral CT Applications while filming, saving and annotating images | Software design | A Field Safety Notice (Attachment #1 Field Safety Notice FSN 16-061-72800659 and FSN 88100051), will be sent to US consignees via certified mail. Philips Key Market representatives will distribute the notice internationally, based on local regulations. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 20 | 20 | Software was distributed in the following states: AR, IN, LA, MA, MN, OR and TN.Softward was also shipped to the following countries: Belgium, Denmark, France, Germany, Japan, Korea, Switzerland and United Kingdom. | Terminated | July 02, 2018 | 595 | 44
Z-2340-2018 | 80262 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CARESCAPE Monitor B650 |  |  | May 25, 2018 | 2018 | GE Healthcare Finland Oy | When multiple CARESCAPE Monitor B650 units are connected to the same network and a network overload occurs for a prolonged time, the monitors may simultaneously restart as designed. The monitor restart will not be completed until the network issue has been corrected. | Software design | The firm notified its consignees of the problem by letter on 05/25/2018. The letter stated the following safety instructions:"If a prolonged network overload occurs and CARESCAPE Monitor B650 on the network does not resume normal functionality in approximately 90 seconds, take the following actions: 1. Temporarily switch the monitor from central monitoring to local-only (bedside) monitoring: - Disconnect the LAN cables from IX/MC ports on backside of the monitor. The monitor will now be in local-only mode. - Keep the patient under close observation and continuous surveillance at the local (bedside) monitor. Disconnect used LAN cables from the monitor, see Figure 1, below, showing the cables to disconnect. 2. Contact your IT department or other appropriate personnel in your facility responsible for the Patient Monitoring Network, and provide them with the further instructions on network configuration provided.3.After the network issue has been resolved, switch the monitor back into central monitoring by re-connecting the LAN cables to IX/MC ports.4. Verify correct monitoring state and alarm function.If a transient network overload occurs in which the CARESCAPE Monitor B650 on the network restarts within approximately 90 seconds (i.e., all functionality resumes and the monitor reconnects to the network automatically), proceed with instruction step #2 above to check network functionality. If your network is not maintained by GE, please ensure that the network in your facility meets the network configuration requirements. GE Healthcare will provide revised network installation and user instructions documentation at no cost to you. Complete and return the Customer Response Form attached to this notification. Once the Customer Response Form is received, GE Healthcare will arrange for the documentation delivery. "NOTE: The letter has an attachment, "Appendix A", which has more detail on the net work configuration. |  | 69231 | 69231 units | worldwide | Open, Classified |  |  | 69231
Z-2342-2018 | 80311 | No | EMF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Safety Sideport Knife 1.0mm 45Âż, REF 378210 |  |  | May 14, 2018 | 2018 | Beaver Visitec | Blades are loose in handles and may fall out due to non cured epoxy | Nonconforming Material/Component | On May 14, 2018, Beaver-Visitec International (BVI) issued PRODUCT RECALL NOTICES to customers. Customers were advised to take the following actions:-Inspect stock, quarantine and destroy product per internal procedures. -Complete and return the response form via email to claimsus@bvimedical.com- Provide objective evidence of product destruction with BVI so that credit to customer accounts can be applied. Customers with questions may contact the firm by calling 866-906-8080. |  | 1354 | 1354 units | US Nationwide in the states of CA, IA, CO, HI, IA, MD &amp; TX | Open, Classified |  |  | 1354
Z-2339-2018 | 80195 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ratcheting Cannulated Driver Handle, Catalog #: 210-00-004, Rx Only, Non Sterile |  |  | March 27, 2018 | 2018 | Trilliant Surgical Ltd. | Firm is voluntarily recalling cannulated driver handles (210-00-003 and 210-00-004) that may demonstrate a degradation of function with the adapter collar sticking, and/or not retaining the mating adapter male shaft. | Other | On 03/27/18, the firm sent Medical Device Recall notices to customers informing them to: 1) Check inventory stock for identified devices; 2) Locate and remove and/or discontinue use of identified devices; 3) Return all related product(s) to the firm in accordance with an included return response / acknowledgement form. Customers were instructed to contact the firm at 1-800-495-2919 to organize product returns. |  | 90 | 90 | U.S. States: NE, MA, GA, TX, OR, AZ, FL, OH, WA, NJ, UT, VA, CA, MN, PA, DE, OK, KS, WI, MT, IA, MO | Open, Classified |  |  | 184
Z-2341-2018 | 80249 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci X Surgical System, Model IS4200. |  |  | March 07, 2018 | 2018 | Intuitive Surgical, Inc. | Specific da Vinci X Surgical Systems (IS4200) were shipped with defective Arm Controller Joint (ACJ) boards that were found to have a high rate of failure due to a defective component. | Nonconforming Material/Component | Intuitive Surgical, Inc. sent a Recall letters dated 3/7/2018 to affected customers via FedEx. The letter identified the affected product problem, risk to health, affected serial numbers, the action the firm will be taking, and the actions to be taken by the customer. |  | 143 | 108 ACJ's on 35 surgical systems | US Nationwide AR, CA, FL, GA, IL, IA, MN, MS, NE, OK, TN, and TX.Foreign distribution was made to Australia, Austria, Finland, France, Germany, Italy, Netherlands, Spain, and United Kingdom. | Completed |  |  | 143
Z-2343-2018 | 80231 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | 26-Lead Transfer Set for use with APEXÂż Compounding System, Product Catalog Number 2112550 |  |  | June 04, 2018 | 2018 | B. Braun Medical, Inc. | There is a potential for valves on the 26-Lead Transfer Set to be incorrectly oriented. When a transfer set with an incorrectly oriented valve is installed on the system, there is potential for unintended admixture. | Under Investigation by firm | Urgent Medical Device Correction - Recall Notification letters were distributed on 6/4/18. The letters instruct customers to perform the following:1. Review this Medical Device Correction Notification in its entirety and ensure that all users of the above mentioned products in your organization and other concerned persons are informed about this medical device correction - recall notification. If you are a distributor, please forward this recall notification to your customers. 2. Perform a visual inspection of valve orientation prior to installation onto APEX Compounding System in accordance with the Visual Inspection Instructions contained in Attachment 1. 3. If any misaligned valves are identified, do not use the transfer set. Please contact Customer Support at 1-800-227-2862 for product return and replacement.4. Utilizing the attached "Medical Device Correction - Recall Notification Acknowledgement" form, please acknowledge that you have received and reviewed this information.5. Return the completed "Medical Device Correction - Recall Notification Acknowledgement" form to B. Braun Medical Inc Quality Assurance department by faxing or e-mailing the form within two (2) weeks of receipt of this notice. It is important this form is returned to B. Braun Medical Inc can meet regulatory requirements of the United States Food and Drug Administration. |  | 12320 | 12320 | The products were distributed to the following US states: AL, CA, CO, CT, FL, GA, IL, LA, MA, MD, MI, MN, MO, NC, NJ, NV, OH, PA, TX, and WA. | Open, Classified |  |  | 12320
Z-2344-2018 | 80163 | No | DIP | Diphenylhydantoin test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products PHYT Slides, Catalog Number 8298671. in vitro diagnostic |  |  | May 09, 2018 | 2018 | Ortho-Clinical Diagnostics | There is a potential for imprecise results when using the slides. Increase in imprecision when using PHYT coating 0165. This leads to higher and lower than expected results for both quality control samples and patient samples. | Under Investigation by firm | Urgent Product Correction Notification letters were distributed on 5/9/18. The letters instruct customers to do the following:" Immediately discontinue using and discard all remaining inventory of VITROS PHYT Slides for the affected lots." Complete the Confirmation of Receipt form no later than May 23, 2018. Ortho will replace your remaining inventory of the affected lots or credit your account." Post this notification by each VITROS System in your facility or with the user documentation." Forward this notification if the product was distributed outside of your facility.If the customer is a distributor, they were also instructed to:" Send the letter to all customers who were shipped an affected lot of VITROS PHYT Slides. |  | 5712 | 5712 | Worldwide distribution. US nationwide, Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Colombia, Denmark, France, Germany, India, Italy, Japan, Mexico, Netherlands, Norway, Poland, Portugal, Russia, Singapore, Spain, and Sweden | Open, Classified |  |  | 5712
Z-2377-2018 | 80320 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Veo Insulin Pump |  |  | April 24, 2018 | 2018 | Medtronic Inc. | The MiniMed Paradigm Veo insulin pump has an error that impacts the Arabic language translation.This translation error occurs in the Predictive Alerts setting screen, which allows user to program alerts that will sound if users are predicted to reach their pre-set low or high sensor glucose values. | Under Investigation by firm | Medtronic Inc. conducted this field corrective action via Urgent Field Safety Notification to customers: patients and health care providers. |  | 239859 | 239,859 | International only. No U.S. distribution | Open, Classified |  |  | 239859
Z-2356-2018 | 80148 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | microTargeting" Drive System Distributed by Medtronic, Product Number 9033G0601 |  |  | May 10, 2018 | 2018 | FHC, Inc. | The FHC Depth Stop Adapter may cause damage to the lead and stylet when the depth stop screw is over-tightened onto the lead. This damage can manifest as low impedances or a short circuit between one or multiple electrode pairs in the lead, and has resulted in the need to remove and replace the implanted lead intraoperative or in a follow-up procedure. | Device Design | Urgent Medical Device Recall notification letters were sent to customers on 5/10/18. The letters included the following instructions:FHC, Inc. Field Representatives will review your inventory and replace effected Depth Stop Adapters you have on hand with an improved design. In the interim, we advise that you do not over-tighten the Depth Stop Adapter screw. If low impedances or a short circuit is found, physicians should replace the lead prior to use and report the issue to FHC, Inc. or your local field rep. |  | 2952 | 2952 total | The devices were distributed to the following US states: AL, AR, AZ, CA, CO, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, and WY.The devices were distributed to the following foreign countries: Argentina, Australia, Canada, Columbia, Denmark, Ecuador, France, Germany, Great Britain, Italy, Japan, Norway, Romania, South Africa, Switzerland, Saudi Arabia, Taiwan. | Open, Classified |  |  | 2952
Z-2360-2018 | 80322 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Vulcan Fast Red Chromogen Kit 2, part of an immunohistochemistry (IHC) procedure using an alkaline phosphatase (AP) detection system. Catalog numbers FR805-5L and FR805S |  |  | April 24, 2018 | 2018 | Biocare Medical, LLC | Chromogen failed to develop due to a stability issue caused by environmental factors with a raw material in the Vulcan Fast Red Buffer. | Material/Component Contamination | On April 24, 2018 the recalling firm issued a notification letter to their customers with instructions to return a response form. |  | 43 | 43 units | CA, VT, KY, LA, OR, RI, GA, IL, MA, TX, KS, FL, Japan, and Germany | Open, Classified |  |  | 43
Z-2376-2018 | 80281 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | REPLANT 5.0mmD ABUTMENT, Part number 6050-52-60 |  |  | May 18, 2018 | 2018 | Implant Direct Sybron Manufacturing, LLC | Due to out of specification, the clinician would not be able to fit and engage the abutment with the implant | Process control | On May 18, 2018, Implant Direct Sybron Manufacturing LLC issued URGENT MEDICAL DEVICE RECALL notices to their customers. Customers were advised to take the following action:1. Inspect your inventory for the affected product.2. Complete and return the Acknowledgement and Recall Return Form within 48 hours for Quarantine product and return products listed. 3. Please return the product and the firm will send you a replacement part. 4. Authorized Implant Direct Sybron Manufacturing distributors, the firm requests that you identify those customers that may have been shipped the affected product lot and contact these customers to inform them of this issue within fortyeight (48) hours. Customers with questions with questions may contact Implant Direct Sybron Manufacturing LLC Customer Care at 1-888-649-6425. Customer Care is available Monday through Friday, between the hours of 5am-5pm PST. |  | 60 | 60 units | CA, FL, GA, IL, LA, MI, MT, NC, NJ, NY, OK, OR, SC, SD, TN &amp; WA; International: Jordan, Great Britain, Germany, Spain, Italy &amp; Poland | Open, Classified |  |  | 60
Z-2370-2018 | 80188 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ZimmerÂż M/L Taper Hip Prosthesis With KinectiveÂż Technology Cementless, Size 16.25; 00-7713-016-00 |  |  | May 30, 2018 | 2018 | Zimmer Biomet, Inc. | M/L Taper Hip Prosthesis and M/L Taper Hip Prosthesis with Kinectiv Technology manufactured prior to October 2015 are being recalled due to the potential for black residue to be observed within the polyethylene bag containing the implant. The black residue is a combination of the titanium plasma spray used to create the porous surface of the implant and polyethylene bag containing the implant. The recalling firm states the residue is biocompatible. Risks include delay of surgery while locating a replacement device, and surgical intervention due to wear caused by debris. | Packaging | On May 30, 2018, the firm emailed Urgent Medical Device Recall letters to its Japanese subsidiary, who then further distributed the letters locally to Distributors and Risk Managers in Japan. The letter informed affected consignees of the product issue. Customers were instructed to do the following:Distributor Responsibilities:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. Request a Return Authorization Number by emailing zimmerbiometintlirarequests@zimmerbiomet.com.c. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.d. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.e. Mark RECALL on the outside of the returned cartons.Risk Manager Responsibilities:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.If you have further questions or concerns after reviewing this information, please call Zimmer Biomet customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call r |  | 665 | 665 | Products were distributed solely to Japan. | Open, Classified |  |  | 3931
Z-2378-2018 | 80197 | No | OFF | Venous blood pressure manometer. | Cardiovascular | Cardiovascular | Enforcement Discretion | BARD MYPICC Kit, REF CK000095B, 5F French Size, packaged 5 kits/case, Sterile, RX, For use in vascular access procedures. |  |  | January 12, 2018 | 2018 | Medline Industries, Inc. | One of the cases of product might not have been sterilized. | Process control | The recalling firm's sales representative telephoned their one customer on 1/12/2018 requesting return of the product. |  | 6 | 6/5-kit cases | The product was distributed to UT. | Completed |  |  | 6
Z-2379-2018 | 80046 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | BCI Advisor Vital Signs Monitor, Model 9200, Reorder numbers 920654235, 920654235, 920654335A, and 920674235A. (The differences in the reorder numbers have to do with certain functions being installed or not installed.) |  |  | November 29, 2017 | 2017 | Smiths Medical ASD Inc. | Some printing modes on the devices may not function properly (tabular data printing, real-time continuous printouts, and automatic interrupt print feature). | Software design | Smiths Medical sent an Urgent Medical Device Letters dated 11/29/2017 via FedEx tracked mailing. The letter identified the affected product, problem and actions to be taken. For questions call Smith Medical Customer Service at 1-800-258-5361. |  | 10 | 10 devices | Worldwide Distribution - US Nationwide in the states of CO and MT. Foreign distribution was to Armenia. | Terminated | July 11, 2018 | 224 | 10
Z-2380-2018 | 80318 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | da Vinci Xi EndoWrist Stapler 45 Reload Blue, Model |  |  | April 23, 2018 | 2018 | Intuitive Surgical, Inc. | Intuitive Surgical, Inc. (ISI) has recently become aware of an issue with the da Vinci Xi Stapler 45 Blue and Green Reloads that may be missing a component of the knife mechanism. Without this component, the reload knife will not move and cut tissue when the stapler is fired. | Nonconforming Material/Component | Intuitive Surgical, Inc. (ISI) conducted the recall and notified customers via letter "Field Safety Notice Urgent Medical Device Notification" sent the week of April 23, 2018. The letter notified customers of the recall, reason for field action, risk to health, affected products, Actions to be taken by the Customer/User, Actions to be taken by Intuitive Surgical, and Further Information and Support: If you need further information or support concerning this Medical Device Notification , please contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:"North and South America: (800) 876-1310, Option 3 (4 AM to 5 PM PST) or mail: customerservice@intusurg.com."Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8AM to 6 PM CET) or ics@intusurg.com"South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)"Japan: 0120-56-5635 or 03-5575-1362 (9 AM to 6 PM JST). |  | 7337 | 7337 | U.S., foreign and government accounts.Foreign: Australia, Austria, Belgium, Brazil, Canada, China (Hong Kong), Czech Republic, Denmark, Finland, France, Germany, Ireland, Italy, Japan, Luxembourg, Netherlands, Norway, Qatar, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom. | Open, Classified |  |  | 14153
Z-2383-2018 | 80299 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | VPS Rhythm ECG Accessory Pack, REF RHY-100-ECGO |  |  | May 17, 2018 | 2018 | Arrow International Inc | Product pouch label may not easily identify thelot number, expiration date and associated symbols.. | Labeling False and Misleading | On May 17, 2018, Arrow International issued Urgent Medical Device Recall Notices and Response Forms to impacted customers via FedEx 2-day. Customers were advised to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the product codes and lot numbers listed, so that the affected products can be returned to Arrow.2.To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. Soon after, a customer service representative will contact customers with a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to Arrow International.3.If you have no affected stock, please complete the Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow the firm to document your receipt of the letter.Customers with questions, feel free to contact your local sales representative or Customer Service at 1-866-396-2111. |  | 2070 | 2070 units in total | US Nationwide Distribution in the states of AR, AZ, CT, FL, GA, IL, IN, KS, KY, MA, MI, MN, MO, NC, NJ, NY, OH, PA, SC, TX, VA &amp; WA | Open, Classified |  |  | 2070
Z-2389-2018 | 80008 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Elecsys Vitamin D total II, Vitamin D Test System, Material number 07028148190 |  |  | April 27, 2018 | 2018 | Roche Diagnostics Corporation | The device may give a falsely elevated result that is non-reproducible. If vitamin D level is monitored by laboratory tests and adapted based on the result, a dosage reduction of vitamin D supplementation cannot be excluded, and as a consequence, vitamin D deficiency may develop over time. | Under Investigation by firm | On April 27, 2018, the firm issued an Urgent Medical Device Correction letter to all affected customers. The letter advised customers in the following workaround to improve sample quality prior to measurement with the Elecsys Vitamin D total II assay. The firm stated these steps are to be carried out temporarily until further information is provided:1. Perform a thorough inspection of your pre-analytical handling to ensure a correct assay performance of Elecsys Vitamin D total II and to comply with the individual specifications of the primary tube manufacturers for all tubes in use (centrifugation conditions are particularly important as well as the elimination of foam/bubbles from the surface of the sample).2. If the issue persists, switch from plasma to serum.3. If you still encounter the issue for plasma or serum, re-centrifuge the sample in a secondary tube for 10 minutes at 2000 x g prior to measurement with the Elecsys Vitamin D total II assay.If you are still experiencing issues after completing the steps above, discontinue the use of the Elecsys Vitamin D Total II assay. Contact the Roche Support Network Customer Support Center at 1-800-428-2336. For the MODULAR ANALYTICS E 170 module, cobas e 411 analyzer, and cobas e 601 and 602 analyzers, the availability of the Elecsys Vitamin D Assay (catalog number 06506780160) has been extended and may be used as an alternate method until this issue is resolved.Customers were asked to do the following: " Follow the workaround outlined in this Urgent Medical Device Correction (UMDC) to improve sample quality prior to measurement with the Elecsys Vitamin D total II assay until further information is provided." If none of the workaround actions listed on the previous page successfully resolve the issue, please discontinue use of the Elecsys Vitamin D total II assay and contact the Roche Support Network Customer Support Center at 1-800-428-2336." For all analyzers except the cobas e 801 module, the V |  | 347 | 347 units | US Nationwide distribution. | Open, Classified |  |  | 3938
Z-2391-2018 | 80110 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | NOVAPLUS(R) Pregnancy Cassette Device with Droppers, Urine hCG Pregnancy Cassette, Cat. No. VPT1-CASS |  |  | October 20, 2016 | 2016 | Ekla Corporation | Incorrect hCG results | Process control | Emails were sent to Distributors and Medical Facilities on 09/21/2016. Follow up phone calls also were made. |  | 295 | 295 kits | US National Distribution in the states of CA, IL, MN and RI | Open, Classified |  |  | 295
Z-2392-2018 | 80255 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Reliance 1227 Cart and Utensil Washer/Disinfector, FW03101, FW03102, FW03S003 |  |  | May 23, 2018 | 2018 | Steris Corporation | The firm has become aware that the Reliance 1227 Cart and Utensil Washer/Disinfector s Chemical Low Level alarm, intended to prevent the user from initiating a cycle when a low chemical level situation occurs, does not operate as intended. Currently, if a low chemical level situation occurs, the alarm will only be generated at the unit s next power-up or when accessing service mode. The failure of the Reliance 1227 to identify that the chemicals used for cleaning are low or empty could result in bedpans and urinals, basins, case carts, beds, theater shoes and other miscellaneous reusable patient care items not being properly cleaned before disinfection or reuse. | Software Design Change | On May 23, 2018, the firm notified affected customers by Urgent Device Field Correction letters distributed via FedEx. The letter stated that the system software had been updated to correct the issue, and that a STERIS service technician would be contacting each facility to perform the system software update.In the interim of the software update, users were instructed to continue to verify the units chemistry levels daily as part of the facilitys Quality Management System. Chemistry containers with low chemistry volume should be replaced prior to initiating the next cycle.If you have questions regarding this matter, please contact Katilin Slaper, Product Manager, at (440) 392-7431, STERIS Customer Service at 1-800-548-4873, or your local STERIS Representative. |  | 118 | 118 | Worldwide Distribution - US Nationwide in the states of AL, AK, CA, CO, CN, DE, FL, GA, IL, IA, KS, KY, LA, MA, MI, NJ, NY, ND, OH, OK, OR, PA, RI, SC, TN, TX, VT, VA, WI, and WV. Foreign distribution to Canada. | Open, Classified |  |  | 118
Z-2410-2018 | 80294 | No | KDP | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Vacuum Regulator, Push-T-Set Intermittent Suction Unit (PTS-ISU) |  |  | March 12, 2018 | 2018 | Ohio Medical Corporation | Final quality control testing was not completed before the devices were distributed.. | Employee error | Ohio Medical notified customers via phone call on about 03/12/2018. The notification included identification of the devices distributed to that customer, explanation that due to the need to the missed additional testing they were removing the devices, and providing a replacement until the testing can be completed. |  | 11 | 11 units | US Nationwide Distribution in the states of NC, IN, TN, LA,, FL MI, IL. | Open, Classified |  |  | 11
Z-2393-2018 | 80131 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Tibial Nail, Standard T2 Tibia Âż12x315 mm, Product Number 18221231S |  |  | April 27, 2018 | 2018 | Stryker GmbH | The manufacturer has discovered that potentially out-of-specification products may have left the factory. Non-conforming cannulation of the nails may result in reduced component strength and potentially premature nail breakage. | Employee error | Urgent Product Recall letters dated 4/27/18 were distributed to customers. The letters instructed customers to perform the following actions:1. Inform individuals within your organization who need to be aware of this device recall.2. Immediately check all stock areas and/or operating room storage to determine if any devices from the affected product list are at your facility. Response is required, even though you may not have any physical inventory on site anymore.3. Quarantine and discontinue use of the recalled devices.4. Maintain awareness of this notice internally until all required actions have been completed within your facility5. Inform Stryker if any of the subject devices have been distributed to other organizations.a) Please provide contact details so that Stryker can inform the recipients appropriately.b) If you are a Distributor, note that you are responsible for notifying your affected customers. 6. Branches/Agencies: Return all affected products available at your location to:Stryker Orthopaedics/PFA Product ReturnsAttn: Distribution Inventory Team325 Corporate DriveDock M-EastMahwah, NJ 07430PR 1774317Event 60207. Hospitals/Branches/Agencies: Complete and sign the enclosed Business Reply Form and fax to 1-888-230-2125 or to Email: strykerortho6020@stericycle.com |  | 12 | 12 | The products were distributed to the following US states: AL, Co, FL, KY, NC, TN, UT, and WA. | Open, Classified |  |  | 48
Z-2412-2018 | 80433 | Yes | NKE |  |  | Cardiovascular | PMA | Percepta CRT-P MRI SureScan, REF W1TR01 |  |  | June 21, 2018 | 2018 | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | Potential for a device reset to occur in Percepta CRT-P MRI SureScan and Percepta Quad CRT-P MRI SureScan due to a timing interaction between the EffectivCRT Diagnostic and the Ventricular Safety Pacing feature (VSP). | Device Design | On 21-Jun-2018, Medtronic began notifying US physicians and health care professionals of this issue and provided Patient Management recommendations to ensure patient safety. A Software Update to SW040 Version 8.1 will be deployed to all Medtronic 2090 and Encore clinic and hospital programmers via the Software Distribution Network or USB. Field Representatives with customers who have Percepta CRT-P patients will be sent a USB device loaded with the updated software. In consultation with an Independent Physician Quality Panel, Medtronic recommends the following actions for physicians following patients:- Contact your local Medtronic Representative and schedule installation of the updated Percepta CRT-P Application Software (SW040 Version 8.1) onto Medtronic 2090 and Encore clinic and hospital programmers.- For a patient whose Percepta CRT-P device has experienced a Reset Alert or Observation: Consider scheduling an in-clinic device interrogation as soon as possible for the patients device to receive the automatic update.- For a patient whose Percepta CRT-P device has not experienced a Reset Alert or Observation: At their next scheduled in-clinic device interrogation, the patients device will receive the automatic update.Please complete the enclosed Clinician Confirmation Certificate and return via email to RS.CFQFCA@medtronic.com.For questions, contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 1842 | 1842 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 13718
Z-2411-2018 | 80293 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | FDI FORT DEFIANCE INDUSTRIES P2131 AUTOMATED FIELD STEAM STERILIZER, NSN: 6530-01-641-4641 |  |  | May 24, 2018 | 2018 | Fort Defiance Industries, LLC | During long-term storage (i.e. military depot), the interaction of stagnant water and dissimilar metals in the plumbing of the P2131 device produces deposits in the water system that can potentially interfere with the rear sight gauge and / or the water liquid level switch. | Device Design | Fort Defiance Industries (FDI) notified the consignees by email on 05/24/2018. The letter explained the issue. FDI has revised the (1) Installation, Operational, and Performance Qualification (IQ/OQ/PQ) activities, (2) preparation for transport / storage, (3) preventive maintenance, and (4) device warning statements. FDI has revised the procedures for safety switch verification during IQ/OQ/PQ and complete water drainage before transport / storage of the device. |  | 199 | 199 units | US Nationwide Distribution in the states of PA. VA. TX | Open, Classified |  |  | 199
Z-2409-2018 | 80284 | No | GMD | Cryptococcus neoformans serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Cryptococcal Antigen Lateral Flow Assay (CrAg LFA), Model # CR2003 |  |  | March 30, 2018 | 2018 | Immuno-Mycologics, Inc | Potential for LFA strips to improperly wick which may cause false negative cryptococcosis test results | Nonconforming Material/Component | On March 30, 2018, IMMY issued URGENT MEDICAL DEVICE RECALL notices to their customers via email. The notice instructed customers how to identify a device which may fail. The following are actions to be taken by the customer/user:" You may continue using the affected lots. Any testing performed on defective strips should be considered invalid and should be repeated." Please immediately notify lab personnel who are running the test to look for defective strips (as seen by fast wicking and/or very weak control lines)." Please immediately complete the attached Acknowledgement and Receipt Form even if you do not have any affected stock remaining and return via email to: customerservice@immy.com. Before contacting customer service, please have the following information available: approval to receive a no charge replacement and/or a no-charge PO and the shipping information, including an attention line. |  | 875 | 875 units | On March 30, 2018, IMMY began issuing URGENT MEDICAL DEVICE RECALL notices to customers via email. Customers were advised on how to recognize that the device has failed. The following are Actions to be taken by the customer/user:" You may continue using the affected lots. Any testing performed on defective strips should be considered invalid and should be repeated." Please immediately notify lab personnel who are running the test to look for defective strips (as seen by fast wicking and/or very weak control lines)." Please immediately complete the Acknowledgement and Receipt Form even if you do not have any affected stock remaining and submit via email to: customerservice@immy.com. " Ensure relevant staff members are informed of this recall, including relevant clinicians." If you have supplied any potentially affected product to another organization, please advise that organization of this recall and notify IMMY. " In case product is in transit, display this letter in a prominent place for one month.Should you need a replacement kit, you will need to return your affected inventory. Instructions will be provided by customer service on howto return kits to IMMY.If you have any questions, contact IMMYs Quality Assurance Department by calling 1-405-360-4669 Monday through Friday 8:30 AM to 5:00 PM Central Standard Time. | Open, Classified |  |  | 875
Z-2421-2018 | 80379 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic EnTrust Escudo D144DRG (OUS distribution) |  |  | June 21, 2018 | 2018 | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | Medtronic has become aware of a potential for loss of high voltage and anti-tachycardia pacing therapy in EnTrust and Escudo implantable cardioverter defibrillators (ICDs) as they near elective replacement indicator (ERI) voltage. Under certain circumstances, the device may display an immediate End of Life (EOL) Observation with no prior ERI alert. Though no ERI alert is triggered, there may not be enough remaining battery capacity to charge the high voltage circuits, resulting in an excessive charge time EOL Observation, leading to a loss of high voltage and anti-tachycardia pacing therapy. | Device Design | On 21-Jun-2018, Medtronic began notifying US follow-up physicians, and Risk Managers of health care facilities of this issue via hand delivery by Field Representatives. In consultation with an Independent Physician Quality Panel, MDT recommends the physicians following patients with affected devices: "Consider scheduling an in-office patient follow-up as soon as possible to assess the potential for this issue per the steps described below. "Use a programmer to ensure the Excessive Charge Time EOL&and; the Low Battery Voltage ERI&amp; Patient Alerts have been programmed to On-High "Instruct patients to contact your office if they hear device alert tones. "If this issue has occurred, an EOL: replace device immediately Observation will be displayed on the programmer QuickLook report. Schedule device replacement immediately. The notification provided Patient Management recommendations to ensure patient safety and effective high voltage therapy remain as the device battery voltage approaches its 2.61V ERI threshold. |  | 1867 | 1867 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 24741
Z-2422-2018 | 80282 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | AperFix AM Femoral Implant, Model numbers CM-2409 (9mm x 24 mm) and CM-2410 (10 mm x 24 mm). For orthopedic use in ligament reconstruction. |  |  | July 11, 2013 | 2013 | Cayenne Medical Inc. | Potential for AperFix 24 mm implant breaking inside the drilled femoral socket. | Device Design | Customers were contacted on approximately 07/11/2013 and were instructed to return any affected products on hand. |  | 976 | 976 units | US nationwide distribution. | Terminated | September 14, 2018 | 1891 | 976
Z-2425-2018 | 79602 | No | PLS |  |  | Pathology | PMA | PD-L1 IHC 22C3 pharmDx is a companion diagnostic used to determine PD-L1 levels for determination of patient treatment options. |  |  | February 02, 2018 | 2018 | Dako North America Inc. | To correct the kit's instructions for use as the storage time for gastric and gastroesophageal junction adenocarcinoma cut sections will be changed from six months to five months. | Process control | On February 2, 2018 the firm sent letters to the consignees with the following instructions; Description of the issue:The Field Correction has been initiated to correct the kit's Instructions for Use (I FU) as the storage time for Gastric and Gastroesophageal Junction (GEJ) Adenocarcinoma cut sections will be changed from six months to five months. This means that section 8.2: Cut Section Storage Recommendation will bechanged to the following: 'To preserve antigenicity, tissue sections, once mounted on slides, should be held in the dark at 2-8 ÂżC (preferred), or at room temperature up to 25 ÂżC. Slide storage and handling conditions should not exceed 25 ÂżC at any point post-mounting to ensure tissue integrity and antigenicity. 8.2.1 NSCLC Cut Section Storage RecommendationCut sections must be stained within 6 months when stored at 2-8 ÂżC (preferred), or at 25 ÂżC. 8.2.2 Gastric or Gastroesophageal Junction (GEJ) Adenocarcinoma Cut Section Storage Recommendation Cut sections must be stained within 5 months when stored at 2-8 ÂżC (preferred), or at 25 ÂżC."Note that only Gastric and Gastroesophageal Junction (GEJl Adenocarcinoma cut section storage time has been updated. Storage time for Non-Small Cell Lung Cancer (NSCLCl cut sections remains the ~ You may continue the use of PD-Ll IHC 22C3 pharmDx, Code No. SK006. However, we recommend thatyour laboratory follow the instructions noted to preseive the antigenicity, as this may decrease after five months of cut section storage. Potential patient risk:Based on our investigations, there is a remote possibility that staining Gastric and Gastroesophageal (GEJ) Adenocarcinoma cut sections may result in a weak staining or a false negative staining after five months of storage. However, there have been no reported incidents based on the previous storage time recommendation.Actions to be taken by the user:Within 10 calendar days, please take the following actions 1. Discard the IFU provided within the PD-Ll |  | 1051 | 1051 kits | AZ, CA, CO, CT, DC, FL, GA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, NH, NJ, NV, NY, OH, OK, PA, TX, UT, VA, WA, and WI | Completed |  |  | 1051
Z-2423-2018 | 80291 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Hyalomatrix, Product nos. 651105, 651103, 651104, and 651301 |  |  | May 10, 2018 | 2018 | Anika Therapeutics, Inc. | Hyalomatrix products are susceptible to degradation as a result of the decomposition of PVC in the packaging material during sterilization, which resulted in degradation of the film. Over the long term it is possible that product integrity may be compromised. | Package design/selection | On May 10, 2018, the firm notified its affected distributor via emailed memorandum of its voluntary recall. The distributor was informed of the affected lots and quantities, and asked to discontinue distribution and sales activities related the affected products. |  | 1711 | 1711 | US Nationwide Distribution in the state of IL. | Open, Classified |  |  | 1711
Z-2426-2018 | 80324 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Coronary Sinus (CS) Uni-Directional Diagnostic Electrophysiology (EP) Catheter, catalog numbers D135303 and D135304. |  |  | March 29, 2018 | 2018 | Stryker Sustainability Solutions | Thrombogenicity test results demonstrated that both the test devices (SSS) and control devices (OM) did not meet the internal specification of non-thrombogenic . | Use error | The recalling firm issued a notification letter to their sales representatives via email on 3/29/2018 with a response form. For customers who did not return the response form, their will be a subsequent email or fax to their sales rep. The company will continue to follow up via phone and email as needed. The letter identified the affected product, problem and actions to be taken. |  | 2722 | 2,722 units | Worldwide Distribution - US Nationwide and Canada | Open, Classified |  |  | 2722
Z-2427-2018 | 79735 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LigaSure" Exact Dissector Nano Coated 20.6mm-21cm |  |  | March 15, 2018 | 2018 | Covidien LLC | Devices were found with a sterile breach (puncture) in the pouch from a sort of 8,064 packaged devices. | Other | Medtronic Regional sent a notification letter to customers via overnight delivery. the letter identified the affected product, problem and actions to be taken. The letter request that you quarantine and return any unused products of the affected item codes and production lots. Unused products from the affected item codes should be returned. If you have distributed Covidien LigaSure" Exact Dissector listed, please promptly forward the information from this letter to those recipients. All unused products from the affected item codes must be returned. The potentially affected product was shipped to your facility between November 2017 and February 2018. |  | 936 | 936 | Worldwide Distribution in the countries of Australia, Austria, Belgium, Denmark, Finland, France, Germany, Ireland, Italy, Netherlands, Norway, Poland, Saudi Arabia, Spain, Swede, Switzerland and United Kingdom | Open, Classified |  |  | 936
Z-2577-2018 | 80404 | No | NQM | Urinary pH (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AmniSure ROM Test under the following labels: 1) Catalog Number FMRT-1-25 and 2) Catalog Number FMRT-1-25-CA |  |  | May 18, 2018 | 2018 | Qiagen Sciences, Inc. | Potential for control line not appearing on the test strip | Nonconforming Material/Component | On May 18, 2018 QIAGEN issued Urgent Voluntary Medical Device Recall notices and response forms to customers. Customers were advised to take the following actions:1) Immediately discontinue use of all impacted lots and return all unused product to QIAGEN.2) Share the Voluntary Medical Device Recall Notice with all users of the AmniSure ROM Test in your facility to ensure they are aware of this recall.3) Use the Return Response form to confirm receipt of the notification and to record quantity of material to return.If you have any questions regarding this notice, please contact Technical Services at 1-800-344-3631. |  | 1620 | 1620 kits | Worldwide - US Nationwide distribution an in the country of , Canada | Open, Classified |  |  | 1620
Z-2433-2018 | 80319 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AesDex, LLC Cardica C-Port xA PLUS Anastomosis System. Catalog Number FG-000100. |  |  | April 02, 2018 | 2018 | AESDEX | Following pressurization of the devices, damage to the device handle mechanism may occur, accompanied by a loud noise and potential dispersion of plastic fragments and components. | Under Investigation by firm | AesDex, LLC initiated a voluntary removal via Advisory Notice that was sent to all consignees in the United States on April 2 and 3, 2018. The Advisory Notice describes the issue, and instructs the consignees to stop any use of any product they may have in their inventory and return any unused product. The notice was sent via FedEx and a response form has been included with the notice in which the consignee is requested to fill-out the response form and return the response form to AesDex via Fax or email. |  | 587 | 587 units | U.S. and international: Belgium, Germany, Italy and Netherlands. No gov accounts. | Open, Classified |  |  | 987
Z-2430-2018 | 80069 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Apollo XT Suture Anchor, 5.5 x 15 |  |  | January 22, 2018 | 2018 | Valeris Medical, LLC | The label on Apollo XT 5.5mm Suture Anchor (part # XTS-5515; lots 2358, 2362, and 2363) indicated that the anchor screw material was titanium, however that part number is made of PEEK. | Process control | Valeris Medical notified customers on approximately 01/20/2018 via phone call, text, or email. A formal letter was mailed to customers on 02/09/2018. The letter identified the affected product, problem and actions to be taken. Customers were instructed to remove affected products from inventory to quarantine, affix replacement labels provided by Valeris over the current labeling, and notify any customers if product was further distributed to so the product can be quarantined and returned for labeling. Additionally, Valeris requested confirmation in writing that the field correction action has been completed. |  | 113 | 113 units | US Nationwide Distribution in the states of California and Oklahoma. | Open, Classified |  |  | 113
Z-2431-2018 | 80256 | No | LGW |  |  | Neurology | PMA | Vectris(TM) Trial Screening Lead Kit for Spinal Cord Stimulation, Models: |  |  | May 25, 2018 | 2018 | Medtronic Neuromodulation | This voluntary recall is being conducted due to the curved tip introducer needle which is included in the kit, to have a potential manufacturing defect. This could result in difficulty advancing or withdrawing the Vectris lead through the curved tip introducer needle. | Process control | Beginning 25 May 2018, Medtronic sent a letter via mail to affected customers. The notification included the customer letter, confirmation form (reply form), and a list of affected product serial numbers. The letter informed customers of the recall, and provided instructions on the identification and return of affected products.The firm expanded the recall, notifying their consignees by letter on 07/13/2018.Customers with device related questions can contact Technical Services at 1-800-707-0933 weekdays 7am-6pm CT. |  | 10800 | 10,800 units | Worldwide | Open, Classified |  |  | 18982
Z-2598-2018 | 80388 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACS Universal Loader (Optional) to the BD FACSLyric Clinical System, Model 662879 (Universal Loader), Rx Only. The firm name on the label is Becton, Dickinson and Company, BD Biosciences, San Jose, CA. |  |  | May 16, 2018 | 2018 | Becton, Dickinson and Company, BD Biosciences | The solenoid plunger tip length is too short resulting in insufficient engagement to the Universal Loader enclosure to lock during use. | Nonconforming Material/Component | The firm, BD Life Sciences, produced the "URGENT PRODUCT RECALL-CORRECTION" letter dated 4/6/2018 to its customers during an on-site visit. by the field service engineer made beginning 5/16/2018 to correct the issue. The visits were completed 6/15/2018.If you require further assistance, contact BD Customer/Technical SupportBD Customer Support at 877.232.8995 (prompt 2) for instruments in the United States. For customers outside the US, contact your local BDrepresentative or distributor. |  | 8 | 8 devices | US Distribution was made to FL, KS, KY, MA, NJ, and SD. | Terminated | August 13, 2018 | 89 | 8
Z-2440-2018 | 80374 | No | FPD | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MED-RX, Ref 54-1650R 5Fr X 16 inch (40cm) Polyurethane Feeding Tube For Pediatric Use with Orange Radiopaque Stripe and Enteral Connector, packaged 10/case, Sterile, RX. The firm name on the label is chs, Oakville, ON, Canada. |  |  | June 11, 2018 | 2018 | Canadian Hospital Specialties | The product may be packaged with the incorrect length of feeding tube as labeled. | Process control | The recalling firm issued a letter dated 6/1/2018 via email on 6/11/2018 to formally notify their customer of the reason for recall, discontinue distribution of the product, and complete the enclosed Product Recall Response Form for return via email. |  | 10 | 10/10-tube cases | Distribution was made to GA. The foreign recalling firm also distributed product to Canada. | Terminated |  |  | 10
Z-2436-2018 | 79784 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension VistaÂż CTNI Sample Diluent, KD692, SMN# 10445205 |  |  | March 21, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | The diluent may have an incomplete slit on the septum in the cap of the vial. A diluent cap with a defective septum has the potential to cause erroneously elevated results if a sample is autodiluted onboard the system, due to a reduced volume of diluent being pipetted from the vial. | Under Investigation by firm | Notification letters were distributed 3/21/18. The letters instructed customers to:Âż Inspect stock, discontinue use of and discard the Dimension Vista Diluent lots listed in Table 1.Âż Please review this letter with your Medical Director.Âż Review your inventory of these products to determine your laboratorys replacement needs and to provide information to Siemens for reporting to the authorities.Âż Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days.Âż If you have received any complaints of illness or adverse events associated with the product listed in Table 1, immediately contact your local Siemens Customer Care Center or your local Siemens Technical Support Representative.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product. |  | 1446 | 1446 | The products were distributed to the following US states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. The products were distributed to Guam, Saipan, and Virgin Islands. The products were distributed to the following countries: Australia, Austria, Bahamas, Belgium, Brazil, Canada, Canary Islands, Croatia, Czech Republic, Denmark, France, Germany, Italy, Japan, Lebanon, Netherlands, Poland, Portugal, Qatar, Russian Federation, Slovakia, Slovenia, Saudi Arabia, South Korea, Spain, Switzerland, and United Arab Emirates. | Open, Classified |  |  | 2838
Z-2439-2018 | 80301 | No | LJG | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx QuickFISH Slides, a component of the QuickFish BC Fixation Kit - BC, REF QFFIXBC1-25, Catalog Number: CS012 |  |  | May 14, 2018 | 2018 | Opgen Inc | Possible diminished performance prior to its established expiration date. | Under Investigation by firm | On May 9, 2018, OpGen, Inc. notified all affected customers by email. On May 14, 2018, OpGen, Inc. (formerly AdvanDx, Inc.) issued Medical Device Recall notices and acknowledgement forms to their customers via email. Customers were advised to take the following actions:- Inspect stock, quarantine the impacted product and dispose of it in theproper manner.- Complete and return the Acknowledgement of Receipt and Destruction of Materials Form via mail to: OpGen, Inc. Attn: Regulatory Affairs, 708 Quince Orchard Dr., Suite 205, Gaithersburg MD 20878; Fax to: 1 (30) 869-9684 or email to: regulatory@opgen.com.- Should you encounter any false positive or false negative results ineither control well of QuickFISH Slides Lot 17306401, please contact OpGen Technical Support immediately at techsupport@opgen.com.Customers with questions may contact the firm via telephone 1-866-376-0009 or by e-mail: techsupport@opgen.com. |  | 20000 | 800 total (32 kits/25 slides per kit) | Worldwide Distribution: US (nationwide) to states of: AZ, CA, IL, TX &amp; WA; and countries of: Italy &amp; Denmark. | Open, Classified |  |  | 20000
Z-2435-2018 | 80394 | Yes | NPT |  |  | Cardiovascular | PMA | Edwards SAPIEN 3 Transcatheter Heart Valve (THV), Model: 9600TFX, Size: 26 mm |  |  | June 11, 2018 | 2018 | Edwards Lifesciences, LLC | SAPIEN 3 THV was packaged with the incorrect valve size with respect to the label. | Labeling False and Misleading | On 06/11/2018, the firm sent an Urgent Product Recall notice, via sales representative hand delivery, to the customer requesting they return the customer response form and the affected device. The customer was advised to call Customer Service at 1-800-424-3278, if they had any additional questions. |  | 1 | 1 Valve | NJ | Open, Classified |  |  | 1
Z-2316-2018 | 80392 | No | REP |  |  |  |  | MX-10 surveillance laser system |  |  | May 08, 2018 | 2018 | Wescam Inc | Product has a nominal ocular hazard distance (NOHD) that is longer than specified and persons could be exposed to light in excess of the exposure limit. | Radiation Control for Health and Safety Act | Firm sent a Service Bulletin dated June 2018 to customers. The Bulletin identified the affected product, problem and actions to be taken. For question call 888-5-WESCAM or email support.wescam@L3T.com. |  | 29 | 29 | US Nationwide Distribution | Open, Classified |  |  | 29
Z-2467-2018 | 80165 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | T2100 Micro flex Drive Treadmill, powered |  |  | April 02, 2018 | 2018 | GE Medical Systems Ultrasound &amp; Primary Care Diagnostics, LL | A performance issue with customer owned spare parts, T2100 Microflex drive (2026182-002 or 2026182-004), was not addressed with a previous safety correction. If these parts were installed from customer owned stock on the T2100 Treadmill, uncontrolled walking belt motion during a stress exercise test could occur. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated March 30, 2018 to customers. The letter identified the affected product, problem and actions to be taken. The letter stated that GE Healthcare will correct all affected products at no cost. Customers were asked to complete and return the attached "Customer Response" form via email to DCAR.30085@ge.com and a GE Healthcare representative will contact you to arrange for a correction if necessary. For questions contact GE Healthcare Service at 1-800-437-1171 or your local service Representative. |  | 3721 | 3,721 devices total | Worldwide Distribution - US Nationwide in the states of AZ, CA, CO, FL, GA, ID, IL, IN, KS, LA, MA, MD, MI, MO, NC, NE, NJ, NV, NY, OH, OK, PA, SC, TN, TX, WA, WI. International distribution to Kenya DEMOCRATIC REPUBLIC OF CONGO, ALGERIA, ARGENTINA, Australia, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BHUTAN, BOLIVIA, BOSNIA-HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, BURMA(MYANMAR), CAMBODIA, CAMEROON, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DEMOCRATIC REPUBLIC OF CONGO, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, GEORGIA, GERMANY, GREECE, GUAM, Hong Kong, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JORDAN, AZAKHSTAN, KENYA, KOREA, KUWAIT, LEBANON, LIBYA, MACEDONIA, ALAYSIA, MAURITIUS, MEXICO, MOLDOVA, MOROCCO, NAMIBIA, NEPAL, NEW CALEDONIA, NEW ZEALAND, Nigeria, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIA, SERBIA--MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SUDAN, SWITZERLAND, SYRIA, TAIWAN, TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, UZBEKISTAN, VENEZUELA, VIETNAM, Virgin Islands, YEMEN. | Open, Classified |  |  | 3721
Z-2523-2018 | 80486 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Brivo XR385, model 5215463, Digital Diagnostic Radiographic System |  |  | September 16, 2016 | 2016 | GE Healthcare, LLC | The placing of weight on the VRAD detector, of the Brivo XR385 digital radiographic system, can introduce image artifacts, which can result in patient exam retakes. | Component change control | An Important Electronic Product Radiation Warning letter was sent to consignees beginning 9/16/2016. The letter identified affected product, described the issue, and stated that GE will correct the devices without charge.Questions can be directed to 1-800-437-1171. |  | 58 | 58 (55 US, 3 OUS) | AK, AL, AZ, CA, CT, FL, GA, HI, IA, IL, KS, KY, LA, MA, MD, MS, NJ, NM, OR, PA, TN, TX, VA, and Guam, and Canada. | Terminated | July 26, 2018 | 678 | 58
Z-2445-2018 | 80312 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | ADVANSYS MLP/DLP, Model Numbers: 181051S, 181052S, 181041S, 181042S, 181031S, 181032S, 181033S, 181021S, 181022S &amp; 181023S |  |  | May 22, 2018 | 2018 | NewDeal SA | Use of the impacted product may cause a superficial infection requiring PO antibiotics and wound care or deep infection requiring IV antibiotics and device removal. | Packaging change control | On May 22, 2018 Newdeal SAS, a company of Integra LifeSciences, issued URGENT VOLUNTARY MEDICAL DEVICE RECALL notices to their customers. On October 22, 2018 URGENT VOLUNTARY MEDICAL DEVICE RECALL notices were issued to additional customers. Customers were advised to take the following actions:1. Examine your inventory to determine if you have any affected lots in original product packaging identified within the notice.2. If you do have any of the affected lots in original product packaging, stop using the product immediately, remove them from service and place in quarantine until the product can be returned.3. Complete the Acknowledgement and Return Form and check the box: I do have affected products on the list and record the lot number.4. Or, complete the attached Acknowledgement and Return Form and check the box: I do not have affected products.5. Complete the other information on the form as indicated on the form. Keep a copy of the form for your records.Return the completed Acknowledgement and Return Form by email or fax indicating your receipt and review of the notification.Customers with questions regarding the instructions, please contact Customer Service at 1-800-654-2873 Monday to Friday 8:00 AM  8:00 PM EST or custsvcnj@integralife.com . |  |  |  | CA, CO, CT, ID, KS, MA, MD, MN, MO, MS, NC, NY, OH, OR, PA, SD, VA, WA &amp; WI | Open, Classified
Z-2442-2018 | 80382 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Health Harmony Mobile application software |  |  | December 06, 2017 | 2017 | Intel-GE Care Innovations LLC | It was discovered that in certain situations, including partial sessions and when taking adhoc measurements, the patient data was not synchronizing in a timely manner with the backend database, resulting in the patient's clinician not getting patient data tor one or two days. | Under Investigation by firm | The firm sent e-mails to their 4 consignees during the first week of December 2017. with the following information:As you may know, we have discovered a technical issue in our Health Harmony product that may result in some patient data not being transmitted to the Maestro 2 clinical dashboard in a timely manner (i.e. vital signs, health sessionanswers). This can also manifest as a patient unable to execute their health session. Our investigation shows that in some cases, your clinical team may not receive patient readings in the normal timeframe. We take this issue very seriously. We are focused on investigating the root cause while also mitigating the effects on patient management as our engineering team solves theissue. While we continue to investigate this issue, we are developing a mitigation and will begin deploying shortly. This includes a progress notice and a message on the patient tablet when the system is synchronizing and transmittingresults. The message will remind patients to stay in the program until their results have successfully uploaded and transmitted.Please be assured we are urgently working on addressing this issue.Your Care Innovations program management team will follow up with you to discuss potential impact on your standard of care, the next steps and answer any questions you may have about this situation.Thank you for your continued patience,Marcus GrindstaffCHIEF OPERATING OFFICERCare InnovationsCell: 916.548.3013 |  | 1302 | 1302 | US Nationwide Distribution in the states of CO, ID, PA, and LA | Open, Classified |  |  | 1302
Z-2470-2018 | 80384 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Universal Nozzle with snout 300x400; PBS dedicated nozzle with snout XL |  |  | June 27, 2018 | 2018 | Ion Beam Applications S.A. | Gantry rotation or proton irradiation can be allowed even if apertures or accessories are not correctly positioned in snout 300x400 or snout XL. This can lead to risk of crushing for patient and/or user, and risk of mistreatment for the patient. | Device Design | Recall was initiated on 6/27/18. Firm plans to provide the Urgent Medical Device Correction notifications to IBA site managers present on customer sites. The site managers will deliver the notification to customers by hand or email.The recall letter alerts users to the potential issue and instructs users to verify that apertures and accessories are fully inserted with both sides in the guide. The firm plans to deploy a solution on or before May 31, 2019, that will prevent occurrences of the issue. |  | 14 | 14 | US Nationwide Distribution in the states of FL, VA, PA, OK, NJ, WA, and TX. | Open, Classified |  |  | 14
Z-2469-2018 | 80302 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Fx SOLUTIONS OFFSET HEAD CoCr 50x20, REF 106-5000 |  |  | June 04, 2018 | 2018 | Fx Solutions | Breach in the external blister of certain lots of Humelock Offset Head 50 X 20 and Centered Head 50 X 19. | Package design/selection | The firm initiated the recall by email on 06/04/2018 in France, on 06/05/2018, to distributors in the US and elsewhere. The distributor (US Agent) sent the notice on to the division agents who store the product. The firm requested return of the units. |  | 145 | 145 units | Worldwide Distribution - US Nationwide and the countries of: France, Argentina, South Africa, Spain, Israel, Germany, Lebanon, Italy, Chile, Belgium, Greece, Mexico, Poland, Hong Kong, Switzerland | Open, Classified |  |  | 276
Z-2448-2018 | 80306 | No | GAW | Nonabsorbable polypropylene surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PROLENE SUTURE MONO BLUE - BV1755 Taper Point 3/8C, D10031 |  |  | May 07, 2018 | 2018 | Ethicon, Inc. | The needle associated with certain suture product codes has the potential to perforate the packaging due to a packaging design issue. A perforated package has the potential to cause a soft tissue injury to a healthcareprofessional handling the subject product or a patient infection due to a breach of sterility within the packaging ifit goes unnoticed by the healthcare professional. | Package design/selection | On May 24, 2018, the firm distributed Urgent: Medical Device Recall letters to affected U.S. customers.Customers were instructed to do the following:1. Examine your inventory immediately to determine if you have product subject to this recall (removal) onhand and quarantine such product(s).2. Remove the product subject to this recall (removal) and communicate the issue to relevant operatingroom or materials management personnel, or anyone else in your facility who needs to be informed.3. If any product subject to this recall (removal) has been forwarded to another facility, contact that facilityto arrange return.4. Complete the Business Reply Form (BRF) confirming receipt of this notice andfax or email it to Stericycle at 1-888-473-8015 or Ethicon3681@stericycle.com within three (3) businessdays. Please return the BRF even if you do not have product subject to this recall (removal).5. Keep this notice visibly posted for awareness until all product subject to this recall (removal) has beenreturned to Stericycle. While processing your returns, please maintain a copy of this notice with theproduct subject to this recall (removal) and keep a copy for your records.6. Customers are required to return all unused D-Special PROLENEÂż Polypropylene Suture &amp; D-SpecialPRONOVAÂż Poly (Hexafluoropropylene - VDF) Suture products that are in their inventory immediately.Only D-Special PROLENEÂż Polypropylene Suture &amp; D-Special PRONOVAÂż Poly (Hexafluoropropylene- VDF) Suture product subject to this recall (removal) returned by September 21, 2018 will be eligiblefor credit reimbursement. Any product subject to this recall (removal) returned after September 21,2018 will not be eligible for credit reimbursement.7. To return product subject to this recall (removal), photocopy the completed BRF, place it in the box withthe product, and affix the pre-paid authorized shipping label included with the recall notification letter.Ethicon will pay for |  | 1752 | 1752 | Accounts in AZ, CA, CO, FL, GA, IL, IN, KS, MA, MN, MO, NC, NE, NJ, NY, OH, OR, SC, SD, TN, TX, VA, WA, WI, WV. | Open, Classified |  |  | 43092
Z-2491-2018 | 80288 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hydroline Trumpet Valve, Trumpet Valve, 5mm x 33cm Probe, Double Spike, Ref ASU1201, packaged 1/box, 12 boxes/case, single use, RX, Sterile. |  |  | June 11, 2018 | 2018 | Cardinal Health 200, LLC | The suction valve may not close properly which could cause continuous suction. | Device Design | The recalling firm issued letters dated 6/11/2018 via FedEx Next Day Priority Mail requesting return of the product. |  | 63020 | A total of 63,020 valves were distributed. Specific product quantities were not provided. | Distribution was made to AL, AR, AZ, CA, CO, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MT, NC, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, and WY. Government and military distribution was also made. Foreign distribution was made to Belgium, France, Germany, Israel, Italy, Netherlands, Singapore, and United Kingdom. | Open, Classified |  |  | 63020
Z-2471-2018 | 79887 | No | OYC |  |  | Clinical Chemistry | PMA | t:slim G4 Insulin Pump with Dexcom G4 Platinum CGM |  |  | April 23, 2018 | 2018 | Tandem Diabetes Care Inc | The fuel gauge, the component that reads and reports the battery parameters to the pump, could provide inaccurate readings, which present to the user in one of two ways: A) by triggering a Malfunction 4 Alarm, or B) by triggering a succession of notifications prior to the pump shutting off, including Alert 2 (Battery Low), Alert 3 (Battery Very Low) and Alarm 12 (Battery Very Low). The alarms in both scenarios notify the user that the pump has stopped delivering insulin. | Device Design | On 04/23/2018, customers were sent Urgent Medical Device Voluntary Recall letters via overnight mail delivery and email. The letter instructed customers to call Customer Support at 1-877-801-6901 to arrange for replacement pumps. Customers were advised to continue using effected devices provided they kept Emergency Kits with backup insulin supplies with them. If customers noted unexpected battery behavior, malfunction 4 alarm, or the following alerts: Alert 2 (Battery Low), Alert 3 (Battery Very Low) and Alarm 12 (Battery Very Low), then customers were advised to discontinue use and use backup insulin supplies, while they waited for replacement pumps. The firm advised customers to fill out and return via fax, email, or mail Medical Device Recall Return Response forms. |  | 55 | 55 | U.S. Consignees: IN, AZ, UT, TX, NY, NJ, MD, MN, IL, CO, AL, WI, CA, PA, OH, OR, MS, FL, NM, ID, NV, KY, VA, | Open, Classified |  |  | 55
Z-2480-2018 | 80289 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Presource PBDS, Kit, Common Diabetes. Kit includes a exerciser volumetric 5000 ML device, two items of cloth, skin prep 2% CHG, a packet of 4 chlorhexidine gluconate antiseptic cloths, and Gatorade pouches. |  |  | February 20, 2018 | 2018 | Cardinal Health 200, LLC | A defect in the seal of the Gatorade pouch contained in kits caused the pouch to leak within a limited portion of the kit. | Packaging | Cardinal Health notified customers via Fed Ex overnight, on about 02/21/2018. The URGENT PRODUCT RECALL letter instructed customers to: 1) Inspect your inventory for the affected product code and lot number; 2)Segregate and quarantine all product on hand; 3) Return enclosed acknowledgment form via fax to 847-689-9101 or 614-652-9648, whether or not you have affected product, indicating the product code, lot and quantity of product youve quarantined; 4) Notify any customers to whom you may have distributed productaffected by this recall; 5)Contact Customer Service at 800-964-5227 to arrange return and credit. Customers were also instructed to return any affected kit on site. |  | 8928 | 8,928 kits total | US Nationwide Distribution in the states of California. | Open, Classified |  |  | 8928
Z-2479-2018 | 80464 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Oasis MRI system |  |  | June 18, 2018 | 2018 | Hitachi Medical Systems America Inc | The stainless steel belt which runs the length of the table underneath the tabletop is exposed when the tabletop is extended into the magnet for scanning. The firm has received two complaints where technologists have cut their fingers on this belt. | Device Design | On June 18, 2018, the firm distributed an Urgent Medical Device Correction Notices to affected customers. Customers were informed of the issue and advised that the firm would be adding protective guards to both sides of the belt to prevent accidental contact. The firm expects to begin installing the guards around fall 2018.In the meantime, customers are advised to avoid direct contact with the stainless steel belt.Questions, comments, or requests may be sent to recalls@hitachihealthcare.com. |  | 264 | 264 | Worldwide Distribution - US Nationwide distribution.in the states of Mexico. | Open, Classified |  |  | 264
Z-2484-2018 | 80329 | No | NLQ |  |  | General &amp; Plastic Surgery | 510(k) | Reprocessed HARMONIC AceÂż +7, 5mm Diameter Shears with Advanced Hemostasis (HARH), Item Number HARH36 |  |  | May 08, 2018 | 2018 | Stryker Sustainability Solutions | Stryker Sustainability Solutions has received an increase in reports indicating that the affected Reprocessed HARH36 devices may display an error code (No Instrument Uses Remaining) upon initial connection to the generator. When this error code is present, the device is not able to be used. | Device Design | Firm notified sales reps via email on 5/8/18 and instructed to have the customers complete response forms. |  | 180 | 180 units | US Nationwide Distribution in the states of PA, FL, MD, MA, CA, NY, NC, TX, MT, AL, LA, TN | Open, Classified |  |  | 180
Z-2441-2018 | 80418 | No | RGW |  |  |  |  | Disinfection unit UV (432332-9020-000) for the Celldiscoverer 7 microscope |  |  | May 29, 2018 | 2018 | Zeiss, Carl Inc | Under certain circumstances, the firmware makes it possible for the Disinfection unit UV (432332-9020-000) to activate outside of the Celldiscoverer 7 housing. This may result in exposure of the users to harmful UV radiation. | Radiation Control for Health and Safety Act | The Company will immediately notify by direct contact (when available) and/or express courier each first line purchaser and/or known user site with informing them of the risk, how to prevent it and the notice of action.The Company will dispatch factory trained service representatives to perform following actions:oInstallation of Hot Fix # 14 (or newer) for ZEN 2.3oInstallation of Firmware 1 .209 (or newer) on nodes x l 5 and x l 9 (Main Controller 2) |  | 8 | 8 | US Nationwide Distribution | Open, Classified |  |  | 8
Z-2497-2018 | 80459 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation stand-alone software treatment planning system, RayStation 4.5, RayStation 4.7, RayStation 4.9 (RayPlan 1), RayStation 5, RayStation 6 (RayPlan 2), RayStation 7 (RayPlan 7) and RayStation 8A (RayPlan 8A) |  |  | June 29, 2018 | 2018 | RAYSEARCH LABORATORIES AB | If the beam model has a highly asymmetric primary source, it is not correctly taken into account in the calculation of DMLC fields whenthe collimator is rotated. This could lead to potentially significant overdosage at delivery. The user must be aware of the issue to avoid incorrect dose calculations during treatment planning. | Under Investigation by firm | On June 29, 2018, the firm issued an Urgent Field Safety Notice, Medical Device Correction #32484 to affected customers via email. The firm notified customers of the issue with the affected models: If the beam model has a highly asymmetric primary source, it is not correctly taken into account in the calculation of DMLC fields when the collimator is rotated.The user was instructed to do the following: 1. Inspect the beam model for all LINACs with jaw movement per beam, x-jaws and where the MLC is not above both jaws (i.e., Varian style LINACs) that allow DMLC planning.2. If either the x- or y width of the primary source is below 0.01 cm or if they differ with more than a factor 2, do not use this beam model for DMLC. Contact RaySearch support for further assistance in adjusting the beam model and identifying potentially affected patients.3. Please educate physics staff about this issue.4. Inspect your product and identify all installed units with the above software version number(s), then confirm that you have read and understood this notice by replying to the notification email.The firm plans to resolve the issue in the next version of RayStation/RayPlan, scheduled for market release in 2018 (subject to market clearance in some markets). |  | 746 | 746 | US Nationwide Distribution to accounts in AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, LA, MA, MD, ME, MI, MO, MS, MT, NC, NJ, NV, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, VT, WA, and WI. | Open, Classified |  |  | 746
Z-2502-2018 | 80452 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | VOLISTA StandOP Surgical Light, |  |  | June 14, 2018 | 2018 | Maquet SAS | Certain VOLISTA StandOP Surgical Lights have been identified as having a potential light head detachment issue. Improper assembly of the bracket that connects the light head to the axle fork could result in the light head bracket breaking, causing the light head to detach from the arm and become suspended only from the cabling inside the arm. Use of an affected device may result in serious injury to patient and/or medical staff if the light head detaches and drops, impacting patient and/or medical staff during demonstration, service, or use. | Process control | On June 14, 2018, the firm issued an Urgent Medical Device Field Correction letter via FedEx to affected customers. Customers were advised of the potential light head detachment on VOLISTA StandOP Surgical Lights.Customers were instructed to do the following:1. Ensure that all users were made aware of the recall; distributors were asked to forward the recall letter to their customers' attention for appropriate action.2. Check your VOLISTA StandOP serial number versus the list of affected devices on page 5 andensure it is available for a Getinge Service technician visit.3. Per the VOLISTA Instructions for Use*, Section 11.2 (General Maintenance) - Daily Checks (performed by the user), examine your affected VOLISTA StandOP surgical light immediately. If the bracket in the device is improperly assembled, mechanical play will be observed when thelight head is manipulated in a way that is not expected to be encountered in clinical use. Further instructions for testing the device are included in the letter.4. Complete and sign the enclosed Customer Response Form (Page 5) and return to Getingeusing the fax number or email address provided on the form.A Getinge Service person will contact the person you listed on the Customer ResponseForm to schedule service for an update of your device, free of charge. You can also arrangea visit by a Getinge-trained or authorized service technician by calling Getinge CustomerSupport at (800) 950-9912 (select option 1). Customer Service hours: 8:00 AM and 5:00 PM(EST) Monday to Friday. |  | 151 | 151 | Worldwide Distribution - US Nationwide in the states of AR, AZ, CA, CO, CT, FL, GA, IA, ID, KS, KY, LA, MD, ME, MO, NE, NY, OH, OR, TN, TX, UT, and WA. Foreign distribution to Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil Cameroon, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Ghana, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Italy, Japan, Jordan, Kenya, Korea, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Namibia, Netherlands, Nigeria, Oman, Pakistan, Panama, Paraguay, Poland, Portugal, South Korea, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Tanzania, Yemen, and Zimbabwe. | Open, Classified |  |  | 151
Z-2501-2018 | 80481 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Visualase Cooled Laser Applicator System (VCLAS) labeled as: |  |  | March 23, 2018 | 2018 | Medtronic Navigation, Inc. | Medtronic has become aware that certain lots of the VCLAS 3MM, 10MM, and 15MM laser diffusing fibers were not tested at intended laser power levels, due to an issue related to calibration of the test equipment, which can lead to the potential for low laser energy output from fiber or fiber burnout at lower wattage. | Equipment maintenance | Customers were notified on about 03/23/2018 via URGENT: MEDICAL DEVICE RECALL letter. Instructions included to examine inventory for affected devices, contact Medtronic Technical Services at 1-888-826-5603 to arrange for the return of affected devices, and complete and return the Consignee Response Form. The recall was later expanded and customers were notified on about 05/16/2018 URGENT: MEDICAL DEVICE RECALL letter. Instructions did not change from the initial letter, but informed customers that the recall had been expanded to include two (2) additional lots. |  | 1283 | 1,283 devices total | Worldwide Distribution - US Nationwide in the states of AL, AR, AZ, CA, CO, DC, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI. International distribution to Israel. | Open, Classified |  |  | 1283
Z-2505-2018 | 80423 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Endo GIA 60 mm Extra Thick Black Articulating Reload with Tri-Staple Technology, Item Code EGIA60AXT |  |  | May 22, 2018 | 2018 | COVIDIEN MEDTRONIC | Device may be missing a sled component. The sled component is responsible for staple deployment. Thismay result in failure to form a staple line when tissue is divided, leading to bleeding or leakage of luminalcontents. | Process control | On May 22, 2018, all U.S. consignees received an UPDATED URGENT MEDICAL DEVICE RECALL and RETURN FORM notice via Federal Express or certified. The initial notice was sent to impacted customers on April 26, 2018.Customers are advised to take the following required actions:1. Immediately discontinue use, and quarantine affected products listed2. Return affected product as indicated.3. Complete the Recalled Product Return Form even if you do not have inventory and submit by either method below:* Product purchased directly from Medtronic: feedback.customerservice@Covidien.com or fax to (800) 895-6140.*Product purchased through distributor: FCAMITG@Covidien.com or fax it to (203) 492-7719. |  |  |  | Nationwide U.S.A. Internationally to: Albania, Algeria, Austria, Belgium, Bulgaria, Croatia,Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece,Guadeloupe, Hungary, Iceland, Ireland, Israel, Italy, Ivory Coast, Jordan,Kazakhstan, Kuwait, Latvia, Lebanon, Luxembourg, Macedonia, Malta,Martinique, Moldova, Morocco, Namibia, Netherlands, New Caledonia,Norway, Pakistan, Poland, Portugal, Reunion, Romania, Russian federation,Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden,Switzerland, Tunisia, Turkey, United Arab Emirates &amp; United Kingdom | Open, Classified
Z-2496-2018 | 80400 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity Digital PET/CT System, Model 882446 with the following UDI's: |  |  | May 21, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | The block used to connect the patient pallet to the couch may be assembled incorrectly. A screw may be misaligned in the patient pallet. | Software design | On May 21, 2018, Philips issued Customer Information notices to their clients. A Philips Field Service Engineer will contact customers directly to schedule the updates at your specific site. If you need any further information or support concerning this issue, please contact your local Philips representative. For SUPPORT in North America and Canada, contact the Customer Care Solutions Center 1-800-722-9377, follow prompts. |  | 32 | 32 units | Worldwide Distribution - US Nationwide in the states of CA, GA, LA, MD, MI, NY, OH, TN, TX, VT, Washington DC &amp; WIInternationally - Canada, Mexico, Belgium, France, Israel, Luxembourg, Netherlands, Singapore, Spain &amp; Switzerland | Open, Classified |  |  | 32
Z-2517-2018 | 80498 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 2M insert 15degree for MUTARS RS cup and LUMiC TiN orthopedic hip implant acetabular component. |  |  | December 28, 2017 | 2017 | IMPLANTCAST GMBH | Possibility that the retaining ring in the 2M inlay 15degree for MUTARS RS Cup and LUMiC TiN is incorrectly mounted. | Process design | Implantcast sent customers an Urgent Product Safety Information / Product Recall letter on December 28, 2017. The letter recommends that physicians follow their patients by radiological exam. There is a return form for the hospital or physician to confirm receipt of the letter. |  | 307 | 304, 3 in US | Distribution in Europe and US states of Texas, Rhode Island and Georgia. | Completed |  |  | 307
Z-2526-2018 | 80476 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis zee systems with Laird Cooling Unit for SSFD: Material # 7555118. Artis is an angiography system developed for single and biplane diagnostic imaging and interventional procedures. |  |  | May 25, 2018 | 2018 | Siemens Medical Solutions USA, Inc | Due to a defective sealing, coolant may inflow into the electrical parts of the system cabinet, which could cause a short circuit and cause the system to shut down suddenly during an ongoing procedure | Process control | Siemens issued on May 25, 2018 an URGENT: MEDICAL DEVICE CORRECTION customer notification to affected customers and advised a hardware update will be initiated . Users advised of the problem, potential risks and action to be taken. Siemens will correct the error by replacing all affected cooling units with Update Instruction AX076/17/S. Users may continue to use the Artis system while awaiting implementation of the corrective action. |  | 3 | 3 US | HI , IL, MI, NC, NY, VA, WI | Open, Classified |  |  | 9
Z-2520-2018 | 79701 | No | FGB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Uretero-reno fiberscope URF-P6RP6 |  |  | January 17, 2018 | 2018 | Olympus Corporation of the Americas | Breakage of the endoscope's insertion tube bending section during surgical procedures. | Device Design | On 1/17/18 an Urgent Medical Device Removal Action Notification letter was shipped to affected customer via Fed-EX. The notification informed customers the recall was initiated in response to customer complaints regarding the breakage of the endoscopes insertion tube bending section during surgical procedures.Customers were instructed to complete the following: 1. Inspect your inventory and identify any URF-P6/P6R models; 2.Olympus will contact your facility to make arrangements for return of your URF-P6/P6R fiberscope(s) for the device exchange. You will be provided instructions on returning the URF-P6/P6R for this exchange; 3.Olympus has discontinued previously distributed copies of the URF-P6/P6R Reprocessing Manual and the Operation Manual Inspect your inventory of the Reprocessing Manual and the Operation Manual, and discard any existing inventory of the URF-P6/P6R Reprocessing Manual and the Operation Manual; 4.Implement use of the enclosed Reprocessing Manual which recommends only sterilization methods, and the enclosed Operation Manual; 5. Ensure all reprocessing personnel are completely knowledgeable and thoroughly trained on the new reprocessing instructions in the new Reprocessing Manual; 6. If you may have further distributed the URF-P6/P6R, please identify your customers, notify them at once of this product recall, and appropriately document your notification process. Your notification to your customers may be enhanced by including a copy of this recall notification letter; 7. Please indicate on the enclosed questionnaire that you have received this notification. Return of the enclosed questionnaire will be deemed to be a request for a new replacement URF-P6/P6R fiberscope(s). Fax the completed form to 484-896-7128. |  | 2399 | 2399 | US Nationwide. Canada and Mexico | Open, Classified |  |  | 3461
Z-2522-2018 | 80230 | No | MOM |  |  | Cardiovascular | PMA | CardioMEMS HF System Hospital and Patient Electronics Units: |  |  | June 14, 2018 | 2018 | Abbott Laboratories, Inc. | Abbott is advising customers that a small number of CardioMEMS(R) Hospital Electronics Systems (Model CM3000) and Patient Electronics Systems (Model CM1100) may deliver a system error, known as Error 5. While this error message is intended to present if the electronics system exceeds a certain temperature, these units may deliver a false Error 5 message due to an incorrectly configured component within the device electronics. | Process control | The firm notified the hospital and clinics of the recall on 06/14/2018 and the patients on 06/15/2018 by letter. The letter explained that the units may deliver a system error, known as Error 5. While this error message is intended to present if the electronics system exceeds a certain temperature, these units may deliver a false Error 5 message due to an incorrectly configured component within the device electronics. The firm is requesting return of the units. Patients may contact Abbott Remote Care Technical Support at 1-844-692-6367 for assistance and further information. |  | 2521 | 2521 units | US, Canada, Netherlands, Germany, Great Britain, France, Ireland, Belgium | Open, Classified |  |  | 2521
Z-2518-2018 | 79622 | No | OCS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Uretero-reno videoscope URF-V2 |  |  | January 17, 2018 | 2018 | Olympus Corporation of the Americas | Potential breakage of the endoscope s insertion tube bending section during surgical procedures | Device Design | On 1/16/18 an Urgent Medical Device Removal Action Notification letter was shipped to affected customer via Fed-EX. The notification informed customers the recall was initiated in response to customer complaints regarding the breakage of the endoscopes insertion tube bending section during surgical procedures. Customers were instructed to complete the following: 1. Inspect your inventory and identify any URF-V2/V2R models; 2.Olympus will contact your facility to make arrangements for return of your URF-V2/V2R u videoscope(s) for the device exchange. You will be provided instructions on returning the URF-V2/V2R for this exchange; 3.Olympus has discontinued previously distributed copies of the URF-V2/V2R Reprocessing Manual and the Operation Manual. Inspect your inventory of the Reprocessing Manual and the Operation Manual, and discard any existing inventory of the URF-V2/V2R Reprocessing Manual and the Operation Manual; 4.Implement use of the enclosed Reprocessing Manual which recommends only sterilization methods, and the enclosed Operation Manual; 5. Ensure all reprocessing personnel are completely knowledgeable and thoroughly trained on the new reprocessing instructions in the new Reprocessing Manual; 6. If you may have further distributed the URF-V2/V2R, please identify your customers, notify them at once of this product recall, and appropriately document your notification process. Your notification to your customers may be enhanced by including a copy of this recall notification letter; 7. Please indicate on the enclosed questionnaire that you have received this notification. Return of the enclosed questionnaire will be deemed to be a request for a new replacement URF-V2/V2R videoscope(s). Fax the completed form to 484-896-7128. |  | 2548 | 2548 | US Nationwide and Canada | Open, Classified |  |  | 2548
Z-2524-2018 | 80417 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Pheno, Fluoroscopic X-Ray system, Model Number 10849000 |  |  | May 30, 2018 | 2018 | Siemens Medical Solutions USA, Inc | During manual LAO/RAO rotation at maximum speed or a combination of LAO/RAO and Cran/Caud rotation at maximum speed, the possibility exists for the safety monitoring within the system control to become activated. As a result, the C-arm stops the rotation and can drop up to 40 mm (1.6 ). | Device Design | An Urgent Medical Device Correction notification was distributed to customers on 5/30/18. The notification provided the following information to customers:Siemens will correct the error by inspecting all affected ARTIS Pheno systems. As an immediate measure to prevent movement stops during a C-arm rotation followed by a C-arm drop, the maximum rotation speeds will be reduced. This correction limits the maximum manual LAO/RAO rotation from 25Âż/sec to 20Âż/sec and Cran/Caud rotation from 20Âż/sec to 10Âż/sec.All other system parameters as well as 3D acquisition and Perivision/Peristepping remain unaffected. You may continue to use the ARTIS Pheno system while awaiting implementation of the corrective action.Our service organization will contact you shortly to arrange a date to perform this corrective action with UpdateInstruction AX004/18/S. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436.Following the corrective action, the cause is eliminated and a recurrence of the fault is prevented. Additionally, Siemens is developing a solution to restore the system back to the original settings. This solution will be distributed to all affected customers upon completion. |  | 40 | 40 in the United States | The devices were distributed to the following US states: AL, CA, CO, FL, IL, KY, MA, MI, MN, MS, ME, NJ, NY, OH, TN, TX, UT, and VA. | Open, Classified |  |  | 40
Z-2527-2018 | 80287 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MR Surgical Suite Option |  |  | March 30, 2018 | 2018 | GE Medical Systems, LLC | There is a potential safety issue with the table transfer release of the MR Surgical Suite System which may cause an unexpected loss of function of the patient transfer mechanism that allows the patient to be moved from the GE MR Surgical Table to the Maquet Magnus Surgical Table. | Nonconforming Material/Component | GE Healthcare notified customers on about 03/30/2018 via URGENT MEDICAL DEVICE CORRECTION letter. The letter informed customers that the MR Surgical Suited System could still remain in use, it referred them to the GE MR Surgical Suite Operator's Manual for detailed descriptions of the various safety interlocks and operating mechanisms, encouraged users to monitor for any issues releasing the transfer board from the GE MR Surgical Table, and instructed users to test the patient transfer features as part of a quality check prior to each use. The letter also informed customers that a GE Healthcare representative will contact customers to ensure they have the latest operator manual at their site, update preventative maintenance documentation to include a procedure that focuses specifically on this equipment, and GE Healthcare representative will contact customers for scheduling a visit at which time GEHC will conduct this preventative maintenance procedure for their device and install new hardware. Questions or concerns can be directed to GE Healthcare Service at 1-800-437-1171 or your the Service Representative. |  | 54 | 54 units total | US Nationwide distribution to AZ, CA, CO, FL, IL, KS, MI, MN, NY, OH, PA, TX, UT, VA, WI. Worldwide distribution to Austria, Brazil, China, Czech Republic, France, Germany, India, Italy, Japan, Korea, Kuwait, South Africa, Turkey, Turkey, United Arab Emirates, United Kingdom. | Open, Classified |  |  | 54
Z-2529-2018 | 80393 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Nobel Biocare 17Âż Multi-Unit Abutment, Conical Connection NP, 2.5mm, REF 36614. Product Usage:Intended to be used in the upper and lower jaw and used for supporting tooth replacements to restore chewing function. The Multi-Unit Abutments, in combination with endosseous implants, are indicated for multiple unit reconstructions when screw retrained prosthetics are preferred. |  |  | March 16, 2018 | 2018 | Nobel Biocare Usa Llc | The product label has missing information, such as the lot number and expiration date, as well as other information. | Nonconforming Material/Component | Nobel Biocare sent an Urgent Quality Message letter dated March 13, 2018 via email to the U.S. consignees. The letter identified the affected product, problem and actions to be taken. For question contact customer service department at 800-322-5001. |  | 74 | 74 units were distributed to the U.S. | Worldwide Distribution - US Nationwide in AR, CA, CO, DC, FL, IL, IN, KS, MD, MI, NE, NJ, NY, ND, OH, OK, PA, SC, TX, VA, WA, WI, and WY.Foreign distribution was made to Austria, France, Germany, Hong Kong, Italy, Japan, Lebanon, Mexico, Spain, and Switzerland. | Terminated |  |  | 74
Z-2528-2018 | 80078 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Invia Motion Negative Pressure Wound Therapy System. Labeled as: |  |  | February 07, 2018 | 2018 | Medela Inc | Device may display a battery missing error. | Device Design | Medela sent an Notification letter dated February 1, 2018 to customers via Certified Mail. The letter identified the affected product, problem and actions to be taken. Instructions included to place all affected items on hold until an insert outlining additional charging instructions were included with the devices. Those instructions included to provide a printed copy of the charging instructions in each device package, review those insert instructions with patients and care-givers, and complete and return the confirmation letter. For questions contact Medela Partner Services at info-healthcare@medela.com or via phone 877-735-1626. |  | 778 | 778 units | US Nationwide Distribution domestically to AR, CA, FL, GA, IL, IN, LA, MI, MN, MO, NC, NJ, NY, OH, OK, PA, TX. Distributed internationally to Spain, Italy, Denmark, Sweden, Switzerland, Netherlands. | Terminated | December 04, 2018 | 300 | 778
Z-2530-2018 | 80296 | Yes | NOU | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Vivo 65, Continuous Ventilator, Home Use, Cat. No. 204000 |  |  | May 08, 2018 | 2018 | Human Design Medical Llc | Some Vivo 65 devices have an unreleased version of the Firmware upgrade tool. | Employee error | On May 8, 2018, the firm notified all affected customers via phone to quarantine all affected devices. On May 18, 2018, formal Field Safety Corrective Action letters were sent to all affected customers via FedEx. Customers were asked to remove affected units from use and return to the firm, and to complete and return the Acknowledgement and Product Replacement forms. Customers were informed that the firm would be shipping replacement devices. |  | 45 | 45 | US Nationwide Distribution in the states of CA, FL, GA, MI, SC, and TX. | Open, Classified |  |  | 45
Z-2531-2018 | 80164 | No | LGX | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | Access Hi-Low Exam Tables, Model 6501-XX |  |  | March 29, 2018 | 2018 | Brewer Company, LLC | One lot, the wire harness was over-crimped on some of the connectors that go to the two vertical side switches on the exam table. | Nonconforming Material/Component | The Brewer Company notified customers, via "MEDICAL DEVICE CORRECTION" letter, on about 03/21/2018. Instructions included to cease distribution of the product, coordinate with Brewer to schedule rework of the cable in affected tables as quickly as possible, and to notify Brewer of customers if the product was further distributed to coordinate rework of those tables. Additionally secondary controls were provided that the operator must press and continually hold the foot control pedal to operate the table, the foot control includes a stop button, depressing any pedal on the foot control will stop the table's automatic movement, and the manual instructs the operator to remove any obstructions from under the exam table prior to movement. Questions can be directed to (888) 273-9371. |  | 7 | 7 devices | Devices distributed in WI, TX, MN and IL. | Open, Classified |  |  | 7
Z-2535-2018 | 80471 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | HexaPOD evo RT Couch Top (also known as HexaPOD evo Module), a component of the HexaPOD evo RT system. |  |  | May 16, 2018 | 2018 | Elekta, Inc. | Upon completion of a 6D workflow, the HexaPOD evo Module may be in a tilted state if it is moved to the position used for 3D (non-iGUIDE) treatments resulting in incorrect patient positioning. | Under Investigation by firm | The letter from Elekta dated May 2018 was issued via email 5/16/2018 and was flagged "URGENT IMPORTANT FIELD SAFETY NOTIFICATION. The letter identified the affected product, problem and actions to be taken. |  | 725 | 725 devices | Worldwide Distribution - US Distribution of the device was made nationwide. There was government distribution but no military distribution. Foreign distribution was made to Canada, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Botswana, Brazil, Chile, China, Colombia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Democratic People's Republic of Korea, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Peru, Poland, Portugal, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Sweden, Thailand, Turkey, United Kingdom, and Viet Nam. | Open, Classified |  |  | 725
Z-2536-2018 | 80375 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Arctic Sun Temperature Management System, Model 5000. |  |  | June 13, 2018 | 2018 | Bard Medical Division | An improper wiring connection on the chiller assembly may result in overheating of the connection and potential failure of certain components in the system. | Process control | The firm,Bard Medical, sent a "MEDICAL DEVICE CORRECTION" letter dated 6/13/2018 to its customers. Bard Medical advised their customers to examine inventory and identify any ARCTIC SUNÂż 5000 Temperature Management Systems that match the attached Serial Number List they sent with the letter. The firm also asked their customers to complete the online acknowledgement form(link was provided in the letter) if a customer had an ARCTIC SUNÂż 5000 Temperature Management Systems that matches the serial number list. If a customer had further distributed or transferred an ARCTIC SUNÂż 5000 Temperature Management System, Bard Medical advised the customer to please identify the respective organizations and notify them at once of this communication, including a copy of this communication letter and accompanying attachments and to have the organization promptly complete the online acknowledgement form in the link provided. Upon completion of the online acknowledgement form, the customer will be contacted by a customer service representative to schedule the unit(s) for repair.The link to the Online Acknowledgement Form: www.novasyte.com/as5000 Bard Medical stated they partnered with Novasyte to assist in this action. Novasyte contact information is as follows: Phone: (760) 507-1908Email: AS5000@novasyte.com |  | 5608 | 5608 Units Total | Domestic Distribution: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL,GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY.Worldwide Distribution: Canada, Austria, Bahrain, Belgium, Brazil, BrusselsChina, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Gibraltar, Greece, Guam, Hungary, Iran, Italy, Jordan, Korea, Luxemburg, Malaysia, Netherlands, Norway, Oman, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE, UK, Vietnam. | Open, Classified |  |  | 5608
Z-2786-2018 | 80535 | No | CDM | Urinary urobilinogen (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clarity Urocheck 10SG Urinalysis Strips, Model # DTG-10SG |  |  | June 22, 2018 | 2018 | Clarity Diagnostics Llc | Clarity Diagnostics LLC discovered specific lots of Clarity 10SG Strips are not compatible for use in the Clarity Urocheck 120 Urine Analyzers, due to lack of Canister Code Bar Code Label. | Labeling design | The firm notified the major distributor who received this product on or about 06/14/2018 via a letter entitled "URGENT: Voluntary Medical Device Recall for Clarity 10SG Urinalysis Reagent Strips".The distributor letter stated effective immediately, the purpose of this letter is to advise that Clarity Diagnostics LLC is voluntarily recalling Clarity 10SG Strips Lot# URS7100027EX, Exp:11/2019 , URS7100140EX, Exp:11/2019 and URS8010016EX, Exp: 03/2020. The affected lots are not compatible for use in the Clarity Urocheck 120 Urine Analyzers which is why the recall is being performed.The letter also stated to ensure that communication is relayed to Clarity Diagnostics customers. Once the three lots have been secured, Clarity Diagnostics will follow up with the distributor regarding the disposal of the product. Clarity Diagnostics provided a telephone number for any customer concerns. Clarity Diagnostics has taken appropriate steps to assure that this problem will not reoccur in the future.Clarity Diagnostics also provided a letter to the distributor on or about 06/22/2018 entitled "URGENT PRODUCT RECALL". This letter was provided for distributor to notify its customers. This letter provided notification of a product recall regarding specific lots of Clarity Diagnostics - 10SG Urinalysis Reagent Strips. This notice has been issued because the affected lots are not compatible for use in the Clarity Urocheck 120 Urine Analyzers. Affected product first shipped April 5, 2018. Questions can be directed to the Clarity Diagnostics individual handing the recall---Ashish Parikh at Clarity Diagnostics (877) 722-6339. |  | 7536 | 7536 Bottles | Worldwide Distribution - US Nationwide in the states of AZ, MA, SC, WA. and the country of Austria. | Open, Classified |  |  | 7536
Z-2534-2018 | 80401 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PDS II (Polydioxanoe) Suture Dyed &amp; Undyed-PDSII Undyed 45cm (18") 5-0 Single Armed PS-6 Prime Needle |  |  | June 08, 2018 | 2018 | Ethicon, Inc. | Products do not meet a tensile strength specification | Under Investigation by firm | Johnson &amp; Johnson Affiliate Japan notified consignees on 6/08/18 by customer letter through post or email or sales representative visits. Accounts requested to Examine inventory, remove and quarantine such product(s) and complete the Business Reply Form (BRF). Product to requested to be returned. |  | 72 | 72 | Japan | Open, Classified |  |  | 5820
Z-2548-2018 | 80415 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | AperFix Femoral Implant Coring Removal Drill, labeled as the following: |  |  | April 28, 2010 | 2010 | Cayenne Medical Inc. | A review of product complaints in 2010 identified a trend for the Coring Removal Drill of premature wear and/or breakage. This recall occurred in 2013. | Component design/selection | Customers were contacted via email on approximately 10/10/2013 and were instructed to return any affected products on hand. |  | 1444 | 1,444 units total | Worldwide distribution. US nationwide, Denmark, Finland, Malaysia, Netherlands, Poland, Switzerland, and Turkey. | Terminated | February 26, 2019 | 3226 | 1444
Z-2537-2018 | 80510 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Patient Positioning System with KRC2 controller using software versions 2.3.1 2.3.10 |  |  | September 05, 2017 | 2017 | Forte Automation Systems Inc | Communications between the Patient Positioning System and the accuracy filter can periodically fail with no clear indication to the operator. | Software design | On September 5, 2017 Forte has distributed an Urgent Medical Device Correction letter with the following information, Subject: Robotic Treatment Table Positioning Error Commercial Name of Affected Product: Patient Positioning System Affected Version(s) / Lot(s): KRC2 Controllers with software version 2.3.1  2.3.10 Reference / FSCA Identifier: 3003900998-2017-00001Type of Action: Notification and CorrectionThis letter provides users with a Description of the Problem, Details, Recommended User Action and Forte Automation Systems, Inc. Actions. This document also Pending resolution with a permanent solution, Forte Automation Systems, Inc. recommends reimaging the patient prior totreatment, especially after any table movement. This includes movement initiated by the image position verification system. It also advises users that a solution for the issue is in development. |  | 13 | 13 | US in the state of California | Open, Classified |  |  | 13
Z-2540-2018 | 80406 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OMNI Hip System, ApeX-LNK Acetabular Inserts 0 degree, 4+mm offset, Product Codes: H6-14436, H6-13928 &amp; H6-14836 |  |  | May 18, 2018 | 2018 | OMNIlife science Inc. | The Internal packaging, the pouch and patient labels are improperly labeled. | Employee error | On May 18, 2108, all US Agents holding the affected inventory were notified via telephone. On May 23, 2018, OMNILife Science, Inc. issued URGENT MEDICAL DEVICE NOTICES to their customers advising them that the product is being recalled, should be quarantined and returned for repackaging. For questions call 774-226-1847 or 800-448-6664. |  | 16 | 16 units | US Nationwide Distribution - in the states of AZ, CA, CO, IL, FL &amp; MA | Terminated | January 28, 2019 | 255 | 16
Z-2539-2018 | 80280 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | AD-TECH Medical Instrument Corporation Drill Sleeve Guides DSG-6-3-090-2.4N utilizing raw material (RM0677) from the following lots from vendor 00-0001062., AD0021502, AD0021628. |  |  | May 02, 2018 | 2018 | Ad-Tech Medical Instrument Corporation | The diameter of the drill sleeve guide raw material was found to be under tolerance, potentially resulting in the drill bit seizing in the guide during surgery. | Nonconforming Material/Component | Customers were notified on 05/02/2018 via a MEDICAL DEVICE RECALL letter. Instructions included to immediately examine your inventory and quarantine any product subject to recall and replacement. Customers/distributors who received affected devices will be contacted to return their unopened/unused and/or reused stock of the affected product. |  | 65 | 65 Total | US Nationwide distribution in the states of CA, FL, IL, MI, NH, NY, OH, OR, PA, SC, UT, WV. Worldwide distribution to Canada, Denmark, Germany, India, Israel, Japan, Russia, South Korea, Taiwan, United Kingdom. | Terminated | February 08, 2019 | 282 | 65
Z-2538-2018 | 80157 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | BLAST Bandage, sterile. |  |  | April 17, 2018 | 2018 | Tactical Medical Solutions, Inc. | Blast Bandages (part# BLAST) with Lot# 0617 were not terminally sterilized. | Mixed-up of materials/components | Customers were notified on 04/17/2018 via letter and email. Instructions included to review bandages in inventory for affected lot, contact Tactical Medical to arrange replacement, complete and return the Acknowledgement and Receipt Form. |  | 1097 | 1,097 units | Domestic distribution to AL, AP, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MO, MS, NC, ND, NM, NY, OH, PA, SC, TX, VA, WI. International distribution to Austria, Canada, France, Kenya, Latvia, Netherlands, New Zealand, Poland, Singapore, South Africa, Switzerland, Turkey, United Kingdom. | Open, Classified |  |  | 1097
Z-2541-2018 | 80403 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PR4100 Microplate Reader US, Catalog Number 45591587 |  |  | April 02, 2018 | 2018 | Bio-Rad Laboratories Inc | The connector of the power supply was capable of fitting other components provided with the system. However, the voltage output of the reader power supply is incompatible with the other component and, if incorrectly connected, could result in irreparable damage to the system. | Device Design | An Urgent Product Correction notification letter dated 4/2/18 was distributed to customers. The letter instructed customers to check the system and confirm it is labeled as shown in the letter. Then complete the attached form, indicating if the system includes a D-Link Brand USB hub, and fax or email back the form to Bio-Rad as soon as possible. If needed, a Bio-Rad Technical Support Representative will contact you shortly to arrange for replacement of the hub. |  | 153 | 153 | US Nationwide Distribution inn the states of : AL, AR, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OR, PA, SC, TN, TX, VA, VT, WA, WI, and Puerto Rico. | Open, Classified |  |  | 153
Z-2547-2018 | 80380 | No | NDS | Tissue culture media for human ex vivo tissue and cell culture processing applications. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gibco Dulbecco's Modified Eagle Medium (DMEM), high glucose, pyruvate, Model Number 11995123 |  |  | May 10, 2018 | 2018 | Life Technologies, Corp. | It has been determined that the fill port tube seal integrity was compromised during the manufacture of the product lot. This may cause media to leak out of the primary sterile packaging into outer non-sterile dust cover packaging, and subsequently mold may grow in the pooled, leaked media. | Unknown/Undetermined by firm | Urgent Medical Device Recall Notification letters dated 6/28/18 were distributed to customers to notify them of the recall of GIBCO CTS AIM V Serum-Free Media (SFM); Gibco Dulbecco's Modified Eagle Medium (DMEM), high glucose, pyruvate; Gibco Dulbecco's Modified Eagle Medium (DMEM), high glucose; and Gibco RPMI 1640 Medium. The letters instructed customers to do the following:For the above mentioned lots of product, please complete an inspection of the seal integrity of your inventory.There should not be visual evidence of media within the space between the media bag itself and the dust cover. If you are using product that was manufactured between the dates noted above and you have confirmed media between the media bag and dust cover, we ask that you dispose of these bags per local and state requirements and complete the attached Customer Response Form. As noted above, for bags that do not have any visible media beyond the seal, we have confirmed through sterility testing that these units meet sterility specifications and are acceptable for use.Please inform all affected users within your organization of this notification. If you have used this product in manufacturing, you must notify impacted customers and impacted facilities of this notice and our concern for the potential of fill tube seal integrity loss.Please complete the following Customer Response Sheet if you have product that was manufactured between September 2017 and December 2017 and you have confirmed visual evidence of media between the media bag and dust cover. We will work with you to ensure that you receive credit for the discarded material.North American customers: Please complete the following Customer Response Sheet if you have confirmed visual evidence of at the bag-port or media between the media bag and dust cover in your material. Email a scanned copy of the response sheet along with this letter to customer.relations@thermofisher.com. If you have any questions, please ca |  | 70 | 70 | The products were distributed to the following US states: CA, CO, CT, FL, GA, MA, MD, PA, RI, UT, and WA. The products were distributed to the following foreign countries: Canada, France, Great Britain, Japan, Korea, Netherlands, and Switzerland. | Open, Classified |  |  | 1157
Z-2498-2018 | 80402 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper |  |  | June 11, 2018 | 2018 | Philips Healthcare | The first time an operator selects a new procedure type during a single examination,the shutter position resets to the open position for the new procedure type. If the shutters had previously been changed during the examination, that setting is not retained after the first time the procedure type is changed during a single examination. | Software design | An Urgent Medical Device Correction letter dated June 11, 2018 was sent to customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers, until a software revision that corrects this issue becomes available, users should verify that the desired shutter position is set when performing the run after the first time the procedure type is changed during a single examination. This can be accomplished by first selecting a different procedure and then reselect the original procedure on the Xper Module or on the Data Monitor.Customer shall ensure that all staff with access to the affected systems are informed of the contents of this Field Safety Notice.A copy of this Field Safety Notice shall be placed together with the documentation of the system until the system has been corrected by Philips.For questions contact your local Philips representative Technical Support Line 1-800-722-9377. |  | 459 | 459 affected systems | US Nationwide Distribution | Open, Classified |  |  | 459
Z-2555-2018 | 80229 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Leica Microsystems products labeled as the following: Leica M320 Microscope Systems with Article Numbers 10448419, 10448420, 10448485, 10448486 |  |  | June 04, 2018 | 2018 | Leica Microsystems, Inc. | Possible failure of screws which hold the gas spring fixation bracket in place. Due to this failure, the gas spring, which is fixed to the upper part of the parallelogram, cannot hold up the weight of the optics carrier. | Nonconforming Material/Component | Leica Microsystems notified customers on about 06/04/2018 via URGENT FIELD SAFETY NOTICE letter. The letter informed customers that they would be contacted by a Leica Microsystems representative to arrange for the field update in the next 12 months. Customers were instructed that they could continue to use the system pending the update if they ensure a functional check is performed prior to surgery and follow user manual directions. If any anomalies are noticed prior to the correction, customers should contact their Leica representative. Customers were also instructed to provide the notification to all appropriate personnel , notify customers if the unit was transferred, and complete and return the Medical Device Correction Form. A follow-up recall notification was sent to customers on or about 06/07/2018. The follow up letter discusses the functional check prior to use and provides additional compliant information. |  | 8564 | 8,564 Total Units | Worldwide Distribution US Nationwide in the states of : CA, CO, CT, FL, HI, ID, IL, KS, KY, LA, MD, MI, MN, MO, NC, NJ, NY, OH, PA, SD, TN, TX. International Distribution: Canada, AUSTRALIA, BELGIUM, Bolivia, BRAZIL, BULGARIA, CHILE, CHINA, DENMARK, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, HONGKONG, HUNGARY, INDIA, ISRAEL, ITALY, JAPAN, KUWAIT, MALAYSIA, PHILIPPINES, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, TAIWAN,THAILAND, TURKEY, UK | Open, Classified |  |  | 8564
Z-2549-2018 | 80524 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Prevent G Safety Winged Blood Collection Set Glide - Sterile/EO |  |  | September 22, 2017 | 2017 | Guangzhou Improve Medical Instruments Co., Ltd. | Leaks complaints. | Under Investigation by firm | Several e-mails in early 2017 were sent to their only US consignee. |  | 200000 | 200,000 | Kansas | Open, Classified |  |  | 200000
Z-2556-2018 | 80276 | Yes | MUQ |  |  | Cardiovascular | PMA | BioGlue Spreader Tip, 12mm, Sterile, RX Only. |  |  | June 04, 2018 | 2018 | CryoLife, Inc. | The CryoLife L6318 (Label, BioGlue 12mm Spreader Tip, Indicator Label 2 1/8" x 1 1/4") was incorrectly applied to the Syringe Applicator Tip pouches and Syringe Spreader Tip pouches were not included in the boxes. | Error in labeling | CyroLife notified their customers on/about 06/04/2018 via Urgent - Important Product Removal letter. instructions indicated that affected pouches have been returned to CryoLife, Inc. Any questions/clarifications regarding this removal are directed to contact 800-438-8285 or by email to fieldassurance@c ryolife.com. |  | 855 | 855 Total | Distribution is to Japan | Terminated | February 08, 2019 | 249 | 855
Z-2557-2018 | 80376 | No | JOY | Automated cell-locating device. | Hematology | Hematology | 510(k) | cobas m 511 integrated hematology analyzer, Model Number 07261691190 |  |  | May 07, 2018 | 2018 | Roche Diagnostics Hematology | Discrepant results have been reported in patients with severe microcytic anemia (e.g.iron deficiency, thalassemia) and the following parameters are affected: RBC (red blood cell count), HGB (hemoglobin concentration), MCH (mean corpuscular hemoglobin), HCT (hematocrit), and MCV (mean corpuscular volume). | Software design | An Urgent Field Safety Notice was distributed to customers via email on 5/7/18. The notification included the following actions for to be taken by the customer/user: Until this rule is implemented, when cobas m 511 integrated hematology analyzer displays a HGB value of less than or equal to 9g/dL, Roche advises the user to not report results for the RBC, HGB, MCH, HCT, and MCV parameters. Roche advises the user to perform laboratory confirmatory testing before making transfusion decisions. Delta check against previous results might additionally be used to evaluate the results.For customer evaluating the instrument Roche kindly advises to not report values for diagnostic use until the rule is configured by your Roche Application Specialist. |  | 17 | 17 | No distribution in the United States.The devices were distributed to the following foreign countries: Austria, France, Germany, Hong Kong, Japan, Netherlands, Pakistan, Singapore, South Korea, Switzerland, Thailand, and Vietnam. | Open, Classified |  |  | 17
Z-2551-2018 | 79855 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | ImplantDirect, Dental implants intended for use in partially and fully edentulous upper and lower jaws in support of single or multiple-unit restorations and terminal or intermediate abutment support for fixed bridgework. |  |  | March 03, 2017 | 2017 | Implant Direct Sybron Manufacturing, LLC | The main vial label lists the incorrect part number, but the cap label is correct. | Process control | Recall notifications dated 2/13/2017 were issued to U.S. consignees via UPS overnight on 3/3/2017. Foreign consignees received letters on 3/13/2017 and 3/14/2017. |  | 124 | 124 units | Distribution was made to AZ, CA, CO, CT, FL, ID, IL, NC, NJ, NV, NY, PA, TX, UT, VA, and WA. There was no government/military distribution.Foreign distribution was made to Switzerland, Germany, Lithuania, and Spain. | Terminated | August 06, 2018 | 521 | 124
Z-2554-2018 | 80515 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Radiation Therapy Treatment Planning System, Model nos. 2.5, 3.5, 4.0, 4.3, 4.5, 4.7, 4.9, 5.0, 6.0, 6.1, 6.2, 7.0 |  |  | June 06, 2018 | 2018 | RAYSEARCH LABORATORIES AB | The firm has learned that some RayStation/RayPlan users have commissioned machines with erroneous Beam profile correction parameters. These parametersaffect the dose calculated in corners of large or off-axis fields. This effect cannot be seen in the Beam Commissioning module and dose in large or off-axis fields needs to be validated using the Beam 3D Modeling module in RayPhysics/RayPlan Physics. The user mustbe aware to avoid incorrect dose calculations during treatment planning. | Software design | On June 6, 2018, the firm issued an Urgent Field Safety Notice, Medical Device Correction letter to affected customers via email. The notice was a reminder to follow the instructions listed in the product documentation when creating aphoton beam model in RayStation/RayPlan.ACTIONS TO BE TAKEN BY THE USER" Always validate the beam model for all clinically relevant fields, including, but not limited to, largeand off-axis fields, before the system is used to create clinical treatment plans." Carefully review all beam model parameters before commissioning." Be aware that the Beam profile correction values may need to be adjusted in order to correctlymodel the primary collimator that limits the dose in the corners of large or off-axis fields." Review any existing photon beam models to ensure that the actions above have been properlyperformed.Please educate physics staff about these user responsibilities.Inspect your product and identify all installed units with the above software version number(s), then confirmthat you have read and understood this notice by replying to the notification email.Safe use relies on careful validation of the beam model for all clinically relevant fields before the systemis used to create clinical treatment plans. This information is already present in the Instructions For Use,the Reference Manual and the RayPhysics/RayPlan Physics Manual for the existing product versions. Tofurther reduce the risk of using improper beam models, the safety warnings within the product documentationwill be updated in the next version of RayStation/RayPlan, scheduled for market release in June 2018. |  | 718 | 718 | US Nationwide distribution. | Open, Classified |  |  | 718
Z-2550-2018 | 80283 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Tubing Pack Convenience Kit, Product Code 65217 |  |  | May 12, 2017 | 2017 | Terumo Cardiovascular Systems Corp | Kit was labeled with the incorrect Expiration Date of April 30, 2019. | Incorrect or no expiration date | On May 12, 2017 Terumo Cardiovascular Systems Corporation telephoned their one client and notified them of the labeling error and asked them to hold product. All product was immediately quarantined. |  | 60 | 30 cases/2 units (Total 60 units) | US state of FL | Terminated | October 26, 2018 | 532 | 60
Z-2564-2018 | 80021 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TVS LPA, Catalog #: DIL-059-00 |  |  | May 02, 2018 | 2018 | Galt Medical Corporation | The products listed might contain unsafe levels of bacterial endotoxins (Pyrogens) that were introduced during a manufacturing step. Bacterial Endotoxins also called pyrogenic bacteria can activate the inflammatory process and produce fever, chills, and hypotension in a patient. | Other | On approximately 05/01/2018 recall letters and customer response forms were mailed to all consignees. |  | 3530 | 3530 | U.S.: MN, PA, VA, CA, UT, FL, NY, GA, LA, MO, TX, MS, OH, WA, MA, IL, KY, NC, and Puerto Rico;OUS: Canada, Australia, France, Germany, Ireland, Italy, Netherlands, Poland, Slovakia, Switzerland, Belgium, United Kingdom, Korea | Open, Classified |  |  | 72327
Z-2574-2018 | 80441 | No | GNP | Epstein-Barr virus serological reagents. | Microbiology | Microbiology | 510(k) | Diamedix Is-EBV-EA-D IgG Enzyme Immunoassay Test Kit for In Vitro Diagnostic Use Catalog No. 720-640 |  |  | January 05, 2018 | 2018 | Diamedix Corporation | Contains a gel-like contaminant that may affect product performance. | Under Investigation by firm | Diamedix sent a letter dated January 5th, 2018 was sent to all their customer with the following: Products contains a gel-like contaminant that may affect product performance.An inspection of the product in the Diamedix inventory confirmed this observation. Although, no failed runs have been directly attributedthis issue, we feel that the best course of action is to replace the affected diluent. If you are using the reagent lot specified above, please discontinue use immediately. In order to process your replacement, please complete the attached Certificate of Destruction and return this document to Diamedix Customer Service via email: CustomerService@erbadiagnostics.com or fax to (305) 418- 2321. If you have further questions or concerns you may contact our Technical Service Department at (305) 324-2307 or (800) 433-0945.Please be assured that Diamedix Corporation is working diligently to resolve the issues with this product. We apologize for any inconvenience that this may have caused and we thank you for your continued patience and support.Sincerely,Tracy Chadwrick Director, QA/QC |  | 147 | 147 | AZ, CA, FL, IL, IN, KY, MA, MI, MN, MO, NH, NJ, NY, OH, OK, PA, RI, TN, TX, and WI | Completed |  |  | 652
Z-2572-2018 | 80424 | No | LYJ |  |  | Neurology | PMA | VNS Therapy Programming System, Rx Only, Model 3000, v1.0.2.2 |  |  | January 24, 2018 | 2018 | LivaNova USA Inc | Unintended warning message displayed on generators programmed with a Model 3000 v.1.0.2.2 programmer. | Software design (manufacturing process) | Consignees were notified on about 01/24/2018 via URGENT: MEDICAL DEVICE CORRECTION letter. Instructions for physicians with new devices (interrogated for the first time) include to 1) apply the changes and proceed with programming as usual. The warning message requires a mandatory action to Apply Changes . After applying changes, therapy will automatically be programmed to 0 mA. The device may be immediately programmed back on to re-enable therapy. 2) To complete and return the attached Customer Response Form (by fax to (281) 853-1248 or by email to M3000FieldAction@livanova.com. Instructions to physicians with devices that have been implanted (previously interrogated) include 1) Restore the normal and magnet mode stimulation parameters from the Summary screen immediately after the warning message screen. If desired, additional programming changes may be made. 2) For AutoStim (M106 only), detection must be re-enabled first, prior to re-enablingAutoStim therapy. The previous AutoStim parameters are not retained and must be reentered manually. If needed, the previous parameters may be found by ending the session and viewing the Session Report for the generator. 3) Please complete and return the attached Customer Response Form by fax to (281) 853-1248 or by email to M3000FieldAction@livanova.com. Additionally customers were instructed to ensure that this notice is communicated to all relevant personnel within the organization. For questions regarding the information in the letter, please contact Clinical Technical Supportat (866) 882-8804 (Monday to Friday, 8 AM to 5 PM CST) or e-mail at cservices@livanova.com. |  | 243 | 243 devices total | Distributed domestically to AZ, CA, GA, IA, IL, IN, MO, MT, NC, NY, PA, SD, TN, TX, WA, WI. | Open, Classified |  |  | 243
Z-2571-2018 | 80420 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Drill (Fluted-Hard Bone), for Quattro Link 2.9mm, labeled as: |  |  | October 10, 2013 | 2013 | Cayenne Medical Inc. | A review of complaints identified a trend for drills breaking. | Device Design | Customers were contacted on approximately 10/10/2013 and were instructed to return any affected products on hand. |  | 229 | 229 devies total | US distribution to AK, AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI. | Terminated | February 26, 2019 | 1965 | 229
Z-2569-2018 | 80353 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | FlowGate2 Balloon Guide Catheter, UPN 90485 (8Fx85cm size) and 90495 (8Fx95cm size) |  |  | April 11, 2018 | 2018 | Stryker Neurovascular | There may be product manufactured with the outer diameter of the balloon catheter exceeding the defined specification at a customer location. If the user experiences resistance between the FlowGate and an interfacing device, the device would be removed and a second device would be used, which could potentially prolong a procedure. | Employee error | The recalling firm sent a notification letter on 4/11/2018 with instructions to return a response form. |  | 968 | 968 units | Nationwide | Open, Classified |  |  | 968
Z-2568-2018 | 80023 | No | OLV | Electroencephalograph. | Neurology | Neurology | 510(k) | NicoletÂż Ambulatory EEG |  |  | February 01, 2018 | 2018 | Natus Neurology Inc | Potential to be able to import an ambulatory exam for a patient for which the exam was not started. | Software design | Customers were notified via letter on about 02/01/2018. Instructions to ensure the empty exam is removed appropriately after starting the ambulatory recording included: 1) Only shutdown Recorder via the Recorder UI, either from the File menu or the X button on the top right of the application; 2) If any of the four conditions above occurs and the 2.60 MB exam is not removed, the emptyexam can be deleted from NicVue prior to importing the ambulatory exam, doing so will allow the patient ID check to occur normally. Customers were also instructed to provide the recall communication to all applicable users and complete and return the verification form. |  | 171 | 171 units | Distributed domestically to AR, AZ, CA, CO, DC, FL, GA, ID, IL, IN, MA, ME, MI, MO, N Y, NC, ND, NJ, NV, NY, PA, TN, TX, VA, WA, WI. Distributed internationally to Australia, Belgium, China, Germany, Hong Kong, Israel, Italy, Jordan, Kuwait, Russia, Slovenia, Tunisia, United Arab Emirates, Vietnam. | Open, Classified |  |  | 171
Z-2567-2018 | 79485 | No | JDO | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Custom Made Implant System with Minimum Invasive Grower (MIG) component |  |  | December 21, 2017 | 2017 | Howmedica Osteonics Corp. | From March 2007 to October 2014, product did not contain updated instructions for use (IFU) to clarify the function of locking mechanism and warnings related to the locking screw. | Labeling design | On December 21, 2017 a FIELD SAFETY NOTICE was issued to customers notifying of the corrective action. An attachment was supplied with the letter titled "Minimally Invasive Grower - Instructions For Adjustment that is effective on and after October 14, 2014. Questions or concerns can be directed to your local sales representative. |  | 5 | 5 | NY, FL, KS | Terminated | October 05, 2018 | 288 | 5
Z-2570-2018 | 80239 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | XLTEK EMU40EX EEG Headbox |  |  | May 15, 2018 | 2018 | Natus Neurology DBA Excel Tech., Ltd. (XLTEK) | Possible degradation of an internal electronic component over time which may cause the XLTEK EMU 40EX Breakout box to fail. It may feel hot and plastic case may start melting. | Other | Natus Neurology notified customers on or about 05/15/2018, via Advisory Notification Letter to stop using the device immediately and contact Natus regarding replacement options. Customers were advised that if the malfunction occurs, to immediately cease using the device and contact Natus for replacement. Customers were also instructed to communicate the notification to all applicable users and to complete and return the advisory verification form. |  | 462 | 462 Units Total | USA Distribution: AL, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MIMN, MO, NC, NE, NJ, NV, NY, OH, OR, PA, ,RI, SC, TN, TX, UT, VA, WA, WI.World wide Distribution: Canada, China, Cyprus, Germany, Hong Kong, Italy, Mexico, Saudi Arabia, Spain, United Kingdom | Open, Classified |  |  | 462
Z-2576-2018 | 80378 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Percutaneous Sheath Introducer Kit for use with 7 - 7.5 Fr. Catheters (8 Fr. 10 cm sheath length .035 inch dia. spring-wire guide), REF ES-09807. The percutaneous sheath introducer permits venous access and catheter introduction to the central circulation. |  |  | May 23, 2018 | 2018 | Arrow International Inc | Product contains dry natural rubber latex. Label states Latex Free. | Process control | On May 23, 2018, Arrow International issued Urgent Medical Device Recall notices and response forms to their customer. Customers are advised to take the following actions:1. Immediately discontinue distribution and quarantine affected product.2. Using the provided customer letter and Recall Acknowledgement Form templates, customers who have further distributed product should contact those individuals.3. To return affected products from your inventory, complete and return the Recall Acknowledgement Form via to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. 4. If you and your customers have no affected stock, please complete and return the Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com.5. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-396-2111. |  | 60 | 60 units | Puerto Rico | Open, Classified |  |  | 60
Z-2575-2018 | 80475 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid Cardiac Control - Level 3, cat. no. CQ5053 |  |  | June 08, 2018 | 2018 | Randox Laboratories Ltd. | The device is not suitable for the control of the Troponin T assay due to unacceptable variation between vials. The deviation in precision observed is up to 28%. IQC that is reported as out of range could lead to a delay in reporting Troponin T results. If the failure occurs the operator may observe the quality control result for Troponin T falling outside range and/or increased imprecision for Quality Control monitoring. | Under Investigation by firm | On June 8, 2018, the firm issued an Urgent Medical Device Recall letter to affected consignees. The letter stated that Randox has confirmed Liquid Cardiac Control CQ5053 Lot: 4245CK is not suitable for the control ofthe Troponin T assay due to unacceptable variation between vials. The deviation in precision observed is up to 28%. Customers were instructed to do the following:" Review results generated with the affected batches in line with the clinical profile of the patient." Discuss the contents of this notice with your Medical Director." Update kits with corrected IFU's excluding Troponin T values." Complete and return the vigilance response section of this form totechnical.services@randox.com within five working days.)In addition, distributors were instructed to send a copy of the letter to all affected customers and to those who need to be aware within their organization. |  | 1 | 1 kit (US), 602 kits (OUS) | Worldwide Distribution - US in the state of WV. Foreign distribution to Australia, Chile, China, Germany, Hungary, Ireland, Italy, Moldova, Netherlands, Peru, Poland, Saudi Arabia, Slovakia, Spain, Sultanate of Oman, UAE, UK, and Vietnam | Terminated | November 05, 2018 | 150 | 1
Z-2588-2018 | 80087 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VIDAS LH, Reference Numbers 30406 and 30406-01. |  |  | March 01, 2018 | 2018 | bioMerieux, Inc. | Several complaints were received for invalid calibration with low calibrator S1 while using VIDAS LH. Internal tests confirmed a decrease of the strips signal overtime for specific lots. | Nonconforming Material/Component | bioMerieux notified customers on about 03/01/2018, via letter. Instructions included to distribute the recall information to all appropriate personnel in the laboratory, forward the information to individuals to which the product many have been transferred to, destroy any impacted lots on site, and complete and return the Acknowledgement Form. Questions and concerns can be directed to local bioMerieux Customer Service representatives or for Clinical Technical Support Center by email at CustomerService-ImmunoMolecular@biomerieux.com. |  | 5429 | 5,429 units | Worldwide Distribution - US Nationwide in the state of Arizona and North Carolina. Distributed internationally to Algeria, Algeria, Angola, Argentina, Bahrain, Bangladesh, Belarus, Bosnia-Herz., Brazil, Bulgaria, Burkina Faso, Cameroon, Canada, Chad, Colombia, Congo, Costa Rica, Cote d'Ivoire, Czech Republic, Djibouti, Dutch Antilles, Ecuador, Egypt, Estonia, Ethiopia, France, Gabon, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Macedonia, Mali, Mauretania, Mexico, Morocco, Mozambique, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Palestine, Panama, Paraguay, Philippines, Poland, Portugal, Republic of Congo, Romania, Russia, Saudi Arabia, Senegal, Serbia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Syria, Tanzania, Thailand, Tunisia, Turkey, Turkmenistan, Ukraine, Uruguay, United Arab Emirates. | Open, Classified |  |  | 5429
Z-2580-2018 | 80484 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery MI, labeled as Discovery DMI 15 cm 3 ring, and Discovery DMI 20cm 4 ring |  |  | May 24, 2018 | 2018 | GE Healthcare, LLC | Some CT and CT/PET systems recently installed with a GE supplied A! electrical panel may not be properly wired to the partial system UPS (Uninteruptible Power Supply). | Component change control | GE Healthcare notified customers on about 05/24/2018 via URGENT MEDICAL DEVICE CORRECTION letter. The letter outlined the reason for recall and notified customers that a GE Healthcare representative would be contacting them to arrange for the correction of the devices. Any questions or concerns can be directed to GE Healthcare Service at 1-800-437-1171 or a local Service Representative. |  | 42 | 42 units total | Nationwide distribution to AL, CA, CO, FL, IL, IN, MA, MD, MI, MS, NC, NJ, NY,OH, OK, SC, TN, TX, VA. | Terminated | February 08, 2019 | 260 | 42
Z-2591-2018 | 80321 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ENVOY 500 ISE CALIBRATOR KIT, reference 55117 (contained 6 x 20 mL Calibrator High level and 6 x |  |  | April 18, 2017 | 2017 | ELITech Clinical Systems SAS | ELITech Clinical Systems SAS initiated this recall because some users of ENVOY 500 ISE CALIBRATOR Kit (Part number: 55117), for Envoy 500 systems are observing trouble while calibrating with some vials. Users contacted the firm when they received the messages "ISE OUT OF REPRODUCIBILITY" or "ISE SLOPE OUT OF RANGE" (i.e., failed calibration). This has resulted in delay in obtaining patient results until the calibrator lot is replaced. | Under Investigation by firm | ELITech Clinical Systems SAS notified consignees on April 18, 2018 via USPS certified letter, dated April 10, 2018 titled Product Recall Notice. The Product Recall Notice provided customers/consignees reason for recall, actions to be taken, and included a confirmation of receipt of product recall notice form to be filled out and returned by customer/consignee. |  |  |  | U.S. only. No gov. accounts. No international. | Open, Classified
Z-2590-2018 | 80489 | No | OAF |  |  | Ear Nose &amp; Throat | PMA | Final Pack, Esteem II, Sound Processor Model 2001, containing 2 Sterile Wipes 002866-001 (which expired in April 2018). An Implantable Middle Ear Hearing Device. |  |  | June 28, 2018 | 2018 | Envoy Medical Corporation | Sterile Wipes packed in these four Esteem II Sound Processors had an earlier expiration date than the final pack and the Esteem II Sound Processors. The Sterile Wipes expired prior to the date labeled on the final pack. | Packaging process control | The firm notified its customers via email on 06/18/2018 and 06/19/2018. The 06/18/2018 email to a customer stated Envoy Medical would like to have the devices the customer has on site returned, as apparently, it was confirmed that each device is packaged with a wipe that shows expired just on the wipe only. Envoy will replace the recalled devices by overnight shipping replacements.The 06/19/2018 email to a customer stated it was discovered that the sound processors (primary and backup) were previously sent to the site have a wipe in them that show an expired date on the packaging. The email also stated this is something that doesn't affect the actual sound processor itself as those are good, but just to be on the safe side, make sure the 2 sound processors currently at the site (serial #s 213658 and 213729) are returned to Envoy Medical ASAP so we can have them repackaged with wipes within their expiration dates added to the sound processor packaging. Envoy will send out 2 new Sound Processors (primary and backup). Both emails contained Envoy Medical contact information for the consignees in case any questions or concerns arose. |  | 4 | 4 Units | US Distribution to states of: FL, OR | Open, Classified |  |  | 4
Z-2594-2018 | 80421 | No | GDJ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MAYFIELD INFINITY XR2 LOW PROFILE BASE UNIT, Part Number A2079A |  |  | June 29, 2018 | 2018 | Integra LifeSciences Corp. | There is a potential for the knob assembly in the swivel adapter to fracture/break during use. | Device Design | The firm, Integra, sent an "Urgent Voluntary Medical Device Correction Field Safety Notice" dated 6/29/18 to customers. The notification letter instructed customers to perform the following: 1) Review your MAYFIELD XR2 instructions for use. Please pay particular attention to the CAUTIONS listed on page 5 of the accompanying instructions for use section as noted below." Always make sure the base unit is properly secured to the operating table" The base unit must not be used if the device appears to be damaged or functioning incorrectly" Over-tightening the mechanisms adjustment screws may result in damage to the unit" Over extending or overloading the base unit may result in unintended movement, shortened product life and/or damage to the unit" Do not alter the construction of this device as it may result in serious patient injury" Do not steam sterilize; components may be damaged by high heat leading to device damage and reduced performance2) Visit Integra website and review video highlighting instructions for use at the following link: https://www.integralife.com/mayfield-imaging-cranial-stabilization-system/category/cranial-stabilization-mayfield-imaging-cranial-stabilization-system3) Complete and return the attached Field Safety Notice Acknowledgement Form promptly via email: FCA@integralife.com or Fax to: 1-609-750-4220.Distributors were also instructed to: "We also kindly ask you to contact the final customers who may have the affected products and provide them with this letter."On July 20, 2018, an additional Urgent Voluntary Medical Device Recall letter was distributed to customers. The letters ask customers to perform the following:1. Allow your sales representative to check the current knob assembly on your MAYFIELD Infinity XR2 Radiolucent Base Unit (A2079), Standard; MAYFIELD Infinity XR2 Low Profile Base Unit (A2079A); MAYFIELD Infinity XR2 Radiolucent Base Unit, Extended (A2079E) and MAYFIELD Spine Table Adaptor ( |  | 1 | 1 units | Worldwide Distribution: US (nationwide) to states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, and WV; and to countries of:: Australia, Canada, China, Dominican Republic, Hong Kong, India, Japan, Korea, Malaysia, Mexico, Pakistan, Panama, Peru, Singapore, Sri Lanka, Taiwan, and Thailand. | Open, Classified |  |  | 1044
Z-2388-2018 | 80177 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Discovery IGS 740 Interventional Fluoroscopic X-ray Systems |  |  | March 29, 2018 | 2018 | GE Medical Systems, SCS | The Discovery IGS and Innova IGS systems may experience X-ray abort errors which could result in the loss of imaging mode before or during fluoroscopic use. | Process control | GE Healthcare notified customers on approximately 03/29/2018, via an URGENT MEDICAL DEVICE CORRECTION letter. Safety instructions included that before each use, to ensure the IGS system is fully functional as indicated in the GE product labeling and follow each facility's established procedures before and during each use to manage the patient if a repeated X-ray abort error occurs during the procedure or in the case a complete loss of fluoroscopic viewing capability may occur. If the problem persists, customers are instructed to contact their local GE Healthcare Service Representative. Additionally, customers are instructed that GE Healthcare will correct all affected products at no cost and a GE Healthcare representative will contact each customer to arrange for the correction. Questions or concerns regarding the notification can be directed to GEHealthcare Service at 1-800-437-1171 or the local Service Representative. |  | 102 | 102 devices total | US distribution: AR, CA, CA, FL, GA, IA, IL, NC, NH, NJ, NM, NY, PA, SC, TN, A,WI, WV. International Distribution: Argentina, Brazil, China, Colombia, Egypt, France, India, Indonesia, Iran, Italy, Japan, Kuwait, Lithuania, Myanmar, Palestine, Poland, Russia, Saudi Arabia, Singapore, Spain, Tunisia, Turkey, United Kingdom, Vietnam. | Open, Classified |  |  | 102
Z-2566-2018 | 80570 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Unity software, formerly known as DR Systems Unity PACS software. The firm name on the label is Merge Healthcare, Hartland, WI. Merge Unity is a medical image and information management system that allows viewing, selection, processing, printing, telecommunications, and media interchange of medical images from a variety of diagnostic imaging systems. |  |  | August 07, 2017 | 2017 | Merge Healthcare, Inc. | The software is not identifying the patient as having atypical hyperplasia, resulting in an incorrect Gail Risk calculation. | Process design | MERGE sent a Medical Device Recall letter on about 08/07/2017, to all affected customers. Customers were advised to discontinue using the affected product and to reply using the enclosed form. Customers with questions were instructed to send an email to recall@merge.com. For questions regarding this recall call 262-912-3514. |  | 11 | 11 sites have potentially affected software | Nationwide distribution to CA, KY, ME, OH, OR, TN, TX, WA, WV. | Open, Classified |  |  | 11
Z-2296-2018 | 80059 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LFIT Anatomic CoCr Femoral Heads, Catalog No. 6260-9-140 |  |  | May 22, 2018 | 2018 | Howmedica Osteonics Corp. | The recalling firm has received a higher than expected number of complaints documenting femoral head/hip stem disassociation for certain sizes of femoral heads manufactured prior to March 4, 2011. | Process control | On May 22, 2018, Stryker issued a Product Safety Notification to its affected customers. The letter explained the issue and instructed customers to do the following:Follow-upPatients implanted with the Products should continue to be followed per the normal protocol established by his/hersurgeon(s). There are no recommended changes to the frequency of the standard follow-up care protocol.Required actions1. Hospitals/Surgeons: Please inform product users of this notification and forward this notice to all those individualswho need to be aware within your organization. Please also complete and sign the enclosed Business Reply Formand fax a copy to 888-861-8582 or email to Stericycle at strykerortho5940@stencycle com. Product is not required tobe returned as part of this notification.if you have any questions, feel free to contact the firm at (201)831-5151. |  |  |  | Distributed nationwide. Foreign distribution to Australia, Canada, Chile, China, Brazil, France, Germany, India, Italy, Japan, Malaysia, Netherlands, New Zealand, Poland, Portugal, Romania, Singapore, South Africa, Spain, Sweden, Switzerland, and United Kingdom. | Open, Classified
Z-2521-2018 | 80240 | No | KWL | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CORAIL Cementless Femoral Stem HA Coated Standard Size 11, Catalog No. 3L92511 |  |  | June 21, 2018 | 2018 | DePuy Orthopaedics, Inc. | Two lots of femoral stems were labeled as Size 11. However, they are actually Size 12. Potential adverse effects include poor joint mechanics, improper fit, loosening, and additional surgical procedure to correct. It may also result in an intraoperative fracture which could represent a surgical delay and additional procedures. | Process control | On June 21, 2018, the firm issued an Urgent Information: US Recall Notice via email to affected customers. Medical professionals were asked to do the following: " Cease using the affected components immediately." Immediately return unused affected units: Medical facilities are to determine if any of the recalled components are still on hand, and return affected devices immediately to their DePuy Sales Consultant or return them to DePuy for credit following normal purchasing procedures. Reconciliation forms should be faxed to 574-371-4939." Reconciliation Form: Complete the Reconciliation Form and return to your DePuy Synthes Sales Consultant or fax to 574-371-4939 within five (5) days of this notice." Records: Retain a copy of the completed Reconciliation Form in your files along with this notice." Additional Notifications: Notify surgeon users at your facility by providing them with a copy of this notice to ensure surgeon users are aware of this recall notice. Forward this notice to others in your facility that need to be informed." Maintain a copy of this notice with the affected devices.For questions about device recall information provided, please contact Kim Long, Senior Recall Coordinator, at 574-371-4917 or klong16@its.jnj.com.U.S. Distributors and Sales Offices received a separate notice with specific instructions on how to identify, remove, and return the affected devices. |  |  |  | Accounts in AL, CA, CO, GA&lt; IL, LA, ME, NC, NH, NY, OR, TX, UT, and WA. Foreign distribution to Belgium, Croatia, Cyprus, Denmark, Germany, Indonesia, Korea, Sweden, and Tunisia. | Open, Classified
Z-2600-2018 | 80445 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Instrument trays used to store AltiVate Small Shell Trials (FA S ALTIVATE RSP HUM TRLS SM). |  |  | June 26, 2018 | 2018 | Encore Medical, Lp | A black, aluminum oxide coating may appear on the tray after cleaning. The black, aluminum oxide coating is not biocompatible. | Process control | The firm, DJO LLC, sent an "URGENT FIELD SAFETY NOTICE" letter dated 6/26/2018 to its direct accounts by email on 06/26/2018. The letter described the product, problem and actions to be taken. The customers were instructed to return product; contact Customer Service as soon as possible to request an RMA for return and exchange of the full instrument set; and to coordinate the exchange of products with Customer Service.If you do not have these instrument sets or the number of sets identified above, you are required to complete the attached form and email it to sandra.busick@djdoglobal.com.ALL ACTIVITIES OUTLINED ABOVE SHOULD BE COMPLETED BY JULY 30, 2018.If you have any questions, please call (512) 834-6255 or by email to teffany.hutto@djoglobal.com. |  | 30 | 30 units | US Distribution to states of; AZ, CA, CO, FL, IN, NJ, NY, NV, MI, MN, MO, OH, RI, SC, TN, TX, UT and WA and country of; Japan. | Completed |  |  | 30
Z-2597-2018 | 80522 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | power cord for: (a) AMIA Automated PD Cycler, Product Code: 5C9320; (b) Kaguya PD System, Product Code: T5C8500 (Japanese distribution only) |  |  | July 12, 2018 | 2018 | Baxter Healthcare Corporation | There were instances where the power cord socket detached from the back of the AMIA cycler when the power cord was unplugged from the socket. | Under Investigation by firm | Method of Notification: A Safety Alert communication will be sent to affected customers via U.S.P.S., first class mail.Mitigating Instructions:Baxter is asking customers to:1.Unplug the devices from the wall before unplugging the power cord from the back of the cycler.2.Confirm the receipt of the letter by completing the enclosed customer reply form and returning it to Baxter by either mail, fax, or scanned email. |  | 5445 | 5445 units | US and Japan | Open, Classified |  |  | 5445
Z-2599-2018 | 80330 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Pos Combo Panel Type 43, Catalog number B1017-217for use in determining antimicrobial agent susceptibility and/or identification to the species level of rapidly growing aerobic and facultative gram-positive cocci, some fastidious aerobic gram-positive cocci and Listeria monocytogenes. |  |  | May 04, 2018 | 2018 | Beckman Coulter Inc. | The drug dilution sequence for Moxifloxacin printed on the panel is incorrect. As a result, incorrect MIC results for Mxf could be reported. | Labeling design | A recall notification letter was sent the week of 05/01/2018 via email, mail, and phone if needed. |  | 686 | 686 boxes | NJ, TX, IA, PA, FL, NE, OH, NY, CA, MO, OK, LA | Open, Classified |  |  | 686
Z-2596-2018 | 80430 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EZ1 DSP Virus Kit (48), IVD, REF 62724 intended for in vitro diagnostic use. |  |  | May 28, 2018 | 2018 | QIAGEN | Underfilled reagent well that could lead to incorrect results | Under Investigation by firm | On May 28, 2018 QIAGEN issued an Urgent Notice of Product Field Action and a Acknowledgement of Receipt form to customers. Actions to be taken by the customer/user/commercial partner are as follows:- If you have any remaining stocks, do not use them, contact our technical service for a free-of-charge replacement and dispose of the product in accordance with your national, local safety and environmental regulations. - Review the notice with your laboratory/medical director.- Forward the information to all individuals and departments within your organization. - If product was further distributed, forward a copy of the notice tothe product end user.- QIAGEN also recommends a review of the corresponding results. Consider the possibility of under-quantification and false negative results and evaluate the need for retesting with different EZ1 DSP Virus Kit lots e.g. depending on the clinical presentation and history of the patient- Complete and return the Acknowledgement of Receipt form and email to: quality.communications@qiagen.com. |  | 287 | 287 pcs | Domestic - MDInternationally - AE, AU , CZ, DE, DZ, ES, FR, GB, HK , IE, IT, NO, RO &amp; SA | Open, Classified |  |  | 287
Z-2602-2018 | 80466 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Dermatone AN, Model No. 88710100 |  |  | June 08, 2018 | 2018 | Zimmer Surgical Inc | Devices may have a loose control bar, which could compromise the ability to control the thickness of the graft. Devices could also have a loose width plate that if undetected, could result in an imperfect, yet still usable graft. Potential risks include harvesting a skin graft of a thickness much greater than intended, described as patient gouged, minor delay of surgery, and sub-optimal (but usable) graft. | Under Investigation by firm | On June 8, 2018, the firm distributed an Urgent Medical Device Recall letter to affected consignees. The letter alerted customers to the issue. Customers were asked to do the following:1. Review this notification and ensure that affected personnel are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your facility.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.c. Request an RGA number for each unit by calling 800-830-0970 or emailing dover.repairs@zimmerbiomet.com.d. Include a copy of Attachment 3  Certificate of Sterilization with returned instruments.e. Mark RECALL and RGA Number on the outside of the returned cartons.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5. If you have further questions or concerns after reviewing this notice, please call customer service at 800-830-0970 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 140 | 140 | Distributors in AL, AZ, CA, FL, IA, IL, MA, MI, MN, MS, NH, NJ, NY, OH, RI, SD, and TX. Foreign distribution to Canada, Australia, China, and Netherlands. | Open, Classified |  |  | 140
Z-2601-2018 | 80542 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Polysorb Braided Absorbable Suture 1 36" 90 cm |  |  | June 15, 2018 | 2018 | Covidien LLC | Potential proplylene glycol contamination | Nonconforming Material/Component | All U.S. consignees were notified via Federal Express or certified mail on June 15, 2018, and the letter informs customers of the potential that the suture was contaminated with propylene glycol during the manufacturing process and the actions they should take. Customers are advised to discontinue use of the product and return all inventory of the affected lot to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. |  | 972 | 972 units | CA, NY &amp; TX | Open, Classified |  |  | 972
Z-2603-2018 | 80541 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Medtronic CD HORIZON Spinal System is intended to help provide immobilization and stabilization of spinal segments as an adjunct to fusion of the thoracic, lumbar, or sacral spine.MULTI-AXIAL SCREW FOR 5.5/6.0MM RODS, REF 55840005045 |  |  | July 09, 2018 | 2018 | Medtronic Sofamor Danek USA Inc | This lot has incorrectly been laser marked both on the screw head itself as well as the data carrier tag attached to the screw. The description on face two of the screw head and on the data carrier tag both incorrectly indicate a 5.5 mm diameter screw when the actual screw diameter aligns with 5.0 diameter screw size requirements. | Employee error | Medtronic distributed a voluntary removal notice via FedEx courier service to impacted U.S. consignees, delivered 09 July 2018.The notice stated provided a list of Loaner kits where the screws may also be found and provided the following instructions:" Instructions to Customers:1.) Please locate and remove the impacted product from normal storage locations. Do not use this product. 2.) Complete the enclosed Customer Confirmation Form and return via email to rs.sdmriskmgt@medtronic.com.3.) As applicable, contact your Medtronic Sales Representative to coordinate the return and replacement of affected product.Transmission of this Field Safety Notice:1.) This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred.2.) Please transfer this notice to other organizations on which this action has an impact.3.) Please maintain a copy of this notice in your records." |  | 47 | 47 units | US | Open, Classified |  |  | 47
Z-2604-2018 | 80499 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Modera Modular Pedicle Screw System Shank Insertion Instrument, 3-107-07 |  |  | February 13, 2017 | 2017 | New Era Orthopaedics, LLc | This is a retrospective recall of an instrument used in spine fusion procedures with Modera Modular Pedicle Screw system implants. The firm received a customer complaint in which excessive forces were applied to the shank insertion instrument and the laser weld failed. | Device Design | On February 13, 2017, the firm emailed its distributors to notify them that their Rev. A devices would be replaced with Rev. B devices, and asked the distributors to mail back the Rev. A devices. |  | 6 | 6 pieces | 2 distributors in CO and TX. | Terminated | November 07, 2018 | 632 | 6
Z-2605-2018 | 80091 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Think Surgical Cutter, Flat, Âż6.2mm x 144mm. For use only with the TCAT Computer-Assisted Tool. Part Number 106428 |  |  | March 30, 2018 | 2018 | Think Surgical, Inc. | There is a potential for the cutter head to detach from the cutter shaft. | Process control | The firm, Think Surgical Inc., sent an "Urgent Medical Device Recall" notices dated 3/30/2018 to customers on 3/30/18. The notices instructed customers to perform the following actions:Examine your inventory, locate any unused devices noted above and quarantine them immediately.If you have received any accessory kits from THINK surgical Incorporated with the noted cutter part numbers included, do not use any affected product and return any unused product to the company.THINK Surgical Incorporated has provided you with a Returned Material Authorization Form (RMA), a FedEx Envelope and return FedEx AirBill to retun any unused devices. Your account will receive a credit and replacement units will be sent. Fill out the accompanying Acknowledgement Form.Email the completed Acknowledgement Form to Ms. Meliha Mulalic at mmulalic@thinksurgical.com.Contact THINK Surgical Incorporated QA/RA Group at 510-249-2337, attention Ms. Meliha Mulalic, should you have any questions.Urgent Medical Device Recall Information: Lot Number Correction letters dated 4/4/18 were distributed as well. The letters corrected an error in the initial notification letter which alerted customers to a typographical error in which lot 9159 should read 9519. |  | 951 | 951 | The devices were distributed to the following US states: NJ, NY, OH, and TX. The devices were also distributed to Korea. | Open, Classified |  |  | 951
Z-2652-2018 | 80427 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 125" (318 cm) Y-Type Blood Set w/170 Micron Filter, Hand Pump, 3 ClaveÂż, Remv 3 Gang 4-Way Stopcock w/Baseplate, Rotating Luer, 3 Ext, Sterile |  |  | July 31, 2017 | 2017 | ICU Medical, Inc. | Some devices were labeled as Does not contain DEHP even though the product may contain small amounts of DEHP. | Labeling mix-ups | The firm, ICUmedical, notified their customers of the recall on about 07/31/2017 via"URGENT: Medical Device Recall Notification letter". Instructions included to inspect inventory on hand for affected devices, remove and return any affected devices, and complete and return the Recall Response Form to ICU Medical RecallCoordinator; email:.Recall@icumed.com; including labeling "Incorrect labeling" in subject or FAX: (801) 264-1755. Instructions also included to notify customers of the recall if devices were further distributed. Questions can be directed to ICU Customer Service Monday through Friday, 7 AM CST - 4:30 PM CT at 877-946-7747, option 9 or e-mail at ProductReturns@icumed.com |  | 160016 | 160,016 units total | Worldwide Distribution: US (nationwide) distribution to AL, AR, CA, CT, DE, FL, IA, IL, IN, KS, MA, MD, ME, MI, MO, NC, NJ, NY, OH, OR, PA, RI, TN, TX, UT, VA, WA, WV; and countries of:: Canada, Cyprus, Japan, Kuwait, MALAYSIA ,and South Africa. | Completed |  |  | 160016
Z-2659-2018 | 80497 | [6 empty] | Leica M530 OHX Microscope System Supplied with 100 - 120 Vac. |  |  | June 25, 2018 | 2018 | Leica Microsystems, Inc. | Recall is due to a design weakness of the power insert module. | Device Design | Leica Microsystems notified customers on about 06/25/2018 via URGENT FIELD SAFETY NOTICE letter. Instructions included advising customers of the issue, informing them that Leica Microsystems representative will contact them to arrange the fiel update to the device, distribute the recall notification to all individuals that should be aware of the recall or if the device was further distributed and a response form was provided to complete and return. The notification letter also instructed customers to report any anomalies with the system to Leica Microsystems if they occur prior to the correction being made and to ensure a functional check is performed prior to each surgery as described in the user manual. |  | 123 | 123 devices total | United States distribution: AR, AZ, CA, CO, CT, DE, FL, IL, IN, KS, KY, LA, MO, NJ, NY, NV, OH, OR, PA, TX, UT, PR. International distribution to Japan, Brazil, Columbia, Ecuador, Mexico, Taiwan | Terminated | February 26, 2019 | 246 | 123
Z-2661-2018 | 80465 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Q-Connect 272 KG Scale used with Maxi Sky 2 ceiling lift |  |  | February 12, 2018 | 2018 | Arjohuntleigh Magog | The spreader bar can detach from the scale attachment under certain conditions. | Nonconforming Material/Component | Thefirm, ARJO, sent an "Urgent-A Field Safety Notice" and Customer Response Forms to affected customers by common carrier and signature required. All affected customers need to ensure that all caregivers and users of the affected devices a remade aware of the Field Safety Notice and all listed devices at their facility are available to be corrected during service technician visit. Affected customers need also to review the Field Safety Notice, acknowledge receipt and understanding of the Field Safety Notice and the required actions by completing, signing and returning the Customer Response Form to the local Arjo office. The Arjo Sales and Service office shall arrange Arjo service technician to visit the facilities to carry out the upgrade of the scales attachment on Maxi Sky 2 ceiling lifts.If you have any further questions or require assistance completing the Customer Response Form, please contact Arjo via email at ArjoCustomer.Response-US@Arjo.com |  | 874 | 874 Units (both styles) | Worldwide distribution: US (nationwide) to states of: CO, FL, IA, IL, KY, MN, NY, PH, and WA; and countries of: Australia, Belgium, Brazil, Canada, China, France, Germany, Hong Kong, Ireland, Italy, Japan, Netherlands, Norway, Singapore, Swiss, Turkey, United Kingdom, and United Arab Emirates. | Open, Classified |  |  | 874
Z-2736-2018 | 80483 | No | GZO | Galvanic skin response measurement device. | Neurology | Neurology | 510(K) Exempt | iTOVi Tracker/Scanner, Model Nos. ITS10 and ITS15 |  |  | July 02, 2018 | 2018 | Hudson Scientific LLC | The device and its accessories may not have been manufactured according to current Good Manufacturing Practices. In addition, the firm has received three complaints regarding electric shocks received when using the device. | Nonconforming Material/Component | On July 2, 2018, the firm notified customers of the recall via an emailed Urgent Medical Device Removal letter. Customers were informed that the company had received three complaints of electrical shock when using the device, and that the device and its accessories may not have been manufactured according to current good manufacturing practices.ACTIONS TO BE TAKEN BY THE CUSTOMER1. Dispose of the iTOVi Scanner and its power adapter and charger in the trash.2. Fill out and return to us the enclosed Attachment A: Urgent Removal Response Letter. The bottomof the form identifies three different methods of returning this form to iTOVi.3. Forward this letter to any users to whom you may have distributed the device.4. If you or others have encountered any illnesses or injuries during or after use of any iTOViTracker/ Scanner products, iTOVi recommends that you arrange for yourself and other productusers to seek medical attention.TYPE OF ACTION BY THE COMPANY:As of June 15, 2018, all iTOVi Tracker/ Scanner software applications (apps) will be deactivated via thefollowing steps:1. Using the server-based software that handles iTOVi account logins, iTOVi will disable the iTOVicomputer software application that runs on iOS and Android operating systems for all iTOVidevices.2. iTOVi will issue an update to the iTOVi software application which displays a message saying thatthe iTOVi Scanner device is no longer usable and should be disposed of in the trash.iTOVi is in the process of creating a more reliable, innovative iTOVi Scanner, and will be sending out anew device to replace older models. |  | 35949 | 35,949 units | Nationwide distribution. Worldwide distribution. | Open, Classified |  |  | 35949
Z-2552-2018 | 80500 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | AirLife Resuscitation Device, Adult labeled as the following: |  |  | May 08, 2018 | 2018 | Vyaire Medical | An error in manufacturing of the AirLife Resuscitation Devices that has the potential to obstruct airflow of the supplemental oxygen delivery during use. | Equipment maintenance | Vyaire Medical notified customers on about 05/08/2018 via URGENT RECALL NOTIFICATION letter. Instructions included to conduct 100% physical inventory for any affected devices, destroy any affected devices according to each facility's destruction protocol, notify customers of the recall if affected devices have been further distributed, and complete and return the Customer Response Form. Questions and comments can be directed to Vyaire Medical's Customer Support Team at 1-833-327-3284 (Option #1) or email: customersupport@Vyaire.com Monday-Friday 8:00 am - 5:30 pm CST. |  | 15714 | 15,714 units total | US distribution only to: AL, AR, AZ, CA, CA, CO, FL, GA, HI, IA, IL, KY, LA, MA, MD, MI, MN, MO, MS, NC,NJ, NY, OH, OR, PR, SD, TX, VA, WA. | Open, Classified |  |  | 15714
Z-2737-2018 | 80265 | No | ODI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Natus Neurology ErgoJust Carts labeled as the following: ICU Cart Export For PTZ Video Ext, IP PTZ HD ErgoJust Video Extension, EuroJust ICU Cart Export For PTZ Video Ext. |  |  | May 18, 2018 | 2018 | Natus Neurology DBA Excel Tech., Ltd. (XLTEK) | Natus Neuro has discovered a possible manufacturing defect in the ErgoJust mobile cart that may cause the cart s video pole to fail. | Nonconforming Material/Component | The firm, Natus Neurology, notified its customers on approximately 05/18/2018 via an "URGENT: Medical Device Correction Notification" letter. Natus Neuro Technical Service will contact customers regarding sending this kit to their bio-medical engineering department. Kits are expected to be available the first week of June.If you have any questions regarding this correction, please contact our Technical Support Administrator by email at ErgojustFCA@natus.com. |  | 797 | 797 Total | United States Distribution: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV.Worldwide Distribution: Austria, Canada, Colombia, Costa Rica, Czech Republic, France, Germany, Hong Kong, Israel, Japan, Jordan, Korea, ,Republic of Latvia, Malaysia, Oman, Philippines, Portugal, Saudi Arabia, Serbia, Slovenia, Switzerland.Sold To Location StateAKALARAZCACOCTDCDEFLGAIAIDILINKSKYLAMAMDMIMNMOMSMTNCNDNENHNJNMNVNYOHOKORPARISCTNTXUTVAVTWAWIWVSold To Location StateAKALARAZCACOCTDCDEFLGAIAIDILINKSKYLAMAMDMIMNMOMSMTNCNDNENHNJNMNVNYOHOKORPARISCTNTXUTVAVTWAWIWVAKALARAZCACOCTDCDEFLGAIAIDILINKSKYLAMAMDMIMNMOMSMTNCNDNENHNJNMNVNYOHOKORPARISCTNTXUTVAVTWAWIWV | Open, Classified |  |  | 797
Z-2727-2018 | 79315 | No | LTF | Electrical salivary stimulatory system. | Dental | Dental | 510(k) | radial stem implant #2 (Sterile packed). Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  |  | October 24, 2017 | 2017 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery. | Under Investigation by firm | On October 27, 2017 an URGENT MEDICAL RECALL NOTIFICATION was issued to customers titled Product Field Action PR 1642784. The letter requests that recipients notify all necessary individuals within their organization of the recall and that a Stryker trauma representative will collect any affected product from the customer location. Questions or concerns can be directed to 201-831-6693 |  | 544 | 544 | Distribution US Nationwide. | Open, Classified |  |  | 19930
Z-2733-2018 | 80555 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | Mallinckrodt Pharmaceuticals INOflo DS, NO Gas Control System. |  |  | July 09, 2018 | 2018 | INO therapeutics LLC/dba ikaria | INOmeters affixed to the INOflo cylinders in Lot 17FP-0570 were programmed with an incorrect expiration date (11/2017). For Lot 17FP-0570, the label affixed to each INOflo cylinder has the correct expiration date noted (11/2020). | Process control | Mallinckrodt Manufacturing LLC communicated to the Market Authorization Holder, Air Water Inc. of Japan the same day the US Food &amp; Drug Administration was notified on July 09, 2018. However, Air Water Inc. of Japan communicated directly with the Healthcare Providers on July 06, 2018 with no formal letters of notification, known to Mallinckrodt. |  | 108 | 108 units | In the country of JAPAN | Open, Classified |  |  | 108
Z-2734-2018 | 80474 | pick, massaging | JET | Massaging pick or tip for oral hygiene. | Dental | Dental | 510(K) Exempt | Waterpik Sonic-Fusion Professional, Model: SF-02 |  |  | July 02, 2018 | 2018 | Water Pik, Inc. | Charging base may overheat with localized melting and sparking, possibly causing fire, shock or burns. | Under Investigation by firm | On 07/02/2018, the firm sent customers an Urgent Medical Device Recall notice. Customers were instructed to stop using affected devices, immediately unplug the units, and to return these devices using firm-provided prepaid return labels. For questions, consumers should call 1-800-674-7718 from 7:00 a.m. to 5:00 p.m. MDT, Monday through Friday. On 07/05/2018, Urgent Medical Device Recall notices were mailed to all customers via overnight mail.On 07/13/2018, the firm issued a press release, which also stated to call 1-800-674-7718 to receive a product return kit. Customers were encouraged to report adverse events to MedWatch via http://www.fda.gov/MedWatch/report.htm (online), http://www.fda.gov/MedWatch/getforms.htm (mail), 1-800-332-1088 (phone), or 1-800-FDA-0178 (fax). |  | 3797 | 3797 | U.S. Nationwide Distribution in the states of TX, KY, WA, NY, VA, CA, NC, IA, MA, OR, OH, GA, NV, CT, IL, SC, PA, MD, IN, WV, ND, MS, CO, AZ, WI, KS, FL, HI, OK, VT, MO, ID, RI, MN, MI, AR, LA, TN, DC, NJ, NH, NE, AL, UT, ME, NM, AK, DE, SD, MT, WY and country of Canada | Open, Classified |  |  | 3797
Z-2662-2018 | 80490 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medfusion Syringe Infusion Pump, containing the Barrel Clamp Guide mechanism |  |  | April 13, 2018 | 2018 | Smiths Medical ASD Inc. | Certain Barrel Clamp Guides may contain a ridge that could potentially lead to spring slippage, resulting in the inability of the pump to recognize a loaded syringe or the pump, misidentifying the size of a loaded syringe. | Process change control | Customers were notified of the recall on about 04/13/2018 via URGENT MEDICAL DEVICE RECALL NOTICE letter. Instructions included to evaluate current inventory for affected Barrel Clamp Guide, review and complete the Recall Response Form, return affected Barrel Clamp Guides to Stericycle for destruction, and if the affected Barrel Clamp Guides have already been installed in pumps, continue to use the pumps, but utilize the Syringe Verification Reference Tool provided with the Recall Notice until the pumps are repaired. Customers were also instructed to notify customers if the affected devices were further distributed. Questions regarding the recall can be directed to Stericycle at SmithsMedical5120@stericycle.com. |  | 276 | 276 units | US distribution to AL, AZ, CA, DC, FL, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NJ, NM, NY, OH, OK, PA, TN, TX, UT, VA, WA. International distribution to Canada and Lebanon. | Open, Classified |  |  | 276
Z-2735-2018 | 80513 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Guider Softip XF Guide Catheter |  |  | December 11, 2017 | 2017 | Boston Scientific Corporation | Potential polymer material degradation. | Process change control | On December 11, 2017, Stryker Neurovascular sent notification to all their customers informing them of the recall. Included is a letter to be sent to each affected account which includes instructions to return product to Stryker and Form to be completed by customers to document products which have been consumed and products to be returned to Stryker.Return the completed form to your nominated Stryker Representative or to NVFieldActions@stryker.com. |  | 31005 | 31,005 | Worldwide - US Nationwide Distribution - AL, FL, IL, IN, KY, MA, MI, MN, MS, NC, NV, NY, OH, OR, TN, TX, UT, WA and the country CANADA | Open, Classified |  |  | 31005
Z-2738-2018 | 80458 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | The Neonatal Flow Sensor Cable used on devices CARESCAPE R860, ENGSTROM CARESTATION, ENGSTROM PRO and contained in Neonatal Software Upgrade Kit 2080496-010. |  |  | June 28, 2018 | 2018 | Datex-Ohmeda, Inc. | The Neonatal Flow Sensor Cable may be missing the connector housing exposing the wires in the cable. | Under Investigation by firm | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated 6/28/2018 to its customers on 06/28/2018. The letter describes the product, problem and actions to be taken. The customers were instructed to complete and return the attached "Customer Response" for via email to Recall34092.NFSCable@ge.com. The firm requested the return of cables that lacked the connector housing.GE Healthcare will replace all affected products at no cost to you.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 417 | 417 units | Worldwide Distribution: US (nationwide) to: Washington, DC and countries of: Algeria, Bangladesh, Brazil, Chile, China, Colombia, Germany, Greece, Guatemala, Honduras, India, Indonesia, Iran, Iraq, Italy, Kenya, Lebanon, Malaysia, Maldives, Mexico, Morocco, Myanmar, Nepal, Netherlands, Nigeria, Pakistan, Poland, Romania, Russia, Saudi Arabia, Singapore, South Africa, Spain, Tunisia, Turkey, United Arab Emirates, and Vietnam. | Terminated | February 22, 2019 | 239 | 417
Z-2739-2018 | 80495 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | VitalGo Total Lift Bed VG-TLB425T-TQR, Version 71064V0-27003 |  |  | April 12, 2018 | 2018 | Vitalgo Inc. | Potential for unintentional movement of the support surface of the bed from tilted position to horizontal position or any position in between. | Component change control | The customer was notified of the recall on about 04/12/2018 via letter. The letter was follow up to previous communications on 02/19/2018, 03/06/2018, 03/07/2018, 03/26/2018, and 04/09/2018 and reminded the customer that all affected beds at each service center are to be placed on Quality Hold. Further directions will be provided at a later time to conduct a correction on the beds. |  | 85 | 85 units | US distribution to Texas. | Open, Classified |  |  | 85
Z-2740-2018 | 80565 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Slit Knife 2.6mm 45degree Bevel Up, Model Number 373726. for use in ophthalmic surgery. |  |  | June 18, 2018 | 2018 | Beaver Visitec | Mislabel of product 373726 Lot 6001700 with 374891 Lot 6001701 on individual packages within the correctly labeled shelf carton. | Packaging process control | BVI issued PRODUCT RECALL NOTICES to customers dated June 11, 2018. Customers were instructed to inspect stock, quarantine and destroy product per internal procedures. BVI is kindly asking the customers to provide objective evidence of product destruction so that a credit can be applied to accounts. Customers are also encouraged to complete and return the response form. If you have questions regarding this recall please call Customer Service at 866-906-8080. |  | 440 | 440 | Distribution to US states of CA, IL, MD, MN, NJ, OH &amp; PA; and Israel | Open, Classified |  |  | 440
Z-2743-2018 | 80079 | No | LTI |  |  | Gastroenterology/Urology | PMA | Apollo Lap-Band AP System Small with RapidPort EZ |  |  | November 22, 2017 | 2017 | Apollo Endosurgery Inc | Component of the Lap-Band system was missing from the package | Under Investigation by firm | Consignees were initially contacted via telephone and later via letter. The firms letter requested the hospital risk manager dispose of the defective product. |  | 54 | 54 | Worldwide Distribution - FL, IL, LA, MI, MO, NC, NY, TN, TX, VA and WV and the countries of France, Italy, Spain, and United Kingdom | Open, Classified |  |  | 54
Z-2589-2018 | 80462 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Arkon Anesthesia Delivery System with Arkon Software Version 2.70, Display Unit Assembly 650-1769-00 installed, and Model #: 99999. This system includes a ventilator. |  |  | July 11, 2018 | 2018 | Spacelabs Healthcare, Ltd. | Arkon Anesthesia Delivery System may go into a failed state (mechanical ventilation ceases) while the machine is in use or while idle. | Nonconforming Material/Component | The firm, Spacelabs Healthcare, sent "Field Corrective Action  Display Unit Resets" letters dated 7/3/2018 and 7/9/2018 to its customers via priority service on 07/11/2018. The letters informed them to circulate this notification to all affected persons and to add copies to appropriate user manuals. The firm will contact customers to replace the display assembly.. Customers with questions are encouraged to call Technical Support at 1-425-363-5212. |  | 328 | 328 | Worldwide Distribution: U.S (nationwide) to states of.: LA, NE, CO, WY, MT, NM, IN, ME, NY, IL, MA, TX, PA, PR; and countries (OUS) of: PERU BOLIVIA, PANAMA, MALAYSIA, POLAND, IRELAND, MOROCCO, MEXICO, CANADA, UNITED KINGDOM, FRANCE, and CHINA, ECUADOR. | Open, Classified |  |  | 328
Z-2742-2018 | 80650 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | 1mL Medallion Syringes |  |  | February 02, 2018 | 2018 | Merit Medical Systems, Inc. | Supplier manufacturing defect with the syringe plunger tip. | Nonconforming Material/Component | On February 2, 2018 Merit Medical sent a "URGENT PRODUCT RECALL NOTICE" letter to all their consignees requesting the following actions;Actions required of you:1. Please immediately determine if any of the devices identified in the attached Customer Response Form (CRF) are within your facility, quarantine them, and discontinue use and distribution.2. Ensure that applicable personnel within your organization are made aware of this field action.3. If the product has been further distributed to other facilities, institutions, or manufacturers, please ensure this notice is immediately shared with them and note the quantity distributed on the CRF. Additional distribution details may be required by health authorities.4. Please fill out, scan and email the completed CRF to Merit Customer Service at response@merit.com or fax at 801.316.4880 within 10 days.5. Please immediately return all affected lots in your possession to Merit, per the instructions found in the CRF.If you have any questions concerning this communication, please dont hesitate to contact your Merit Sales Representative or Merit Customer Service at 801.208.4381. |  | 22234 | 22,234 | Worldwide Distribution: US (nationwide) to states of: AL, AR, AZ, CA, CO, FL, GA, IL, IN, KS, LA, MD, ME, MI, MN, MO, NC, ND, NJ, NM, NY, OH, OR, PA, RI, SC, SD, TB, TN, TX, UT, VA, WI, and WV; and countries of; Canada, Colombia, France, and Italy. | Open, Classified |  |  | 22234
Z-2744-2018 | 80539 | No | FEN | Urodynamics measurement system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | AbViser AutoValve IAP (intra- abdominal pressure) Monitoring Device, Patient mount: a) REF ABV300, b) REF ABV301 |  |  | July 03, 2018 | 2018 | ConvaTec, Inc | Bubble Leak Testing on AbViser AutoValve IAP Monitoring Devices without Transducer confirmed the potential for a pinhole breach in the sterile barrier. | Nonconforming Material/Component | The firm, ConVaTec, sent an "URGENT: VOLUNTARY MEDICAL DEVICE RECALL" letter dated July 2, 2018, to its customers via United Parcel Service of America (UPS) Ground with tracking on July 3, 2018, with emails sent to distributors on July 9, 2018. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1. Inspect your inventory to confirm if you have any of the affected product codes.2 Perform a count of affected product currently in inventory. Complete the enclosed response form and return it to the address on the response form. Return the attached Recall Response Form even if no affected product is in inventory.3 If you have any affected product in inventory please contact ic@convatec.com or call the Customer Interaction Center at the toll free number 1-800-422-8811. You will be provided further instructions regarding how to return affectedproduct and arrange credit. |  | 3881 | 3881 units | worldwide Distribution: US (nationwide) and countries of: Australia, Austria, Canada, Bahrain, Bangladesh, Belgium, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Hungary, India, Israel, Italy, Kenya, Kuwait, Malta, Mauritius, Netherlands,Norway, Portugal, Qatar, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, United Arab Emirates, and Vietnam. | Open, Classified |  |  | 3881
Z-2745-2018 | 80257 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | G8 Automated HPLC Analyzer: HLC-723G8-ST, 021560; and HLC-723G8-LA, 021674 |  |  | June 06, 2018 | 2018 | Tosoh Bioscience Inc | HbAE is known to interfere with the HbA1c assay on the current version of software, Ver. 5.23. Customers should exercise caution when reviewing chromatograms and ensure that Flag 43 is enabled on their device to avoid reporting invalid test results in the presence of HbAE. | Software design | The firm issued an Urgent Medical Device Recall letter to its affected customers on June 6, 2018. In the notice, Tosoh is currently investigating a resolution to this issue to eliminate interference from hemoglobin variants. In the interim, customers should carefully review the chromatography to determine whether or not a variant peak is detected and follow the recommendations included in the letter.Customers were asked to complete and return the attached Acknowledgement Form and the flag setting confirmation to Tosoh Bioscience, Inc., within 14-days of receiving this notification by FAX, email, or postal mail as indicated on the form.If you have obtained unexpected test results or received any complaints of illness or adverse events associated with the use of the analyzer, please contact Tosoh Technical Support 24 hours a day, seven days a week at (800) 248-6764. |  | 665 | 665 | Worldwide Distributed - US Nationwide. and the countries of Foreign distribution to Brazil, Caribbean, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Panama, Peru, and Uruguay. | Open, Classified |  |  | 665
Z-2746-2018 | 80615 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250/MEVION S250i Proton Beam Radiation therapy system |  |  | June 29, 2018 | 2018 | Mevion Medical Systems, Inc. | QFix has advised Mevion Medical Systems that they are identifying unauthorized combinations of QFix Couch Tops with non-QFix devices in clinical use. The kVue Oen Touch latch consists of a precision latch mechanism that can be damaged when used with incompatible (third-party) accessories. This can create unsafe conditions for use and may compromise the subsequent use of any kVue inserts, whether authorized or not. Use ofunauthorized or unvalidated inserts may lead to an incorrect patient treatment or physical injury. | Device Design | On June 29, 2018, the firm sent out a User Notice to consignees to alert customers to the Advisory Notice issued by QFix regarding the use of an Orfit HP Pro Solution combined with the QFix kVue couch top.Customers using third-party inserts are advised to immediately inspect the latching mechanism for damage and contact the third-party insert vendor for advice on the use of the insert and confirmation of compatibility.Mevion will update user manuals to include the following warning about using alternative third party vendor inserts:WARNING: Interchangeable Couch Top inserts are integral to the immobilization,treatment beam planning, and physical safety of each patient. Use ofunauthorized or unvalidated inserts may lead to an incorrect patienttreatment or physical injury. Only use Mevion-provided accessories ordevices that are confirmed by the third party suppliers to be compatiblewith the Mevion Couch for safe and effective patient treatment.Please contact Rob Cessac at rcessac@mevion.com or call 636-485-4962 for anyquestions or concerns. |  | 7 | 7 | US Nationwide Distribution in the states of DC, FL, MO, NJ, OH, and OK. | Terminated | October 02, 2018 | 95 | 7
Z-2749-2018 | 80584 | No | IFB | Cytocentrifuge. | Pathology | Pathology | 510(K) Exempt | EZ Single Cytofunnel Brown, A78710004 |  |  | June 04, 2018 | 2018 | Richard-Allan Scientific Company | Mislabeled product | Labeling mix-ups | On June 4, 2018, ThermoFisher Scientific issued URGENT MEDICAL DEVICE RECALL notices to customers. Customers were advised to inspect stock, quarantine, safely dispose or return impacted product. Evidence of disposal is required to be sent to Thermo Shandon Limited. If affected product has been used with no issues, please notify Thermo Shandon Limited of the quantity that is unavailable for disposal. Complete and return the Recall Return Response Acknowledgement &amp; Receipt form via email to marganne.miller@thermofisher.com |  | 41560 | 1,039 cases/40 boxes per case | US Nationwide Distribution | Open, Classified |  |  | 41560
Z-2748-2018 | 80586 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Pheno Model 10849000. Interventional Fluoroscopic X-Ray System |  |  | June 22, 2018 | 2018 | Siemens Medical Solutions USA, Inc | Potential hardware issue which may cause thermal effects and possible damage to system parts. | Nonconforming Material/Component | On June 22, 2018, Siemens Healthineers issued URGENT MEDICAL DEVICE CORRECTION notices to customers via courier service. In the event the problem occurs, the system can be operated normally, however, in the case of a leakage current, X-ray will not be possible. This may result in a situation where it is necessary to cancel or restart clinical treatment or to continue treatment on an alternative system. Please follow the standard emergency procedures you have in place in the event of system failures. It is recommended to consider reinforcing these emergency procedures until the update has been completed.Siemens will correct the error by updating all affected generators with a RCD (Residual Current Device) to detect leakage currents via Update Instruction AX060/16/S.Additionally, Siemens recommends installing a RCD (Residual Current Device) in the electrical system of the hospital as an appropriate backup system and general electrical protection. You may continue to use the Artis Q and Artis Pheno systems while awaiting implementation of the corrective action.Siemens Healthineers will contact customers to arrange a date to perform this corrective action. Please feel free to contact the service organization for an earlier appointment at 1-800-888-7436. |  | 605 | 605 units in total | US Nationwide Distribution | Open, Classified |  |  | 605
Z-2751-2018 | 80530 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Monitor with one Blood Parameter Module and one HSat Probe; 500AHCT |  |  | June 26, 2018 | 2018 | Terumo Cardiovascular Systems Corporation | H/Sat Hybrid Intensity Shift has been seen in a subset of the H/Sat population, which has resulted in an increase of Hybrids failing the H/Sat Color Chip Test which is an automated self-test when the device is powered on. This will not result in device inaccuracy, only Color Chip Test Failure which disables the H/Sat module. | Vendor change control | On June 26, 2018, the firm began notifying customers of the market withdrawal via telephone calls. The call alerted customers of the issue and scheduled an expedited return of affected CDI System 500 units to the firm's service depot so that Service Technicians may replace the affected H/SAT with a replacement H/SAT device. Customers then received a copy of the completed phone script by email along with a customer response form. Customers were asked to review the phone script, and complete and return the response form. Customers were also asked to assure that all users were notified of the issue, and to prominently display the notice where all users may access it.To contact Terumo CVS with any questions or concerns:Terumo CVS Customer Service: 1.800.521.2818 (Monday  Friday, 8 a.m.  6 p.m. ET)Recall Fax: 1.734.741.6149 |  | 128 | 128 in total | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, FL, GA, IL, KY, LA, MA, MD, MI, MN, MO, MS, NC, NM, NY, OH, SC, TN, TX, WA, and WV. Foreign distribution to Middle East, Latin America, Europe, Japan, Canada, and Mexico. | Open, Classified |  |  | 128
Z-2787-2018 | 80443 | No | DIS | Barbiturate test system. | Toxicology | Toxicology | 510(k) | TCube One Step Multi-Drug Oral Fluid Test Cube (AMPSO,COC20,METSO, OPl40,OXY20,THC2S) |  |  | May 18, 2018 | 2018 | Guangzhou Wondfo Biotech Co., Ltd. | GUANGZHOU WONDFO BIOTECH discovered on the TCube One Step Multi-Drug Oral Fluid Test Cube the potential for failure of the saliva collection device. The use of the saliva collection device could cause the foam tip to become dislodged in the user's mouth creating a choking hazard. | Nonconforming Material/Component | The firm notified its customers on or about 05/22/2018 via a letter entitled "URGENT MEDICAL DEVICE RECALL". The firm advised the end user to examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, please identify the customers and notify them at once of this product recall. The notification to customers may be enhanced by including a copy of the recall notification letter.The firm advised the end user to follow the directions indicated in the "Customer Required Action" section below and return the completed response form as soon as possible by faxing. to 123-456-7890,or by scanning and e-mailing to TOX.FCA@alere.com. They also stated to complete and return the form even if the customer may have consumed or distributed all affected product. Any questions should be directed toward Alere Toxicology Services Phone: 123-456-7890, Email:FCA.TOX@alere.com. |  | 16000 | 16,000 units | US Distribution: in the state of VA. | Open, Classified |  |  | 126000
Z-2791-2018 | 80492 | No | NHC | Ventricular catheter. | Neurology | Neurology | 510(k) | VentriClear Ventricular Drainage Catheter Set, Cat. No. N-VVDC-01-ABRM |  |  | May 31, 2018 | 2018 | Cook Inc. | This lot of VentriClear failed endotoxin testing. Potential adverse events include immune responses ranging from non-specific febrile reaction to life-threatening systemic inflammatory response syndrome (SIRS). | Release of Material/Component prior to receiving test results | On May 31, 2018, the firm contacted the single customer (a distributor) who received the entirety of the affected lot, via email. The email alerted the customer that the lot failed endotoxin testing and instructed the customer to quarantine the lot immediately. The firm then conducted a teleconference with the distributor to further discuss interim controls for future lot release. The recalling firm requested return of the affected lot. The distributor returned the lot to the firm. The recalling firm stated that none of the product was distributed to end users. |  | 250 | 250 | One distributor in Indiana; product was not further distributed to end users. | Open, Classified |  |  | 250
Z-2790-2018 | 80519 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Improve Blood Collection Sets with Multiple Sample Luer Adapter - Sterile/EO |  |  | October 19, 2016 | 2016 | Guangzhou Improve Medical Instruments Co., Ltd. | Complaints about the needle is not smooth, with discoloration, a kind of gel agglomeration on the needle tip. | Nonconforming Material/Component | Several e-mails were sent between the manufacturer to the US distributor in 2016 |  | 700000 | 700,000 | Puerto Rico | Open, Classified |  |  | 700000
Z-2792-2018 | 80327 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MegaPower Electrosurgical Generator, Product code 1000, GTIN 10614559103715 |  |  | May 15, 2018 | 2018 | Megadyne Medical Products, Inc. | If two devices are plugged into the same channel/port in the generator, both devices activate. If one device is left laying on the patient, the result could be burning of the patient. | Use error | The firm sent letters in June 2018, informing their customers of the issue. The firm will send labels that can be applied to the product to remind them not to plug 2 active devices into the same channel at a time. |  | 7673 | 7,673 units | Worldwide distribution. US Nationwide, Mexico, Europe, Middle East, Latin America, Africa, and Asia Pacific Markets | Open, Classified |  |  | 7673
Z-2796-2018 | 80467 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-C3 Ventilator Unit |  |  | March 09, 2018 | 2018 | Hamilton Medical AG | An out of specification component on the control board may cause the ventilator to open the inspiratory channel and the expiratory valve. The ventilator will shut down and an emergency buzzer providing an acoustic alarm and the blinking alarm LED on the device s front side. | Nonconforming Material/Component | Hamilton Medical notified customers on about 03/09/2018 via letter. Instructions included that a Hamilton Medical Field Service Representative will be in contact with each customer to schedule the correction. Questions or concerns can be directed to Hamilton Technical Support at 1-800-426-6331, option #2. |  | 2 | 2 devices | US distribution to CA, CO, IL, LA, MN, NY, OK, VA, WI. Worldwide distribution to China, Germany, India, Indonesia, Israel, Kazakhstan, Korea, Pakistan, Taiwan, Poland, Armenia. | Open, Classified |  |  | 18
Z-2821-2018 | 80534 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | ECHO Bi-Metric Hip System, Reduced Proximal Profile, Standard 135o neck, item number 192414. orthopedic hip prosthesis femoral stem |  |  | June 08, 2018 | 2018 | Zimmer Biomet, Inc. | Two lots of the Echo BI-Metric Hip Stem and ARCOS Modular Revision Hip Stem may have been comingled. Potential health consequences include extension of surgery time to find a replacement component. | Process control | The firm sent an initial email on June 6, 2018, to US distributors that received affected product. Distributors were asked to quarantine any affected product on hand. The formal recall notice was then prepared and emailed to distributors on June 8, 2018. On the same day, hospital risk managers and distributors received notification via courier."The distributors notice identified the issue and their responsibilities. These responsibilities included locating and removing the product in their territory, as well as identifying hospitals who have previously used the product."Distributors will return on-hand product to Zimmer Biomet and ensure all of their products are accounted for using the form provided in the letter."The hospital risk manager notice identified the issue and their responsibilities. These responsibilities include:-Assisting the Zimmer Biomet sales representative with the quarantine of the product -Returning Certificate of Acknowledgement to Zimmer Biomet.If you have further questions or concerns after reviewing this information, please call Zimmer Biomet customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 12 | 12 | US distribution to CA, GA, IL, MO, NY, and VA. | Open, Classified |  |  | 12
Z-2822-2018 | 80540 | No | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | NICO Shepherd' s Hook - Greenberg, sterile. An accessory to NICO s BrainPath device |  |  | June 19, 2018 | 2018 | Nico Corp. | Mislabeled hook pouch included a transport/storage label rather than the product label containing key product information such as part number, lot number, and expiration date, etc | Packaging process control | On June 19, 2018 , NICO issued letter by mail and/or email letters to consignees that have received hooks from the affected lot. The letter provides instructions to inspect all of its current stock of NN-8024, Greenberg Hooks for the proper pouch label. if an incorrect label is discovered, the hook should beremoved from stock and returned to NICO.Questions or need additional information, please contact jay.dittman@.niconeuro.com or 317-660-7118 ext. 104. |  | 204 | 204 devices | US Nationwide distribution | Terminated | November 09, 2018 | 143 | 204
Z-2818-2018 | 80682 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | HANDS-OFF Multi-Lumen Central Venous Catheter with Blue FlexTip for use only with Arrow One-Piece Percutaneous Sheath Introducer System, Material HO-14702 |  |  | July 10, 2018 | 2018 | Arrow International Inc | The packaging may not be sealed. If the packaging is compromised in this manner, the sterility of the product cannot be guaranteed. | Under Investigation by firm | The firm, Teleflex, sent an "Urgent Medical Device Recall Notification" letter dated 7/10/18 to customers. The letter included the following instructions:1. If you have affected stock, immediately discontinue use and quarantine any products with the product codes and lot numbers listed above, so that the affected products can be returned to Arrow.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419- 8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to Arrow International.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document your receipt of this letter.Distributors were instructed to do the following:1. Immediately discontinue distribution and quarantine any products with the product codes and lot numbers listed above.2. Using the provided customer letter and Recall Acknowledgement Form templates, communicate this recall to any of your customers who have received product included within the scope of the recall.3. Have the customers return any affected product to you, together with a completed Recall Acknowledgement Form, for consolidation and return to Arrow International. In the event that an alternative approach is needed, contact Arrow International Customer Service for more information at 1-866-396-2111.4. To return affected products from your inventory, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer s |  | 7024 | 7024 units | Worldwide Distribution: US (nationwide) to the following states of: AK, AL, CA, FL, HI, IL, IN, KS, KY, LA, MA, MI, MN, MO, MT, NC, NE, NY, OH, OR, PA, SC, SD, TX, WI, and WV; to countries of:: Finland, Germany, Hungary, Netherlands, South Africa, United Arab Emirates and the United Kingdom. | Open, Classified |  |  | 12531
Z-2794-2018 | 80552 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Rio System Irrigation Clip - Catalog # 111690 |  |  | May 10, 2018 | 2018 | Mako Surgical Corporation | Out of tolerance. | Nonconforming Material/Component | Stryker sent an Urgent Medical Device Removal letter dated May 14, 2018 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to sign the enclosed Recall Notification Business Reply Form and fax a copy to 1-855-741-4992 or email to strykerortho7035@stericycle.com. |  | 10606 | 10,606 | Worldwide Distribution - US Nationwide in the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WYand the countries of Australia, Hong Kong, India, Italy, Korea, Midrand Allandale, Netherlands, Singapore, and United Kingdom | Open, Classified |  |  | 10606
Z-2793-2018 | 80512 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage Express Fluid Warmer (FW-300, FW-300-A, FW-300-WW, FW-301, FW-301-A) |  |  | June 11, 2018 | 2018 | NxStage Medical, Inc. | There is a potential risk of electrical fire when fluid leaks into the power receptacle on the back of the warmer. The risk exists any time the warmer is plugged in. The firm has received 7 reports of smoke, sparks, or flames coming from the Express Fluid Warmer since 2016. No adverse events have been reported. | Use error | On June 11, 2018, the firm distributed Medical Device Product Safety Notices to affected customers and dialysis centers. Patients were contacted either via email or UPS second day air. Dialysis centers were contacted via email.The recall communication alerted customers that the firm has received 7 reports of smoke, sparks, or flames coming from the Express Fluid Warmer since 2016. No adverse health consequences or medical interventions have been reported. The firm informed customers that they were updating the User Guide to add the following new warning language:"Warning: Do not use any fluid bags that are leaking, and use caution not to spill fluids near thewarmer. If the warmer is plugged in, leaking or spilled fluids in the warmer power receptacle can causean electrical short even if the warmer is not being used. This can lead to an electrical fire and risk ofserious injury or death. To minimize this risk, do the following:" When not used during treatment, do not plug the Express Fluid Warmer into the cycler." Check all fluid bags for leaks prior to and during use." If fluid spills onto the warmer, unplug the warmer power cord from the rear power receptacleand contact Technical Support."Customers were asked to do the following: 1. Please follow the instructions on the attached reply form to confirm that you received this notice.2. Please observe the above warning and unplug your Express Fluid Warmer when not in use.If you have any questions or comments,please feel free to contact Customer Service at 1-800-NXSTAGE. |  | 14975 | 14975 | Worldwide Distribution - US Nationwide in the countries of Australia, Canada, France, Germany, Italy, Kuwait, Netherlands, Saudi Arabia, Spain, Sweden, and UK. | Open, Classified |  |  | 14975
Z-2823-2018 | 80651 | No | FJS | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Flex-Neck Classic Peritoneal Dialysis Catheter (Pediatric) |  |  | December 20, 2017 | 2017 | Merit Medical Systems, Inc. | Mix-up between adult and pediatrics PD catheter. | Packaging process control | Communication to consignees began on December 20, 2017.Written notification to Merit Accounts included product identification information,instructions to immediately quarantine any devices and discontinue use and return unopened product to Merit.Written notification to Merit Sales Representatives included product identification information and instructions to immediately return remaining Peritoneal Dialysis Catheter devices.An email notification was also sent to the Merit sales representatives of the affected accounts. The sales representatives were provided a copy of the Customer Response Form(s) applicable to their account(s). The sales representatives were instructed to assist their account(s) with the completion of the Customer Response Form(s) and the return of the affected units back to the Merit Field Assurance Department. |  |  |  | Worldwide Distribution - US Nationwide in the states of AL, CA, DE, FL, KY, MA, MO, NC, NY, OH, PA, and UT and the countries of Germany | Open, Classified
Z-2827-2018 | 80482 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tosoh Bioscience AIA Analyte Application Manual (AAM) CD |  |  | June 05, 2018 | 2018 | Tosoh Bioscience Inc | Distribution of unapproved product labeling | Error in labeling | The firm, Tosoh Bioscience, notified customers via a phone call on May 29, 2018 and May 31, 2018, reference "Product Labeling-Removal". The phone call described the product, problem and actions to be taken. The firm instructed to confirm receipt of the CD containing AAMs; return the unapproved labeling; if printouts were made of any information on the CD then destroy it, if not, thank customer for response and complete the Customer Call form.The firm will provide the Customers updated CD following approval and release.Any further questions, contact Sr. QA Manager at 614-317-1909 or email: Levent.durukan@tosoh.com. |  | 9 | 9 CDS | US Distribution to states of: FL, MA, MT, OK, OR and WI. | Open, Classified |  |  | 9
Z-2824-2018 | 80636 | Yes | NIK |  |  | Cardiovascular | PMA | Sorin Platinium CRT-D 1711 DF-1, Sterile EO. Implantable Cardiac Defibrillator and Resynchronization Therapy Defibrillator |  |  | July 09, 2018 | 2018 | Sorin CRM SAS | Hardware configuration may lead to overconsumption, followed by loss of pacing and sensing capabilities. The device then cannot identify an arrhythmia that would require defibrillation shock therapy. | Component design/selection | On 07/09/2018, Urgent Field Safety Notices were mailed via certified postal service to physicians informing them about the removal of the non implanted Platinium devices that may present with this hardware failure and providing recommendations for managing implanted patients, which includes patient follow-up one month after discharge, and then every three months until the device nears the replacement date. Follow-up should include checking for battery status and system warnings. The firm recommends that patients be enrolled in remote monitoring. Physicians were instructed to fill out Customer Reply forms and ensure that all personnel involved in the management of patients implanted with affected devices are aware of the information outlined in the notice. For further assistance the firm provided 877-663-7674. |  | 16 | 16 | US distribution to: NJ, AZ, MS, OK, LA, PA, FL, MI, IL, and MS | Open, Classified |  |  | 26
Z-2800-2018 | 80460 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | DISTALE - 203 cm (80") Bifuse Ext Set w/6 Gang 1o2 Manifold, 8 MicroClave Clear (Green Rings), Back Check Valve, 3-Port NanoClave Manifold, Check Valve, Rotating Luer, Filter Cap is a single use, sterile, non-pyrogenic device which provides access for the administration of fluids from a container to a patients vascular system. |  |  | March 01, 2018 | 2018 | ICU Medical de Mexico, S.A. de C.V. | ICU Medical is expanding the scope of their previously initiated recall to include additional I.V. Administration Sets with NanoClave manifold due to the potential to develop a leak or for separation between the female luer and the NanoClave manifold. | Other | ICU Medical notified customers of the recall on about 03/01/2018 via URGENT: Medical Device Recall Notification. The notification informed customers that this was an update and expansion to previously initiated recall. Instructions included to inspect inventory for affected devices, arrange for the return of affected devices, and complete and return the Urgent Medical Device Recall Notification Response Form to recall@icumed.com. |  | 122 | 122 pouches | US distribution to AZ, MA, NC, OH, TX, VA, WA. Worldwide distribution to Belgium, France, Germany. | Open, Classified |  |  | 3591
Z-2829-2018 | 80533 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Nokia BPM+ Wireless Blood Pressure Monitor |  |  | June 11, 2018 | 2018 | Withings Sas | Device could not meet the requirements for systolic pressure | Component change control | Beginning on June 11, 2018, All end-users who have installed a device will be notified of the recall both by a CRM email and a warning content showing in the companion mobile application. End-users who start using the device after the recall initiation date will be notified upon installation by email and app content as well. Retailers, distributors and partners will also be notified.Upon notification of the recall, the users will be presented with the alternative of either choosing to keep their device or send it back and get a refund. If they chose to get a refund, end-users will be asked to fill-out a form to receive a return waybill and get a refund in exchange of their product. Customers having purchased their device from a retailer will be eligible for a refund as well. |  | 1269884000 | Total 44,000 units (15,139 US/ 28,861 OUS) | US Nationwide Distribution | Open, Classified |  |  | 1269884000
Z-2830-2018 | 80558 | No | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | O-SPOT REF 791, 791S |  |  | June 13, 2018 | 2018 | Beekley Corporation | There is a potential for imaging artifact to mimic calcifications on patient's mammograms | Nonconforming Material/Component | On June 14, 2018, Beekley Medical issued an "URGENT MEDICAL DEVICE RECALL" notice dated June 13, 2018 to customers. The notice described the products, problem and actions to be taken. Customers are advised to do the following:- Inspect stock and quarantine affected product- complete and return the acknowledgement response form via fax to: 1-800-735-1234 or mail.- contact firm if inventory of affected product remains with customers at 1-800-233-5539 between 8:00am - 5:30pm EDT.If you have any issues occurred or for further question, contact Quality Assurance &amp; Regulatory Manager at 1-800-233-5539x434. |  | 3384024 | 3,384,024 pieces in total. | Worldwide Distribution: US (Nationwide) and countries of; AUSTRALIA, AUSTRIA, BAHRAIN, CANADA, GREAT BRITAIN, ISRAEL, SAUDI ARABIA, SOUTH AFRICA &amp; UNITED ARAB EMIRATES. | Terminated | November 13, 2018 | 153 | 3384024
Z-2836-2018 | 80575 | No | PEJ |  |  | Clinical Chemistry | PMA | Salivary Estradiol ELISA, REF SLV-4188 in vitro diagnostic quantitative measurement of active free Estradiol, an estrogenic steroid, in saliva. |  |  | July 14, 2016 | 2016 | DRG Instruments GmbH | Lower recovery of salivary estradiol levels due to poor differentiation between kit Standard 0 and Standard 1. | Nonconforming Material/Component | On July 16, 2016, DRG Instruments GmbH emailed customers informing them that they might face a bad differentiation of Standard 0 and Standard 1 Salivary Estradiol kits. Customers were encouraged inspect stock, quarantine and discontinue use of the product. In case the kit was already used, customers were advised to re-check the data and re-run the samples as prevention. |  | 169 | 169 kits | NJ; Germany, Belgium, Spain &amp; Germany | Completed |  |  | 169
Z-2828-2018 | 80669 | No | LYZ | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Powder Free Vinyl Exam Gloves, Item Nos. 14-116, 14-118 |  |  | July 02, 2018 | 2018 | Cypress Medical Products LLC | The firm is recalling one lot of McKesson Powder Free Vinyl Exam Gloves because a sample of product from that lot did not meet certain release specifications. It is possible that certain gloves from this lot may not be of medical exam quality and, therefore, may not provide an effective barrier to infectious material. | Nonconforming Material/Component | On July 2, 2018, Cypress Medical Products LLC notified its customer, McKesson Medical-Surgical, that it was planning to recall distributed products. McKesson Medical-Surgical then notified the affected distribution center as well as end users. End users were sent an Urgent Voluntary Medical Device Recall letter. The letter informed users that the firm is voluntarily recalling one lot of gloves because a sample of product from that lot did not meet certain release specifications. Customers were instructed to do the following:1. Check stock for affected product.2. Immediately cease distribution or use of the affected product, and quarantine immediately.3. Complete the reply form included with the letter and return to Rachel Weaver at Rachel.Weaver@McKesson.com, even if you do not have any affected product in your possession.4. If you have further distributed this product, please notify your customers who received this product about this recall and inform them how to return affected product. Affected product on hand at customer sites should be returned to the distribution center. Counts of returned, quarantined quantities should be forwarded to Cypress for tracking.If you have questions concerning the recall, please contact Ms. Jill Early: 804-553-2146, Jill.Early@McKesson.com. |  | 66 | 66 cases | US Nationwide Distribution in the states of in AL, AR, IA, LA, OK, TX, UT, and VA. | Open, Classified |  |  | 66
Z-2835-2018 | 80691 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | The da Vinci Xi Surgical System (IS4000) Model Name: IS4000, Endoscope Controller. |  |  | May 17, 2018 | 2018 | Intuitive Surgical, Inc. | An Endoscope Controller has been improperly calibrated during manufacturing. This improper calibration may result in the Endoscope Controller showing surgical images that are dimmer, have color imbalance or have reduced fluorescence and grayscale when using Firefly. There is also potential for increased fogging of the endoscope. | Use error | Intuitive Surgical notified customers via a "Field Safety NoticeUrgent Medical Device Correction  ISIFA2018-06-Cda VinciÂż Xi Miscalibrated Endoscope Controller" letter. The letter provided: Reason for Field Action, Risk to Health, Affected Products (description, part number, serial number), Action taken by Intuitive Surgical, Actions to be taken by the Customer/ User, contact information for Further Information &amp; Support, and an Acknowledgement Form. |  | 1 | 1 | U.S. | Open, Classified |  |  | 1
Z-2871-2018 | 80563 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima CT660. Sold under the following product names: |  |  | April 27, 2018 | 2018 | GE Healthcare Japan Corporation | GE Healthcare has identified that some CT systems may have a damaged cable that can expose 120VAC toService Engineers working inside the gantry. | Employee error | The firm, GE Healthcare, sent an " URGENT MEDICAL DEVICE CORRECTION" letter dated 4/27/2018 to its Customers on about 04/27/2018. The letter describes the product, problem and actions to be taken. The customers were Instructed to: download and provide the updated Service Manual to the organization that services the equipment on site and informed customers that GE Healthcare would be correcting all affected devices at no cost. Questions or concerns can be directed to GE Healthcare Service at 1-800-437-1171 or the local Sales Representative. GE Healthcare expanded the recall to include an additional product and model number. Customers were notified on about 08/09/2018 via URGENT MEDICAL DEVICE CORRECTION letter. Instructions were the same as the original notification. |  | 4007 | 4,007 devices total (products 1, 2, 3) | Worldwide Distribution: US (Nationwide) distribution to states of: AK, AL, CA, CO, CT, DE, FL, GA, GU, HI, IA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, RI, SC, SD, TN, TX, UT, WI, WV, WY; and countries of:: ALGERIA, ARGENTINA, ARMENIA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOVINA, BRAZIL, BULGARIA, CAMEROON, Canada, CHILE, CHINA, COLOMBIA, Cote D'Ivoire, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ETHIOPIA, FINLAND, FRANCE, French Guiana, GERMANY, GHANA, GREECE, GUADELOUPE, GUATEMALA, HONDURAS, Hong Kong, HUNGARY, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, Korea (Republic of), KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIA, MALAYSIA, Mali, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW , ZEALAND, NIGERIA, NORWAY, OMAN, PAKISTAN, Palestine, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, REUNION, ROMANIA, Russia, RWANDA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, UZBEKISTAN, VENEZUELA, VIETNAM, and YEMEN. | Open, Classified |  |  | 4099
Z-2875-2018 | 80543 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Connecting Bolt, Model 14-442093 |  |  | June 18, 2018 | 2018 | Zimmer Biomet, Inc. | Product was manufactured from the incorrect material, which shows a potential decrease in the fatigue strength of the connecting bolt that can lead to fracture. | Employee error | Zimmer Biomet issued URGENT MEDICAL DEVICE RECALL notices to customers via courier service. Risk Manager were advised to take the following actions:1. Review the notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete and return the Certificate of Acknowledgement form via email to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5. If you have further questions or concerns after reviewing the notice, call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com.Distributors are encouraged to do the following:Review th notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. Request a Return Authorization Number through Fast SMS or by emailing rgarequest@zimmerbiomet.comc. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.d. Include a hardcopy of Attachment 1 in each carton of your return shipment for |  | 247 | 247 units | AR, AZ, CA, FL, GA, IN, KY, LA, MA, MI, MN, MO, MT, NJ, NY, OH, SC, TX, UT &amp; WI | Open, Classified |  |  | 247
Z-2874-2018 | 80491 | No | JST | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | bioMÂżrieux VITEK 2 NH ID Test Kit |  |  | January 23, 2018 | 2018 | Biomerieux Inc | Missing desiccant. | Nonconforming Material/Component | bioMÂżrieux sent to customers letter requesting the following actions;- Please confirm this letter has been distributed and reviewed by all appropriate personnel within your organization.- Please dispose of any VITEKÂż 2 NH ID test kits, Lot 2450570403, in your stock that were received prior to 25-JAN-2018 and do not have a blue sticker on the carton. Please contact your local bioMÂżrieux representative for replacement of destroyed test kits.- Please store this letter with your bioMÂżrieux VITEKÂż 2 System documentation.- Complete the attached Acknowledgement Form and return it to your local bioMÂżrieux representative.bioMÂżrieux, Inc. is committed to providing our customers with the highest quality products, and we apologize for any inconvenience this has caused your business. If you have any questions or concerns, please contact your local bioMÂżrieux representative. |  | 57 | 57 cartons | US Nationwide Distribution in the states of AK, AL, AR, AZ, CA, CO, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT. VA. VT. WA. WI, and WV | Open, Classified |  |  | 57
Z-2840-2018 | 80690 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ELEKTA Digital Accelerator under the following brand names: Elekta Synergy, Synergy S, Synergy Platform, Precise Digital Accelerator, Infinity, Axesse, and Versa HD |  |  | December 01, 2017 | 2017 | Elekta, Inc. | There is a potential for an uncontrolled extension of iViewGT / XVI detector arms. | Component design/selection | On December 6, 2017 Elekta sent notification to their customers informing them the following:Description of Problem:Elekta has identified that if the microswitch within the middle arm is not set up correctly, it is possible for the MV or kV detector arm to extend due to gravity.When the detector arm is fully open there is no risk.Details:In order for an uncontrolled extension to occur on a detector arm, the following conditions must be present: Âż The microswitch (within the Solenoid Limit Switch of the PCB assembly in the middle arm) is activatedincorrectly. Âż The detector-locked solenoid is not fully engaged in the interface bracket.This is most likely to lead to a potential collision in the following scenarios:iViewGT - The detector arm is opened or closed above the patient, when the gantry angle is at 180Âż ( 45Âż). XVI - The detector arm is opened or closed above the patient, when the gantry angle is at 90Âż ( 45Âż)Clinical Impact:It is possible that a collision with the extended detector arm can cause serious injury.Recommended User Action:Elekta recommends that the user does not open or close the detector arm when it is positioned above the patient, as specified in the Instructions For Use (IFU).Elekta Corrective Actions:Elekta Field Service Representatives will do checks to make sure that that the microswitch activates only when the detector-locked solenoid is fully engaged in the interface bracket. If you have any concerns, contact your Elekta Field Service Representatives who are aware of how to do these checks.This notice has been provided to the appropriate Regulatory Authorities.We sincerely apologize for any inconvenience this action may cause and thank you in advance for your cooperation. |  | 2998 | 2,998 | Worldwide Distribution - US Nationwide in the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, Alaska, District of Columbia, Hawaii and Puerto Rico Albania, Algeria Antigua and Barbuda Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Botswana, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cuba, Curacao, Cyprus, Czech Republic, Denmark,Ecuador, Egypt, Finland, France, Gabon, Georgia, Germany, Ghana, Greece,Honduras, Hong Kong, Hungary, India, Indonesia, Iran (Islamic Republic of),Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of,Lao People's Democratic Republic, Libyan Arab Jamahiriya, Lithuania, Malaysia, Malta, Mexico, Morocco, Mozambique, Myanmar, Namibia, Nepal, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Paraguay,Peru, Philippines, Portugal, Republic of Serbia, Romania, Russian Federation,Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Venezuela, Viet Nam and Zimbabwe | Open, Classified |  |  | 2998
Z-2838-2018 | 80735 | No | IYB | Diagnostic x-ray tube mount. | Radiology | Radiology | 510(K) Exempt | 3D V, model. no. |  |  | June 29, 2018 | 2018 | Siemens Medical Solutions USA, Inc | In rare cases of insufficient maintenance or high clinical workload, the first rope of the ceiling stand, which is designed to take the load, could break without triggering the safety lock and this could lead to overloading the safety rope which is not designed for continuous load and ongoing movement under load. This may lead to mechanical fatigue and cause the arm to drop down during patient positioning. Patients and users may be seriously injured. | Device Design | On 6/29/18, the firm emailed its four affected original equipment manufacturers (OEMs), alerting them to the potential malfunction of the ceiling stand's lifting column that could cause the arm to drop down during patient positioning.Manufacturers were asked to distribute an attached Safety Notice to affected customers. In addition, the firm urged manufacturers to order retrofit kits for all customers and to take any steps needed to ensure that customers install the kits. The kits are free of charge. Manufacturers were also asked to complete a response form and return to the firm.Affected customers were advised of the product problem and instructed to install retrofit kits to avoid risk to patients and to users.If you have any questions, please contact Mr. Wolfgang Bohrer at wolfgang.bohrer@siemens-healthineers.com. |  | 667 | 667 | US Nationwide Distribution to accounts in IL, FL, and NY. | Open, Classified |  |  | 1366
Z-2839-2018 | 80671 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi Scenaria CT system |  |  | June 29, 2018 | 2018 | Hitachi Medical Systems America Inc | The firm received a complaint where, during patient positioning, the technologist was adjusting the patient table height and the table failed to stop when the control was released and contacted the CT gantry. The problem could not be duplicated at the site, so to avoid further issues while a root cause investigation was pursued, Hitachi choose to add an optional touch sensor to the gantry to add a redundant level of protection. It is conceivable that a patient s extremity may be caught between the tabletop and the gantry surface during uncontrolled motion or that the table contact of the gantry might cause internal damage to the unit. | Nonconforming Material/Component | On 6/19/18, the firm was notified of the problem at a customer site. The firm is currently investigating the underlying cause to formulate a planned corrective action. On 6/29/18, the firm took a preventive measure at the site which originated the complaint: Hitachi added an optional touch sensor to the unit to lessen the chance of injury. The firm may take additional corrective action on other units distributed to customers once root cause investigation is complete. |  | 1 | 1 | US Distribution to one customer in TX. | Open, Classified |  |  | 1
Z-2869-2018 | 80573 | No | JTY | Culture medium for pathogenic Neisseria spp. | Microbiology | Microbiology | 510(k) | BD BBL GC-Lect Agar; Cat. No. 297715 |  |  | June 21, 2018 | 2018 | Becton Dickinson &amp; Co. | A portion of this lot was manufactured using Gentamicin instead of Vancomycin. Gentamicin is an inhibitory agent of Neisseria organisms. Thisproduct is a selective plated medium to provide enhanced growth and recovery of Ngonorrhoeae. Inhibited growth may cause a false negative result or a delayed result leading toincorrect or delayed treatment. | Under Investigation by firm | The firm, BD Diagnostic System, sent an "URGENT MEDICAL DEVICE PRODUCT RECALL" letter dated June 21, 2018, to its customers via UPS, to the attention of Microbiology Manager/Supervisor. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:I. Complete the Customer Response Form and return to the BD contact noted on the form whether or not you have any of the impacted material so that BD may acknowledge your receipt of this notification.2. Discard any remaining inventory of lot 8110781 . BD will process replacement of any discarded inventory upon receipt of the Customer Response Form.3. Share this Recall Letter with all users of the product within your facility to ensure awareness.BD has identified the root cause of this situation and has initiated corrective actions to preventrecurrence in the future.If you require further assistance, please contact: 800-638-8663Customer/Technical Support Monday- Friday between 7:00am and 7:00pm (EST) in the United States. |  | 10360 | 10,360 plates | US Distribution in states of: AK, AL, AR, AZ, CA, CO, DC, DE, FL, GA, HI, ID, IN, KS, KY, MA, MI, MS, MT, NC, JH, NY, OK, OR, PA, SC, TN, TX, WA, and WI. | Open, Classified |  |  | 10360
Z-2854-2018 | 80577 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DigitalDiagnost 4.1 Flex / Value (Stitching Patient Support) 712028 |  |  | June 27, 2018 | 2018 | Philips Medical Systems Gmbh, DMC | The hook does not securely hold the footplate in the vertical position. Therefore, the footplate can fall or swing down and injury operator or patient. | Device Design | On June 27, 2018, Philips Medical Systems issued URGENT FIELD SAFETY NOTICES, Medical Device Recall letters to customers. Customers were encouraged to always check the right position of the hook and ensure that the hook is locked according tot he IFU. Press hook down until it is locked. Make sure that the hook is in position. A Philips Service Engineer will contact impacted customers as soon as the Field Action Kit is ready to be implemented. Should you need to communicate with Philips in regards to this program, please reference FCO-71200185. If you need further information or support concerning this issue, please contact your local Philips representative. |  |  |  | US Nationwide; International to 66 countries | Open, Classified
Z-2876-2018 | 80548 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Persona Partial Knee System Spacer Block Alignment Tower, Item No. 42-5399-006-35 |  |  | June 22, 2018 | 2018 | Zimmer Biomet, Inc. | The PPK System Spacer Block Alignment Tower mating feature that mates with the spacer block was undersized, which could potentially lead to mating issues between the devices. This could cause a delay in surgery of less than 30 minutes. There is an option in the surgical technique for performing limb alignment without the alignment tower. | Process control | On June 22, 2018, the firm distributed Urgent Medical Device Recall letters to distributors and hospital risk managers.Actions for Risk Managers:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.Actions for Distributors:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.c. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.d. Include a copy of Attachment 2  Certificate of Sterilization with returned instruments. Mark RECALL on the outside of the returned cartons.4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com.a. Review the list of hospitals included with the email notification sent to your facility, which includes a list of hospitals that have already been notified of this recall.b. Identify whether there are any additional hospitals that Zimmer Biomet has not notified and list these |  | 385 | 385 | Worldwide Distribution - Nationwide distribution and the countries of Japan and Netherlands. | Open, Classified |  |  | 385
Z-2877-2018 | 80487 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ulnar Osteotomy Compression Plate, Short |  |  | April 16, 2018 | 2018 | TriMed Inc. | Two instrument holes located on the side of the plate, which act to align the saw guide are located on the wrong side of the plate | Process control | Tri-Med issued notification on April 16, 2018 requesting: Determine the lot number of all Ulnar Osteotomy Compression Plates, Short (UOCP-S) .If you have UOCP-S plates with the lot numbers listed (10003, 10017):o Remove the UOCP-S plates from the tray and quarantine it. Do not use this plate. Contact TriMed at quality@trimedortho.com or call 800-633-7221. TriMed will arrange have the plate(s) returned to the company and replaced.Return the Recall Response card attached to this letter to TriMed as soon as possible. Contact TriMed via telephone or email as per below if you have any questions: Telephone*: 800-633-7221, Email: quality@trimedortho.com.TriMed will replace the potentially defective plates with plates that have been manufactured and confirmed to meet TriMeds specifications. |  | 722 | 722 | Worldwide Distribution in the state of TX and countries of Australia, China Sweden | Open, Classified |  |  | 722
Z-2741-2018 | 80545 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | CPAP Mask Cushion, Model Nos. PB781S, PB781M, PB781L, 781S, 781M and 781L |  |  | June 18, 2018 | 2018 | Compass Health Brands (Corporate Office) | A May 2015 design change in the cushion seal replacement part and accompanying elbow replacement part resulted in incompatibility of the new cushion seal with the previous design of the elbow. Use of the new cushion seal with the previous design elbow could lead to unacceptable CO2 porting. As there was no model number change for this design change, customers may not be aware of the incompatibility. | Component change control | On June 19. 2018, the firm sent Urgent Recall Field Correction notices via certified mail. The notice alerted customers to the following product issue:Compass Health Brands Corp. has recently identified the need to clarifycompatibility information for the replacement cushion seals for the Model# PB7801 Zzz-Mask SG Series Full Face CPAP Mask.&gt;- A design change made to the cushion seal replacement part (and accompanying elbowreplacement part) in May 2015 (18 months prior to Compass Health acquisition of the Pro basicsproduct line) resulted in incompatibility of the new cushion seal with the previous design of theelbow.&gt;- While no complaints or injuries have been reported to-date, use of the new cushion seal withthe previous design of elbow could result in unacceptable CO2 porting.&gt;- Since a model number change did not accompany this design change, customers and theirconsumers would likely not be aware of this incompatibility.Customers were asked to do the following:Ensure all affected personnel are fully informed of this notice." Forward this notice to your Regulatory Compliance Manager, Purchasing Manager, CustomerService Manager and Field Technicians." Advise your customers regarding the proper combination of cushion and elbow. The attachedconsumer communication letter can be used to communicate this information. As noted on theattached consumer communication letter, any consumers who are currently using anincompatible cushion/elbow combination should be instructed to discontinue use of thecushion and elbow and should be instructed to contact their dealer to obtain replacement kits." Dispose of any current stock you may have of the affected cushion model numbers (PB781S,PB781M, PB781L, 781S, 781M and 781L manufactured after May 2015) and contact CompassHealth for replacement kits (NOTE: Full mask model numbers are NOT affected by this recallfield correction)." Please contact Compass Health Brands Corp. Customer S |  | 742 | 742 | US nationwide distribution to AL, CT, FL, GA, KY, MS, NC, NY, TN, and TX. | Open, Classified |  |  | 742
Z-2753-2018 | 80556 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Handicare P600 Series Patient Lift |  |  | November 03, 2017 | 2017 | Handicare Usa Inc | Premature strap wear and breakage at maximum weight conditions (625 lbs.). | Nonconforming Material/Component | HandiCare USA sent an Urgent notification letter dated November 3, 2017 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were informed new units will be sent to replace the affected units identified at no charge. For questions contact your Customer Service |  | 4 | 4 | Worldwide Distribution - US Nationwide in the states of AK, CA, CT, DE, GA, IA, IL, KY, LA, MA, ME, MI, MN, MO, NE, NV, NY, OH, OR, PA, RI, TN, TX, WI, and Puerto Rico and the countries of Canada and Australia, | Open, Classified |  |  | 720
Z-2881-2018 | 80639 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Dymaxeon 5.5 Curved Rod Size 5.5 x 110mm (Yellow), Model Number 06.6441B |  |  | June 29, 2018 | 2018 | Back 2 Basics Direct LLC | Packages labeled as containing 110mm rods may contain 100mm rods. | Labeling mix-ups | Customers were contacted via phone the week of 3/5/18 and 7/12/18. The customers were informed of the recall and asked any of the affected remained in their possession, if the products have already been implanted, and if any patients experienced any adverse events. |  | 10 | 10 | The products were distributed to the following US states: OH. | Terminated | November 15, 2018 | 139 | 10
Z-2882-2018 | 80680 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vertebral Motion Analyzer (VMA) Version 2.3.252. |  |  | January 11, 2017 | 2017 | Ortho Kinematics, Inc | Ortho Kinematics Inc. sent a Notice of Correction to Released Testing Results, Radiological Read Report for the Vertebral Motion Analyzer (VMA) test because it contained an error. The error occurred due to a software bug that has been corrected. | Software Manufacturing/Software Deployment | Ortho Kinematics Inc sent notification to customers on 1/11/2017 via Letter and electronic notification through web app. The company deployed a software upgrade (VMA version 2.3.253) on January 11, 2017. |  | 1 | 1 | US Nationwide Distribution | Open, Classified |  |  | 1
Z-2833-2018 | 80174 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Liebel-Flarsheim Hydra Vision Urology X-Ray System (DR) |  |  | February 27, 2017 | 2017 | LIEBEL-FLARSHEIM COMPANY LLC | Software issue. The difference between the display and dosimeter readings In the Child/Pediatric automatic exposure mode is in the range of 52-65%. The display in all other modes deviated by less than 35% from the dosimeter readings. | Radiation Control for Health and Safety Act | The firm, Liebel-Flarsheim, notified its customers by phone on or about 2/27/2017 to inform them of the product, problem and actions to be taken. The firm plans to have these Sedecal software updates implemented over the course of the next 12 months.For additional information, please contact me at 314-376-4788. |  | 313 | 313 | US Distribution to: AL, AZ, CA, CT, CO, FL GA, HI, IA, IL, ID, KS, KY, MD, MI, Michigan, MN, NC, NH, NV, NY, OK, OH, PA, SC, TN, TX, VA, WA, WI, and WV, | Open, Classified |  |  | 341
Z-2880-2018 | 80618 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | CoLink(TM) Lapidus Plate XP, +2 mm, Left, REF P40 ST265, QTY 1, In2Bones |  |  | July 16, 2018 | 2018 | In2bones USA, LLC | CoLink(TM) Lapidus Plate +2mm Right and Left, were laser marked incorrectly. In these lots, right plates were laser marked as left plates and the left plates were laser marked as right plates. | Process control | In2Bones sent a Medical Device Recall letter dated July 18, 2018 to affected customers (agents and hospitals). The letter identified the affected product, problem and actions to be taken. The Agent notice requested the consignee to inform and distribute the notice to all relevant person within the organization. The firm is seeking the return of the productsThe notice to the hospitals and healthcare professionals explained the problem and recommended that patients be monitored for any adverse events within the normal standard of practice.For questions contact Quality and Regulatory Affairs Team at 901-260-7931.. |  | 48 | 48 units | Worldwide Distribution - US Nationwide and the countries of France | Terminated | February 22, 2019 | 221 | 95
Z-2826-2018 | 80508 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | The Tri-Flo Subglottic Suction System |  |  | June 20, 2018 | 2018 | Vyaire Medical | Vyaire Medical has discovered potential patient safety risk with associated with the Tri-Flow Subglottic Suction System. The distal soft tip of the catheter has been identified as having the potential to break off during patient use and result in aspiration of a solid foreign body. | Under Investigation by firm | The firm notified its customers on/about 06/20/2018 via a letter entitled " URGENT: Tri-Flo" Subglottic Suction System RECALL NOTIFICATION". This letter described the affected part, and lot numbers as well as the distribution dates for affected part numbers. The letter described the following actions to be taken by the customer which include:1) Inspect current inventory on-hand. A 100% physical inventory should immediately be performed to identify and remove the Tri-Flo TM Subglottic Suction Systems (CM28010) from distribution due to the identified potential patient safety risk.2) If a facility has further distributed affected product, Vyaire states a entity must further notify customers of this URGENT RECALL NOTIFICATION .3) Return all in-stock inventory affected by this recall by calling Vyaire's Customer Service Support team at 833-3BREATH (833-327-3284) (Option #1) Monday-Friday 7:00 am CST- 5:30 pm CST to obtain a pre-paid shipping Return Goods Authorization (RGA). Requests for credit can be made by providing the original P.O. number for in-stock inventory.4) Complete the Customer Response Form (Appendix 1) and return to GMB-GLBÂż VSFieldActions@Vyaire.com.The letter also included instructions if an end-user customer/ facility that did not purchase the product directly through Vyaire Medical, to work with the distributor to return or receive credit for in-stock inventory affected by this recall.Any additional questions concerning this voluntary Recall, the firm recommended contacting Lindy Schenning at 1.872.757.0109 or via e-mail at Lindy.Schenning@Vyaire .com. |  | 2150 | 2150 Units Total | Domestic Distribution: AZ, CA, CT, FL, GA, IA, IL, IN, MA, MD, MN, NJ, OH, OK, SD, TX, WA, WI. | Open, Classified |  |  | 2150
Z-2886-2018 | 80631 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Pediatric Two-Lumen Central Venous |  |  | July 03, 2018 | 2018 | Arrow International Inc | One lot of product is from product code AK-22502 so it contains a 2-lumen, 5 Fr x 5 cmcatheter. However, this product is incorrectly labeled as AK-25502, so the lidstock says itcontains a 2-lumen, 5 Fr x 8 cm catheter. No patient injuries have been reported related to thisissue. | Employee error | On July 3, 2018, the firm sent out Urgent Medical Device Recall letters to customers. The letter communicated the labeling error.Affected consignees are instructed to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products withthe product code and lot number listed above, so that the affected products can be returnedto Arrow.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allowus to document the amount of product you have on hand for return. A customer servicerepresentative will contact you with a Return Goods Authorization (RGA) Number and willprovide instructions for the return of products to Arrow International.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Formand fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. Thiswill allow us to document your receipt of this letter. |  | 166 | 166 | Distributors in MO and UT. Medical facilities in AL, CA, IL, MA, NH, NM, NY, TX, VA, and WI. | Open, Classified |  |  | 166
Z-2897-2018 | 80583 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Artholine Punch B 4, 6MM H 1, 6MM, used for the removal of soft tissue, cartilage tissue or other non-osseous tissue via natural or surgically created passages. |  |  | July 24, 2017 | 2017 | Richard Wolf GmbH | Punch may have a burr in one section of the joint and when the punch is open, there is an increase potential for injury. | Process control | The firm notified customers on about 07/24/2018 via a Urgent safety information letter. Customers were instructed to check inventory for the affected device, return the product to Richard Wolf GmbH, and complete and return the response form. The letter also instructed customers to ensure all users in their organization are notified of the recall and to notify customers if the product was further distributed. Safety-related questions can be directed to: Tel.: 0049 7043 35 1013 Questions concerning handling can be directed to: Technical Service Dpt., Tel.: 0049 7043 35 1046 |  | 12 | 12 units | Worldwide distribution to Germany, Hong Kong, Ireland, Romania, and Turkey. | Open, Classified |  |  | 12
Z-2885-2018 | 80414 | Yes | BZR | Breathing gas mixer. | Anesthesiology | Anesthesiology | 510(k) | O-Two EquinoxÂż Relieve N2O/O2 Analgesic Gas Mixing and Delivery System |  |  | May 14, 2018 | 2018 | O-Two Medical Technologies, Inc. | Higher levels of nitrous oxide delivered than specification | Device Design | O-Two Medical Technologies Inc. issued letter on May 14. 2018titled: Notice of Product Upgrade of the 0-Two Equinox Relieve Analgesic Gas Delivery System , requesting customers to return their units for inspection and upgrade. Contact via email to david@otwo.com. |  | 3 | 3 units | AK, CA, FL | Open, Classified |  |  | 3
Z-2887-2018 | 80340 | Yes | BZR | Breathing gas mixer. | Anesthesiology | Anesthesiology | 510(k) | CAREstream ProNox Nitrous Oxide Delivery System labeled as the following: |  |  | May 15, 2018 | 2018 | CAREstream Medical LLC | Two analgesic gas delivery system oxygen/nitrous oxide mixers were reported to have delivered higher levels of nitrous oxide than specification during routine testing. | Under Investigation by firm | CAREstream Medical Ltd. notified its customers on or about 05/15/2015 via a Medical Device Product Correction letter. The letter informed customers to return their units, (model number: CMI-0100-PNX-xxx) for inspection and upgrade at the firms service center in Florida. So as to allow the customer's continued use of the devices safely during this process (and until all units are upgraded), CAREsteam recommends the device only be used with an oxygen monitor with alarms, attached to the device in accordance with our product manual (refer to Users Manual (revision 6) for instructions). Customers were also instructed to fax or email CAREstream Medical Ltd. to acknowledge receipt of the Medical Device Product Correction letter. |  | 919 | 919 Units Total | Domestic Distribution: AK, AL, AR ,AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV. International Distribution: Canada and Morocco. | Open, Classified |  |  | 919
Z-2890-2018 | 80768 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vertebral Motion Analyzer (VMA), VMA DataLink Software.2.3.231, 2.3.232. |  |  | August 12, 2016 | 2016 | Ortho Kinematics, Inc | Ortho Kinematics, Inc. notified customers of erroneous results of VMA tests due to a software bug that has been corrected. | Software design | Ortho Kinematics, Inc. notified customers via email dated August 19, 2016, Notice of Correction to VMA Results. The notification included the following information: reason for notice of correction, error information, and required action of acknowledgement. |  | 1 | 1 | U.S. Nationwide Distribution | Open, Classified |  |  | 1
Z-2884-2018 | 80613 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker T2 ANKLE ARTHRODESIS NAIL RIGHT 10MMX200MM Item Number: 1819-1020S |  |  | June 27, 2018 | 2018 | Stryker GmbH | Incomplete seal on the Tyvek lid may compromise the sterility of the device | Packaging process control | Stryker notified consignee via UPS (with return receipt) on June 27, 2018 advising user of the problem, quarantine product and return. Complete and sign the Recall Notification Business Reply Form. Contact (201) 831-6693. |  | 38 | 38 | US distribution to KY | Open, Classified |  |  | 38
Z-2891-2018 | 80771 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vertebral Motion Analyzer (VMA), 2.3.239 - 2.3.254. |  |  | March 02, 2017 | 2017 | Ortho Kinematics, Inc | Ortho Kinematics notified customers that errors were contained in Radiological Read Report and VMA Report for the VMA test. The error resulted from an administrative error of VMA user account settings and has been fixed. | Software design | Ortho Kinematics notified customers via telephone, email, and letter. Letter dated August 23, 2017 contained the following information: reason for notification, error report information, and requested acknowledgement of receipt. |  | 1 | 1 | U.S. Nationwide Distribution | Open, Classified |  |  | 1
Z-2883-2018 | 80752 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS II NON-POROUS CRUCIATE RETAINING FEMORAL COMPONENT, SIZE 6 RIGHT, REF 71420026 Knee prosthesis |  |  | August 01, 2017 | 2017 | Smith &amp; Nephew, Inc. | Two lots of Genesis II Non Porous C/R Femoral Size 6 right packaged were distributed containing Genesis II C/R Femoral Size 4 left. | Employee error | The firm notified its consignees of the problem and the recall by email on 08/01/2018. The firm requested the return of the product. |  | 27 | 27 units | Australia, Austria, Germany, Great Britain, Poland, Spain | Terminated | February 19, 2019 | 567 | 27
Z-2889-2018 | 80769 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vertebral Motion Analyzer (VMA), 2.3.250. |  |  | May 23, 2018 | 2018 | Ortho Kinematics, Inc | This email is to provide notification that, due to a software bug that has been corrected, the Radiological Read Report for the VMA test(s) contained an error. | Software design | Ortho Kinematics notified customers via Email and notification through web app with a notification dated December 6, 2016. The notification contained the following information: reason for correction, report information, and required action for acknowledgement. |  | 1 | 1 | U.S. Nationwide Distribution | Open, Classified |  |  | 1
Z-2896-2018 | 80645 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips SureSigns VS2 monitor, 863079 |  |  | July 13, 2018 | 2018 | Philips Electronics North America Corporation | Philips has received several reports in which monitors with lithium ion batteries that had exceeded their battery life expectancy have overheated or ignited.These batteries should be replaced every 3 years or upon reaching 300 charge-discharge cycles.Although the Philips SureSigns VS &amp; VM Monitors and View Station (VSV) can display actual information on either or both battery age and charge-discharge cycles, the existing labeling does not include full instructions on how to use this information to determine when to replace the battery. An overheated battery may in turn cause the device case to overheat and possibly melt or cause the device to ignite, which can cause injury to a patient, nearby users, or cause damage to property. | Labeling design | On July 13, 2018, the firm issued an "URGENT-Medical Device Correction" letter dated July 10, 2018, to affected customers to advise them that that monitors with batteries that have exceeded their specified replacement interval may overheat or ignite. The firm issued a Service Guide addendum included with the letter, which contains information to assist with managing the battery and its replacement.Customers were advised to do the following:Upon receipt of this notification, carefully read the enclosed SureSigns Addendum to the Service Guide.Âż VS Series Customers -Addendum to the Service Guide: 453564807541Âż VM Series Customers -Addendum to the Service Guide: 453564807551Âż VSV Customers- Addendum to the Service Guide: 453564815301Promptly perform the Battery Maintenance that is called out in the Addendum for each of your affected Philips SureSigns VS &amp; VM Monitors and View Station (VSV) by determining if the Battery Information Screen indicates if the battery cycle count exceeds the limit of 300 cycles, or the battery is older than three (3) years. If so, the battery needs to be replaced. The replacement battery can be ordered using the standard Philips replacement processes.Details on how to replace the battery can be found in the Philips SureSigns VS &amp; VM Monitors and View Station (VSV) Service Guide.Review this information with all staff members who are responsible for device management of the Philips SureSigns VS &amp; VM Monitors and View Station (VSV).Please store the below Service Guide Addendum with your Philips SureSigns VS &amp; VM Monitors and View Station (VSV) Service Guide documentation.Âż VS Series Customers -Addendum to the Service Guide: 453564807541Âż VM Series Customers -Addendum to the Service Guide: 453564807551Âż VSV Customers- Addendum to the Service Guide: 453564815301Complete and return the response card provided.In addition to providing the Service Guide Addendum and this notice, Philips plans to release a customer ins |  | 207439 | VM 96,324 Units; VS 110,424 Units; VSV 691 Units | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Afghanistan, Albania, Andorra, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bhutan, Bolivia, Brazil, Bulgaria, Canada, Cambodia, Chile, China, Colombia, Costa Rica, Cote D'Ivoire, Croatia, Czech Republic, Dominican Republic, El Salvador, Egypt, Estonia, Ethiopia, France, French Guiana, French Polynesia, Finland, Guadeloupe, Germany, Ghana, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Israel, Italy, Japan, Jordan, Kenya, Korea, Kosovo, Republic of, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Lybia, Macao, Malaysia, Maldives, Malta, Martinique, Mauritius, Mexico, Morocco, Mozambique, Myanmar, Namibia, Nepal, Netherlands, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Palestine, State of, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, Zambia and Zimbabwe. | Open, Classified |  |  | 207439
Z-2888-2018 | 80704 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO lnfiniti Needle Guide Bracket. used for performing needle-guided (or catheter) procedures with the use of diagnostic ultrasound transducers. |  |  | February 23, 2018 | 2018 | Civco Medical Instruments Co. Inc. | A cover failed leak testing during validation testing of new needle guides from a replacement tool. | Device Design | Telephone Script to be read:CIVCO is calling to notify our customers regarding a recall action being initiated.CIVCO has recently discovered a dimensional issue in our Infiniti bracket which may allow the needle guide to a puncture a cover. You are being notified, as your facility has been identified as having received product from one of the affected manufacturing lots. Advise on Action to be taken by Customer or Distributor:To return Infiniti needle guide, contact your Account Representative at 800-445-6741 or 319-248-6757. Alternately you may also use email at order@civco.com to schedule the return of your device.Additional information to obtain:Verification of Email contact to send notification letter and response letter. |  | 177 | 177 boxes | Worldwide Distribution - US Nationwide in the states of GA, KY, CA, NY , NJ, TX, IL, MA, RI, AZ, and IA and countries of Ukraine, Hong Kong, Italy, Netherlands, Switzerland, United Kingdom, Denmark, Germany, Belgium, Spain, Japan, China, and Canada | Open, Classified |  |  | 177
Z-2917-2018 | 80764 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | August 03, 2018 | 2018 | Elekta, Inc. | Order Status Remains Approved When Should Indicate Complete. | Unknown/Undetermined by firm | On August 2018 the firm sent letter to all their customer with the following instructions; Recommended User Action:A user with appropriate pharmacy rights can unlock the record. This blanks out the details requiring the user to re-enter the administration details previously entered.To avoid this locking issue, close the MAR window before navigating to the CWS. This document contains important information for the continued safe and proper use of your equipment.Âż Please post this notice in a place accessible to all users, e.g. Instructions for Use, until this action is closed.Âż Advise the appropriate personnel working with this product the content of this letter. Elekta Corrective Actions:The issue is resolved in MOSAIQ Release 2.65 Beta 02 and in release 2.64 SP10.In order to meet regulatory requirements, you are required to complete this the Acknowledgement Form and return it to Elekta immediately upon receipt but no later than 30 days. |  | 185 | 185 | AK, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, KS, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, WA, WI, WV, and WYAustralia, Canada, Ireland, New Zealand, Portugal, Singapore, South Africa, Spain, Sweden, and United Kingdom | Open, Classified |  |  | 185
Z-2916-2018 | 80654 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Johnson and Johnson Vision Care 1-Day ACUVUE MOIST Daily Wear, 8.5 Base Curve, -3.25 Power. Master Lot, containing 12 split lots, repackaged in 30 and 90-packs . |  |  | July 23, 2018 | 2018 | Johnson &amp; Johnson Vision Care, Inc. | Johnson and Johnson Vision Care received a number of reports of a partially formed contact lens, which appears to the customer or consumer as visibly smaller in diameter, thicker/stiffer to the touch and more darkly tinted blue than the normal contact lens. | Equipment maintenance | On July 23, 2018, the Johnson &amp; Johnson notified customers of the recall via an "URGENT FIELD SAFETY NOTICE" letter date July 23, 2018. The letter to US and international customers included the following instructions: review inventory and determine if any affected product is on hand, cease using any affected product, provide the recall notification to anyone within the organization who needs to be aware of the recall, return any affected product and arrange for replacement product, and complete and return the enclosed Customer Reply Form. The distributor notification included a copy of the appropriate recall notification depending on the country they are located in as well as the following instructions: to cease distribution of any affected inventory, remove from inventory any affected product, distribute the included RECALL NOTIFICATION to all customers where product was further distributed, ensure customers understand the importance of completing and returning the Customer Reply Form, arrange for the return of any affected product on hand.If you have any questions, contact Sr. Director, Worldwide Quality Systems Strategy at 904-443-3794 or aalvar56@its.jnj.com. Also contact Project Manager, Escalation &amp; Field Actions, phone 904-443-3516, email amyers@its.jnj.com. |  | 195660 | 195,660 Total; 89,640 US &amp; Puerto Rico, 106,020 OUS | Worldwide distribution: US (Nationwide) distribution to CA, CT, FL, GA, IA, IL, MA, MD, ME, MO, NC, NE, NH, NJ, NY, OH, PA, TN, TX, UT, and VA, Puerto Rico; and countries of: Canada, Barbados, Japan, and Korea. | Open, Classified |  |  | 195660
Z-2919-2018 | 80684 | No | LKO | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | is a quantitative/qualitative enzyme-linked immunosorbent assay (ELISA) for the detection of autoantibodies specific for Anti-Ro (SS-A) antigen in human serum or plasma. EDTA, lithium heparin, citrated plasma. It is intended to aid in the diagnosis of systemic rheumatic diseases, particularly SjÂżgren's syndrome and is not definitive in isolation. Autoantibody levels represent one parameter in a multi-criterion diagnostic process. |  |  | October 21, 2016 | 2016 | Euro Diagnostica AB | Incorrectly labeled Manufactured by DiaSorin, The correct term should be Distributed by DiaSorin. | Error in labeling | none |  | 4 | 4 | MN | Open, Classified |  |  | 20
Z-2918-2018 | 80689 | No | NHX | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | Immunoscan CCPlus test kit is an enzyme-linked immunosorbent assay (ELISA) for qualitative and semiquantitative determination of IgG antibodies to Cyclic Citrullinated Peptides (CCP) in human sera. |  |  | February 14, 2016 | 2016 | Euro Diagnostica AB | Decreasing performance overtime. | Nonconforming Material/Component | None |  | 25 | 25 kits (US only) | Minnesota and TexasForeign: Iran | Open, Classified |  |  | 25
Z-2923-2018 | 80554 | No | NHG | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cortisol Saliva Luminescence Immunoassay |  |  | January 12, 2018 | 2018 | Tecan US, Inc. | Product is not meeting specification for traceability to NIST cortisol standard, generating 1.3-1.8 times higher cortisol values with laboratory samples in comparison to NIST cortisol standard. | Nonconforming Material/Component | The firm, IBL International, sent an "Urgent Field Corrective Action" letter dated 1/12/2018 via registered mail to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to examine your stock and affected lots; contact all customers and provide the letter; if you have any of the affected product, stop using and quarantine/scrap all remaining product from lots LCL129 up to LCL 133. Review results generated with affected lots and follow your laboratory standard operational procedures to review any potential aberrant results that may have occurred; complete and return the Customer Response Form by fax to:1-919-287-2961; email to:QA@tecan.com or mail to: Attention: Tecan QA, 9401 Globe Center Dr., Suite 140, Morrisville, North Carolina 27560,United States to confirm receipt and actions.The firm is expect to have a new lot of Cortisol Saliva LUM available to ship by the end of February 2018.if you have any questions, please contact your local Tecan Helpdesk or IBL or Sr. Director Global Regulatory Affairs at 919-943-7118 or email: Laura.nea@tecan.com. |  | 870 | 870 (USA) 180 (OUS) | Worldwide distribution; US distribution to states of: GA, IA, LA, MA, MN, OR, PA, SC, and UT; and countries of: Germany, Finland, Switzerland, France, Austria, Denmark, Poland, Norway, Belgium, Netherlands, Italy, Canada, South Africa, Russia, and Japan. | Open, Classified |  |  | 870
Z-2927-2018 | 80516 | No | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | RAPIDPoint 405 Blood Gas Analyzer; 10282093, 10310464, 10314817, 10317193, 10318999, 10320055, |  |  | July 16, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | Therapeutic levels (1mg/mL and 2mg/mL) of Hydroxocobalamin may interfere with tHb and some of the CO-Ox fractions that are reported on the RAPIDPoint and RAPIDLab Blood Gas systems indicated. This may cause lower than expected values for carboxyhemoglobin (fCOHb) and methemoglobin (fMetHb). A negative interference with fCOHb has the potential to alter the medical assessment of the patient and may withhold necessary follow-up treatment and/or initiate cessation of medication in response to elevated fMetHb levels. | Device Design | The firm, Siemens Healthineers, sent an "Urgent Field Safety Notice" letter dated 7/2018 to its customers via FedEx on 7/18/18. The recall communication alerts customers that therapeutic levels of Hydroxocobalamin may interfere with tHb and some of the CO-Ox fractions that are reported on the RAPIDPoint and RAPIDLab Blood Gas Systems identified in the letter. Customer are advised to review the letter with their Medical Director, and complete and return the effectiveness check form within 7 days. Customers may continue to use their devices. Siemens is not recommending a review of previously generated results.The firm stated that they will be revising the Operators Guides for the 500 and 1200 series Operator's Guides with information on the interfering substance. The RapidPoint 405 Operator's Guide will not be revised due to the discontinuation of the manufacture of this analyzer. However, users should remain aware of the issue with hydroxocobalamine interference.Customers with questions were directed to contact their local Siemens technical support representative or Safety Officer (POC) at 800-441-9250 or email to: steven.andberg@siemens-healthineers.com. |  | 2248 | 2248 | Nationwide distribution. Worldwide foreign distribution. | Open, Classified |  |  | 14942
Z-2926-2018 | 80562 | No | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | Stanbio AC power adapter for the HemoPoint H2 meter, Model G3000 series. The firm name on the label for the HemoPoint H2 shows Manufactured by EKF-diagnostic GmbH, Germany. |  |  | February 08, 2018 | 2018 | Stanbio Laboratory, LP | The US prong (type A) adapter plate may crack, break, or detach and remain in the electrical outlet exposing live electrical contacts. | Nonconforming Material/Component | The recall was initiated via letter dated 2/7/2018 by first class mail letter on 2/8/2018 instructing the consignee to stop using the adapter if it is affected and it will be replaced. |  | 2804 | 2,804 devices | Distribution was nationwide. There was military distribution and no government distribution. Foreign distribution was made to Bolivia, Malaysia, Mexico, Papua New Guinea, Paraguay, Philippines, and Suriname. | Open, Classified |  |  | 2804
Z-2924-2018 | 80621 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Stryker Mako Vizadisc Knee Procedure Tracking Kit |  |  | July 24, 2018 | 2018 | Mako Surgical Corporation | Specific lots of the Vizadisc Knee Procedure Tracking Kit and Vizadisc Hip Procedure Tracking Kit have the potential to be damaged on the Vizadisc reflective material causing an inability to be detected by the camera. | Under Investigation by firm | On about 07/24/2018, Stryker sent letters to all their consignees via URGENT MEDICAL DEVICE REMOVAL letter. Instructions for risk mitigation included to perform a visual inspection the discs during pre-surgery setup as the defect can be recognized due to apparent discoloration and in the event that a Vizadisc is damaged, risk may be mitigated by testing the Vizadiscs by attaching them to the probes and arrays during pre-surgery setup, as per Stryker's recommendations in "Surgical Technique Guide." This provides an opportunity for the user to show each instrument to the camera after assembly to confirm that all discs are detectable. Additional instructions included to inform all individuals of the recall within the organization, complete and return the enclosed Recall Notification Business Reply Form, and arrange for to return all affected devices to Stryker in Mahwah, NJ.On about 11/27/2018, Stryker initiated an expansion of the recall and notified customers via "UPDATE URGENT MEDICAL DEVICE REMOVAL" letter. Instructions are the same as in the original notification letter. |  | 19906 | 19,906 total devices | All devices were distributed in the US in the following states: AK, AL, AR, AZ, CA, CO, CT, FL, IA, ID, KS, KY, LA, MA, MD, MI, MN, MO, NE, NC, NH, NJ, NY, NV, OH, PA, RI, SC, SD, TN. TX, UT, VA, WA, WI. | Open, Classified |  |  | 19906
Z-2931-2018 | 80431 | No | NAO | Gamma-glutamyl transpeptidase and isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | N Latex CDT Kit |  |  | May 31, 2018 | 2018 | CSL Behring GmbH | Siemens Healthcare Diagnostics has observed a negative bias for the carbohydrate-deficient transferrin (CDT) measurement when using the affected N Latex CDT lots in comparison to the HPLC method.CDT results in absolute concentrations may be influenced by patients transferrin levels and, therefore, results are reported as ratio of CDT to total transferrin, called %CDT.As the transferrin determination with N Antiserum to Human Transferrin is not affected, the observed effect leads to calculated %CDT values that show a negative bias of approximately 15% compared to the HPLC method. This could result in a shift of weak positive patient samples into the reference range of 1.19  2.47 %CDT which was derived from a study population of healthy adults. There are no control or calibrator materials that would catch this bias. | Under Investigation by firm | On May 31, 2018 Siemens Healthcare Diagnostics issued Urgent Medical Device Correction notices (UMDC) PP18-002.A.US to the 4 affected customers in the U.S. Actions to be Taken by the Customer are as follows:Siemens Healthcare Diagnostics has assigned optimized values to the affected lots of standards and controls included in the kit lots. Those lots of standards and controls can be used with the revised assigned values and acceptance ranges.As per Urgent Medical Device Correction (UMDC) PP18-001.A.US, customers were instructed to: - Recalibrate your CDT method with N CDT Standard lots listed in the UMDC letter, using the revised assigned values provided in the UMDC letter.- Use the revised assigned values and respective ranges of the N CDT Controls provided in the UMDC letter.- For BN ProSpec and Atellica NEPH 630 customers, a new Lot Data CD as well as new Secure Download Files are expected to be available in June 2018, including the revised CDT target values for N CDT Standard, N CDT Control 1 and N CDT Control 2.- Please make sure that the values you manually entered according to this information are not overwritten by using an older Lota Data CD/Secure Download File afterwards.Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter and the Urgent Field Corrective Action letter within 30 days. |  | 6635 | 6635 units | Domestic - DE, IN, KS, MO &amp; VAInternationally - AustraliaAustriaBelgiumCanadaChinaCzech RepublicFinlandFranceGermanyHungaryIndiaItalyJapanLatviaLithuaniaNetherlandsNorwayPolandPortugalSaudi ArabiaSlovakiaSloveniaSouth AfricaSouth KoriaSpainSwedenSwitzerlandTaiwanTurkeyUnited KingdomUruguay | Open, Classified |  |  | 6635
Z-2942-2018 | 80581 | Yes | MPB | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DUO FLOW DOUBLE LUMEN CATHETER, XTP114CT, XTP114MT, ; |  |  | July 12, 2018 | 2018 | Medical Components, Inc dba MedComp | The suture wings have the potential to crack or break. Potential patient injury includes catheter dislocation, bleeding, and air emboli. The firm has received six complaints (29 broken suture wings) to date for malfunction, without serious injury or death. | Process control | On July 12, 2018, the firm sent a Product Recall letter to its customers. Customers were provided with a full list of product codes and lot numbers and informed that those items had been recalled due to a potential for the suture wings to crack or break.Customers were informed that implanted catheters with intact suture wings do not need to be removed. However, an alternative method of catheter securement should be initiated for all implanted catheters, due to the potential for the suture wings to crack or break.Customers were instructed to return all unused recalled product to Medcomp. If product was further distributed, customers were asked to notify their sub-accounts.Contact your customer service representative for a Returned Goods Authorization (RGA)number if necessary. The phone number to call is 215-256-4201. Customers are asked to complete and return a customer response form included with the recall letter by fax (215-256-9191) or email (complaints@medcompnet.com). |  |  |  | Nationwide distribution; worldwide distribution. | Open, Classified
Z-2943-2018 | 80624 | No | FMH | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | Bard Midstream Urine Collector with Rigid Funnel and Iodophor Prep Pads. Product Packaging: Three (3) PVP-I prep pads are packed into each urine collection kit carton along with a pre-connected drainage funnel, specimen container, funnel lid, specimen lid and a room label. |  |  | March 21, 2018 | 2018 | Bard Medical Division | Medline notified BMD that the Aplicare Povidone Iodine (PVP-I) Prep Pads were not meeting the iodine assay level requirements to support 36 month expiration dating. | Unknown/Undetermined by firm | On March 21, 2018 Bard medical Division sent all their customers a recall letter with the following instructions: "Immediately examine your inventory and quarantine product subject to this recall. Please refer to the affected reorder number / lot list in Attachment 1 to help you locate the product. Do not use or further distribute any affected product."Please complete and return the accompanying Recall &amp; Effectiveness Check Form attached to this letter regardless of whether or not you have any of the affected products. Additional instructions for product return are contained on the form."If you have further distributed any units from the affected lots, please identify your customers / end users and notify them at once of this product recall by forwarding a copy of this letter. Your notification should include a copy of the this notification and the accompanying enclosuresIf you have affected product, please contact the Recall Coordinator via phone at 1-800-793-8110 or email BMD.FieldAction@crbard.com to obtain a recall authorization number (RCL #). |  | 131950 | 131,950 | AR, CA, CO, FL, GA, IA, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NY, SC, TN, TX, UT, and WA, | Open, Classified |  |  | 131950
Z-2944-2018 | 80617 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche cobas p 612 pre-analytical system (63x) |  |  | July 16, 2018 | 2018 | Roche Diagnostics Corporation | Sample material may potentially come in contact with the pipetting nozzle during operation of cobas p 612 pre-analytical system (63x) (LCP1) with non-filter tips. The possible presence of biological material on the nozzle may cause contamination resulting in potential false positive results, depending on the sensitivity of the analytical technology. | Under Investigation by firm | The firm, Roche, issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated 7/6/18 to its customers on 7/6/18. The letter described the product, problem and actions to be taken. The Customers were instructed to do the following: Complete all sections of the enclosed faxback form and fax it to number 1-866-469-2231 or email it to Roche6840@stericycle.com." Ensure you have at least a two-week supply of disposable filter tips before the Roche FES arrives to modify your system." Provide copies of this Urgent Medical Device Correction (UMDC) to other clinicians who may need to be aware of this issue." If your facility has distributed the affected product to another site, please ensure this UMDC is provided to that site." File this UMDC for future reference.If you are using cobas p 612 pre-analytical (63x) (LCP1) systems for aliquoting in combination with non-filter tips, you must use filter tips until the root cause is identified and appropriate corrective actions have been implemented.Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 if you have questions about the information contained in this UMDC. |  | 12 | 12 units | US Distribution to states of: CA, NY,TX, WA, and WI. | Open, Classified |  |  | 12
Z-0202-2019 | 80501 | Yes | OAD |  |  | Cardiovascular | PMA | Boston Scientific Maestro 4000 Cardiac Ablation System, Firmware Version 5.14, Model Numbers: M00440000, M0044000D0, M0044000H0,M0044000R0, M0044000Z0, |  |  | June 20, 2018 | 2018 | Boston Scientific Corporation | Reports of unintended continuation of radiofrequency (RF) energy delivery from the Cardiac ablation system after the foot switch has been released. |  | On 06/20/2018, a system notice was sent to customers, informing them that releasing the foot switch may not stop RF delivery. If RF delivery does not stop as intended, press and release the RF Power Control button. Customers were advised to report occurrence and clinical outcome of unintended RF delivery to the firm:United States (1-800-949-6708), Europe, Middle East, and Africa (+800-5555-7707), Asia Pacific (+65-6418-8878). Notification outside of the U.S. started 06/21/18. |  | 793 | 793 | U.S.:PA, IL, NY, CA, SC, AR, SD, AZ, KY, FL, TX, MA, MI, OH, LA, VA, NC, WA, DE, IN, TN, NJ, CT, AL, ME, GA, NH, MN, MD, DC, OK, HI, WI, CO, MO, IA, GU, NE, OR, ID, UT, RI, KS, NV, MS, WVForeign (OUS): Canada, Albania, Austria, Belgium, Bulgaria, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Greece, Hungary, Iran, Ireland, Israel, Italy, Jordan, Kuwait, Lithuania, Namibia, Netherlands, Lebanon, Pakistan, Poland, Portugal, Russian Fed., Saudi Arabia, South Africa, Spain, Sudan, Sweden, Switzerland, Turkey, Unit.Arab Emir., Argentina, Australia,Bangladesh, Chile, China, Colombia, Japan, Malaysia, Mexico, Nicaragua, Panama, Philippines, Sri Lanka, Taiwan, Thailand, Vietnam | Open, Classified |  |  | 793
Z-2945-2018 | 80688 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Ceterix Disposable Knot Pusher / Suture Cutter; NovoCut Suture Manager. Model# CTX-C001. orthopedic arthroscopic accessory. |  |  | May 16, 2018 | 2018 | Ceterix Orthopedics, Inc. | Possible nonconforming blades may result in a higher than specified cut force required to cut the suture. | Nonconforming Material/Component | Ceterix Orthopaedics notified customers via an "Urgent Field Safety Notice: for NovoCut Suture Manager. The letter dated May 16, 2018 provided the following: Details on affected devices, Description of the problem, Risks to Health, Actions to be taken by the customer, and Contact reference person. |  | 96 | 96 units | United States distribution. | Open, Classified |  |  | 96
Z-2947-2018 | 80656 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Needle Guide, REF 676-150, QTY 24, STERILE EO |  |  | July 30, 2018 | 2018 | Civco Medical Instruments Co. Inc. | The assembly of the needle guides were found to be either missing adhesive or did not have adequate adhesive to hold the assembled cannula in place on the needle guide. | Process control | The firm initiated communication on 07/30/2018 by email and telephone calls. The firm requested return of the recalled product.Questions may be made to the Account Representative by calling CIVCO Medical Solutions, Multi-Modality Imaging at 800-445-6741 or 319-248-6757 7:00 AM - 5:30 PM CST or by email at order@civco.com |  | 3 | 3 units | United States, Austria, Belgium, Denmark, France, Germany, Poland, Spain, Switzerland and United Kingdom | Open, Classified |  |  | 110
Z-2904-2018 | 80371 | No | GWO | Preformed alterable cranioplasty plate. | Neurology | Neurology | 510(k) | QUIKFLAP, 3 X 2-HOLE DOG BONE PLATES WITH 1.5X4MM SELF DRILLING SCREWS, Part Number 12-01530S, for use on the cranium |  |  | June 22, 2018 | 2018 | Stryker Leibinger GmbH &amp; Co. KG | After a three year term, there is a potential for the external packaging (sterile barrier) to become compromised. | Package design/selection | URGENT MEDICAL DEVICE RECALL letters dated 6/22/18 were sent to customers with the following requested actions to be taken:1. Please inform users of this Medical Device Removal and pass this notice to all appropriateindividuals within your organization.2. Return all affected products available at your location to:Stryker CMFAttn: Recall Coordinator6300 Sprinkle RdPortage, MI 49002REF: PFA #1739180 Or, Contact Stryker Customer Service and refer to PFA #1739180 with any questions about returning the product to us.3. Complete and sign the enclosed Business Reply Form and fax a copy to: (877) 648-7114 or email a copy to CMF-custserv@stryker.com.4. Keep a copy of the completed and executed Customer response Form for your records. |  | 22831 | 22,831 total | US distribution in the states: AZ, FL, NH, and OH, Australia, Canada, France, New Zealand, and UK. | Open, Classified |  |  | 22831
Z-2948-2018 | 80702 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Medline 0.9 % Sodium Chloride Injection, USP 5 mL in 10 mL Syringe ZR |  |  | February 01, 2018 | 2018 | Medline Industries Inc | Packaged with an incorrect overwrap. | Employee error | The firm, Medline Industries, Inc., sent an "URGENT RECALL IMMEDIATE ACTION REQUIRED" letter dated February 1, 2018 to their consignees with the required actions to be followed; 1. Immediately check your stock for item number EMZ10051240 and the affected lot number, 3131591, and quarantine all affected product.2. Please return the completed enclosed destruction form listing the quantity of affected product on hand. Even if you do not have any affected product please complete and return the form, as indicated on the destruction form.When we receive your completed destruction form your account will receivereplacement(s).3. If you have transferred possession of this product to another individual, to another department or another location in your organization you are required by FDA regulations to notify them of this recall communication. Have the customers return any affected product to you for collection and return to Medline Industries, Inc.Should you have any questions, contact the Recall Coordinator at 1-866-359-1704. |  | 90870 | 90,870 units | US Nationwide Distribution to states of: CA, CT, DE, FL, GA, IN, LA, MA, MD, MN, MS, NC, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, VT, WA and WV | Completed |  |  | 90870
Z-2949-2018 | 80666 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | P.F.C. SIGMA Revision Knee System Distal Augment, Model Number 960830 |  |  | July 18, 2018 | 2018 | DePuy Orthopaedics, Inc. | This unit may be missing the screw/collet Assembly | Employee error | The notification letter distributed in July 2018 instructed the customer to perform the following actions:" Please cease using the affected device immediately." Medical facility to return affected implant immediately to their Sales Consultant." Review this notice and complete the Acknowledgement section (Attachment A) to signify that your facility has been informed of this recall. Return the completed Acknowledgement to your Sales Consultant within five (5) working days of this notice." Retain a copy of the completed Acknowledgement Form in your files along with this notice." Notify surgeons at your facility by providing them with a copy of this notice to ensure surgeons are aware of this recall." Share this notice with others in your facility who need to be informed. |  | 1 | 1 | The devices were not distributed in the United States.The devices were distributed to the following foreign countries: Great Britain. | Open, Classified |  |  | 1
Z-2817-2018 | 80619 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | bioMerieux VITEK 2 Gram Positive Cefoxitin Screen, VITEK 2 Gram Positive AST for Oxacillin, VITEK SYSTEMS General Susceptibility Card |  |  | January 26, 2018 | 2018 | bioMerieux, Inc. | Customer reports indicated an increase in the rate of non-detected MRSA in association with the VITEK 2 AST-P631 product. | Under Investigation by firm | bioMerieux sent an Urgent Product Correction Notice letter to their French customers on January 26, 2018. Customer letters in the Americas were sent August 21, 2018. The letters identified the affected product, problem, and actions to be taken. Customers should follow steps outlined in the letter and contact bioMerieux for any additional questions regarding these corrective actions. Note: the actions do not include removal of the product. For any questions or concerns, contact your local bioMerieux representative. |  | 13961320 | 13,961,320 US | Worldwide distribution. US nationwide including Puerto Rico, Bermuda, Brazil, Canada, China, Curacao, Costa Rica, Dominican Republic, Ecuador, France, Guatemala, Indonesia, Israel, Hong Kong, South Korea, South Africa, Mexico, Myanmar, Nicaragua, Peru, Philippines, Paraguay, Singapore, Thailand, El Salvador, and Vietnam. | Open, Classified |  |  | 13961320
Z-2957-2018 | 80757 | No | PMX | Absorbable hemostatic agent and dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | FLOSEAL Special Applicator Tips in cartons labeled in part, "Floseal SPECIAL APPLICATOR TIPS***6 x Malleable Tip Trimmable Tip***Baxter***." The pouches inside the cartons are labeled in part, "Floseal SPECIAL APPLICATOR TIPS***Baxter***"." |  |  | May 30, 2018 | 2018 | Baxter Healthcare Corporation | Baxter r recalled this device due to the use of an incorrect plastic formulation in the manufacture of the blue luer component of the FLOSEAL malleable tips. | Mixed-up of materials/components | Baxter Healthcare notified customers of the recall on 05/30/2018 via Urgent Product Recall letter to the Directors of Pharmacy. Instructions included to locate and remove all affected product from inventory, contact Baxter Healthcare Center for Service to arrange for return of affected devices, complete and return the enclosed Baxter customer reply form, and notify customers if further distributed. Adverse reactions to the product can be reported to FDA MedWatch Adverse Event Reporting program and/or Baxter Corporate Product Surveillance at 800-437-5176 between the hours of 8;00 AM and 5:00 PM Central Time, Monday through Friday. |  | 46908 | 46,908 pouches | Nationwide distribution to AK, AL, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. International distribution to Australia, Korea, Taiwan, Columbia. | Open, Classified |  |  | 46908
Z-2956-2018 | 80693 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE 9100c NXT anesthesia systems |  |  | July 24, 2018 | 2018 | GE Medical Systems China Co., Ltd. | On systems equipped with the Auxiliary Common Gas Outlet (ACGO), there is a potential for the auxiliary gas switch to unexpectedly switch from the ACGO position back to the circle position. On systems equipped with the Aux O2 + Air function the switch could unexpectedly switch from the Aux O2 + Air position back to the circle position. These situations could lead to hypoxia. | Nonconforming Material/Component | The firm issued a medical device correction notice on 07/24/2018. The notice explained the issue and provided safety instructions should the use observe a switch change. The units will be corrected. Questions or concerns may contact 1-800-437-1171 or the local service representative. |  | 401 | 401 units | Worldwide | Open, Classified |  |  | 963
Z-2953-2018 | 80716 | stains, hematology | KQC | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Sysmex XN-Series Fluorocell PLT, Catalog #CD994563. Distributed in the U.S.A. by Sysmex America, Inc., Lincolnshire, IL. |  |  | July 27, 2018 | 2018 | Sysmex America, Inc. | The fluorescent platelet counts (PLT-F) and immature platelet fraction (IPF) values can be falsely low when using affected lots of Fluorocell PLT. | Nonconforming Material/Component | The firm, Sysmex, issued a "Product Notifications" letter dated 7/2018 via U.S. mail on 7/27/2018 to the consignees who received the affected lots. The letter was also issued to all other XN-Series customers with registered email addresses with the Sysmex Customer Resource Center via email blast on 7/27/2018 who did not directly receive the affected lots but could have received them through other channels. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1. follow your internal procedures for evaluation of flagged results.2. Distribute this Product Notification to the appropriate personnel to take the following actions:3. Upon receipt of unaffected lots of replacement reagent, immediately exchange any affected Fluorocell PLT cartridges currently in use on your analyzer.4. Discard any remaining affected lots in inventory according to your laboratorys protocols and procedures.5.Please proceed to the URL: http://pages.sysmex.com/fluorocellrecall.html to complete the Fluorocell Recall Response Form. Confirm receipt and understanding of the Product Notification. Sysmex will replace the quantity of affected lot(s) identified in this Product Notification that are confirmed to be currently in your inventory.If you have any questions concerning this information, please contact Sysmex Quality Assurance Department at quality@sysmex.com.br. |  | 5168 | 5,168 units | Worldwide Distribution: US (Nationwide) including Puerto Rico; and countries of: Argentina, Brazil, Canada, Chile, Costa Rica, Ecuador, Honduras, and Nicaragua. | Open, Classified |  |  | 5168
Z-2958-2018 | 80804 | No | OLW | Electroencephalograph. | Neurology | Neurology | 510(k) | Li-ion battery packs used in BIS Vista and BIS View Monitoring Systems. |  |  | July 26, 2018 | 2018 | Covidien LLC | Revised replacement instructions for the Li-ion battery packs used in BIS Vista and BIS View Monitoring Systems | Device Design | Medtronic U.S. consignees were notified via Federal Express or certified mail on July 26, 2018, and the letter informs customers to inspect their batteries. Batteries with date code 3114 or older (Week 31 of 2014) exceed 4 years of age and should be removed from the monitor and disposed of according to facilitys standard process. Included with this notification are stickers to be placed on BIS" monitors when the battery pack is removed the stickers No Battery Installed should be placed on the monitors. The BIS Vista and View monitoring systems may be used without a battery pack only if connected to AC power source with emergency power backup system (Red Outlet). |  | 6937 | BIS 46,961; Spare Batteries: 6,937 | NationwideForeign: Australia, New Zealand | Open, Classified |  |  | 6937
Z-2951-2018 | 80740 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | HaemoneticsÂż OrthoPATÂż System-OPAT QuickConnect Reservoirs. |  |  | July 13, 2018 | 2018 | Haemonetics Corporation | OrthoPATÂż reservoirs are missing the cannula inside the reservoirs. The absence of this cannula prevents the emptying of blood from the reservoirs. The inability to empty the reservoir means that the collected blood cannot be processed. | Process control | The firm, Haemonetics, notified customers by an email titled "URGENT: MEDICAL DEVICE RECALL" on July 13, 2018 . The email described the product, problem, and actions to be taken. The customers were instructed to inspect product for defect and return units of the affected lots for replacement with non-affected units; complete the attached customer acknowledgement form via FAX TO +1-781-356-3558 OR SCAN AND E-MAIL TO: CORPORATEREGULATORY@HAEMONETICS.COM.You can contact your local customer service representative at: North America 1.800.537.2802 or email:customerservicena@haemonetics.com. |  | 348 | 348 units | US Nationwide Distribution | Open, Classified |  |  | 780
Z-2878-2018 | 80720 | No | NGV |  |  | Cardiovascular | PMA | WaveCrest Left Atrial Appendage (LAA), Occlusion System (32 mm), Catalog Numbers WCR1523 and WCR2523. |  |  | May 07, 2018 | 2018 | Biosense Webster, Inc. | When attempting to recapture the 32-mm device, the tip of the delivery sheath may fold or buckle, resulting in increased retraction force and difficulty or failure to recapture the device. | Component design/selection | Coherex Medical notified customers via "URGENT FIELD SAFETY NOTICECoherex WaveCrest Left Atrial Appendage (LAA) Occlusion System (32 mm)Catalog No: WCR1523, Lot Number: All" letter dated May 7, 2018.The letter provided an overview, and important information about the product description and codes, lots subject to removal, why customers were being contacted, What actions are required/instructions, available assistance/contact information, and Reply Form. |  | 25 | 25 units | No U.S. Distribution. International Distribution only. | Open, Classified |  |  | 25
Z-2968-2018 | 80723 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology Hemo, Release SW version 13.0. |  |  | March 12, 2018 | 2018 | McKesson Israel Ltd. | Change Healthcare has identified an issue where, under certain circumstances, the Real Time Monitor (RTM) may not display physiological signals. | Software Design Change | The firm, Change Healthcare Israel Ltd., sent an "Urgent Field Safety Notice" dated on 03/12/2018 to its customers. The notice explained the product, issue and actions to be taken. The customers were instructed to: distribute the notice to all those within your organization who need to be aware; alert other parties affiliated with your organization, which are affected by this field safety notice;and maintain awareness of this notice and resulting action until the defect has been corrected, to ensure effectiveness of the corrective action.Change Healthcare has developed software updates to address the problem.If you have any questions, contact Change Healthcare Customer Support (US) 1-800-654-4366 or (UK) 440(0) 208-952-799 to arrange for the installation of software update/upgrade. |  | 31 | 31 units | US Distribution to states of: CA, CO, CT, FL, GA, KS, IN, LA, NJ, NY, OK, OH, SC, and TX; and internationally to: UK. | Open, Classified |  |  | 31
Z-2959-2018 | 80805 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO V-PRO 1, V-PRO 1 Plus, and V-PRO maX Low Temperature Sterilization Systems |  |  | July 26, 2018 | 2018 | Steris Corporation | There is the potential for a scenario where a punctured cup of sterilant expires while the system is powered down, upon power-up the unit will not detect the expired cup of sterilant; allowing for subsequent cycles to be run. | Software design | The firm, Steris, sent an "Urgent Device Field Correction" letters dated 7/25/18 were distributed to customers. The letters included the following instructions for customers:STERIS Action  The system software for V-PRO 1, V-PRO 1 Plus, and V-PRO maX will be updated to ensure the VAPROX HC Sterilant expiration date is detected regardless of whether the unit is powered on or off. STERIS Service Technicians will contact all affected Customers and update the system software in the affected V-PRO Sterilization Systems.User Action  In the interim of the software update, as is normal practice, V-PRO 1, V-PRO 1 Plus, and V-PRO maX operators should place the unit into Standby mode when not in use in lieu of fully powering down the unit. As is stated in the Operator Manual, the unit should only be shut down during unit service activities as required:"Control Power Disconnect Switch - Located behind Operating End (OE) lower access door. This switch is used to disconnect power to the Sterilization Unit. Under normal operation, this switch remains in ON position at all times and is only accessed during sterilization unit servicing. This switch is designed to enable lock-out for maintenance purposes."In addition, users should continue to review their cycle printouts upon completion of each cycle to verify the VAPROX Sterilant was not expired.If you have questions regarding this matter, please contact Associate Product Manager, at (440) 392-7140, STERIS Customer Service at 1-800-548-4873, or your local STERIS Representative. |  | 2889 | 2889 units | Worldwide Distribution: US(nationwide) and to countries of: Albania, Algeria, Australia, Austria, Bahrain, Bangladesh, Belgium, Belarus, Brazil, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Curacao, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, Hungary, Iceland, India, Indonesia, Iran (Islamic Republic of), Iraq, Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Malta, Mexico, Morocco, Nepal, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Suriname, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 2889
Z-2960-2018 | 80727 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Overhead Counterpoise Systems (OCS), OCS CONF,STD CEIL MNT, STD HORZ ARM, Catalog # OCS 115A, Material Number 59880253 |  |  | June 28, 2018 | 2018 | Bayer Medical Care, Inc. | The torque wrench used on the ceiling column assembly during a recent service visit was outside of the torque specification, potentially overtightening the bolts. The overtightening of bolts may compromise the structural integrity of the system which could lead to the OCS disengaging from the ceiling mount and this may cause injury to the patient or the operator. | Equipment maintenance | On June 28, 2018, Urgent Medical Device Field Safety Corrective Action letters were distributed to customers. The letters informed the customers of the following:A Bayer Service Representative will be scheduling a service visit to replace the ceiling column assembly of your OCS system(s) at no cost to you. |  | 2 | 2 | The devices were distributed to the following US states: OH, TX, and WV. | Terminated | November 26, 2018 | 151 | 7
Z-2963-2018 | 80609 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Outer Package Label: Tapered Screw-Vent Implant System, 4.1mm X 10mm, 3.5mm, REF:TSVM4B10, Sterile:R, Rx Only; Inner Package Label: Implant, MTX, 3.7mm x 10mm,3.5mm, REF: TSVB10. |  |  | May 10, 2018 | 2018 | Zimmer Dental Inc | Dental implants have incorrect item/lot numbers and size on the inner package labeling. Correct item/lot numbers and size are on the outer package labeling. | Labeling mix-ups | The firm, Zimmer Biomet, sent an "Medical Device Recall" notice dated 5/10/2018 via FedEx/UPS courier to its customers. OUS distributors received a "distributor" notice. The notice described the product, problem and actions to be taken. The Customers were instructed to immediately return affected product to Field Action, PM Regulatory Compliance, Zimmer Biomet, 4555 Riverside Dr., Palm Beach Gardens, FL 33410 US and asked to complete and return a Certificate of Acknowledgement. and send to CorporateQuality.PostMarket@zimmerbiomet.com. If you have further questions or concerns, please call the Recall Department at 561-776-6700 between 8:00 am and 5:00 pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a prompt to leave a voicemail. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 299 | 299 units | Worldwide Distribution: US (nationwide) in states of: AL, CA, CO, CT, FL, GA, ID, IL, MA, MD, MI, MN, MO, NH, NJ, NY, PA,TN, TX, UT, VA, and WA; and to countries of: Australia, Bulgaria, Canada, Croatia, France, Germany, Italy, Portugal, and Spain. | Open, Classified |  |  | 299
Z-2964-2018 | 80642 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DRAD-3000E (Radrex-i) TFP-4336W (Wireless FPD) |  |  | April 18, 2018 | 2018 | Canon Medical System, USA, INC. | It was found when an operator performs a radiography using the wireless flat panel detector (FPD), a message window was displayed on the monitor stating image transmission was not completed and there was no image. It also showed the OK button to reacquire image data form the FPD, and the Cancel button to cancel the reacquisition.When the operator selects the OK button, the same message window appears. The operator then repeated the same operation several times and finally selected the Cancel button to quit the reacquisition mode. | Software design | On 04/18/2018, the firm mailed Urgent: Medical Device Correction letters via USPS return receipt mail. The letter contains a return reply form that is to be faxed. After one month contact will be made with each non-responding location and a copy of the letter will be faxed to that location and the fax receipt will serve as proof of notification.Upon release of the modified software it will be assigned to a firm employee with instructions for implementation at each site. Each incident is assigned unique service incident (case) and the completion of the service incidents are tracked through completion. In the event that a customer refuses the installation the customer would be requested to sign a service case record stating that the corrected software was refused. |  | 43 | 43 | Domestic: AR, FL, GA, PA, IL, LA, MI, NC, NJ, NY, OH, TN, TX, UT, VT, WI, and WVForeign: Australia, Canada, The Netherlands, South Korea, and Malaysia | Open, Classified |  |  | 43
Z-2965-2018 | 80731 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Peg Fully Threaded, Locking, 2.5mm x 10mm, Item Number 131212610 Orthopedic device used for stabilization and fixation of small bone fragments |  |  | July 12, 2018 | 2018 | Zimmer Biomet, Inc. | A size 22mm threaded peg was found in the package of a 10mm unit. | Under Investigation by firm | The firm, Zimmer Biomet, sent an "Urgent Medical Device Recall " letter dated 7/12/18, to customers. The described the product, problem and actions to be taken.The letter instructed Risk Managers to perform the following actions:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com.Distributors were instructed to perform the following:Your Responsibilities Review this notification and ensure that affected team members are aware of the contents.1. Immediately locate and quarantine affected product in your inventory.2. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. Request a Return Authorization Number through zimmerbiometintlirarequests@zimmerbiomet.com for international returnsc. For each return, send a copy of Att |  | 25 | 25 units | International Distribution to countries of: Japan and Netherlands. | Open, Classified |  |  | 25
Z-2967-2018 | 80692 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Foundation Modular Femoral Stem, Part Number, 220-10-110. Knee prosthesis component. |  |  | May 14, 2018 | 2018 | Encore Medical, Lp | After receipt of a product complaint for the labeling, it was determined that this device does not have FDA clearance. | No Marketing Application | The firm, DJO, sent an "URGENT FIELD SAFETY NOTICE" letter dated 5/14/2018 to it's customers via email, on May 14, 2018 and contained the following information: Description of Issue, Quantity in Field, product information and code, Action required, and an Acknowledgement form with contact information. The customers were instructed to: check your inventory for part number 220-10-110; Customer Service has already requested that these be quarantined by your agency; Contact Customer Service to request an RMA for return of these devices. As there are no currently cleared 10x100 femoral stem extensions, there is no product available for replacement; and complete the attached acknowledgment form, even if you do not have the affected product, and email the form to sandra.busick@djoglobal.com by May 18, 2018.Modular Femoral Cemented Stem -- 220-10-110 -determined that this device does not have FDA clearance.If you have any questions, please contact Manager, Regulatory Affairs at (512) 834-6255 or by email at teffany.hutto@djoglobal.com. |  | 12 | 12 | U.S. Distribution to states of: AZ, MO, MS, NY, OK, and UT; internationally to: Japan. | Open, Classified |  |  | 12
Z-2952-2018 | 80517 | No | PCH | Gastrointestinal microorganism multiplex nucleic acid-based assay. | Microbiology | Microbiology | 510(k) | Verigene Enteric Pathogens Nucleic Acid Test (EP) for detecting and identification of pathogenic enteric bacteria, viruses, and genetic virulence markers from liquid or soft stool preserved in Cary-Blair medium, collected from individuals with signs and symptoms of gastrointestinal infection. |  |  | May 08, 2018 | 2018 | Luminex Corporation | Elevated false positive results Yersinia enterocolitica (Yersinia) from customers using the EP Nucleic Acid Test Cartridges component (20-006-023) from the Verigene EP Nucleic Acid Test Kit (20-005-023). | Under Investigation by firm | The firm notified customers on or about 05/08/2018 via a letter entitled "URGENT: MEDICAL DEVICE RECALL". The Luminex informed the customer reason why they are receiving a recall notice---elevated false positive Yersinia enterocolitica (Yersinia) results when using the EP Nucleic Acid Test Cartridges component (20-006-023) from the VerigeneÂż EP Nucleic Acid Test Kit (20-005-023). The letter described the intended use of the test kit and the distribution dates. The letter provided instructions for destroying the product if found in the inventory by a customer at a customer site. Luminex advised the customer to complete and return an enclosed Acknowledgment and Receipt form whether a customer had a recalled product on site or not. Luminex Technical Support are available assist in completing this form. Luminex Customer Experience Team contact information is as follows:1-877-785-2323 (U.S. and Canada) (24 hours/7days) +1-512-381-4397 (Outside U.S. and Canada) (24 hours/7days) support@luminexcorp.com |  | 3950 | 790 Cartridge Carriers containing 3160 Cartridges | Domestic Distribution: Arizona, Colorado, Connecticut, Florida, Georgia, Illinois, Indiana, Kansas, Michigan, Minnesota, Missouri, New Hampshire, Tennessee, Texas, Utah, West Virginia, Wisconsin. | Open, Classified |  |  | 3950
Z-2966-2018 | 80667 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Tibial Alignment Guide, REF 351-10-00 Vantage Fixed Ankle system components for ankle replacements. |  |  | July 13, 2018 | 2018 | Exactech, Inc. | Exactech is recalling the Tibial Alignment Guide because the Global Trade Item Number (GTIN) on the device is incorrect. The GTIN on the device reads 10885862525629, when it should read 10885862525628. | Employee error | The firm notified their direct consignees by email on 07/13/2018 and requested the return of the product. |  | 7 | 7 units | CA, OH, NY, OK, FL, TX, SC, Germany, Spain, Switzerland, and UK | Open, Classified |  |  | 7
Z-1220-2018 | 79198 | No | JNO | Trypsin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoDELFIA Neonatal immunoreactive trypsin(ogen) (IRT) kit, intended for the quantitative determination of human immunoreactive trypsin(ogen) (IRT) in blood specimens. (UDI: (01)06438147298716(17)190430(10)656479); Product Number: B005-212 |  |  | September 25, 2017 | 2017 | PerkinElmer Life and Analytical Sciences, Wallac, OY | The Quality Control Certificate for certain AutoDELFIA Neonatal IRT Kits is missing information required to assess the acceptability of measured control values which may result in a delay in reporting IRT results. | Process control | Actions to be taken: Inspect your inventory of AutoDELFIA Neonatal IRT kits for batch number 1065647901 and replace the Quality Control Certificate included in the kit packaging with the revised certificate attached to this letter. If you have already used the kit lot there is no need to retest the samples reported if the results of the controls included in the assay were within your established acceptance range. Complete and return the enclosed response form identifying the quantity of affected kits in your possession. Please inform others affected in your organization accordingly. To comply with regulatory requirements we require that you complete the enclosed Response Form and return it. |  | 240 | 240 | Distributed to CT and TX. | Terminated | September 06, 2018 | 346 | 240
Z-2969-2018 | 80372 | No | MNQ |  |  | Anesthesiology | PMA | Inspire II Model 3024 Implantable Pulse Generator (IPG) for Obstructive Sleep Apnea, Inspire Part #900-003-020, |  |  | June 12, 2018 | 2018 | Inspire Medical Systems Inc. | Incorrect use-by date on the device registration/patient file labels. | Error in labeling | The firm's sales team members called all of the affected implanting surgeons between 6/12-14/2018 to discuss the recall. The recalling firm issued "Dear Health Care Provider" letters dated 6/15/2018 via certified mail on 6/15/2018 to all implanting surgeons and hospital customers. The letter identified the affected product, problem and actions to be taken. For questions call 1-763-205-7974. |  | 93 | 93 devices | US Nationwide Distribution in the states of AL, AZ, CA, FL, GA, IL, KY, MA, MI, MN, MS, NY, OH, PA, TN, TX, VA, WA, and WI. | Completed |  |  | 93
Z-2973-2018 | 80496 | No | OHS | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Baby Quasar Plus, Light Therapy System, packaging, box, sleeve and user manual. Handheld Wrinkles Reduction Device. |  |  | June 29, 2018 | 2018 | Quasar Bio-Tech, Inc. | Quasar did not receive FDA clearance for claims of collagen/elastin production as stated on the packaging and in the user manuals. | Under Investigation by firm | Quasar Bio-Tech notified customers on about 06/20/2018 via URGENT PRODUCT RECALL letter. Instructions included place any current, affected inventory on hold and arrange for return to Quasar Bio-Tech, Inc. Customers were also instructed to complete and return the reply form. Questions or concerns should be directed to Quasar Bio-Tech, Inc., at 800-944-1523 or 941-306-5812. |  | 6783 | 6,783 units total | Nationwide distribution to CA, FL, GA, HI, IL, LA, MD, NE, NY, OH, TX, VA, WA, WI. Worldwide distribution to Australia, China, Netherlands, United Kingdom | Open, Classified |  |  | 6783
Z-2981-2018 | 80800 | No | HQD | Rigid gas permeable contact lens. | Ophthalmic | Ophthalmic | 510(k) | Zenlens Toric-Gas Permeable Contact lenses for Daily Wear in a 10-cc polycarbonate vial (Rx) as follows: |  |  | July 03, 2018 | 2018 | Alden Optical | Vials labeled for the prescriptions contained incorrect lenses | Process control | The firm, Alden Optical, notified customers of an "URGENT VOLUNTARY MEDICAL DEVICE RECALL" by phone and follow-up letter on 7/3/18. The letter described product, problem and actions to be taken. The customers were instructed to check your inventory; return the affected lenses via email with return shipping labels provided by email. Customers requested to notify patients if the lens has been dispensed or distributed, and have the patient return the lens(es), and complete the enclosed Business Response Card and return it to Alden Optical via U.S. Mail within 5 days. If you have any questions regarding this recall, please contact Alden Optical Customer Service at 1-800-253-3669. |  | 36 | 36 vials | US Distribution in the states of: CA, GA, IL, MD, NC, NJ, NV, OH, OK, PA, TN, TX and UT. | Open, Classified |  |  | 40
Z-2980-2018 | 78411 | [6 empty] | LASER CDS |  |  | April 12, 2017 | 2017 | Medline Industries, Inc. | Multiple surgical packs were assembled with a non-sterile CHG applicators labeled as "ChloraPrep" (kit component numbers: 29183, 52115, 82279, and 82281). The non-sterile components 29183, 52115, 82279, and 82281 should have been placed inside the sterile portion of the pack, where it would subsequently be exposed to the ethylene oxide sterilization process. However, the components were not placed in the sterile portion of the pack during assembly and were never exposed to the sterilization process. Components 29183, 52115, 82279, and 82281 were inadvertently released non-sterile. The remainder of the pack components were properly sterilized and only components 29183, 52115, 82279, and 82281 were affected. Components 29183, 52115, 82279, and 82281 are manufactured by Becton, Dickinson (BD) and are labeled "ChloraPrep". ChloraPrep" applicators contain 2% w/v chlorhexidine gluconate (CHG) in 70% v/v isopropyl alcohol (IPA). The applicators are used to facilitate skin preparation for various preoperative and pre-injection procedures. | Under Investigation by firm | Firm contacted customers via phone and letter beginning April 12, 2017. Customers were instructed to remove the Ethicon BioPatch Protective Disk with CHG from the kit(s) and return it to the firm. The firm would provide a sterile replacement BioPatch. Affected product will be destroyed. |  | 144 | 144 kits | CA, NM | Open, Classified |  |  | 3380
Z-2990-2018 | 80928 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Vivid E7 ultrasound system |  |  | March 22, 2018 | 2018 | GE Medical Systems Ultrasound &amp; Primary Care Diagnostics, LL | The system side of the power cord may break and expose the electrical conductors leading to the risk of shock. | Component design/selection | On March 23, 2018 GE Healthcare sent letters to their consignees informing them about the recall of LOGIQ 9, LOGIQ E9, Vivid 7, Vivid E7, Vivid E9, Vivid E80, Vivid E90, and Vivid E95 ultrasound systems which have been installed since April 2015 or which have received a power cord replacement since April 2015. Please ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. Your safety and the safety of your patients is of our utmost concern. SafetyIssue The system side of the power cord may break and expose the electrical conductors leading to the risk of shock. This can lead to injuries up to and including cardiac arrhythmia or cardiac arrest. A minor injury has been reported as a result of this issue. Safety Instructions You may continue to use your system. If you need to remove the power cord from the rear of the unit (see picture of location below), before doing so be sure to do the following:1. Turn the system off2. Unplug it from the wallIf the system end of the power cord (see picture below) is damaged, discontinue use of the system until a replacement cord is provided. If you experience any power issues with your unit, discontinue use of the system and contact GEHC Service. GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.Contact InformationIf you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 12426 | 12,426 (4,717 US; 7,709 OUS ALL PRODUCTS) | Nationwide and Puerto RicoFOREIGN: Algeria, Antigua and Barbuda, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Plurinational State of Brazil, Brazil, Brunei Darussalam, Bulgaria, Burkina Faso, Canada, Chile, China, Colombia, Congo, The Democratic Republic of the Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Ethiopia, Finland, France, Germany, Greece, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Luxembourg, Macedonia, The Former Yugoslav Republic of, Malaysia,Mexico, Morocco, Myanmar, Nepal, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, State of, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation,San Marino, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, and Viet Nam | Open, Classified |  |  | 12426
Z-2992-2018 | 80896 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | VidiStar(TM) PACS &amp; DICOM Viewer Software system |  |  | August 28, 2018 | 2018 | Hitachi Healthcare Americas Corp Informatics Division | The secure filesystem client software used in the interface between the Vidistar PACS and an EHR system may cause intermixed images from multiple patients showing in a single study. | Software in the Use Environment | The firm initiated the recall by email on 08/28/2018. The firm will be removing the 3rd party software from the interface and replacing it with new software and notify the affected customers when the action is completed. |  | 12 | 12 units | SC, CO, IL, AZ, MT, OH, IN, TX, GA, MA | Open, Classified |  |  | 12
Z-3017-2018 | 80813 | No | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | remel THIOGLYCOLLATE MEDIUM W/O IND. W/ DEXTROSE W/ HEMIN &amp; VIT.K |  |  | August 13, 2018 | 2018 | Remel Inc | Products have intermittently failed performance testing during stability studies with various anaerobic microorganisms. | Under Investigation by firm | The firm initiated the recall by letter on 08/13/2018. The firm directed the consignee to you inspect their stock and destroy any remaining inventory of the lots listed above and contact Customer Services or your local distributor regarding replacement. They also requested that the consignee review results and laboratory reporting associated with use of the listed lots and consider retesting and/or seek appropriate expert advice for further action.The firm instructed the consignee to further distribute the notice to all who need to be aware within your organization or to any organization where the potentially affected products have been transferred. For questions, please contact Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 2730 | 2730 units | US, Canada, Singapore | Open, Classified |  |  | 3162
Z-2993-2018 | 80826 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Diagnostica Stago STA NÂżoplastineÂż CI Plus, 10 ml vial. In-Vitro Diagnostic for determination of the prothrombin time (PT) in plasma. |  |  | July 24, 2018 | 2018 | Diagnostica Stago, Inc. | QC values outside of the assigned ranges (prolonged Prothrombin time / decreased PT %) | Device Design | Diagnostica Stago issued notification on July 24, 2018 to US Customers provided the reason for the notification, actions to be taken to mitigate the risk. Action: If you have in your laboratory, any of the lots among those listed in the appendix, we are asking you to: If it is not already done, run a Quality Control test at every change of vial.Return to your Stago affiliate, by fax or e-mail, the enclosed form completed and confirming that you have read this letter.Questions contact Stago Hotline (1-800-725-0607) |  | 24043 | 24043 kits | Nationwide | Open, Classified |  |  | 24048
Z-3020-2018 | 80204 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Wingspan Stent System, Rx Only, Sterile EO, Catalog #: M003WE0250090, M003WE0300090, M003WE0350090, M003WE0400090, M003WE0450090, |  |  | June 05, 2018 | 2018 | Stryker Neurovascular | Use-related coating damage on delivery catheter shafts of Wingspan delivery catheters were reported | Use error | On 06/05/2018 Urgent Medical Device Voluntary Customer Notification letters were sent to consignees via courier restating information present in the Directions for Use and additional information. Consignees were instructed to complete Acknowledgement Forms acknowledging receipt and review of the customer notification. |  | 11585 | 11585 | China | Open, Classified |  |  | 11585
Z-2995-2018 | 80785 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Portable Suction Pump, (a) Style 758010, Model CARE-E-VAC 3, Ohio Medical LLC; (b) Style 758000, Model CARE-E-VAC 3, Ohio Medical LLC |  |  | August 14, 2018 | 2018 | Ohio Medical Corporation | The care-e-vac(R)3 Battery/AC Powered Portable Aspirator battery terminal may come in contact with the pump causing an electric short leading to damage to wiring and resulting in loss of function or an inoperable unit under battery power. | Nonconforming Material/Component | The firm, Ohio Medical, sent an "URGENT  MEDICAL DEVICE CORRECTION NOTIFICATION" letter dated 8/7/2018 to its customers on 8/14/2018. The letter described the product, problem and actions to be taken. The customers were instructed to take the affected care-e-vacÂż3 units out of service and follow the instructions, site: https://www.ohiomedical.com/4897450, provided with this notification to change the orientation of the battery; If, upon opening the device, you identify damage to wiring, please contact Ohio Medical Customer Service at 866.549.6446 or email to customer.service@ohiomedical.com for additional instructions to replace the wiring harness; and complete the attached Customer Acknowledgement Form and return it to our Quality Assurance Department at 847/855-6304 (fax), email to: CEV3_FCA@ohiomedical.com, or mail in the envelope provided to: Ohio Medical LLC, ATTN: Quality Assurance Department, 1111 Lakeside Drive, Gurnee, IL 60031.Please direct questions to Sr. Director, QA/RA at 847-855-6326 or email: barbara-ann.conway-myers@ohiomedical.com. |  | 1443 | 1443 units | US Nationwide Distribution and internationally to countries: Canada, Saudi Arabia, and United Arab Emirates. | Open, Classified |  |  | 1443
Z-3023-2018 | 80529 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | fMRI Hardware System VisualSystem allows video signals from the stimulus presentation PC to enter the shielded scanner room and presented by a set of coil-mounted displays or by an in-room LCD monitor. - LCD Monitor Power supply (LPS) for LCD Monitor, Versions 1.0 and 1.1 |  |  | August 20, 2018 | 2018 | NordicNeuroLab AS | Complaints of this version of the power supply unit (LPS) (1.0 and 1.1) being dragged into the magnet bore (MR Scanner). | Labeling design | The firm sent a letter explaining the issue by email to the consignees on 08/20/2018. The firm requested that the consignee check the stock without removing the device from the magnet. It is was not in the magnet room, safely quarantine the device, and respond to the recall notice. The firm also directed any consignee who further distributed the product to notify those entities and to contact NNL. NNL will follow-up with them directly. NNL will upgrade all existing models with a safer design of the LCD Power supply unit. |  | 163 | 163 units | worldwide | Open, Classified |  |  | 163
Z-3022-2018 | 80831 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | ev3 TurboHawk PERPHERAL PLAQUE EXCISION SYSTEM FOR SURGICAL USE, LARGE VESSEL XTENDED TIP FOR CALCIUM, Model THS-LX-C for peripheral vascular use. |  |  | August 10, 2018 | 2018 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | The Distal Flush Tool (DFT) accessories that were packaged in this lot number are not compatible with the TurboHawk LX-C catheter. This may result in the inability to remove excised material from the tip of the catheter. | Process control | Medtronic began notifying the Consignees on 10-Aug-2018. Consignees were asked to quarantine and return affected product. |  | 67 | 67 units | Distribution to US states of KS, CA, FL, GA, AZ, MI, LA, TX, NE, KY, and China | Open, Classified |  |  | 67
Z-3021-2018 | 80765 | No | MCX |  |  | Cardiovascular | PMA | CSI, Orbital Atherectomy System (OAS) Saline Infusion Pump, Rx only, Part Number: 7-10014, Model Number: SIP-3000. |  |  | April 13, 2018 | 2018 | Cardiovascular Systems Inc | The pumps may switch to stand-by during use requiring the pump to be reset prior to continuing treatment. | Device Design | Consignees were sent a CSI "Urgent Medical Device Recall" letter dated 13 APRIL 2018. The letter listed the Affected Product, Recall Description, SIP Replacement and Return &amp; contact information. Consignees were notified the pumps will be replaced. Consignees may continue to use their pumps however if a yellow light fault is observed and the troubleshooting per the IFU does not resolve the fault, discontinue use of the pump and notify CSI. |  | 18 | 18 | US Nationwide Distribution in the states of AZ, CA, FL, GA, IL, MO, and TX | Open, Classified |  |  | 18
Z-3026-2018 | 80527 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | VTI 8 MHz Surgical Doppler Transceiver Selectabable Channel |  |  | June 19, 2018 | 2018 | Vascular Technology, Inc. | Device maybe mislabeled as a 20 MHz Doppler transceiver instead of an 8 MHz Doppler transceiver | Labeling Change Control | The firm, Vascular Technology, Inc., "URGENT: MEDICAL RECALL DEVICE" letter dated June 18, 2018, to its customer on June 19, 2018. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine inventory and return all product subject to recall. In addition, if you have further distributed this product, please identify your customers and notify them at once of this product recall.If you have any questions, contact VP Sales person at 603-594-9700. |  | 10 | 10 units | US Distribution to state of: CA. | Open, Classified |  |  | 10
Z-3025-2018 | 80696 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci X/Xi EndoWrist 8mm Monopolar Curved Scissors Model No. 470179-14 |  |  | August 15, 2018 | 2018 | Intuitive Surgical, Inc. | Monopolar Curved Scissor (MCS) instruments may be susceptible to developing cracks near the proximal edge of the extension tube. These cracks may fracture the plastic surrounding the instrument shaft, and may cause fragments of the extension tube to fall, including potentially into the patient. | Use error | On August 15, 2018, the firm sent an Urgent Medical Device Notification to customers via Federal Express overnight service informing them that certain lots of monopolar curved scissors may be more susceptible to developing cracks near the proximal edge of the extension tube that may cause fragments of the extension tube to fall, including potentially into patients. Customers were advised of the following 1) Always inspect the instruments prior to use for any damage; 2) Use of an instrument for a task other than the instruments designed use may result in breakage or damage; 3) During reprocessing, make certain that the MCS has been adequately rinsed and flushed, and all cleaning chemicals have been removed. Customers were instructed to read and complete actions on the firms website or to fill out and return Acknowledgement Forms to the firm. Customers with questions are encouraged to call Customer Service: - North and South America: 800-876-1310 Option 3 (4 am to 5 pm PST) - Japan: 0120-56-5635 or 03-5575-1362 (9 am to 6 pm JST) - South Korea: 02-3271-3200 (9 am to 6 pm KSTJ) - Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 am to 6 pm CET)Field Corrective Actions: All customers who received a-16 version of S/Si MCS instrument and the -14 version of X/Xi MCS instrument will receive a customer letter, as a notification of the issue and to reiterate guidelines for proper inspection and use of the MCS instrument. Preventive Actions: The -16 version of S/Si MCS instrument and the -14 version of X/Xi MCS instruments are no longer being shipped to customers. A new design has been introduced to prevent the extension tube from being susceptible to cracking. This new material in has been shown to be robust to environmental stress cracking. |  | 42274 | 42,274 (X/Xi) | Worldwide Distribution - US Nationwide in the states of :AL, AK, AZ, AR, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MA, ME, MD, MI, MN, MS, MO, NB, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WV, WI, and WY and the countries of :Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Finland, France, Germany, Guadeloupe, Iceland, India, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Luxembourg, Martinique, Mexico, Monaco, Netherlands, Norway, Pakistan, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Venezuela | Open, Classified |  |  | 65338
Z-3027-2018 | 80317 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Aortic Perfusion Cannula, 24F x 30CM, resterilized. |  |  | March 07, 2018 | 2018 | SPS Sterilization, Inc | The expiration date on the device labeling exceeds its actual validated shelf life because the sterilization process used does not have adequate support for the extended shelf life. | Labeling False and Misleading | Customers were notified via letter on about 03/02/2018. Instructions included to review inventory for affected devices, arrange for their return to SPS Sterilization, and complete and return the Buisness Response Card. A follow up letter was sent to customer on about 03/07/2018 and included a description of the risk and hazard involved. Instructions remained the same. |  | 10 | 10 units total | US distribution to AZ, CA, FL, OH, PA, TX, VA, WV | Terminated | November 27, 2018 | 265 | 10
Z-3045-2018 | 80846 | No | HQP | Thermal cautery unit. | Ophthalmic | Ophthalmic | 510(k) | Beaver Accu-Temp High Temperature Cautery, 2 inch Shaft with Fine Tip, Model Number 8445000 |  |  | July 19, 2018 | 2018 | Beaver Visitec | A small number of devices in this lot may have a melted cap and a hole formed through the packaging. The risk is potentially the product will no longer be sterile for use and may cause harm to a patient. | Nonconforming Material/Component | Urgent Medical Device Recall notification letters dated 7/19/18 were distributed to customers. The letters instructed customers to perform the following actions:1. Identify the product.2. Immediately examine your inventory and quarantine product subject to recall. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.3. Consider all potential users of this product in your wholesale and/or user level supply chain. Your assistance is appreciated and necessary to prevent patient harm.4. Follow the attached Destruction Instructions to disable the wire tip to safely destroy this product to prevent activation or future use of this product.5. Complete and return the enclosed response form as soon as possible. Please provide objective evidence of the product destruction and we will credit your account.Customers were also contacted via phone beginning 8/1/18 |  | 800 | 800 units | Worldwide Distribution - US Nationwide in the states of: CA, DC, FL, IA, ID, IL, MA, MI, MN, NY, OK, PA, TN, TX, UT, and WA and the countries of Australia, France, Netherlands, Spain, and United Kingdom. | Open, Classified |  |  | 800
Z-3061-2018 | 80755 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | AffixusÂż Hip Fracture Nail Right 125Âż 11 mm x 380 mm, Item Number: 814311380 |  |  | July 17, 2018 | 2018 | Zimmer Biomet, Inc. | The set screw is not able to advance or reverse from the original position in the nail. | Under Investigation by firm | On 7/17/18, Urgent Medical Device Recall letters were sent to Risk Managers and Distributors. The letters instructed customers to do the following:Risk Manager Responsibilities:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com.Distributor Responsibilities:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. Request a Return Authorization Number through Fast SMS or by emailing rgarequest@zimmerbiomet.comc. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.d. Include a hardcopy of Attachment 1 in each carton of |  | 5947 | 5947 total | The products were distributed to the following US states: AL, AR, AZ, CA, CO, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WY, and Puerto Rico.The products were distributed to the following foreign countries: Chile, Ecuador, Japan, Netherlands. | Open, Classified |  |  | 5947
Z-3177-2018 | 80843 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Gentle Thread PLGA Full Thread Interference Screw, 10x30mm, Item Number 905629 |  |  | July 20, 2018 | 2018 | Zimmer Biomet, Inc. | The Gentle Thread PLGA Round Head, 7x20mm and the Gentle Thread PLGA Full Thread 10x30 have been commingled. | Process control | On 7/20/18, Urgent Medical Device Recall notification letters were sent to customers. The letters instructed customers to do the following:Risk Managers: 1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com.Surgeons:1. Review this notification for awareness of the contents.2. There are no specific patient monitoring instructions related to this recall that are recommended beyond your existing follow-up schedule.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be tra |  | 52 | 52 | The products were distributed to the following US states: IN, TN.The products were distributed to the following foreign countries: Australia, Chile, Netherlands. | Open, Classified |  |  | 52
Z-2998-2018 | 80807 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | iView DAB Detection Kit, Catalog Number 05266157001, model 760-091 |  |  | August 02, 2018 | 2018 | Ventana Medical Systems Inc | Leaking and sticking reagent dispensers for horseradish peroxidase (HRP) dispensers. A false negative IHC result may occur. | Under Investigation by firm | The firm disseminated their field safety notices by letter on 08/02/2018. The firm requested the destruction of the product. |  | 3068 | 3068 units | worldwide | Open, Classified |  |  | 53232
Z-2991-2018 | 80867 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | AIRO Mobile CT System |  |  | August 01, 2018 | 2018 | Mobius Imaging, LLC | The Tube Current Modulation feature (Modulated Scans) is not working in AIRO systems with software version 2.0.0.0, and operators would not be able to detect this fault until after a scan is completed. | Radiation Control for Health and Safety Act | Mobius Imaging, LLC sent an URGENT MEDICAL DEVICE RECALL letter to customers advising that a voluntary recall related to the AIRO Mobile CT System was initiated. The letter included the reason for the recall and actions to be taken by the customers. |  | 147 | 147 | US and foreign distribution. | Open, Classified |  |  | 147
Z-3175-2018 | 80979 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Leica M525 F20 Surgical Microscope System |  |  | September 04, 2018 | 2018 | Leica Microsystems, Inc. | The firm received complaints that the instrument tipped over when the swing arm was fully extended. | Device Design | The firm disseminated their Field Corrective Action Notices by mail beginning on 09/04/2018 to the US and Canadian consignees. The letter explained the problem and informed the consignee that an update package with enhanced instructions for use and additional marking will be provided within the next 3 months.Additional consignees found were notified by email on 10/01/2018. |  | 160 | 160 units | US and Canada | Open, Classified |  |  | 160
Z-3042-2018 | 80627 | No | HBT | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | AnastoClip GC Closure System, REF# 4007-08 a closure system designed to penetrate the outer layers of tissue during the forming of an anastomosis or closing of a vessel with possible intraluminal penetration. |  |  | June 19, 2018 | 2018 | LeMaitre Vascular, Inc. | The wrong instructions for use (IFU) were supplied with the AnastoClip GC closure system. The IFU provided lists some indications which are not cleared by the FDA for devices distributed in the United States: 1) closure of arteriotomies and veinotomies, 2) attachment of synthetic vascular prostheses, and 3) approximation of dural tissue. | Under Investigation by firm | On June 19, 2018, the firm sent Urgent Medical Device Recall letters to affected consignees. The letters informed customers of the product labeling issue. Customers were asked to complete the customer reply form and return it to the firm, and to forward the recall to any other facilities where product had been distributed. After the customer service department receives the completed form, they will contact the customer to arrange for a return of the recalled product at no charge to the customers. Customer service will also arrange for customers to receive replacement product.If you have any questions concerning this recall, please contact Ms. Rose Lerer, Senior Quality Engineer, 781-425-1671 or rlerer@lemaitre.com. |  |  |  | Distributed to accounts in FL, NV, OK, SC, and TX. | Open, Classified
Z-3043-2018 | 80697 | No | CGZ | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products Cl- Slides, Product Code 844 5207 |  |  | July 11, 2018 | 2018 | Ortho-Clinical Diagnostics | The slides do not meet the maximum allowable interference (MAI) claims for triglyceride interference when using serum samples. Triglyceride interference on the Cl- Slide was observed around a triglyceride concentration 600mg/dL, however the current triglyceride interference claim states interference should not occur below 800mg/dL of triglyceride. | Under Investigation by firm | On 7/11/18, Important Product Correction Notification letters were distributed to customers. The letters instructed customers to perform the following actions:VITROS Cl- Slides and VITROS CREA Slides Instructions for Use will be updated to contain the revised information described in this letter; we will notify you upon availability. In the interim, retain a copy of this customer letter for your records.Calibration Diskette and ADD Data Release Version (DRV) 5992 (and above) will support the revised measuring range for VITROS CREA Slides. We anticipate that DRV 5992 will be available beginning on July 28, 2018 by eConnectivity or will be sent via your normal Cal Disk or ADD mailing.Prior to the availability of the revised IFU, be aware that if a serum or plasma sample contains a Triglyceride concentration of 600 mg/dL or higher, the Chloride result may be positively biased by approximately 2.1 mmol/L. Follow your normal laboratory procedures as you would for other known assay interferences. Note: It is acceptable to continue using VITROS Cl- Slides.Upon availability, install the Calibration Diskette or ADD to implement the revised Measuring Range:For VITROS 4600, 5600 or 5,1 FS Systems: Install ADD DRV 5992 or above, Select All Assay Data optionFor VITROS 250/350 Systems: Install Calibration Diskette DRV 5992 or aboveRefer to the enclosure for additional instructions to update the Measuring Range.Update laboratory procedures and your Laboratory Information System (LIS), as necessary.Retain this letter until we send a notification to alert you when the revised IFUs are available.Complete the Confirmation of Receipt form and return by July 25, 2018.Post this notification by each system that processes VITROS Cl- or CREA Slides.Forward this notification if the product was distributed outside of your facility. |  | 128196 | 128196 | Worldwide Distribution - US Nationwide and in the countries of: Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Columbia, Denmark, France, Germany, India, Italy, Japan, Mexico, Norway, Poland, Portugal, Russia, Singapore, Spain, Sweden, The Netherlands, and United Kingdom. | Open, Classified |  |  | 548301
Z-3178-2018 | 80686 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Permobil powered wheelchairs, Model Numbers: (a) M300 PS, (b) C300 PS, (c) C350 PS, (d) C500 PS, (e) C400 PS |  |  | August 15, 2018 | 2018 | Permobil, Ab | This correction decision has been made because of a potential failure of one or both backrest bars providing back support. If one or both back support telescoping bars fail the user could fall back leading to a potential injury. | Device Design | The firm initiated the recall on 08/15/2018 by letter to the direct accounts (dealer/distributors). The letter instructed the consignee to contact the end users and perform the correction. |  | 3134 | 3134 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 3134
Z-3180-2018 | 80860 | No | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | VITEK(R) 2 AST-P655, For susceptibility testing of Staphylococcus spp. Enterococcus spp., ad S agalactiae against specified antimicrobials, REF 421913 |  |  | August 13, 2018 | 2018 | bioMerieux, Inc. | Routine internal review of Device History Records identified six VITEK(R) 2 AST card lots were manufactured with less drug (gentamicin) than required for the Gentamicin High-Level Resistance (ghlr01n) screen test due to an error when making 1500mL and 2000mL media well preparations (solutions). | Process control | The firm initiated the recall by email on 08/13/2018. The firm requested the consignee discard the impacted product and contact the local bioMerieux representative for credit. |  | 10919 | 10919 cartons | Worldwide Distribution US Nationwide and the countries of: Turkey, Austria, Switzerland, Czech Republic, Germany, Netherlands, | Open, Classified |  |  | 13457
Z-3181-2018 | 80894 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | IsoFlex LAL Support Surface, Model Number 2860 |  |  | August 13, 2018 | 2018 | Stryker Medical Division of Stryker Corporation | Certain 2860 IsoFlex LAL support surfaces were assembled with a non-conforming fire barrier. | Mixed-up of materials/components | Urgent Medical Device Correction notification letters dated 8/13/18 were distributed to customers. The letters included the following instructions:1. Locate the units listed on the attached business reply form and identify theaddress where they can be serviced.2. Return the enclosed business reply form to confirm receipt of this notification byfax 269 488 8691 or email productfieldaction@stryker.com.3. Upon receipt of the completed business reply form, a member of Strykers fieldservice personnel will contact you to arrange for the replacement of the firebarrier on your affected IsoFlex LAL units.4. If you have questions regarding the continued safe use of your products, pleasecontact Technical Support by calling +1 800 327 0770 and selecting opt 2.5. If you have loaned or sold any of the products listed in this letter, please forwarda copy of this notice to the new users and advise us of their new location in thespace provided on the business reply form.6. If you have disposed of any of these units and they are no longer in use, please,advise us of their obsolescence by providing us with their serial numbers in thespace provided on the business reply form. |  | 243 | 243 | US Nationwide in the states of: KY, NH, OH, PA, and TX. | Open, Classified |  |  | 243
Z-3182-2018 | 80762 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Makoplasty RIO Standard System; Ethernet to Fiber Optic Converter Catalog # 200933 |  |  | December 27, 2017 | 2017 | Mako Surgical Corporation | Communication-connection error. | Nonconforming Material/Component | A customer Notification letters and Product Correction Acknowledgement Forms will be sent to affected customers via UPS Day Air return receipt. Customers where informed that the affected products will be replaced. |  | 736 | 736 (US) and 174 (OUS) | Worldwide Distribution - US Nationwide and the countries of Australia, New Zealand, Taiwan, China, Hong Kong, Greece, United Kingdom, England, Netherlands, Luxembourg, Germany, Ireland, Switzerland, Italy, India, Japan, South Korea, Vietnam, South Africa, and Turkey, | Open, Classified |  |  | 736
Z-3183-2018 | 80832 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | PR1000 Pivot Rehab Bed AC powered adjustable bed. |  |  | July 25, 2018 | 2018 | Med-Mizer, Inc. | Under certain conditions, when pivoting bed to the extreme left there are situations where the wire/cables may become pinched. A failure would result in the resident requiring to be transferred to a different bed until repairs can be made | Device Design | Customers were contacted via telephone on 7/25/18 to notify them of the issue and to make them aware of the correction. Customers were then mailed a notification letter along with necessary parts. |  | 230 | 230 units | The products were distributed to the following US states: AR, AZ, FL, GA, IA, ID, IL, IN, KY, ME, MI, MN, MO, NC, ND, NE, NM, NY, OH, OK, PA, SC, SD, TN, TX, VA, WA, and WI.The products were distributed to the following foreign countries: Canada, China, and Denmark. | Open, Classified |  |  | 230
Z-3184-2018 | 80712 | No | CGZ | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products Cl- Slides, Product Code 684 4471 |  |  | July 11, 2018 | 2018 | Ortho-Clinical Diagnostics | The slides failed to meet current claims for Limit of Blank (LoB), Limit of Detection (LoD) and Limit of Quantification (LoQ) for urine samples. The measuring range (15-300 mmol/L) for urine samples when tested with the Cl- slide is not affected. The current Cl- LoQ is 5 mmol/L and the LoQ determined by this investigation is 15 mmol/L. | Under Investigation by firm | On 7/11/18, Important Product Correction Notification letters were distributed to customers. The letters instructed customers to perform the following actions:VITROS Cl- Slides and VITROS CREA Slides Instructions for Use will be updated to contain the revised information described in this letter; we will notify you upon availability. In the interim, retain a copy of this customer letter for your records.Calibration Diskette and ADD Data Release Version (DRV) 5992 (and above) will support the revised measuring range for VITROS CREA Slides. We anticipate that DRV 5992 will be available beginning on July 28, 2018 by eConnectivity or will be sent via your normal Cal Disk or ADD mailing.Prior to the availability of the revised IFU, be aware that if a serum or plasma sample contains a Triglyceride concentration of 600 mg/dL or higher, the Chloride result may be positively biased by approximately 2.1 mmol/L. Follow your normal laboratory procedures as you would for other known assay interferences. Note: It is acceptable to continue using VITROS Cl- Slides.Upon availability, install the Calibration Diskette or ADD to implement the revised Measuring Range:For VITROS 4600, 5600 or 5,1 FS Systems: Install ADD DRV 5992 or above, Select All Assay Data optionFor VITROS 250/350 Systems: Install Calibration Diskette DRV 5992 or aboveRefer to the enclosure for additional instructions to update the Measuring Range.Update laboratory procedures and your Laboratory Information System (LIS), as necessary.Retain this letter until we send a notification to alert you when the revised IFUs are available.Complete the Confirmation of Receipt form and return by July 25, 2018.Post this notification by each system that processes VITROS Cl- or CREA Slides.Forward this notification if the product was distributed outside of your facility. |  | 168919 | 168919 | Worldwide Distribution - US Nationwide- The products were distributed to the following foreign countries: Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Columbia, Denmark, France, Germany, India, Italy, Japan, Mexico, Norway, Poland, Portugal, Russia, Singapore, Spain, Sweden, The Netherlands, and United Kingdom. | Open, Classified |  |  | 168919
Z-3185-2018 | 80724 | No | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PATHFAST Myoglobin II reagent cartridge |  |  | April 30, 2018 | 2018 | Polymedco, Inc.,/Polymedco Cancer Diagnostic Products, Llc | A pinhole caused by the laser printer in the aluminum sheet over the ALP reagent wells caused leakage through the pinhole. If the cartridge with leakage is used, higher measurement value as low as +12% and no higher than +20% bias may be obtained. | Process control | The firm, Polymedco, initiated a recall by telephone and followed with email on 4/30/18 to inform users to inspect the reagent cartridges for evidence of leakage from the ALP well and to discard cartridges that had evidence of leakage. Polymedco will replace any affected reagent cartridges. If you have any questions, contact Polymedco Technical Support at the1.800.431.2123 Ext 285. |  | 272 | 272 KITS in total | US Distribution to states of: AR, CA, FL, ID, LA, MN, NY and TX. | Open, Classified |  |  | 272
Z-2971-2018 | 80628 | No | HBL | Neurosurgical head holder (skull clamp). | Neurology | Neurology | 510(k) | DORO LUCENT Locking Transitional Member, REF 1101.031 Component of headrest system used in open and percutaneous craniotomies and spinal surgeries |  |  | July 05, 2018 | 2018 | Pro-Med Instruments Gmbh | The combination of two different design revisions of DORO LUCENT Base Unit, Locking Transitional Member, and Transitional Member within one set up might lead to a malfunction and failure of the headrest system. | Component change control | The firm initiated the recall by email on 07/05/2018.This stage is only relevant for hospitals / users who are in possession with concerned products of two different revisions (mix-up):a)Discontinue the use of the combination and quarantine the concerned products b)Contact pro med instruments for further instructions relating to temporary replacement possibilities |  | 33 | 33 units | Worldwide distribution to US states of CA, TX, TN, MS, GA, WV, and OH, Qatar, India, UK, Singapore, Australia, Germany, South Africa, Italy, and Japan | Open, Classified |  |  | 95
Z-3188-2018 | 80938 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | sterEOS workstation |  |  | August 01, 2018 | 2018 | EOS Imaging | 3D projections may be incorrect whenthe pair of images used is a secondary capture generated from a series containing more than two images, where one of the images has been manually flipped. Projection errors may lead to errors in calculation of 3D clinical parameters whenadjusting the position, size and shape of the 3D objects used. Incorrect clinical parameters can be a contributing factor in incorrect clinical decisions, or sub-optimal surgical interventions. | Software design | On August 1, 2018, a Field Safety Notice was issued to all impacted customers. Measures to be taken by users: EOS system and sterEOS review workstation users are requested not to apply manual symmetry to an image before saving a 2D secondary capture if more than two images are displayed in the EOS viewer mode or in 2D sterEOS.Users of the 3D sterEOS review workstation are requested to check that 3D projections associated with pelvic parameters match when a 3D image is created from a pair of images belonging to a series of more than two images.INSTRUCTIONS FOR END USERS:1) Please read the field safety notice.2) Follow the instructions provided in this field safety notice.3) Please transfer this notice to whomever needs to be informed within your organization or in any organization to which the sterEOS review workstation may have been transferred.INSTRUCTIONS FOR DISTRIBUTORS:1) Please read the field safety notice.2) Please transfer all the documentation relating to the field safety notice to your customers in the field and keep proof of delivery of this documentation. Please regularly submit your customers' response rate to EOS imaging by email (regulatory@eos-imaging.com). EOS imaging needs this proof to monitor the evolution of corrective measures. |  | 17 | 17 units | US nationwide | Open, Classified |  |  | 17
Z-3191-2018 | 80981 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | ddR Formula B X-ray System, Model ddR Formula B |  |  | November 18, 2016 | 2016 | Swissray | Potential for bucky (the part that holds the grid and is moveable to position the patient) moving unexpectedly. | Device Design | The firm, Swissray Medical, issued official communication to customers as a Action Notification Report/work Instructions on November 18 2016. The Work Instructions state the affected devices, the problem symptoms, causes of the problem, remedy, any tools required, and the step-by-step corrective actions. The manufacturer, Swissray Medical AG, distributed the Work Instructions for the FCO and the Action Notification Report Form to all distributors and service entities who would be performing corrective actions in the field. The Swissray Field Engineer then presented the Report to the customer (end user facility) and communicated with the customer in-person while performing the corrective action on-site. When corrective actions were completed, both the customer and service personnel then sign the Report to confirm that the customer was informed about the corrective action and that corrective action was completed. The Report Form states the title and unique designation of the corrective action, whether it was mandatory by the manufacturer or was an optional product improvement, and the unique document ID number of the Work Instructions for the FCO. The customers were instructed to Complete the FCO Action Notification Report and send to:Swissray Medical AGexpert CenterTurbistrasse 25-27CH-6280 HochdorfSwitzerlandOr scan and email to service@swissray.com with FCO in the subject line.If you have any questions, contact QA/RA Coordinator at 908-372-6408 or email: Deniz.Kortan@swissray.com. |  | 27 | 27 devices in US | Worldwide Distribution: US (Nationwide) and to countries of: Austria, Dubai, China. Germany, Libia, Malaysia, Russia, and South Africa. | Open, Classified |  |  | 27
Z-3190-2018 | 80905 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Pheno Interventional Fluoroscopic X-Ray System, Model Number 10849000 |  |  | July 29, 2018 | 2018 | Siemens Medical Solutions USA, Inc | Temperature of the detector may fall out of its designed working range (cool down) due to incorrectly configured temperature control unit. This unintentional system behavior may occur while the unit is powered down for long periods (e.g. at night). | Device Design | Urgent Medical Device Correction notification letters dated 7/27/18 were distributed to customers. The letters informed the customers of the following planned actions:Our service organization will contact you shortly to arrange a date to perform this corrective action with Update Instruction AX001/18/S. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436.Following the corrective action, the cause is eliminated and a recurrence of the fault is prevented. Additionally, Siemens is developing a solution to restore the system back to the original settings. This solution will be distributed to all affected customers upon completion. |  | 429 | 429 total | US nationwide distribution. | Open, Classified |  |  | 429
Z-3192-2018 | 80761 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Xper Flex Cardio Physiomonitoring System Model # 453564669081 Software Version 1.6.0.0388 |  |  | August 03, 2018 | 2018 | Invivo Corporation | Released with an incorrect, older firmware version installed. May result in intermittent loss of ECG monitoring, inaccurate ECG amplitude display and or auxiliary output, and inaccurate heart rate (HR) displayed due to QRS detection fault. | Software Manufacturing/Software Deployment | Philips is asking customers to follow the Action to be Taken by Customer/User section of the August 3, 2018, Field Safety Notice: The Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Philips will replace all affected devices. Users should discontinue use of affected devices immediately. A Philips representative will contact you regarding your affected device. |  | 38 | 38 | Distribution in the Netherlands | Open, Classified |  |  | 38
Z-3195-2018 | 80875 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008T HEMODIALYSIS SYSTEM W/BIBAG, Part Number 190766, Applicable only with 2008T BlueStar Premium Upgrade Kit, Part Number 191138 |  |  | July 13, 2018 | 2018 | Fresenius Medical Care Renal Therapies Group, LLC | In the hemodialysis machine during the cleaning/disinfection program, when there is a power failure or interruption, the dialysis program button should be greyed out when the machine is powered up and a mandatory rinse should be required. However, if a power failure or interruption occurs during the cleaning/disinfection program with the auto-start feature enabled, the machine enters the dialysis program even though the dialysis program button is disabled and the mandatory rinse has not been performed. As a result, disinfecting chemical may not be completely removed from the machine before starting a treatment and disinfectant could exist in the dialysate delivered to the patient. | Under Investigation by firm | On 07/13/2018 phone calls were immediately conducted to the affected clinics. Contact was made to the Technical Service Directors explaining issue and to disable the Auto Start Function on their BlueStar SW v2.71 installed 2008T Hemodialysis Machines. |  | 40 | 40 | US Nationwide in the states of : CA, CT, MI, NC, and NY. | Open, Classified |  |  | 120
Z-3198-2018 | 80407 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | URO-3 Triple Stain (CD44 + p53) with CK20, Catalog #: PM370TSAA, Package Type: Bag, which contains components: a) CD44 + p53, Catalog #: PM373DSAA, Package Type: Bottle, and b) CK20, Catalog #: PM062AA, Package Type: Bottle |  |  | March 15, 2018 | 2018 | Biocare Medical, LLC | Immunohistochemistry device kits found to display incorrect staining pattern due to mislabeling of two components. | Other | On 03/15/2018, an initial email notification was sent to customers who had received the affected lot of product. Customers were asked to return unused product. On 03/29/2018, the firm mailed and emailed Urgent: Medical Device Recall Notification letter to customers. Customers were instructed to discontinue affected product use and distribution, complete and return the Return Response Form. Telephone number 800-799-9499 was provided. |  | 11 | 11 | Worldwide Distribution - Unites States Nationwide in the states of NY AND FLand the countries of: Morocco and Italy | Open, Classified |  |  | 11
Z-3196-2018 | 80632 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Portex Pulsator Arterial Blood Sampling Kit with Dry Lithium Heparin for Gases and Electrolytes, REF 4080P-2, Sterile, RX. The firm name on the label is Smiths Medical ASD, Inc., Keene, NH. |  |  | July 13, 2018 | 2018 | Smiths Medical ASD Inc. | The product was mislabeled with the incorrect label. | Labeling Change Control | Recall notifications dated 7/13/2018 were issued via FedEx overnight mail on 7/13/2018 explaining the reason for recall, risk to health, and instructing the customer to check their inventory for the affected lot numbers. |  | 29600 | 29,600 kits | US Nationwide Distribution in the states of CT, MA, MI, NC, NJ, NY, RI, TX, UT, and WA. | Terminated | October 19, 2018 | 98 | 29600
Z-3199-2018 | 80839 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | 4.5 Fr x 40 cm single-lumen PICC Kit, Product Code CDC-44041-VPS2, 5.5 Fr double-lumen PICC Kit, Product Code CDC-44052-VPS2 |  |  | July 17, 2018 | 2018 | Arrow International Inc | The products may contain the incorrect catheter. Kits that should contain a 4.5 Fr single-lumen catheter may contain a 5.5 Fr double-lumen catheter. Kits that should contain a 5.5 Fr double-lumen catheter may contain a 4.5 Fr single-lumen catheter. This issue may cause a delay in treatment while a replacement catheter is obtained. | Under Investigation by firm | On 7/17/18, Urgent Medical Device Recall notification letters were sent to customers. Customers were instructed to do the following:1. If you have affected stock, immediately discontinue use and quarantine any products with the product codes and lot numbers listed above, so that the affected products can be returned to Arrow.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to Arrow International.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document your receipt of this letter.Distributors were instructed to do the following:1. Immediately discontinue distribution and quarantine any products with the product codes and lot numbers listed above.2. Using the provided customer letter and Recall Acknowledgement Form templates, communicate this recall to any of your customers who have received product included within the scope of the recall.3. Have the customers return any affected product to you, together with a completed Recall Acknowledgement Form, for consolidation and return to Arrow International. In the event that an alternative approach is needed, contact Arrow International Customer Service for more information at 1-866-396-2111.4. To return affected products from your inventory, complete the enclosed RecallAcknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with |  | 1257 | 1257 total | The products were distributed to the following US states: AL, AZ, CA, GA, MA, MD, ME, MI, NY, PA, RI, SC, VA, and WA. | Open, Classified |  |  | 1257
Z-3193-2018 | 80715 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products CREA Slides, Product Code 680 2584 |  |  | July 11, 2018 | 2018 | Ortho-Clinical Diagnostics | The slides did not meet the current claims for Limit of Blank (LoB), Limit of Detection (LoD) and Limit of Quantification (LoQ) for serum and urine samples and therefore affecting the low end of the measuring range for this assay. The current low end measuring range is 0.05mg/dL for serum and 1.2mg/dL for urine. Based on the CREA testing results for the LoQ, the revised low end CREA measuring range is 0.15mg/dL for serum and 3.2mg/dL for urine. | Process control | On 7/11/18, Important Product Correction Notification letters were distributed to customers. The letters instructed customers to perform the following actions:VITROS Cl- Slides and VITROS CREA Slides Instructions for Use will be updated to contain the revised information described in this letter; we will notify you upon availability. In the interim, retain a copy of this customer letter for your records.Calibration Diskette and ADD Data Release Version (DRV) 5992 (and above) will support the revised measuring range for VITROS CREA Slides. We anticipate that DRV 5992 will be available beginning on July 28, 2018 by eConnectivity or will be sent via your normal Cal Disk or ADD mailing.Prior to the availability of the revised IFU, be aware that if a serum or plasma sample contains a Triglyceride concentration of 600 mg/dL or higher, the Chloride result may be positively biased by approximately 2.1 mmol/L. Follow your normal laboratory procedures as you would for other known assay interferences. Note: It is acceptable to continue using VITROS Cl- Slides.Upon availability, install the Calibration Diskette or ADD to implement the revised Measuring Range:For VITROS 4600, 5600 or 5,1 FS Systems: Install ADD DRV 5992 or above, Select All Assay Data optionFor VITROS 250/350 Systems: Install Calibration Diskette DRV 5992 or aboveRefer to the enclosure for additional instructions to update the Measuring Range.Update laboratory procedures and your Laboratory Information System (LIS), as necessary.Retain this letter until we send a notification to alert you when the revised IFUs are available.Complete the Confirmation of Receipt form and return by July 25, 2018.Post this notification by each system that processes VITROS Cl- or CREA Slides.Forward this notification if the product was distributed outside of your facility. |  | 627543 | 627543 | Worldwide distribution. US nationwide, Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Colombia, Denmark, France, Germany, India, Italy, Japan, Mexico, Norway, Poland, Portugal, Russia, Singapore, Spain, Sweden, The Netherlands, and United Kingdom. | Open, Classified |  |  | 627543
Z-3201-2018 | 80685 | No | QBJ |  | Clinical Chemistry | Clinical Chemistry | 510(k) | Contour Plus Link 2.4, Product Catalog Number: MMT-1151SK. wireless blood glucose monitoring system |  |  | May 16, 2018 | 2018 | Medtronic Inc. | CONTOUR PLUS LINK 2.4 meter with an incorrect unit of measure was included into meter kits and distributed. This meter contains the incorrect unit of measure mg/dL instead of mmol/L. | Nonconforming Material/Component | The firm, Medtronic, sent an "URGENT MEDICAL DEVICE RECALL- CONTOUR"PLUS LINK 2.4 Meter (MMT-1151SK - Batch# MD04M051P) Incorrect Unit of Measure" letter to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:A. Confirm if the Model (REF) of your meter is MMT-1151SK and the batch (LOT) number is MD04M051P.B. If the REF and LOT match this notice, stop using the CONTOUR"PLUS LINK 2.4 meter immediately.C. Call your local Helpline and let the customer service team know you received a notice stating you may have a CONTOUR"PLUS LINK 2.4 meter with the incorrect unit of measure.D. Return your recalled CONTOUR"PLUS LINK 2.4 meter to Medtronic by following the instructions provided by the Helpline representative. Medtronic will provide a free replacement.If you have any questions, contact the Program Manager, Field Corrective Action at 818-576-4700. |  | 2 | 2 | International Distribution to Slovakia only. | Open, Classified |  |  | 2
Z-3197-2018 | 80872 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | ARROWÂż Bipolar Electrode Catheter with shrouded pins , 6 Fr, 110 cm |  |  | July 30, 2018 | 2018 | Arrow International Inc | Products are labeled with an incorrect expiration date | Error in labeling | Arrow International notified US consignee notification by letter on 30 JUL 2018 via FedEx 2-day mail. Account action: immediately discontinue use and quarantine any products with the product codes and lot numbers listed above, so that the affected products can be returned to Arrow. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. |  | 7 | 7 eaches | US in the state of PA | Open, Classified |  |  | 7
Z-3209-2018 | 80763 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Optetrak Logic CC Stem Extension, 22mm x 160mm |  |  | May 08, 2018 | 2018 | Exactech, Inc. | Potentially have non-conforming internal threads. | Under Investigation by firm | The firm, Exactech, contacted the consignee on May 3, 2018 by telephone to notify the consignee to quarantine the device pending recall notice distribution. Exactech distributed an "URGENT MEDICAL DEVICE RECALL" notice dated May 7, 2018 to the consignee in possession of affected inventory by electronic mail on May 7, 2018.In order to comply with the notice, consignees have been instructed to:" Immediately cease distribution or use of this product." Extend the information to their accounts that may have this product in their possession." Identify and quarantine any of the subject devices in their inventory" Complete and return the attached Recall Acknowledgment Form and Recall Inventory Response Form to Exactech at recalls@exac.com.On May 15, 2018, the recalled device was returned to Exactech and quarantined.If you have any questions, contact Sr. Manager, Regulator Affairs at 352-327-4613. |  | 1 | 1 | US Distribution to Ohio only. | Completed |  |  | 1
Z-3202-2018 | 80944 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Disposable Insulin Syringe 0.3mL in 3 different sizes 29Gx1/2, 30Gx5/16, and 31G x 5/16. |  |  | August 24, 2018 | 2018 | ShinChang Medical Co., Ltd. | Defect consists of a small crack in the top end of the barrel near the needle. | Under Investigation by firm | The company has been requesting that the importers immediately cease all distribution of affected Lot Numbers and notify its customers to do the same. The customers shall include wholesale distributors, retail pharmacies and mail order pharmacies or patients. The customer who received the affected products shall quarantine. The customers may return the affectedproducts and the company will replace. Any and all questions pertaining to the recall and disposition of recalled product can be addressed to Yushin Kim, CEO ShinChang Medical Co.,Ltd. By email at yushin0229@scmedical.co.kr or by phone at +82-54-463-2400 |  | 6570800 | 6,570,800 | FL, TN, and VA | Open, Classified |  |  | 6570800
Z-3211-2018 | 80933 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Ashland(TM) Box Film Ball Cube II EBT3, P/N: 027863-002 rev E |  |  | August 31, 2017 | 2017 | Accuray Incorporated | the Ball Cube II phantom calibration film was not manufactured within necessary dimensions and may introduce up to 0.5 mm of positional inaccuracy to the CyberKnife System | Nonconforming Material/Component | The firm, Accuray, disseminated their "Urgent Field Safety Notices" dated 08/31/2018 by email and letter to its customers. The letter described the product, problem and actions to be taken. The letter requested the customers cease use and quarantine the product for recovery and replacement by an Accuray Representative and complete and return the Acknowledgement Form via email at: astsauver@accuray.com or mail to: Accuray Incorporated, one Erdman place, Madison, WI 53719.If you have any questions, contact Vice President, Regulatory Affairs at 608-824-2800 or email: dbiank@accuray.com or the Regulatory Affairs Specialist at 408-716-4660 or jysingh@accuray.com. |  | 259 | 259 units | Worldwide Distribution: United States (nationwide), and countries of: United Kingdom, Germany, France, Finland, Netherlands, Portugal, Hong Kong, Poland, Australia, United Arab Emirates, Thailand, and Japan. | Open, Classified |  |  | 259
Z-2754-2018 | 80641 | No | EZL | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Temperature sensor Foley catheter 18FR, Catalogue Number 102205101880AB |  |  | June 21, 2018 | 2018 | Degania Silicone, Ltd. | The catheters are suspected to contain defective sensors; if used, defective catheter will show lower body temperature than the actual body temperature of a patient. | Employee error | Urgent Field Safety Notice letters were distributed to customers on 6/21/18. The letters instructed customers to perform the following:1. Please forward this information to the immediate attention of the persons responsible for usage of the affected parts. Sub-recall must be initiated by your company to collect any distributed product from the suspected list.2. Please fill in the attached Recall Response form and indicate the quantity of the affected product in your possession.3. Promptly return the completed Recall Response form to Degania Silicone by e-mail: Katia@ds-il.comAny affected product should be returned to Degania Facility at Rhode Island:Degania Silicone, Inc.14 Thurber Blvd, Suite ASmithfield, RI 02917, USContact name: Jackie Dubone-mail: jackie@deganiasilicone.com |  | 153 | 153 | Worldwide Distribution - US Nationwide in the states: IL, NC, NY and the countries of : Belgium, Israel, Japan, and Mexico. | Open, Classified |  |  | 367531
Z-3205-2018 | 80747 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 400mL, Premium, Urine Meter, 14 FR, Silicone Temperature Sensing Foley Traywith Cath-Secure Plus, Sterile, Item Code P4P14TSD |  |  | July 11, 2018 | 2018 | COVIDIEN LLC | The temperature sensor catheters may be defective in that they will show a lower body temperature measurement than the actual temperature of the patient. | Nonconforming Material/Component | On 7/11/18, an Urgent Medical Device Recall letter was distributed to customers. The letter instructed customers to do the following: 1. Please quarantine any unused product of the Covidien item codes and lot numbers listed above (see Attachment A to locate).2. If you have distributed any of the affected Covidien catheters or kits, then please promptly forward a copy of this letter to those recipients.3. All unused product of affected item codes and lot numbers must be returned.4. Please return unused affected product as explained below. You must complete the Recalled Product Return Form (Attachment B) regardless of whether or not you have inventory of affected product. |  | 27515 | 27515 total | The products were distributed to the following US states: CA, CO, FL, GA, IA, IL, IN, KY, LA, MA, MD, MO, NC, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA, and WA. | Open, Classified |  |  | 27515
Z-3212-2018 | 80657 | No | HBO | Scalp clip. | Neurology | Neurology | 510(k) | Medtronic Clip Gun Magazines are packaged in individually sealed pouches and distributed in boxes containing 5 individually pouched Clip Gun Magazine. |  |  | May 03, 2018 | 2018 | Medtronic Neurosurgery | The packaging sterile barrier does not comply with the applicable packaging standard for terminally sterilized medical devices. While the Instructions for Use (IFU) state Do not use if package has been previously opened or damaged , it is possible that the seal breach may go undetected by the user's naked eye. | Under Investigation by firm | Consignee Notification URGENT: MEDICAL DEVICE RECALL has been sent to US customers on 03-May-2018. For foreign customers, the Consignee Notification is in process. The Consignee letter informs customers of the situation and the required course of action which includes:1. Cease use of the listed product2. Account for number of product units used or discarded and return any units by completing the attached individualized Customer Confirmation form provided. Please follow the instructions described in the Customer Confirmation form.3. Clinicians and their Patients who may have received the listed devices should monitor the patients condition to determine if there are any signs of infection. If this is the case, the patient should seek medical attention immediately to treat the infection.In addition, the Notification Letter should be provided to all those who need to be aware of this matter within your organization or to any organization where the potentially affected products may have been transferred. Fax the Customer Acknowledgement Response Form to Medtronic Neurosurgery at (805) 571-8787 or take a clear picture or scan of signed form and email it to rs.mnsfca@medtronic.com with "Magazine Recall" in the subject line. |  | 2616 | 2616 | Domestic:AL, AZ, AR, CA, CO, DC, FL, GA, ID, IL, IN, IA, KS, LA, MD, MI, MN, MS, MO, NB, NV, NJ, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WYOUS:Australia, Germany, Kenya, Russian Federation, South Africa | Open, Classified |  |  | 2616
Z-3213-2018 | 80922 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | INTEGRIS Allura 9, System Code 722018 |  |  | July 16, 2018 | 2018 | Philips Electronics North America Corporation | There is a potential for the Monitor Ceiling Suspension, which holds a large screen monitor, may detach from the actuator rotor shaft causing the monitor to fall from the ground. When a Monitor Ceiling Suspension detaches from the actuator rotor shaft and the monitor falls, there is a risk ofinjury for the patient, user and bystander. | Device Design | The firm, BU IGT System, sent an "Urgent -Medical Device Correction" notification letter dated 07/16/2018 to customers. The letter describes the product, problem and actions to be taken. The letter instructs customers to do the following: In order to reduce the risk for patients, users and bystanders if this problem would reoccur, we recommend the following actions until the correction has been implemented.-Avoid unnecessary movements of the Monitor Ceiling Suspension.-For those movements that are necessary, avoid that the user, patient or bystander are in close proximity to the monitor.-When moving the Monitor Ceiling Suspension, ensure that no body partsof the staff or patient are underneath the monitor.-Do not move the monitor above the patient.Please ensure that all staff with access to the affected systems are informed of the content of this Medical Device Correction.All affected products will be corrected by means of a Field Change Order (FCO) free of charge. This FCO will be available mid-August.You will be contacted by our local Philips representative to schedule this corrective action.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722- 9377, 8am to 8pm (EST) Monday thru Friday. |  | 5115 | 5115 total | Worldwide Distribution: US (nationwide) to states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries of: Algeria, Argentina, Australia, Austria, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, British Virgin Islands, Brazil, Brunei Darussalam, Blugaria, Cambodia, Canada, Chile, China, Columbia, Croatia, Cyprus, Czech Republic, Denmark, Dominic Republic, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea South, Republic of Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mauritius, Mexico, Moldova, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestine, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Reunion, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Tunisia, Turkey, Ukraine, UAE, UK, Uzbekistan, Venezuela, Vietnam, and Yemen. | Open, Classified |  |  | 5115
Z-3240-2018 | 80863 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Hoffmann LRF Hexapod Software |  |  | July 26, 2018 | 2018 | Stryker GmbH | When creating a correction plan with the Correct Axial First button selected and the No. of corrections per day set to more than 1 (e.g., 2, 3, or 4), the Hexapod Software computes an erroneously accelerated correction plan for the axial portion of the plan. | Under Investigation by firm | Stryker contracted Third Party: Stericycle Contact Center, 6026 Lakeside Blvd., Indianapolis, IN 46278 notified affected hospitals via 2-day UPS (with return receipt) on July 26, 2018 All affected surgeon users notified via email with Request read Receipt and Request delivered Receipt options selected. Provided the following: if you have selected the Correct Axial First button, with multiple corrections per day, redo the correction plan by either of the following methods:a.Calculate the correction plan with one correction per day selected on the Correction Plan page orb.If more than one correction per day is desired, the Correction time may be overridden by entering the desired number of days in the Override correction time field on the Limiting Anatomical Structure (LAS) page. If the correction time is overridden, the correction plan will be accurate regardless of the number of corrections. Continue to evaluate the patients through routine follow-up radiographs for alignment and distraction. |  | 233 | 233 | Nationwide | Open, Classified |  |  | 233
Z-3241-2018 | 80841 | No | MOL | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Durex Real Feel Polyisoprene Condoms, 10 count and 24 count, SKU Numbers 02340-89456-14 and 02340-89457-00 |  |  | July 23, 2018 | 2018 | Reckitt Benckiser LLC | The products failed the filed specification of 1.0kPa for Burst Pressure. | Under Investigation by firm | On 7/23/18, an Urgent Medical Device Recall notification letter was distributed to customers. The letter instructed customers to do the following:1. Examine your inventory in your distribution center immediately to determine if you have the specified batch in your warehouse inventory.2. Segregate and quarantine the specified batch.3. If you have product of the specified batch, contact our Stericycle representative immediately for a RETURN AUTHORIZATION number.a. Contact Information:Phone: 877-448-5308Fax: 877-884-94104. Contact Stericycle Expert Solutions at 877-448-5308 for assistance with coordination of pickup and removal of product from your facility. If arranging pick up of product, freight will be prepaid by Stericycle. A pre-paid return label has been included for your convenience. Alternatively, you may arrange shipment of product directly to Stericycle at the following address:Attn: Event 42412670 Executive Drive, Suite AIndianapolis, IN 46241Dock Hours are SAM -4:30PM Monday- FridayPlease call to schedule a dock appointment5. Finally, notify Stericycle representative via email RB4241@stericycle.com that you have:a. Checked your inventory in your distribution centers and confirm you have NOproduct of the specified batchb. You have collected all units in your inventory and product has been shipped back in accordance with the instructions from your Stericycle representative.On 7/26/18, a second Urgent Medical Device Recall notification letter was sent to customers which included two additional batches included in the recall. This letter included the same customer instructions as the 7/23/18 letter. |  | 487980 | 487980 condoms | US Nationwide. | Open, Classified |  |  | 487980
Z-3247-2018 | 80865 | No | PPN | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | The Shockwave S4 IVL Catheters are intended for lithotripsy-enhanced balloon dilatation of lesions, including calcified lesions, in the peripheral vasculature, including the iliac, femoral, ilio-femoral, popliteal, infra-popliteal, and renal arteries. Shockwave S4, Peripheral Intravascular Lithotripsy Catheter, CE 0086, STERILE R, |  |  | July 18, 2018 | 2018 | Shockwave Medical, Inc. | Catheter unable to maintain inflation pressure during treatment (4ATM) and/or during IVL post dilatation (at 6ATM). | Process control | On 7/19/2018, the firm sent an Urgent Medical Device Recall Notice to distribution chain/consignees via e-mail informing them that, in some cases, the S4 Peripheral IVL Catheter is not performing as intended. We have seen instances of leaks in the balloon which prevent the balloon from staying inflated at 4 atm for the full course of lithotripsy application. There have been no reports of adverse clinical events related to this issue. This issue does not impact patients who were previously treated with the S4 device.Customers were instructed to discontinue use of all S4 IVL products listed in the recall and place it into quarantine until a Shockwave representative was able to arrange for the S4 IVL product to be returned to Shockwave Medical.The Shockwave representative retrieving the product from the customer is to complete the Acknowledgement and Receipt Form.Recall ContactNora HaddingVice President, Quality and Regulatory AffairsEmail: nhadding@shockwavemedical.comPhone: 510-279-4262Monday through Friday, 8:00 AM to 4:30 PM Pacific time. |  | 612 | 612 | Domestic: CO, TN and TXForeign: Austria, Germany and New Zealand | Open, Classified |  |  | 612
Z-3243-2018 | 80957 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo 220/240V AC, Advanced Perfusion System Platform (APS) for use in cardiopulmonary bypass procedures. |  |  | August 28, 2018 | 2018 | Terumo Cardiovascular Systems Corporation | Screws on a TerumoÂż Advanced PerfusionSystem 1 Power Manager Board (the base unit circuit board) were not torqued to specification. It is possible that a screw which has not been properly torqued can go undetected during manufacturing testing and may loosen over time, resulting in the potential for an electrical short in the system | Process control | Terumo CVS issued letter on 8/2/818 via Fed'X 2 day air to inform users of the problem for the Power Manager boards with unverifiable or out of specification screws, potential hazard and field correction. A Terumo Field Service Representative will contact users to schedule the field correction activities. Terumo CVS Customer Service: 1.800.521.2818, Monday  Friday, 8 a.m.  6 p.m. ET |  | 10 | 10 units | NationwideForeign: Mexico, BELGIUM, CANADA, India, Indonesia, Singapore, South Korea , Vietnam | Open, Classified |  |  | 57
Z-3248-2018 | 81045 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | ddR Formula B X-ray System, ddR Formula B |  |  | August 09, 2016 | 2016 | Swissray | Possible injury due to movement of the arm, calibration loss and communication loss. | Device Design | November 2016, Swissray Medical issued official communication to customers as a Action Notification Report. The manufacturer, Swissray Medical AG, distributed the Work Instructions for the FCO and the ActionNotification Report Form to all distributors and service entities who would be performing corrective actions in the field. The Swissray Field Engineer then presented the Report to the customer (end user facility) and communicated with the customer in-person while performing the corrective action on-site. When corrective actions were completed, both the customer and service personnel then sign the Report to confirm that the customer was informed about the corrective action and that corrective action was completed. |  | 26 | 26 units | US Nationwide | Open, Classified |  |  | 26
Z-3245-2018 | 81046 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | DX-D 600 |  |  | August 15, 2018 | 2018 | Agfa Healthcare NV | A supplier for the Agfa DX-D 600 system reported to Agfa a safety issue with the cables of a similar device. A non-Agfa customer had reported an incident caused by the breakage of the two steel cables that support the equipment to its roof anchor. The supplier made a decision to communicate to all customers a mandatory review of the entire install base, changing the pulley and steel cables in case of detecting any deterioration in the cables or irregular or excessive wear of the pulley. This issue is also applicable to the DX-D 600 system due to having the same mechanical design for the involved parts. If the two steel cables that support the equipment to its roof anchor break, the equipment might fall and cause harm to the patient, user, or third parties. | Component design/selection | On August 15, 2018, an "URGENT FIELD SAFETY NOTICE" letter was emailed to the US and Canadian consignees. The letter described the product issue: Breakage of the steel cables that support the equipment to its roof anchor may cause the equipment to fall down and potentially harm the patient, user, or third parties.Actions undertaken by Agfa NV:Soon Agfa NV or its representative will visit your site to do a verification/inspection ofyour unit related to this topic. Despite the fact that this issue potentially affects oldersystems with high workload, Agfa NV or its representative will verify/inspect all DX-D600 systems.Recommended actions to be taken by the customer:If you notice an abnormal displacement in the vertical axis or any suspicious noiseduring the vertical movement, please notify your Agfa NV Service contact at once.Customers were asked to complete the attached feedback form to confirm receipt. |  | 108 | 108 | Worldwide Distribution - US Nationwide in the states of CA, DC, FL, IL, IN, KS, LA, MA, MN, MO, NV, NY, OH, OR, SC and TN. Foreign distribution to Canada. | Open, Classified |  |  | 108
Z-3250-2018 | 80904 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard 360 Revision Knee System 75x 5mm Universal Tibial Block Augment, Model Number 185224 |  |  | August 02, 2018 | 2018 | Zimmer Biomet, Inc. | The implant only has one bolt in the sterile package but it should have had two. | Under Investigation by firm | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL" notification letters dated 8/2/18 to customers on 8/2/18. The letters described the product, problem and actions to be taken. The letters instruct customers to perform the following actions:Risk Manager Responsibilities:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com.Distributors:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.c. Include a hardcopy of Attachment 1 in each carton of your retur |  | 16 | 16 | Worldwide Distribution: US (nationwide) to the following US states: AZ, KS, NM, NV, and UT; and to countries of: Australia and Netherlands. | Open, Classified |  |  | 16
Z-3246-2018 | 80879 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural Nail Tibial Nail - Yellow 8.3mm, Model Number 47249534008 |  |  | July 30, 2018 | 2018 | Zimmer Biomet, Inc. | The item is labeled as an 8.3mm, but it is actually a 12 mm. | Mixed-up of materials/components | On 7/30/18, Urgent Medical Device Recall notification letters were distributed to customers. Distributors were instructed to do the following:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals and surgeons within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. Request a Return Authorization Number through zimmerbiometintlirarequests@zimmerbiomet.com for international returnsc. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.comd. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.e. Mark RECALL on the outside of the returned cartons.4. Retain a copy of your Inventory Return Certification and product return forms for your records in the event of a compliance audit of your facility.5. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com.Risk Managers and Surgeons were asked to do the following:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement an |  | 1 | 1 unit | The product was not distributed in the United States. The product was distributed in the Netherlands. | Open, Classified |  |  | 1
Z-3244-2018 | 80585 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | G8 Automated HPLC Analyzer - 723G8 |  |  | July 13, 2018 | 2018 | Tosoh Bioscience Inc | Analyzers were distributed with software which lacks a 510K | No Marketing Application | On July 13, 2018, Tosoh Bioscience started to telephone customers to set up on-site visits. On July 16, 2018, Tosoh Bioscience began to issue URGENT MEDICAL DEVICE RECALL notices to customers via hand delivery. Immediate Actions to be taken by the Customer/User:- Thoroughly review the content of this letter with the Tosoh field service representative.- Allow the Tosoh field representative to verify the software version currently on your analyzer- If the software is v5.23, your analyzer needs no further action- If the software is v5.24, allow the Tosoh field representative to download all patient results from your analyzer, remove software version v5.24, and upload software version v5.23.- Continue to use your analyzer and report results with software v5.23 once these activities have been completed.- Complete and return the Acknowledgement Form to the Tosoh field representative.- File the Urgent Medical Device Recall notice with your laboratory records and forward the information to others in your laboratory.If you have obtained unexpected test results or received any complaints of illness or adverse events associated with the use of the analyzer, please contact Tosoh Technical Support 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to contact Bernadette OConnell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will be available to answer any questions Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 8 | 8 units | US Nationwide Distribution in states of - CA, OK, VA 7 WI | Open, Classified |  |  | 8
Z-3254-2018 | 80947 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si/X/Xi Surgical System Redundant Medical Grade Power Supply (RMGPS) Units |  |  | July 16, 2018 | 2018 | Intuitive Surgical, Inc. | Certain Redundant Medical Grade Power Supply (RMGPS) units specific to Surgeon Console and Patient Side Carts were manufactured incorrectly and may fail prior or during use. | Nonconforming Material/Component | On 07/16/2018, the firm, Intuitive Surgical, sent an "Urgent Medical Device Correction" Notification to customers vi a Federal Express and email, and posted on the US Customer Portal . The letter described the product, problem and actions to be taken. The notice states to note that solely having the affected power supply module does not mean the system will encounter a power loss event. To encounter the related power loss scenario above, both of the following two (2) conditions must be met.1. The system contains the affected power supply.2. A failure of one of the power supply modules is experienced.The notice instructs customers to follow their specific system user manuals for guidance on how to proceed in the scenario where a loss of power is experienced on a console or cart. If the failure occurs to the Surgeon Console and the site has a dual console set up, the site may continue with the planned surgery robotically with one console.To correct this issue, an Intuitive Surgical Representative will schedule a site visit to provide a replacement RMGSP.The notice also instructs customer/user to, please take the following actions to ensure all affected personnel are fully informed of this Field Safety Notice. Forward this letter to your Risk Manager, OR Director, Purchasing, Biomedical Engineering staff, and other members of your medical staff who perform da Vinci procedures.1. You may continue using the system until the Intuitive Surgical representative replaces the unit. Please follow your specific system user manuals for guidance on how to proceed in the scenario where a loss of power is experienced on a console or cart.2. Ensure surgeons and patient side assistants using the da Vinci Si, Xi, and X Surgical System read and understand the contents of this letter.3. Inform affected personnel when corrective action has been completed.4. Complete the attached Acknowledgement Form and return it via fax or email to Intuitive Surgical per the instr |  | 1014 | 1,014 currently installed on systems | Worldwide Distribution: US (nationwide) and countries of: Austria, Australia, Belgium, Brazil, Canada, Switzerland, China, Czech Republic, Germany, Denmark, Ecuador, Spain, Finland, France, United Kingdom, Greece, Hong Kong, Ireland, Israel, India, Italy, Japan, Korea, Kuwait, Luxembourg, Mexico, Netherlands, Norway, New Zealand, Romania, Russia, Saudi Arabia, Sweden, Singapore, Turkey, Taiwan, and Venezuela. | Open, Classified |  |  | 1014
Z-3252-2018 | 80918 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet Sports Medii cine Juggerknot Long Flex Drill Bit with Sleeve Nitinol intended to be used for soft tissue to bone fixation with indications for use in: Shoulder, Foot and Ankle, Knee, Hand and Wrist and Hip repair.Stainless Steel, Sterile Item Number: 110016992 |  |  | August 03, 2018 | 2018 | Zimmer Biomet, Inc. | Expiration date incorrectly listed on the label | Device Design | Zimmer notified distributors on 8/3/18 distributors notified via email. Hospital risk managers, as well as distributors with product, notified via Fed'X.The letter identifies the issue and their responsibilities, locating and removing the product in their territory, as well as identifying hospitals who have fielded inventory. Product to be returned to Zimmer Biomet.Question customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday or to :CorporateQuality.PostMarket@zimmerbiomet.com |  | 217 | 217 units | AZCOFLFLGALAMAMAMONYSCTX | Open, Classified |  |  | 217
Z-3253-2018 | 80919 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic DxTerity(TM) TRA 5F, Diagnostic Catheter, REF U5TRAN35 |  |  | August 16, 2018 | 2018 | Medtronic Vascular | Medtronic has determined that the French size indicator on the inner pouch may incorrectly reflect a 6 French size, rather than the correct 5 French size. The catheters inside the pouch are the correct U5TRAN35 item, the outer carton has all the correct labeling, and all the other information and configuration details on both the carton and the pouch are correct. | Error in labeling | Medtronic initiated consignee communications on 16-Aug-2018 to Risk Managers and Health Care Professionals. US notifications were sent via 2-Day UPS delivery. Consignees are asked to quarantine and return any unused potentially affected product to Medtronic. |  | 250 | 250 units | CA, IA, NE, TN, TX, VA, WA, Jamaica, New Zealand, Thailand | Open, Classified |  |  | 250
Z-3251-2018 | 80308 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Visualase Thermal Therapy System, Software: Visualase Software versions 3.1.1 - 3.2, Product #: 9735542 |  |  | June 01, 2018 | 2018 | Medtronic Navigation, Inc. | Inaccuracy of MR thermometry during MRI-guided laser ablation procedures using the Medtronic Visualase Thermal Therapy System may result in unaccounted spread of thermal energy to the surrounding tissue. | Under Investigation by firm | On 06/01/2018, Urgent Medical Device Correction letters and Consignee Response Forms were mailed to customers via overnight mail service. Customers were provided with additional information intended to supplement the system manual and instructed to call the firm at 1-888-826-5603 if they had any questions. |  | 161 | 161 | U.S.: MN, NY, PA, FL, AZ, MA, VA, CA, CO, AL, WI, TX, MO, DC, OH, NH, GA, SC, NJ, WA, CT, MI, OK, IN, MD, IL, KY, LA, OR, RI, UT, AR, MS, TN, NCForeign: United Kingdom, Germany, Switzerland, Netherlands, Israel | Open, Classified |  |  | 161
Z-3255-2018 | 81001 | Yes | LWS |  |  | Cardiovascular | PMA | ESSENTIO Pacemaker |  |  | September 10, 2018 | 2018 | Boston Scientific Corporation | Boston Scientific has identified a subset of active pacemaker and CRT-Ps systems with anelevated potential for early pacemaker replacement due to hydrogen-induced accelerated battery depletion. | Component design/selection | On September 2018 Boston Scientific sent letters to their customers with the following information: Dear Doctor,Boston Scientific has identified a subset of approximately 2,900 active ACCOLADE", PROPONENT", and ESSENTIO" pacemakers and VISIONIST" and VALITUDE" cardiac resynchronization therapy pacemakers(CRT-Ps) with an elevated potential for early pacemaker replacement due to hydrogen-induced accelerated battery depletion. There have been no reports of injury related to this behavior. You are receiving this letter because you or your center has implanted or follow one or more patients with a pacemaker in this subset. This device behavior can be identified via the regular pacemaker follow-up process, either in-clinic or through the LATITUDE" NXT Remote Patient Management System (LATITUDE). Therefore, a follow-up interval of nomore than six (6) months is recommended and is consistent with existing international societal guidelines1. If accelerated depletion is suspected, Boston Scientific recommends consultation with Technical Services toreview the pacemaker diagnostic data available from either a recent data upload in LATITUDE or a Save to Disk to confirm accelerated depletion and determine an appropriate timeframe for pacemaker replacement.DescriptionThe pacemaker manual describes how increases in pacing requirements or changes in programmed parameters may result in an expected longevity reduction and are considered normal battery depletion. However, hydrogen exposure within the pacemakers circuitry may compromise the electrical performance of low voltage capacitors causing current leakage and a moderate acceleration in the rate of battery depletion. Because this accelerated depletion does not occur rapidly, a follow-up interval of no more than six months is recommended. Boston Scientific has determined a liner component to be the source of hydrogen and identified a subset of previously distributed pacemakers that have an elevated potential for exh |  |  |  | AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, W A, WI, WV, and WY | Open, Classified
Z-3263-2018 | 80921 | No | JDO | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | HipLOC Hip Fracture Fixation System-Hiploc System 135Âż Plate 38 mm Barrel 4 Hole Plate |  |  | August 09, 2018 | 2018 | Biomet UK Ltd. | HipLOC System Plate Barrel due to an undersized bore on the HipLOC plate barrel, can result in the lag screw not fitting through the barrel | Process design | Zimmer Biomet issued US recall notices issued on August 9, 2018 via email and FedEx to Distributors and Hospitals advising of the problem and instruction to remove and return.Your Zimmer Biomet sales representative will remove the affected product from your facility. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.comnistructions.mailto:CorporateQQuality.PostMarket@zimmerbiomet.com. |  | 44 | 44 units | US Nationwide in the states of AZ, MO, NJ and NM | Open, Classified |  |  | 61
Z-3279-2018 | 80264 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Discovery MR750w Nuclear Magnetic Resonance Imaging System labeled as the following product name sold under: |  |  | June 06, 2018 | 2018 | GE Medical Systems, LLC | GE Medical Systems has discovered a small area on the bore surface of the Discovery MR750w Nuclear Magnetic Resonance Imaging System, normally accessible to the patient, can get warm enough to cause a serious burn in the event the patient touches the bore and proper padding is not used. | Nonconforming Material/Component | GE Medical Systems notified customers on or about 06/06/2018 via a customer letter entitled "Urgent Medical Device Correction". The letter directed its customers to continue to use the system by following the existing Operator Manual Safety Information for proper patient padding and observation during scanning. Safety information included the following:WARNING: Place appropriate non-conductive padding between the patient and the bore wherever a portion of the body may come into contact with the magnet opening.CAUTION: Continuous patient observation and contact are required in all modes of operation.CAUTION: Extra attention should be utilized when scanning patients who are unconscious, sedated, or may have loss of feeling in any body part or are physically or mentally impaired and unable to alert the personnel.If a brown discoloration is seen on the inner surface of the patient bore, or if a burning smell is noticed, please contact your local Service Representative.The letter also stated GE Healthcare will correct all affected products at no cost. A GE Healthcare representative will contact customers to arrange for the correction. The firm could also be contacted by calling a GE Healthcare Service at 1-800-437-1171 or a local Service Representative. |  | 516 | 516 Units Total | Worldwide Distribution - US Nationwide and in the states of AK, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, US, UT, VA, WA, WI, and WY.and the countries of Algeria, Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bermuda, Brazil, Bulgaria, Canada, China, Egypt, France, Germany, Guam, Hungary, India, Indonesia, Ireland, Italy, Japan, Kazakhstan, Korea, (Republic of), Kuwait, Latvia, Lebanon, Martinique, Mexico, Netherlands, New Zealand, Norway, Poland, Qatar, Romania, Russia, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. | Open, Classified |  |  | 516
Z-3278-2018 | 81037 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Retractor System Elastic Stays (Part 3314-lG; 3316-lG; 3550-lG) |  |  | September 04, 2018 | 2018 | CooperSurgical, Inc. | There is a possibility that the seal of the sterile pouch may be compromised, thereby increasing the risk of infection. | Nonconforming Material/Component | Urgent Medical Device Recall Notification letters dated 9/4/18 were distributed to customers. The letters instruct customers to do the following: Please discontinue use of the product with any packaging irregularities, quarantine the product, and complete the attached Acknowledgement and Receipt Form to arrange for either a product replacement or credit to your account through CooperSurgical. Once completed and returned to CooperSurgical, a Customer Service Representative will contact you with a Return Merchandise Authorization (RMA) number and provide instructions for the return of product to CooperSurgical. If replacements are requested, a replacement order will be placed immediately. If you have no affected stock, please complete and return the enclosed Acknowledgement and Receipt Form, in order for us to document receipt of this letter. |  | 18000 | 18,000 units | Worldwide Distribution n- US Nationwide in the states of: AL, AR, AZ, CA, CT, DC, FL, GA, IA, ID, IL, KS, KY, LA, MA, MD, MN, MO, NC, NE, NJ, NY, OH, OR, PA, SC, TX, VA, WA and WV, and the following foreign countries of: Australia, Canada, Colombia, Czech Republic, Germany, Ireland, Israel, Italy, Lebanon, Netherlands, Russian Federation, Saudi Arabia, Singapore, Spain, Switzerland, Turkey, and the United Kingdom. | Open, Classified |  |  | 18000
Z-3268-2018 | 80837 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | Bausch + Lomb Stellaris Elite 23 GA Vit Cutter, Model BL5627, 7500 cpm, packaged in sterile pouches, 6/shipper. |  |  | August 08, 2018 | 2018 | Bausch &amp; Lomb Inc | The back cap separates from the body of the vitrectomy cutter. | Nonconforming Material/Component | Recall letters flagged URGENT MEDICAL DEVICE VOLUNTARY RECALL dated 8/8/2018 were issued 8/8/2018 via certified mail return receipt to the U.S. consignees. Letters for foreign consignees were dated 8/7/2018 and were flagged URGENT FIELD SAFETY NOTICE. |  | 103.33 | 103.33 cases | Distribution was made to AL, AZ, CA, CO, FL, GA, HI, IL, IN, KS, KY, MA, ME, MO, MT, NC, NJ, NY, OH, OK, OR, PA, TN, TX, UT, and WV. There was government distribution but no military distribution. Foreign distribution was made to Russia, Romania, Poland, The Netherlands, Jordan, Italy, Israel, Switzerland, Hungary, Greece, Guadeloupe, Great Britain, France, Spain, Singapore, Malaysia, and Thailand. | Open, Classified |  |  | 699
Z-3260-2018 | 81006 | Yes | LWS |  |  | Cardiovascular | PMA | Model 3300 LATITUDE(TM) Programming System |  |  | September 10, 2018 | 2018 | Boston Scientific Corporation | There is a potential for the Pacing System Analyzer (PSA) to exhibit unintended cross-chamber stimulation. | Device Design | The firm initiated the recall on 09/10/2018 to the US sales representatives by email. The sales representatives were asked to hand deliver the notice to the hospitals when the programmer is retrieved for that location. The sales representatives will electronically confirm the hand delivery of the letter to each affected customer after it occurs. |  | 74 | 74 units | Worldwide distribution - US Nationwide in the State of FL, UT, and VA, and countries of Austria, Czech Republic, Germany, Greece, Netherlands, Norway, Portugal, Spain, Sweden, United Kingdom, Switzerland | Open, Classified |  |  | 74
Z-3282-2018 | 80943 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | C-flex 570C +19.0D Intraocular Lens |  |  | August 09, 2018 | 2018 | Rayner Intraocular Lenses Limited | Firm become aware of reports of post-operative refractive errors following implantation of lenses. | Error in labeling | On August 31 the firm sent letters to all customers (health care facilities) that have received lenses in this batch and instructed to quarantine any remaining unused lenses from the C-flex 570C +19.0D batch 017100825. Replacement, reimbursement or substitution is offered (as appropriate). |  | 49 | 49 | US in the states of MO | Open, Classified |  |  | 49
Z-3281-2018 | 80652 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | AXS Infinity LS Long Sheath:((6F/.088in)/ x 70cm GEN-10800-70, (6F/.088in)/ x 80cm GEN-10800-80, (6F/.088in)/ x 90cm GEN-10800-90) Made in Ireland, Sterile EO. The AXS Infinity LS Long Sheath system contents are: One (1) AXS Infinity LS Long Sheath, One (1) Hemostasis valve with extension tubing, One Long Dilator. |  |  | May 02, 2018 | 2018 | Stryker Neurovascular | The AXS Infinity LS Long Sheath Dilator outer diameter (OD) specification was exceeded. | Nonconforming Material/Component | On 5/2/2018, "Urgent Medical Device Customer Communication" notifications were mailed via courier. Customers were informed to complete and return a Customer Acknowledgement Form. Customers with questions are encouraged to call (510) 413-2100. |  | 15 | 15 units | Only U.S. Consignees: CO,NJ,NY, OH,PA, and WV. | Open, Classified |  |  | 15
Z-3280-2018 | 80436 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Sodium Slides, Catalog # 837 9034 |  |  | May 05, 2018 | 2018 | Ortho-Clinical Diagnostics | Potential for positively biased results | Nonconforming Material/Component | On 05 May 2018, Customer letter (Ref. CL2018-123) was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all VITROS customers, who have a analyzer capable of running VITROS Na+ slides, to inform them of the issue and to request that they immediately discontinue use of this VITROS Na+ Slides, GENs 8 and 13. Continued use of VITROS Na+ Slides, GENs14,16,17&amp; 18 is acceptable provided the conditions detailed in the communication are met.Foreign affiliates were initially informed by email on 05 May 2018 of the issue and instructed to notify their consignees of the affected product with this issue and the requirement to immediately discontinue use of the VITROS Na+ slides GENs 8, 13, 14, 16, 17 and 18.Due to the amount of time required to ship replacement product to some worldwide locations, a revised communication to the foreign affiliates occurred on 08 June 2018. Due to the health risk associated with potentially leaving the laboratory without a Na+ method, the revised letter instructs customers that they may continue to use GENs 14, 16, 17 and 18 until replacement product is received. |  |  |  | US Nationwide, Bermuda, Austria, Brazil, Canada, Colombia, China, India, Japan, Mexico, Singapore, UK, France, Germany, Italy, Spain, Portugal, Poland, Russia, Denmark, Norway, Sweden, Belgium and the Netherlands | Open, Classified
Z-3283-2018 | 80925 | No | EOQ | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Bronchovideoscope BF-160 |  |  | August 08, 2018 | 2018 | Olympus Corporation of the Americas | The attachment of non-Olympus accessories to the bronchoscope s instrument channel port may have resulted in more applied force than expected and lead to loosening of the instrument channel port. | Other | Urgent Medical Device Safety Notice letters dated 8/8/18 were distributed to customers. The letters instructed customers to perform the following:Our records indicate that your facility has purchased one or more of the above-referenced endoscopes. OLYMPUS requires that you take the following actions:Inspect your inventory for the referenced Olympus models. Please keep this customer letter in your inventory and conduct the following action prior to every use. Step 1:Wearing sterile gloves, attempt to rotate or turn the instrument channel port in counter-clock direction as shown in picture 01. If the instrument channel port is able to be rotated or turned, do not use the endoscope and return it to Olympus for repair. If the instrument channel port does not rotate or turn, go to step 2.Step 2:Visually inspect the rubber part around the instrument channel port of the referenced endoscope. If the rubber part is in a normal condition as shown in picture 02, the endoscope can be used in a patient procedure. If the rubber part is lifted from the molding parts as shown in picture 03, do not use the endoscope and return it to Olympus for repair. If you have any questions about how to conduct this important inspection or you are not sure of the results of your inspection, you should contact the Olympus Technical Assistance Center (TAC), prior to using the device again. The phone number for TAC is 1-800-848-9024 option 1. To request service for your endoscope, contact Olympus at 1-800-848-9024, option 3.If you may have further distributed this product, please identify your customers and inform them at once of this action.Please complete and return the enclosed Reply Form to Olympus at facsimile (484) 896-7128. |  | 3948 | 3948 | The products were distributed US nationwide and Canada. | Open, Classified |  |  | 74372
Z-0004-2019 | 81010 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam Radiotherapy Delivery System version 2.0 and 2.5 |  |  | August 20, 2018 | 2018 | Varian Medical Systems | Reports have been received of an anomaly that can result in a treatment without intended gating (respiratory tracking/monitoring). This issue occurs when a patient planned with gating is treated on more than one system. | Software design | On 08/20/2018 the firm sent an URGENT MEDICAL DEVICE CORRECTION FIELD SAFETY NOTICE to customer via courier service (US and Canada via FedEx on 8/23/2018, and Rest of World via third party courier on 8/24/2018) informing them the firm had received reports of an anomaly that can result in a treatment without intended gating, and that this issue occurs when a patient planned with gating is treated on more than one system. Customers were advised, for treatment plans with intended gating, the firm recommends users monitor and pay attention to the system status message. When a patient with a gated treatment plan is moved to a second system, users must verify that intending gating was not overridden. If gating is overridden, users should reselect it from the Tools menu. The customer were informed that they will be contacted by a Customer Service Representative to schedule upgrade to the correction when available and if they have any questions regarding selecting gating from the Tools menu to contact customer support.Help Desk Contact Information:Phone: +1.888.827.4265Email: support @varian.comIn addition, the URGENT MEDICAL DEVICE CORRECTION FIELD SAFETY NOTICE, informs the customers to:1) Please retain a copy of this document along with your most current product labeling.2) Advise the appropriate personnel working in your radiotherapy department of the content of this letter.3) For future reference, this document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com.and In order to satisfy regulatory requirements, we request that you complete the attached Recall Return Response form provided. Kindly return the completed form to returnresponse@varian.com. |  | 387 | 387 | Worldwide Distribution - US nationwide in the states of AL, AR, AZ, CA, CO, FL, Guan, IL IN, KS, KY, LA, MA, MD, MI, MO, NH, NJ, NY, OH, OK, PA, Puerto Rico, TN, TX, VA, WI and the countries of :Algeria, Australia, Azerbaijan, Belgium, Brazil, Bulgaria, Canada, China, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, India, Ireland, Israel, Italy, Japan, Kazakhstan, Lebanon, Morocco, Nepal, The Netherlands, New Zealand, Norway, Panama, Poland, Reunion Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey and United Kingdom | Open, Classified |  |  | 391
Z-0003-2019 | 80313 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | JOURNEY BCS Knee Conventional Polyethylene Inserts: 74023111, 74023112, 74023113, 74023114, 74023115, 74023116, 74023117, 74023118, 74023121, 74023122, 74023123, 74023124, 74023125, 74023126, 74023127, 74023128, 74023131, 74023132, 74023133, 74023134, 74023135, 74023136, 74023137, 74023138, 74023141, 74023142, 74023143, 74023144, 74023145, 74023146, 74023147, 74023148, 74023211, 74023212, 74023213, 74023214, 74023215, 74023216, 74023217, 74023218, 74023221, 74023222, 74023223, 74023224, 74023225, 74023226, 74023227, 74023228, 74023231, 74023232, 74023233, 74023234, 74023235, 74023236, 74023237, 74023238, 74023241, 74023242, 74023243, 74023244, 74023245, 74023246, 74023247, 74023248, 74023251, 74023252, 74023253, 74023254, 74023255, 74023256, 74023257, 74023258, 74023261, 74023262, 74023263, 74023264, 74023265, 74023266, 74023267, 74023268, 74023271, 74023272, 74023273, 74023274, 74023275, 74023276, 74023277, 74023278, 74023281, 74023282, 74023283, 74023284, 74023285, 74023286, 74023287, 74023288, 74023311, 74023312, 74023313, 74023314, 74023315, 74023316, 74023317, 74023318, 74023321, 74023322, 74023323, 74023324, 74023325, 74023326, 74023327, 74023328, 74023331, 74023332, 74023333, 74023334, 74023335, 74023336, 74023337, 74023338, 74023341, 74023342, 74023343, 74023344, 74023345, 74023346, 74023347, 74023348, 74023411, 74023412, 74023413, 74023414, 74023415, 74023416, 74023417, 74023418, 74023421, 74023422, 74023423, 74023424, 74023425, 74023426, 74023427, 74023428, 74023431, 74023432, 74023433, 74023434, 74023435, 74023436, 74023437, 74023438, 74023441, 74023442, 74023443, 74023444, 74023445, 74023446, 74023447, 74023448, 74023451, 74023452, 74023453, 74023454, 74023455, 74023456, 74023457, 74023458, 74023461, 74023462, 74023463, 74023464, 74023465, 74023466, 74023467, 74023468, 74023471, 74023472, 74023473, 74023474, 74023475, 74023476, 74023477, 74023478, 74023481, 74023482, 74023483, 74023484, 74023485, 74023486, 74023487, 74023488, 74023648, 74203326 |  |  | June 13, 2018 | 2018 | Smith &amp; Nephew, Inc. | The data in the National Joint Registry of England, Wales and Northern Ireland (NJREWNI) and Australian Orthopaedic Association National Joint Replacement Registry (AOANJRR) indicate that the first generation JOURNEY BCS Knee System has a revision rate over 1.5 times the primary total knee arthroplasty device class average revision rates reported in those registries. | Other | The Field Safety Notice (FSN) sent to risk managers on 06/13/2018 asks customers to inspect inventory, locate and quarantine affected, unused devices. Affected product subject of the voluntary removal is to be returned to Smith &amp; Nephew. All consignees are asked to complete an FSN Response Form, even if they do not have product to return.A similar notification was addressed to implanting surgeons, requesting their receipt and acknowledgement of the FSN. |  | 64642 | 64642 | Worldwide distribution, including US nationwide and Puerto Rico, Austria, United Arab Emirates, Australia, Belgium, Canada, China, Switzerland, Germany, Denmark, Finland, France, Great Britain, India, Spain, Italy, Netherlands, Portugal, Sweden, Venezuela, South Africa. | Open, Classified |  |  | 117484
Z-0011-2019 | 80569 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | ellex I.Science iTRACK 250A, Canaloplasty Microcatheter Kit, for single use, REF iT-250A, Rx. The firm name on the label is Ellex iScience, Inc., Fremont, CA. Used in ophthalmic surgery. |  |  | January 11, 2018 | 2018 | Ellex iScience, Inc. | Sterility failure found in one of the units in the lot. | Process control | The recalling firm called their consignees on 1/11/2018 instructing them to quarantine unused devices and monitor the patients operated with the devices in this lot. The firm issued a letter dated 1/16/2018 via FedEx Express on 1/17/2018 informing them a potential issue had been identified and to segregate the product immediately. |  | 90 | 90 units | Distribution US nationwide and Switzerland. | Terminated | October 02, 2018 | 264 | 90
Z-3370-2018 | 80811 | No | BSO | Anesthesia conduction catheter. | Anesthesiology | Anesthesiology | 510(k) | CE17TKPS CONTINUOUS EPIDURAL ANES TRAY, Material Number 332233 |  |  | July 20, 2018 | 2018 | B. Braun Medical, Inc. | The catheter connectors may not stay closed during in and in some cases, leakage or disconnection of the catheters from the catheter connectors has been observed. | Under Investigation by firm | The firm, B. Braun, sent an "URGENT MEDICAL DEVICE CORRECTION  RECALL NOTIFICATION" letters dated 7/20/18 to its customers. The letters described the product, problem and actions to be taken. The letters instruct customers to do the following: 1. Review this Medical Device Correction Notification in its entirety and ensure that all users of the mentioned products in your organization and other concerned persons are informed about this Medical Device Correction  Recall Notification. If you are a distributor, please forward this recall notification to your customers.2. For continued safe use of the Perifix Catheter Connector as part, review and follow the instructions described in Attachment 2, Step-by-Step Correction of the Perifix Catheter Connector using Label .3. For continued safe use of the Perifix Catheter Connector, review and follow the instructions described in Attachment 3, Step-by-Step Correction of the Perifix Catheter Connector using Cloth/Silk Medical Tape . This method can be applied to the Perifix Catheter Connector in Procedural Trays.4. Utilizing the attached Medical Device Correction  Recall Notification Acknowledgement form, please acknowledge that you have received and reviewed this information.5. Return the completed Medical Device Correction  Recall Notification Acknowledgement form to B. Braun Medical Inc. Quality Assurance department by faxing to: 610-849-1197 or e-mailing to: PA_QualityAssurance.BBMUS_Service@bbraunusa.com the form within two (2) weeks of receipt of this notice. It is important this form is returned so B. Braun Medical Inc. can meet regulatory requirements of the United States Food and Drug Administration.6. If you have any affected products as identified in attachment 1 that you choose to return instead of following the continued safe use instructions within Attachment 2 and 3, please identify on the Medical Device Correction  Recall Notification Acknowledgement form that you will be returning t |  | 2400 | 2400 | Worldwide Distribution: US (nationwide) and countries of: Australia, Canada, Guatemala and Kuwait. | Open, Classified |  |  | 1208477
Z-0010-2019 | 80718 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABL90 FLEX analyzers with software versions 3.3MR1 or below. The ABL90 FLEX Analyzer is a portable, automated Analyzer that measures pH, blood gases, electrolytes, glucose, lactate, bilirubin and oximetry in whole blood. |  |  | May 25, 2018 | 2018 | Radiometer America Inc | The ABL90 FLEX can aspirate Cal 2 solution instead of rinse solution, when the following take place:Âż The analyzer is in ready modeÂż Operator lifts the inlet and then closes it again, thereby initiating a rinseÂż Immediately after lifts the inlet and quickly closes it again | Device Design | The firm, Radiometer America, Inc., notified customers via letter dated 05/25/2015. The letter provided the following: Reason for recall, Affect product description and codes, Risk for the patient, What customers must do, Countermeasure provided by Radiometer, contact information and Response Form. The customers were instructed to do the following:1. Inform users of the ABL90 FLEX analyzer that the inlet is not to be repeatedly lifted and closed. If this does happen, the easiest way to reset the analyzer flow selector manually is to re-install the solution pack. Follow the procedure described in the IFU for installation of a solution pack. You may re-use the same solution pack.2. We recommend that you review previous results, where cNa+ and cCl- results are off from the expected results with the values found in table 1.3. Complete the recall response fax form (last page of this customer advisory letter) and submit to your Radiometer representative via Fax to:714 582 0806.If you have any questions, please contact your Radiometer representative or Quality Assurance/ Regulatory Affairs Manager at 1-800-736-0600 opt.4. |  | 12102 | 12,102 | US Distribution and Internationally to Canada. | Open, Classified |  |  | 12102
Z-3210-2018 | 80758 | No | MNL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Endoscopic Cleaning Brush, single use, sold in kits and labeled as the following: |  |  | June 15, 2018 | 2018 | Medline Industries Inc | There is potential risk of brush detachment during cleaning. If brush detachment is not identified, there is the potential for the brush to be left within the scope. | Under Investigation by firm | Medline notified customers on about 06/15/2018 via URGENT RECALL MEDLINE INDUSTRIES, INC. IMMEDIATE ACTION REQUIRED letter. Accounts that purchased a custom kit or standard Pack were instructed to immediately check inventory for affected devices, quarantine any affected devices, affix the provided stickers on affected kits and continue use but were instructed to not use the Endoscopic Cleaning Brush in the kit and to instead pull another approved product from inventory. They were also instructed to complete and return the Verification Form. Accounts that just purchased the Endoscopic Cleaning Brush were instructed to immediately check inventory for affected devices, destroy any affected devices, notify customers if the product has been further distributed, and complete and return the provided destruction form listing the quantity of affected products destroyed or that there was no product on hand to destroy. |  | 183606 | 183,606 devices total | Nationwide distribution to AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, VA, VT, WA, WI, WV, WY | Open, Classified |  |  | 183606
Z-3187-2018 | 80672 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Freedom EV0-2 with any version of EVOware and EVOware Sample Tracking Add-On V1.0 or higher |  |  | July 23, 2018 | 2018 | Tecan US, Inc. | There is a software issue that can potentially impact sample identification and traceability under specific conditions when using EVOware Add On Software: Sample Tracking -Multi Instrument V1.0 or higher, Part Number 30019971. | Device Design | On July 23, 2018, Tecan sent Field Corrective Action Letters to their consignees advising them of the problem. Field corrections will be made in the near future. |  | 19 | 19 EVOware Sample Tracking Software Add-on | Nationwide distribution to CA, MA, NC, OH, TX, and VA. Worldwide distribution to Australia, Canada, China, EU, and Japan. | Open, Classified |  |  | 19
Z-0368-2019 | 80730 | No | LBH | Cavity varnish. | Dental | Dental | 510(k) | 3M Fast Release Varnish, 5% Sodium Fluoride, Mint Flavor, Rx Only, single use. 200 pack, with each package containing 1 0.40 ml varnish and 1 applicator brush. |  |  | July 06, 2018 | 2018 | 3M Company - Health Care Business | Product has an unpleasant flavor due to a manufacturing process error resulting in an extraneous material being introduced into the product. | Mixed-up of materials/components | On or about 07/06/2018, 3M Health Care notified consignees via phone call or "URGENT Medical Device Recall" letter. Distributors were instructed to examine inventory for affected product, complete and return the Distributor Inventory Checklist to 3M Oral Care, return any affected product on hand, and provide the attached Dentist Letter and Dentist Inventory Checklist to any dentists that may have received the affected product. Dentists were instructed to examine inventory for affected product, complete and return the Dentist Inventory Checklist to 3M Oral Care, and return any affected product on hand.If you have any questions, please contact Regulatory Affairs Manager by telephone at 651-575-8052. |  | 73 | 73 units (200 pack) | US distribution to states of: CA, FL, IN, NV, PA, TN, TX and WA. | Terminated | February 26, 2019 | 235 | 73
Z-0372-2019 | 80927 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | EVO Excel, Emission Computed Tomographic System, Material Number 10910501 |  |  | October 05, 2018 | 2018 | Siemens Medical Solutions USA, Inc. | The Evo Excel External-ECG Power-Connection current limit is mis-labeled as 1.0A. It should be labeled 0.55A., | Other | Siemens disseminated the notices by letter on 10/05/2018. Firm will deliver a Symbia Evo Excel Addendum and DVD addendum with this updated information at a later date. |  | 264 | 264 units | Worldwide | Open, Classified |  |  | 264
Z-0360-2019 | 81093 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | CoaguChekÂż XS System Prothrombin time test: |  |  | September 13, 2018 | 2018 | Roche Diagnostics Corporation | Abnormally high INR test results with the affected CoaguChek test strips | Nonconforming Material/Component | Roche issued URGENT MEDICAL DEVICE RECALL AMENDMENT lettersdated 11/5/18 advising customers patients, HCPs, IDTFs, HCP distributors and HCP/patient distributors.Distributors are advised to take the following actions:- Read the Urgent Medical Device Recall Amendment for Healthcare Professionals and Clinics Furnishing PT/INR Test Strips to Patients for Home Testing Use (TP-00454).- Distribute a copy of the Urgent Medical Device Recall Amendment TP-00454 to healthcare professionals to whom you have directly shipped affected test strips.- Stop distribution of all affected CoaguChek XS PT Test Strip lots.- Complete all sections of the fax form (TP-00462) and fax it to number 1-888-627-2279 or email it to roche3866@stericycle.com.- Return all affected product following the instructions outlined in this Urgent Medical Device Recall Amendment.Independent Diagnostic Testing Facilities should take the following action:- Read the Urgent Medical Device Recall Amendment for CoaguChek Patients (TP-00453) and Urgent Medical Device Recall Amendment for Healthcare Professionals and Clinics Furnishing PT/INR Test Strips to Patients for Home Testing Use (TP-00454).- Distribute a copy of the Urgent Medical Device Recall Amendment for CoaguChek Patients (TP-00453) to your CoaguChek patient self-testers.- Distribute a copy of the Urgent Medical Device Recall Amendment for Healthcare Professionals and Clinics Furnishing PT/INR Test Strips to Patients for Home Testing Use (TP-00454) to your Healthcare Professional customers.- Stop distribution of all affected CoaguChek XS PT Test Strip lots.- Complete all sections of the fax form (TP-00461) and fax it to number 1-888-627-2279 or email it to roche3866@stericycle.com.- Return all affected product following the instructions outlined in the Urgent Medical Device Recall Amendment.Healthcare Professionals and Clinics Furnishing test strips to Patients for Home Use should take the following action:- Stop usin |  | 1103860 | 1,103,860 | Nationwide | Open, Classified |  |  | 1103860
Z-0373-2019 | 81192 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Sterile Knee Instruments Disposable Fluted Headless Pins, DIA 3.2mm LNTH 89mm, REF 7650-2038A - Product Usage: Howmedica Osteonics Corp. (Stryker, Joint Reconstructive Division) Sterile Instrumentation consists of manual surgical instrumentation intended for use during hip, knee, shoulder, elbow or wrist arthroplasty, or trauma surgeries. |  |  | September 10, 2018 | 2018 | Howmedica Osteonics Corp. | Potential for outer pouch sterile barrier to be compromised | Under Investigation by firm | All affected consignees were notified via 2-day UPS letter (with return receipt). The Urgent Product Recall notice was sent on September 11, 2018 and an updated version dated November 26, 2018 was distributed by Strykers third party recall company Stericycle.Actions to be Taken by Customer:1. Please inform users of this Urgent Product Recall UPDATE and forward this notice to all those individuals who need to be aware within your organization.2. Hospitals/Branches/Agencies: Complete and sign the enclosed Recall Notification Business Reply Form and fax a copy to 1-888-312-5960 or email toStrykerortho5807@stericycle.com |  | 15 | 15 pouches | US Nationwide in the states of AL, FL, IL, MI, NJ &amp; OH. | Open, Classified |  |  | 15
Z-0367-2019 | 81240 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | EXACTAMIX Empty EVA Bags - 3000 mL Product Usage: A plastic container used to hold large-volume sterile solutions to be administered to a patient through an intravascular administration set. |  |  | September 12, 2018 | 2018 | Baxter Healthcare Corporation | Potential ability to leak once used for compounding. | Process control |  |  | 39400 | 39400 units | Worldwide distribution - US Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WV. Countries of Canada, Australia, New Zealand, China, Hong Kong, Brazil, Chile, and Uruguay. | Open, Classified |  |  | 212000
Z-0362-2019 | 81208 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Volumat MC Agilia Volumetric Infusion Pump, Model No. Z021135, Software version 1.7. Product Usage - The Volumat MC Agilia is a piece of transportable medical equipment intended to be used in healthcare facilities environment by healthcare trained professionals according to hospital protocols on adults, pediatrics and neonatal human patient to administer via a single channel or mounted on a multiple channels rack accessory: Intermittent or continuous delivery of parenteral fluids (solutions, colloids, parenteral nutrition) and medications (including but limited to diluted drugs, chemotherapy) through clinically accepted IV routes of administration). Transfusion of blood and blood derivatives products. Note that for transfusion procedures, use of a special administration set is required). |  |  | December 02, 2017 | 2017 | Fresenius Vial Sa | Fresenius Kabi initiated a mandatory software upgrade of the Volumat MC Agilia Volumetric Infusion Pump because it was designed to improve customer experience while preventing common software malfunctions. | Software design | A recall notification dated 11/17/17 was sent to customers to inform them that Fresenius Kabi initiated a mandatory software upgrade of the Volumat MC Agilia Volumetric Infusion Pump. The new software version 1.9a is currently available and designed to improve customer experience while preventing common software malfunctions. Customers will be contacted by a Fresenius Kabi service representative or distributor to schedule a service appointment for upgrading the software. Customers with any questions or require additional information are instructed to contact the Fresenius Kabi Technical Hotline at (800) 937-5060. |  | 7082 | 7082 units | Worldwide distribution - US Nationwide and countries of Canada and China. | Open, Classified |  |  | 7082
Z-0287-2019 | 81127 | No | LOE |  |  | Orthopedic | PMA | SpF PLUS-Mini (60 ÂżA/W) Implantable Spinal Fusion Stimulator |  |  | February 19, 2018 | 2018 | Zimmer Biomet, Inc. | Lack of adequate validation and controls to ensure that product cleanliness parameters were consistently met. | Under Investigation by firm | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 2/22/2018 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 4879 | 4,879 units (2,982 units have been implanted and 537 have been scrapped) | US Nationwide Distribution: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV ; and country of:Australia. | Open, Classified |  |  | 4879
Z-0376-2019 | 81065 | No | ELR | Root canal post. | Dental | Dental | 510(K) Exempt | ParaPost Plus 25 Titanium Alloy Spiral Vented and |  |  | August 21, 2018 | 2018 | Coltene Whaledent Inc | ParaPost Plus Titanium Posts contain stainless steel post that are a size 4 | Process control | The firm, Coltene, issued an "Urgent Medical Device Recall" Letter dated August 20, 2018 to the customers on August 21st via email with a delivery and read receipt requested. The letter described the product, problem and actions to be taken. The customers were instructed to return the product to Coltene/Whaledent Inc.; return any unused product and we will replace your product at no cost, and complete and return the Return Response Form via email: tricia.cregger@coltene.com or Phone: 330-916-8817.If you should have additional questions, please contact Regulatory Affairs Manager, at 330-916-8904 and Tricia.Cregger@coltene.com. |  | 30 | 30 packets | US Distribution to states of: CA, IL, NV, NY, PA, and WA' and to country of: Canada. | Open, Classified |  |  | 30
Z-0374-2019 | 81418 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS-IW with Universal Viewer version 5.0.x with PACS-IW foundation displays medical images (including mammograms) and data from various imaging sources. |  |  | October 11, 2018 | 2018 | GE Healthcare, LLC | Synchronization failure | Software design | On October 12, 2018 GE Healthcare send recall notification to all their consignees with the following instructions;If you do not use a Diagnostic viewer (e.g. Centricity Universal Viewer Zero Footprint Client or any 3rd party dicom viewer) connected directly to the same Centricity EA to display images for diagnostic purposes, your system is not affected by the safety issue.If you do use a Diagnostic viewer (e.g. Centricity Universal Viewer Zero Footprint Client or any 3rd party dicom viewer) connected directly to the same Centricity EA to display images for diagnostic purposes, please contact GE Healthcare service representative immediately. GE Healthcare will assist you in determining if your system is affected and, if necessary, will assist you with a solution to confirm that study updates are made in the Enterprise Archive, until GE Healthcare can install a correction on your system.GE Healthcare will provide a correction that will be installed by a GE service engineer. A GE Healthcare representative will contact you to arrange for the correction. This activity will be performed at no cost to you.Contact InformationIf you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 13 | 13 | CA, CO, DE, FL, HI, NY, OK, PA, SD, TX, WA, WI, and WVBelgium, Brazil, Canada, Colombia, France, Germany, Ghana, Ireland, Italy, Mexico, Netherlands, Panama, Poland, Saudi Arabia, South Africa, Spain, Switzerland, and Turkey. | Open, Classified |  |  | 86
Z-0379-2019 | 81161 | No | KGN |  |  | General &amp; Plastic Surgery | 510(k) | Endoform Dermal Template 4x5, SKU 529313 |  |  | November 27, 2018 | 2018 | AROA BIOSURGERY | Potential for pouch seal failure | Employee error | The firm sent a Customer Notification letter dated November 27, 2018. The letter identified the affected product, problem and actions to be taken. Customers are encouraged to visually inspect the seals on devices from these lots prior to use with particlar care. Should you observe a comprimised seal, please discard the device and select a new, fully sealed device for use. |  | 8853 | 8,853 total boxes | US Nationwide Distribution | Open, Classified |  |  | 8853
Z-0383-2019 | 80862 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EZ1 Advanced XL, Model Number: 9001492, GTIN: 04053228010760 |  |  | July 23, 2018 | 2018 | Qiagen Sciences LLC | An issue with the power branching printed circuit board (PCB) on the instrument that has the potential to trigger a heater malfunction was discovered. If this failure were to occur, nucleic acid extractions could be compromised and extractions need to be repeated. This could cause a delay in obtaining results from downstream applications. | Nonconforming Material/Component | The firm sent an Urgent Notice of Product Field Action letters dated July 23, 2018 to customers. The letter identified the affected product, problem and actions to be taken by the customer/user/commercial partner in order to assure that your instrument is currently functioning properly. Customers were asked to complete and sign the included Acknowledgement of Receipt form and either email it to QIAGEN Technical Service at Techservice-eu@qiagen.com or fax it to +49 (0)2103-29-22400. For questions contact your local QIAGEN Technical Services Department. |  | 298 | 298 | Worldwide Distribution - US Nationwide in the states of: AZ, CA, FL, GA, IA, ID, IL, MA, MD, MI, MN, NE, NJ, NY, NV, OH, PA, PR, SD, VA, and WA.The products were distributed to the following foreign countries: Alger, Argentina, Australia, Austria, Canada, China, Dubai, Egypt, France, Germany, Greece, Hungary, India, Italy, Japan, Jordan, Korea, Latvia, Mexico, Morocco, Norway, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Tunisia, Turkey, UK. | Open, Classified |  |  | 298
Z-0381-2019 | 81159 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid Cardiac Control, CQ5052 |  |  | October 02, 2018 | 2018 | Randox Laboratories Ltd. | Randox has previously issued a recall for Liquid Cardiac Control CQ5053 lot 4245CK under REC334 on June 8, 2018. The firm has now confirmed the Liquid Cardiac Controls CQ5051, CQ5052 and CQ5053 are no longer suitable for the control of the Troponin T assay due to unacceptable variation between vials. | Under Investigation by firm | On October 2, 2018, the firm notified distributors of the recall via Expanded Urgent Medical Device Correction email. Customers were informed that the recall was related to previous recall REC334 on June 8, 2018. The firm has confirmed that Liquid Cardiac Controls CQ5052 and CQ5053 are no longer suitable for the control of the Troponin T assay due to unacceptable variation between vials.Customers were advised to take the following actions:- Discontinue use of these products for the Quality Control (QC) monitoring of Troponin T assays.- Discuss the contents of this notice with your Medical Director.- Update kits with revised IFUs excluding Troponin T values and the attached important notice to prevent further use of the device in the QC of Troponin T assays.- Complete and return the vigilance response section of this form to technical.services@randox.com within five working days.- Contact your local Randox sales representative for alternative product details.You may contact the firm at technical.services@randox.com. |  | 1738 | 1738 | Distributed to IN, NY, PR, and WV. | Open, Classified |  |  | 5609
Z-0385-2019 | 80956 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Alaris Infusion Pump administration sets labeled as: |  |  | April 20, 2018 | 2018 | CareFusion 303, Inc. | Complaints where users were unable to prime the administration set. | Process control | A "Medical Device Safety Notification" Letter dated April 20, 2018 was to all affected customers, stating that the firm has become aware of issues with infusion pump administration sets. Customers are experiencing difficultly or failing to prime the pump administrative sets. Infusion pump administration sets, various models with check valves manufactured during January 2017 through March 2018.Issue:The Infusion pump administration sets require priming prior to patient use as described in the directions for use. It is during the priming process that this issue can be observed. During priming, when attempting to fill the drip chamber, the drops into the drip chamber will not be observed once the roller clamp is opened. This issue is highly detectable.Steps for detecting this issue during priming:-Open administrative set package, remove set and close roller clamp.-Insert administrative set spike into the IV bag or bottle and hang container 20 inches above the Pump module.-Fill drip chamber to 2/3 full.-Open the roller clamp.-Drips into the drip chamber will be very slow or absent.Potential Risk:An occlusion in the administration set line can prevent the ability to prime the set, rendering the administration set unusable. The inability to prime can contribute to the delay of the start of the infusion. There have been no reports of patient harm as a result of this issue.Recommended Steps for Users:You can continue to use the product as this issue is highly detectable. If during priming you experience a difficultly or inability to prime, please notify CareFusion Customer Advocacy at 888-812-3266 or customerfeedback@BD.com. Meantime, BD has the following options if you experience this issue:In some instances clinicians may be able to increase the pressure to assist the check valve by gently squeezing the IV bag following these steps:1. Begin by removing the IV set from the package and close the roller clamp.2. Spike the bag of IV flu |  | 118254342 | 118,254,342 | Worldwide Distribution - US Nationwide including GuamInternational: Canada, Australia, New Zealand, EU, Middel East, South Africa, Okinawa, Saipan, Buenos Aires, Lima, Selangor, Hong Kong, Taiwan, and Budapest | Open, Classified |  |  | 118254342
Z-0386-2019 | 81181 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | PCEA ASV Yellow Microbore, Catalog Numbers 30893, 30893-07 |  |  | September 26, 2018 | 2018 | Becton Dickinson &amp; Company | The products have the potential to leak between the connection of the male luer lock and the yellow striped tubing. | Under Investigation by firm | The initial recall letters were sent out to customers on September 26, 2018. An update was sent to customer on October 5, 2018 that included additional lots. The letters identified the affected product, problem and actions to be taken.Contact Information:For BD Support Center contact Phone: 888-562-6018 Email: SupportCenter@CareFusion.comFor Customer AdvocacyPhone: 888-812-3266Email: customerfeedback@bd.com |  | 566737 | 566,737 total | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MO, MT, NC, NE, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and Guam.The products were distributed to the following foreign countries: Canada, United Arab Emirates and Saudi Arabia. | Open, Classified |  |  | 566737
Z-0384-2019 | 81040 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | BD NEXIVA Closed IV Catheter System Dual Port 18GA 1.25 IN (BD Nexiva Catheter), Catalog Number 383539 |  |  | August 23, 2018 | 2018 | Becton Dickinson &amp; Company | One lot of the specified needles is not performing as intended due to a failure in the needle tip shield/safety mechanism, which may lead to an increased risk of a contaminated needlestick injury. | Under Investigation by firm | Urgent Medical Device Recall notification letters dated 8/23/18 were sent to customers. The letters instruct customers to perform the following actions:1. Immediately review your inventory for the specific Catalogue (Ref) and lot number listed above and destroy all product subject to the recall following your institutions process for destruction.2. Share this recall notification with all users of the product to ensure that they are also aware of this recall.3. Complete the attached Customer Response Form and return to the BD contact noted on the form whether or not you have any of the impacted material so that BD may acknowledge your receipt of this notification and process your product replacement.4. Report any adverse health consequences experienced with the use of this product to BD. Events may also be reported to the FDA's MedWatch Adverse Event Reporting program.Web: MedWatch website at www.fda.gov/medwatch Phone: 1-800-FDA-1088 (1-800-332-1088)Mail: MedWatch, HF-2, FDA, 5600 Fishers Lane, Rockville, MD 20852-9787 Actions TakenDistributors were instructed to:1. Immediately review your inventory for the specific Catalogue (Ref) and lot number listed above and return all product subject to recall according to the Packing Instructions enclosed.2. Identify all your customers that purchased any affected product, as defined in this recall notification. E-mail an excel file listing of all customers to BDMPSRC@bd.com, within 72 hours of receipt of this letter.3. Complete the attached Wholesaler/Distributor Response Form and return to the BD contact noted on the form whether or not you have any of the impacted material so that BD may acknowledge your receipt of this notification and process your credit accordingly.4. Report any adverse health consequences experienced with the use of this product to BD. Events may also be reported to the FDA's MedWatch Adverse Event Reporting program.Web: MedWatch website at www.fda.gov/medwatchPhone: 1 |  | 133230 | 133,230 units | Worldwide Distribution -US Nationwide in the states of: CA, CO, FL, GA, IL, LA, MI, ND, NE, NY, OH, PA, RI, SC, VA, WA and WV.The products were distributed to the following foreign countries: Argentina, Australia, Belgium, Brazil, Canada, Colombia, Costa Rica, Mexico and New Zealand. | Open, Classified |  |  | 133230
Z-0389-2019 | 81160 | No | LBR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific AutoHDL/LDL Cholesterol Calibrator, Catalog Number H7545-CAL - Product Usage: For the calibration of Pointe Scientific s autoHDL and autoLDL Cholesterol Reagent Set in serum or plasma. For in vitro diagnostic use only. |  |  | August 10, 2018 | 2018 | Medtest Holdings, Inc. | Vials of the autoHDL/LDL Calibrator were shipped without a vial label. | Error in labeling | Notification letters dated 8/14/18 were sent to customers. Customers were also contacted via telephone on 8/10/18. Additional letters containing labels dated 8/17/18 were sent to customers. |  | 22 | 22 | US Nationwide in the states of CA, FL, GA, MI, and OH. | Open, Classified |  |  | 22
Z-0390-2019 | 81294 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis POWERFLEX PRO PTA Dilatation Catheter - Product Usage: Is intended to dilate stenoses in iliac, femoral, ilio-femoral, popliteal, infra popliteal and renal arteries and for the treatment of obstructive lesions of native or synthetic arteriovenous dialysis fistulae. The device is also indicated for postdilatation of balloon-expandable and self-expanding stents in the peripheral vasculature. |  |  | September 05, 2018 | 2018 | Cordis Corporation | Product may not meet the internal manufacturing shaft subassembly burst strength specification (Failure of the Burst Shaft Test). | Process control | A letter dated September 5, 2018 was sent to all their customers by Cordis with the following request:1) Read this Urgent Medical Device Recall letter.2) Immediately check your inventory to confirm that you do not have any unitsfrom the affected lots in your possession. Identify and set aside any unitsfrom the identified lots in a manner that ensures the affected product will notbe used. Check all storage and usage locations.3) Review, complete, sign and return the enclosed Acknowledgement Formdirectly to Cordis at the fax number on the form.4) Return any affected product to the address listed on the form, with referenceto your Customer Number which is listed on the form.5) Share this letter with others in your facility who need to be made aware ofthis recall.6) Please contact any other facility that may have been sent the affected unitsof POWERFLEXÂż PRO product from your facility. If any units of the affectedlots are found to be at the other facility, please arrange the return of theunits.7) Maintain awareness of this notice until all affected product has been returned to Cordis8) Keep a copy of this notice with the affected product.For any health care professional with medical concerns, please contact theCordis Endovascular Medical Information Line 877-DEVICE5 (877-338-4235,Monday through Friday from 9:00 AM to 5:00 PM EST.For questions related to the recall that are not adequately addressed in thisletter, please contact the market action team at:GMB-FieldCorrectiveAction@cardinalhealth.comFor questions related to the Acknowledgement Form and its return, please call (800)292-9332. |  | 162 | 162 units total (74 units in USA) | Worldwide distribution - US Nationwide in the states of AL, CA, FL, GA, IL, IN, KS, KY, LA, MA, MD, MO, MS, NC, ND, NJ, OH, PA, SC, TN, TX, and WI. Countries of Japan and Netherlands. | Open, Classified |  |  | 162
Z-0393-2019 | 81072 | No | GXN | Preformed nonalterable cranioplasty plate. | Neurology | Neurology | 510(k) | Biomet Microfixation RapidFlap System 12mm SpinDown Clamp, sterile |  |  | August 22, 2018 | 2018 | Zimmer Biomet, Inc. | Spin Down RapidFlap nut that interfaces with the outer plate was assembled incorrectly. | Process control | Zimmer Biomet conducted recall to the consumer/user level, including any intermediate wholesale or retail consignees. Distributors notified initially via email on 8/22/18. The firm expanded the recall on 10/5/18 to include all lots with lot expiry before 09/18/2023, letters issued via email and FedEx to Distributors/ Risk Managers/Surgeons.Letter identifies issue and requesting removing and return of product.Surgeons will be provided with a letter that identifies the issue and their responsibilities, which include:- Following any specified patient monitoring instructions- Returning Certificate of Acknowledgment to Zimmer Biomet.Hospital risk managers will be provided with a letter identifying the issue.Questions or concerns after reviewing this notice, please call customer service at 907-741-4400 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to Benjamin.Blystone@zimmerbiomet.com . |  | 12825 | 12825 units | Worldwide distribution: US Nationwide, CanadaArgentinaAustraliaBelgiumBrazilChileColombiaCosta RicaCzech RepublicDominican RepublicEcuadorFinlandFranceChinaEcuadorGermanyGreeceHong KongIndonesiaItalyJapanJordanKoreaLebanonMalaysiaNetherlandParaguayPeruSingaporeSouth AfricaSpainSwedenSwitzerlandThailandTrinidad and TobagoTurkeyUnited Arab EmiratesUnited KingdomUruguayVietnam | Open, Classified |  |  | 69293
Z-0392-2019 | 81066 | No | CDD | Vitamin B[bdi1][bdi2] test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Immunodiagnostic Products Vitamin B12 Reagent Pack 1/2, Catalog Number 1453489 |  |  | August 28, 2018 | 2018 | Ortho-Clinical Diagnostics | The reagent pack exhibits the potential for an increased frequency of calibration failures. | Under Investigation by firm | The firm, Ortho Clinical Diagnostics, sent an "IMPORTANT PRODUCT CORRECTION NOTIFICATION" letter dated 8/28/18 to customers. A second "IMPORTANT UPDATE: PRODUCT CORRECTION NOTIFICATION" letter dated 10/22/18 was distributed to customers to inform them of additional affected lots. The letters described the product, problem and actions to be taken. The customers were instructed to do the following:If this lot was successfully calibrated and your quality control results are acceptable; you may continue using your existing inventory of VITROS Vitamin B12 Reagent Pack 1/2 Lot 2601 until your replacement order arrives. We advise that you perform quality control on every reagent pack." If this lot cannot be successfully calibrated, contact our Ortho Care" Technical Solutions Center." Complete the Confirmation of Receipt  Response Required form and return by September 12, 2018 (so that we can credit your account or expedite your replacement order) via Fax: to: 1.888.557.3759 or 1.585.453.4110 or e-Mail to: RA-OCDUS-CONFIRMAD@ORTHOCLINICALDIAGNOSTICS.COM." Discontinue using Lot 2601 upon receipt of your replacement order." Post this notification by each system that processes VITROS Vitamin B12 Reagent Pack 1/2." Forward this notification if the product was distributed outside of your facilityThe distributors were instructed to do the following:Discontinue distributing and discard all inventory of VITROS Vitamin B12 Reagent Pack 1/2 and Calibrators for Lot 2610 and/or 2620." Send the customer letter to all customers who were shipped VITROS Vitamin B12 Reagent, Lots 2610 and/or 2620 from your facility." Complete and return the Confirmation of Receipt form no later than November 5, 2018." Forward this notification if you have distributed this product outside of your facility.If you have further questions, please contact our Ortho Care" Technical Solutions Center at 1-800-421-3311. |  | 6630 | 6630 | Worldwide Distribution: US (nationwide) to the following US states: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, PA, SC, SD, TN, TX, VT, WA, WI, WV, and WY; and countries of: Belgium, Bermuda, Brazil, Canada, Chile, Columbia, Denmark, France, Germany, India, Italy, Portugal, Spain, and UK. | Open, Classified |  |  | 6630
Z-0399-2019 | 81185 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | StageOne Hip, Models 431207 &amp; 431209Usage: |  |  | September 21, 2018 | 2018 | Zimmer Biomet, Inc. | There is a potential for comingling of the 43MM and 51 MM molds | Error in labeling | On September 21, 2018, Zimmer Biomet issued URGENT MEDICAL DEVICE RECALL notices and Return Certification Forms to customers via courier service. Distributors are encouraged to take the following actions:Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete the Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. For each return, send a copy of Return Certification Form to CorporateQuality.PostMarket@zimmerbiomet.com.c. Include a hardcopy of the Return Certification Form in each carton of your return shipment for immediate processing.d. Mark RECALL on the outside of the returned cartons.Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com.a. Review the list of hospitals included with the email notification sent to your facility, which includes a list of hospitals that have already been notified of this recall.b. Identify whether there are any additional hospitals that Zimmer Biomet has not notified and list these accounts on the Additional Accounts form. Please provide the form in Excel format.c. If there are no additional accounts to notify, please indicate that there are no additional accounts, or indicate None or NA on the form.5. Retain a copy of your Inventory Return Certification and product return forms for your records in the event of a compliance audit of your facility.Risk Managers are encouraged to take the following actions:Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove |  | 72 | 72 units | US Nationwide Distribution in the states of: AZ, FL, MI, OH, VA &amp; WI | Open, Classified |  |  | 72
Z-0269-2019 | 81164 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | REF VV19F Reinforced Dual Lumen Cannula, 19Fr., Sterile EO, DISPOSABLE: DO NOT RE-USE CE 0473 |  |  | March 30, 2015 | 2015 | OriGen Biomedical, Inc. | Potential for clear extension tube to separate from the hub, which could result in blood loss and may require intervention to prevent permanent impairment/damage | Mixed-up of materials/components | On April 1, 2015 the firm sent an MEDICAL DEVICE MARKET WITHDRAWAL to customers/distributors via email to inform them the catheters may start to show signs of failure at the junction of the clear extension tube and blue hub connection that can be recognized by air entrainment or blood leakage. If the preliminary failure is identified, the catheter should be replaced immediately. Serious injuries could occur due to the failure mode. The customers are instructed to discontinue use of all products from the affected lot and to return any unused products immediately and to complete and return the acknowledgement and receipt form even if the customer does not currently have products in its inventory.The firm is conducting an investigation to determine root cause of the failure and will notify the customers when the investigation has been completed and when replacement products are available. If customers have any questions regarding the return of product or the status of the investigation, they can contact:OriGen BiomedicalAttn: Bernie R Silvers7000 Burleson Road, Building DAustin, TX 78744Phone Number +1 512-474-7278 |  | 134 | 134 | Worldwide Distribution - US Nationwide in the states of: AR, AZ, CA, FL, GA, IL, IN, LA, MA, MI, MO, NC, NY, OH, OR, PA, RI, TX, WA, and WIOUS: Belgium, Canada, Germany, Kingdom of Saudi Arabia, Poland, Spain Sweden, Thailand and United Kingdom | Open, Classified |  |  | 134
Z-0403-2019 | 81293 | Yes | NTG |  |  | Orthopedic | PMA | DT Medi Tech Hintermann Series Talar Implant LEFT-SIZE 4 |  |  | September 25, 2018 | 2018 | DT MedTech, LLC | Incorrect package labeling identifying the device | Packaging process control | DT Med Tech LLC initially notified accounts by email notification on 9/25/18 followed by notification letter; The letters explain the process for theidentification and return to HealthLink BV (distribution warehouse in The Netherlands) of the products affected (lot AAADP/AAAEM). |  | 1 | 1 | Australia, Ireland, Korea, Germany | Open, Classified |  |  | 55
Z-0404-2019 | 81355 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic Navigation 0-arm 02 Imaging System interventional fluoroscopic x-ray system: |  |  | September 17, 2018 | 2018 | Medtronic Navigation, Inc.-Littleton | O-arm 02 Imaging Systems correction to software version 4.1.0; new version of the user manual and an Image Acquisition System ("IAS") product label; replacing blown fuse of the new AC power input circuit | Software design | Medtronic Navigation issued Urgent Medical Device Correction letter on 9/27/18, notification letter advising, Medtronic field service personnel will install 4.1.0 software onto the designated consignee systems, update IFU and product labels as well as install a new designed AC power input circuit within the Mobile View Station {MVS).Questions regarding contact Medtronic Technical Support at 1-800-595-9709. |  | 641 | 641 units | Worldwide Distribution - US NationwideForeign:AUSTRALIAAUSTRIABAHRAINBANGLADESHBELGIUMBRAZILCANADACHILECOLOMBIADENMARKFINLANDFRANCEGERMANYHRICELANDINDIAIRELANDISRAELITALYJAPANJORDANKOREA, REPUBLIC OFLUXEMBOURGMARTINIQUEMEXICONETHERLANDSNEW ZEALANDPOLANDPORTUGALPUERTO RICOQATARSAUDI ARABIASINGAPORESOUTH AFRICASPAINSWEDENSWITZERLANDTAIWANTHAILANDTURKEYUNITED ARAB EMIRATESUNITED KINGDOM | Open, Classified |  |  | 641
Z-0405-2019 | 81247 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | FlareHawk Interbody Fusion System (previously XCage Interbody Fusion System) Instructions for Use (IFU's) only. Product Usage: Indicated for spinal intervertebral body fusion with autogenous bone graft in skeletally mature individuals with degenerative disc disease. |  |  | February 27, 2018 | 2018 | Integrity Implants Inc. | Recently Integrity Implants uncovered a discrepancy in the cleaning parameters in several of their labeling material. The firm is initiating the field correction to ensure that updated information is provided to all accounts. | Under Investigation by firm | The firm is instructing customers to verify their inventory and return the recalled labeling material and replace them with the revised instructions and send the tracking form back to the firm. |  | 75 | 75 sets | US nationwide in the states of AZ, CA, CO, DC, FL, GA, IA, IL, IN, KS, LA, MA, MS, NJ, NV, NY, OH, OR, PA, RI, SC, TX, and VA. | Open, Classified |  |  | 75
Z-0401-2019 | 81150 | No | NAE | Hydrogel wound dressing and burn dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Simpurity HYDROGEL ABSORBENT SHEET WOUND DRESSING ADHESIVE BORDER, 4" X 5", 12 Pads, Catalogue number SNS58345 |  |  | August 22, 2018 | 2018 | Safe N Simple, LLC | The sterilization process used was unqualified. | Process control | Safe n Simple extended the recall to the only customer that ordered the product in the lot numbers (see table). The firm provided a written notice to the customer mentioned in the table with specific instructions along with a phone conversation. Safe n Simple instructed the customer to provide the inventory level of SNS58315 and SNS58345 specific to lot numbers 18060564 and 18060563 respectively and to return such product to Safe N Simple LLC at Safe N Simple LLC expense. The effectiveness of the recall was good since the customer immediately responded by shipping all their inventory back to us. |  | 125 | 125 cases | CA | Open, Classified |  |  | 229
Z-0391-2019 | 80984 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | TB Series Bipolar Pacing Leads, French Size 4F, 5F, 6F for Curve Type Straight, Atrial J, 60 degree Curve, and Right Heart. TB is packaged for sterile single use in a double tray. The lead is placed in a PETG inner tray that is sealed with a Tyvek inner lid. The sealed inner tray is placed in an outer PETG tray. The outer tray is sealed with an outer Tyvek lid. An inner label is placed on the outer Tyvek lid. The sealed trays are placed in a labeled white shelf box with on Instruction for use. Packed on lead per outer white shelf box. |  |  | September 27, 2018 | 2018 | Oscor, Inc. | During the use of some TB - Temporary Bipolar Pacing Leads, featuring the 2mm unshrouded connectors, the connector cap housing may slide and potentially expose the connection wire. In some instances, this may cause the wire to be more susceptible to loss of connectivity or breakage during movement of the cables causing interruption of the pacing system. | Device Design | A press release for the recall was issued on 09/27/2018. Additionally, on about 09/27/2018, Oscor notified customers via email and mail with a "PRODUCT RECALL - ACTION REQUIRED" letter. For distributors, instructions included to examine inventory for affected devices, quarantine any affected devices, notify customers if the affected devices have been further distributed, retrieve those affected devices from customers, and complete and return the attached Exhibit A. Any affected devices on hand and at customers should not be destroyed, but returned to Oscor. For hospitals, instructions included to examine inventory for affected devices, quarantine any affected devices, and complete and return the attached Exhibit A. Any affected devices on hand and at customers should not be destroyed, but returned to Oscor. Questions can be directed to TB Recall Coordinator at 727-937-2511 or e-mail TB@Oscor.com. Oscor business hours are Monday to Friday from 8:30AM to 5:30PM Eastern Time. On 11/07/2018 Oscor expanded the recall and notified customers via "PRODUCT RECALL - ACTION REQUIRED" letter and press release. The expansion includes expired devices distributed by Oscor from 12/21/2011 to 04/30/2015, to result in affected devices distributed from 12/21/2011 to 05/17/2018. Instructions for distributors and hospitals remain the same as the original notification. |  | 36688 | 36,688 units | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, CO, FL, HI, IN, KS, KY, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PR, RI, TN, TX, UT, VA, WA. Countries of AUSTRIA, BENIN REPUBLIC, BOLIVIA, BRAZIL, BULGARIA, CANADA, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, GEORGIA, GERMANY, GREECE, GUATEMALA, INDIA, IRAN, IRAN, ISRAEL, ITALY, JORDAN, KAZAKHSTAN, KUWAIT, LATVIA, LEBANON, MACEDONIA, MAURITIUS, MEXICO, NEW ZEALAND, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, RUSSIA, SAUDI ARABIA, SENEGAL, SOUTH AFRICA, TAIWAN, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIETNAM. | Open, Classified |  |  | 36688
Z-0407-2019 | 81300 | No | OSR |  |  | Cardiovascular | PMA | Medtronic CareLink Encore 29901 Programmer |  |  | October 02, 2018 | 2018 | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | Vulnerabilities have been identified in the SDN download process that may allow an individual with malicious intent to update the programmers with non-Medtronic software during an SDN download. | Software design | The firm began disseminating the Field Corrective Action notices by hand delivery on 10/02/2018 and followed with the posting of Medtronic Security Bulletin and the Urgent Medical Device Correction customer letter online on 10/11/2018.The notice explained vulnerabilities have been identified in the SDN download process that may allow an individual with malicious intent to update the programmers with non-Medtronic software during an SDN download and beginning October 11th, 2018, Medtronic will be disabling the SDN for programmer updates and will rely solely on the USB update method. |  | 4072 | 4072 units | Worldwide | Open, Classified |  |  | 34244
Z-0408-2019 | 81352 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife M6, Part Number 054000-004 |  |  | September 26, 2018 | 2018 | Accuray Incorporated | A robotics supplier notified Accuray of two manufacturing variations in the fastening of the in-line wrist and the casting of the wrist for the robot. The variation may result in premature failure of the component. | Process control | The firm, KUKA, issued a Field Safety Notice dated 9/13/2018 and disseminated it beginning on 09/26/2018 by mail, email, and hand delivery. The Field Safety Notice explained the issues and stated that all units would be inspected. Should a unit be faulty, the firm will arrange for remediation.If you have any technical questions, please feel free to get in touch with KUKA Account Manager or contact our hotline via phone 1-800-459-6691 or email: FocusCenter@Kuka.com. |  | 93 | 93 units | worldwide distribution. | Open, Classified |  |  | 93
Z-0443-2019 | 81323 | general purpose reagent | PPM | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | bioMÂżrieux NucliSENS easyMAG Lysis Buffer ref. 280134 |  |  | October 08, 2018 | 2018 | bioMerieux, Inc. | Some eluates become colored because of remaining hemoglobin in the eluate. | Nonconforming Material/Component | Required ActionsWe request you to take the following actions at this time:-Please distribute this information to all appropriate personnel in your laboratory, retain a copy in your files, and forward this information to all parties that may use this product, including others to whom you may have transferred our product.-Stop using the NUCLISENS Lysis Buffer, Ref. 280134, Lot: Z019EA1LB, exp. 28-MAY-2019, and destroy any stock of this lot Z019EA1LB you might have in your laboratory.-Discuss any concern you may have regarding previously reported results with your Laboratory Medical Director to determine the appropriate course of action.-Contact your local bioMerieux, Inc. Customer Service representative if you have observed the issue and if you have a doubt regarding your results.-Complete the attached Acknowledgement Form in Attachment A to confirm receipt of this notice and return it to your local bioMÂżrieux, Inc. representative.bioMÂżrieux, Inc. is committed to providing our customers with the highest quality product possible. We sincerely apologize for any inconvenience that this may have caused your business. If you require additional assistance or have any questions, please contact your local bioMÂżrieux, Inc. Clinical Customer Service representative at (800) 682-2666 or Industry Customer Service representative at (800) 634-7656. For non-urgent inquiries, please contact the Clinical Technical Support Center by email at CustomerService-ImmunoMolecular@biomerieux.com. The anticipated turnaround time for email inquiries is 24-48 hours during normal business hours (7AM  7PM EST), and up to 72 hours dependent upon weekends and company observed holidays. |  | 268 | 268 | CA, FL, GA, MD, MI, MN, NC, NJ, NM, NY, OH, OR, RI, SC, TN, TX, VA, WA, and WIAustria, Australia, Belgium, Switzerland, Colombia, Germany, Spain, France, UK, India, Italy, Japan, Netherlands, Portugal, Sweden, Singapore, South Africa, Israel, Dutch Antilles, Angola, Estonia, and Bangladesh | Open, Classified |  |  | 268
Z-0442-2019 | 81119 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 3.2mm Proximal Reamer/Cannulated Drill provided as part of the Fibulock Fibular Nail System Pack. FibuLock Nail Procedure Pack: Contains instrumentation for the implantation of the FibuLock implant. Contents include: Actuation Driver, 6.2mm Reamer, 3.2mm Reamer, 2 mm Drill, Spade Tip Guide Wire, and 1.6mm x 12 K-Wire. |  |  | September 13, 2018 | 2018 | Arthrex, Inc. | There is potential to break during use. | Software change control | Arthrex sent letters to their consignees with the following instructions.1. Immediately discontinue use, sale, and distribution of the affected product.2. Please contact Arthrex Product Surveillance at 866-267-9138 orcomplaints@arthrex.com as soon as possible to arrange for return of theaffected product. Our Product Surveillance Specialists can provideassistance regarding alternative solutions and are available to answerquestions regarding credit for affected devices in your possession.3. If you are an Arthrex agency, you do not need to notify any customers asArthrex will be notifying customers directly.4. If you are an Arthrex customer, you do not need to notify any patients ofthis product recall.5. If you have any questions about this product recall, please contact ArthrexProduct Surveillance at 866-267-9138 or complaints@arthrex.com. |  | 832 | 832 | Distributed throughout the U.S. to the following states: NC, CA, NJ, CO, NV, UT, LA, KY, OH, AZ, MO, MN, TN, NY, WI, MA, PA, TX, IL, PA, AR, FL, IA, MI, AL, SD, OR, WA, MD, KS, and GA. | Open, Classified |  |  | 832
Z-0413-2019 | 81182 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Touchscreen display kit. Sold under the following product names: TOUCHSCREEN DISPLAY KIT 11.X ENGLISH FOR USA, TOUCHSCREEN DISPLAY KIT 11.X ENGLISH, TOUCHSCREEN DISPLAY KIT 11.X GERMAN, |  |  | September 27, 2018 | 2018 | GE Healthcare, LLC | Upgraded Aisys anesthesia devices and Aisys CS2 anesthesia devices with PSVPro Spontaneous Breathing Modes do not provide the PSVPro ventilation mode feature of transitioning to SIMV PCV backup ventilation mode. PSVPro Spontaneous Breathing Modes do not provide the feature of transitioning to SIMV PCV backup ventilation mode for a patient who stops spontaneous breathing if the PSVPro ventilation mode is in use prior to the clinician starting the Cycling Procedure. | Software design | GE Healthcare notified customers on about 09/27/2018 via "URGENT MEDICAL DEVICE CORRECTION" letter. The letter informed customers that they could continue to use the anesthesia device. The issue can only occur if the patient stops spontaneous breathing in PSVPro ventilation mode after completion of lung recruitment breaths. Stopping the Cycling Procedure resolves the issue. Continue to monitor patient as is normal for Apnea. A GE Healthcare representative will contact each customer to arrange for the correction. Questions or concerns can be directed to GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 3700 | 3,700 units total | Worldwide distribution - US nationwide in the states of CA, CO, DC, FL, GA, HI, IA, IL, IN, KY, LA, ,MA, MD, MI, MN, MO, MT, NC, NE, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV. Countries of Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea (Republic of), Kuwait, Macedonia, Mexico, Netherlands, New Zealand, Nigeria, Norway, Peru, Portugal, Qatar, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam. | Open, Classified |  |  | 3700
Z-0410-2019 | 81121 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | R&D; systems, a bio-techne brand Human Total 25-OH Vitamin D IVD ELISA Kit, REF RDKAP1971 |  |  | September 11, 2018 | 2018 | R &amp; D Systems, Inc. | The kit was provided without a copy of the lot-specific kit booklet and the Calibrator 0 vial 18D1/4 label incorrectly specified reconstitution volume of 2mL. | Unknown/Undetermined by firm | The firm contacted the consignees (end users) by telephone on 09/11/2018 to alert them of the issue. The firm followed with an email on 09/21/2018 restating the issue and offering to replace any unused kits. |  | 8 | 8 kits | IA and CO | Open, Classified |  |  | 8
Z-0409-2019 | 81278 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Thermal Chamber Motor Cable located in the following models: |  |  | September 28, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | Thermal chamber motors have electrical cables with single insulation rather than the required double insulation. There is no exposed current due to the single insulation that could result in a potential electrical shock | Device Design | The firm, Siemens Healthcare, issued on 09/28/2018, "Urgent Medical Device Correction" DI-18-02.A.US (US customers)/Urgent Field Safety Notice DI-18-02.A.OUS ( for OUS customers), informing them of the issue and providing them with the steps Siemens is taking to correct the issue. A Siemens Service Engineer will schedule a Service visit to inspect the instrument and install a protective cover, if necessary, to correct the issue. Customers were instructed to perform the following actions:"Please arrange a time for your system upgrade when contacted by Siemens. "Return completed Field Correction Effectiveness Form to Siemens Healthcare Diagnostics via fax to the Customer Care Center  Technical Solutions at 312-275-7795 or please send a scanned copy of the completed form via email to uscctsfcaecfax.team@siemens-healthineers.com. "retain this letter with your laboratory records, and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center, your local Siemens Technical Support Representative or the Quality Engineer QT, LD, Commercial Product Quality at 302-631-0528 or email: richard.andersen@siemens-healthineers.com. |  | 277 | 277 units | Worldwide Distribution: US (Nationwide) and countries of: Algeria, Argentina, Australia, Brazil, Canada, China, Columbia, Croatia, Egypt, France, Germany, Great Britain, Greece, Hungary, India, Italy, Japan, Malaysia Mexico, Netherlands, Poland, Portugal, Russia, Saudi Arabia, Slovenia, South Africa, South Korea, Spain, Tanzania, Thailand, Tunisia, Turkey, and United Arab Emirates. | Open, Classified |  |  | 277
Z-0414-2019 | 81282 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | BioMend Extend, REF# 0142Z, Synthetic bone grafting material. |  |  | August 16, 2018 | 2018 | Collagen Matrix, Inc. | There is a possibility that the product was packaged in the wrongbox. Specifically, the product may have been placed inside a BioMend box instead of aBioMend Extend box. The only difference between the two boxes is theproduct trade name. The label placed on the box and on the product packaged inside the box iscorrect. The Instructions for Use inside the box are correct. | Employee error | On August 16, 2018, Zimmer Biomet Dental issued Urgent Medical Device Recall Notices to affected customers. The letter described the issue and specified that there was no projected risk to patient health. CLINICIAN RESPONSIBILITIES:Complete the attached Certificate of Acknowledgement within 15 days of receipt and email thecompleted form to postmarket@zimmerbiomet.com or return the form to:Product RecallZimmer Biomet4555 Riverside Dr.Palm Beach Gardens, FL 33410If any of the items noted in the scope of this letter remain in your possession, please returnthem along with a copy of the form to the address shown above. Replacement product will beprovided upon receipt of the return.If you have any questions after reviewing this notice, please call 800-342-5454 between 8:00am and 5:00 pm EST, Monday through Friday. Calls received outside of call center operatinghours will receive a prompt to leave a voicemail. |  | 569 | 569 | Distributed to one distributor located in Florida. | Open, Classified |  |  | 569
Z-0416-2019 | 81166 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | OviTex Reinforced BioScaffold 6x10cm, Part Number F10254-0610G |  |  | April 06, 2018 | 2018 | AROA Biosurgery, LTD. | Degradation of the PGA suture material used in the manufacture of the resorbable mesh devices was observed during an on-going product stability study. Further investigation indicated that devices over 18-months showed evidence of no longer meeting the pre-defined ball burst specification respective to the number of tissue layers. | Other | Beginning in April 2018, customers were visited by representatives and the affected units were replaced. |  | 881 | 881 total | The products were distributed to the following US states: AL, CA, FL, IN, MA, MI, NH, and NY. | Open, Classified |  |  | 881
Z-0439-2019 | 81211 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | The Spirit Select Bed, Model 5700 Spirit Plus /with high rail siderail option |  |  | October 08, 2018 | 2018 | Stryker Medical Division of Stryker Corporation | The beds with the high rail side rail option may not be complaint with IEC 60601-2-52/C1:2010. There is a potential risk of entrapment. | Device Design | Medical Device Correction notification letters were distributed to customers on 10/8/18. |  |  |  | The products were distributed to the following US states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY.The products were distributed to the following foreign countries: Canada. | Open, Classified
Z-0441-2019 | 81217 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Heron toilet and bath chair Model 880505-01 &amp; 880505-11 |  |  | June 30, 2017 | 2017 | R82 A/S | Replacement campaign on all units that have a gas spring head in aluminum. | Other | The firm sent a letter ask you to replace the gas spring heads on these Herons  the replacement kits are being sent to your attention. Please carry out the required retrofitting yourself or forward the kits and instructions to those that maintain the product on your behalf. This replacement must take place as soon as you possibly can and no later than 31-12-2017 and the reporting form documenting this must be filled in for each unit and returned to R82 A/S. |  | 687 | 687 | NationwideDenmark, Germany, Spain, Finland, France, Great Britain, Island, Netherlands, Norway, Poland, Switzerland, Japan, Australia and New Zealand | Open, Classified |  |  | 687
Z-0440-2019 | 81171 | No | NEK |  |  | Orthopedic | PMA | INFUSE Bone Graft X SMALL KIT REF 7510100 |  |  | September 27, 2018 | 2018 | Medtronic Sofamor Danek USA Inc | The Extra Small INFUSE Bone Graft kit from lot M111806AAA, may contain two absorbable collagen sponges (ACS), but should only contain one ACS. | Error in labeling | The firm initiated the recall by letter on 09/27/2018. The firm will coordinate the return of the recalled product. |  | 138 | 138 units | US | Open, Classified |  |  | 138
Z-0444-2019 | 81003 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Espiner Tissue Retrieval System-Eco-Sac |  |  | June 20, 2018 | 2018 | Espiner Medical Ltd | Label was incorrect. The label states that thevolume of the retrieval bag was 1200ml, but it should state 1000ml | Error in labeling | Espinar Medical notified the distributor to notify accounts to review inventory, product batches should not be used, and should be segregated if deemed necessary by the hospital/healthcare facility. Complete the acknowledgement form. Contact Global Medical Partners LLC via email at cs@globalmedllc.com or by calling 888.384.8490 to obtain an RGA number for the return of the product. |  | 230 | 230 units | US Nationwide in the states of CA, IL, MT, NC, NJ, | Open, Classified |  |  | 230
Z-0445-2019 | 81220 | No | PAM | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | VERIGENE Blood Culture Gram-Positive Nucleic Acid Kits |  |  | September 14, 2018 | 2018 | Luminex Corporation | Potential to lead to a false negative results. | Nonconforming Material/Component | On September14, 2018, Luminex sent Urgent Customer Advisory Notifications to their consignees. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Experience Team. |  | 12700 | 12,700 trays | Worldwide Distribution - US Nationwide in the states of United States - AZ, CA, CO, CT, DC, FL, GA,, IA, ID, IN, KA,, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NV, NJ, NY, OK,,OH, OR, PA, SC, TN, TX, UT, VT, WA, KS, and VA..France | Open, Classified |  |  | 12700
Z-0446-2019 | 81197 | No | PLO | Device to detect and identify microbial pathogen nucleic acids in cerebrospinal fluid. | Microbiology | Microbiology | 510(k) | -Outer Kit Box (P/N RFIT-LBL-0140) - FilmArrayÂż Meningitis/Encephalitis (ME) Panel REF RFIT-ASY-0118 30 For FilmArray systems LOT KIT UDI (01)00815381020123 CE 0086 IVD |  |  | September 10, 2018 | 2018 | BioFire Diagnostics, LLC | Firm identified an issue in the manufacturing process which might contribute to elevated rates of false positive results for HSV-1 while using the ME Panel. Only specific kits are affected. | Material/Component Contamination | On 09/10/2018, the firm sent an Urgent Medical Device Recall Letter via FedEx to the Recall Coordinator/Laboratory Director for all consignees.Consignees are instructed to:-Immediately examine their inventory for products with the affected Lot numbers and to discontinue use and to discard any remaining product in their possession that is subject to the recall.. BioFire will replace discarded product at no charge.-If the consignee has further distributed the affected product, they are to identify these customers and to notify them at once of the product recall.-Complete and return the Acknowledgement of Receipt Form that was enclosed with the Customer Advisory Notice letter, as soon as possible, and to return the Acknowledgement Form to the firm.Consignees were provided with firm contact information including phone numbers: 801-736-6354 and 1-800-735-6544 |  | 1350 | 1350 pouches | Worldwide Distribution - US Nationwide in the states of: CA DC, FL, IN, KY, MN, MO, NB, NJ, OH, TX and WI.International: Canada, Columbia and Mexico | Open, Classified |  |  | 1350
Z-0456-2019 | 80754 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General Hospital | 510(k) | Powdered Surgical Gloves, labeled as: |  |  | December 28, 2016 | 2016 | American Contract Systems, Inc. | Powdered Surgical Gloves are an FDA banned substance. | Other | The firm, ACS, notified Customers via phone call and email on about 12/28/2016. The customers were informed about the product, problem and actions to be taken. The customers were instructed to identify and remove all the affected kits for replacement of the recalled gloves; tray insert label must reflect the omission of the powered gloves; and ban, quarantine and return affected product.If you have any questions, contact that Quality Manager at 816-920-5846 or email to: tbakke@amconsys.com. |  | 1180 | 1,180 kits | US Distribution to states of: IL, MO, NE, and OH. | Terminated | November 14, 2018 | 686 | 1180
Z-0462-2019 | 81092 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Xper Flex Cardio Physiomonitoring System, Software version 1.6.0.0388 |  |  | September 24, 2018 | 2018 | Invivo Corporation | The Invasive Blood Pressure (IBP) numeric values will freeze on the display when any of the following functions are performed: change an IBP Site Label using the Site Label icon or the keyboard, activate or deactivate the mean pressure waveform display option (On/Off) using the Invasive Pressure Options Dialog or the keyboard, activate or deactivate the waveform hide option (On/Off) using the Invasive Pressure Options Dialog, activate or deactivate the pressure waveform filter option (On/Off) using the keyboard, and/or perform a Pullback using the Pullback icon or the keyboard | Software Design Change | Customers were notified on about 09/20/2018 via an "URGENT - Medical Device Correction" letter. An initial letter was sent to distributors with instructions to provide the Field Safety Notice to all customers the devices were further distributed to. The "URGENT - Medical Device Correction" letter provided guidance on how to identify the Service Number and serial number on the Flex Cardio device, which is on the bottom right corner of the back of the device. Actions to be taken by the customer/user until Philips contacts the customer to arrange for a software update to address this issue, is if the IBP numeric values become frozen and do not update in real-time, opening the Waveform Setup Dialog and then immediately pressing the OK button will restore the proper display of the IBP numeric values. Alternative methods of performing the functions listed in the Problem Description section that do not cause the problem to occur include the following: 1. Change an IBP Site Label using the Pressure Setup Dialog or a macro; 2. Activate or deactivate the mean pressure waveform display option (On/Off) by using the Waveform Setup Dialog, the Pressure Setup Dialog or a macro; 3. Activate or deactivate the waveform hide option (On/Off) using the Waveform Setup Dialog, the Pressure Setup Dialog, the keyboard or a macro; 4. Activate or deactivate the pressure waveform filter option (On/Off) using a macro. A Philips representative will contact each customer regarding their affected Flex Cardio devices. A software update will be required on both the affected Flex Cardio device and on the host workstations driver. These actions will be implemented free of charge by Philips. |  | 76 | 76 units | Distributed to the Netherlands | Open, Classified |  |  | 76
Z-0458-2019 | 81514 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Q biplane with Material 10848282- |  |  | October 19, 2018 | 2018 | Siemens Medical Solutions USA, Inc | Vibrations of the C-arm of the plane B may result in reduced image quality caused by the screw connection to the C-arm in the ceiling support | Device Design | Siemens issued on 10/19/18 an Urgent Medial Devoce Correction notification to via AX033/18/S advising users of the problem, health risk, steps to be taken. Siemens has initiated screw connection exchange of the ceiling bearing via Update Instruction AXOSS/17 /S. |  | 54 | 54 units | US Nationwide Distribution | Open, Classified |  |  | 83
Z-0460-2019 | 81444 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | RadPRO(R) Mobile 40kW Digital X-Ray System - Model: SM-40HF-B-D-C |  |  | November 02, 2018 | 2018 | Virtual Imaging, Inc. | There is a potential for the RadPRO Mobile 40kW Digital X-Ray System Model SM-40HF-B-D-C to separate from the Tube Head Mount and potentially fall off the system. | Component change control | The firm initiated the field correction by issuing a field safety notice on 11/03/2018. The letter identified the affected product, problem and action to be taken. |  | 253 | 253 units | Worldwide Distribution - US Nationwide and the countries of, Canada, Italy | Open, Classified |  |  | 253
Z-0450-2019 | 81223 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Affixus Hip Fracture Nail Left 125 Deg 9x165mm, Item Number 816709165 |  |  | October 02, 2018 | 2018 | Zimmer Biomet, Inc. | The set screw is not able to advance or reverse from the original position in the nail. | Mixed-up of materials/components | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL  LOT SPECIFIC" notification letters dated 10/2/18 to customers. The letters described the product, problem and actions to be taken. The customers were instructed to1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send toCorporateQuality.PostMarket@zimmerbiomet.com or fax to 574-372-4265 . This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 2163 | 2,163 in total (expanded) | Worldwide Distribution: US (nationwide) to states of: AL, AR, AZ, CA, CO, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and Puerto Rico; and foreign countries of: Chile, Costa Rica, Ecuador and Netherlands. | Open, Classified |  |  | 2163
Z-0455-2019 | 81456 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | NeodentÂż GM Mini Conical Abutment - Product Usage: Mini Conical Abutments are intermediary prosthetic components to be installed onto GM implants to support the final prosthesis. They are provided in a rotational shape for the coupling with the prosthesis and in different gingival heights to match the variations in mucosal thickness. The GM Exact Mini Conical abutment is provided angled. They are indicated for screw-retained multiple-unit prostheses onto implants |  |  | October 02, 2018 | 2018 | Straumann Manufacturing, Inc. | Laser engraved label does not match with the item in the package | Employee error | Straumann sales reps. visited each account, retrieved and returned product to Straumann USA. |  | 21 | 21 units | US nationwide distribution in the states of GA, TX , VA. | Terminated | January 28, 2019 | 118 | 21
Z-0370-2019 | 81339 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Millennium MG Multi-Geometry Digital CSE Dual Detector Gamma Camera H3000ZL Acquisition System |  |  | September 14, 2018 | 2018 | GE Healthcare, LLC | A detector can detach and fall. | Other | GE Healthcare will inspect and if required correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the inspection. |  | 1166 | 1,166 | United States: AK,AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,Guam,MA,MD,ME,MI,MO,MS,MT,NC,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,PR, RI,SC,SD,TN,TX,UT,VA,VT,WA,WI AND WVInternational:Albania, Argentina, Australia, Azerbaijan, Belgium, Bosnia ,Herzegovina, Brazil, Canada, Chile, China, Costa Rica, Colombia, Cyprus, France, Germany, Greece, Guatemala, Honduras, Iceland, India, Iraq, Italy, Jamaica, Japan, Kenya, Korea, Macedonia, Mexico, Namibia, Panama, Peru, Philippines, Portugal, Russia, South Africa, Spain Switzerland, Taiwan, United Kingdom, Turkey, Venezuela and Vietnam | Open, Classified |  |  | 1460
Z-0461-2019 | 81417 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Parietex Composite Parastomal Mesh (Product Number PCOPM15 / PCOPM20) |  |  | October 05, 2018 | 2018 | Covidien LLC | The firm is issuing this voluntary removal following receipt of reports of parastomal mesh failure identified several years following parastomal hernia repair using the modified Sugarbaker repair technique. In these reports, product failure led to hernia recurrence requiring additional surgical treatment. Symptoms of hernia recurrence may include discomfort, localized pain-free or painful bulging, and possible changes in the overlying skin. Medtronic has received, worldwide, a total of ten reports of mesh failure following use of the product in the last five years. Patients who have received a Parietex" composite parastomal mesh for the treatment of a parastomal hernia should continue to receive ongoing monitoring by their healthcare providers for the recurrence of a parastomal hernia. | Unknown/Undetermined by firm | On October 5, 2018, the firm notified its customers of the product issue via Urgent Medical Device Voluntary Removal.Required Actions:1. Please quarantine and discontinue use of the affected item codes.2. Complete the Product Return Form even if you do not have inventory. If no inventory, fax to 203-492-7719 or email FCAMITG@covidien.com.3. If purchased directly from Medtronic, fax the return form to 800-895-6140 or email feedback.customerservice@covidien.com. Ship affected product with RGA# provided by customer service to:MedtronicAttn: Field Returns Department195 McDermott RoadNorth Haven, CT 06473 USA4. If purchased from a distributor, contact distributor directly to arrange for return of product. Fax completed form to 203-492-7719 or emailFCAMITG@covidien.com.If you have any questions regarding this communication, please contact your Medtronic representative or Customer Service at 800-962-9888, option 2. |  | 7333 | 7333 | Worldwide distribution--US Nationwide | Open, Classified |  |  | 7333
Z-0469-2019 | 81187 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA-900 Analyzer, Product code 022930, 022930R |  |  | October 05, 2018 | 2018 | Tosoh Bioscience Inc | Tosoh Bioscience has become aware of a potential issue with the pickup arm and pickup arm motor on the AIA-900 Analyzer. If the pickup arm or pickup arm motor malfunctions, the analyzer will cease to function and result in an inability to test patient samples. This may cause a delay in testing and reporting patient results. Tosoh has received fifty-one (51) complaints related to this issue with no serious injuries reported. | Device Design | On October 5, 2018, the firm notified its affected customers via an Urgent Medical Device Recall letter. The letter informed customers of the potential product issue with all active AIA-900 analyzers.The firm's service procedures have been updated to proactively troubleshoot and correct the pickup arm andpickup arm motor during installations and scheduled preventive maintenance visits. At the next scheduledpreventive maintenance visit a Tosoh field service engineer will verify the pickup arm motor current andadjust as needed. Adherence to preventive maintenance schedules provides effective mitigation againstpotential failure of these components.Immediate Actions to be taken by the Customer/User" Thoroughly review the content of this letter." Continue to use the instrument. A Tosoh representative will contact you to schedule your nextpreventive maintenance visit.o If a failure of the pickup arm is observed or the analyzer stops functioning, contact TosohTechnical Support for an expedited service visit." Complete and return the attached Acknowledgement Form by any of the listed methods." File this notice with your laboratory records and forward this information to others in your laboratory.If you have obtained unexpected test results, delays in test result reporting, or received any complaints ofillness or adverse events associated with the use of the analyzer, please contact Tosoh TechnicalSupport 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to contact BernadetteO'Connell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will be available toanswer any questions Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 393 | 393 | Distributed nationwide. Foreign distribution to Canada and Latin America. | Open, Classified |  |  | 393
Z-0467-2019 | 81321 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Stryker Sustainability Solutions LASSO NAV eco Reprocessed Diagnostic EP Catheters designed to obtain electrophysiological mapping of the cardiac structures of the heart. |  |  | August 23, 2018 | 2018 | Stryker Sustainability Solutions | Electrophysiology catheters were identified to be item # D134903 and distributed. It was subsequently determined that these catheters are item # D134909, which is not approved for reprocessing. | Under Investigation by firm | On 10/30/2018 an Urgent Medical Device Recall notice was hand-delivered by a sales representative to the customer. The customer was notified that the affected devices had already been returned via the normal collections process. If affected devices were forwarded to additional facilities, the customer was asked to communicate this recall to them and to provide them with the recall effectiveness check form. Customers were asked to complete and return the recall effectiveness check form. |  | 2 | 2 | US: KSOUS: None | Open, Classified |  |  | 2
Z-0465-2019 | 81237 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Lipase Flex Reagent Cartridge, Model Number 10460277 |  |  | September 07, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | There is lot to lot variability for QC and patient results when using the identified reagent cartridge lots. | Under Investigation by firm | Urgent Medical Device Correction letters dated 9/7/18 were distributed to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product. |  | 18396 | 18396 | Worldwide - US Nationwide DistributionsThe products were distributed to the following foreign countries: Argentina, Australia, Bangladesh, Brazil, Canada, Chile, China, Colombia, Germany, Ecuador, Honduras, Indonesia, India, Japan, Republic of Korea, Mexico, Malaysia, Nicaragua, New Zealand, Peru, Philippines, Paraguay, El Salvador, Thailand, Taiwan, Uruguay and South Africa, Austria, Bahamas, Belgium, Czech Republic, Denmark, , France, , Italy, Netherland, Norway , Portugal, Kuwait, Slovenia, Spain, Switzerland, United Kingdom, Qatar, Republic of Korea, Saudi Arabia, Slovakia, Algeria, Bahrain, Bulgaria, Egypt, , Hungary, Latvia, Lebanon, Lithuania, Macedonia, Nepal, Poland, Romania, Russia Fed, Serbia, Turkey, U.A.E, United Kingdom. | Open, Classified |  |  | 22352
Z-0468-2019 | 81242 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow AGB+ Multi-Lumen CVC Kit, 7 Fr., 3 Lumen, 20cm catheter length, .025 inch dia. spring-wire guide, Product Code CDC-45703-B1A - Product Usage: The multiple-lumen catheter permits venous access to the central circulation by the way of the femoral, jugular, or subclavian veins. The ARROWg+ard Blue Plus antimicrobial catheter is intended to help provide protection against catheter-related infections. It is not intended to be used as a treatment for existing infections nor is it indicated for long-term use. |  |  | September 21, 2018 | 2018 | Arrow International Inc | The products labeling contains inconsistencies. In certain instances, the lidstock, banner card and shipper label state that the Spring Wire Guide (SWG) within is .032 rather than the correct SWG size .025 . For other lots in scope the lidstock correctly states the SWG size .025 however the banner card states the incorrect SWG size .032 . | Employee error | Urgent Medical Device Notification letters dated 9/21/18 were distributed to customers. |  | 3640 | 3640 | US Nationwide distribution in the states of CA, FL, ID, IL, MI, NC, NJ, PA, and WI. | Open, Classified |  |  | 3640
Z-0464-2019 | 81408 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Abbott TECNIS 1-Piece Aspheric Acrylic IOL with the TECNIS iTEC Preloaded Delivery System, Lens Model PCB00, Sterile, Rx - Product Usage: The lenses are indicated for the visual correction of aphakia in adult patients in whom a cataractous lens has been removed by extracapsular cataract extraction. These devices are intended to be placed in the capsular bag. |  |  | June 21, 2018 | 2018 | Johnson &amp; Johnson Surgical Vision Inc | Expired intraocular lenses were distributed. | Process control | The recalling firm issued letters dated 6/21/2018 via Federal Express requesting return of the IOL's. |  | 17 | 17 IOL's | US Nationwide distribution in the states of IN, MN, NC, NV, TX, and WA, WV. There was no foreign/military/government distribution. | Completed |  |  | 17
Z-0475-2019 | 81194 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA-2000 Analyzer, Model Nos. AIA-2000LA, AIA-2000ST - Product Usage: The AIA-2000 is capable of performing three methods of immunoassay: an immunoenzymetric (IEMA) or sandwich immunoassay, a competitive binding (EIA) immunoassay, and a two-step immunoenzymetric immunoassay. An antigen-antibody reaction begins by combining a patient sample, control, or calibrator with a diluent in an immunoreaction test cup from the AIA-PACK reagent series. |  |  | September 28, 2018 | 2018 | Tosoh Bioscience Inc | There is a possibility that the liquid suction pump on some AIA-2000 analyzers may not perform as intended and cause an insufficient volume of wash solution to be removed during the washing step. Incomplete washing of the sample may result in falsely elevated test results for the sandwich assays and falsely decreased test results for the competitive binding assays. Insufficient washing is detectable through review of the calibration data and QC trending. For example, if incomplete washing occurs, the reported values for the zero- calibration rate will be above the acceptance range. The calibration value may shift and the instrument will not calibrate successfully. If this issue occurs, there may be a delay in testing and reporting patient test results. There have been no user or patient injuries reported related to this issue. | Component design/selection | On September 28, 2018, the firm distributed Urgent Medical Device Recall letters to affected customers. Customers were informed of the possibility that the liquid suction pump on some AIA-2000 analyzers may not perform as intended and cause an insufficient volume of wash solution to be removed during the washing step. Tosoh has taken the required actions to inspect the liquid suction pump during preventative maintenance andcorrect the instruments where this issue has been observed.Immediate Actions to be taken by the Customer/User" Thoroughly review the content of this letter. If additional information is needed, please contact Tosoh TechnicalSupport." Continue using your instrument.o If there is an unexplained rise in the calibration rate and the AIA-2000 Analyzer serial number is listedabove, please contact Tosoh Technical Support.o If the assay calibration is as expected, then at the next scheduled Preventive Maintenance (PM) visit,a Tosoh field service engineer will inspect the liquid suction pump and replace it as needed." File this notice with your laboratory records and forward this information to others in your laboratory." Complete and return the attached Acknowledgement Form by any of the listed methods.If you have obtained unexpected test results, delays in test result reporting, or received any complaints of illnessor adverse events associated with the use of the analyzer, please contact Tosoh Technical Support 24 hours aday, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to contactBernadette O'Connell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will be available toanswer any questions Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 7 | 7 | US Nationwide distribution in the states of Florida, Kentucky, Maryland, Ohio, Pennsylvania, and Texas. | Open, Classified |  |  | 7
Z-0474-2019 | 81471 | Yes | LWS |  |  | Cardiovascular | PMA | The SQ-RX Pulse Generator a component of the Subcutaneous Implantable Cardioverter Defibrillator (S-ICD) System |  |  | November 01, 2018 | 2018 | Boston Scientific Corporation | Experienced accelerated battery depletion and a shortened replacement interval due to latent internal battery malfunctions. | Device Design | The firm, Boston Scientific, sent an "Important Medical DeviceInformation" letter to customers on November 5, 2018 by hand, email via our sales representatives or Federal Express or Courier. The letter described the product, problem and actions to be taken. We will reconcile that 100% of the identified customers /customer representatives signed for the letter, or the letter has been confirmed to be delivered. Boston Scientific will post the product advisory letter on our website, within the Product Performance Resource Center located @ www.bostonscientific.com/en-US/pprc/product-advisories.html. The customers were instructed to do the following: Your doctor has been notified to look for this possible change, so please go to all scheduled defibrillator checks and clinic visits. As always, if you or your family hear beeping from your device, please contact your heart clinic as soon as possible. Your doctor can confirm your ability to hear and recognize this sound during an appointment.If you have additional questions regarding this information or would like to report a clinical event, please contact your Boston Scientific representative orTechnical Services, or the Recall Contact at 651-582-3494 or email: Heather.Young@bsci.com. |  | 12900 | 12,900 | US Nationwide Distribution: Alabama, Alaska, Arizona, Arkansas, California, Colorado, Connecticut, Delaware, District of Columbia, Florida, Georgia, Hawaii, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Mexico, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Puerto Rico, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Vermont, Virginia, Washington, West Virginia, Wisconsin, and Wyoming. | Open, Classified |  |  | 12900
Z-0473-2019 | 81391 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre Vascular Pruitt F3Âż-S Polyurethane Outlying Carotid Shunt, Sterile |  |  | October 01, 2018 | 2018 | LeMaitre Vascular, Inc. | Pouches from this lot are not sealed compromising the sterility | Packaging process control | Lemaitre Vascular initiated recall by letter on October 1, 2018, provided with reason for recall action to be taken, risk to health. Complete the enclosed customer reply form, and return it to LeMaitre Vascular, Inc. Return all devices from lot PFP1055, whether they appear to be unsealed or not. If you distributed any affected product to other facilities, please forward this recall to the recipients.Contact Information: rlerer@lemaitre.com 781-425-1671 Monday through Friday, 8:00 AM to 4:30 PM, |  | 46 | 46 units | FLForeign: Austalia, Korea | Open, Classified |  |  | 46
Z-0476-2019 | 81376 | No | MRR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂż Chemistry Products dLDL Reagent-GENs 26 &amp; 27. In Vitro Diagnostic to quantitatively measure LDL Cholesterol (LDLC) |  |  | October 08, 2018 | 2018 | Ortho-Clinical Diagnostics | VITROS Chemistry Products dLDL Reagent GENs 26 and 27, negatively biased results | Under Investigation by firm | Ortho Clinical Diagnostics issue the customer notification on October 8, 2018, provides information regarding the potential for biased results when using VITROS dLDL Reagent, Generations (GENs) 26 and 27. Actions to take:For GEN 26: Discontinue using GEN 26 and switch to an alternate GEN of product as soon as possible. Please specify full or partial sales units remaining and indicate credit or replacement on your Confirmation of Receipt form." For GEN 27: Install ADD DRV 6003 on your VITROS System upon receipt. Following a successful recalibration, resume using VITROS dLDL Reagent, GEN 27 remaining in your inventory." Complete the Confirmation of Receipt form and return by October 22, 2018.Questions contact our Ortho Care" Technical Solutions Center at 1-800-421-3311 |  | 5850 | 5850 units | NationwideForeign: AustraliaBelgiumBermuda, HM11BrazilCanada L3R 4G5 ChileChinaColombiaDenmarkFranceGermanyIndiaItalyJapanMexicoNorwayPolandPortugalRussiaSingapore SpainSwedenThe NetherlandsUnited Kingdom | Open, Classified |  |  | 5850
Z-0478-2019 | 81508 | No | NDC | Predictive pulmonary-function value calculator. | Anesthesiology | General Hospital | 510(k) | Insulia Diabetes Management Companion |  |  | October 08, 2018 | 2018 | Voluntis | The firm identified an issue with the basal calculator identified on the Android version of the Insulia application. The bug could lead from low-impact to high-impact hypoglycemia depending on the circumstances. A new update of Insulia is now available (version 1.6.42), which resolves this issue. | Software in the Use Environment | On October 15, 2018, patients were notified that they could download the new version of the application from the Google Play Store. The notification was made through an automated message displayed on the application screen of all impacted patients requesting them to update their Insulia application. Patients were informed that the new version of the app resolved a an issue with the basal calculator identified on the Android version of the application. Patients who do not upgrade to the new version by November 9, 2018 will be blocked from using the app until they upgrade.For technical assistance or if you have any questions, please contact the Insulia support hotline. |  | 14 | 14 | Distributed to accounts in Arizona, Masachusetts, Michigan, Ohio, Tennessee, and Vermont. International distribution to Canada and France (however, French customers are not affected by this recall). | Open, Classified |  |  | 14
Z-0477-2019 | 81409 | No | CFN | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | The device consists of reaction buffer, sheep anti-IgM antiserum, six calibrators and |  |  | July 25, 2018 | 2018 | The Binding Site Group, Ltd. | After running the calibration curve, the recovery of the IgM kit control results decrease demonstrating a reduction in accuracy within the same day; control results may be out of range one day after running the assay calibration curve. Samples have been shown to be similarly affected. | Under Investigation by firm | On 07/25/2018, the Urgent Field Safety Notice and Field Correction E-back Form were emailed to customers. Customers were advised to cease use of the affected product and to dispose of the kits locally. This notice states that a witness signature of kit disposal is required on the E-back form. This notice needs to be passed on to all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. Customers with additional questions are encouraged to contact their local firm representative, or customer service via phone at +44(0) 1214569696, or email at Technical.support@bindingsite.co.uk. The E-back Form requests that each customer and witness sign and date confirming they have both received and have followed the advice issued in the Notification. It also asks for the number of product in stock and number of kits disposed. |  | 108 | 108 | U.S.: IL, WV, NE, SD, NY, IN, WA, MD, NJ, PA, NC, AR, FL, VA | Open, Classified |  |  | 108
Z-0485-2019 | 81219 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tosoh AIA-900 immunoassay Analyzer |  |  | September 28, 2018 | 2018 | Tosoh Bioscience Inc | A software issue may cause the liquid volume detection entries for 2nd and 3rd reagent bottles to be overwritten with the data from the 1st reagent bottle. Tosoh has released software version 2.51 (or later) to fix the problem. This failure is mitigated by the hardware design. The consequence would be a slight delay in testing patient sample or reporting patient result. | Software design | On September 28, 2018, the firm issued an Urgent Medical Device Recall letter to affected customers. Customers were informed of the software issue. The firm has released software version 2.51 (or later) which corrects the problem. Three analyzers were identified that do not have the most current software version and are at risk for the error: 10313605, 10503612, and 10634103. A field service engineer will be contacting customers with these units in their possession to schedule a service visit in order to upgrade those analyzers to the current version of the operating software.Immediate Actions to be taken by the Customer/User-Thoroughly review the content of this letter.-Complete and return the attached Acknowledgement Form.-File this notice with your laboratory records and forward this information to other supervisors ormanagers in your laboratory to ensure that they are aware of the potential issue.If you have obtained unexpected test results, delays in test result reporting, or received any complaintsof illness or adverse events associated with the use of the analyzer, please contact Tosoh TechnicalSupport 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to contactBernadette O'Connell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will beavailable to answer any questions Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 3 | 3 | Distributed to accounts in GA and NY. | Open, Classified |  |  | 3
Z-0479-2019 | 81224 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Giraffe OmniBed Carestation. Neonatal incubator. Device sold under the following product names: Carestation and Carestation CS1 |  |  | September 25, 2018 | 2018 | GE Healthcare, LLC | Potential electric leakage affecting the user. | Process design | GE Healthcare notified customers on about 09/25/2018 via URGENT MEDICAL DEVICE CORRECTION letter. The letter indicated that the device could continue to be used along with good clinical practices. GE Healthcare will visit each location to perform the necessary testing. Until GE Healthcare can visit each site, instructions are to disconnect devices that use the USB (Giraffe OmniBed Carestation and Giraffe Incubator Carestation only) and the RS232 Nurse call ports, and inspect all power cords to ensure there is no damage to the cord or ground pin. GE Healtchare will perform necessary electrical safety test on all affected products and correct if required. A GE Healthcare representative will contact each customer to arrange for this test. For questions or concerns, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 1724 | 1,724 devices total | Worldwide distribution. US Nationwide, Algeria, Australia, Austria, Belgium, Canada, Chile, China, Costa Rica, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Honduras, India, Indonesia, Ireland, Italy, Japan, Kazakhstan, Korea, Lithuania, Mexico, Netherlands, New Zealand, Norway, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam. | Open, Classified |  |  | 1724
Z-0487-2019 | 81367 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | PremierPro Aluminum Adjustable Walker |  |  | October 01, 2018 | 2018 | SVS LLC | This recall is due to a mandated order from the FDA for stock recovery and return of the imported item. | Other | SVS LLC, dba S2S Global notified customers initially via telephone on about 10/01/2018 and followed up with a "Notice of Product Recall" letter on about 10/03/2018. The letter requested customers check inventory for affected walkers, quarantine any on hand, make arrangements to return the affected walkers, and complete and return the Medical Device Acknowledgement &amp; Return Response Form. |  | 256 | 256 devices | US distribution to OH, NY,TX, PA | Open, Classified |  |  | 256
Z-0490-2019 | 81399 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | Irrisept OR, Finished Bottle Assembly 450 ml STEP 1 |  |  | September 05, 2017 | 2017 | Medex Cardio-Pulmonary Inc., d.b.a. Smiths Medical Company | Lack of sterility assurance: leaking containers which could lead to exposure to infectious agents. | Other | Firm sent recall notification letters to customers on 9/5/2017. |  | 10440 | 10,440 | AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OR, PA, PR, SC, TN, TX, UT, VA, VT, WA, and WI | Open, Classified |  |  | 10440
Z-0488-2019 | 81377 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Sedecal NOVA FA DR System. for diagnostic radiography. |  |  | July 18, 2018 | 2018 | Sedecal S.A. | A safety problem with installation and maintenance could lead to the breakage of the two steel cables that support the equipment to its roof anchor. Due to this, the equipment may fall and cause harm to the patient, user, or third parties. | Device Design | Initial communication from SEDECAL to all our distributors was performed on 07/18/2018 by email. The letter explained the problem and requested the consignee cease use of the unit if they observe an abnormal displacement in the vertical axis or any suspicious noise during the vertical movement. |  | 145 | 145 units | US and Canada | Open, Classified |  |  | 145
Z-0486-2019 | 81546 | No | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ascom Telligence Nurse Call System |  |  | November 02, 2018 | 2018 | Ascom (US) Inc. | There is an issue that could cause the Telligence System Gateway II, NGGTWY2 H, to stop communicating to the hallway devices connected to it. When this occurs, the dome lights connected to the Gateway II Ports A and B will start flashing, indicating that they are no longer communicating to the Gateway II. If this occurs, device failures that WOULD otherwise be annunciated WILL NOT be indicated at staff consoles that are configured to display failures. | Software Manufacturing/Software Deployment | Ascom Partners / Distributors were contacted by e-Mail, USPS Mail and Phone as necessary on 11/02/2018 to determine where the Telligence 5.0 Product is installed. |  | 57 | 57 units | AL, CA, CO, GA, IN, KY, LA, ME, MI, MO, MN, MS, NE, NV, OH, PA, RI, TX UT, VA, WI | Open, Classified |  |  | 57
Z-0472-2019 | 81039 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER EPICS XL Flow Cytometer |  |  | January 12, 2018 | 2018 | Beckman Coulter Inc. | A manufacturing defect on the amplifier board of the Flow Cytometers my cause signal loss and/or signal drifting resulting in absence of data or a population shift in the data plots. This can impact patient results for any application. | Nonconforming Material/Component | On 01/12/2018 the firm sent an Urgent Medical Device Recall Letter to affected customers via email/mail informing them that as a result of customer complaints and subsequent internal investigations, Beckman Coulter has determined that an internal electronic component on the circuit Amplifier boards in the FC 500 system may be affected by a manufacturing defect. Each FC 500 system contains seven (7) of the potentially affected Amplifier boards. All instruments are potentially impacted. Customers were informed to implement the following actions for the application used:1. For All Applications including Laboratory Developed Tests:a. As per the product documentation all data must be reviewed by a laboratoryprofessional prior to the release of reported results from the lab.b. Immediately implement the collection of Time as a parameter and create Time versus Parameter plots which will allow the monitoring of signal integrity during data acquisition as instructed in Attachment 2.c. Review data as described below: i. Review of all Time plots for each parameter. ii. Monitor consistent Forward Scatter, Side Scatter and all fluorescence data over time as shown in Attachment 2. iii. Unexpected fluctuations in the events over time may indicate compromised fluidics, signal integrity or data acquisition conditions. iv. All data must be reviewed prior to the release of any results from the laboratory, via an LIS or any other mechanism.2. For tetraCXP and stemCXP:a. As it is not possible to add time versus parameter, discontinue use of theautomated tetraCXP and stemCXP applications.b. You can continue to use the tetraCHROME reagents and Stem Kit reagentswith the manual gating instructions provided in the product labeling. Refer toCYTO-STAT tetraCHROME IFU, PN B90108 (for tetraCHROME CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5, PN 6607013 and tetraCHROME CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5, PN 6607073) and the Stem-Kit IFU,PN B6 |  | 2549 | 2549 | Worldwide distribution. US Nationwide including Puerto Rico.OUS:FC 500:Japan, South Korea, China, Italy, Spain, United Kingdom, Germany, Bangladesh, Czech Republic, France, Pakistan, India, South Africa, Austria, Saudi Arabia, Algeria, Turkey, Belgium, Netherlands, Bulgaria, Sweden, Australia, Hungary, Sri Lanka, Uruguay, Canada, New Zealand, Switzerland, Gabon, Portugal, Mexico, Singapore, Monaco, Hong Kong, Guadeloupe, Zambia, United Arab Emirates, Ireland, Reunion, Poland, Egypt, Slovakia, Croatia, Brazil, Chile, Bahrain, Slovenia, Bosnia and Herzegovina, Greece, Myanmar, Morocco, Serbia, Philippines, Lebanon, Romania, Moldova, Qatar, Barbados, Iraq, Oman, Namibia, Indonesia, Thailand, The Bahamas, Macau, El Salvador, Panama, Denmark, Venezuela, Azerbaijan, Israel, Colombia, Peru, Malaysia, Vietnam, Tunisia, Kuwait, Kenya, Libya, Latvia, Russia, Belarus, Ukraine, Kazakhstan, TaiwanEPICS XL:Albania, Algeria, Angola, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Botswana, Bulgaria, Burundi, Cameroon, Canada, China, Congo, Croatia, Cyprus, Czech Republic, Denmark, Dominican, Republic, Egypt, Ethiopia, France, Germany, Greece, Guinea, Guinea-Bissau, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Lebanon, Lesotho, Libya, Macau, Malawi, Mexico, Morocco, Mozambique, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Swaziland, Sweden, Switzerland, Syria, Taiwan, Tanzania, The Bahamas, Tunisia, Turkey, Turkmenistan, Uganda, Ukraine, United Arab Emirates, United Kingdom, Zambia, Zimbabwe | Open, Classified |  |  | 5664
Z-0388-2019 | 81454 | No | REM |  |  |  |  | DeltaVision OMX SR Imaging System Model Number: 29115476 |  |  | September 10, 2018 | 2018 | GE Healthcare Biosciences | Non-conformance with the DeltaVision OMX SR, the interlocks are not wired in a redundant fashion to make the interlock system single-fault tolerant. | Device Design | The firm, GE Healthcare, sent an "Important Electronic Product Radiation Warning" letter to affected customers on September 10, 2018. The letter identified the affected product, problem and actions to be taken. GE Healthcare will, without charge, remedy the defect or bring the product into compliance. Your GE Service Representative will contact you to implement the correction. For questions contact your local GEHC service representative or call 425-657-1512 or email: trisha.koenke@ge.com. |  | 39 | 39 | Worldwide - US Nationwide Distribution | Open, Classified |  |  | 39
Z-0489-2019 | 81122 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Minicon Clinical Sample Concentrators for In Vitro Diagnostic Use 9031 (Model: Minicon B-15) &amp; 9051 (Model: Minicon CS-15) |  |  | May 14, 2018 | 2018 | Merck Millipore Ltd. | In vitro diagnostic devices, used greater than 6 months after date of manufacture, may have increased concentration times of greater than the specification of 3 hours and up to 24 hours, which may lead to delays in carrying out the subsequent analysis tests. | Process control | On May 14, 2018, the firm sent an "Urgent Field Safety Notice" to customers via email and following up with a hard-copy letter on June 6, 2018 via United States Postal Services. The letter described the product, problem and actions to be taken. The corrective action being taken by the firm is to reduce the shelf life of the device from 12 months to 6 months.The customers are being instructed to immediately examine their inventory to determine if they have any remaining impacted product and forward a copy of this letter to the personnel in your organization that utilize this product and/or need to be made aware of this issue.In addition, if the customer has further distributed this product, please identify its customers and notify them at once of this issue. The notification to their customers should include a copy of this notification and attachment 2Customers are instructed to please stop use of these lots if older than 6 months and destroy any remaining inventory older than 6 months, and to complete and return the receipt confirmation form to QACorkGroup@merckgroup.com.If you would like more information regarding this notification, please contact us at QACorkGroup@merckgroup.com or go to our website to find the phone number of your local office at www.emdmillipore.com/offices. |  | 1360 | 1360 units | Worldwide Distribution: US (nationwide) to states of: AZ, CA, CT, DC, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NJ, NY, OH, OK,OR, PA, RI, TN, TX, VA, VT, WA and WI; and countries of: Canada and Australia. | Open, Classified |  |  | 1360
Z-0491-2019 | 81148 | No | JTW | Microbiological specimen collection and transport device. | Microbiology | Microbiology | 510(k) | BD Vacutainer Urine Collection Cups, Cat. No. 364941 |  |  | August 30, 2018 | 2018 | Becton Dickinson &amp; Company | BD Vacutainer Urine Collection Cups may contribute to a slight elevation of Red Blood Cell (RBC) counts when used with the Sysmex UF Series Analyzers. BD recommends following Good Clinical Practice (GCP) by using microscopy to confirm an elevated RBC Count in a Urine sample as per CLSI GP16-3A. 29(4) and CAP Checklist on Urinalysis. Product will continue to ship to not impact critical patient care. | Under Investigation by firm | On August 30, 2018, BD notified Sysmex via an emailed Urgent Medical Device Product Safety Alert letter. Sysmex was informed that BD had confirmed the product issue with BD Vacutainer Urine Collection Cups were used with Sysmex UF Series Analyzers.Sysmex was instructed to take the following actions:- Notify Sysmex customers of the potential for elevation of RBC counts- Complete the customer Resopnse form and email it to BD at Gail_Griffiths@bd.com to confirm receipt of the Safety Alert.Product will continue to ship so as to not impact critical patient care.BD recommends following Good Clinical Practice (GCP) by performing microscopy to confirm an elevated RBC Count in a urine sample as per CLSI GP16-3A.29(4) and CAP Checklist on Urinalysis. |  | 29982150 | Total: 29,982,150 BD cups (US); 97,800,900 BD cups (OUS) | Illinois | Open, Classified |  |  | 29982150
Z-0496-2019 | 81472 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Knee System -PS Open Box Femoral Right, 65 mm |  |  | October 22, 2018 | 2018 | Zimmer Biomet, Inc. | Label on the outer carton or the patient labels does not match label on the inner sterile packaging | Labeling mix-ups | Zimmer issued recall letter on 10/22/18 via email and FedEx. Acountts requested to immediately locate and quarantine affected product in your inventory. Return affected product from your distributorship and from affected hospitals. Complete Inventory Return Certification Form. uestions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. |  | 8 | 8 units | AR, MO, MN, NJ, NY, OH, MN , NJ, VA | Open, Classified |  |  | 8
Z-0470-2019 | 81317 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | BrightMatter Guide, Product Number SYN-0026, Versions: 1.5.0, 1.5.1 and 1.5.2 |  |  | October 09, 2018 | 2018 | Synaptive Medical Inc | This recall has been initiated due to a software defect found in the Guide System software when used with NICO BrainPath ports under certain circumstances. When one trajectory is set with a blue port and another trajectory is set with a gold port, the software defect is triggered when the user switches between these trajectories during a surgical procedure. | Software design | Synaptive Medical sent an Urgent Medical Device Recall letter dated October 9, 2018 via registered priority mail . The letter identified the affected product, problem and actions to be taken until a software patch is released. Customers were requested to complete and return the enclosed Acknowledgement and Receipt form to Synaptive Medical Inc. For questions call 1-844-462-7246. |  | 37 | 37 units | Worldwide - US Nationwide Distribution and the country of Canada | Open, Classified |  |  | 37
Z-0503-2019 | 81445 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow B Solution-RFP-403, Premixed Dialysate for Hemodialysis |  |  | October 17, 2018 | 2018 | NxStage Medical, Inc. | PureFlow B Solution smaller chamber of the two chamberbag can burst electrolyte fluid and cause injury if it comes in contact with the eyes or skin | Packaging process control | NxStage initiated a recall on October 17, 2018 via " URGENT MEDICAL DEVICE REMOVAL" Recall Notice to its customers. The notice provided reason for recall, health risk, action to be taken. The customers were instructed to do the following: 1 Check all boxes in your Pureflow B solution inventory, separate them from your other inventory and do not use them.2. Follow the instructions on the attached reply form or click on the following link to respond electronically, https://www.nxstage.com/rfprecalloct2018, to acknowledge your receipt of this notice.3. Contact NxStage Customer Service to arrange for return of all affected product and for replacement product to be sent,If you have any questions or comments, please feel free to contact NxStage Customer Service 1-866- NXST AGE (1-866-697-8243). |  | 2189 | 2189 | US Nationwide Distribution. | Open, Classified |  |  | 423353
Z-0508-2019 | 81480 | No | HIF | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | PNEUMOCLEAR Heated High-Flow Tube Set |  |  | October 22, 2018 | 2018 | WOM World of Medicine AG | It is possible that the packaging of the product can be damaged by the prongs on the tube set. | Packaging process control | WOM sent a Recall Notification letter dated October 22, 2018 to the their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to send all in the US recalled products back to:W.O.M. WORLD OF MEDICINE USA, Inc.4531 36th StreetOrlando, FL 32811-6527WOM contacts for coordinating the shipment of the requested products to Orlando:Customer Service -W.O.M. OrlandoPhone: +1 407 472 0372Customer Service -W.O.M. OrlandoPhone: +1 407 472 0371Please send all recalled products returned to Stryker Venlo/Netherlands back to:W.O.M. WORLD OF MEDICINE GmbHAlte Poststr. 1196337 Ludwigsstadt - GermanyCustomer Service: +49 30 39981 524 |  | 124180 | 124,180 tube sets | Worldwide Distribution - US in the state of California and country of Netherlands | Open, Classified |  |  | 124180
Z-0498-2019 | 81430 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Reprocessed Xcel Bladeless Trocar (H12LP) Model # H12LP Reprocessed Endoscopic Trocars12x100mm - Product Usage: Physical Product Description: The Reprocessed Blunt Tip Trocar consists of a radiolucent cannula and obturator sized 12mm in diameter (refer to Figure 1). The Reprocessed Blunt Tip Trocar has an obturator that has a blunt plastic tip that gently moves aside any internal viscera that may be adjacent to the abdominal or thoracic wall. The trocar cannula consists of two seals, an outer integrated removable self-adjusting seal that may accommodate instruments ranging from 5mm to 12mm in diameter, and an internal seal. These seals minimize gas leakage when instruments are inserted or withdrawn through the seal. The stopcock valve provides attachment for gas insufflation and desufflation. The stability plug component, is specific only to the H12LP catalog number. The stability plug (adjustable plug) is placed on the outside of the trocar sleeve (cannula). The stability plug provides suture clips (suture anchors), which can be used to suture the Trocar sleeve in place. Intended Use: Reprocessed Endoscopic Trocars are indicated for use to establish a port of entry for endoscopic instruments in patients requiring minimally invasive surgical procedures. |  |  | June 19, 2018 | 2018 | Stryker Sustainability Solutions | Through investigation, it was determined that H12LP Trocars in scope of a previous H12LP recall were distributed to facilities within the distribution chain following the recall execution activities initially performed by distributor. | Other | On June 29, 2018, the firm emailed an Urgent Medical Device Recall Notification to customers informing them that the firm is voluntarily recalling specific lots of the reprocessed Xcel blunt tip trocar, item number H12LP that were previously recalled under Product Field Action Z-0765-2017 initiated on11/2/16. It has been identified through complaints that recalled product was actively distributed within the O&M; network after initiation of Product Field Action Z-0765-2017. PLEASE DISTRIBUTE THIS INFORMATION TO ALL PERSONNEL RESPONSIBLE FOR REPROCESSED H12LP XCEL BLUNT TIP TROCAR.Item No: H12LPPlease note: Only devices distributed from SSS to O&M; from the affected lots are in scope. The affected lots were distributed from SSS to O&M; from June 29, 2015 August 15, 2016.Affected Lot Numbers are:3727435 3811778 3890089 4245630 4530457 4597997 4616746 46722933732833 3821740 3894678 4274018 4533116 4598476 4618135 46745123732851 3829338 3896022 4282444 4548738 4602957 4659478 46748573752003 3833206 4074165 4386368 4551779 4603588 4659568 46808533762495 3834196 4113667 4458288 4557026 4607573 4660191 46812453762972 3849330 4118604 4459232 4566912 4608496 4662392 46832683765649 3853141 4133825 4466103 4571871 4608514 4663276 46869423775681 3854774 4138769 4466644 4572487 4609747 4664364 46880313779516 3861804 4147801 4467337 4574911 4610175 4667846 46883263782202 3863303 4154685 4473956 4576086 4610894 4668260 46888873783776 3866398 4158429 4474932 4576443 4611167 4668273 46888943784430 3870240 4170275 4476349 4577407 4611179 4668287 46907863786146 3873578 4173170 4477829 4580431 4612670 4669208 46957073794419 3880942 4187693 4479578 4581177 4613359 4670139 46987463803918 3881364 4198853 4512285 4583936 4614603 4670712 47044373809891 3886277 4202332 4513266 4589373 4615246 46710553810037 3890070 4235633 4515596 4593587 4615536 4672065The O&M; distribution centers known to have received shipments of product in scope of this |  | 4848 | 4848 | US nationwide distribution in the states of OK, GA, NJ, TX, IA, CO, TX, CA, LA, PA, VA, MD, WA, and IL.OUS: None. | Open, Classified |  |  | 4848
Z-0499-2019 | 81433 | No | CFJ | Lactate dehydrogenase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista Lactate Dehydrogenase (LDI) Flex reagent cartridge, an in vitro diagnostic test for the quantitative measurement of lactate dehydrogenase in human serum and plasma on the Dimension VistaÂż System.Cat. No. K2054, Mat. No. 10464323 |  |  | October 11, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed that Dimension Vista Lactate Dehydrogenase(LDI) may exhibit an erroneously elevated result when the preceding assay in the reactioncuvette is Lactic Acid (LA). The magnitude of the elevation may be variable. Both quality controland patient samples may be impacted. Incomplete removal of residual lactate dehydrogenase, acomponent of the LA reagent, has been identified as the cause of the potentially elevated result. A falsely elevated result for lactate dehydrogenase may lead to additional investigations todetermine the etiology of tissue damage. Results would be correlated with clinical history andpresentation in addition to other diagnostic laboratory testing. | Process control | An Urgent Medical Device Correction were issued on October 11, 2018 to all Dimension Vista customers to notify them of the issue, the potential risk to health, and actions to be taken by the customer. Customers were instructed to: -Configure Dimension Vista System to eliminate the potential for cuvette carryover from lactic acid.-If unable to separate the assays due to lack of additional server based on the the Dimension Vista model, run Lactate Dehydrogenase method in duplicates and contact Siemens if either result is discordant.-Complete and return the Field Correction Effectiveness Check Form attached to the letter within 30 days.-If you have received any complaints of illness or adverse events associated with the product listed in Table 1 of the letter, immediately contact your local Siemens Customer Care Center or their local Siemens Technical Support Representative |  | 13435 | 13435 | Worldwide-domestic nationwide and foreign distribution. | Open, Classified |  |  | 13435
Z-0514-2019 | 81475 | No | GGM | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | HemoTrol Level 2- In Vitro Diagnostic Hematology quality control mixture |  |  | October 18, 2018 | 2018 | EUROTROL INC | Incorrect measurement results caused by microbial contamination of the product, results in values lower than assigned values | Material/Component Contamination | Eurotrol distributors and initial consignees contacted via e-mail with the recall notification letter on 10/16/18. The recall notification letter sent via overnight mail in case e-mail is not deliverable.The recall notification letter instructs distributors how to perform the subrecall towards end-users and inform Eurotrol about the performed subrecall actions Primary instruction will be to stop using the product (discard according local regulation on medical waste) and contact Eurotrol for replacement material.Questions, contact (Monday through Friday, 9:00 AM to 5:00 PM, Eastern Time):Name: Eurotrol Inc. Address 850 N Black Branch Rd. Zipcode &amp; city: Elizabethtown, KY 42701 USA Phone: T +1 502 501 1180 E-mail: officeusa@eurotrol.com. |  | 2000 | 2000 | Worldwide Distribution: US (Nationwide) and countries of: Austria, Australia, Bahrain, France, Hong Kong, Indonesia, Ireland, Japan, Jordan, Kenya, Republic of Korea, Kuwait, Maldives, New Zealand, Mozambique, Pakistan, Poland, Slovenia, Africa, Spain, Thailand, Finland, Norway, and UK. | Open, Classified |  |  | 4122
Z-0513-2019 | 81266 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | PerkinElmer QSight 210 MD Mass Spectrometer Simplicity 3Q MD v1.0 with Hot Fix 2 Software intended to identify compounds in human samples by ionizing the compound and separating the resulting ions. Part Number: BC004872 |  |  | September 24, 2018 | 2018 | PerkinElmer Health Sciences Canada, Inc. | Under certain conditions when using PerkinElmer QSight 210 MD mass spectrometer with Simplicity 3Q MD and Hotfix 2 software, it is possible for the original data from the last acquired sample to be over-written when the automatic flush is enabled. | Software design | Perkin Elmer issued notification on 9/24/18 customers are requested to follow instructions contained within the Urgent Medical Device Correction letter and disseminate the information to appropriate functions within their organization.Inform relevant personnel and execute one of the following two options:Qption 1 (Users who do not require Autamatic Flushinq):Deactivate the automatic flushing option in your software:1) Navigate to the Batch Acquisition Queue tab2) Click the Queue Options button3) Under the "System Flush and Standby" section, DESELECT the "Enable automatic flushing" option.Option 2 (Users who require Automatic Flushinqt:Add a non data (blank or "dummy") sample to the acquisition queue to be run as your final sample(before going to Flush/standby). This will ensure that a true sample or control data are not overwritten.Contact PerkinElmer Heatth Scbnces C,arda, Bolton facility Quality Manager- SandiOrdog for further information (647. 505. 8682). The issue identified above will be addressed through a software update. PerkinElmer will contactyou when the fix is available and upgrade your software. |  | 18 | 18 units | FL, MN, PAIndia, Italy, Korea, Spain | Open, Classified |  |  | 18
Z-0509-2019 | 81212 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA-900 Automated Immunoassay Analyzer, product code 022930 /022930R |  |  | October 05, 2018 | 2018 | Tosoh Bioscience Inc | Tosoh Bioscience has become aware of a potential issue with the axis base of the turntable rotation drivemotor on the AIA-900 Analyzer. If the tension of the timing belt between the turntable rotation pulley anddrive motor is not adjusted correctly, it is possible to increase the lateral load on the motor axis resultingin the inclination and damage of the motor axis. Since January 12, 2017, Tosoh has been inspecting thetension of the timing belt during scheduled preventative maintenance visit and correcting the analyzerswhere this issue has been observed. If the drive motor fails, the AIA-900 analyzer will cease to function and will be unavailable for use. If thisissue occurs, there may be a delay in testing and reporting patient test results. Tosoh has received five (5)complaints related to this issue with no serious injuries reported. | Nonconforming Material/Component | On October 5, 2018, the firm issued an Urgent Medical Device Recall - Correction letter to affected customers. The letter described the product issue and advised customers that since January 12, 2017, Tosoh has been inspecting the tension of the timing belt during scheduled preventative maintenance visit and correcting the analyzerswhere this issue has been observed. Serial numbers of affected AIA-900 analyzers which have not yetbeen inspected were provided as a list.Immediate Actions to be taken by the Customer/User" Thoroughly review the content of the letter." File this notice with your laboratory records and forward this information to other supervisors ormanagers in your laboratory to ensure that they are aware of the potential issue." Complete and return the attached Acknowledgement Form by any of the listed methods." Identify if the serial number of your AIA-900 analyzer is listed in the letter, contact TechnicalSupport to schedule a preventative maintenance service visit. At the scheduled visit, allow the Tosohfield service engineer to inspect and verify that the axis base of the turntable rotation drive motor isoperating as expected or allow the Tosoh field service engineer to correct the potential issue." If the serial number of your analyzer is not listed, Tosoh service records indicate that theturntable rotation drive motor on your instrument(s) has already been inspected.If you have obtained unexpected test results, delays in test result reporting, or received any complaintsof illness or adverse events associated with the use of the analyzer, please contact Tosoh TechnicalSupport 24 hours a day, seven days a week at (800) 248-6764 for an expedited service visit.Should you have any questions regarding this medical device recall, please feel free to contactBernadette O'Connell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will beavailable to answer any questions Monday - Friday, from 9:00 AM |  | 247 | 247 | Worldwide Distribution - US Nationwide in the states of FL, MD, NV, NY, SC, and TX. and Foreign distribution to Cayman Islands, Chile, Columbia, Dominican Republic, Ecuador, Guatemala, Honduras, Peru, Puerto Rico, Uruguay, and Venezuela. | Open, Classified |  |  | 247
Z-0510-2019 | 81460 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Multi-Track Angiographic Catheter (MMTA) |  |  | October 17, 2018 | 2018 | Numed Inc | Instructions for Use booklets were not included on the outerpouch | Labeling design | NuMED issued recall notification letter and response form on 10/17/2018 via email to the one distributor and followed up by a hard copy of the letter and response form sent via UPS Next Day Air. The letter identified the affected product, problem and actions to be taken. The customer was instructed to return affected product.. |  | 100 | 100 units | US distribution in the state of PA | Terminated | January 28, 2019 | 103 | 100
Z-0436-2019 | 81382 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | QUADROX-i Neonatal without integrated arterial filter and VHK 11000 (VKMO 10000-USA#QUADROX-i Neo.+Reservoir), Part Number 701055120 |  |  | October 25, 2018 | 2018 | Maquet Cardiovascular Us Sales, Llc | There is a potential for blood leakage at the blood inlet connector when using the reducing screw adapter 1/4 x 3/16 that may result in clinically relevant loss of prime before patient connection or blood loss during patient connection. | Nonconforming Material/Component | The firm, GETTINGE, sent an "URGENT MEDICAL DEVICE RECALL" notification letters dated 10/25/18 to customers. The letters described the product, problem, and actions to be taken. The customers were instructed to do the following:-Please immediately examine your inventory to determine if you have any of the QUAD ROX-i Neonatal Oxygenator with the product codes/lot numbers listed in this notice.-If you wish to use the product, please remove and do not use the enclosed reducing screw adapter 1/4" x 3/16" (Ref. No. 70104.8593).-If you do not wish to use the product as packaged, please contact Maquet/Ge tinge Customer Support at +l (888) 627-8383 (press option 2, then option 2), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. (Eastern Time Zone) to obtain instructions on returning the affected productfor a full credit. Please note that this product is no longer being offered in the United States, and therefore product replacement is not available.-Please complete and sign the attached MEDICAL DEVICE RECALL- RESPONSE FORM (page 4) to acknowledge that you have received this notification. Return the completed form to Maquet/Ge tinge by e-mailing a scanned copy to neoguadroxi2018@getinge.com or by f axing the form to +l (800) 859-8371.If you are a distributor who has shipped any affected products to customers, please forward this document to their attention for appropriate action.If you have any questions, please contact your Maquet Cardiopulmonary /Ge tinge representative or call the Maquet Cardiopulmonary /Ge tinge Customer Support at +l (888) 627-8383 (press option 2, then option 2), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. (Eastern Time Zone). |  | 490 | 490 total in US | Worldwide Distribution: US (nationwide) to states of: AZ, CA, FL, IA, LA, NC, and TX and countries of: Canada, Argentina, Brazil, Chile, Colombia, Mexico, Uruguay, United Arab Emirates, Belgium, Switzerland, Germany, Egypt, Spain, France, United Kingdom, Ireland, Italy, Lebanon, Netherlands, Pakistan, Portugal, Reunion, Saudi Arabia, Sweden, South Africa, Poland, Serbia, Russia, Turkey, Ukraine, China, India, Japan, Republic of Korea (South Korea), Malaysia, Thailand, and Taiwan. | Open, Classified |  |  | 490
Z-0519-2019 | 81058 | No | MHZ |  |  | Neurology | Contact ODE | Dash 3000 Patient Monitoring System is used to monitor physiologic parameter data on adult, pediatric and neonatal patients. |  |  | August 31, 2018 | 2018 | GE Healthcare, LLC | The patient monitors may simultaneously restart as designed if all are connected to the same network and a network overload occurs for a prolonged time. | Software design | The recalling firm issued letters dated 8/31/2018 via FedEx on 8/31/2018 informing the customer of the issue and providing safety instructions for mitigation of the issue if a prolonged network overload occurs and the Dash and/or Solar patient monitor on the network does not resume normal functionality in approximately 90 seconds. |  | 547920000 | Approximately 182,640 monitoring systems for Dash 3000/4000/5000 | Distribution was made nationwide, including Puerto Rico, Guam, and Samoa. There was also government and military distribution.Foreign distribution was made to Canada, Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bahamas, Bahrain, Bangladesh, Belarus, Belgium, Bermuda, Bolivia, Bosnia and Herzegovina, Brazil, Brunei Darussalam, Bulgaria, Cambodia, Cameroon, Chile, China, Colombia, Costa Rica, Cota D'Ivoire, Croatia, Cyprus, Czech Republic, Denmark, Djibouti, Dominican Republic, Ecuadore, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, French Guiana, Gabon, Georgia, Germany, Ghana, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea (Republic of), Kuwait, Kyrgyzstan, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mauritius, Mexico, Moldova, Morocco, Namibia, Nepal, Netherlands, New Zealand, Nicaragua, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Republic of Serbia, Reunion, Romania, Russia, Saint Lucia, Saudi Arabia, Senegal, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syria, Taiwan, Tajikistan, Tanzania, Thailand, Trinidad, Tunisia, Turkey, Turkmenistan, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, Viet Nam, Yemen, Yugoslavia, and Zimbabwe. | Open, Classified |  |  | 548071210
Z-0516-2019 | 81451 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Granuflo Dry Acid Concentrate- 2K 2.0 Ca 1 Mg 16.5 GAL |  |  | October 09, 2018 | 2018 | Fresenius Medical Care Renal Therapies Group, LLC | Discolored powder has confirmed the presence of a foreign substance | Under Investigation by firm | Fresenius issued on 10/9/18 via Certified Mail An Urgent Medical Device Recall Customer Notification to affected consignees asking the consignees to check their stock immediately. If they have any listed GranuFlo Dry Acid Concentrate lot. numbers discontinue use immediately and place the product in a secure, segregated area.The Urgent Medical Device Recall Customer Notification ask for the return of the affected lots and to contact Stericycle on instructions how to return the recalled lots. Contact Stericycle at 1-888-671-8851 for instructions on how to return the recalled product. This is EVENT# 10021. |  | 4752 | 4752 | Distribution US nationwide and Mexico | Open, Classified |  |  | 119297
Z-0525-2019 | 81555 | No | NAB | Nonresorbable gauze/sponge for external use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Family Wellness First Aid Sterile Gauze Pad 3 in x 3 in, 10 CT. White, bleached, non-woven, rayon/polyester sterile Gauze Pad 3 inch x 3 inch, 12-ply 10 count cardboard box |  |  | October 16, 2018 | 2018 | ASO, LLC | Potential that gauze pads may not be fully sterilized | Under Investigation by firm | ASO, LLC notified customers of the recall on about 10/16/2018 via "URGENT: MEDICAL DEVICE RECALL" letter sent to corporate. Instructions were to immediately stop using/distributing the affected gauze pads, place them in quarantine, and return to ASO LLC. Customers were also instructed to complete and return the response form. |  | 218 | 218 cases (72, 10-ct boxes per case) | Distributed nationwide to AR, FL, IA, IN, KY, NY, OK, TX, UT, VA. | Open, Classified |  |  | 218
Z-0543-2019 | 81569 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64, Model 728231 |  |  | October 18, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | There is a potential for separation of fiberglass resin which may cause the gantry front cover and the metal support on the cover to fail/fall | Process control | On October 18, 2018, Philips notices to their customers via courier service. Customers were advised to take the following actions:If any unusual noise is heard, discontinue use of the system immediately and contact your Philips representative for support.Field Order Change (FCO) 72800703 will be deployed to initiate an inspection of the front gantry cover. If the inspection finds evidence of weakening at this connection point, a temporary repair solution will be put in place to secure the cover for continued system use. FCO 72800704 will then be deployed to replace the affected front cover. A Philips Field Service Engineer will contact you to schedule the inspection of your system. Customers in need of further information or support concerning this issue, please contact your local Philips representative (North America &amp; Canada) 1-800-722-9377. |  | 6343 | 6,343 units in total | US Nationwide Distribution | Open, Classified |  |  | 6343
Z-0463-2019 | 81341 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE HEALTHCARE CARESCAPE R860 Inspiratory Safety Guard, Part Numbers: |  |  | October 05, 2018 | 2018 | GE Healthcare, LLC | The ventilator Inspiratory Safety Guard (ISG) may disconnect from the breathing circuit pathway. As a result, this could create a loss of ventilation which may lead to inadequate oxygenation for patients, increasing the possibility of hypoxia. | Process control | The firm, GE Healthcare, initiated an "URGENT MEDICAL DEVICE CORRECTION" on 10/05/2018 by letter to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: Follow the safety instructions and If you have any incorrect ISGs, return to GE Healthcare or destroy on site, and complete and return the attached Customer Response form checking box #2 to indicate that you do have affected ISGs. E-mail the completed form to Recall34095.InspiratorySG@ge.com. Please ensure that all potential users in your facility are made aware of this safety notification and the recommended actions.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 1140 | 1140 units (307 units in US) | Worldwide Distribution: US (Nationwide) including Puerto Rico, and to countries of: Albania, Argentina, Australia, Austria, Bangladesh, Bolivia, Bosnia and Herzegovina, Brazil, Cambodia, Canada, Chile, China, Colombia, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, Germany, Greece, Honduras, Hungary, India, Indonesia, Iraq, Italy, Japan, Kazakhstan, Korea (Republic of), Kuwait, Latvia, Lebanon, Libya, Malaysia, Maldives, Mexico, Myanmar, Nepal, Netherlands, Nicaragua, Nigeria, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. | Open, Classified |  |  | 1140
Z-0528-2019 | 81572 | No | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | BD BBL Trypticase Soy Agar with 5% Sheep Blood, Catalog No. 221261, (non-selective, non-differential) |  |  | October 22, 2018 | 2018 | Becton Dickinson &amp; Co. | a portion of a single lot was not manufactured in accordance with the internal formulation resulting in inhibited growth of Gram negative organisms. | Mixed-up of materials/components | Distributors were notified via email on October 22, 2018, requesting their customer distribution list. Customer letters were sent via UPS beginning on October 29, 2018. **Distributors are Advised to do the following:1.Discard any remaining inventory of the above referenced product. 2.Complete the attached Distributor Response Form and return to the BD contact noted on the form whether or not you have any of the impacted material so that BD may acknowledge your receipt of this notification and process your product replacement. 3.Within three days, provide an Excel file or a list with name, address and telephone number of end-users who were shipped this product from your distribution branches so we may inform them of this issue. Fax or email to: BD Regulatory ComplianceFax: 410-316-4258Email: BDRC2@bd.com4.A copy of the customer communication is attached for your records.**Customers area Advised to Take the Following Action(s):1. Complete the attached Customer Response Form and return to the BD contact noted on the form whether or not you have any of the impacted material so that BD may acknowledge your receipt of this notification. 2. Discard any remaining inventory of lot 8221508. BD will process replacement of any discarded inventory upon receipt of the Customer Response Form.3. Share a copy of the Recall Letter with all users of the product within your facility to ensure awareness.Customers with questions may call 800-638-8663, Monday through Friday 7 am - 7 pm, EST. |  | 54300 | 54,300 plates | US Nationwide distribution in the states of FL, MA &amp; NJ. | Open, Classified |  |  | 54300
Z-0527-2019 | 81239 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Optional MDRC-1119 Front Cover Attachment Kit for review display monitors. |  |  | March 30, 2017 | 2017 | Barco N.V. | If metal strips are removed and re-applied to another display, there is potential for the metal strips to be deformed resulting in a smaller contact surface with the front cover magnet strips. There is a risk that the protective front cover could come loose and may fall down. | Device Design | Customers were notified via email on about 03/30/2017. Customers were alerted that a safety related update has been made to the installation manual. A safety notice has been shipped to distributors with the updated installation instructions manual and an acknowledgement letter. Distributors were asked to do the following:a) Inform all your customers who have purchased the Front Cover Kit by sending them the Safety Notice;b) Send the updated installation instructions to all your customers who have purchased the Front Cover Kit;c) Fill out the acknowledgment letter; attach and send in reply to this email. |  | 2087 | 2,087 units | Worldwide - US Nationwide Distribution to CA, IL, GA, TX, and WI. In the countries of Belgium, Brazil, Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Poland, Saudi Arabia, South Africa, South Korea, Spain, Sweden, Taiwan, Turkey. | Open, Classified |  |  | 2087
Z-0529-2019 | 81303 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Gel Alkaline Phosphatase Isoenzyme Control Kit: |  |  | November 07, 2016 | 2016 | Helena Laboratories, Inc. | Due to a customer experiencing complaints on bone fraction of assay not being in range. The complaint: ~12% lower enzyme activity was noted. | Unknown/Undetermined by firm | On 11/02/2016, the firm, Helena Laboratories, Inc., sent a "Withdrawal Notification" to customers via email. The notice described the product, problem and actions to be taken. The customers were instructed remove lot from inventory and discard; contact technical services at 1-800-231-5663 ext.1600 for MRA#; and complete Withdrawal Notification form and return copy to Helena Laboratories via FAX at:409-842-1874 or email to: jjones@helena.com.The firm is doing a Withdrawal of this lot, and will replace this material, if they have not already done so, with a new kit lot. If the customer indicates how many kits it has remaining, the firm will forward replacement, under Material Return Authorization and to use an enclosed form to report replacement needs.Customers are informed to contact Technical Services is 1-800-231-5663, Ext. 1600; for MRA#, and for questions, to contact Jessica Jones at extension 1177, or email jjones@helena.com.The firm's public contact may be reached at (409)842-3714 extension 1177. |  | 1591 | 43 kits (6 kits Domestic/37 Internationally) | Worldwide Distribution: US (nationwide) to states of: GA, IN, and MI, and countries of: Australia, Brazil, Canada, Japan and Hong Kong. | Open, Classified |  |  | 1591
Z-0526-2019 | 81452 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NEOCONNECT, Enteral Extension with ENFit connector, REF EXT-60NC |  |  | October 24, 2018 | 2018 | Neomed Inc | The sterile pouch label states the expiration date is 2016-11-19, but the correct expiration date is 2019-11-09. The case label and the shipper label have the correct date | Under Investigation by firm | The firm, NeoMed, sent an URGENT: MEDICAL DEVICE RECALL" letter dated 10/24/2018 to its consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to do the following:Please contact NeoMed Quality Assurance by email to recall@neomedinc.com to confirm receipt of this letter and obtain a Return Materials Authorization for any units identified above that remain in your inventory. Or, you may mail theenclosed reply card to NeoMed Inc., 100 Londonderry Ct #112, Woodstock, GA 30188.If you have any questions, call 770-516-2225 or email: bryan@brosseauconsult.com. |  | 654 | 654 boxes (32700 units) | US Distribution to states of: California, Florida, Illinois, Michigan, New Jersey, Ohio, Tennessee, and Texas. | Open, Classified |  |  | 654
Z-0539-2019 | 81394 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Natural NailÂż System - Tear Drop Guide Wire, 2.4 mm Diameter, 100 cm Length |  |  | October 22, 2018 | 2018 | Zimmer Biomet, Inc. | Various Trauma Guide Wires- creases in the sealing area on the end opposite of the product's chevron opening compromising product sterility | Packaging process control | Zimmer Biomet conducted recall to distributors and hospital risk managers on 10/22/18, distributors notified via email. Hospital risk managers, as well as distributors with product, will be notified via Fed'X." Distributors responsibilities include:- Locating and removing the product in their territory, as well as identifying hospitals with fielded inventory. Distributors will return on-hand product to Zimmer Biomet and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers will be provided with a letter identifying the issue and their responsibilities. These responsibilities include: Ensuring affected personnel are aware of the notice contents- Assisting the Zimmer Biomet sales representative with the quarantine of the affected products. Certificate of Acknowledgement complete and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility. Questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday |  | 3764 | 3764 | Worldwide - US Nationwide DistributionForeign: CANADAAUSTRALIABRAZILCHILEHONG KONGJAPANMALAYSIANETHERLANDSNICARAGUASINGAPORETAIWAN | Open, Classified |  |  | 202318
Z-0549-2019 | 81071 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Invacare 9805 and 9805P Hydraulic Patient Lifts, Model Nos. 9805 and 9805P |  |  | November 16, 2018 | 2018 | Invacare Corporation | Invacare has identified via customer complaints, the potential for the mounting bolt connecting the caster to the base frame of the patient lifts to become loose. Loose hardware can cause wear to the hardware and housings. If left unresolved, the caster may separate from the base of the lift causing the Lift to tilt and a patient fall. Failure does not occur right away but happens over time if the lift is not properly maintained.Patient falls are inherently dangerous and may result in a range of injuries that are dependent on physical condition of the patient and the nature of the fall. When injuries occur, they are typically non-serious injuries such as bruising, lacerations, or abrasions. In rare cases, more serious injuries such as bone fracture, head trauma, or death can occur.For caregivers using the lift, attempts to manually stabilize a lift if it becomes unstable could result in sprains, serious back injuries, or bone fractures. | Device Design | On November 16, 2018, the firm, INVACARE, distributed "URGENT - MEDICAL DEVICE FIELD ACTION" letters with an attached Technical Bulletin to providers to inform them of the product issue, problems and actions to be taken.Providers are instructed to do the following: 1. Review the included Technical Service Bulletin.2. Inspect your existing stock for loose fasteners or hardware as required in the user manual.3. Tighten any loose hardware after inspecting for damage that require replacement of the Lift base.4. You are responsible for ensuring that this field action is conducted to the customer level. Your customers with affected units should be contacted and informed about this field action. We recommend that you provide them with the enclosed Consumer Letter.5. Complete and return the Provider Response Card within 5 business days via Fax to: 440-326-3544 or email to: recall@invacare.com.The consumer letters instructed end users to do the following:1. Continue to have your patient lift(s) inspected and maintained by a qualified service provider at the intervals defined within your User Manual.2. Continue to check for loose fasteners or hardware prior to each use.3. Should you find loose fasteners or hardware, contact a qualified service provider to perform maintenance. Do not attempt to perform maintenance yourself.If you have any questions concerning these instructions, please call Invacares recall and field action support line at (877) 413-6008, Monday  Friday, 8 a.m.  5 p.m. EST, for assistance. |  | 73703 | 73,703 units | Worldwide Distribution: US (Nationwide) and countries of: Canada, Colombia, Germany, Guam, India, Northern Mariana Islands, Philippines and South Africa. | Open, Classified |  |  | 73703
Z-0553-2019 | 81513 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet 20mm Compress Device Short Anchor Plug With Drill Set / With Pin Inserter Ion Implant |  |  | October 24, 2018 | 2018 | Zimmer Biomet, Inc. | Potential metal burrs in the holes of the device preventingthe appropriate drill and transverse pins from being inserted | Employee error | Zimmer Biomet initiated recall by letter on 10/24/18 to distributors and hospitals advising of the problem and action to be taken-locate quarantine and return product. Complete Inventory Return Certification Form. Questions or concerns call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. |  | 6 | 6 units | US Nationwide Distribution in the states of IN, DC, MD | Open, Classified |  |  | 6
Z-0551-2019 | 81605 | No | IZD | Nuclear uptake probe. | Radiology | Radiology | 510(K) Exempt | Captus 700 Thyroid Uptake System, Model # 5430-3137 and 5430-3138 |  |  | November 05, 2018 | 2018 | Capintec Inc | There is a potential for the spring arm failure as a result of a broken tension rod (a component inside the arm). A break in the tension rod can casue the collimator to fall downward to its lowest point of travel, approximately 25 inches from the ground. | Under Investigation by firm | Notification letters were sent to customers on 11/5/18. A Revised Product Alert - Corrective Action letter was sent to customers beginning 1/15/19 due to an expansion of the event. |  | 65 | 65 | Worldwide distribution - US nationwide distributed in the states of AR, AZ, CA, CO, CT, DC, FL, IA, IN, LA, MD, MI, MN, NC, NJ, OH, OK, PA, SC, SD, TX, VA, WI, WV, and countries of Argentina, Brazil, Canada, China, Cyprus, Greece, Hong Kong, India, Japan, South Korea, Vietnam. | Open, Classified |  |  | 450
Z-0558-2019 | 81545 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis Zee and Zeego, model no. 10272462 |  |  | October 26, 2018 | 2018 | Siemens Medical Solutions USA, Inc | This recall has been initiated due to a possible leakage of a hose in the cooling unit of the X-ray tube.The Artis zee system may lose coolant due to a possible leakage in a hose of the cooling unit in the X-ray tube. This will cause the X-ray tube not to be sufficiently cooled down. As a result, the system will display a message Tube hot, have a break . The X-ray will be blocked several minutes later and to prevent any damage to the machine the following message will be displayed NO XRAY, TUBE TOO HOT .This issue is very sporadic and may occur during an ongoing clinical procedure.Use of this product may have a potential impact on planned procedures as they may be delayed as well as in emergent cases, as clinical treatment may need to be terminated and transferred to a functioning system. | Process control | On October 26, 2018, the firm, Siemens Healthineers, distributed "URGENT: MEDICAL DEVICE CORRECTION" letters to affected customers. The letters informed customers of the product, problem and actions to be taken. The customers were instructed to do the following: promptly notify and instruct accordingly all the staff at your organization who need to be aware of this notice and Immediately examine your inventory and quarantine product subject to recall (if applicable). In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall.If the device has been sold and is therefore no longer in your possession, please forward this notice to the new owner and inform us of the identity of the device's new owner where possible.The letter also stated that Siemens will replace the coolant pump on affected lots of cooling units to correct the issue. Siemens' service organization will contact customers to arrange a date to perform the corrective action with Update Instruction AX013/18/S. To request an earlier appointment, feel free to contact the service organization at 1-800-888-7436.If you have any questions, contact Regulatory Technical Specialist at 610-219-4834 or email to: anastasia.sokolova@siemens-healthineers.com. |  | 64 | 64 (16 in US) | US Distribution to states of: AR, CA, KS, KY, MD, MN, NM, OK, PA, RI, TX, and UT. | Open, Classified |  |  | 64
Z-0557-2019 | 81582 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | STEPHANIX D2RS Digital Dynamic Remote Systemhttp://cts.fda.gov/division-tracking/images/trash.png |  |  | November 02, 2018 | 2018 | Stephanix | The firm has detected a potential risk using the command. After releasing the command, the movement of the table may continue instead of stopping. | Software design | The firm initiated the recall by email on 11/02/2018 to the sole US distributor. This email included a letter to be provided to the end user. The letter directed the user as follows:"As a precautionary measure and before your usual maintenance team intervention, the user should look at the equipment during all of its movement (as stated in the user manual). Immediately release the command if an abnormal movement is detected. If necessary, use the closest emergency switch to stop the system."f you need more information or technical assistance, you can phone your usualcontact. |  | 17 | 17 units | US Nationwide Distribution in the states of FL, IN, WV, TN, MO, WI, CA, FL, GA, PA | Open, Classified |  |  | 17
Z-0555-2019 | 81178 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Thermacare HeatWraps Menstrual |  |  | October 02, 2018 | 2018 | Pfizer Inc. | The firm received complaints of the wrap coming apart and leaking granular material | Device Design | Pfizer, Inc. initiated the recall by letter on 10/02/2018 to the direct accounts. The notice instructed the direct consignees to conduct a sub-recall by notifying any of their accounts to determine if they had any of the affected lot(s) in their inventory at any of their locations. If so, the recalling firm requested that they discontinue distributing the product lot and promptly return the affected product to Pfizer Inc. (c/o Stericycle Inc.).The firm expanded the scope of their recall to the consumer level and disseminated a notice on 11/27/2018 to its direct accounts. The firm issued a press release on 11/26/2018. |  | 293880 | 293880 units | US Nationwide Distribution and PR | Open, Classified |  |  | 418068
Z-0559-2019 | 81054 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA-360 Analyzer, product code 019945/019945R |  |  | August 23, 2018 | 2018 | Tosoh Bioscience Inc | The firm has implemented new product labeling changes for the Tosoh AIA-360 Analyzer. These labeling changes affect the Tosoh Quick Reference Guide, Training Manual and Training DVD previously provided by Tosoh Bioscience.These changes are being made to align the Tosoh Quick Reference Guide and associated training materials to the Operators Manual and provide clarification to the requirement that the maintenance of the B/F probe is to be performed by Tosoh a field service representative, the requirement for use of 70% Ethanol during daily shutdown procedures and the requirement for use of CAP Class 1 reagent grade water for dilution of the concentrated wash and diluent solutions and reconstitution of lyophilized reagents. Failure to follow these requirements may negatively affect system performance, lead to corrosion of instrument components over time, or result in bacterial contamination, respectively. | Labeling Change Control | On August 23, 2018, the firm, Tosoh Bioscience, notified affected distributors and customers via "URGENT: MEDICAL DEVICE RECALL" letter. The described the product, problem and actions to be taken. Summary and location of labeling changes:Âż- Elimination of monthly B/F probe maintenance.Âż- Added requirement for use of 70% Ethanol.Âż- Added requirement for use of CAP Class l reagent grade water per CLSI Guidelines and Standards.Âż- Added clarification of substrate stability. Substrate, when stored at 2-8ÂżC between uses is stable for 30 days.Customers were asked to do the following: -Discontinue maintenance activities on the B/F probe. If issues with the B/F probe are observed, call Technical Support for assistance.Âż- Please continue to perform all your other maintenance procedures at the required intervals as per the instructions in the Operators Manual.Âż- Discard the Training DVD, it is now obsolete. An updated version of the Training DVD, QRG and Training Manual will be sent to you in a separate mailing.Âż- At daily shutdown, use 70% Ethanol to replace the substrate solution.Âż- Use only CAP Class l reagent grade water.Âż- Complete and return the attached Acknowledgement Form to Tosoh Bioscience, Inc., within 15-days of receiving this notification.Âż- Maintain this notification with your laboratory records and forward this information to others who may have received this product.Âż- If you have obtained unexpected test results or received any complaints of illness or adverse events associated with the use of these products, contact Tosoh Technical Support 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please contact call (800) 248-6764 or email at bernadette.oconnell@tosoh.com Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 1228 | 1228 | Worldwide Distribution: US (Nationwide) and countries of: Mexico, Canada, Colombia, Chile, Venezuela, Ecuador, Uruguay, Peru, and Honduras. | Open, Classified |  |  | 1228
Z-0560-2019 | 81524 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Atrium 28Fr PVC Thoracic Catheter RIGHT ANGLE, Part No. 8128 |  |  | October 12, 2018 | 2018 | Atrium Medical Corporation | This recall has been initiated in response to a seal defect found in certain sterile barrier pouchescontaining the device. Affected pouches are un-sealed at the bottom. If the pouch defect is not detected it could represent a serious hazard to the patient.A sterility breach could potentially lead to product contamination. The potentially affected device isused in contact with critical anatomical structures and deep into the body. If contaminated, theconsequences for the patient could be severe. The severity of the possible clinical complicationsdepends on the degree of contamination of the device, the microorganisms involved and the patient srisk factors. To date, Getinge/Atrium has not received any reports of patient injury related to thisissue. | Vendor change control | The firm notified affected customers on October 29, 2018 via an Urgent Medical Device Recall letter. The letter alerted customers to the issue of the product's sterile barrier pouch not being properly sealed at the bottom.Customers were asked to do the following:1. Examine your inventory to determine if you have the affected product listed in thisNotification. If so, remove affected products, quarantine them, and place in a securelocation.2. If you have affected product, contact Atrium Medical Customer Service at 1-800-370-7899,Monday through Friday between 9:00 am to 5:00 pm (EST), for a Return Authorization (RA) toreturn the product for replacement or credit.3. Complete and sign the enclosed Medical Device - Recall Response Form on Page 4 of thisletter to acknowledge receipt of this notification. If you have affected product to return,enter the quantity and your RA number in the space provided.4. Return the completed Medical Device - Recall Response Form to:EMAIL catheter2018.us@getinge.com or FAX 1-973-807-1654.If you have any questions, please contactyour Getinge/Maquet representative or call Maquet Customer Service at 1-603-880-1433, Mondaythrough Friday between the hours of 9:00 am to 5:00 pm Eastern Time Zone. |  | 2930 | 2930 units | Worldwide distribution - US Nationwide distribution, and country of Saudi Arabia. | Open, Classified |  |  | 2930
Z-0561-2019 | 81378 | No | JPB | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | Cascade Abrazo c-ACT-LR Test Card |  |  | February 24, 2017 | 2017 | Helena Laboratories, Inc. | Due to some cards in this lot reported to have experienced variations in results, (verified by wave form), resulting in some cards with shorter or longer times than expected. | Device Design | On 02/24/2017, the firm sent an email to affected users, with a Withdrawal Notification. The firm informed the customers/distributors that Due to some cards in this lot experiencing variations in wave form, resulting in some cards with shorter or longer times than expected. The firm instructed the customers/distributors to disposes of affected lot, as a preventive quality measure. On 04/12/2017, the firm mailed formal letters to all customers.PUBLIC CONTACT: "Name: Dr. Jessica Jones Hanka"Title: Technical Services Manager"Address: 1530 Lindbergh Drive, Beaumont, TX 77707"Telephone #: (409) 842-3714 x1177"Email: jhanka@helena.com |  | 27 | 27 kits | US: Products under control of Manufacturer in TXOUS: Four International Distributors - Australia, Netherlands, Italy, and England | Open, Classified |  |  | 27
Z-0590-2019 | 81528 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Orthopedic Salvage System (OSS) Porous Stem, Item No. 150394 |  |  | October 10, 2018 | 2018 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a lot specific medical device recall for two lots of the Orthopedic Salvage System (OSS) and the Comprehensive Shoulder System. The investigation determined that the Mini Humeral Stem is potentially labeled as the Bowed IM Stem lot. The associated risk of this product issue is a potential extension of surgery &lt;30 minutes if a replacement is readily available. | Under Investigation by firm | On October 10, 2018, the firm notified affected customers via Urgent Medical Device Recall letters. Customers were advised of the product issue. Distributors were asked to take the following actions:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. If the affected product is located at a hospital, remove the product from the hospital and provide the Risk Manager Recall Notice to the hospital and request that they sign the Certificate of Acknowledgement.4. Immediately return all affected product from your distributorship and from affected hospitals within your territory along with hospital Certificate of Acknowledgement.a. For each return, complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.c. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.d. Mark RECALL on the outside of the returned cartons.5. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com.a. Review the list of hospitals included with the email notification sent to your facility.b. Identify whether there are any additional hospitals that Zimmer Biomet is not aware of and list these accounts on the Additional Accounts form. Please provide the form in Excel format.c. If there are no additional accounts to notify, please indicate that there are no additional accounts, or indicate None or NA on the form.Hospital Risk Managers were asked to take the following actions:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your |  | 3 | 3 total | Distributed to accounts in IN, MD, MO, MS, and NM. | Open, Classified |  |  | 3
Z-0582-2019 | 81609 | No | DSQ |  |  | Cardiovascular | PMA | Medtronic HeartWare Ventricular Assist HVAD System Battery Charger Model 1600US - Product Usage: The HeartWare Ventricular Assist System (HVAD) is indicated in the United States (US) and European Union (EU) for use as a bridge to cardiac transplantation (BTT) as well as an alternative to transplantation as destination therapy (DT) for patients who are at risk of death from refractory end-stage left ventricular heart failure. The HVAD Pump Controller is a microprocessor unit that controls and manages HeartWare System operation. It sends power and operating signals to the blood pump and collects information from the pump. The controller requires two power sources for safe operation: either two batteries, or one battery and an AC adapter or DC adapter. While active, patients will typically use two batteries. While relaxing or sleeping, patients should use power from an electrical outlet (AC adapter) because it provides power for an unlimited period of time. The HeartWare Battery Charger is part of the HeartWare Ventricular Assist System (HVAD). Patients receive 4 total batteries with their system along with a battery charger and are instructed to have spare, fully charged batteries always available. The battery charger can charge up to 4 batteries at a time using 4 charging bays. Each battery slides into a bay and is connected to the battery charger. |  |  | November 19, 2018 | 2018 | Heartware, Inc. | HVAD Battery Charger units manufactured with wrong inductors. | Other | On November 2018 Medtronic sent the following letter to all their consignees:Dear Physician or Healthcare Professional:HeartWare, now a part of Medtronic, has been notified by an outside vendor that a quantity of HVAD battery charger units (see image above) may have been manufactured with an incorrect circuit component. This subset may exhibit one or more of the battery charging bays not fully charging the battery and/or battery charger indicators (LED) not lighting.The behaviors may occur in any individual bay, independent of the other charging bays. In analyzed units, up to 2 charging bays exhibited one of the above behaviors. In all cases, at least 2 of the remaining bays were able to continue to charge batteries as needed. We are not able to predict if, or when, one of the identified battery charger units may malfunction. The potential exists, however, that all circuits responsible for charging batteries and LED function may fail due to the incorrect component.In the population of potentially affected product, Medtronic has observed a malfunction rate of 16.7%. The number of potentially affected battery chargers remaining in distribution is estimated to be 2031 worldwide. These units were distributed between June 2017 and June 2018. Charging units outside of the serial numbers listed in Appendix A, all new manufactured battery chargers, and the individual batteries are not in scope of this advisory.Through November 02, 2018, Medtronic has received zero (0) reports of patient injuries associated with this issue.Patient Management RecommendationsWe realize that each patient requires unique clinical considerations. In consultation with Medtronics Independent Practitioner Quality Panel (IPQP), the following recommendations are provided:For patients with a HeartWare Battery Charging Unit in the identified list of Serial Numbers (refer to Appendix A):" Instruct patients to monitor the battery charging unit for any of the following beha |  | 2031 | 2031 | Worldwide distribution - US nationwide distribution in the state of AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI. Countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, Hong Kong, India, Italy, Latvia, Lebanon, Netherlands, Poland, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 2031
Z-0588-2019 | 81165 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE LOGIQ E10 Ultrasound System, Model #530000E10. |  |  | September 14, 2018 | 2018 | GE Healthcare, LLC | There is a potential for an undetected over temperature condition on the trans-esophageal (TEE) probe leading to a possible burn injury. | Process change control | The recalling firm, GE Healthcare, issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated 9/14/2018 via traceable means to each of their customer locations. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: Do not use a TEE probe (GE model 6Tc-RS) from another device on the LOGIQ E10 system until your LOGIQ E10 is corrected. You may scan normally with any other probes.GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 103 | 103 devices(64 US; 39 OUS) | Worldwide Distribution: US (nationwide) to states of: AZ, IL, MA, ME, MN, MT, PA, TX, and WA and countries of: Foreign distribution was made to Austria, Denmark, Finland, and Norway. | Completed |  |  | 103
Z-0583-2019 | 81077 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | PROCARE B20 Patient Monitor, V1; and B20 Patient Monitor, V2. |  |  | September 11, 2018 | 2018 | GE Healthcare, LLC | Patient monitors may restart due to network overload caused by network configuration. | Software design | The recalling firm, GE Healthcare, issued "URGENT MEDICAL DEVICE CORRECTION" letters dated 9/11/2018 via trackable mail on 9/11/2018 to its customers. The letters described the product, problem and actions to be taken. The letters informed the customer of the issue and providing safety instructions for mitigation of the issue if a prolonged network overload occurs and the Bx0 and/or B1x5 patient monitor on the network does not resume normal functionality in approximately 90 seconds. The customers were instructed to follow the safety instructions.GE Healthcare will provide revised network installation and user instructionsdocumentation at no cost to you in a future mailing.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 28004 | 28,004 patient monitors | Worldwide Distribution: US (nationwide) including Puerto Rico; and countries of: Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Bangladesh, Belgium, Bolivia, Bosnia-Herzegovina, Brazil, Brunei, Bulgaria, Burma (Myanmar), Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Gabon, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Ivory Coast, Jamaica, Jordan, Kampuchea, Kazakstan, Kenya, Korea, Kuwait, Lebanon, Lithuania, Macedonia, Malaysia, Mexico, Moldova, Morocco, Nepal, Netherlands, New Zealand, Nicaragua, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Senegal, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, Viet Nam, and Zambia. | Open, Classified |  |  | 79907
Z-0581-2019 | 81519 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Laser Fiber Assy-Single Use-LFS-200-0.22NA-SMA905-Extension Sleeve Silver-Sterile ETO-ForTec Fiber SU-200-TRUE-RT Label Smartscope packaging. |  |  | October 17, 2018 | 2018 | Leoni Fiber Optics GmbH | A percentage of the products shipped in Smartscope single-box packaging may exhibit a broken package seal, thus having compromised sterility. There have been two reports of compromised sterility; however, all products packaged in Smartscope single-box packaging are at risk of a compromised product seal.Broken seal pouches compromise sterility and could increase the risk of infection if used. | Packaging | The firm, LEONI, notified its distributor of the product issue on 10/16/18 via phone call and a subsequent "URGENT: MEDICAL DEVICE RECALL" letter. The distributor was informed of the product issue, and asked to disseminate recall information to customers.Actions to be taken by the Customer/User:End users should discontinue usage of the product. Product should be immediately quarantined. The Distributor should complete the acknowledgement of receipt form, to arrange for replacement and or credit to their account. A Leoni Customer Service Representative shall contact you with a Return material Authorization. If a replacement order is needed please inform customer service at the time of RMA issuance. - Discontinue usage if packaging irregularities exist- Return RMA units to Leoni Fiber Optics Inc. - Inform Customer Service if replacement product is required.- Complete acknowledge form and return back to Leoni Fiber Optics Inc. Attention Roger West, LFOI Quality Manager, Roger.West2@leoni.comIf you have any questions, please contact Quality Manager at 757-258-4805 orEmail: Roger.West2@leoni.com. |  |  |  | US Distribution to the state of: Ohio. | Open, Classified
Z-0589-2019 | 81562 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | ARJOHUNTLEIGH GETINGE GROUP |  |  | October 22, 2018 | 2018 | ARJOHUNTLEIGH POLSKA Sp. z.o.o. | Inspection process at the manufacturer found nonwoven material was used in specific lots of product that may break wile in use during the patient/resident transfer. Patient/resident might fall on the floor. IN consequence, the patient might sustain a major trauma, head injury - damage to tissue that is irreversible or may lead to death. | Nonconforming Material/Component | On 10/10/2018 the firm informed its US consignee via telephone of an Urgent Medical Device Removal Action. On 11/10/2018, an Urgent Field Safety Notice Letter was sent via United Parcel Services (UPS) to consignees. The firm informed its consignee that according to inspection and test results at the manufacturer, it was found that nonwoven material used in this specific lot of the product might break on the main body while in use during the patient/resident transfer. If the product fails as described, the patient/resident might fall on the floor. In consequence, the patient might sustain a major trauma, head injury - damage to tissue that is irreversible or may lead to death. The notification also instructed the consignee to distributed to those individuals who need to be aware within your organization or to any organization where the potentially affected devices have been transferred.Customers were instructed to contact the firm if potentially affected products were found.Customers were also instructed to maintain awareness of this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action.If you have any further questions or require assistance completing the Customer Response Form, please contact Arjo via email at ArjoCustomer.Response-US@Arjo.com.In addition to the notification, a Customer Response Attestation requiring customers to complete the form and return via email to arjocustomer.response-us@arjo.com or fax to 1-833-fax-arjo. |  | 10 | 10 | US nationwide distribution in the state of WV.OUS: Information not provided due to the product being manufactured OUS and imported to US | Open, Classified |  |  | 10
Z-0566-2019 | 81361 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | SonoSite Needle Guide Bracket, REF numbers P06013-06 and P06013-07, Reusable non-sterile bracket for use with SonoSite L25 series transducers. |  |  | November 02, 2018 | 2018 | CIVCO Medical Instruments Co., Inc. |  | Other | The recalling firm began issuing letters dated 10/30/2918 via email on 11/2/2018 requesting return of the boxes of product or reusable brackets that may be in their inventory. |  | 297 | 297 boxes | Foreign distribution was made to Canada, Argentina, Australia, Austria, Belgium, Brazil, Chile, China, Colombia, Croatia, Cyprus, Denmark, Estonia, France, Germany, Hong Kong, India, Israel, Italy, Japan, Latvia, Lebanon, Lithuania, Netherlands, New Zealand, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Taiwan, Trinidad/Tobago, Turkey, Ukraine, United Arab Emirates, and the United Kingdom. | Open, Classified |  |  | 5391
Z-0593-2019 | 81461 | No | MHY |  |  | Neurology | PMA | St. Jude Medical InfinityTM DBS System 8CH Directional Lead, 30 cm, 0.5 REF 6170 - STERILE EO Rx ONLY |  |  | October 11, 2018 | 2018 | St. Jude Medical, Inc. | The most proximal unsegmented electrode of the Deep Brain Stimulation leads, may be constructed with MP35N instead of the required 90/10 platinum-iridium. | Mixed-up of materials/components | On October 11, 2018, the firm sent an Urgent Medical Device Recall Notice to customers (i.e, Hospital Administrators) via overnight mail service informing them that the firm's records indicate that the institution has/had received the affected 8-Channel Lead that may be impacted by an isolated manufacturing issue. The notification explains that the firm believes that some of the listed leads may contain an electrode that does not conform to material specification. In total, there have been four (4) reports from (2) centers describing that the most proximal stimulation tip electrode (electrode position 4) appeared lighter than the more distal electrodes when radiographed. There have been no reports of patient harm to date. Further communication with additional details about this action, inclusive of management recommendations for any patients that may have been implanted with an affected lead, will be forthcoming as soon as possible.The notification also informs the customer that the firm's representative will assist with removing these leads from the customer's account and will facilitate replacing them. The firm is requesting the customer's support on the removal and return of these leads. The FDA has been notified of this activity.In the meantime, if hospitals or it's physicians have any questions, they can contact their local Sales Representative or Abbott Support at 1-800-727-7846 (Opt3).Customers were also provided with an Acknowledgement Form and instructed to complete all requested information and to return immediately. The Acknowledgement Form included a list of the models and serial numbers of the devices that have been delivered to the facilities, but that records do not yet reflect an implant. The customer is instructed to review and indicate if the device(s) were removed or implanted.Customer are instructed to please return the Acknowledgement Form to Abbott at email: DBSNotices@Abbott.com.On October 29, 2018, the firm sending out |  |  |  | Worldwide distribution - US nationwide distribution in the states of AZ, CA, CO, FL, GA, IL, MD, MN, MO, MT, NC, NJ, NY, OH, PA, and WA. International distribution in the countries of Argentina, Australia, Austria, Belgium, Brazil, Colombia, Czech Republic, Finland, France, Germany, Hungary, Italy, Mexico, Poland, Russian Federation, Saudi Arabia, South Korea, Spain, Sweden, Taiwan, Turkey, and United Kingdom | Open, Classified
Z-0592-2019 | 81531 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IntelliVue MX40 Patient Monitor, Model Nos. 865350, 865351, 865352, 867146 - Product Usage: Indicated for use by health care professionals whenever there is a need for monitoring the physiological parameters of patients. Intended for monitoring and recording of, and to generate alarms for, multiple physiological parameters of adults and pediatrics in hospital environments and during transport inside hospitals. |  |  | October 10, 2018 | 2018 | Philips Electronics North America Corporation | A problem has been detected in the Philips IntelliVue MX4O that, if it were to occur, could affect theperformance of the equipment. The volume of the MX4O speaker in these devices may be diminished or not audible above 4,500 feet when operating in Monitor Mode. These devices mayintermittently exhibit Speaker Malfunct INOP messages after the Power OnSelf-Test.The issue is only apparent when the device is being used in Monitor Mode at analtitude of greater than 4,500 feet. | Vendor change control | On November 9, 2018, the firm notified customers of the issue via Customer Information Letters. The letter discussed the product problem. Customers are asked to confirm speaker behavior. It is recommended that the clinical user reboot the MX40 prior to enteringMonitoring Mode to confirm proper operation of the speaker. Customers are also asked to complete the attached Reply Card and return to Philips as soon as possible. The firm will replace affected devices used at above 4,500 feet (1,372 meters).If you need any further information or support concerning this issue, please contact Philips Customer Support at 1-800-722-9377. |  | 35469 | 35469 units | US Nationwide distribution. | Open, Classified |  |  | 35469
Z-0594-2019 | 81356 | No | JPB | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | Cascade Abrazo c-ACT-LR Test Card |  |  | April 20, 2016 | 2016 | Helena Laboratories, Inc. | Due to some cards in this lot experiencing variations in wave form, resulting in some cards with shorter or longer times than expected | Device Design | On 04/20/2016 the firm, Helena Laboratories, contacted affected users via email, and then sent a "WITHDRAWAL NOTIFICATION" (dated April 25, 2016) on 04/27/2016. The notification described the product, problem and actions to be taken. The firm instructed the customers/distributors to disposes of affected lot, as a preventive quality measure.The Withdrawal Notification for Removal of Product from Inventory stated,This lot is to be handled by Distributors, as follows:___ Prevent any shipments or scheduled demonstrations with this lot___ Remove this lot from Available Inventory and discard the cards___ Place cards in Quarantine  pending disposal___ Contact International Sales for credit and delivery of replacement cardsVerify amount REMOVED FROM STOCK: Total number pulled: ____________________ Date _____________________________ The following signifies written certification that this lot has not been sent to an end user:Verified by: _____________________________________________________________Please complete this form and return a copy to Helena Laboratories, via: FAX: 409-842-1874 or e-mail to jjones@helena.com  For questions, contact  Mgr., Technical Services at: 1 (409) 842-3714 ext. 1177.*Reason for removal: Some cards in this lot are experiencing variations in wave form, resulting in possible shorter, or longer, times than expected. While this issue, if present, should be found via use of the Quality Control materials, due to its intermittent nature in the cards, this alone cannot be relied upon. |  | 12 | 12 Kits | US Distribution to state of TX and internationally to: Australia, Netherlands, Italy, and England. | Open, Classified |  |  | 12
Z-0595-2019 | 81477 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Xcela Power Injectable PICC with PASV (XCELA PASV 4F SL 55CM Catheter Kit CE PG), Cat no. 55-711 |  |  | October 16, 2018 | 2018 | Angiodynamics Inc. (Navilyst Medical Inc.) | Specific lots of Xcela Power Injectable and BioFlo PICCs are labelled with incorrect expiration dates. The labelled expiration date extends the shelf life of the product beyond the date supported by validation testing. Although the product under recall has not yet reached its validated expiry date, removal of the affected product is warranted to prevent use beyond the validated expiry date. | Under Investigation by firm | On October 16, 2018, the firm issued Urgent Voluntary Medical Device Recall letters to affected customers. The letter identified the product issue. Customers were asked to do take the following immediate actions:1. Actions to be taken IMMEDIATELY:- Stop using the product subject to recall. - Remove any affected (recalled) product from your inventory (whether in labs, Central Supply, Shipping and Receiving or ANY other location). - Segregate this product in a secure location for return to AngioDynamics, Inc.- Forward a copy of this recall notification to all sites to which you have distributed affected product. - No actions are required on product that has been used or is presently in-use; the affected product is within its validated expiry2. Complete and return the Reply Verification Tracking Form. - For any affected product to be returned, please call AngioDynamics Customer Service at 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday  Friday: Eastern Standard Time) to obtain a replacement or credit for your returned product. - Promptly complete, sign, and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return); following the directions on this page and the Reply Verification Tracking Form. - Email Reply Verification Tracking Form (preferred): recall@angiodynamics.com - Fax Reply Verification Tracking Form: Attn: Xcela and BioFlo Recall CoordinatorFax number 1-800-782-1357 3. Package and Return the Recalled Product. - Package any product that is being returned in an appropriate shipping box. - Write the RMA number on the RMA/Address label (provided on the Recall Verification Tracking Form) and affix the label to the outside of the shipping box. - Please use our FedEx Account Number (284750594) to return this package via second day delivery. - Seal the box and return to: AngioDynamics, Inc. 603 Queensbury Avenue Queensbury, NY 12804 Attn: Xcela and BioFlo Re |  | 1 | Box of 1 Units | Distributed to accounts in Florida. Foreign distribution to Great Britain. | Open, Classified |  |  | 5
Z-0597-2019 | 81560 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Salem Sump Dual Lumen Stomach Tube 10 Fr/Ch (3.3 mm) x 36" (91 cm), Product Code 8888264911 |  |  | October 31, 2018 | 2018 | COVIDIEN LLC | This voluntary recall is being conducted due to incorrect packaging. In one lot of Salem Sump Dual Lumen Stomach Tubes the 10Fr product was packaged in a 16Fr pouch. The use of a 10 Fr Salem Sump Tube when a 16Fr is needed could result in decreased fluid collection, with the potential to affect patient care, particularly in an emergency situation where the correct size is not readily available. | Under Investigation by firm | On October 31, 2018, the firm distributed Urgent Medical Device Recall letters to affected customers. Customers were informed of the incorrect packaging issue. Action Required:1. INSPECT your inventory for the affected product code and lot number (listed on package)2. SEGREGATE and QUARANTINE all on-hand product3. RETURN the enclosed acknowledgment form via fax to 847-689-9101 or 614-652-9648, whether or not you have affected product, indicating the product code, lot and quantity of product youve quarantined or discarded.4. NOTIFY any customers to whom you may have distributed, or forwarded product affected by this recall5. CONTACT the appropriate Customer Service group to arrange for return and credit of anyaffected product: Hospital800-964-5227 Federal Government800-444-1166 Distributor800-635-6021 All other customers888-444-54401 |  | 1522 | 1522 | Worldwide Distribution - US Nationwide and foreign distribution to Canada. | Open, Classified |  |  | 1522
Z-0598-2019 | 81576 | No | PEW | Implantable transprostatic tissue retractor system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | UroLift System |  |  | July 06, 2018 | 2018 | Neotract Inc | Upon implant deployment, the Capsular Tab may not be delivered as the needle is retracted. In this failure mode, the needle is deployed into the prostate and retracted, leaving no implant behind. Failure to deliver a capsular tab may result in a delay in completing a treatment or an inability to complete a treatment for the patient. Use of a recalled device may lead to increased frequency of exiting known risks associated with the product, including bleeding and tissue trauma associated with delivering the needle. | Finished device change control | On July 6, 2018, the firm sent an Urgent Medical Device Recall Notification to customers via Federal Express next day Service informing them that upon implant deployment, the capsular tab may not be delivered as the needle is retracted. In this failure mode, the needle is deployed into the prostate and retracted, leaving no implant behind. Failure to deliver a capsular tab may result in a delay in completing a treatment or an inability to complete a treatment for the patient. Use of a recalled device may lead to increased frequency of existing known risks associated with the product, including bleeding and tissue trauma associated with delivering the needle. It is important to note that this issue impacts only the delivery device. The implant is not impacted. Any implants that have been delivered with the device are not affected.The notification also provides instructions and the following actions to take:1) If you HAVE affected stock:a. Immediately discontinue use and quarantine any products with the lot numbers listed above, so that the affected products can be returned to the firm.b. Please return the product as follows:i. Complete the enclosed Recall Acknowledgement Form and email to uroliftrecall@teleflex.com or fax it to (925) 401-0699, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will issue you a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to the firm.ii. Write the RGA Number on each box and follow the shipping instructions provided by the firm's Customer Service.2) If you HAVE NO affected stock:a. Please complete the enclosed Recall Acknowledgement Form and email to uroliftrecall@teleflex.com or fax it to (925) 401-0699, Attn: Customer Service.b. This will allow us to document your receipt of this letter.If customers have any other questions, they can contact their local Urology Consu |  | 8804 | 8,804 | US: AK, AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, | Open, Classified |  |  | 8804
Z-0599-2019 | 81190 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA-2000 Analyzer, Product code 022100, 022100R, 022101, 022101R |  |  | September 28, 2018 | 2018 | Tosoh Bioscience Inc | Tosoh Bioscience has become aware of a potential issue with the AIA-900 and AIA-2000 analyzers that could result in slippage of the incubator unit contained within the system. If the tension of the timing belt is not adjusted correctly, it is possible for a misalignment that will result in a mechanical error. Tosoh has taken the required actions to inspect the tension of the timing belt during preventative maintenance and correct the instruments where this issue has been observed. If incubator slippage occurs, the unit will cease to function, and a system error will be reported. If this occurs, there may be a delay in testing and reporting patient test results. There have been no reports of injury related to this issue. | Process control | On September 28, 2018, the firm, Tosoh Bioscience, notified consignees via an "URGENT: MEDICAL DEVICE RECALL" letter. The letter described the product, problem and actions to be taken. The letter informed customers that the AIA-900 and AIA -2000 analyzers had a potential issue that could result in slippage of the incubator unit within the system.Immediate Actions to be taken by the Customer/User" Thoroughly review the content of this letter." Continue to use your Instrument.o Contact Tosoh Technical Support if mechanical errors occur.o The tension of the timing belt will be assessed and adjusted as needed at the next scheduled preventive maintenance visit." Complete and return the attached Acknowledgement Form via Fax to: 1-605-636-8651; email: biorecallresponse@tosoh.com or mail to: Recall Coordinator, Tosoh Bioscience, Inc., 3600 Gantz Road, Grove City, OH 43123." File this notice with your laboratory records and forward this information to others in your laboratory.If you have obtained unexpected test results, delays in test result reporting, or received any complaints of illness or adverse events associated with the use of the analyzer, please contact Tosoh Technical Support 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to call (800) 248-6764 or email: bernadette.oconnell@tosoh.com Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 95 | 95 | US Nationwide distribution and International distribution to Colombia, Chile, Honduras, Cayman Islands, Panama, Costa Rica, Guatemala, Venezuela, Uruguay, Peru, and Ecuador. | Open, Classified |  |  | 490
Z-0601-2019 | 81188 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Hemoglobin A1c Control, Part no. 220232 for use as quality control material to monitor the precision of laboratory testing procedures for HbA1c quantitation. |  |  | October 05, 2018 | 2018 | Tosoh Bioscience Inc | Tosoh Bioscience has become aware of potential degradation of HbA1c control lot # 7055 before the 60-days product stability claim. If degradation of these controls occurs, P00 peaks may be seen and the % HbA1c control values may drift higher than the assigned range. This may result in delays when performing routine instrument quality control. If this issue occurs, there may be a delay in testing and reporting patient test results. Tosoh has received two (2) complaints related to this issue with no serious injuries reported. | Device Design | The firm sent notification to consignees on October 5, 2018, via Urgent Medical Device Recall letter. The letter informed consignees of the potential degradation of HbA1c control lot 7055 prior to the 60-days product stability claim.Immediate Actions to be taken by the Customer" Review the content of this letter with your laboratory staff." Inform laboratory staff to avoid storing the control material in direct exposure to light." Complete and return the attached Acknowledgement Form to Tosoh Bioscience, Inc., within 15-daysof receiving this notification." In the event of an early degradation of the control material, new controls should be used to verify theinstrument performance prior to use." Maintain this notification with your laboratory records and forward this information to others in yourlaboratory." If you have obtained unexpected test results, delays in test result reporting, or received anycomplaints of illness or adverse events associated with the use of the analyzer, please contact TosohTechnical Support 24 hours a day, seven days a week at (800) 248-6764." Should you have any questions regarding this medical device recall, please feel free to contactBernadette O'Connell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will beavailable to answer any questions Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 1000 | 1000 | Distributed nationwide. Foreign distribution to Canada. | Open, Classified |  |  | 1000
Z-0604-2019 | 81421 | No | CFN | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | Binding Site Human IgM Kit (for SPAPLUS), REF: NK012.S |  |  | November 06, 2018 | 2018 | The Binding Site Group, Ltd. | Firm is offering advice to customers experiencing challenges when validating the calibration curve of the IgM kits. This is indicated by kit control values reporting outside the specified ranges. | Under Investigation by firm | On 11/06/2018, the firm, The Binding Site, emailed an "Urgent Medical Device Correction notice" and Field Correction E-back Form to customers. The notice described the product, problem and actions to be taken. Customers were instructed to do the following:1) Reagents are to be loaded on-board the analyser and left for a minimum of 18 hours prior to running the calibration and validating the curve. Please leave the reagents on-board the analyser until empty.2) Once calibration is established and validated, reagents must remain on the instrument until the reagent has been fully consumed.3) When not in use the control vials and calibrator sets should be removed from the instrument, be capped to avoid evaporation and stored at 2-8 degrees C.4) To be able to accept the calibration curve and to use the assay to report test results, the control materials to be tested at a minimum once a day should return results within Âż15% of the concentration(s) as stated on the accompanying QC certificate. If a control measurement is out of range when assayed with a stored curve the assay must be recalibrated. If on recalibration the control values measured with the new curve are still out of range, the instrument and the assay parameters should be checked before repeating the assay.5) If calibration issues are still experienced when following this advice please do not hesitate to contact a firm representative.The E-back Form requests that each customer sign the form confirming they have both received and have followed the advice issued in the Notification; and return completed E-back form to technical.support@bindingsite.com or your local Binding Site Representative.This notice needs to be passed on all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. Customers with additional questions are encouraged to contact their local firm representative, or customer service via phon |  | 60 | 60 | U.S Distribution to states of: CA, IL, PA, FL, NE, WA, IN, NY, SD and VA, | Open, Classified |  |  | 60
Z-0608-2019 | 81680 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery NM/CT 670 Pro, model 5376204-70-54 |  |  | May 24, 2018 | 2018 | GE Healthcare, LLC | When your system was installed, the Main disconnect switch, EPO button on the A1 panel, and Remote EPO button(s) (outside of the A1 panel) may not have been wired correctly which could result in these not shutting off all power to the entire system as intended, resulting in a potential electric hazard. | Component change control | The firm initiated the recall by letter on 05/24/2018. The letter explained the issue and requested the UPS(s) be disconnected from the system until the system could be repaired. |  | 1 | 1 unit | CA, DC, OH, UT, WI, Canada | Completed |  |  | 7
Z-0603-2019 | 81443 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx, Rechargeable Li-Ion (lithium ion battery), REF: M3538A, which is a component of HeartStart MRx Monitor/Defibrillators. |  |  | November 23, 2018 | 2018 | Philips North America LLC | Affected lithium-ion batteries may contain a defective component (Thermal Cut-off or TCO). Should the component fail, the battery will no longer charge or deliver power to the Monitor/Defibrillator, which may prevent operation of the device if it is not connected to AC or DC power. Should a second, unaffected battery be present in the Monitor/Defibrillator, the battery failure may go unnoticed. | Nonconforming Material/Component | On 11/23/2018 Urgent Medical Device Recall notices were mailed via USPS certified mail to customers. Customers were instructed to do the following. Inspect the batteries installed in your monitor/defibrillator to determine if they are affected by this recall. If any affected batteries are found, remove them from service immediately, along with any other affected batteries in your inventory. Further, customers were asked to complete and fax the Customer Reply Form to: +1.877.499.7223 or email to recall.response@philips.com.Your monitor/defibrillator can only be returned to service once a charged battery, unaffected by this recall, is installed.Firm will replace affected Rechargeable Lithium-Ion Batteries at no charge to the customer. Firm will contact you when replacement batteries are available.Customers with additional questions are encouraged to call 1-800-722-9377.Distributors were advised to send a copy of the Urgent Medical Device Recall notice and Customer Reply form to any customer to whom they have distributed affected batteries. |  | 1880 | 1880 Batteries | U.S.: AZ, WA, PA, UT, NY, CO, MD, AR, TX, MO, MI, FL, VA, MA, GA, TN, NV, OH, CA, NH, KY, KS, OR, WI, NC, IL, LA, NJ, WV, AL, WY, MSForeign (OUS): United Kingdom, Korea Rep, Singapore, Australia, Taiwan, Sri Lanka, Brazil, France, Switzerland, Germany, Austria, Spain, Italy, Trinidad and Tobago, Estonia, Belgium, Portugal, Korea, Republic of, Argentina, India, Puerto Rico, Qatar, Netherlands, Mexico, South Africa, Ireland, Saudi Arabia, New Zealand, Bolivia, Costa Rica, Chile, Canada, Russian Federation, Slovenia | Open, Classified |  |  | 1880
Z-0605-2019 | 81587 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Tritanium Posterior Lumbar (PL) cage intervertebral body fusion implant. Intended for an open, less invasive, and minimally invasive PLIF or TLIF surgical procedure. |  |  | November 28, 2018 | 2018 | Howmedica Osteonics Corp. | The surgical technique is being updated to caution against misuse due to reports of cage fractures occurring intra-operatively and post-operatively. | Under Investigation by firm | Urgent Medical Device Correction letters were sent to customers on 11/28/18. |  | 28461 | 28,461 | US Nationwide | Open, Classified |  |  | 28461
Z-0602-2019 | 81172 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA-900 Automated Immunoassay Analyzer performs three methods of immunoassay: an immunoenzymetric (IEMA) or sandwich immunoassay, a competitive binding (EIA) immunoassay, and a two-step immunoenzymetric immunoassay. |  |  | September 28, 2017 | 2017 | Tosoh Bioscience Inc | The firm has become aware of a potential issue that may affect the mixing motor drive harness and mixing home detection sensor relay harness on your AIA-900 Analyzer. It is possible that the harnesses could rub against the Plarail chain during operation and result in premature wear. If observed, this issue will cause the following system error to be reported D.lane Mix home overrun error [4093] . If observed, this issue will cause the AIA-900 Analyzer not to function and the D.lane Mix home overrun error [4093] will be reported by the system. Replacement of the motor drive harness and mixing home detection sensor relay harness will mitigate the potential for premature wear and effectively prevent this issue from causing the system error. If the sensor relay harness fails, the AIA-900 analyzer will cease to function and will be unavailable for use. If this issue occurs, there may be a delay in testing and reporting patient test results. Tosoh is aware of two complaints associated with the unexpected wearing of the harness. There have been no user or patient injuries reported that are related to this device malfunction. | Component design/selection | On September 28, 2018, the firm notified customers via Urgent Medical Device Recall letter. Customers were informed of the potential issue.Immediate Actions to be taken by the Customer:Thoroughly review the content of this letter." Identify if the serial number of your analyzer is listed in this notification.o If the serial number of your analyzer is listed:1. At the next service call or scheduled PM, whichever occurs first, a Tosoh field serviceengineer will verify the motor drive harness and mixing home detection sensor relayharness are bundled together and placed into a double-coated cable or repair theharness bundle as applicable.2. Complete and return the attached Acknowledgement Form by any of the listedmethods.o If the serial number of your analyzer is not listed:1. Tosoh records indicate that the harness bundle on your instrument(s) has alreadybeen replaced. Complete and return the attached Acknowledgement Form by any ofthe listed methods." File this notice with your laboratory records and forward this information to other supervisors ormanagers in your laboratory to ensure that they are aware of the potential issue. |  | 258 | 258 | Nationwide distribution. Foreign distribution to Chile, Colombia, Ecuador, Guatemala, Honduras, Panama, Peru, Puerto Rico, Uruguay, and Venezuela. | Open, Classified |  |  | 258
Z-0615-2019 | 81151 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott i-STAT CG8+ Cartridges intended for use in the in vitro quantification of potassium in arterial, venous, or capillary whole blood. |  |  | September 13, 2018 | 2018 | Abbott Point Of Care Inc. | Inaccurate potassium results may be generated. When testing with control fluid and/or patient samples with low or reference range potassium concentration, elevated potassium results may be generated. When testing with samples with a high concentration of potassium, decreased potassium results may be generated | Manufacturing material removal | Abbott Point of Care issued recall letter on 9/13/18 via FedEx , advising of the problem, recommend action to return all unused cartridges to APOC . If you have any questions regarding this information, please contact Abbott Point of Care Technical Support at 844-256-9531 or your Abbott Point of Care representative. |  | 17550 | 17550 cartridges | WI | Completed |  |  | 17550
Z-0617-2019 | 81478 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista 1500, Model No. 10444801 an in vitro diagnostic device intended to duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. |  |  | October 16, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has identified an issue with Dimension Vista software v.3.7 and v.3.8.1 that may result in aliquot well contamination.If a test is cancelled by the operator while the instrument is aliquotting the sample, the Dimension Vista instrument may continue to deliver the cancelled sample or vial contents to the aliquot well(s). Future samples may be assigned to the same aliquot well(s) and become contaminated with the existing contents or overflow of the aliquot well may occur potentially contaminating a surrounding well(s).This issue only occurs when an operator manually cancels a test(s) for calibration, quality control (QC) or patient sample while the instrument is in the process of aliquotting the sample. | Under Investigation by firm | On October 16, 2018, the firm notified customers via Urgent Medical Device Correction letters. The letters informed customers of the product issue. Actions to be Taken by the Customer1. Whenever possible, order/test(s) for calibration, QC or patient samples should only be cancelled by the operator when the system has finished aliquotting the sample." Once the Sample Rack has been unloaded from the Dimension Vista the test(s) may be cancelled following normal procedure." For automation systems once the tube has been released back to the track the test(s) may be cancelled following normal procedure.2. If the order/test(s) for calibration, QC or patient samples must be cancelled while the sample is onboard the system, please follow the instructions below:a. From the Operation menu select Pause .b. When System Paused displays in the status area, the test(s) may be cancelled.c. From the Operation menu select Reset to reset the Dimension Vista.3. If order/test(s) for calibration, QC or patient samples were cancelled by the operator while the sample is onboard the system and the system was not paused:a. From the Operations menu select Stop . All in-progress tests will be abortedb. From the Operations menu select Shutdown c. The console menu is displayed. To restart system, select Start Vista. In addition, please perform the following:" Review this letter with your Medical Director." Please retain this letter with your laboratory records, and forward this letter to those who may have received this product." Return the completed Field Action Effectiveness Form to Siemens Healthcare Diagnostics as instructed in the last page of this letter.If you have any questions, please contact your Siemens Customer Care Center or your local Siemens Technical Support Representative. |  | 1628 | 1628 | Nationwide/worldwide distribution. | Open, Classified |  |  | 3277
Z-0610-2019 | 81678 | No | GFO | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | StaclotÂż LA (REF 00600) |  |  | November 14, 2018 | 2018 | Diagnostica Stago, Inc. | There have been reports of shortened (T1-T2) results, where some instances of result recovery outside of the range of less than or equal to 15 seconds. | Nonconforming Material/Component | The firm, STAGO, sent a "Field Safety Notice" to its customers on 11/14/18. The notices describe the product, problem and actions to be taken. The customers were instructed to do the following: If you have any lots listed in the appendix in your laboratory, stop using and destroy the concerned kits; contact our sales administration team (1-800-222-2624 prompt #3) as soon as possible in order to obtain a replacement lot and complete the attached coupon acknowledging receipt of the letter.For additional information, please contact the STAGO hotline (1-800-725-0607) or email: Donald.kraft@us.stago.com or 1-973-631-1200 ext 4104. |  | 22342 | 22342 total | US Nationwide Distribution | Open, Classified |  |  | 22342
Z-0609-2019 | 81521 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Carousel MAS1025-01, part of the Brain Surgical Accessories Intra-operative Kit (KIT0295), accessory used with the Renaissance System (TPL0038). Indicated for precise positioning of surgical instruments or spinal implants during general spinal and brain surgery. It is intended to be used whenever the clinician and/or patient benefits from generated 3D imaging of, high contrast objects. |  |  | November 26, 2014 | 2014 | Mazor Robotics Ltd | In some situations the carousel position pin (locker) may cause the carousel upper plate to pop out and the carousel may then inadvertently move. | Device Design | All Mazor Brain Clinical Sales Representatives (support during clinical cases) were notified of the issue on 11/26/2014 via email as well as received training. |  | 5 | 5 units | Affected devices were distributed in the US to four different medical facilities in four states | Open, Classified |  |  | 5
Z-0614-2019 | 81527 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Xper Flex Cardio Physiomonitoring System, Software versions 1.5.16.0322 and 1.6.1.0390, Model Numbers 453564241901, 453564241911, 453564483321, 453564483331, 453564621791, 453564621801, 453564669081 |  |  | November 15, 2018 | 2018 | Philips Electronics North America Corporation | There are 3 issues:1) Under certain conditions, real time waveforms may not be accurately plotted or displayed. 2) The oxygen saturation (SpO2) numeric value displayed on the device may freeze.3) In auto cycle mode, if the NIBP communication is lost, the pump would not cycle and the non-invasive blood pressure (NIBP) numeric value displayed on the device would not update. | Software Design Change | The firm initiated the recall on X by letter. The letter explained the issues, potential risk, and provided workarounds for the consignee to use until the new software is ready for release. |  | 651 | 651 units | worldwide | Open, Classified |  |  | 651
Z-0613-2019 | 81170 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA-360 Analyzer is capable of performing two methods of immunoassay: an immunoenzymetric (IEMA) or sandwich immunoassay, a competitive binding (EIA) immunoassay. An antigen-antibody reaction begins by combining a patient sample, control, or calibrator with a diluent in an immunoreaction test cup from the AIA-PACK reagent series. |  |  | March 22, 2017 | 2017 | Tosoh Bioscience Inc | The firm has become aware of a potential issue with the AIA-360 analyzer that could prevent the table home sensor from detecting the table home position. If the malfunction with the table home sensor occurs on the AIA-360, the sensor board may need to be replaced. If the home sensor fails to detect the home position, the AIA-360 analyzer will cease to function and will be unavailable for use. If this issue occurs, there may be a delay in testing and reporting patient test results. Tosoh has received three (3) complaints related to this issue with no serious injuries reported. | Process control | On October 5, 2018, the firm notified customers of the recall via Urgent Medical Device Recall letter. Customers were informed of the potential issue with the analyzer that could prevent the table home sensor from detecting the table home position. Customers were advised that the table home sensor board will be evaluated and replaced if needed at the next scheduled preventive maintenance visit.Actions to be taken by User:" Thoroughly review the content of this letter." Continue to use your analyzer. A Tosoh representative will contact you to schedule your next preventivemaintenance visit.o If you experience any problems, please contact Tosoh Technical Support for an expedited servicevisit." Complete and return the attached Acknowledgement Form." File this notice with your laboratory records and forward this information to others in your laboratory.If you have obtained unexpected test results, delays in test result reporting , or received any complaints of illnessor adverse events associated with the use of the analyzer, please contact Tosoh Technical Support 24 hoursa day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to contactBernadette O'Connell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will be availableto answer any questions Monday- Friday, from 9:00 AM to 5:00 PM (PST). |  | 438 | 438 | Nationwide distribution. International distribution to Honduras, Caribbean, Chile, Colombia, Costa Rica, Uruguay, Ecuador, Panama, Mexico, Canada, and Venezuela. | Open, Classified |  |  | 438
Z-0619-2019 | 81615 | No | MRZ | Infusion pump. | General Hospital | General Hospital | 510(k) | SAPPHIRE Epidural Set Yellow-Striped Microbore, 0.2 Micron Filter, PAV, 123 Inch, List Number 16373-02 |  |  | November 08, 2018 | 2018 | ICU Medical Inc | There have been reports of distal occlusion alarms during the use of SAPPHIRE Epidural Set Yellow-Striped administration sets. | Device Design | On 11/08/18, Stericycle, Inc. sent a "URGENT MEDICAL DEVICE CORECTION" letter to its consignees. The letter described the product, problem and actions to be taken via traceable mailing. Delivery will be confirmed via traceable mailing. Consignees who further distributed the product were asked to notify their customers.For further inquiries, please contact ICU US Field Action Manager at 224-706-2487 or veasna.seth@icumed.com. |  | 31050 | 31,050 units | US distribution and Internationally to Canada. | Open, Classified |  |  | 31050
Z-0618-2019 | 81415 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK 2 Antimicrobial Susceptibility Testing (AST) using one or more of these susceptibility products: |  |  | October 22, 2018 | 2018 | bioMerieux, Inc. | False resistant results for Streptococcus anginosus and Streptococcus constellatus strains were reported. | Component change control | The Firm, Biomerieux, issued a " URGENT Product Correction Notice" to their consignees informing of the matter and providing certain actions for them to take. Each Notice will be accompanied with an Acknowledge Form which is to be returned to bioMerieux.. The consignees were instructed to take the following action at this time:" Please confirm this letter has been distributed and reviewed by all appropriate personnel within your organization." Reference your Laboratory Standard Operating Procedures to identify card types and breakpoints used in your facility to determine if your laboratory is impacted by the issue described in this letter." Reference your Laboratory Standard Operating Procedures to identify if retrospective analysis is required for reported results associated with the referenced limitation." For systems utilizing EUCAST breakpoints and operating with VITEKÂż 2 Systems Software version 8.01, create a bioART Rule for the referenced limitation associated with Ceftriaxone (cro01n)." Please store this letter with your bioMÂżrieux instrument documentation.-Please complete the attached Acknowledgement Form (Attachment A) and return it to bioMÂżrieux, Inc.If you have any questions or concerns, please contact your local bioMÂżrieux representative or call , Regulatory Quality Compliance Specialist, at 314-731-8805 or email: debra.broyles@biomerieux.com. |  | 184374 | 184,374 | Worldwide Distribution: US (nationwide) to states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, WV, WY including Puerto Rico and to countries of:Algeria, Angola, Azerbaijan, Bahrain, Bangladesh, Belarus, Sarl, Benin, Bosnia-Herz, Cameroon, Chad, Cote d'Ivoire, Croatia, Dem. Rep. Congo, Egypt, , Georgia, Guam, Ethiopia, Iraq, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, UAB, Lithuania, Macedonia, Malta, Mongolia, Morocco, Mozambique, Niger, Oman, Pakistan, Palestine, Qatar, Romania, Saudi Arabia, Senegal, Serbia, Slovenia, Sri Lanka, Tunisia, Ukraine, Gabon, and Zimbabwe. | Open, Classified |  |  | 184374
Z-0627-2019 | 81595 | No | NEY | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Emprint Standard Percutaneous Antenna with Thermosphere Technology, CA20L2 intended for use in percutaneous, laparoscopic, and intraoperative coagulation (ablation) of soft tissue, including partial or complete ablation of nonresectable liver tumors. |  |  | November 15, 2018 | 2018 | Covidien LLC | The firm received consumer reports of the ceramic trocar tip of the Emprint ablation antenna disengagingfrom the needle shaft post-ablation. This can result in the tip remaining in the patient. | Nonconforming Material/Component | On November 15, 2018 Medtronic issued UPDATED URGENT MEDICAL DEVICE RECALL notices to customers. Customers are advised to take the following actions:1. Please immediately quarantine and discontinue use of the affected products.2. Return affected product as indicated.3. If you have distributed the products, please promptly forward the information from the letter to those recipients.4. Complete the Recalled Product Return Form even if you do not have inventory.If you have any questions or concerns, please do not hesitate to contact your Medtronic representative or Technical Support at (800) 448-3644, Option 3, Option 8. |  |  |  | US Nationwide | Open, Classified
Z-0630-2019 | 81604 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂż Complete Knee Solution Femoral Augment Block Distal only, 15mm, Size E with Screw |  |  | October 31, 2018 | 2018 | Zimmer Biomet, Inc. | Screw is missing from the package | Process control | Zimmer Biomet issued "URGENT MEDICAL DEVICE RECALL" Letter on 10/31/18 via email and FedEx to the consumer/user level, including any intermediate wholesale or retail consignees, distributors and Risk managers. Distributors responsibilities include locating and removing the product in their territory, as well as identifying hospitals that have previously used the product. Hospital risk managers responsibilities include the following: 1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send toCorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5. If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 19 | 19 units | Worldwide Distribution: US (nationwide) to states of: OH, FL, NJ, NY, OH, PA, TX, WV and to countries of: CANADA, AUSTRALIA, JAPAN, GERMANY, and NETHERLANDS. | Open, Classified |  |  | 19
Z-0628-2019 | 81600 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | enVista One Piece Hydrophobic Acrylic Intraocular lens - |  |  | November 05, 2018 | 2018 | Bausch &amp; Lomb Surgical, Inc. | Cosmetic imperfections on the surface of some lenses. | Process control | On Nov 5, 2018 Bausch &amp; Lomb sent letters to all their consignees, requesting the following;1. Review and quarantine your inventory of all the impacted lots for this recall.2. Complete the enclosed Medical Device Voluntary Recall Acknowledgement Form and contact B&L; to obtain a return material authorization number and arrange for a pick up of the identified product. Contact 800-338-2020 option 2.If you have any questions, contact the Quality Manager at 727-724-6695. |  | 70440 | 70,440 units | US Nationwide Distribution: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL IN, KS, KY, la, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY. | Open, Classified |  |  | 70440
Z-0497-2019 | 81167 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 980 (PB980) ventilator series: |  |  | September 18, 2018 | 2018 | COVIDIEN LLC | Software Update: External USB Drive performance and its impact on Graphic User Interface (GUI) functionality and labeling of the scalar waveform displayed on GUI during ventilation | Device Design | Medtronic notified U.S. consignees notified via Federal Express or certified mail on September 19, 2018, and the letter informs customers that Medtronic has developed a software update, Maintenance Release (MR) 5.4. Medtronic Service Engineers will load the MR5.4 software during the next scheduled or planned maintenance visit. Immediately notify all care environments in which the PB980 ventilators are used about this notification." If your facility has distributed PB980 ventilators to other persons or facilities, please promptly forward a copy of this letter to those recipients." Complete the attached form and return it as directed to confirm your receipt and understanding of this information.Any incidents related to this issue or if you have any questions, please contact our Technical Support Department immediately at 1-800-255-6774.Out--Side US accounts notified by Medtronic Regulatory Partners worldwide on 9/17/18 for translation in their respective country. |  | 9593 | 9,593 | Worldwide Distribution - US NationwideForeign:CanadaArgentinaChileJapanHong KongTaiwanUruguayAustriaAustriaAzerbaijanBelgiumBosnia And HerzegovinaBulgariaCroatiaCzech RepublicFranceGermanyGreeceHungaryIrelandIsraelItalyKazakhstanKenyaKuwaitLatviaLebanonLithuaniaNetherlandsOmanPakistanPolandPortugalRomaniaRussian FederationSaudi ArabiaSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTunisiaTurkeyUnited Arab EmiratesIranUnited Kingdom | Open, Classified |  |  | 9593
Z-0629-2019 | 81709 | No | DSQ |  |  | Cardiovascular | PMA | Jarvik 2000 Ventricular Assist System external components: JHI-212 Y Cable, JHI-202 Lithium-ion Battery Cable, JHI-204 Extensions Cable, JHI-201 Pedestal Cable |  |  | October 15, 2018 | 2018 | Jarvik Heart Inc | There is a potential for the male connectors on external cables to become damaged. This could include a detached male connector housing, misaligned connector, cracked connector housing, and bent pins. | Device Design | Urgent Field Safety Notice letters dated 10/5/18 were distributed to customers. |  | 274 | 274 | Products were distributed to the following US states: FL, NY, and UT.Products were distributed to the following foreign countries: Croatia, Czech Republic, France, Hungary, India, Italy, and Japan. | Open, Classified |  |  | 274
Z-0623-2019 | 81488 | No | GKA | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | Genesys 2 Reagent, REF 01-03-0022 intended to be used with automated high performance liquid chromatography (HPLC) systems employing ion exchange chromatography. For in vitro diagnostic use only. |  |  | November 01, 2018 | 2018 | Clark Laboratories, Inc. (dba,Trinity Biotech USA) | lack of 510K | No Marketing Application | The firm initiated the recall by letter on 11/01/2018. The letter explained the problem and requested the return of the recalled items. |  | 216 | 216 units | IL, WV, NJ, AZ, MD, UT | Open, Classified |  |  | 1317
Z-0643-2019 | 81537 | No | IWJ | Manual radionuclide applicator system. | Radiology | Radiology | 510(K) Exempt | Theragenics Applicator Needle (18g x 20cm) |  |  | October 09, 2018 | 2018 | Theragenics Corporation | Some of the pouches containing Brachytherapy Applicator Needles (18g x 20cm) were not sealed | Under Investigation by firm | Theragenics notified customers on about 10/09/2018 via phone call and follow up email. Instructions included to return any affected product to Theragenics for replacement. |  | 12000 | 480 packs (5 needles each)  NOTE: The multipacks are part of a box of 25 needles, containing five (5) single packs and four (4) 5-packs. | Nationwide distribution to FL, CA, AL, DE, NC, TX, UT, MS. International distribution to South Africa. | Open, Classified |  |  | 12000
Z-0639-2019 | 81589 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEKÂż 2 Systems Software Version 9.01 Update Kit. |  |  | November 13, 2018 | 2018 | bioMerieux, Inc. | The VITEKÂż 2 Systems Software Version 9.01 software may display incorrect organism identification information as well as the associated Antimicrobial Susceptibility Testing (AST) results when using the VITEKÂż 2 FLEXPREP" feature. | Software design | bioMerieux notified customers of the recall on about 11/13/2018 via "URGENT PRODUCT CORRECTION NOTICE." Instructions included taking the following actions: 1) Confirm the letter has been distributed to, and reviewed by, all appropriate personnel; 2) If appropriate in accordance with consignee's laboratory's Standard Operating Procedures, perform retrospective review of any patient results potentially impacted in relation to the described VITEK FLEXPREP anomaly; 3) Store this letter with the bioMÂżrieux, Inc. VITEK 2 instrument documentation; 4) Complete and return the Acknowledgement Form in Attachment A to confirm receipt of the notice and return it to your local bioMÂżrieux, Inc. representative. Regardless of each consignee's current use of the VITEK FLEXPREP" feature, your local Field Service Engineer has implemented the software solution on your VITEK 2 System.If you require additional assistance or have any questions, please contact your local bioMÂżrieux, Inc. Clinical Customer Service representative at (800) 682-2666. For non-urgent inquiries, please contact the Technical Support Center by email at CustomerService-IDAST@biomerieux.com. The anticipated turnaround time for email inquiries is 24-48 hours during normal business hours (7AM  7PM EST), and up to 72 hours dependent upon weekends and company observed holidays. |  | 2 | five (5) VITEKÂż 2 Systems Software version 9.01 | Nationwide distribution to CO, IL, MO, TX. | Open, Classified |  |  | 2
Z-0641-2019 | 81635 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity Universal Viewer Zero Footprint Client, Model Number 2090258-001, Software Versions 6.0 SP7, SP7.0.1, SP7.0.2, SP7.0.3, SP8, SP8.0.1, SP8.0.a and SP8.0.2, SP9 and SP9.0.1 display o orientation markers can become is incorrect. |  |  | November 09, 2018 | 2018 | GE Healthcare, LLC | Issue #1 affects Software Versions 6.0 SP7, SP7.0.1, SP7.0.2, SP7.0.3, SP8, SP8.0.1, SP8.0.a and SP8.0.2.When an image is flipped horizontally and vertically, the orientation markers appear correctly on the image to which they were applied. However, once new image rendering happens on this image or series (with scrolling; playing cinÂż; maximizing or changing viewports/layouts; or synchronizing viewports and resizing the browser), the new image orientation is correctly maintained but the display of orientation markers can become incorrect.Issue #2 affects Software Versions 6.0 SP7, SP7.0.1, SP7.0.2, SP7.0.3, SP8, SP8.0.1, SP8.0.a, SP8.0.2, SP9 and SP9.0.1. When an image in a series is flipped and then rotated in Zero Footprint Viewer (ZFP), the orientation markers appear correctly on the image to which they were applied. However, once you navigate to another image in the series and the image you navigated to has a different orientation than the prior image from which you navigated, the new image orientation is correctly maintained but the display of orientation markers can become incorrect. | Software design | The firm issued the medical device correction notices on 11/09/2018 by letter. The letter outlined the problems and provided work-around strategies to be employed until a product modification is available to correct the issues. |  | 508 | 508 units | Worldwide | Open, Classified |  |  | 508
Z-0638-2019 | 81579 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General Hospital | 510(k) | Protexis Neoprene Surgical Glove, Size 9.0 |  |  | October 31, 2018 | 2018 | Cardinal Health 200, LLC | Labeling issue in that the breakthrough (i.e., permeation) times for two chemotherapy drugs (Carmustine and Thiotepa) stated on the wallet (primary packaging) of the Protexis Neoprene surgical gloves are incorrect. Additionally, breakthrough times for Mitomycin were added to the same label in error. | Labeling mix-ups | Cardinal Health notified customers on about 10/31/2018 via "URGENT: PRODUCT CORRECTION NOTICE." Customers were instructed to examine inventory for affected products, affix the provided warning label on the principal display panel on the affected dispenser packs, notify internal parties and/or customers if the product was further distributed of the recall, and complete and return the provided Acknowledgement Form. If the wallets containing the gloves are removed from the dispenser packs, customers were also instructed to provide a warning label to accompany the affected product inform users of the correct breakthrough times. If customers need additional labels, they were instructed to contact 800-292-9332 or gmb-fieldcorrectiveaction@cardinalhealth.com for assistance. For further questions regarding the Product Correction Notice, please reference the following Customer Service telephone numbers: Hospital (800-964-5227), Federal Government (800-444-1166), Distributor (800-635-6021), All other customers (888-444-5440). |  | 1475095 | 1,475,095 pairs total | Worldwide Distribution: US (Nationwide distribution) to states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and to countries of: Austria, Belgium, Canada, Chile, Costa Rica, France, Germany, Italy, Luxembourg, Martinique, Monaco, Netherlands, Poland, Reunion, Spain, Sweden, Switzerland, Taiwan, and United Kingdom. | Open, Classified |  |  | 1475095
Z-0640-2019 | 79584 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | GloCyte RBC Reagent - Product Usage: The GloCyte Automated Cell Counter System is intended for use by trained healthcare professionals in clinical laboratories to provide quantitative determination of fluorescence labeled total nucleated cells and erythrocytes in cerebrospinal fluid collected from adult and pediatric patients. The GloCyte Low and High Level Controls are assayed hematology controls designed to monitor the performance of the GloCyte Automated Cell Counter System. Assayed parameters include total nucleated cells and erythrocytes. |  |  | February 07, 2018 | 2018 | Advanced Instruments, LLC | RBC Reagent failure. High fluorescent counts during RBC Reagent + Diluent QC check that will result in failed QC check, preventing user from testing patient samples for RBC. | Device Design | Customers have been sent a recall notification communicating the reason for recall. Once a root cause is identified, recalling firm will notify customers and make available new RBC Reagent lots to allow customers to resume use of reagent in RBC counting. |  |  |  | US Nationwide distribution in the states of IN, WA, NC, CA, GA, CO. | Open, Classified
Z-0642-2019 | 81404 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | Zimmer Compress Devices and Instruments: |  |  | October 12, 2018 | 2018 | Zimmer Biomet, Inc. | Correction to update the surgical technique for the Compress System. To make users aware of the changes that were made only to the selection criteria section on page 42 regarding which array to use in the placement of the anti-rotation pins. | Device Design | Zimmer notified accounts on 12/13/18 via email and FedEx (both original and expanded accounts) as a follow up to the medical device correction notification dated October 12, 2018. Zimmer Biomet identified additional item numbers that are within the scope of this issue. .Distributors letter identify the issue and their responsibilities, which include: Reviewing the correction notification and ensuring affected team members are aware of the contents. Offering assistance in ensuring users are aware of the updated surgical technique and how to locate it. Destroying previous copies of the surgical technique. Completing the Certification of Acknowledgement.Updated surgical techniques 1110.2-GLBL-en REV0818 and BMET0188.2 REV111218 for the Compress and Mini Compress systems are used, effective immediately, for surgeries that require the affected items.a. Access the full surgical techniques at www.zimmerbiomet.com by navigating to the Limb Salvage specialty and then Bone Compressive Devices under the Medical Professionals menu. |  | 531 | 323 devices US; 208 distributed OUS | NationwideForeign:CANADAARGENTINAAUSTRALIABELGIUMCHILE FINLANDJAPAN NETHERLANDS | Open, Classified |  |  | 531
Z-0650-2019 | 81563 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System, Version 3.3.1 &amp; below |  |  | October 19, 2018 | 2018 | Ortho-Clinical Diagnostics | Software anomaly may incorrectly extend the on-analyzer stability time for a loaded VITROS Reagent. | Software Manufacturing/Software Deployment | On October 19, 2018, Ortho Clinical Diagnostics issued "URGENT PRODUCT CORRECTION NOTIFICATION" notices via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers to inform them of the issue and actions to take to prevent occurrence of the anomaly. On November 19, 2018 Ortho Clinical Diagnostics re-issued correction notices to customers pertaining to the 5,1 FS system (CL2018-194_5,1FS). Customers may continue using their VITROS Systems and have been informed that the software issue will be mitigated with the next version of software. The expected dates of these software versions are included in the communication. Foreign affiliates were informed by email approximately on 21 October 2018 of the issue and instructed to notify their consignees of the issue and required actions to prevent the occurrence of the anomaly.Please refer to the Customer letter (Ref. CL2018-194_51FS, CL2018-194_3600, CL2018-194_4600, and CL2018-194_5600) for specific details on Required Actions by User. If you have further questions, please contact Ortho Care" Technical Solutions Center at 1-800-421-3311. |  | 2413 | 2413 | US Nationwide Distribution | Open, Classified |  |  | 4928
Z-0644-2019 | 81608 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AvenirÂż Muller Stem Standard, uncemented HA, 1, Taper 12/14 Item Number: 01.06010.001 - Product Usage: Avenir@ Muller Stems are intended to reduce pain and increase hip mobility through long term cementless fixation of total or hemi hip arthroplasty in the femur of patients with an adequate bone stock to support the component. A system consisting of a stem, a ball head and a cup is used for the treatment of degenerative diseases or trauma of the hip. All Avenir Muller Stems are intended to provide connection to articulation with specified ball heads through the 12/14 taper. The stems may be used for total hip and hemi-hip arthroplasty. During total hip arthroplasty, the stems may be combined with constrained or semi-constrained acetabular systems. |  |  | October 31, 2018 | 2018 | Zimmer GmbH | Mislabeled Labeled Avenir MÂżller Stem size 1, the stem inside the packaging is an Avenir MÂżller Stem size 2 . | Process control | Zimmer Biomet issued Notices to the OUS consignees on October 31, 2018. Letter identifies action to be taken: Review this notice and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement |  | 13 | 13 units | International distribution in the countries of Germany and Switzerland. | Open, Classified |  |  | 13
Z-0649-2019 | 81614 | No | PQP |  |  | Pathology | PMA | FoundationOne CDx Test Results |  |  | October 31, 2018 | 2018 | Foundation Medicine, Inc. | An incorrect test result for a single analyte may have been reported for up to 800 physician reports | Other | On October 31, 2108, an outreach program was implemented by Foundation Medicine's Medical Affairs team to advise ordering physicians who received reports of possibility that the relevant reports could contain an incorrect test result. |  | 800 | 800 reports | US nationwide | Open, Classified |  |  | 800
Z-0648-2019 | 81574 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | IMRIS ORT 300, Removable Operating Room Table, Part Numbers: (a) 114093-000, (b) 114093-006 - Product Usage: The ORT200 and ORT300 are intended for use during diagnostic examinations, radiologic procedures, or surgical procedures to support and position a patient. The tables are indicated for imagining from head to sacrum with MR/CT imaging modalities. The ORT100 table is intended for use during diagnostic examination or surgical procedures to support and position the patient. |  |  | November 06, 2018 | 2018 | Deerfield Imaging, Inc. | Table may drift in the roll position while in use | Under Investigation by firm | The firm initiated the recall by letter on 11/06/2018. The letter explained the issue, requested the consignee follow the recommended daily test in all positions until the roll cylinders are exchanged. |  | 14 | 14 units | Worldwide distribution - US Nationwide and countries of China, Qatar, Canada, Australia, Japan, France, Germany, Sweden, Italy. | Open, Classified |  |  | 56
Z-0657-2019 | 81645 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Reliance LITe Nerve Hook Bayoneted, 90, Blunt Tip, REF # 48066103 |  |  | November 26, 2018 | 2018 | Howmedica Osteonics Corp. | Instead of having a spherical diameter, the probe was observed to have 2 radii and a flat surface. Additionally, upon further investigation, it was observed that some parts exhibited an edge on the intersection between the radius and the diameter of the probe. | Under Investigation by firm | The firm, Stryker, sent an "Urgent Medical Device Recall" letter dated 11/14/2018 to its customers on 11/26/2018. The letter described product, problem and actions to be taken. The customers were instructed to: Examine your inventory; Immediately quarantine the affected product; and complete customer response form to Stryker Spine at email: Spine-RegulatoryAction@stryker.com.For questions contact Manager-Regulatory Compliance &amp; Quality System at 866-45-SPINE (77463).. |  | 39 | 39 | Worldwide distribution: US Distribution to the state of: PA and countries of:: Australia, Canada, France, Italy, Netherlands, Poland, and Spain. | Open, Classified |  |  | 39
Z-0656-2019 | 81607 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 8Fr SAFE GRAD CAT W/Valve (Argyle Suction Catheter Kits with Chimney Valve), Model Number 36826 |  |  | November 05, 2018 | 2018 | COVIDIEN LLC | The product was released into commercial distribution while subject to importation refusal. | Under Investigation by firm | Urgent Medical Device Recall letters dated 11/5/18 were sent to customers. |  | 32071 | 32,071 units | The products were distributed to the following US states: AL, AR, AZ, CA, CO, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, and WI. | Open, Classified |  |  | 32071
Z-0660-2019 | 81664 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SP CAL (LIQ), REF IT2692 |  |  | November 19, 2018 | 2018 | Randox Laboratories Ltd. | Calibrator value sheets have been modified for assigned value of Alpha-1-Acid-Glycoprotein (AGP), IgA and IgG to improve the alignment. IgG measurement reference has also been corrected. | Under Investigation by firm | On November 19, 2018, Randox Laboratories Ltd. issued "Urgent Medical Device Correction" notices and New Calibrator Value sheets (Revised 15 Nov 18 pl) to customers. Customers were advised to take the following action:- Review all remaining stock. Add the notice and replace the calibrator value sheet with the revised document provided.- Discuss the contents of the notice with your Medical Director. Review results generated with the affected batches in line with the clinical profile of the patient. - Retain a copy of the letter with laboratory records and forward a copy of the letter to those who may have received this product.- Return the completed response form via email to technical.services@randox.com.If you have any questions, contact RA Specialist by email at maria.kelly@randox.com. |  | 54500 | 250 kits (32 kits US / 218 kits OUS) | US Distribution to states of : IN, NJ, NY, TX and WV and internationally to Ireland. | Open, Classified |  |  | 54500
Z-0655-2019 | 81529 | refractor, manual, non-powered, including phoropter | HKN | Manual refractor. | Ophthalmic | Ophthalmic | 510(K) Exempt | RT-5100 Refractor, a component of the Epic-5100 System. |  |  | June 14, 2018 | 2018 | Nidek Inc. | The lens bonding was incomplete on the refractor. | Employee error | The recalling firm called the sole distributor on 6/14/2018 to determine the quantity of instruments still in their control. During the phone call, the distributor was also notified of the invoice numbers that show the serial numbers involved, was asked to segregate the affected product, and a schedule was discussed for the recalling firm's technicians to visit the distributor for correction of the product. |  | 85 | 85 instruments | Distribution was made to FL. | Completed |  |  | 85
Z-0654-2019 | 81616 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Empower FastLoad CT Syringe Pack -sterile disposable syringe kit for use with the EmpowerCTÂż, EmpowerCTAÂż and EmpowerCTAÂż+ Injector Systems |  |  | November 12, 2018 | 2018 | Bracco Injeneering S.A. | Insufficient seal on the sterile barrier of the device, compromising the sterility of the device | Packaging process control | The firm, Bracco Injeneering, issued "URGENT: MEDICAL DEVICE RECALL" letter via third party recall firm, on 11/12/18 to customers and distributors identifying the problem, and requesting return of unused product., Complete and return the Removal Recall Response Form even if you have no recalled stock on hand.Questions: contact Stericycle at 888-759-6907 (Monday- Friday 8 a.m. to 5 p.m. Eastern Time) or acistmedical3021@stericycle.com. |  | 33350 | 33,350 packs (667 cases) | US Nationwide Distribution | Open, Classified |  |  | 33350
Z-0659-2019 | 81649 | No | IXY | Wall-mounted radiographic cassette holder. | Radiology | Radiology | 510(K) Exempt | Bucky Stand SU-4000 Input 2017.11.17. |  |  | November 08, 2018 | 2018 | Konica Minolta Healthcare, Americas, Inc. | Unintentional U-Arm movement | Device Design | On 11/08/2018, the firm send a letter titled;" IMPORTANT PRODUCT INFORMATION", regarding the Konica Minolta KDR U-Arm. Potential unintentional U-Arm movement due to the tablet enclosure buttons, located on the tablet enclosure, remaining engaged after being depressed.The firm included a document titled, KDR U-ARM Tablet Enclosure FAQ as part of the IMPORTANT PRODUCT INFORMATION. The firm is preparing to issue an Interim Corrective Action; Disabling the buttons contained on the KDR tablet enclosure and a Final Corrective Action; A modification to the mounting of the PCB - which contains the drive buttons will reduce the potential of the event -this will help assure the buttons make the proper contact without staying depressed.In addition, the firm issued a "Technical Bulletin-Business Partner for the Product: KDR U-Arm Tablet Enclosure." The preliminary notification was sent to dealer partners which require all Partners to take immediate action.There are two critical actions needed to mitigate this risk:1.Interim Action (Immediate Action Required):The firm has released an immediate field corrective action to temporarily disconnect the tablet enclosure buttons. Instructions can be found in KON-000127 .2.Final Action (Anticipated Release January 2019):A modification to the tablet design will help ensure the buttons make the proper contact without staying depressed.The Interim Action/Field Correction is required to be completed by December 15, 2018. Contact Konica Minolta Dispatch department at 1- 800-945-0456 to schedule a Konica Minolta Engineer to complete Corrective action. For each affected unit, it is important that your engineer performs the following:1. Complete the Tablet Button Disconnection procedure - KON-0001272.Fill out form KON-000157 (Customer Signature required)3.Take a picture of the completed form KON-000157 and email to ServiceBilling@gcp.konicaminolta .com. |  | 151 | 151 | US Nationwide Distribution: AL; CA; CO; FL; GA; IA; IL; IN; KY; LA; MA; MI; MN; MO; NC; NH; NY; OH; OK; PA; SC; TN; TX; and VA. | Open, Classified |  |  | 151
Z-0658-2019 | 81598 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Optiflux F160NR Capillary High Flux Dialyzer Single Use Only, Product Code 0500316E |  |  | November 05, 2018 | 2018 | Fresenius Medical Care Renal Therapies Group, LLC | Potential for external blood leaks from the dialyzer header | Other | On November 6, 2018, Fresenius Medical Care issued URGENT MEDICAL DEVICE RECALL notices to customers. Customers were advised to take the following actions:1. Check your inventory for the affected lot numbers.2. If found, discontinue use of these lots.3. Please all affected units in a secure, segregated area.4. Contact Stericycle at 1-888-731-7986 for instructions on how to return the recalled product or if replacement product is needed right away. (Refer to EVENT # 6273).5. Complete and return the Reply Form.Customers with questions may contact Medical Information and Communications at 855-616-2309 or thru the website www.fresenius-medinfo.com. |  | 109164 | 9,097 cs x 12 units each | US Nationwide | Open, Classified |  |  | 109164
Z-0682-2019 | 81547 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL800 analyzer, model numbers 393-800 and 393-801. |  |  | July 03, 2018 | 2018 | Radiometer America Inc | The analyzer software may cause a mis-match of patient demographics and test results when a certain sequence of events occur. | Software design | The recalling firm issued letters dated 7/3/2018 on 7/3/2018 via FedEx. |  | 2131 | 2,131 analyzers | Distribution was nationwide. There was also military/government distribution. Foreign distribution was made to Canada. | Open, Classified |  |  | 2131
Z-0678-2019 | 81641 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | HydroSet XT Injectable HA Bone Cement, REF 897005, STERILE |  |  | November 15, 2018 | 2018 | Howmedica Osteonics Corp. | Incorrect expiry date of 9-Feb-20Correct expiry date is 30-Aug-19 | Labeling mix-ups | On November 15, 2018, Stryker issued Urgent Medical Device Recall notices to customers via Fedex. Customers were advised to take the following actions:1. Inform users of the medical device recall and forward a copy of the notice to all those individuals who need to be aware within the organization.2. Examine inventory to identify impacted HydroSet kits.3. Immediately quarantine any affected product.4. Complete and return the Business Reply Form via email to Strykerortho6629@stericycle.com or via fax to 1-888-345-3845.5. Return all affected product available at your location to the following address:Stryker Orthopaedics/PFA Product ReturnsAttn: Distribution Inventory Team 325 Corporate Drive, Dock M-East, Mahwah, NJ 07431, Ref. PFA 1878253 |  |  |  | CO, FL, GA, IA, MA, ME, MI, MN, NC, ND, NH, NY, PA, SC, TN, TX, VA, WI &amp; WV | Open, Classified
Z-0681-2019 | 81640 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Laparotomy Sponge, XR No Loop, Single-Use, Sterile, 18'' x 18'' pre-washed with bands without handles, 5 sponges/pack, 40 packs/case. X-Ray Detectable, 1 Non-Absorbent outer wrap. |  |  | November 27, 2018 | 2018 | Medical Action Industries Inc | Potential compromise of product sterility due to breach of sterile barrier. | Under Investigation by firm | Medical Action Industries notified customers on about 11/27/2018 via "URGENT MEDICAL DEVICE NOTICE OF RECALL" Letter. Consignees were instructed to immediately examine inventory, quarantine all affected product, notify customers if the affected product was further distributed, and complete and return the recall response form. For consignees with affected product in inventory, Medical Action Industries will issue a Returned Materials Authorization (RMA) to return affected product. |  | 13 | 13 cases (2,600 sponges) | Devices were distributed in 5 U.S states: MN, WA, NY, MA and MI. No product was distributed outside of the US. | Terminated | February 08, 2019 | 73 | 13
Z-0669-2019 | 81632 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alinity i 25-OH Vitamin D Reagent Kit, List Number 08P4532 |  |  | November 21, 2018 | 2018 | Abbott Ireland Diagnostics Division | Abbott has become aware of a manufacturing issue that can result in damage to a bottle neck in some of the reagent cartridges found in the lots listed in Appendix A. If a bottle neck is damaged, it may cause a failure during the reagent loading process on the Alinity i instrument. | Nonconforming Material/Component | The firm initiated the recall by letter on 11/21/2018. The notice explained the problem and provided instruction on how to inspect the reagent cartridges for defect. The consignee was directed to discard the defective units. The units will be replaced. |  | 760 | 760 units | AR, FL, GA, MD, OK, VA | Open, Classified |  |  | 1995
Z-0670-2019 | 81655 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | NeoSpan SuperElastic Compression Staple w/Instruments, Size 10x10x10, Part Number T50 SN010 |  |  | November 14, 2018 | 2018 | In2bones USA, LLC | The NeoSpan Compression Staple (10x10x10) kits from two lots contain the wrong size drill bit. They contain a 3mm drill bit instead of a 2mm drill bit. | Employee error | The firm initiated the recall by letter on 11/14/2018. The letter requested the return of the product. |  | 30 | 30 units | TX, VA, PA, MO, NJ, FL, TN, NC, OH, WI, LA | Open, Classified |  |  | 30
Z-0674-2019 | 81597 | No | CDZ | Testosterone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ST AIA-PACK Testosterone (TES) is designed for IN VITRO DIAGNOSTIC USE ONLY for the quantitative measurement of testosterone in human serum on specific Tosoh AIA System analyzers. Measurement of testosterone is used to aid in the diagnosis and management of conditions involving excess or deficiency of this androgen. |  |  | November 30, 2018 | 2018 | Tosoh Bioscience Inc |  | Device Design | On November 30, 2018, Tosoh Bioscience issued Urgent Medical Device Recall notices to customers. Customers are advised to take the following actions:" Do not use the assays identified in notice for those patients who have recently had a fluorescein angiography procedure. For these patients, use an alternative testing method that does not use a fluorescein-based technology." Inform medical professionals that the Tosoh AIA assays in the notice utilize a fluorescein-based technology and should not be used to test patient samples for patients who have recently undergone fluorescein angiography." Continue to use the Tosoh AIA assays listed in notice only for patients who have not undergone a fluorescein angiography procedure exam, or on patient samples which have been taken prior to receiving the fluorescein angiography procedure." Complete &amp; return the Urgent Medical Device Recall Acknowledgement Form within 15-days of receiving the notification." Maintain a copy of the notification with your laboratory records and forward this information to others who may have received this product.Should you have any questions, please contact Bernadette OConnell at 1(800) 248-6764 or by email at Bernadette.oconnell@tosoh.com. She will be able to answer any questions Monday-Friday from 9 AM to 5 PM (PST). |  | 26745 | 26745 units | US Nationwide | Open, Classified |  |  | 101916
Z-0680-2019 | 81581 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | GEM Coupler Forceps, GEM4183C, packaged individually. |  |  | November 02, 2018 | 2018 | Baxter Healthcare Corporation | Potential presence of rust on Coupler Forceps (GEM4183C). | Process control | Synovis Micro Companies Alliance (MCS), Inc. a subsidiary of Baxter International Inc., notified customers on about 11/02/2018 via "Urgent Product Recall letter sent via U.S.P.S., first class mail. Instructions included to review inventory for affected product, arrange for return of all affected product to Synovis MCA, complete and return the provided Baxter customer reply form, and notify customers if the affected product was further distributed. General questions can be directed to Synovia MCA at 800-510-3318 or 205-941-0111 Monday through Friday, between 8:00 am and 5:00 pm Central Time. |  | 123 | 123 devices | Nationwide distribution to CA, FL, GA, KS, LA, MA, MS, NE, NJ, NV, NY, OH, PA, TX, UT, VA, WA, WV. International distribution to United Kingdom, Germany, South Africa, Japan, andAustralia | Open, Classified |  |  | 123
Z-0663-2019 | 81210 | No | KRB | Thermodilution probe. | Cardiovascular | Cardiovascular | 510(k) | Iced Thermoset Closed-Loop Injectate Delivery System for cold Injectate. 1 unit per Cavity #2-4"x9" Cav/20 pieces per case R1-15645 |  |  | October 05, 2018 | 2018 | ICU Medical, Inc. | Potential for leakage due to defect in the thermoset check valve component. | Equipment maintenance | On October 5, 2018, ICU Medical URGENT MEDCAL DEVICE RECALL letters were sent to ICU Medical's consignees advising them of this matter. Each letter contained a response form and a return label. |  | 25640 | 25,640 units | United States - AL, CA, CO, DC, FL, GA, IN, KY, LA, MA, ME, MD, MI, MO, MS, MT, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV, WY | Open, Classified |  |  | 27400
Z-0683-2019 | 81246 | No | OZO |  |  | Clinical Chemistry | PMA | Medtronic MiniMed 630G 6025805-003_C, Rx Only |  |  | October 03, 2018 | 2018 | Medtronic Inc. | Reports of occurrences in which insulin pumps with version 4.10 software have failed to make expected audio sounds during alerts, alarms or sirens. This could cause a user to miss system notifications, alarms or sirens associated with how the pump is working, and with high and low glucose alerts. | Software design (manufacturing process) | On 10/3/2018, the firm sent an URGENT FIELD SAFETY NOTIFICATION to customers and distributors via email and/or US Postal Service First Class to inform them of reports of occurrences in which insulin pumps withversion 4.10 software has failed to make expected audio sounds during alerts, alarms, or sirens. This failure could either cause the alarm volume to be stuck at a 4 (out of 5) level regardless of the personal setting, or it could make no sound. Either of these occurrences could cause the user to miss system notifications, alarms or sirens associated with how the pump is working, andwith high and low glucose alerts. Consignees are being informed that they are receiving this letter because the firm's records indicate they may have one of the potentially affected insulin pumps. The notification further explains, this issue occurs when a specific electronic component in the pump malfunctions. This potential loss of audio could delay the user's response to the underlying reason for the alert, alarm, or siren, which could then lead to possible health and safety risks such as hypoglycemia or hyperglycemia. If this issue occurs with the pump, the audio cannot be permanently repaired or regained. The only way to permanently regain the audio capabilities is to exchange the affected pump with a replacement. Even if the pump has the audio issue, itwill continue to deliver insulin as expected. Loss of the Vibrate feature has not been reported to occur as part of the above audio failure mode. If a customer has a pump that contains the specific electronic component, and the pump passes the Audio Beep test described below, there is still a chance that the pump may lose its audio capabilities at some future date.Customers are instructed to:What you should do:1. If you have not already done so, enable the Vibrate feature on your pump. Even if the Audio Beep test passes, you should enable the Vibrate feature (in addition to your Audio feature) since it will add |  | 142889 | 142889 | US: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, Puerto Rico, RI, VT, SC, SD, TN, TX, UT, VA, Virgin Island, VT, WA, WI, WV, and WY | Open, Classified |  |  | 142889
Z-0685-2019 | 81700 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | I-ViewÂż Contrast Enhanced Digital Mammography, ASY-08109 |  |  | December 06, 2018 | 2018 | Hologic, Inc. | Calibration issue not possible to visualize contrast uptake in the subtracted mammography images | Software design | Hologic issued letter on 12/6/2018 USPS Certified Mail to customers to inform them of this product correction, and Hologic asks customers to only use I-View for patients with known lesions and confirm the image quality is of acceptable diagnostic quality.Hologic will scheduling appointments to recalibrate the systems for contrast enhanced digital mammography. |  | 337 | 337 units | NationwideForeign:AustraliaAustriaBelgiumBrazilCanadaChinaColombiaEcuadorEgyptFranceGermanyGreeceHong KongIndiaIrelandIsraelItalyJapanKorea, Republic ofKuwaitMalaysiaMexicoNetherlandsPakistanPeruPolandPortugalSaudi ArabiaSingaporeSlovakiaSouth AfricaSpainSwitzerlandSwitzerlandTaiwanThailandTurkeyUnited Arab EmiratesUnited Kingdom | Open, Classified |  |  | 337
Z-0686-2019 | 81705 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System Version 2.64. |  |  | December 04, 2018 | 2018 | Elekta, Inc. | If a Setup field (CT, kV, MV or MVCT) is created by COPYING a treatment field or another setup field, then thecopied field will have the same isocenter as the parent and that isocenter may not be the correct isocenter forimaging when it is sent to the machine. | Software design | Elekta notified customers on about 12/04/2018 via letter titled, "Urgent Important Field Safety Notification." The letter contains recommended user action and that appropriate personnel working with the device should be advised on the content of the letter and the letter should be placed in an area accessible to all users. The letter stated that the issue will be resolved in Service Packs to MOSAIQ Releases 2.64, 2.65 and 2.70. Customers will be informed when the fix is available through a Product Bulletin.An acknowledgement Form was sent as part of the letter to be completed and returned no later than 30 days. The firm can be contacted at 408-830-8000 or 855-693-5358. |  | 551 | 551 units | Distributed in the United States and outside of the Unites States: Algeria; Antigua and Barbuda; Argentina; Australia; Belgium; Botswana; Brazil; Canada; China; Colombia; Denmark; France; Germany; India; Indonesia; Iran; Ireland; Israel; Italy; Japan; Latvia; Lebanon; Malaysia; Mexico; Morocco; Myanmar; Netherlands; New Zealand; Norway; Peru; Philippines; Poland; Portugal; Russian Federation; Rwanda; Saudi Arabia; Senegal; Singapore; Slovenia; South Africa; Spain; Sweden; Switzerland; Taiwan; Thailand; Turkey; United Kingdom; Viet Nam | Open, Classified |  |  | 551
Z-0552-2019 | 81214 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | StealthStation Cranial and Synergy Cranial Depth Gauge: |  |  | September 26, 2018 | 2018 | Medtronic Navigation, Inc. | An incorrect display of the Biopsy Depth Gauge graphical representation in the Guidance View using Cranial software, which resulted in the biopsy of healthy tissue. | Under Investigation by firm | On 09/26/2018, the firm sent an Urgent Medical Device Correction Notification to all US customers informing them of potential inaccuracy during biopsy procedures using the StealthStation S7 Depth Gauge feature on StealthStation S7 and i7 systems running Synergy Cranial and StealthStation Cranial software that could result in healthy tissue being biopsied. The notification letter stated, In navigated biopsy procedures, where the Navigate Projection feature is used and the projection is longer than the length of the plan, the graphical Biopsy Depth Gauge feature can display inaccurate information. The Biopsy Needle Depth Gauge is a numerical and graphical representation of the needle cutting window in the Guidance view quadrant to assist in visualization of the cutting window position.During a biopsy procedure, the monitor shows four quadrants (Trajectory 1, Trajectory 2, Guidance, and Probes Eye views) as well as the Biopsy Depth Gauge. The overlay of the Biopsy Needle within the anatomical views accurately indicates the correct tip location of the instrument. However, if Navigate Projection is selected, and the projection is longer than the length of the plan, then the Biopsy Depth Gauge graphical display inaccurately depicts that the tip of the instrument has not yet reached the plan target. Since 2011, Medtronic has received seven (7) complaints potentially related to this software anomaly, one in which healthy tissue was biopsied.The image on page two contains an example of the error use case. All views correctly show that the instrument is past the target, and the Depth Gauge correctly displays the past-target measurement of -63, which indicates the instrument tip has advanced 63 mm beyond the intended target. However, the graphical Biopsy Needle cut window appears not to have advanced to the target position.If the surgical team relies solely on the graphical representation in the Biopsy Depth Gauge feature and disregards other factors (such as |  | 5487 | 5487 impacted systems | US: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, Guam, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, Puerto Rico, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WYOUS: Albania, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Barbados, Belarus, Belgium, Bosnia, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cote d'Ivoire, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, French Guiana, Germany, Greece, Guadeloupe, Honduras, Hong Kong, Hungary, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Martinique, Mexico, Moldova, Morocco, Myanmar, Nepal, Netherland, New Zealand, Nicaragua, Norway, Pakistan, Peru, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam | Open, Classified |  |  | 5487